## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍 | 專利權人 | IPC
US20110034782 | A1 | 2011-02-10 | APPARATUS FOR BODY SURFACE STIMULATION TREATMENT, PROGRAM FOR BODY SURFACE STIMULATION TREATMENT, AND COMPUTER-READABLE RECORDING MEDIUM HAVING RECORDED THE SAME | 2006-11-29 | [Problems] To provide an apparatus for a body surface stimulation treatment capable of overcoming the uncertainty in the existing body surface stimulation therapy and exhibiting stable therapeutic effects.;[Means for Solving Problems] An apparatus for a body surface stimulation treatment (10) which comprises a body surface stimulation unit (16) whereby the body surface of a patient is physically or chemically stimulated, a vital reaction quantitative evaluation unit (12) whereby the vital reaction of the patient is quantitatively evaluated, and a body surface stimulation regulation unit (14) whereby the stimulation properties imparted to the body surface tissues of the patient by the body surface stimulation unit (16) are regulated depending on the vital data obtained by the vital reaction quantitative evaluation unit (12). | 1. An apparatus for body surface stimulation treatment comprising: a body surface stimulation unit for physically or chemically stimulating the body surface tissue of a patient,a vital reaction quantitative evaluation unit for quantitatively evaluating the vital reaction of said patient, anda body surface stimulation regulation unit for regulating stimulation properties to be imparted to the body surface tissue of said patient by said body surface stimulation unit based on a vital data acquired by said vital reaction quantitative evaluation unit. 2. The apparatus for body surface stimulation treatment according to claim 1, wherein said vital reaction is the autonomic nerve activity, a change of visceral functions caused by the autonomic nerve activity, a vital signal reflecting the change of the visceral functions, or the respiration of said patient. 3. The apparatus for body surface stimulation treatment according to claim 2, wherein said vital reaction quantitative evaluation unit is a neuroelectric meter capable of measuring said autonomic nerve activity, an electrocardiographic meter capable of measuring at least heart rate and electrocardiographic wave form as said vital signal, a local continuous blood flow meter capable of measuring at least pulse wave form as said vital signal, or a blood pressure meter capable of measuring at least blood pressure and pulse rate as said vital signal. 4. The apparatus for body surface stimulation treatment according to claim 1, wherein said body surface stimulation unit is an electric stimulation means for electrically stimulating the body surface tissue of said patient. 5. The apparatus for body surface stimulation treatment according to claim 4, wherein said electric stimulation means is so configured as to enable regulation of at least one factor selected from the group consisting of voltage, electric current, frequency, frequency of stimulation, and duration of stimulation. 6. The apparatus for body surface stimulation treatment according to claim 1, wherein said body surface stimulation unit is a thermal stimulation unit for thermally stimulating the body surface tissue of said patient. 7. The apparatus for body surface stimulation treatment according to claim 1, wherein said body surface stimulation unit is a mechanical stimulation means for mechanically stimulating the body surface tissue of said patient. 8. The apparatus for body surface stimulation treatment according to claim 7, wherein said mechanical stimulation means is so configured as to enable regulation of at least one factor selected from the group consisting of position of stimulation, intensity of stimulation, frequency of stimulation, and duration of stimulation. 9. The apparatus for body surface stimulation treatment according to claim 1 wherein said body surface stimulation regulation unit is so configured as to enable communication with said vital reaction quantitative evaluation unit or said body surface stimulation unit by... | FUJIKIN CORPORATION | A61B005/00 A61N001/00 A61F007/00 A61H001/00
US20110022981 | A1 | 2011-01-27 | PRESENTATION OF DEVICE UTILIZATION AND OUTCOME FROM A PATIENT MANAGEMENT SYSTEM | 2010-07-21 | In a graphical user interface displayed in an electronic display of a computing device, a first device profile and second device profile are presented, the first and second device profiles each comprising at least one device parameter used to configure a medical device of a subject. A user input control is presented to select one of the first or second device profiles to provide a selected device profile. A probabilistic outcome of the subject corresponding to the selected device profile is then presented. | 1. A system comprising: an electronic display;a memory to store a first and second device profile, each profile comprising at least one device parameter used to configure a medical device of a subject; anda processor, coupled to the memory and the electronic display, the processor configured to: present, in a graphical user interface displayed in the electronic display, the first device profile and second device profile;present, in the graphical user interface, a user input control to select one of the first or second device profiles to provide a selected device profile; andpresent, in the graphical user interface, a probabilistic outcome of the subject corresponding to the selected device profile. 2. The system of claim 1, wherein the processor is configured to: present, in the graphical user interface, a user input control to configure the medical device using the selected device profile; andin response to the user input control being activated, configure the medical device with the selected device profile. 3. The system of claim 2, further comprising: a patient monitor operatively coupled to the processor and configured to: determine an efficacy of the medical device when configured with the selected device profile;compare the efficacy of the medical device to a target outcome; andwhen the efficacy of the medical device deviates more than a threshold value from the target outcome, communicating an alert. 4. The system of claim 1, wherein the first device profile is a current device profile of the medical device and the second device profile is a recommended device profile for the medical device. 5. The system of claim 4, wherein the processor is configured to obtain the recommended device profile from a database of device profiles. 6. The system of claim 1, wherein the processor is configured to: present a user input control to adjust a parameter associated with the selected device profile; andin response to the user input control being used, dynamically revise the probabilistic outcome to provide a revised outcome; andpresent the revised outcome in the graphical user interface. 7. The system of claim 1, wherein the processor is configured to: present a user input control to select the probabilistic outcome from a plurality of available probabilistic outcomes, wherein the plurality of available probabilistic outcomes are selected based on at least one of the first or second device profiles. 8. The system of claim 7, wherein the processor is configured to: identify the medical device to provide a medical device identification; anduse the medical device identification to selectively present the plurality of available probabilistic outcomes. 9. The system of claim 1, wherein the processor is configured to: present a user input control to select a second probabilistic outcome; andpresent, in the graphical user interface, the second probabilistic outcome of the subject corresponding to the selected device profile. 10. The system of claim 9... | - | G06F003/048 A61B005/00
US08118756 | B2 | 2012-02-21 | Lancet-integrated sensor | 2010-10-25 | A sensor and a lancet are integrated with each other. A thin strip-shaped sensor and a lancet are integrated so that the lancet moves in parallel, along a longitudinal direction of the sensor. A measuring device to which an integrated lancet and sensor is attached is provided with a function of driving the attached lancet. | 1. A lancet-integrated sensor, comprising: a lancet having a needle with a needle tip at one end of said lancet for lancing skin to collect bodily fluid; anda sensor for analyzing the bodily fluid collected by lancing with said lancet, said sensor comprising a sensor body that has a cavity formed therein that slidably houses a part of said lancet;wherein said sensor body has electrode terminals for making electrical connection with an external measuring apparatus;wherein said lancet has an end part that projects from said sensor body, said end part being located at a second end of said lancet opposite to said needle tip at said one end of said lancet, and said lancet has a connector at said end part projecting outside of said sensor body so that said connector can be grasped by a driving device of an external measuring apparatus;wherein said lancet and said sensor body comprise locking structure that is arranged such that a force applied to said lancet, in a direction that (a) is opposite to a direction in which said needle lances skin and (b) tends to move said lancet into said cavity in said sensor body, will cause said lancet and said sensor body to be locked with each other; andwherein said locking structure comprises a convex locking portion on one of said lancet and said sensor body and a concave locking portion on the other of said lancet and said sensor body, whereby said lancet and said sensor body are locked with each other when said convex locking portion is engaged with said concave locking portion. 2. The lancet-integrated sensor of claim 1, wherein: said lancet comprises a plate member having a shape including two shorter sides and two longer sides;said needle extends from one of said shorter sides of said plate member, in a vicinity of the center of said one of said shorter sides, and in a longitudinal direction of said plate member;said cavity in said sensor body has a shape corresponding to an outline shape that is obtained when said lancet is slid a predetermined amount in the longitudinal direction; andsaid sensor body has one opening at a front end thereof communicating with said cavity through which said needle can project and a second opening at a rear end thereof through which said end part of said lancet projects. 3. The lancet-integrated sensor of claim 1, wherein: said lancet has two plate-shaped projections that project from said two longer sides of said lancet, respectively, in a direction perpendicular to the longitudinal direction;said concave locking portion comprises a concave portion on upper surfaces of said plate-shaped projections;said cavity of said sensor body comprises two hollows that extend the width of said cavity in a direction perpendicular to the longitudinal direction over part of the longitudinal extent of said cavity, said hollows receiving said two plate-shaped projections therein, respectively; andsaid hollows have ceiling surfaces and said convex locking portions comprise a convex portion on sa... | Panasonic Corporation | A61B005/00 B65D081/00
US08615305 | B2 | 2013-12-24 | Implantable medical device with antenna | 2010-10-26 | An implantable medical device for use in a patient management system is described including a housing, a first communications unit, and a second communications unit. The first communications unit can establish a first communications link within short range telemetry communication range of an external local device. The second communications unit can establish a second communications link over a pervasive wireless communications network with a remote computer. A pervasive wireless communications network can include a cell phone network. The device further includes an antenna that is configured to be operatively connected to both the first and second communications units. | 1. An implantable medical device for use in a patient management system comprising: a housing of an implantable medical device;a first communications unit disposed within the housing configured to establish a first communications link within a short-range telemetry communication range of an external local device;a second communications unit disposed within the housing configured to establish a second communications link with a remote computer over at least one pervasive wireless communications network in the group consisting of: a wireless telephone network, a wireless wide area network, and a wireless local area network; anda single antenna operatively configured to be operatively connected to the first communication unit and the second communications unit. 2. The device of claim 1, wherein the first communications unit is configured to deliver notification of one or more significant events to the external local device. 3. The device of claim 2, wherein the second communications unit is configured to deliver notification of the one or more significant events to the remote computer. 4. The device of claim 3, further comprising: a sensor capable of measuring a body characteristic and generating a data signal describing the measured body characteristic; anda processor disposed within the housing of the implantable medical device and in communication with the sensor, wherein the processor is configured to analyze the data signal the identify the one or more significant events. 5. The device of claim 1 wherein the antenna is a fractal antenna. 6. The device of claim 1 wherein the first communication link is a short-range radio frequency link. 7. The device of claim 1 wherein the antenna is configured to: receive and transmit electromagnetic radiation at frequencies from 900 MHz to 950 MHz for the first communication link;receive and transmit electromagnetic radiation at frequencies from 800 MHz to 900 MHz for the second communication link. 8. The device of claim 1 further comprising: a first matching network operatively connected to the first communications unit and the antenna; anda second matching network operatively connected to the second communications unit and the antenna. 9. An implantable medical device for use in a patient management system comprising: a housing of an implantable medical device;a first communications unit disposed within the housing configured to establish a first communications link within a short-range telemetry communication range of an external local device;a second communications unit disposed within the housing configured to establish a second communications link with a remote computer over at least one pervasive wireless communications network in the group consisting of: a wireless telephone network, a wireless wide area network, and a wireless local area network; anda dual mode antenna configured to be operatively connected to the first communication unit and the second communications unit. 10. The device of... | Cardiac Pacemakers, Inc. | A61N001/08
US20110034794 | A1 | 2011-02-10 | BLOOD GLUCOSE METER WITH INDEPENDENT VOICE IC SUPPORTING POST-PRODUCTION ADJUSTMENT OF VOICE FUNCTION | 2010-08-12 | A blood glucose meter with an independent voice IC supporting post-production adjustment of voice function is provided. The blood glucose meter is allowed to work with a replaceable voice carrier so as to provide voice services in different languages that help users to correctly and easily operate the blood glucose meter. The blood glucose meter features that the voice carrier is not built in the blood glucose meter during production of the blood glucose meter. Instead, a driving module is edited in an IC chip of the blood glucose meter. When an actual ordered amount is confirmed, the different voice carriers can be assembled to a voice-IC expansion port of the blood glucose meter, thereby reducing manufacturing and inventory costs for the blood glucose meter. | 1. A blood glucose meter with an independent voice IC supporting post-production adjustment of voice function, the blood glucose meter comprising: a button set located on a panel of the blood glucose meter for being used to input a personal information of a user;a test strip receiving unit for collecting and analyzing blood on a test strip;a micro processing unit for processing and storing an analytic result of the test strip receiving unit;a voice-IC expansion port for receiving an external voice carrier so as to expand the voice function;a driving module for editing a voice driver that supports voice carriers of different languages into a driving module that is then built on an IC for reading the external voice carrier;a voicing unit located on the panel of the blood glucose meter for giving out voice to assist the user in operating the blood glucose meter correctly; anda monitor located on the panel of the blood glucose meter for displaying a blood glucose value, a date, a time or a number of remaining test strips,wherein the blood glucose meter is characterized in that the voice carrier is configured to be assembled to the blood glucose meter after the blood glucose meter is produced and the driving module is built in the blood glucose meter. 2. The blood glucose meter with the independent voice IC of claim 1, further comprising a data transmitting module for transmitting data related to measurement of the user who has used the blood glucose meter to an external computer for storing and updating purposes. 3. The blood glucose meter with the independent voice IC of claim 1, wherein each said voice carrier carries 1 to 6 languages for option. 4. The blood glucose meter with the independent voice IC of claim 2, wherein the data transmitting module is configured to transmit the data to the external computer through infrared ray, RFID, Bluetooth, USB or other wired or wireless transmission means. 5. The blood glucose meter with the independent voice IC of claim 1, wherein the blood glucose meter has plural said voice carriers. | - | A61B005/145
US08409140 | B2 | 2013-04-02 | Injection apparatus | 2008-08-14 | An injection apparatus for making an injection at a predetermined depth in skin comprises: a first skin positioning member, a second skin positioning member, wherein the first and second skin positioning members lie or are moveable to lie in an injection arrangement; an injection needle; and an injection needle movement guide to guide the injection needle for movement from a parking position above the skin, through the lower surface of the second skin positioning member to slide beneath the first skin positioning member to an injection position. A further injection apparatus comprises: an injection needle; a plunger within the injection needle; and a retractor to retract the injection needle such that material to be injected is expelled from the injection needle by the plunger. | 1. An injection apparatus for making an injection at a predetermined depth in skin comprising: a first skin positioning member having a lower surface for positioning on a first patch of skin to hold the first patch of skin in a defined position;a second skin positioning member having a lower surface for positioning on a second patch of skin to hold the second patch of skin in a defined position, wherein the first and second skin positioning members lie or are moveable to lie in an injection arrangement;an injection needle having a discharge opening;an injection needle movement guide to guide the injection needle for movement from a parking position above the skin, through the lower surface of the second skin positioning member to enter the second patch of skin and to slide beneath the first skin positioning member to an injection position in which the discharge opening of the needle lies at a predetermined distance below the lower surface of the first skin positioning member;a plunger within the injection needle; anda retractor to retract the injection needle such that material to be injected is expelled from the discharge opening of the injection needle by the plunger. 2. An injection apparatus as claimed in claim 1, wherein the injection needle is guided for movement of the discharge opening of the needle at a constant distance below the lower surface of the first skin positioning member. 3. An injection apparatus as claimed in claim 1, wherein the first skin positioning member holds the first patch of skin at least substantially flat and/or the second skin positioning member holds the second patch of skin at least substantially flat. 4. An injection apparatus as claimed in claim 1, wherein in the injection arrangement the lower surfaces of the first skin positioning member and the second skin positioning member form a concave arrangement. 5. An injection apparatus as claimed in claim 1, wherein the first skin positioning member and second skin positioning member are adjacent to one another. 6. An injection apparatus as claimed in claim 1, wherein the first skin positioning member and the second skin positioning member are rotatably connected to one another. 7. An injection apparatus as claimed in claim 1, further comprising an actuator to effect relative movement of the first skin positioning member and the second skin positioning member. 8. An injection apparatus as claimed in claim 1, further comprising an attachment member to attach the first skin positioning member to the first patch of skin and/or the second skin positioning member to the second patch of skin. 9. An injection apparatus as claimed in claim 8, wherein the attachment member is adhesive. 10. An injection apparatus as claimed in claim 1, wherein the injection needle is guided for movement using a cam follower mounted to the injection needle engaging in a cam groove in a cam plate mounted for sliding movement with respect to the injection needle. 11. An injection appa... | Medtronic Minimed, Inc. | A61M005/00
US07878990 | B2 | 2011-02-01 | Gait training device and method | 2006-02-24 | A gait training device includes a pair of weight sensors with one of the sensors placed under each foot of an individual. The device also includes a microcomputer and a pair of electrical cables connecting each of the weight sensors to the microcomputer. A control panel/display is connected to the microcomputer and includes a pair of monitors for displaying target loads and/or actual loads averaged over a number of gait cycles for each foot. An audio alarm is also provided and sounds each time that a target load is exceeded. Methods for analyzing and correcting excess forces on an injured limb are also disclosed. | 1. A gait training device for detecting and correcting improper weight distribution during movement of an individual's body, in which said device comprises: two weight sensors with one of said weight sensors positioned under each of an individual's two feet and each of said weight sensors producing an output signal indicative of the weight on the sensor and wherein each of said weight sensors comprises an array of displacement transistors for converting mechanical pressure into an electrical signal;a microcomputer and means for inputting pre-selected target weights into said microcomputer;a control panel including a pair of window displays for displaying the forces applied by each leg of the individual and each of said window displays divided into two areas for displaying the time that each leg spends in a stance phase and a swing phase averaged over a number of gait cycles and said control panel including a pair of target weight selectors wherein each of said selectors include an up and down indicator for increasing or decreasing a target weight;a timer including an indicator, a start button and a reset button to measure the time of a session including multiple gate cycles on said control panel and a second set of buttons on said control panel for storing data and clearing said microcomputer in preparation for a different patient;a pair of cables connecting said weight sensors and said control panel to said microcomputer;said microcomputer including means for comparing the pre-selected weight and the sensed weight and computer memory that allows a clinician to monitor weight bearing data during a selected period of time, and also a speaker or sound buzzer that alerts a clinician and/or patient each time that a target weight is exceeded;means for indicating differences between pre-selected weights and sensed weights; andmeans for transferring the data from the gait training device to a personnel computer;which includes memory means for storing the weight distribution during multiple gait cycles and means for reviewing the memory for analyses by a clinician; andin which said weight sensors are Galium Arsenide displacement transistors. | - | A61B005/103
US07865236 | B2 | 2011-01-04 | Active electrode, bio-impedance based, tissue discrimination system and methods of use | 2005-10-19 | Systems and methods for discriminating and locating tissues within a body involve applying a waveform signal to tissue between two electrodes and measuring the electrical characteristics of the signal transmitted through the tissue. At least one of the electrodes is constrained in area so that localized electrical characteristics of the tissue are measured. Such localized electrical characteristics are determined over a portion of a body of the subject by using an array of electrodes or electrodes that can be moved over the body. A controller may implement the process and perform calculations on the measured data to identify tissue types and locations within the measured area, and to present results in graphical form. Results may be combined with other tissue imaging technologies and with image-guided systems. | 1. A method of discriminating a location of nerve tissue within a subject, comprising: placing a waveform electrode array on skin of the subject, wherein each electrode in the waveform electrode array has an area of no more than approximately 10 mm2;placing a return electrode on skin of the subject at a position on the skin removed from the waveform electrode array at a linear inter-electrode separation distance that falls within a tail region of an impedance v. distance curve for the subject;applying a signal serially to each of the electrodes in the waveform electrode array and the return electrode;measuring a change in a characteristic of the signal resulting from transmission through tissue between each electrode in the waveform array and the return electrode; anddiscriminating a location of nerve tissue located beneath the waveform electrode array by processing the measured change in the characteristics of the signal to identify anisotropic features in underlying tissue. 2. The method of discriminating a location of nerve tissue within a subject of claim 1, wherein the applied signal employs a single frequency between approximately 500 Hz and approximately 2500 Hz. 3. The method of discriminating a location of nerve tissue within a subject of claim 1, wherein the measured change in the characteristics of the signal comprises voltage. 4. The method of discriminating a location of nerve tissue within a subject of claim 1, wherein the measured change in the characteristics of the signal comprises current. 5. The method of discriminating a location of nerve tissue within a subject of claim 1, wherein the measured change in the characteristics of the signal comprises a phase shift. 6. The method of discriminating a location of nerve tissue within a subject of claim 1, wherein: measuring a change in a characteristic of the signal comprises measuring an electrical characteristic at discrete times to generate a data series; andprocessing the measured change in the characteristics of the signal to identify anisotropic features in underlying tissue comprises: using the data series to derive coefficients of a mathematical function approximating the signal resulting from transmission through nerve tissue between each individual electrode in the waveform electrode array and the return electrode; andusing the derived coefficients of the mathematical function to calculate an electrical characteristic of the nerve tissue between each individual electrode in the waveform electrode array and the return electrode. 7. The method of discriminating a location of nerve tissue within a subject of claim 6, wherein the characteristic of the signal resulting from transmission through tissue between each electrode in the waveform array and the return electrode is impedance. 8. The method of discriminating a location of nerve tissue within a subject of claim 6, wherein the characteristic of the signal resulting from transmission through tissue between each electr... | Nervonix, Inc. | A61B005/053
US07869864 | B2 | 2011-01-11 | Methods, systems and devices for detecting and diagnosing heart diseases and disorders | 2007-07-09 | Disclosed herein methods, devices, and systems for detecting and diagnosing a heart disease or disorder in a subject from a prime electrocardiogram which comprises calculating at least one distribution function of the prime electrocardiogram and determining whether the distribution function is indicative of the presence of absence of the heart disease or disorder. | 1. A method for diagnosing or detecting a heart disease or disorder in a subject which comprises creating a prime electrocardiogram from a raw electrocardiogram of the subject,obtaining harmonics from the prime electrocardiogram by performing Fourier analysis;obtaining a plurality of density functions for the harmonics in a range of harmonics of the prime electrocardiogram; wherein the density function refers to the sampling rate multiplied by the harmonic peak energy divided by the length of the period of the prime electrocardiogram;summing the plurality of density functions to give at least one distribution function; anddetermining the presence or absence of the heart disease or disorder based on whether the distribution function is indicative of the heart disease or disorder. 2. The method of claim 1, wherein the prime electrocardiogram is a prime morphological electrocardiogram. 3. The method of claim 2, wherein the prime morphological electrocardiogram is obtained by acquiring a raw electrocardiogram as electronic data from the subject;selecting at least one characteristic segment or at least one characteristic window; andisolating the characteristic segment or the characteristic window by removing information extraneous to the characteristic segment or the characteristic window. 4. The method of claim 3, wherein the characteristic segment is a QT segment, an RT segment, or a PR segment. 5. The method of claim 3, which further comprises refining the characteristic segment or the characteristic window by averaging the characteristic segment or the characteristic window. 6. The method of claim 5, which further comprises repeating the averaged characteristic window or the averaged characteristic segment by N times, wherein N is any positive integer. 7. The method of claim 3, wherein the isolated characteristic segment is a plurality of isolated characteristic segments in series. 8. The method of claim 3, wherein the isolated characteristic window is a plurality of isolated characteristic windows in series. 9. The method of claim 1, wherein the heart disease or disorder is myocardial ischemia. 10. The method of claim 2, which further comprises determining a secondary factor. 11. The method of claim 1, wherein the prime electrocardiogram is created using a beat marker. 12. The method of claim 11, wherein the beat marker is selected using a vectorcardiogram. 13. A system or device for detecting or diagnosing a subject as suffering from a heart disease or disorder comprising an electrocardiograph machine which obtains a raw electrocardiogram from the subject;a data acquisition module which converts the raw electrocardiogram into a raw digital electrocardiogram;a prime electrocardiogram converter which converts the raw digital electrocardiogram into a prime electrocardiogram; andmeans for calculating at least one distribution function of the prime electrocardiogram, andwherein the presence or absence of the heart disease or disorder is b... | Dynacardia, Inc. | A61B005/0456
US07869857 | B2 | 2011-01-11 | Combination magnetic resonance imaging apparatus and patient table | 2003-11-20 | A combination Magnetic Resonance imaging apparatus and patient table, which apparatus has a magnet structure, defining a cavity for accommodating a part of the body under examination, which is supported by a base block, the patient table and the apparatus having a base block and a supporting structure respectively, that are slidable in at least one, or more directions. The patient table and the Magnetic Resonance imaging apparatus have a mechanism for removable connection therebetween and guide thereof along predetermined relative displacement paths. | 1. A system comprising: a magnetic resonance imaging apparatus; anda patient table,wherein the magnetic resonance apparatus has a magnet structure defining a cavity for accommodating a part of a body under examination, which magnetic structure is supported by a magnetic structure base block,the patient table having a supporting structure that is slidable in at least one direction, wherein the patient table and the magnetic resonance imaging apparatus have a curved connection therebetween that allows relative rotation between the patient table and the magnet structure when connected to each other,the curved connection includes a guide for relative displacement between the patient table and the magnetic resonance imaging apparatus,wherein the base block of the magnetic resonance imaging apparatus is disposed on a platform interposed between the magnetic resonance imaging apparatus and a floor, which platform has a base plate and an upper magnetic resonance imaging apparatus supporting plate, which upper supporting plate lies over the base plate, rotary and sliding guide means for rotating and sliding the platform being interposed between said two plates such that the platform is rotatable along an annular path coaxial to an axis of the guide forming the curved connection between the patient table and the magnetic imaging apparatus and the upper support plate is slidable relative to the base plate, andwherein the patient table supporting structure has wheels or rollers for sliding the patient table relative to the magnetic resonance imaging apparatus. 2. The system of claim 1, comprising two or more patient tables that can be simultaneously coupled to the magnetic resonance imaging apparatus and moved in different positions relative thereto. 3. The system of claim 2, wherein the guide further includes a plurality of guide means for displacement of each of the two or more tables relative to each other and to the magnetic resonance imaging apparatus. 4. The system of claim 3, wherein the guide means comprises an arched guide and each table being connected to a carriage that can be coupled to said removable connection, wherein the tables have a supporting structure with wheels or rollers sliding. 5. The system of claim 1, wherein the patient table supporting structure is also disposed on the platform interposed between the magnetic resonance imaging apparatus and the floor. 6. The system of claim 1, wherein the rotary and sliding guide means interposed between the base plate and the upper plate of the platform is oriented along at least one straight axis. 7. The system of claim 1, wherein the guide of the curved connection includes a table sliding guide that has the shape of a sector of a circle, whose axis is oriented perpendicular to the floor or to the apparatus supporting surface. 8. The system of claim 7, wherein the rotary and sliding guide means between the base plate and the upper supporting plate of the platform for the magnetic resonan... | Esaote S.p.A. | A61B005/05
US07894880 | B2 | 2011-02-22 | Measurement of renal extraction fraction using contrast enhanced computed tomography | 2003-10-23 | Renal extraction fraction (EF) is determined through use of computed tomography (CT) measurements of arterial blood before and after injection of a radiographic contrast agent into the blood and CT measurements of renal vein blood after injection of the radiographic contrast agent. | 1. A method of measuring renal extraction fraction (EF) in a living subject using computed tomography comprising the steps of: a) obtaining a CT number (CTPRE) of arterial blood prior to addition of a radiographic contrast agent to the blood,b) providing a radiographic contrast agent to the blood,c) obtaining a CT number (CTA) of arterial blood after addition of the radiographic contrast agent to the blood,d) obtaining a CT number (CTV) of blood in a renal vein after addition of the agent to the blood, ande) determining renal extraction fraction (EF) from the obtained CT numbers. 2. The method as defined by claim 1 wherein renal extraction fraction (EF) is given by:EF=CTA-CTVCTA-CTPRE. 3. The method as defined by claim 2 wherein step b) includes providing iohexol. 4. The method as defined by claim 2 wherein step b) includes providing iothalamate. 5. The method as defined by claim 2 wherein step b) includes providing gadolinium-DTPA. 6. The method as defined by claim 1 wherein step b) includes providing iohexol. 7. The method as defined by claim 1 wherein step b) includes providing iothalamate. 8. The method as defined by claim 1 wherein step b) includes providing gadolinium-DTPA. 9. A method of determining renal extraction fraction (EF) for a kidney in a living subject, using a computed tomography (CT) apparatus comprising the steps of: a) obtaining a measure of x-ray transmission through arterial blood prior to addition of a radiographic contrast agent to the blood, using the CT apparatus,b) providing a radiographic contrast agent to the blood,c) obtaining a measure of x-ray transmission through arterial blood after addition of the radiographic contrast agent to the blood, using the CT apparatus,d) obtaining a measure of x-ray transmission through renal vein blood after addition of the radiographic contrast agent to the blood, ande) determining renal extraction fraction from the measures of x-ray transmission in steps a), and c), and d). 10. The method as defined by claim 9 wherein the measures of x-ray transmission are obtained using computed tomography (CT). 11. The method as defined by claim 10 wherein the measures of x-ray transmission are CT numbers. 12. The method as defined by claim 10 wherein the radiographic contrast agent is selected from the group consisting of iohexol, iothalamate, and gadolinium-DTPA. 13. The method as defined by claim 9 wherein the radiographic contrast agent is selected from the group consisting of iohexol and iothalamate. 14. The method as defined by claim 1 wherein the CT numbers (CTPRE, CTA, CTV) are each calculated as a mean value of measured image intensities for a plurality of voxels at different locations in one or more axial CT scans. 15. The method as defined by claim 14 wherein the CT number (CTV) of venous blood is calculated using voxels selected in a venous region to exclude arterial regions. 16. The method as defined by claim 14 wherein the renal function is calculated based on a diffe... | The Board of Trustees of the Leland Stanford Junior University | A61B005/05
US09445756 | B2 | 2016-09-20 | Dynamic integrated lancing test strip with sterility cover | 2010-09-22 | An integrated lancing test strip includes a test strip and a lancet packet coupled to the test strip. The lancet packet includes a sterility sheet enclosing a lancet to maintain the sterility of the lancet and prevent cross-contamination between the test strip and the lancet. The sterility sheet allows the lancet to be sterilized separately from the test strip. The sterility sheet gives the integrated strip a low profile, which is attractive for packaging multiple integrated strips in cassettes, drums, magazines or the like. In one form, the integrated strip is loaded in a meter that includes an adjustment mechanism that adjusts the position of the test strip relative to the skin being sampled. This allows the user to adjust the position of the test strip so as to not apply excessive pressure against skin, which could hamper bleeding from the incision in the skin. | 1. An apparatus, comprising: a test strip adapted to analyze body fluid; anda packet coupled to the test strip, wherein the packet includes an incision forming member to form an incision in tissue,a guide member with a guide slot in which the incision forming member is slidably disposed, wherein the guide slot defines a guide slot opening through which the incision forming member extends during lancing, the guide member having opposing sides and an end edge located between the opposing sides, wherein the guide slot opens at the opposing sides of the guide member, wherein the guide slot opening opens at the end edge of the guide member, anda sterility sheet having a top layer and a bottom layer, the top layer being sealed over a top surface of the guide member, and the bottom layer being sealed over a bottom surface of the guide member to enclose the incision forming member in the guide slot, wherein the top layer and the bottom layer of the sterility sheet cover the guide slot at the opposing sides and the guide slot opening, wherein at least one of the top layer or the bottom layer of the sterility sheet seals the guide slot opening and the guide slot of the guide member to maintain sterility of the incision forming member, wherein one of the top layer or the bottom layer of the sterility sheet is positioned between the incision forming member and the test strip. 2. The apparatus of claim 1, wherein the incision forming member defines an engagement opening for engaging the incision forming member during lancing. 3. The apparatus of claim 2, further comprising: a meter including a cam arm with a blade that pierces one of the top layer or the bottom layer of the sterility sheet to engage the engagement opening of the incision forming member, anda firing mechanism coupled to the cam arm for firing the incision forming member. 4. The apparatus of claim 3, wherein the meter includes an adjustment mechanism to adjust position of the test strip relative to the tissue during sampling for minimizing flow constriction of the body fluid from the incision. 5. The apparatus of claim 3, wherein the test strip defines a relief opening that is aligned with the engagement opening for allowing the blade to extend completely through the engagement opening in the incision forming member. 6. The apparatus of claim 2, wherein the test strip defines a relief opening that is aligned with the engagement opening. 7. The apparatus of claim 1, wherein the test strip defines a capillary opening in which the body fluid is drawn and a notch positioned proximal the packet to direct flow of the body fluid away from the incision forming member. 8. The apparatus of claim 1, wherein: the incision forming member includes a lancet; andboth the packet and test strip are flat to permit compact packaging. 9. The apparatus of claim 1, wherein the sterility sheet is hydrophobic for wiping the body fluid from the incision forming member during retraction. 10. The apparatus of cla... | Roche Diabetes Care, Inc. | A61B005/00 B65D081/00 A61B005/145 A61B005/15 A61B005/157
US09566018 | B2 | 2017-02-14 | Airway adaptor | 2010-07-15 | An airway adaptor includes: an airway case; and an expired gas guiding portion which is connected to the airway case to introduce a respiratory gas to the airway case. The airway case includes: a first respiratory gas flow path in which the respiratory gas flows in a first direction; and a second respiratory gas flow path in which the respiratory gas flows in a second direction opposite to the first direction. | 1. An airway adaptor comprising: an airway case to which a gas sensor is attachable;an expired gas guiding portion which is connected to the airway case to introduce a respiratory gas to the airway case;a partition wall;an elongated first passage adapted to introduce the respiratory gas from the expired gas guiding portion of the airway case; andan elongated second passage adapted to discharge the respiratory gas from the elongated first passage, which has been introduced to the airway case and returned in the airway case without being communicated with the expired gas guiding portion, to outside of the airway case,wherein the partition wall partitions the elongated first passage and the elongated second passage such that a portion of the elongated first passage is parallel to a portion of the elongated second passage, andwherein the respiratory gas in the portion of the elongated first passage flows in a first direction and the respiratory gas in the portion of the elongated second passage flows in a second direction opposite and in reverse with respect to the first direction. 2. The airway adaptor according to claim 1, wherein the expired gas guiding portion is detachably connected to the airway case. 3. The airway adaptor according to claim 1, wherein the airway case includes a first side face and a second side face which are opposed to each other, and a carbon dioxide sensor is attachable to the first side face and the second side face. 4. The airway adaptor according to claim 1, wherein the elongated first passage and the elongated second passage are juxtaposed to each other. 5. The airway adaptor according to claim 1, wherein one of the elongated first passage and the elongated second passage surrounds the other of the elongated first passage and the elongated second passage. 6. The airway adaptor according to claim 1, wherein a nose cup for surrounding a periphery of nostrils is disposed on one of the expired gas guiding portion and the airway case. 7. The airway adaptor according to claim 1, wherein a part of the airway adaptor is disposed on a nasal mask shell of a nasal mask in which positive pressure ventilation is enabled, and the airway case is projected to an outside of the nasal mask shell. 8. The airway adaptor according to claim 1, wherein the expired gas guiding portion is a nasal tube which is to be inserted into nostrils of a living body. 9. The airway adaptor according to claim 1, wherein the expired gas guiding portion includes: a first guiding portion which guides the respiratory gas from nostrils of a living body; anda second guiding portion which guides the respiratory gas from a mouth of the living body. 10. The airway adaptor according to claim 9, wherein the respiratory gas which is guided by the first guiding portion and the respiratory gas which is guided by the second guiding portion are joined together in the airway case or at an upstream side of the airway case. 11. The airway adaptor according to claim ... | NIHON KOHDEN CORPORATION | A61B005/08 A61B005/083 A61B005/097 A61B005/00
US09901288 | B2 | 2018-02-27 | Methods of detecting gaseous component levels in a breath | 2010-09-21 | Method for detecting gaseous component levels in a breath, comprising: receiving a breath through a breath channel, wherein the breath channel is in fluid communication with a flow rate sensor and an electrochemical fuel cell gas sensor; measuring a flow rate of the breath received through the breath channel; measuring a first time, wherein the first time corresponds to an amount of time elapsed while receiving the breath in the breath channel; and calculating a current gaseous component level utilizing the flow rate, first time and an output from the gas sensor. Methods for detecting an error condition while measuring gaseous component levels in a breath comprising: determining if the peak output occurs while breath is still being received in the breath channel; and if the peak output occurs while breath is still being received in the breath channel, alerting a user of an error condition. | 1. A method for detecting gaseous component levels in a breath, comprising: receiving a breath through a breath channel, wherein the breath channel is in continuous fluid communication with a flow rate sensor and an electrochemical fuel cell gas sensor while the entire breath is received, and wherein the flow rate sensor and the electrochemical fuel cell gas sensor are in electrical communication with a processing device;measuring a flow rate of the breath received through the breath channel utilizing the flow rate sensor and transmitting a corresponding flow rate signal to the processing device;measuring a gas level of the breath received through the breath channel with the electrochemical fuel cell gas sensor and transmitting a corresponding gas level signal to the processing device;measuring a first time, wherein the first time corresponds to an amount of time elapsed while receiving the breath in the breath channel;calculating a current gaseous component level on the processing device utilizing the flow rate signal, the first time and the gas level signal from the gas sensor, wherein the gaseous component level is proportional to the gas level signal, the flow rate and the first time. 2. The method of claim 1, wherein the gaseous component is ethanol. 3. The method of claim 1, wherein the flow rate sensor comprises a pressure sensor. 4. The method of claim 1, wherein the flow rate sensor comprises a temperature sensor. 5. The method of claim 1, wherein calculating the current gaseous component level further comprises calculating a volume of breath received based on the flow rate signal and the first time. 6. The method of claim 5, further comprising measuring the flow rate at a plurality of time intervals; and calculating the volume of breath based on the plurality of flow rates and time intervals by integrating the flow rates over the time intervals. 7. The method of claim 1, further comprising determining if the gas level signal from the gas sensor is a peak gas level signal. 8. The method of claim 7, further comprising: measuring a second time, wherein the second time corresponds to an elapsed time from the time corresponding to when the breath has ceased being received and the time corresponding to the peak gas level signal from the gas sensor;wherein calculating the current gaseous component level further comprises utilizing the second time. 9. The method of claim 7, further comprising: determining if the peak gas level signal occurs while breath is still being received in the breath channel; andif the peak gas level signal occurs while breath is still being received in the breath channel, alerting a user of an error condition. 10. A computer program product comprising a non-transitory computer readable medium carrying instructions for allowing a computer system to detect gaseous component levels in a breath received through a breath channel, the instructions comprising a method of: measuring a flow rate of the breath received t... | OmegaPoint Systems, LLC | A61B005/08 A61B005/097 B60K028/06 G01N033/497 A61B005/087 A61B005/117
US20110046456 | A1 | 2011-02-24 | Assembly Comprising Inserter, Cannula Part and Base Part | 2009-02-06 | The invention relates to an assembly comprising—an inserter (10) for insertion of a penetrating member (7) into a patient,—a penetrating member attached to the inserter (10) before use,—a base part comprises a position shaped to receive the penetrating member where the penetrating member can be attached to the base part during use and fastening means (2) which can provide for the base part to be attached to the patient, where the inserter (10) comprises inserter attachment means (14) locking the inserter (10) to the base part before and during insertion of the penetrating member (7) which inserter attachment means can be unlocked so that the inserted can be removed from the base part after insertion of the penetrating member (7), the inserter attachment means comprise at least one protruding part and at least one corresponding opening. Either the at least one protruding part or the at least one corresponding opening part is positioned on a surface of the base part and a corresponding part either comprising a protruding part or an opening is positioned on a surface of the inserter. | 1. An assembly comprising an inserter (10) for insertion of a penetrating member (7) into a patient,a penetrating member (7) attached to the inserter (10) before use which penetrating member (7) is provided with a through going opening forming a fluid connection between delivery means (8) and a patient when the penetrating member (7) is inserted,a base part comprises a position shaped to receive the penetrating member (7) where the penetrating member (7) can be attached to the base part during use and fastening means (2) which can provide for the base part to be attached to the patient,where the inserter (10) comprises inserter attachment means (14) locking the inserter (10) to the base part before and during insertion of the penetrating member (7) which inserter attachment means (14) can be unlocked so that the inserted can be removed from the base part after insertion of the penetrating member (7), the inserter attachment means (14, 14A, 14B, 14PR/R, 14PL/L) comprise at least one protruding part and at least one corresponding opening,characterized in that either the at least one protruding part or the at least one corresponding opening part is positioned on a surface of the base part and a corresponding part either comprising a protruding part or an opening is positioned on a surface of the inserter (10). 2. An assembly according to claim 1, characterized in that the inserter (10) is released from the base part i.e. the inserter attachment means (14) are disengaged, by the interaction between a moving part (38) of the inserter (10) which moving part guides the insertion of the penetrating member (7) from a first to a final position and a stationary part (29, 115, 14A) of the base part during insertion. 3. An assembly according to claim 2, characterized in that the moving part (38) of the inserter (10) moves forward in a direction which is approximately perpendicular i.e. 90°±45°, normally 90°±15° to the direction of insertion of the penetrating member (7). 4. An assembly according to any preceding claim, characterized in that the penetrating member (7) is attached by a tube shaped connection (16) to the base part before and during insertion. 5. An assembly according to claim 4, characterized in that the connection (16) comprises a flexible material. 6. An assembly according to any preceding claim, characterized in that the inserter (10) is released from the base part by applying a force to the inserter (10) or a part of the inserter (10) in a direction different from the direction of insertion of the penetrating member (7). 7. An assembly according to claim 5, characterized in that the force is applied in a direction which is approximately perpendicular i.e. 90°±45°, normally 90°±15° to the direction of insertion of the penetrating member (7). 8. An assembly according to any preceding claim, characterized in that the release of the inserter (10) from the base part is at least partly provided by the release of a force in direction towards... | - | A61B017/34 A61B005/00
US20110034811 | A1 | 2011-02-10 | METHOD AND SYSTEM FOR SLEEP/WAKE CONDITION ESTIMATION | 2009-04-09 | Insomnia is a prevalent sleep disturbance in the general population. As standard diagnostic method for assessing the nature and the severity of the sleep problem, a so-called sleep log or sleep diary, i.e. a questionnaire usually on paper, is used in most cases. The main drawback of this diagnostic tool is that its accuracy is affected by subjective bias of the patient, e.g. for patients it is often difficult to remember sleep and wake periods during the night correctly. The present invention proposes an automatic sleep log that uses vital body signs as input signals for assessing sleep and wake periods during the night. By using objective data, the diagnosis will be more accurate. Furthermore, this system can also be used to sleep restriction therapy, a non-pharmacological method to treat insomnia. Also in this application it can contribute to patients correctly applying this method and thus leading to a better medical outcome. | 1. Arrangement for the identification of sleep and wake conditions of a subject, comprising at least one sensor, which is integrated in the sleeping environment of the subject, the sensor is provided to generate subject related signals, andcalculation means, provided to receive the sensor signals,to extract at least one feature from these sensor signals,to classify the extracted feature, andto provide a probability indication whether the subject is awake or asleep. 2. Arrangement according to claim 1, wherein the at least one sensor is an ECG sensor, and/or a bed foil sensor sensing signals, which are related to the subject's body movements. 3. The arrangement according to claim 2, wherein the extracting of at least one feature from at least one measured signal derived from the sensor measuring the ECG signal and the bed foil sensor, which senses the body movements includes the cardio-respiratory condition and the coherent power spectrum calculated from both the ECG signal and bed foil sensor respiratory signal. 4. The arrangement according to claim 1, wherein the sensor is an ECG sensor and an ECG signal provided by the ECG sensor undergoes a pre-processing, wherein the preprocessing includes a R-peak detection, ectopic beat removal, linear interpolation and resampling to predefined frequency, and resulting in a RR-interval series. 5. The arrangement according to claim 4, wherein the RR-interval series is further processed using heart rate variability standards in the frequency and in the time domain. 6. The arrangement according to claim 1, wherein a power spectrum is calculated over a predefined first time duration, centred on a second predefined time duration epoch of interest, using a detrending model. 7. The arrangement according to claim 6, wherein the power spectrum is divided into a low-frequency band and a high frequency band and wherein the power spectrum is normalised. 8. according to claim 2, wherein an RR-interval and a breath interval series are combined by estimating the squared coherence function over predefined first time duration centred on a second predefined time duration epoch. 9. The arrangement according to claim 1, wherein the at least one feature form the components of a feature vector for classifying. 10. according to claim 9, wherein the feature classification is based on a standard pattern recognition approach with supervised learning, and wherein the classifier includes Bayesian linear or quadratic discriminant classifier,support vector machine,k-Nearest-Neighbour (kNN) method,Neural Network,Hidden Markov Model (HMM), andwherein the parameters of the classifier are trained on a large database of representative data. 11. A Method for the identification of sleep and wake conditions of a subject comprising the following steps: placing a sensor integrated in a sleeping environment, which is provided to generate subject related signals,receiving the sensor signals,extracting at least one feature from these senso... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/0205
US20110009729 | A1 | 2011-01-13 | APPARATUS FOR MEASURING PHYSIOLOGICAL SIGNALS | 2010-01-18 | Provided is an apparatus for measuring physiological signals by distributing modules according to functions. The physiological signal measuring apparatus includes first and second detection electrodes, a physiological signal detection module, a power supply module, and a connection line. The first and second detection electrodes are adhered to a target skin to detect physiological signals. The physiological signal detection module is installed at the first detection electrode to generate physiological signal data from the physiological signals detected by the first and second detection electrodes. The power supply module is installed at the second detection electrode to supply power to the physiological signal detection module. The connection line connects the physiological signal detection module and the power supply module. | 1. An apparatus for measuring physiological signals, comprising: first and second detection electrodes adhered to a target skin to detect physiological signals;a physiological signal detection module installed at the first detection electrode to generate physiological signal data from the physiological signals detected by the first and second detection electrodes;a power supply module installed at the second detection electrode to supply power to the physiological signal detection module; anda connection line connecting the physiological signal detection module and the power supply module. 2. The apparatus of claim 1, wherein each of the physiological signal detection module and the power supply module comprises a display unit displaying an operation state of the corresponding module and a control unit allowing a user to control the corresponding module. 3. The apparatus of claim 1, wherein the physiological signal detection module comprises a plurality of module connection ports. 4. The apparatus of claim 3, wherein the module connection ports comprise a data input/output port and a power port. 5. The apparatus of claim 1, wherein the connection line has flexibility and elasticity. 6. The apparatus of claim 1, wherein the physiological signal detection module comprises: an analog signal processing unit amplifying and filtering the physiological signal detected by the first detection electrode;an analog-to-digital converting unit converting an analog signal of the physiological signal into a digital signal of the physiological signal data; anda digital signal processing unit processing the physiological signal data through predetermined data operations. 7. The apparatus of claim 1, wherein the first and second detection electrodes have an adhesive material, which is adherable to a skin surface, on one surface and have a fixing member, at which the measurement module is installable, on the other surface. 8. The apparatus of claim 1, wherein the power supply module comprises a disposable battery or a rechargeable battery. 9. The apparatus of claim 1, wherein the physiological signals include electrocardiogram (ECG), electromyogram (EMG), electroencephalogram (EEG), and galvanic skin reflex (GSR). 10. The apparatus of claim 1, further comprising: a third detection electrode adhered to a target skin to detect a physiological signal; anda physiological signal transmission module installed at the third detection electrode to transmit the physiological signal data, which are detected by the physiological signal detection module, to an external device. 11. The apparatus of claim 10, wherein the physiological signal detection module generates the physiological signal data from the physiological signals detected by the first to third detection electrodes. 12. The apparatus of claim 10, wherein the physiological signal transmission module is connected to the physiological signal detection module and/or the power supply module. 13. The apparatus ... | Electronics and Telecommunications Research Institute | A61B005/04
US20110028830 | A1 | 2011-02-03 | DEVICE FOR GUIDING AN INVASIVE MEDICAL INSTRUMENT | 2009-01-15 | The present invention relates to a device (1) for guiding a invasive medical instrument (4), such a needle, a syringe, a biopter, or the like, relative to an imaging apparatus (3). The imaging apparatus is essentially non-invasive. The device of the invention comprises an accommodation (2) for the imaging apparatus. Also the device has a guide (5) adjacent the accommodation for guiding the medical device relative to the imaging apparatus. Thus accurate positioning of the medical instrument is enabled, while keeping the needle in the line of sight of the imaging apparatus. In particular a guide surface (7) is defined, and where opposite the guide surface at least partially resilient press means (8) are provided, which in use press the medical instrument against the guide surface. Thus medical instruments having different sizes or diameters can be employed and maneuvered within a plane of use corresponding with a line or field of sight of an imaging apparatus. The risk of any such medical instrument, like a needle, breaking is greatly reduced, while the freedom of operating the instruments is increased. The medical instrument can even temporarily be released by the physician or operator to perform additional therapeutic or diagnostic steps. | 1. Device (1) for guiding a invasive medical instrument (4), such a needle, a syringe, a biopter, or the like, relative to an imaging apparatus (3), which is essentially non-invasive, the device (1) comprising: an accommodation (2) for the imaging apparatus (3); anda guide (5) adjacent the accommodation (2) for guiding the medical device (4) relative to the imaging apparatus (3),wherein the guide (5) comprises: a guide surface (7) defining a plane for at least a portion of the medical instrument in use to abut there against and defining a larger plane of use for the medical instrument (4), said larger plane at least corresponding with an imaging zone of the imaging apparatus (3); andelongate press means (8) extending along the guide surface in at least one direction which essentially crosses a direction of use of the medical instrument to in use act on the medical instrument (4) at any location along the length of the press means (8) for pressing against the guide surface (7). 2. Device as claimed in claim 1, wherein the guide (5, 7, 8) defines a slot (6) which is oriented in such a direction to pass the medical instrument (4) therethrough for maneuvering the medical instrument (4) essentially in an imaging zone corresponding with the imaging apparatus (3). 3. Device as claimed in claim 1 or 2, wherein opposite the guide surface (7) an additional guide surface is provided, where the additional guide surface forms a carrier for the press means (8). 4. Device as claimed in claim 1, 2 or 3, wherein the elongate press means comprise a leg (8) extending over essentially the same length as the guide surface (7) in a direction away from the accommodation (3). 5. Device as claimed in any one of the preceding claims, wherein the press means comprise a leg (8) extending under an angle relative to the direction of a medical instrument passing along the guide surface. 6. Device as claimed in claim 5, wherein the leg is fixed relative to the guide surface at a position at a larger distance from the guide surface than the angled portions of the leg. 7. Device as claimed in claim 5 or 6, wherein the angled portions of the leg form a cantilever. 8. Device as claimed in any one of the preceding claims, wherein the bottom of the holder is closed with a member made from a material to allow waves from the imaging apparatus to pass there through. | GIESEN DESIGN CONSULTANCY B.V. | A61B005/00 A61B008/00
US20110015538 | A1 | 2011-01-20 | SYSTEM AND METHOD FOR ANALYZING ELECTROENCEPHALOGRAPHY DATA | 2009-07-17 | A system and method for analyzing EEG data. The system comprises data from at least one or all of the 19 (or more) standard scalp locations, a computing device 100 operable to calculate deviance for the data, and a single score for each of a plurality of scalp sites based on the deviance. An output device 10 presents a ranked list of the single scores to identify sites showing abnormal EEG, or a single score for all sites and pairs combined that may be used in assessment or operant conditioning. | 1. A diagnostic system analyzing EEG data comprising: single-site data from at least one single-site;a computing device operable to calculate a single-site deviance for the single-site data and a single score for each of a plurality of scalp sites based on the single-site deviance; andan output device ranking the single scores to identify sites showing abnormal EEG. 2. The system of claim 1, further comprising site-pair data from at least one site-pair, wherein the computing device is further operable to calculate a site-pair deviance for the site-pair data and the single score is calculated for each of a plurality of scalp sites based on the single-site deviance and corresponding site-pair deviance. 3. The system of claim 1, wherein the computing device is operable to export the single scores as output to EEG acquisition software. 4. The system of claim 1, wherein the single-site data is represented by at least one z-score. 5. The system of claim 1, wherein the site-pair data is represented by at least one z-score. 6. The system of claim 1, wherein deviance is calculated by taking the absolute value of the at least one z-score. 7. The system of claim 1, further comprising a deficit input for allowing an operator to identify specific sites where skull deficits exist, wherein the computing device is operable to exclude data from these sites. 8. The system of claim 1, wherein the data is static. 9. The system of claim 1, where the data is in real-time. 10. The system of claim 1, further comprising at least one cross-session report that combines input from multiple types of EEG measures. 11. The system of claim 10, wherein the EEG measures comprise Linked Ears. 12. The method as set forth in claim 10, wherein the EEG measures comprise Average Reference. 13. The method as set forth in claim 10, wherein the EEG measures comprise Laplacian. 14. The system of claim 1, wherein the corresponding site-pair deviance includes calculating deviance for each site-pair data that includes as one of the sites in the site-pair a site corresponding to the single-site for which a score is being calculated. 15. The system of claim 1, wherein the computing device is operable to export the single scores as output for integration with billing software. 16. The system of claim 1, wherein the single scores are operable to measure effects of medications. 17. The system of claim 1, wherein the single scores are operable to measure treatment response. 18. The system of claim 1, further comprising a display operable to graphically display a representation of the scores. 19. The system of claim 1, further comprising a database created and populated with information. 20. The system of claim 19, wherein information comprises the data, the single-site deviance, the site-pair deviance, and the single scores. 21. The system of claim 1, further comprising an interest input for allowing an operator to specify at least one site of interest, wherein scores from the a... | - | A61B005/0476
US07869848 | B2 | 2011-01-11 | Noninvasive measurement system | 2006-10-23 | The noninvasive measurement system provides a technique for manipulating wave data. In particular, wave data reflected from a biological entity is received, and the reflected wave data is correlated to a substance in the biological entity. The wave data may comprise light waves, and the biological entity may comprise a human being or blood. Additionally, a substance may comprise, for example, a molecule or ionic substance. The molecule may be, for example, a glucose molecule.;Furthermore, the wave data is used to form a matrix of pixels with the received wave data. The matrix of pixels may be modified by techniques of masking, stretching, or removing hot spots.;Then, the pixels may be integrated to obtain an integration value that is correlated to a glucose level. The correlation process may use a lookup table, which may be calibrated to a particular biological entity. Moreover, an amplitude and phase angle may be calculated for the reflected wave data and used to identify a glucose level in the biological entity. Additionally, the reflected wave data may be used to determine glaucoma pressure.;The glucose level may be displayed on a monitor attached to the computer. The computer may be a portable, self-contained unit that comprises a data processing system and a wave reflection capture system. On the other hand, the computer may be attached to a network of other computers, wherein the reflected wave data is received by the computer and forwarded to another computer in the network for processing. | 1. A method of correlating captured wave data to a substance in a biological entity, comprising: receiving wave data reflected from the biological entity;forming a plurality of pixels based on the reflected wave data;determining an amplitude for the reflected wave data based on the plurality of pixels;determining a phase angle for the reflected wave data based on the plurality of pixels; andidentifying a glucose level in the biological entity based on the amplitude and phase angle. 2. The method of claim 1, wherein the determination of the amplitude comprises calculating an integrated value. 3. The method of claim 1, wherein the biological entity is an eye. 4. The method of claim 3, wherein the step of determining the amplitude comprises integrating a portion of the plurality of pixels to form an integrated value; and the step of identifying the glucose level comprises correlating the amplitude and the phase angle to a glucose level in a look up table. 5. The method of claim 4, wherein the integrating step further comprises: forming a matrix with the plurality of pixels; and masking a portion of the matrix. 6. An apparatus for correlating captured wave data to a substance in a biological entity, comprising: means for receiving wave data reflected from the biological entity;means for forming a plurality of pixels based on the reflected wave data;means for determining an amplitude for the reflected wave data based on the plurality of pixels;means for determining a phase angle for the reflected wave data based on the plurality of pixels; andmeans for identifying a glucose level in the biological entity based on the amplitude and the phase angle. 7. The apparatus of claim 6, further comprising means for calculating an integrated value to obtain the amplitude. 8. The apparatus of claim 6, wherein the biological entity is an eye. 9. The apparatus of claim 8, wherein the means for determining the amplitude comprises means for integrating a portion of the plurality of pixels to form an integrated value; and wherein the means for identifying the glucose level comprises means for correlating the amplitude and the phase angle to a glucose level in a look up table. 10. The apparatus of claim 9, wherein the means for integrating comprises: means for forming a matrix with the plurality of pixels; andmeans for masking a portion of the matrix. 11. An apparatus for correlating captured wave data to a substance in a biological entity, comprising: a memory comprising a plurality of instructions; anda processing system coupled to the memory and configured to execute the plurality of instructions to:receive data corresponding with wave data reflected from the biological entity;form a plurality of pixels based on the reflected wave data;determine an amplitude for the reflected wave data based on the plurality of pixels;determine a phase angle for the reflected wave data based on the plurality of pixels; andidentify a glucose level in the biological entity ba... | STI Medical Systems, LLC | A61B005/1455
US07887493 | B2 | 2011-02-15 | Implantable device employing movement sensing for detecting sleep-related disorders | 2004-09-13 | Methods and systems for evaluating a pathological condition include acquiring movement information, such as electromyogram (EMG) information, and sleep disordered breathing (SDB) information, and detecting the presence of a pathological condition using both movement and SDB information. Methods may involve sensing physiological signals including at least muscle movement signals. Sleep-related disorders are detected using the sensed physiological signals, the sleep-related disorders including at least an involuntary muscle movement disorder and sleep-disordered breathing. Methods and systems also provide for detecting and treating a sleep-related disorder using movement and SDB information. Cardiac, respiratory, nerve stimulation, drug, or a combination of such therapies may be delivered to treat a detected or diagnosed pathological condition. | 1. A method of detecting sleep-related disorder events, comprising: sensing a plurality of physiological signals including at least muscle movement signals comprising electromyogram signals;detecting the sleep-related disorder events using the sensed physiological signals, the sleep-related disorder events comprising involuntary muscle movement events of non-respiratory muscles and sleep-disordered breathing events, the detection of the sleep-disordered breathing events and the involuntary muscle movement events based at least in part on the electromyogram signals; andcalculating a composite sleep disorder index, the composite sleep disorder index quantifying combined effect of the involuntary muscle movement events and the sleep-disordered breathing events, wherein at least one of calculating and detecting is implemented at least in part patient-internally using a circuit. 2. The method of claim 1, wherein sensing the physiological signals is performed at least in part patient-internally, and detecting the sleep-related disorder events is performed patient-externally. 3. The method of claim 1, wherein sensing the physiological signals is performed patient-externally, and detecting the sleep-related disorder events is performed at least in part patient-internally. 4. The method of claim 1, wherein sensing, detecting, and calculating are performed at least in part patient-internally. 5. The method of claim 1, wherein detecting the sleep-related disorder events comprises: detecting one of the sleep-related disorder events patient-internally; anddetecting another one of the sleep-related disorder events patient-externally. 6. The method of claim 1, wherein detecting the sleep disordered breathing events further comprises patient internally detecting the sleep disordered breathing events using transthoracic impedance signals. 7. The method of claim 6, wherein at least some of the electromyogram signals are sensed from one or more intramuscular locations. 8. The method of claim 1, wherein at least some of the physiological signals comprise transthoracic impedance signals sensed from a patient internal location. 9. The method of claim 1, further comprising detecting one or more sleep stages using the muscle movement signals. 10. The method of claim 1, wherein the sleep-disordered breathing events comprises one or more of sleep apnea, hypopnea, and Cheyne-Stokes respiration. 11. The method of claim 1, wherein the involuntary muscle movement events comprise one or more of a condition associated with bruxism, a condition associated with periodic limb movement disorder, and a condition associated with restless leg syndrome. 12. The method of claim 1, further comprising communicating one or both of the plurality of physiological signals and information associated with the detected sleep-related disorder events to a patient-external processing system. 13. The method of claim 1, further comprising communicating one or more of the plurality of phys... | Cardiac Pacemakers, Inc. | A61B005/08 A61B005/04 A61B005/05
US07869853 | B1 | 2011-01-11 | Analyte monitoring device and methods of use | 2010-08-06 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. An in vivo glucose monitoring system, comprising: an in vivo glucose sensor; anda receiver unit external to a user in signal communication with the in vivo glucose sensor, the receiver unit comprising: one or more processors; anda computer-readable storage medium including instructions which, when executed by the one or more processors, causes the one or more processors to: receive a data stream of sensor data points from the in vivo glucose sensor, the sensor data points indicative of a glucose level of the user;process the data stream of sensor data points indicative of the glucose level; anddetermine the acceleration of the rate of change of the glucose level. 2. The system of claim 1, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the one or more processors to determine whether the acceleration of the rate of change of the glucose level exceeds or meets one or more threshold values. 3. The system of claim 2, wherein the receiver unit further comprises an alarm, and wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the one or more processors to activate the alarm when the acceleration of the rate of change of the glucose level exceeds or meets the one or more threshold values. 4. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 5. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when a single datapoint indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 6. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when a predetermined number of datapoints obtained over a period of time indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 7. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when an average of datapoints obtained over a period of time indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 8. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm t... | Abbott Diabetes Care Inc. | A61B005/00
US07881779 | B2 | 2011-02-01 | Method and apparatus for synthetically detecting ventricular fibrillation | 2006-12-01 | The present application discloses a complexity-based method for synthetically detecting ventricular fibrillation, which centers on complexity calculations while incorporating a plurality of feature values and thus differentiates more effectively among various types of ECG signals. The method further modifies the complexity algorithm, making it more adapted to reflecting characteristics of the VF-related signals, thereby enabling high sensitivity and specificity of detection. Further, the related calculation load is reduced according to the algorithm. As such, the method can fully satisfy the clinical needs and is aimed for solving the problems of low sensitivity, low specificity and weak anti-interference ability present in current medical equipment for detecting ventricular fibrillation, such as monitors, implanted cardioversion defibrillator (ICD), automatic external defibrillator (AED) and so on. Systems for performing the method are also disclosed. | 1. A VF detector, comprising: a data sampler configured to sample a n number of samples of an ECG signal at a predetermined sampling rate;a filter configured to preprocess the samples so as to obtain data y(n);a standardizer configured to perform amplitude standardization to the data y(n) so as to obtain data y1(n);a numeric value calculator, configured to calculate parameters including a slope peak value SLMax, a probability density value PD, and a complexity value C(n) of the data y1(n); andan arbitrator configured to detect a VF in accordance with the calculated slope peak value SLMax, probability density value PD and complexity value C(n), respectively,wherein said numeric value calculator further comprises: a slope calculator, configured to calculate the slope peak value SLMax of said data y1(n);a probability density calculator, configured to calculate the probability density value PD of data y1(n); anda complexity calculation device, configured to calculate the complexity value of data y1(n). 2. The VF detector according to claim 1, wherein said arbitrator further comprises: a first threshold setting device, for setting a first slope threshold SLMaxJ1 and a second slope threshold SLMaxJ2 based on a result of the slope calculation, wherein SLMaxJ1>SLMaxJ2;a second threshold setting device, for setting a first amplitude probability density threshold PDJ1 and a second amplitude probability density threshold PDJ2 based on a result of the amplitude probability density calculation, wherein SLMaxJ1>SLMaxJ2;a third threshold setting device, for setting a complexity high threshold CHigh, a complexity low threshold CLow, and a complexity medium threshold CMid based on a result of the complexity calculation; anda comparator, for comparing the individual parameters to the corresponding thresholds to determine VF. 3. The VF detector according to claim 2, wherein said complexity calculation device further comprises: a sequence constructor, for performing 0-1 conversion to the data y1(n) so as to obtain a converted binary sequence S1, S2, . . . Sn; defining S=S1S2 . . . Sr, and Q=Sr+1 . . . Sr+i, thereby the length of the Q string Qlength=i, wherein r denotes a position where the first character in the present Q string is located in the sequence S1, S2, . . . Sn;a complexity calculator, for calculating a complexity of said sequence by:(a) designating SQ to represent a general character string formed of two concatenated strings S, Q, and SQV to represent a string with the last character of SQ deleted, i.e. SQV=S1S2 . . . SrSr+1 . . . Sr+i−1; and(b) determining whether Q is contained in SQV as follows:if Q is not contained in SQV, determining whether Qlength is larger than Qlengthmax, Qlengthmax being the maximum length of Q string among all previous Q strings: if so, adding Q to S, such that S=S1S2 . . . SrSr+1 . . . Sr+i, while Qlengthmax=Qlength=i, and the complexity C(n) increments;otherwise, adding Q and the subsequent Qlengthmax−Qlength characters t... | Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | A61B005/0402
US07885449 | B2 | 2011-02-08 | System for detecting infectious agents using computer-controlled automated image analysis | 2008-03-21 | A method for recognizing an object in an image is disclosed wherein a fractal map of the image is generated by estimating the fractal dimension of each pixel in the image. The fractal map may be segmented by thresholding and locations of candidate objects are determined. The pixel value of the image pixel corresponding to the same location where the candidate object is found in the fractal map may be compared to a threshold value. If the pixel value is greater than the threshold value, the candidate object is recognized as a valid object. | 1. An apparatus comprising: an array of objective lens components, wherein each of said objective lens component of the array of the objective lens components is configured to provide an image of a portion of a sample comprising cells and one or more infectious agents;an array of cameras, wherein each camera of said array of the cameras is configured to receive a corresponding image of a corresponding portion of the sample from a corresponding objective lens component of the array of the objective lens component and to provide an image signal in response to said corresponding image, wherein all image signals are provided simultaneously by said array of the cameras; andan image processor, responsive to said all image signals provided by said plurality of the cameras, configured to patch said all image signals to provide a composed image signal of said sample, and further configured to generate from said composed image signal one or more signals indicative of corresponding one or more detected infectious agents of said one or more infectious agents. 2. The apparatus of claim 1, wherein said each camera of said array of the cameras is a charge-couple device camera or a complementary metal oxide semiconductor image camera. 3. The apparatus of claim 1, wherein said composed image signal has a red-green-blue color format. 4. The apparatus of claim 1, wherein said cells are treated or untreated animal cells which comprise human or non-human cells of a tissue or of a body fluid, or said cells are plant cells. 5. The apparatus of claim 1, wherein said sample comprises a monolayer of said cells. 6. The apparatus of claim 1, wherein said corresponding one or more detected infectious agents of said one or more infectious agents are treated or untreated; and wherein said one or more infectious agents comprise one or more of: a Chlamydia, a Borrelia, a bacterium, a mycoplasma, a richkettsia, a spirochete, a fungus, a protozoal parasite, a virus and a prion. 7. The apparatus of claim 1, wherein each objective lens component of said array of lens components is configured to be focused individually on said sample. 8. The apparatus of claim 1, wherein said processor is further configured to generate a further signal indicative of animal cells in said samples, wherein said animal samples are treated or untreated. 9. The apparatus of claim 1, wherein said processor is further configured to determine one of or both: a number of infected animal cells relative to a total number of animal cells, and a number of infectious agents of said one or more infectious agents per animal cell. 10. The apparatus of claim 1, wherein said processor is further configured to determine a prognosis of susceptibility to or a diagnosis of a disease or a disorder associated with a presence of said detected one or more infectious agents, by using a relative abundance of said detected one or more infectious agents in combination with information related to a physical or physiological... | Ikonisys Inc. | G06K009/00 C12Q001/00 A61B005/00 A61B008/00 A61B010/00
US07862513 | B2 | 2011-01-04 | Apparatus for minimally invasive calibration of implanted pressure transducers | 2003-05-13 | This invention relates generally to apparatus and methods for the calibration of implanted pressure transducers. It is an object of several embodiments of the present invention to provide apparatus and methods for the calibration of one or more implanted pressure transducers implanted in the body of medical patients. Various embodiments of the present invention are particularly advantageous because they offer a calibration system that is less invasive than the systems currently available. | 1. A pressure measurement calibration system comprising: a first implantable sensor adapted to be implanted at a first location at least partially within a medical patient, wherein the first implantable sensor is configured to provide a first signal indicative of a first pressure at the first location at least partially within the medical patient;a pressure signal receiver for receiving the first signal from the first sensor;a second sensor for measuring a second pressure at a second location with respect to the medical patient and for generating a second signal indicative of said second pressure;a processor configured to determine an adjustment factor based on said first signal, said second signal, and a perturbation, wherein said first sensor is configured to be calibrated based upon said adjustment factor;a mouthpiece coupled to said second sensor; anda communications device configured to deliver a patient instruction to said medical patient when said first pressure falls outside of a previously determined range. 2. The system of claim 1, wherein said first sensor is further operable to be calibrated based upon an absolute pressure. 3. The system of claim 1, wherein said first sensor is further operable to be calibrated based upon a reference pressure. 4. The system of claim 3, wherein said reference pressure is the second pressure. 5. The system of claim 1, wherein said first sensor is further operable to be calibrated based upon a gauge pressure. 6. The system of claim 5, wherein said gauge pressure is the second pressure. 7. The system of claim 1, wherein said second sensor comprises a manometer. 8. The system of claim 1, further comprising a display for displaying the second pressure. 9. The system of claim 1, further comprising a signal conditioning apparatus. 10. The system of claim 1, further comprising a calibrator operable to adjust at least one calibration parameter. 11. The system of claim 1, further comprising a calibrator operable to calibrate said one or more first sensors. 12. The system of claim 1, further comprising a memory coupled to at least one of said first and second sensors, wherein said memory is configured to store at least one calibration parameter. 13. The system of claim 1, wherein said patient instruction is to perform a calibration procedure. 14. The system of claim 1, wherein said patient instruction is to recalibrate said one or more second sensors. 15. The system of claim 1, wherein said patient instruction is to contact a physician. 16. The system of claim 1, wherein said patient instruction is provided when a difference between said first pressure and said second pressure falls outside a previously determined range. 17. The system of claim 1, wherein said patient instruction is provided when a temperature falls outside of a previously determined range. 18. The system of claim 1, wherein said patient instruction is provided when a date falls outside of a previously determined range. 19. The ... | Pacesetter, Inc. | A61B005/00 A61B005/02 A61B005/08
US07894869 | B2 | 2011-02-22 | Multiple configuration medical sensor and technique for using the same | 2007-03-09 | A sensor may be adapted to be placed on multiple tissue sites. A sensor is provided that may have one configuration associated with use on a digit and a second configuration associated with use on another tissue site, such as a forehead. Further, a sensor may be adapted to be a transmission-type sensor or a reflectance-type sensor, depending on its configuration. | 1. A sensor comprising: a conformable sensor body having a transmission configuration adapted to be applied to a first tissue site and a reflectance configuration adapted to be applied to a second tissue site;an emitter and a detector disposed on the sensor body; anda removable portion of the sensor body indicated by a marker, wherein when the removable portion is associated with the sensor body, the sensor body is in the transmission configuration and when the removable portion is removed from the sensor body, the sensor is in the reflectance configuration. 2. The sensor, as set forth in claim 1, wherein the sensor comprises at least one of a pulse oximetry sensor or a sensor for measuring a water fraction. 3. The sensor, as set forth in claim 1, wherein the emitter comprises at least one light emitting diode and wherein the detector comprises at least one photodetector. 4. The sensor, as set forth in claim 1, wherein the first tissue site comprises a digit and wherein the second patient site comprises a forehead. 5. The sensor, as set forth in claim 1, wherein the first tissue site comprises an adult tissue site and wherein the second patient site comprises a pediatric tissue site. 6. The sensor, as set forth in claim 1, comprising an indicator adapted to provide an electronic signal relating to the transmission configuration or the reflectance configuration. 7. The sensor, as set forth in claim 6, wherein the indicator comprises a breakable conductive element, a temperature sensor, a resistor, or a combination thereof. 8. The sensor, as set forth in claim 1, wherein the marker comprises a visible or tactile marker. 9. The sensor, as set forth in claim 1, wherein the marker comprises a perforated marker. 10. A sensor comprising: a sensor body having a first portion and a second portion that is removable from the first portion, wherein the sensor body comprises a single unitary structure prior to removal of the second portion, and wherein the first portion of the sensor body is adapted to be applied to a first tissue site in a reflectance configuration after removal of the second portion and the first portion and the second portion of the sensor body, when together, are adapted to be applied to a second tissue site in a transmission configuration; andan emitter and a detector disposed on the sensor body. 11. The sensor, as set forth in claim 10, wherein the sensor comprises at least one of a pulse oximetry sensor or a sensor for measuring a water fraction. 12. The sensor, as set forth in claim 10, wherein the emitter comprises at least one light emitting diode. 13. The sensor, as set forth in claim 10, wherein the detector comprises at least one photodetector. 14. The sensor, as set forth in claim 10, wherein the emitter and the detector are disposed on the first portion of the sensor body. 15. The sensor, as set forth in claim 14, comprising a second emitter or a second detector disposed on the second portion of the sensor body. ... | Nellcor Puritan Bennett LLC | A61B005/1455
US08419652 | B2 | 2013-04-16 | Non invasive analysis of body sounds | 2009-03-03 | This method for analysing the sounds of body fluid flows includes:—simultaneously acquiring (2) sounds from various locations of a body;—identifying (6) the points of maximum sound intensity (PMIs) of the acquired sounds for each acquisition instant;—determining (10) the source locations of the acquired sounds; and—determining (12, 14) the sound radiation patterns of the acquired sounds. A corresponding device, system and program perform this method. | 1. A method for analysing the sounds of body fluid flows comprising: repeatedly and simultaneously acquiring sounds from a plurality of locations of a body with a plurality of acoustic sensors distributed on the body; andwith a processor, determining characteristics of the body fluid flows, including: identifying points of the maximum sound intensity (PMIs) of the acquired sounds for each of a plurality of sound acquisitions;determining source locations of the acquired sounds; anddetermining sound radiation patterns of the acquired sounds. 2. The method according to claim 1, wherein determining the source locations of the acquired sounds comprises analysing several acquisition instants and determining the maximum of several points of maximum sound intensity as a source location. 3. The method according to claim 2, further comprising: with the processors, defining sound acquisition cycles and/or regions based on a periodicity and repetition of the points of maximum sound intensity. 4. The method according to claim 1, further comprising: with the processor, comparing the determined fluid flow characteristics with existing reference flow characteristics and providing a match likelihood of the characteristics of the acquired sounds with said existing reference flow characteristics. 5. The method according to claim 1, wherein said body fluid flows are human heart blood flows. 6. The method according to claim 1, further including: on a display, displaying an image depicting the points of maximum sound intensity of the sounds acquired from the body. 7. The method according to claim 1, further including: on a display, displaying image frames depicting the identified points of maximum sound intensity, the determined source locations, and the determined sound radiation patterns. 8. The method according to claim 1, further including: forming a plurality of digital image frames, each digital image frame being formed from the determined source locations and the identified points of maximum sound intensity for each of the plurality of simultaneously acquired sound acquisitions. 9. The method according to claim 8, further including: determining the sound radiation patterns from the plurality of digital image frames. 10. A processor programmed to perform the method according to claim 1. 11. A method for analyzing the sounds of body fluid flows comprising simultaneously acquiring sounds from a plurality of locations of a body with an acoustic transducer and determining characteristics of the body fluid flows, wherein determining the characteristics of the fluid flows comprises, with a processor: identifying the points of maximum sound intensity (PMIs) at easch of a plurality of times that the sounds are acquired;determining source locations of the acquired sounds;creating a digital image mapping the points of maximum sound intensity each time the sounds are acquired;tracking the points of maximum sound intensity over a consecutive plurality of the image... | Koninklijke Philips Electronics N.V. | A61B005/08
US08727983 | B2 | 2014-05-20 | Systems and methods for making and using a conductive-fluid detector for a catheter-based medical device | 2010-08-12 | A catheter-based medical device includes a catheter that is configured and arranged for at least partial insertion into a patient and that defines at least one lumen that is configured and arranged to receive a first fluid. A conductive-fluid detector is coupled to the at least one lumen and is configured and arranged to detect when a second fluid is disposed within the at least one lumen that is more conductive than the first fluid. The conductive-fluid detector includes a plurality of axially-positioned bodies, each body defining a lumen. The lumens of the axially-positioned bodies are aligned to form a shared lumen in fluid communication with the at least one lumen of the catheter. Spaced apart electrodes are disposed within the shared lumen. | 1. A catheter-based medical device comprising: a catheter configured and arranged for at least partial insertion into a patient, the catheter defining at least one lumen, the at least one lumen configured and arranged to receive a first fluid; anda conductive-fluid detector coupled to the at least one lumen, the conductive-fluid detector configured and arranged to detect when a second fluid is disposed within the at least one lumen that is more conductive than the first fluid, the conductive-fluid detector comprising a first electrode having a proximal end, a distal end, and comprising an inner surface and an outer surface,a second electrode having a proximal end, a distal end, and comprising an inner surface and an outer surface, anda fluid pathway coupled to the at least one lumen of the catheter, the fluid pathway comprising a plurality of axially-separated bodies, each body having an inner surface and defining a lumen, the lumens aligned to form a shared lumen in fluid communication with the at least one lumen of the catheter, the fluid pathway configured and arranged to receive the first electrode and the second electrode,wherein the fluid pathway comprises a first elongated body having a proximal end disposed over the distal end of the outer surface of the first electrode such that the distal end of the outer surface of the first electrode couples to the inner surface of the first elongated body and the inner surface of the first electrode is disposed in the shared lumen,wherein the fluid pathway comprises a second elongated body having a distal end disposed over the proximal end of the outer surface of the second electrode such that the proximal end of the outer surface of the second electrode couples to the inner surface of the second elongated body and the inner surface of the second electrode is disposed in the shared lumen, andwherein the fluid pathway comprises a connecting body having a proximal end and a distal end, the distal end of the inner surface of the connecting body coupled to the proximal end of the outer surface of the first electrode and the proximal end of the inner surface of the connecting body coupled to the distal end of the outer surface of the second electrode such that the first electrode and the second electrode are electrically separated from one another, and the first and second elongated bodies are axially separated by the connecting body. 2. The catheter-based medical device of claim 1, wherein the outer surface of at least one of the first electrode or the second electrode comprises a plurality of barbs configured and arranged to extend into the inner surface of at least one of the first elongated body, the second elongated body, or the connecting body without piercing the first elongated body, the second elongated body, or the connecting body into which the plurality of barbs extend. 3. The catheter-based medical device of claim 1, wherein at least one of the first electrode or the second electrode is one... | Boston Scientific Scimed, Inc. | A61M005/172 A61M005/44 A61B018/02 A61N007/02
US08706195 | B2 | 2014-04-22 | Method for producing an electrophysiological map of the heart | 2008-05-11 | Method and system for producing an electrophysiological map of a heart of the body of a patient, the method including, for at least one target point, determining a respective target point location and a respective probe orientation, confirming that the tip of a probe is located at the respective target point location, confirming that the tip is oriented at the respective probe orientation, measuring at least one heart parameter value at the respective target point location, and superimposing at least one representation of the at least one heart parameter value on an image of the heart, at the respective target point location, to produce the electrophysiological map. | 1. An apparatus, comprising: a medical positioning system (MPS) comprising circuitry to detect a current position and a current orientation of an MPS sensor coupled to a tip portion of a probe catheter;a first processor comprising circuitry to electrically communicate with said MPS and with an image detector configured to detect a first image of a heart, said first processor being adapted to register the MPS with the image detector;a second processor comprising circuitry to determine a respective tissue target location and a respective target probe orientation, at which at least one heart parameter value is to be measured for each of a plurality of tissue targets within said heart;a third processor comprising circuitry to confirm that said tip portion is located at said respective tissue target location by comparing a currently detected location of said tip portion with said respective tissue target location;a fourth processor comprising circuitry to confirm that said tip portion is oriented at said respective target probe orientation by comparing a currently detected orientation of said tip portion with said respective target probe orientation;a heart parameter measurement unit configured to measure said at least one heart parameter value via a heart parameter sensor coupled to said tip portion; anda fifth processor comprising circuitry to superimpose said at least one heart parameter value at said respective tissue target locations on a second image of said heart to produce an electrophysiological map of said heart. 2. An apparatus according to claim 1, wherein the at least one heart parameter comprises one of: a local activation time (LAT), a respiration artifact, and a change in electrical potential. 3. An apparatus according to claim 1, wherein at least one of the first image of said heart and the electrophysiological map of said heart comprises a two dimensional depiction. 4. An apparatus according to claim 1, wherein at least one of the first image of said heart and the electrophysiological map of said heart comprises a three dimensional depiction. 5. An apparatus according to claim 1, further comprising a sixth processor comprising circuitry to register a plurality of superimposed images. 6. An apparatus according to claim 1, wherein said first, second, third, fourth, and fifth processors comprise a single processor. 7. An apparatus according to claim 1, further comprising a stimulation catheter comprising circuitry to stimulate a portion of the heart. 8. An apparatus according to claim 7, wherein the stimulation catheter comprises one of an elongate temporary pacing lead and an electrophysiology catheter. 9. An apparatus according to claim 8, wherein the pacing lead and the electrophysiology catheter are adapted to couple to a location on or near a coronary sinus structure of the heart. 10. An apparatus according to claim 1, wherein at least a portion of one of the first image of said heart and the electrophysiological map of sa... | MediGuide Ltd. | A61B005/00
US08480969 | B2 | 2013-07-09 | Medical device and cover structure | 2010-10-01 | A medical device includes a casing, a cover member that covers a portion on one-direction side of the same, and a sheet member interposed therebetween. The cover member includes on its periphery a hook part having a hook-like portion and a protrusion. The casing includes a latch part configured to engage with the hook part and a fitted portion for the protrusion to fit in. The sheet member includes an alignment part for aligning itself with the cover member. The cover member preferably has optical transparency. The casing preferably includes a recess shaped to fit with an outer shape of the cover member at a portion to be covered by the cover member. | 1. A medical device that measures a condition of a living body, comprising: a casing including a latch part; anda cover member that includes a hook part to engage with the latch part and covers a portion on one-direction side of the casing, the hook part having a hook-like portion;one of the casing and the cover member including a protrusion, and the other including a fitted portion for the protrusion to fit in, whereinthe casing includes a recess shaped to fit with an outer shape of the cover member at a portion to be covered by the cover member,the latch part is provided in a position of a side wall of the recess corresponding to the hook part,the fitted portion is provided in a position of the side wall corresponding to the protrusion,the hook part is shaped such that a tip of the hook-like portion faces an opposite side from a casing side when the cover member is arranged on the casing,the recess is configured to allow the hook part to move to a bottom surface side of the recess when the hook part of the cover member is pushed,the cover member is in a form of a plate and includes two opposing side surfaces, each of the two side surfaces having the protrusion,the hook part is arranged on a virtual axis being parallel to the two side surfaces and passing between the protrusions, andthe side wall of the recess is shaped such that a portion located on the virtual axis when the cover member is attached and opposing a position corresponding to the hook part is lower than another portion. 2. The medical device according to claim 1, wherein the latch part includes a hole arranged in a position of the side wall of the recess corresponding to the hook part and a projection provided inside the hole,the hole is provided so as to allow insertion of the hook-like portion, andthe projection is configured to engage with the hook-like portion. 3. The medical device according to claim 1, wherein the latch part includes a hole arranged in a position of the side wall of the recess corresponding to the hook part and a pair of projections provided inside the hole,the hole is provided so as to allow insertion of the hook-like portion, andthe pair of projections are configured such that the hook-like portion is located between the pair of projections and at least one of the pair of projections engage with the hook-like portion. 4. The medical device according to claim 1, wherein the fitted portion includes a first groove and a second groove in a position of the side wall corresponding to the protrusion,the first groove is provided from an opening of the recess along the depth of the recess, andthe second groove is provided so as to intersect with the first groove. 5. The medical device according to claim 1, further comprising: a sheet member interposed between the casing and the cover member,the sheet member including an alignment part for aligning itself with the cover member. 6. A cover structure comprising: a base including a latch part; anda cover member th... | Arkray, Inc. | G01N021/75 B65D085/38
US09336352 | B2 | 2016-05-10 | Power injector with help functionality | 2009-02-27 | A power injector (10) with help functionality is disclosed. What may be characterized as a ‘help output’ may be displayed on a graphical user interface (11) in response to a request for help. This help output may be based upon a current state of the power injector (10), may be accompanied by a cross-reference to an operator's manual for the power injector (10), may be expanded to provide additional detail, or any combination thereof. At least one system screen that relates to the help output may be displayed (e.g., simultaneously, sequentially). The help functionality may be in the form of having the power injector (10) incorporate help mode logic (256) that is configured to allow a user to view and obtain information on the various system screens that would be presented when using the power injector (10) to deliver a fluid, but without actually operating the power injector (10) to deliver fluid. | 1. A power injector, comprising: a powerhead, wherein a syringe may be installed on said powerhead;a syringe plunger drive assembly;a first graphical user interface;a fluid delivery mode where fluid is delivered by said power injector;a help mode, wherein said fluid delivery mode is suspended such that fluid is not deliverable by said power injector when said power injector is in said help mode; anda help button, wherein activation of said help button suspends said fluid delivery mode and instead disposes said power injector in said help mode, wherein a selected help topic relates to at least one aspect of said power injector in conjunction with execution of a fluid delivery procedure, and wherein said selected help topic thereafter causes said help mode to generate each of the following on said first graphical user interface: a first help output that relates to said selected help topic;each system screen: i) that is presented on said first graphical user interface when said power injector is in said fluid delivery mode and while said power injector is executing a fluid delivery procedure; and ii) that relates to said selected help topic, wherein said first help output is separate from and in addition to each said system screen that is presented on said first graphical user interface while said power injector is in said help mode, and wherein each system screen that is generated on said first graphical user interface when said power injector is in said fluid delivery mode pertains to said power injector in conjunction with execution of said fluid delivery procedure; andat least one cross-reference that provides additional help information on said selected help topic:wherein said first help output is separate from and in addition to each said system screen that is presented on said first graphical user interface when said power injector is in said help mode;wherein each said system screen that is displayed on said first graphical user interface during operation of said power injector in said help mode is also displayed on said first graphical user interface during operation of said power injector in said fluid delivery mode; andwherein said first help output and one of said system screens are simultaneously presented on said first graphical user interface, from said power injector being in said help mode, and in response to said selected help topic. 2. The power injector of claim 1, wherein said activation of said help button presents a help screen on said first graphical user interface, wherein said help screen presents a plurality of help topics, and wherein said selected help topic is from said plurality of help topics and said selected help topic is generated by user input to said power injector. 3. The power injector of claim 1, wherein said selected help topic is based upon a current state of said power injector when said help button is activated. 4. The power injector of claim 3, wherein said current state is selected from the group cons... | LIEBEL-FLARSHEIM COMPANY LLC | G06F019/00 A61M005/145 A61B006/00
US09386927 | B2 | 2016-07-12 | Implantable blood pressure monitor | 2009-07-28 | A method for monitoring blood pressure includes sensing and storing sympathetic nerve activity data of a patient via a recording lead of an implantable medical device. Changes in sympathetic nerve activity from the nerve activity data are determined. Corresponding changes in blood pressure are determined from the changes in sympathetic nerve activity. An alert signal and/or modification of therapy can be provided. | 1. A method for monitoring blood pressure comprising: implanting a medical device including a processor and a storage module in a patient;coupling a first end of a recording lead to the medical device;inserting a second end of the recording lead adjacent a sympathetic nerve ending;transmitting sympathetic nerve activity via the recording lead to the medical device;recording sympathetic nerve activity in the storage module as stored nerve activity data;determining changes in sympathetic nerve activity from the stored nerve activity data; anddetermining corresponding changes in blood pressure from the changes in sympathetic nerve activity,wherein inserting the second end of the recording lead adjacent a sympathetic nerve ending includes positioning the second end of the recording lead at a post-gangliotic nerve ending for a kidney. 2. The method of claim 1, further comprising: comparing changes in sympathetic nerve activity to a threshold value for the patient; andproviding an alert signal indicative of a blood pressure status. 3. The method of claim 2, wherein providing an alert signal includes providing an alert signal to the patient. 4. The method of claim 2, wherein providing an alert signal includes wirelessly transmitting data to a medical center. 5. The method of claim 1, further comprising: comparing changes in sympathetic nerve activity to a threshold value for the patient; andproviding nerve stimulation therapy via a therapy lead electrically coupled to a pulse generator of the medical device. 6. The method of claim 1, further comprising wirelessly transmitting the stored nerve activity data from the medical device to a medical center via an external handheld medical monitor. 7. The method of claim 1, further comprising wirelessly communicating with the medical device via an external telemetry programmer. | Medtronic, Inc. | A61B005/08 A61B005/0215 A61B005/04 A61N001/36 A61B005/00 A61N001/365
US20110021936 | A1 | 2011-01-27 | Medical data display with 3-D and 2-D color mapping | 2009-07-24 | A device and method for displaying medical data for a stress test monitoring system. A computer-implemented method and apparatus for displaying electrical impulse data of a human heart for analysis includes receiving cardiac electrical impulse data produced during a cardiac stress test by a plurality of electrode leads adapted for placement on a patient and calculating parameters for a plurality of display windows based on the data and providing a main display in which the plurality of display windows are positioned. A three dimensional color mapping plot is displayed in one of the plurality of display windows and a two dimensional color mapping plot is displayed in another of the plurality of display windows. Additionally, a plot of raw lead data from at least one lead is displayed in one of the plurality of display windows. | 1. A computer-implemented method for displaying electrical impulse data of a human heart for analysis, comprising the steps of: receiving data from one or more input files, wherein the data is cardiac electrical impulse data acquired during a cardiac stress test from a plurality of electrode leads adapted for patient attachment;calculating parameters for a plurality of display windows based on said data;providing a main display in which the plurality of display windows are positioned;displaying a three dimensional color plot in one of the plurality of display windows based on said data and said parameters;displaying a two dimensional color plot in one of the plurality of display windows based on said data and said parameters, the two dimensional color plot and the three dimensional color plot correspond with one another and use color to represent ST level; anddisplaying a plot of raw lead data from at least one lead in one of the plurality of display windows. 2. The computer-implemented method of claim 1, wherein the three dimensional color plot may be interactively rotated on the main display such that a user may select a desired viewpoint to observe the three dimensional color plot. 3. The computer-implemented method of claim 1, wherein the lead data is shown on the two dimensional color plot and the three dimensional color plot in anatomic lead order. 4. The computer-implemented method of claim 1, wherein the color shown in the two dimensional color plot and the three dimensional color plot represents ST level. 5. The computer-implemented method of claim 1, further including a two dimensional color map located adjacent to the two dimensional color plot on the main display, the two dimensional color map having an independent color scale. 6. The computer-implemented method of claim 1, further including generating a printable one page stress test report containing the information contained on the main display. 7. The computer-implemented method of claim 1, further including displaying a plot of raw channel data and a representative beat plot in the main display corresponding to a user selected location on the two dimensional plot. 8. A medical apparatus for generating a stress test display, comprising a data acquisition module adapted for connection to a patient undergoing a cardiac stress test;a monitor capable of presenting a real time display of color mapped plots;a computer, comprising at least a processor and a memory, coupled to the data acquisition module and the monitor, said computer containing exercise stress test software, operation of the software including: means for providing a display for reviewing information obtained during a cardiac stress test;means for receiving data, wherein the data is cardiac electrical impulse data produced during the cardiac stress test by a plurality of electrode leads adapted for placement on a patient;means for calculating values for generating plots based on said data;means for displaying a thr... | - | A61B005/044 G09G005/02 G06T011/20
US20110028801 | A1 | 2011-02-03 | CARDIOVASCULAR VESSEL ELASTICITY MONITORING | 2009-07-29 | Elasticity of a cardiovascular vessel of a patient is monitored to provide an indication of whether the patient's health is changing. In some aspects the elasticity of a cardiovascular vessel is determined irrespective of the current blood pressure level of the patient at the time the elasticity is determined. For example, vessel elasticity may be determined based on a defined set of data that maps vessel elasticity with reflectance times for different blood pressure levels. In some implementations, this set of data corresponds to a set of iso-pressure lines. | 1. An implantable medical monitoring apparatus, comprising: a reflectance time detector configured to determine reflectance time of a pressure wave in a cardiovascular system of a patient;a blood pressure detector configured to determine blood pressure of the cardiovascular system; anda cardiovascular vessel elasticity detector configured to determine elasticity of a vessel of the cardiovascular system based on the determined blood pressure, the determined reflectance time, and a data set that correlates different blood pressure values and different pressure wave reflectance time values to different vessel elasticity values. 2. The apparatus of claim 1, wherein the data set comprises a set of iso-pressure lines that map reflectance time to elasticity. 3. The apparatus of claim 1, further comprising a photoplethysmography sensor configured to generate signals indicative of changes in blood flow in the cardiovascular system, wherein the reflectance time detector is further configured to process the generated signals to determine the reflectance time. 4. The apparatus of claim 3, wherein the blood pressure detector is further configured to process the generated signals to determine the blood pressure. 5. The apparatus of claim 1, wherein the cardiovascular vessel elasticity detector is further configured to determine a trend indicative of changes in elasticity of the vessel based on the determined elasticity. 6. The apparatus of claim 5, further comprising a therapy module configured to identify at least one condition associated with the patient based on the trend. 7. The apparatus of claim 6, wherein the at least one condition relates to at least one of the group consisting of: an indication of heart failure, an indication of hypertension, an indication of neurohormonal insult leading to wall stiffness, medical regimen compliance, and salt intake compliance. 8. The apparatus of claim 1, wherein: the data set comprises a set of iso-pressure lines; andthe cardiovascular vessel elasticity detector is further configured to determine a trend indicative of changes in elasticity of the vessel based on the determined elasticity. 9. The apparatus of claim 8, further comprising a photoplethysmography sensor configured to generate signals indicative of changes in blood flow in the cardiovascular system, wherein: the reflectance time detector is further configured to process the generated signals to determine the reflectance time; andthe blood pressure detector is further configured to process the generated signals to determine the blood pressure. 10. A method of monitoring vessel elasticity, comprising: determining reflectance time of a pressure wave in a cardiovascular system of a patient;determining blood pressure of the cardiovascular system; anddetermining elasticity of a vessel of the cardiovascular system based on the determined blood pressure, the determined reflectance time, and a data set that correlates different blood pressure values and d... | PACESETTER, INC. | A61B005/0205 A61B005/02
US20110028821 | A1 | 2011-02-03 | Electrocardiographic Monitoring System and Method Using Orthogonal Electrode Pattern | 2009-07-31 | A system for monitoring a cardiac condition of a patient includes a diagnostic center configured to construct a 12-lead ECG of a patient using a special ECG signals numbering less than twelve by combining the special ECG signals with a transformation matrix, and a wearable device configured to generate the special ECG signals and including. The wearable device includes a belt having one or more belt electrodes, a waistband having one or more waistband electrodes, the belt and waistband electrodes configured to contact the skin of the patient and obtain electrical signals therefrom, and a host unit in electrical communication with the belt and waistband electrodes, the host unit including circuitry for generating the special ECG signals from one or more of the acquired electrical signals and circuitry for special ECG signals to a location remote from the wearable device. | 1. A wearable device configured to generate special ECG signals for constructing a 12-lead ECG, the special ECG signals numbering less than twelve and being combinable with a calibration matrix in order to construct the 12-lead ECG, the wearable device comprising: a belt having one or more belt electrodes;a waistband having one or more waistband electrodes, the belt and waistband electrodes configured to contact the skin of a wearer and obtain electrical signals therefrom; anda host unit in electrical communication with the belt and waistband electrodes, the host unit including circuitry for generating the special ECG signals from one or more of the acquired electrical signals and circuitry for transmitting information based on the special ECG signals to a location remote from the wearable device. 2. The device of claim 1, wherein the total number of skin-contacting electrodes is four. 3. The device of claim 1, wherein the total number of skin-contacting electrodes is five. 4. The device of claim 1, wherein the belt contains three belt electrodes. 5. The device of claim 1, wherein the waistband contains two waistband electrodes. 6. The device of claim 1, wherein the host unit contains at least one host unit electrode configured to contact the skin of the wearer and obtain electrical signals therefrom. 7. The device of claim 1, wherein at least one of the belt electrodes is a common electrode. 8. The device of claim 1, wherein at least one of the belt electrodes is a ground electrode. 9. The device of claim 1, wherein at least one of the waistband electrodes is a common electrode. 10. The device of claim 1, wherein at least one of the waistband electrodes is a ground electrode. 11. The device of claim 1, wherein transmission is conducted in real time. 12. The device of claim 1, wherein transmission is conducted at prescribed time intervals. 13. The device of claim 1, wherein transmission comprises downloading the information using a dedicated connection. 14. A bedside cardiac monitoring system comprising: a diagnostic center configured to construct a 12-lead ECG of a patient using a special ECG signals numbering less than twelve by combining the special ECG signals with a transformation matrix; anda wearable device configured to generate the special ECG signals and including: a belt having one or more belt electrodes;a waistband having one or more waistband electrodes, the belt and waistband electrodes configured to contact the skin of the patient and obtain electrical signals therefrom; anda host unit in electrical communication with the belt and waistband electrodes, the host unit including circuitry for generating the special ECG signals from one or more of the acquired electrical signals and circuitry for transmitting information based on the special ECG signals to a location remote from the wearable device. 15. The device of claim 14, wherein the total number of skin-contacting electrodes is four. 16. The system of claim 14, wh... | NewCardio, Inc. | A61B005/0402
US20110004124 | A1 | 2011-01-06 | IMPLANTABLE MEDICAL DEVICE INCLUDING MECHANICAL STRESS SENSOR | 2010-04-26 | Systems and techniques for determining in vivo mechanical load exerted on an implanted medical device (IMD) are described. The IMD or a device that has a substantially similar form factor as the IMD includes at least one mechanical stress sensor mechanically coupled to a housing or a component within the housing. In some examples, a patient parameter is determined based on a signal indicative of the in vivo mechanical load exerted on the IMD. In addition, in some examples, a processor determines whether a transient or cumulative mechanical load exerted on an IMD exceeded a predetermined threshold. A processor may additionally or alternatively determine whether a pattern in the mechanical loading of the IMD indicates a diversion from a manufacturing, shipping, storage or other handling process. | 1. An implantable medical system comprising: a housing;a therapy delivery module substantially enclosed within the housing, wherein the therapy delivery module is configured to deliver therapy to a patient;a component substantially enclosed within the housing; anda mechanical stress sensor mechanically coupled to at least one of the housing or the component, wherein the mechanical stress sensor generates a signal indicative of a mechanical load exerted on the housing or the component. 2. The system of claim 1, wherein the component comprises a circuit board and the mechanical stress sensor is at least one of mechanically attached to the circuit board or integrated into the circuit board. 3. The system of claim 1, wherein the component comprises at least one of a processor or a power source of the therapy delivery module. 4. The system of claim 1, wherein the mechanical stress sensor is mechanically coupled to at least one of an interior surface or an exterior surface of the housing, or integrated with the interior surface or the exterior surface of the housing. 5. The system of claim 1, further comprising: a telemetry module substantially enclosed within the housing; anda processor substantially enclosed within the housing, wherein the processor is configured to receive the signal from the mechanical stress sensor and transmit information indicative of the signal to an external device via the telemetry module. 6. The system of claim 1, further comprising a processor that is configured to receive the signal from the mechanical stress sensor and automatically determine whether the mechanical load exerted on the housing is greater than or equal to a predetermined threshold based on the signal generated by the mechanical stress sensor. 7. The system of claim 6, wherein the processor is configured to generate an indication if the mechanical load exerted on the housing is greater than or equal to the predetermined threshold. 8. The system of claim 1, further comprising a processor that is configured to receive the signal from the mechanical stress sensor and automatically determine at least one parameter of the patient based on the signal. 9. The system of claim 8, wherein the processor is configured to control the therapy delivery module based on the determined parameter of the patient. 10. The system of claim 1, further comprising a processor and a memory, wherein the processor is configured to determine a mechanical stress signature based on the signal generated by the mechanical stress sensor and store the mechanical stress signature in the memory. 11. The system of claim 10, wherein the processor is configured to compare the mechanical stress signature to a predetermined signature template and generate an indication based on the comparison. 12. The system of claim 1, wherein the mechanical stress sensor comprises at least one of a printed circuit comprising pressure sensitive ink, a piezoresistor, a piezoelectric crystal, a load cell, a... | Medtronic, Inc. | A61B005/103
US20110034827 | A1 | 2011-02-10 | METHOD AND A DEVICE FOR HEARING-RELATED EXAMINATION | 2007-04-17 | Method for hearing-related examination comprising the steps of generating sound signals and/or electrical signals in an examination unit (12). The method includes the steps of removably connecting the examination unit (12) as a modular unit to a base unit (11) and removably connecting a function card (13) to the base unit (11). Further steps are sending control signals from the function card (13) to the examination/module unit (12) for performing a specific examination, and obtaining test result data in the examination unit (12) and collecting said test result data to the function card for storage therein. The device comprises a base unit (11) having a display (15), input means (16; 17; 45) and output means (15) for performing the examination. The base unit (11) comprises a first interface (14) receiving a function card (13) and the function card (13) comprises a first control unit (33), a data memory (31) and a program memory (32) a, said program memory storing software for performing an examination and said data memory (31) storing data related to the examination. The base unit (11) comprises a second in-terface (28) receiving at least one modular unit (12) having a micro controller (29), examination output and input means, and the data memory (31) of the function card is connected to said micro controller (29) of the modular unit for collecting and storing examination data. | 1. Method for hearing-related examination comprising the steps of generating sound signals and/or electrical signals in an examination unit (12), characterised by the steps of removably connecting the examination unit (12) as a modular unit to a base unit (11),removably connecting a function card (13) to the base unit (11) sending control signals from the function card (13) to the examination/module unit (12) for performing a specific examination, andobtaining test result data in the examination unit (12) and collecting said test result data to the function card for storage therein. 2. Method in accordance with claim 1, further including the step of displaying examination related data on a display (15) of the base unit (11). 3. Method in accordance with claim 1, further including the step of receiving control input signals from input means (16; 17; 42; 43) of the base unit (11) and transferring the control input signals to the function card (13). 4. Method in accordance with claim 1, further including the step of connecting the function card (13) to the base unit (11) through a bus (24). 5. Method in accordance with claim 4, further including the step of sending the control signals from the function card (13) to the examination unit (12) through the bus (24). 6. Method in accordance with claim 1, further including the steps of counting in the function card (13) each time a specific examination is performed and limiting the number of examination performances to an adjustable number. 7. A device for hearing-related examination comprising a base unit (11) having a display (15), input means (16; 17; 45) and output means (15) for performing the examination, characterised in that the base unit (11)comprises a first interface (14) receiving a function card (13),that the function card (13) comprises a first control unit (33), a data memory (31) and a program memory (32) a, said program memory storing software for performing an examination and said data memory (31) storing data related to the examination,that the base unit (11) comprises a second interface (28) receiving at least one modular unit (12) having a micro controller (29), examination output and input means, andthat the data memory (31) of the function card is connected to said micro controller (29) of the modular unit for collecting and storing examination data. 8. A device in accordance with claim 7, wherein said first interface (14) comprises a bus (24) and a plurality of control lines (34) allowing said control unit to interact with said at least one modular unit. 9. A device in accordance with claim 8, wherein the bus (24) is extended through the second interface (28). 10. A device in accordance with claim 7, wherein the base unit (11) is provided means (18; 19; 20; 21; 22; 23) for communicating with external units. 11. A device in accordance with claim 10, wherein said means are connected to a digital switch (25). | Entomed AB | A61B005/12
US08157752 | B2 | 2012-04-17 | Posture assessment and feedback instrument | 2009-08-18 | A posture assessment and feedback instrument (FIG. 1A and FIG. 1B) all contained within a single housing (1). The output of a capacitive sensor (4) within the housing (1) varies in response to the distance between the sensor and the body of a user. The housing (1) is strapped to the body of a user (FIG. 3, FIG. 4, FIG. 5A, and FIG. 5B) with a belt (2). The application area is chosen such that the distance between the sensor (4) and the body varies in response to a defined posture change (FIG. 5A and FIG. 5B). Thus the sensor output will vary in response to the defined posture change. The output is compared to a stored threshold value and triggers a feedback signal when the threshold is exceeded. | 1. A method for posture assessment comprising the steps of: choosing a sensing means that provides an output in response to a distance between said sensing means and the body of a user,choosing a suitable area of said body having properties causing said distance to vary in response to a defined posture change, andapplying said sensing means to said suitable area causing said distance and thus said output to vary in response to said defined posture change,whereby no separate signal source and no moveable hardware components are required, thus allowing said sensing means to be durable, cleanable, and usable, andwhereby said user has the user-friendly choice of applying said sensing means either directly to the skin or to the clothing. 2. The method of claim 1 wherein said sensing means is capacitive. 3. The method of claim 1 wherein said output is compared to a stored threshold value and triggers a feedback signal when said threshold is exceeded. 4. The method of claim 1 wherein said defined posture change is spinal. 5. The method of claim 1 wherein said defined posture change is a change in spinal curvature. 6. The method of claim 1 wherein said defined posture change is slumping. 7. The method of claim 1 wherein said sensing means and all other electronic components required in the process of posture assessment are contained within a single housing. 8. A method for posture assessment comprising the steps of:choosing a sensing means whose signal output changes in response to changes of a contact area between said sensing means and the body of a user wearing said sensing means, choosing a suitable location on said body having properties causing said contact area to vary in response to a defined posture change, andapplying said sensing means to said suitable location, causing said contact area and thus said output signal to vary in response to said defined posture change,whereby no separate signal source and no moveable hardware components are required, thus allowing said sensing means to be durable, cleanable, and usable, andwhereby said user has the user-friendly choice of applying said sensing means either directly to the skin or to the clothing. 9. The method of claim 8 wherein said sensing means is pressure-sensitive. 10. The method of claim 8 wherein said output is compared to a stored threshold value and triggers a feedback signal when said threshold is exceeded. 11. The method of claim 8 wherein said defined posture change is a change in spinal curvature. 12. The method of claim 8 wherein said defined posture change is slumping. 13. The method of claim 8 wherein said sensing means and all other electronic components required in the process of posture assessment are contained within a single housing. 14. A method for posture feedback comprising the steps of: choosing a sensing means that provides an output in direct response to a distance between said sensing means and the body of a user,choosing a suitable area of said body havi... |  | A61B005/117 A61B005/103
US20110014954 | A1 | 2011-01-20 | Emergency Medical Diagnosis and Communications Device | 2010-01-07 | A portable emergency medical device capable of communicating with a remote location preferably as a cellular telephone that can measure one or more human vital parameters such as pulse rate, body temperature, skin moisture, blood pressure, ECG or blood chemistry and can receive symptoms from a user either by voice recognition or by keypad and can provide an expert medical diagnosis. The device can store a complete medical history for one or more users and can use an expert system to make the diagnosis. The device can make an emergency medical call either on command or automatically requesting help and optionally supplying medical information and/or GPS location information. | 1. A cellular telephone comprising: a housing;a radio antenna attached to said housing;a microphone and earphone or speaker mounted on said housing;a GPS receiver in said housing;a cellular telephone electrical module contained in said housing in electrical communication with said radio antenna, said GPS receiver and said microphone and earphone or speaker, said cellular telephone electrical module adapted to communicate over a cellular telephone network by placing voice calls or data calls;a processor in said housing;a display on said housing;a keypad on said housing;a medical expert system run-able on said processor, said medical expert system adapted to read keyed-in data from said keypad and write graphical or text data upon said display;a blood pressure sensor in or deployable from said housing;a pulse rate sensor in or deployable from said housing;a body temperature sensor in or deployable from said housing;an ECG interface in or deployable from said housing;a memory in said housing containing one or more particular individuals' medical history;a blood analysis module in said housing, said blood analysis module in fluid communication with a blood drawing module;said processor adapted to control said blood pressure sensor, said pulse rate sensor, said body temperature sensor, said ECG interface, said blood analysis module, said memory and said blood drawing module;wherein said expert system accepts notice from a user of a possible emergency medical condition, and if said user is an individual with a medical history stored in said memory, accesses said user's medical history;and wherein, said processor activates at least one of: the blood pressure sensor, the pulse rate sensor, the body temperature sensor or the ECG interface, the blood drawing module in order to measure a human medical parameter, said processor reporting said human medical parameter to said expert system;and wherein said expert system communicates a diagnosis to said user via at least said display along with particular instructions;and wherein said processor causes said cellular telephone module to place a cellular telephone call if said user fails to respond to said particular instructions, said cellular telephone module also reporting location data of said cellular telephone obtained by said processor from said GPS. 2. The cellular telephone of claim 1 further comprising a voice synthesis module. 3. The cellular telephone of claim 1 further comprising a voice recognition module. 4. The cellular telephone of claim 1 wherein said blood drawing module contains a lance and wash station. 5. The cellular telephone of claim 1 wherein said blood analysis module contains a plurality of capillary reaction chambers. 6. The cellular telephone of claim 1 further comprising a finger insert that includes pulse rate sensor and said body temperature sensor. 7. The cellular telephone of claim 1 further comprising an ECG analysis program. 8. A cellular telephone comprising: a housing;... | - | H04B001/38 A61B005/00
US07862520 | B2 | 2011-01-04 | Body fluid sampling module with a continuous compression tissue interface surface | 2008-06-20 | A sampling module includes a housing and a penetrating member positioned in the housing. An attachment device is coupled to the housing. The attachment device is tiltably attached to a tissue interface that permits the tissue interface surface to tilt. The tissue interface includes an opening for the penetrating member to pass there through. | 1. A sampling module, comprising: a housing;a penetrating member positioned in the housing;an attachment device coupled to the housing, the attachment device being tiltably attached to a tissue interface that permits the tissue interface surface to tilt forward or backward, the tissue interface including an opening for the penetrating member to pass there through;a driver that includes at least: a human interface that communicates at least one output to a user and an input from the user to actuate the penetrating member;a first processor that is fast and high power to drive the penetrating member:a second processor that is slow and low power to process the output to and input from the user and to control the power to the first processor. 2. The module of claim 1, wherein the tissue interface has a contoured geometric configuration. 3. The module of claim 1, wherein the tissue interface has a geometric configuration that matches a geometry of a finger tip. 4. The module of claim 3, wherein the tissue interface applies pressure to the finger tip when urged against the finger tip. 5. The module of claim 1, wherein the tissue interface has an ergonomically contoured pressure surface configured to conform to a finger tip. 6. The module of claim 1, further comprising: a sample reservoir positioned in the housing and adapted to receive a fluid sample from a wound created by the penetrating member as it penetrates a tissue site positioned at the tissue interface surface. 7. The module of claim 1, wherein the tissue interface includes a tissue interface member and a retainer and the attachment device attaches the retainer to the housing. 8. The module of claim 1, wherein the attachment device includes a compliant member. 9. The module of claim 1, wherein the attachment device is configured to provide that the tissue interface is free to tilt to a within a predetermined angle range relative to the housing. 10. A sampling module, comprising: a housing with an input port and a continuous compression surface that applies substantially uniform pressure against a dermal tissue target site;a penetrating member positioned in the housing;a sample reservoir coupled to an input port;a driver that includes at least: a human interface that communicates at least one output to a user and an input from the user to actuate the penetrating member;a first processor that is fast and high power to drive the penetrating member;a second processor that is slow and low power to process the output to and input from the user and to control the power to the first processor. 11. The module of claim 10, wherein the continuous compression surface is an ergonomically contoured surface. 12. The module of claim 11, wherein the ergonomically contoured surface has a geometry that matches that of a finger tip. 13. The module of claim 10, wherein the continuous pressure surface has an elliptical continuous contoured compression surface. | Pelikan Technologies, Inc. | A61B005/00 A61B017/14 A61B017/32
US07865224 | B2 | 2011-01-04 | Method and apparatus for estimating a physiological parameter | 2004-10-12 | In a physiological monitor, a method and an apparatus for determining a patient's pulse rate using data corresponding to a plurality of wavelengths of electromagnetic energy transmitted through the tissue of the patient. The method includes tracking the pulse rate in the data using an adaptive comb filter, the data having signal portions corresponding to the pulse rate and signal portions corresponding to noise, periodically calculating a frequency power spectrum of one of the wavelengths, and using the frequency power spectrum in a pulse rate calculator to determine the pulse rate or to verify the pulse rate calculated by the pulse rate calculator. | 1. In a physiological monitor, a method for determining a patient's pulse rate using data corresponding to a plurality of wavelengths of electromagnetic energy transmitted through the tissue of the patient, the method comprising: tracking the pulse rate in the data using an adaptive comb filter, said data having signal portions corresponding to the pulse rate and signal portions corresponding to noise;periodically calculating a frequency power spectrum of one of said wavelengths; andusing the frequency power spectrum in a pulse rate calculator to determine the pulse rate or to verify the pulse rate calculated by the pulse rate calculator. 2. The method of claim 1 further comprising preprocessing the data by filtering the data to filter out frequencies that do not reflect cardiac blood flow. 3. The method of claim 1 further comprising preprocessing the data by taking the natural logarithm of the data and band pass filtering the data to attenuate the data below approximately 0.5 Hz and above approximately 10 Hz. 4. The method of claim 1 wherein said tracking is configured to track slowly varying heart rate. 5. The method of claim 1 wherein said calculating a frequency power spectrum comprises using a Fourier transform. 6. The method of claim 1 wherein said using a frequency power spectrum comprises determining which of a plurality of peaks in the frequency power spectrum corresponds to the fundamental frequency. 7. The method of claim 1 wherein said tracking comprises defining a comb filter to remove signal energy from the data corresponding to the fundamental frequency and harmonics thereof; and determining a particular harmonic frequency which minimizes noise energy at an output of the comb filter, the particular frequency corresponding with the fundamental frequency. 8. The method of claim 1 further comprising calculating the autocorrelation function of said wavelengths; identifying peaks in the power spectrum and peaks in the autocorrelation function; anddetermining the pulse rate as the frequency corresponding with a peak present in both the frequency power spectrum and the autocorrelation function. 9. The method of claim 1 further comprising using a pattern matching rate calculator to calculate an average pulse rate period over a set number of data samples,wherein the pattern matching rate calculator recognizes sequences of crests and troughs in the data and is configured to reject pulses that are potentially contaminated by motion artifact. 10. The method of claim 1 further comprising using the pulse rate determined using the frequency power spectrum as triggers for a cardiac gated averaging processor configured for averaging the plurality of wavelengths of electromagnetic energy transmitted through the tissue of the patient to generate a plurality of filtered waveforms to render them substantially free of noise. 11. The method of claim 10 further comprising determining oxygen saturation of hemoglobin in arterial blood in the tissu... | Nellcor Puritan Bennett LLC | A61B005/1455 A61B005/02
US07867165 | B2 | 2011-01-11 | System and method for monitoring a physiological condition | 2006-10-19 | Method for treating a medical condition in a human patient comprising choosing a psychological strategy for treating the medical condition, encoding electronic instructions for an interactive video game in such a way that the interactive video game implements the psychological strategy, loading the electronic instructions into a microprocessor-based unit equipped with a display for displaying the interactive video game and with an patient input device for receiving responses to the interactive video game from the human patient, and instructing the human patient how and when to use the microprocessor-based unit to play the interactive video game. The interactive video game contains instructions for a scoring procedure for quantitatively analyzing the medical condition of the human patient, and/or counseling instructions or self-care instructions. The video game can be used in conjunction with a physical parameter measuring device connected to the microprocessor-based unit. | 1. A diabetes self-care system, comprising: (a) a blood glucose meter;(b) a portable microprocessor-based unit signal coupled with the blood glucose meter such as to be capable of advising a user to measure a blood glucose level by downloading blood glucose measurement data into the portable microprocessor-based unit, wherein said portable microprocessor based unit comprises a display screen configured to present to said user an interactive presentation;(c) said blood glucose measurement data being used as input data for a program of instructions running on the portable microprocessor-based unit; and(d) said program of instructions including instructions (i) to advise said user to inject insulin according to an insulin plan, (ii) present to said user on said display screen (a) said blood glucose measurement data and (b) information to educate said user regarding management of diabetes, (iii) receive from said user responses confirming said blood glucose level has been measured, and (iv) positively reinforce compliance with a diabetes management plan by providing incentives in said interactive presentation on said display screen in response to said blood glucose level being within an optimal range. 2. The system of claim 1, wherein said program of instructions comprises (i) said diabetes management plan, and (ii) instructions to send a signal to said portable microprocessor-based unit to advance between steps of said diabetes management plan. 3. The system of claim 1, wherein said program of instructions comprises instructions to display an indicator to check blood glucose level on said display screen of the portable microprocessor-based unit, wherein said display screen comprises a multi-line display configured to (A) present to said user alphanumeric messages of at least one line of text, and (B) visual images. 4. The system of claim 1, wherein said program of instructions comprises instructions to display an indicator to select insulin plan on a display screen of the portable microprocessor-based unit, wherein said display screen comprises a multi-line display configured to (A) present to said user alphanumeric messages of at least one line of text, and (B) visual images. 5. The system of claim 1, wherein said program of instructions comprises instructions to display an indicator to get a menu of foods to eat on said display screen of the portable microprocessor-based unit, wherein said display screen comprises a multi-line display configured to (A) present to said user alphanumeric messages of at least one line of text, and (B) visual images. 6. The system of claim 1, said portable microprocessor-based unit being connectable with a communication unit remotely located from said portable microprocessor-based unit. 7. The system of claim 1, said portable microprocessor-based unit being connectable with a computer in a hospital and remotely located from said portable microprocessor-based unit. 8. The system of claim 7, said portable micropr... | Health Hero Network, Inc. | A61B005/00
US07877132 | B2 | 2011-01-25 | Medical viewing system and method for detecting and enhancing static structures in noisy images using motion of the image acquisition means | 2006-04-13 | A method for use in a medical viewing system for processing for display a sequence of images of a medical intervention comprising manipulating a tool such as a stent (25) within an artery. Two markers (21, 22) are provided on the tool support (10) and the method includes extracting these markers (21, 22) to yield marker location information, from which tool location information can be derived. The image of the tool within the sequence of images is then enhanced relative to the background. The background can be additionally blurred by simulating or effecting relative movement of the image detector and the stent (25) by, for example, zooming in on the stent (25) or rotating the detector relative thereto, during acquisition of the sequence of images. | 1. A method of medical imaging comprising: at least one of moving and positioning a tool in a body organ using a tool support;acquiring a sequence of 2D fluoroscopic images of said tool with an image detector;during the acquiring step, moving said image detector relative to said tool such that the fluoroscopic images of the sequence are acquired at least one of from different directions and with different levels of zoom;automatically extracting markers attached to said tool support and yielding marker location information;at least one of rotating, translating, and scaling the fluoroscopic images of the sequence to bring the markers into registration in the fluoroscopic images of the sequence, background portions of the registered images being different in different images due to being acquired at least one of from different directions and with different levels of zoom;deriving tool location information from said marker location information in the registered images of the sequence; andintegrating the registered images of said sequence to coherently sum portions of the image including the tool to improve the visibility of said tool and sum the background portions of the images to blur the background such that the visibility of the tool is enhanced relative to the background. 2. The method according to claim 1, wherein said movement of the detector relative to the tool is simulated by zooming in on the tool. 3. The method according to claim 1, wherein said movement of the detector relative to the tool is effected by rotating the detector relative to the tool. 4. The method according to claim 3, wherein said detector is rotated substantially about a tool axis. 5. The method according to claim 3, further including: temporal recursive filtering during the integrating step to produce an image of a rotating enhanced tool. 6. The method according to claim 3, wherein the integrating step includes symmetrical temporal integration of images of the sequence that are acquired 180 degrees offset from each other. 7. The method according to claim 6, further including: adapting a window for the symmetrical temporal integration to a rotation speed of the detector. 8. The method according to claim 1, wherein the tool is stationary relative to the body organ during acquisition of the sequence of images. 9. The method according to claim 1, further comprising: receiving system geometry data indicative of detector movement relative to the tool. 10. A medical imaging system comprising: a fluoroscopic image detector which acquires a sequence of fluoroscopic images of a surgical tool which is instrumented with markers that are visible in the fluoroscopic images during a medical intervention, the detector moving relative to said tool during acquisition of said sequence of fluoroscopic images;a processor programmed to: receive the sequence of fluoroscopic images from the detector;identify the markers in each fluoroscopic image;at least one of translate, rotate, and ... | Koninklijke Philips Electronics N.V. | A61B005/05
US08909323 | B2 | 2014-12-09 | System for processing angiography and ultrasound image data | 2010-07-15 | A system provides a single composite image including multiple medical images of a portion of patient anatomy acquired using corresponding multiple different types of imaging device. A display processor generates data representing a single composite display image including, a first image area showing a first image of a portion of patient anatomy acquired using a first type of imaging device, a second image area showing a second image of the portion of patient anatomy acquired using a second type of imaging device different to the first type. The first and second image areas include first and second markers respectively. The first and second markers identify an estimated location of the same corresponding anatomical position in the portion of patient anatomy in the first and second images respectively. A user interface enables a user to move at least one of, (a) the first marker in the first image and (b) the second marker in the second image, to correct the estimated location so the first and second markers identify the same corresponding anatomical position in the portion of patient anatomy. | 1. A method operable on one or more processors for providing a single composite image including a plurality of medical images of a portion of patient anatomy acquired using a corresponding plurality of different types of imaging device, comprising: displaying a single composite image including, a first image area showing a first image of an internal portion of patient anatomy acquired using a first type of imaging device anda second image area showing a second image of said internal portion of patient anatomy acquired using a second type of imaging device different than said first device type, said second image comprising a plurality of rows acquired at corresponding plurality of successive acquisition times and at different anatomical locations, said first and second image areas including first and second markers respectively, said first and second markers identifying an estimated location of the same corresponding anatomical position in said portion of patient anatomy in said first and second images respectively; andmoving at least one of, (a) the first marker in the first image and(b) the second marker in the second image, to correct said estimated location so said first and second markers identify the same corresponding anatomical position in said portion of patient anatomy and said second marker identifies a row in said second image having a known acquisition time. 2. A method according to claim 1, wherein said first type of imaging device is an X-ray imaging device,said second type of imaging device is an intravascular ultrasound imaging device,said portion of patient anatomy comprises a vessel structure. 3. A method according to claim 2, further comprising respectively acquiring, by said X-ray imaging device and said ultrasound imaging device concurrently, first and second sequential series of images at time stamped acquisition times and said estimated location is derived using the timestamps of said first and second sequential series of images. 4. A method according to claim 1, wherein the first type of imaging device is at least one of (a) an X-ray imaging device, (b) an MR imaging device, (c) a CT scan imaging device and (d) a nuclear imaging PET scanner device. 5. A method operable on one or more processors for providing a single composite image including a plurality of medical images of a portion of patient anatomy acquired using a corresponding plurality of different types of imaging device, comprising: storing in at least one repository a first and second sequential series of images concurrently acquired at time stamped acquisition times using an X-ray imaging device and an intravascular ultrasound imaging device respectively;displaying a single composite image including, a first image area showing a first image of a patient internal vessel structure acquired using a first type of imaging device anda second image area showing a second image of said patient internal vessel structure acquired using a second device type of imaging ... | Siemens Medical Solutions USA, Inc. | A61B008/14 A61B006/00 A61B008/12 A61B008/00 A61B008/08 A61B005/00
US08718736 | B2 | 2014-05-06 | Physiological sensor with offset adhesive layer | 2010-07-23 | An exemplary sensor includes a sensor pad defining a perimeter, a light source, a light detector, and an adhesive layer. The light source is configured to generate near-infrared light and transmit the near-infrared light through part of a patient's body. The light detector is configured to receive the near-infrared light generated by the light source after it has traveled through part of the patient's body. The light received by the light detector indicates an amount of oxygen in the part of the patient's body through which the near-infrared light traveled. The adhesive layer is offset relative to the sensor pad to, for example, allow a clinician to easily remove the sensor from the patient. | 1. A sensor comprising: a sensor pad defining a perimeter and having a patient surface configured to be applied adjacent to a patient's skin, said patient surface being substantially free of adhesive;a light source disposed on the sensor pad and configured to generate near-infrared light and transmit the near-infrared light through part of a patient's body;a light detector disposed on the sensor pad and configured to receive near-infrared light generated by the light source after it has traveled through part of the patient's body, wherein the light received by the light detector indicates an amount of oxygen in the part of the patient's body through which the near-infrared light traveled;a cable extending from the perimeter of the sensor pad; andan adhesive layer disposed on the sensor pad, wherein one part of the adhesive layer extends beyond the perimeter of the sensor pad and another part of the adhesive layer at a location spaced from the cable does not extend beyond the perimeter of the sensor pad. 2. A sensor as set forth in claim 1, wherein part of the adhesive layer extends beyond the perimeter of the sensor pad and part of the sensor pad extends beyond an edge of the adhesive layer. 3. A sensor as set forth in claim 1, wherein the sensor pad includes a flexible material. 4. A sensor as set forth in claim 1, wherein the adhesive layer includes a flexible material. 5. A sensor as set forth in claim 1, wherein the adhesive layer includes a pressure sensitive adhesive. 6. A sensor as set forth in claim 1, wherein the adhesive layer includes a hydrocolloid adhesive. 7. A sensor as set forth in claim 1, wherein the light source includes at least one of a light emitting diode and a laser diode. 8. A sensor as set forth in claim 1, wherein the light detector includes a photodiode. 9. A sensor as set forth in claim 1, further comprising a liner disposed on the adhesive layer. 10. A sensor as set forth in claim 1, wherein the sensor pad includes a top layer and a bottom layer, and wherein the light source and the light detector are disposed between the top and bottom layers. 11. A sensor as set forth in claim 10, wherein the top and bottom layers each include a flexible material. 12. A sensor as set forth in claim 10, wherein the adhesive layer is disposed on the bottom layer of the sensor pad and wherein the adhesive layer is offset relative to the bottom layer. 13. A sensor as set forth in claim 12, wherein the perimeter of the sensor pad is at least partially defined by the bottom layer and wherein one part of the adhesive layer extends beyond the perimeter defined by the bottom layer and wherein another part of the adhesive layer is substantially flush with the perimeter defined by the bottom layer. 14. A sensor as set forth in claim 12, wherein part of the adhesive layer extends beyond the perimeter defined by the bottom layer and part bottom layer extends beyond an edge of the adhesive layer. 15. A sensor as set forth in claim ... | Covidien LP | A61B005/00
US08862222 | B2 | 2014-10-14 | Non-invasive measuring of load-induced electric potentials in diarthroidial joints | 2009-03-02 | There is described a method for non-invasively measuring electrical activity in a joint of a subject, the method comprising: removably attaching in a non-invasive manner at least two electrodes to a skin surface around an articulation comprising the joint; generating electroarthrographic potentials within the joint by loading the articulation; capturing the electroarthrographic potentials using the at least two electrodes; discriminating between electroarthrographic potentials originating from joint tissue activity and those from other sources; and generating measurement signals representing the electrical activity of joint tissue. | 1. A method for non-invasively measuring electrical activity in a live joint of a subject, the method comprising: removably attaching in a non-invasive manner at least two electrodes to a skin surface around an articulation comprising the live joint;generating electroarthrographic potentials within the live joint by loading the articulation;capturing said electroarthrographic potentials using said at least two electrodes;discriminating between electroarthrographic potentials originating from joint tissue activity and electrical potentials generated by other sources within the subject; andgenerating measurement signals representing the electrical activity of joint tissue. 2. The method of claim 1, wherein said generating electroarthrographic potentials comprises one of loading the articulation manually, loading the articulation with an electromechanical loading device, and loading the articulation with a piezoelectric loading device. 3. The method of claim 1, further comprising displaying said signals generated in the form of an electroarthrogram. 4. The method of claim 1, further comprising placing at least one of goniometers, accelerometers and sensors adjacent to the live joint to measure at least one of an angle of the live joint and forces exerted by the subject during loading. 5. The method of claim 1, wherein said discriminating between electroarthrographic potentials originating from joint tissue activity and electrical potentials generated by other sources within the subject comprises subtracting at each instant in time a weighted average of electrical potentials originating from the electrodes so as to measure the electroarthrographic potentials with respect to an absolute reference. 6. The method of claim 1, wherein said discriminating between electroarthrographic potentials originating from joint tissue activity and electrical potentials generated by other sources within the subject comprises band pass filtering. 7. The method of claim 1, wherein said discriminating between electroarthrographic potentials originating from joint tissue activity and electrical potentials generated by other sources within the subject comprises low pass filtering. 8. The method of claim 1, wherein discriminating between electroarthrographic potentials originating from joint tissue activity and electrical potentials generated by other sources within the subject comprises high pass filtering. 9. The method of claim 1, further comprising displaying said signals generated in the form of sequences, movies, isopotential maps or 3D potential surfaces. 10. The method of claim 1, further comprising calculating electrical sources within the live joint and representing said electrical sources as electric fields, currents, polarization or charges. 11. The method of claim 1, wherein said discriminating comprises discriminating between electrical potentials originating from cartilage and electrical potentials generated by other sources within the subject. 12.... | Polyvalor, Limited Partnership | A61B005/05 A61B005/04 A61B005/103 A61B005/117 A61B005/00 A61N001/04
US07885439 | B2 | 2011-02-08 | Apparatus and method for creating identification information based on unique components of a body | 2005-02-09 | An identification information creation apparatus and method capable of improving reliability in an identification process. A prescribed number or fewer isolated pixels are eliminated as condition noises from a binarized image created from images obtained by imaging blood vessels unique to a body, and identification information is created based on the blood vessels of the resultant noise-free binarized image. This can reduce a possibility of creating the identification information with the condition noises as a part of the blood vessels, thus making it possible to prevent the identification information from deteriorating due to different imaging conditions and physiologic change of imaging target and thus improving reliability in the identification process. | 1. An information creation apparatus for creating information on a body, comprising: a solid imaging element configured to sequentially capture a plurality of blood vessel images of a same image type based on near-infrared light received from the body;a binarization unit configured to create a plurality of binarized images from the plurality of blood vessel images sequentially received from the solid imaging element;a composition unit configured to sequentially perform an averaging process on the plurality of binarized images to create a composed binarized image;an elimination unit configured to eliminate isolated pixels from the composed binarized image received from the composition unit; anda creation unit configured to create identification information based on the blood vessels of the composed binarized image subjected to the elimination of the elimination unit, whereinthe elimination unit is configured to eliminate a prescribed number or fewer isolated pixels, to determine whether a number of the isolated pixels eliminated is a prescribed threshold value or lower, and when the number of the isolated pixels eliminated is determined to be the prescribed threshold value or lower, to identify the composed binarized image subjected to the elimination as being usable for creating the information on the body; andthe creation unit is configured to create the identification information based on the blood vessels of the composed binarized image identified by the elimination unit. 2. The information creation apparatus according to claim 1, wherein the elimination unit is configured to sequentially extract a search image from the composed binarized image and to eliminate the isolated pixels existing inside outer pixels of the search image extracted. 3. The information creation apparatus according to claim 1, wherein the creation unit is configured to create the identification information based on the blood vessels of the composed binarized image subjected to the elimination of the elimination unit only when the number of the isolated pixels eliminated is determined to be the prescribed threshold value or lower. 4. An information creation method for creating information on a body, comprising: capturing, by a solid imaging element, a plurality of blood vessel images of a same image type based on near-infrared light received from the body;creating a plurality of binarized images from the plurality of blood vessel images sequentially received from the solid imaging element;sequentially performing an averaging process on the plurality of binarized images to create a composed binarized image;eliminating isolated pixels from the composed binarized image created; andcreating identification information based on the blood vessels of the composed binarized image subjected to the eliminating step, whereinin the eliminating step, a prescribed number or fewer isolated pixels are eliminated, whether a number of the isolated pixels eliminated is a prescribed thresho... | Sony Corporation | G06K009/00 A61B005/05
US09612140 | B2 | 2017-04-04 | Support device for sensors and/or actuators that can be part of a wireless network of sensors/actuators | 2009-04-16 | A support device for sensors and/or actuators, of a node element or End Device (10) of a wireless network of sensors/actuators which support device comprises a base body or base module (1) with a plurality of faces (2), to which a sensor and/or actuator module or block (3) is mounted, and an electrical supply block/module (4). The main feature of the base body (1) is the possibility of adding to the node element (10) further sensors and/or actuators (3) as well as adapter elements, bridge interface elements (60) or expansion modules (50). The base body (1) can have a plurality of connection faces (2) with positive engagement means (54), which allow to assemble the above described modules. The positive engagement means (54) ensures an electromechanical connection between the many modules and, furthermore, assists an easy and intuitive accessibility for a user. | 1. A physiological monitoring device comprising: a first base configured to be removably coupleable to a patient, the first base including a first face having a first port,a second face having a second port,a third face having a third port, the first port and the second port having an identical configuration,a controller configured to communicate with a wireless network, anda storage unit configured to communicate with the controller; anda first physiological monitoring module removably coupleable to the first port or the second port, the first physiological monitoring module including an ECG sensor configured to acquire ECG signals from the patient,a first interface compatible with the first port and the second port, anda first identification device on the interface,the first physiological monitoring module configured to communicate with the controller to wirelessly communicate the ECG signals acquired from the patient to the wireless network,the controller configured to identify whether the first physiological monitoring module is in the first port or the second port based on the first identification device,the controller further configured to wirelessly communicate with the wireless network that the first physiological monitoring module is connected to the first base and the type of data to be transmitted;wherein the first identification device comprises at least a pin coupleable to a corresponding pin on either the first face or the second face to exchange electrical communication, wherein the pin on either the first face or the second face receives a signal from the first physiological module, the signal identifying the first physiological module. 2. The physiological monitoring device of claim 1, further comprising a second physiological monitoring module separately removably coupleable to the first port or the second port, the second physiological monitoring module including a second sensor configured to acquire non-ECG physiological signals from the patient, a second interface compatible with the first port and the second port, and a second identification device on the second interface, the second physiological monitoring module configured to communicate with the controller to wirelessly communicate the non-ECG physiological signals acquired from the patient to the wireless network, the controller configured to identify whether the second physiological monitoring module is in the first port or the second port based on the second identification device, the controller further configured to wirelessly communicate with the wireless network that the second physiological monitoring module is connected to the first base and the type of data to be transmitted, wherein the second identification device comprises at least a pin coupleable to a corresponding pin on either the first face or the second face to exchange electrical communication, wherein the pin on either the first face or the second face receives a signal from the second physiological ... | WINMEDICAL S.R.L. | A61B005/02 G01D011/30 A61B005/00
US20110009713 | A1 | 2011-01-13 | Rhythmic percussion exercise garment with electronic interface and method of conducting an exercise program | 2010-08-20 | Hands-free exercise device combining: weights; an electronic vital signs monitoring device with interactive capabilities and a controllable interface with an electronic or instrumental percussive music source, adapted to be worn about the user's body resulting in new method of conducting an exercise program. User varies weight to adjust muscle load during a range of motion. Simultaneously, sound is movement activated: depending on placement, user learns to exercise de-conditioned muscle groups by learning how to produce sound. During exercise, device monitors/records user's vital signs to objectively determine exertion level in real time, and calculate long term progress. Device can automatically adjust tempo, challenging user to respond with concordant change in exercise pace so that actual exertion level will increase or decrease to desired levels. Device can respond to exertion, notifying user when a threshold of exertion is reached by switching the type of instrumental sound, providing indicative audible cue. | 1. A weighted, musical, exercise device comprised of: at least one user attachment means;at least one compartment means attached to the user attachment means;at least one sound producing means containable within the compartment means;at least one vital signs detection means capable of detecting at least one the user's vital signs selected from the group consisting of a user's: pulse, resting heart rate, blood pressure, saturated blood oxygen level, body temperature, respiratory rate, or data necessary to calculate the user's VO2 max, percentage of maximal heart rate or other user selected measure of physical fitness;at least one controller and processor means in electronic communication with the sound producing means and the vital signs detection means;at least one music source means in electronic communication with the controller and processor means; andat least one weight means of a known approximate weight containable within the compartment means. 2. A weighted, musical, exercise device comprised of: at least one user attachment means;at least one compartment means attached to the user attachment means;at least one sound producing means containable within the compartment means, wherein the sound producing means is movement activated;at least one vital signs detection means capable of detecting at least one the user's vital signs selected from the group consisting of user's: pulse, resting heart rate, blood pressure, saturated blood oxygen level, body temperature, respiratory rate, or data necessary to calculate the user's VO2 max, percentage of maximal heart rate or other user selected measure of physical fitness;at least one controller and processor means in electronic communication with the sound producing means and the vital signs detection and vital signs recording means;at least one vital signs recording means in electronic communication with the controller and processor means capable of recording over time at least one of the user's vital signs selected from the group consisting of a user's: pulse, resting heart rate blood pressure, saturated blood oxygen level, body temperature, respiratory rate, or data necessary to calculate changes over time in the user's VO2 max, percentage of maximal heart rate or other user selected measure of physical fitness;at least one digital display means in electronic communication with the controller and processor means;at least one music source means in electronic communication with the controller and processor means; andat least one weight means of a known approximate weight contained within the compartment means. 3. The weighted, musical, exercise device of claim 1 wherein the user attachment means and the compartment means are comprised of: elastomer fabric; at least one attached adjustment strap; andat least one securing means attached to the adjustment strap. 4. The weighted, musical, exercise device of claim 1 wherein the user attachment means is at least one article selected from the group consi... | - | A61B005/00 A63B071/00
US20110038805 | A1 | 2011-02-17 | COMPOUNDS COMPRISING PARAMAGNETIC CHELATES ARRANGED AROUND A CENTRAL CORE AND THEIR USE IN MAGNETO RESONANCE IMAGING AND SPECTROSCOPY | 2009-04-17 | The present invention relates to novel compounds of formula (I) and (II), compositions comprising compounds of formula (II) and their use as contrast agents in magnetic resonance (MR) imaging (MRI) and MR spectroscopy (MRS). | 1. Compound of formula (II) A-(B-L-R-(L-R-(L-R-(L-R-(L′-X′)r)r)r)r)n (II)whereinA denotes a core;B is the same or different and denotes a moiety that constitutes an obstacle for the rotation of the covalent bond between B and L.L is the same or different and denotes a rigid linker moiety, wherein at least one L is present and under the proviso that L never links directly to another L;L′ is present or not and if present is the same or different and denotes a linker moiety;R is the same or different and denotes a branching moiety that reproduces with an individual multiplicity of r, wherein at least one R is present;X′ is the same or different and denotes a paramagnetic chelate consisting of a chelator X and a paramagnetic metal ion M; andr is the same or different and denotes the integer 2, 3 or 4; andn denotes an integer of 3 to 6. 2. Compound according to claim 1 wherein L is selected from the group consisting ofwhereinQ stands for H, C1-C8-alkyl, optionally substituted with one or more hydroxyl or amino groups; and* stands for the possible attachment points to B and R. 3. Compound according to claim 1 wherein A is a cyclic core, a carbon atom or an ethyl group. 4. Compound according to claim 1 wherein A is a saturated or non-saturated, aromatic or aliphatic ring comprising at least 3 carbon atoms and optionally one or more heteroatoms N, S or O, said ring being optionally substituted with one or more of the following substituents: C1-C3-alkyl, optionally substituted with hydroxyl or amino groups, amino or hydroxyl groups or halogen, provided that there are n attachment points left for groups B. 5. Compound according to claim 1 wherein B is a moiety whose rotation is hindered by sterical interaction with L. 6. Compound according to claim 1 wherein B comprises a benzene residue. 7. Compound according to claim 1 wherein the B moieties are interconnected by covalent bonds. 8. Compound according to claim 1 wherein R is a moiety whose rotation is hindered by sterical interaction with L and or L′. 9. Compound according to claim 1 wherein R is selected from the group consisting ofwherein Q is H or CH3. 10. Compound according to claim 1 wherein X is selected from residues of DOTA, DTPA, BOPTA, DO3A, HPDO3A, MCTA, DOTMA, DTPA BMA, M4DOTA, M4DO3A, PCTA, TETA, TRITA, HETA, DPDP, EDTA or EDTP. 11. Compound according to claim 1 wherein M is a paramagnetic ion of a transition metal or a lanthanide metal. 12. Compound according to claim 1 wherein all B are the same and/or all L are the same and/or all R are the same and/or all L′ are the same and/or all X′ are the same. 13. Composition comprising the compound according to claim 1 and at least one physiologically tolerable carrier. 14. Composition according to claim 13 for use as MR imaging agent or MR spectroscopy agent. 15. (canceled) 16. Method of MR imaging and/or MR spectroscopy wherein the composition according to claim 13 is administered to a subject and the subject is subjected to an MR... | - | A61B005/055 C07D257/02 A61K031/555
US20110046451 | A1 | 2011-02-24 | METHOD AND AUTOMATED SYSTEM FOR ASSISTING IN THE PROGNOSIS OF ALZHEIMER'S DISEASE, AND METHOD FOR TRAINING SUCH A SYSTEM | 2009-05-14 | A method and automated system for assisting in the prognosis of the progress and for assisting in the diagnosis of Alzheimer's disease in patients suffering from mild cognitive impairment (MCI). Also provided is a method for training such a system related to identifying discriminant regions of the brain and using these regions to fine tune the assistance method, based on new known cases.;Imaging data (PET or SPECT) is used, representing the cerebral activity in a plurality of spatial zones (voxels). The method then includes a normalization processing of the image data, and analysis of the cerebral activity values read in a selection of voxels forming at least one predetermined discriminant region, defined by its coordinates within a spatial reference system. | 1. An automated method for processing numerical imaging data representing at least one cerebral physiological characteristic measured quantitatively in a plurality of three-dimensional spatial zones, or voxels, within the brain of a patient suffering from mild cognitive impairment, or MCI, with a view to assisting in predicting the appearance of Alzheimer's disease, or conversion, said method comprising the following steps: normalization processing of the image data according to a standard determined as a function of the image data from a population of reference patients of known progress; andanalysis of the values of said physiological characteristic read in a selection of one or more voxels forming at least one predetermined discriminant region, defined by its coordinates within a specific spatial reference system, said analysis providing for said studied patient a value of said physiological characteristic which can be compared with a plurality of reference values read and calculated for said reference patients. 2. The method according to claim 1, characterized in that it relates to numerical data representing a three-dimensional image, the normalization processing comprising the following steps: on the one hand, a spatial normalization of the image data obtained for the studied patient, so as to provide a spatial representation of the brain according to the same specific spatial reference system which is stable from one patient to another; andon the other hand, a quantitative normalization adjusting all of the values of the physiological characteristic read within the brain of said studied patient, so as to provide an overall quantitative value of said physiological characteristic in accordance with a specific quantitative reference system which is stable from one patient to another. 3. The method according to claim 1, characterized in that it uses at least one discriminant region comprising at least two zones situated in different anatomical regions and comprising less than 75% of each of said different anatomical regions. 4. The method according to claim 1, characterized in that it uses at least one discriminant region comprising at least three zones situated in the following anatomical regions: right supramarginal gyrusright angular gyrus, andlower right parietal cortex. 5. The method according to claim 1, characterized in that it uses at least one discriminant region having at least one of the following sets of spatial coordinates:reg.MNIaccording toVoxel-levelcoordinatesthe inventionanatomical regionsPFDR-corrTZscorePuncorrectedxyzreferencename &amp; reference0.0414.214.010.00032−18−12401part of the hippocampus(602)0.0413.323.210.00132−36−6401outside the hippocampus0.0412.392.390.00934−4−22401outside the hippocampus0.0414.043.870.00058−2236501part of the rightsupramarginal gyrus(702)0.0414.013.840.00144−6636501part of the right angulargyrus (703)0.0413.733.730.00154−4656501part of the lower rightparietal cortex (704) 6. The method a... | - | A61B005/00 G06K009/00
US20110021886 | A1 | 2011-01-27 | Device and Method for Reducing the Heat Loss of Infants | 2010-07-21 | A device and method for reducing heat loss of a infant is provided. In one embodiment, the device includes a head portion having a face aperture, a middle portion connected to said head portion and having a belly aperture, wherein said belly aperture is positioned substantially above an umbilical cord of an infant disposed in the device, said middle portion having a closeable opening extending from said face aperture to said belly aperture, and a lower portion connected to said middle portion and having a leg compartment. | 1. A device for reducing heat loss of an infant, comprising: a head portion having a face aperture;a middle portion connected to said head portion and having a belly aperture;wherein said belly aperture is positioned substantially above an umbilical cord of an infant disposed in the device;said middle portion having a closeable opening extending from said face aperture to said belly aperture;a lower portion connected to said middle portion;an eye covering mechanism removably attached to the head portion to at least partially cover the face aperture to thereby cover the eyes of the infant;a temperature sensor attached into the device to assess a temperature inside the device; andan oxygen sensor attached into the device to assess a blood oxygen content of an infant in the device. 2. A device for reducing heat loss of an infant, comprising: a head portion having a face aperture;a middle portion connected to said head portion and having a belly aperture;wherein said belly aperture is positioned substantially above an umbilical cord of an infant disposed in the device;said middle portion having a closeable opening extending from said face aperture to said belly aperture; anda lower portion connected to said middle portion and having a leg compartment. 3. The device of claim 2, wherein said head portion includes a liner. 4. The device of claim 2, wherein said belly aperture includes a closeable flap. 5. The device of claim 2, further comprising an oxygen sensor integrated into the lower portion. 6. The device of claim 2, further comprising an eye covering mechanism removably attached to the head portion to partially cover the face aperture to thereby cover the eyes of the infant. 7. The device of claim 2, further comprising a first portion and a second portion connected to said middle portion and each configured to receive an arm of the infant; and wherein at least one of said first and second arm portions includes an opening and a flap covering said opening. 8. The device of claim 2, further comprising a first and second ear pad removably attached to said head portion and each located to cover an ear of the infant. 9. The device of claim 2, further comprising a temperature sensor for assessing the temperature inside the device and configured to provide temperature data to a display external to the display. 10. The device of claim 2, further comprising a heating element integrated into at least said middle portion. 11. The device of claim 2, wherein said lower portion includes a first leg portion and a second leg portion. 12. A method of forming a device, comprising: providing a head portion having a face aperture;providing a middle portion connected to said head portion and having a belly aperture;wherein said belly aperture is positioned to substantially above an umbilical cord of an infant disposed in the device;said middle portion having a closeable opening extending from said face aperture to said belly aperture and configured to be fi... | - | A61B005/00
US20110009727 | A1 | 2011-01-13 | SYSTEMS AND METHODS FOR PROCESSING, TRANSMITTING AND DISPLAYING SENSOR DATA | 2010-09-10 | Systems and methods for continuous measurement of an analyte in a host are provided. The system generally includes a continuous analyte sensor configured to continuously measure a concentration of analyte in a host and a sensor electronics module physically connected to the continuous analyte sensor during sensor use, wherein the sensor electronics module is further configured to directly wirelessly communicate displayable sensor information to a plurality of different types of display devices. | 1. A system for continuous measurement of an analyte in a host, the system comprising: a continuous analyte sensor configured to measure a concentration of an analyte in a host;a sensor electronics module physically connected to the continuous analyte sensor, wherein the sensor electronics module is configured to process a data stream associated with the analyte concentration measured by the continuous analyte sensor and to generate sensor data based on the data stream, wherein the sensor electronics module comprises a first telemetry module configured to wirelessly transmit sensor data; anda key fob device comprising a second telemetry module and a user interface, wherein the key fob device is configured to display sensor data on the user interface, and wherein the second telemetry module is configured to wirelessly receive sensor data from the sensor electronics module and to transmit sensor data. 2. The system of claim 1, wherein the key fob device is sized to fit on a key chain. 3. The system of claim 1, wherein the key fob device comprises an authentication mechanism. 4. The system of claim 1, wherein the key fob device further comprises a data storage device configured to store sensor data. 5. The system of claim 1, wherein the sensor electronics module is configured to process the data stream to provide displayable sensor information, wherein sensor data transmitted by the sensor electronics module comprises the displayable sensor information. 6. The system of claim 1, wherein the key fob device is configured to receive user input. 7. The system of claim 6, wherein the user interface of the key fob device comprises one or more buttons configured to accept the user input. 8. The system of claim 6, wherein the user input comprises calibration information. 9. The system of claim 6, wherein the telemetry module of the key fob device is configured to wirelessly transmit the user input to the sensor electronics module. 10. The system of claim 9, wherein the telemetry module of the sensor electronics module is configured to receive the user input, and wherein the sensor electronics module is configured to process the user input to provide sensor data comprising displayable sensor information. 11. The system of claim 1, wherein the key fob device is configured to display, on the user interface, a number representative of the analyte concentration based on the sensor data. 12. The system of claim 11, wherein the key fob device is further configured to display, on the user interface, an arrow representative of a rate of change of the analyte concentration based on the sensor data. 13. The system of claim 1, wherein the key fob device comprises a USB connector and a telemetry adapter configured to enable direct wireless access between the sensor electronics module and a personal computer. 14. The system of claim 1, wherein the telemetry module of the key fob device is configured to transmit sensor data wirelessly. 15. The system of clai... | DexCom, Inc. | A61B005/145
US08337422 | B2 | 2012-12-25 | Diagnostic test strip having fluid transport features | 2009-07-14 | A test strip for testing a blood sample is provided with a fluid transport feature to facilitate transport of a blood sample obtained from a lancing operation through a capillary channel to a measurement site. A fluid transport path is defined on the major face of the strip terminating at the mouth of the capillary channel. The fluid transport path includes a depending portion at one end opposite the mouth of the channel. The depending portion extends away from the strip on the side facing the fluid sample, such that a droplet of fluid sample contacting the depending portion is directed toward the mouth of the capillary channel. Thereafter the sample moves by capillary action to the measurement site. | 1. A test strip for testing a fluid sample, comprising: a strip with a first major side for positioning facing a the fluid sample;a capillary channel having a mouth and having a measurement site;a raised fluid transport path on the strip terminating at the mouth of the capillary channel;a raised depending portion that is integral to the raised fluid transport path and positioned at an end of the raised fluid transport path opposite the mouth of the capillary channel, whereby a droplet of the fluid sample contacting the raised depending portion is directed to the raised fluid transport path; anda lancet hole through the strip that is positioned laterally adjacent to the raised depending portion. 2. The test strip according to claim 1, wherein the raised depending portion is formed by two slits in the strip extending from the lancet hole in a direction toward the capillary channel. 3. The test strip according to claim 1, wherein the strip is bent, causing the raised depending portion to extend away from the first major side of the test strip by at least about 100 μm. 4. The test strip according to claim 1, wherein the lancet hole has a widest dimension of about 1 mm to about 2 mm. 5. The test strip according to claim 2, wherein the two slits have an approximately equal length of greater than about 1 mm forming the raised depending portion at the side of the lancet hole. 6. The test strip according to claim 1, wherein a portion of the strip is in a bent state, and moves along the strip during use in a direction opposite to the direction of fluid sample on the strip. 7. The test strip according to claim 1, wherein the measurement site comprises wells containing reagent for an electrochemical detection of glucose in a blood sample. 8. The test strip according to claim 1, wherein a surface of the capillary channel is hydrophilic. 9. The test strip according to claim 1, wherein a surface of the raised fluid transport path is hydrophobic. 10. The test strip according to claim 8, wherein the surface of the capillary channel has sufficient hydrophilicity so that a droplet of blood thereon forms a contact angle of less than about 50 degrees. 11. The test strip according to claim 9, wherein the surface of the raised fluid transport path has sufficient hydrophobicity so that a droplet of blood thereon forms a contact angle of greater than about 50 degrees. 12. The test strip according to claim 1, comprising a plurality of patterned polymeric structural layers that are configured to form when successively stacked the strip, the raised fluid transport path and integral raised depending portion, and the capillary channel. 13. The test strip according to claim 1, wherein the volume of the capillary channel is less than about 1.0 μL. 14. The test strip according to claim 1, further comprising two recesses respectively adjacent to each lateral side of the raised fluid transport path, each recess extending substantially the entire length of the raised f... | Becton, Dickinson and Company | A61B005/00 B65D081/00
US07889329 | B2 | 2011-02-15 | Analysis of optical data with the aid of histograms | 2008-10-08 | A system for determining the concentration of an analyte in a liquid sample comprising a detection unit for detecting light intensities which are radiated from subareas of a detection area of a test element as well as an evaluation unit which determines a frequency distribution for the detected light intensities wherein the frequency distribution has at least one first maximum caused by unwetted subareas or at least one reference area and a second maximum caused by wetted subareas and selects at least one light intensity on the basis of the frequency distribution and determines the concentration of the analyte from the at least one selected light intensity. | 1. A system for determining the concentration of an analyte in a liquid sample comprising: a detection unit for detecting light intensities which are radiated from subareas of a detection area of a test element; andan evaluation unit which determines a frequency distribution for the detected light intensities wherein the frequency distribution has at least one first maximum caused by one of at least one unwetted subarea of the detection area and at least one reference area of the test element and a second maximum caused by subareas of the detection area wetted by a sample and at least one light intensity is selected on the basis of at least one characteristic of the frequency distribution and the concentration of the analyte is determined from the at least one selected light intensity. 2. The system according to claim 1, wherein the concentration of the analyte is determined on the basis of intensity frequencies present at one of after applying the sample to the detection area and before and after applying the sample to the detection area. 3. The system according to claim 1, wherein the test element contains a reagent which is substantially homogeneously distributed throughout the detection area. 4. The system according to claim 1, wherein the sample has a volume of between 0.1 and 500 nl. 5. The system according to claim 1, wherein the concentration of the analyte is determined on the basis of rates of changes in the frequencies of the radiated light intensities after wetting the detection area. 6. The system according to claim 1, wherein the concentration of the analyte is determined on the basis of the maxima of the frequency distribution of the intensities before and after wetting the detection area. 7. The system according to claim 1, wherein an analysis is carried out using multivariate analytical methods to determine the concentration of the analyte. 8. The system according to claim 1, wherein a slope of an intensity gradient between a lowest intensity and a most frequent intensity is used to determine the analyte concentration. 9. The system according to claim 1, wherein the concentration of the analyte is determined on the basis of intensities which exceed a frequency threshold value. 10. The system according to claim 1, wherein an adequate wetting of the detection area is determined on the basis of exceeding a certain frequency of intensities which change after sample application. 11. The system according to claim 1, wherein an analysis of the detection area is carried out when a rate of change of frequencies is below a rate threshold value after application of the sample to the detection area. 12. The system according to claim 1, wherein a frequency distribution determined before the sample is applied to the test element is compared to a frequency distribution determined after the sample is applied to the test element to determine a quality characteristic of the test element. 13. The system according to claim 12, wherein th... | Roche Diagnostics Operations, Inc. | G01N033/48 G01N021/00 G01N021/55 G01N021/86 G01N021/75 G01J003/46 A61B005/00 G01V008/00
US07878987 | B2 | 2011-02-01 | Methods and instruments for assessing bone fracture risk | 2006-05-04 | Methods and instruments for assessing bone, for example fracture risk, in a subject in which a test probe is inserted through the skin of the subject so that the test probe contacts the subject's bone and the resistance of the test bone to microscopic fracture by the test probe is determined. Macroscopic bone fracture risk is assessed by measuring the resistance of the bone to microscopic fractures caused by the test probe. The microscopic fractures are so small that they pose negligible health risks. The instrument may also be useful in characterizing other materials, especially if it is necessary to penetrate a layer to get to the material to be characterized. | 1. A method for assessing bone fracture risk in a subject, the method comprising: inserting a probe assembly through skin and soft tissue of the subject down to a bone, the probe assembly comprising a reference probe having a tubular section that is tapered asymmetrically and a test probe slidably disposed within the tubular section;establishing a stationary contact point directly on the surface of the bone with the reference probe and the test probe, said establishing comprising exposing bone with the reference probe;inserting the test probe into the bone; andassessing fracture risk of the bone by evaluating one or more parameters derived from the insertion of the test probe, wherein the parameters are selected from the group consisting of: (i) extent of insertion of the test probe into the bone, and (ii) resistance of bone to penetration of the test probe. 2. The method of claim 1, wherein the one or more parameters also comprises resistance of the bone to removal of the test probe after it penetrates bone. 3. The method of claim 1, wherein the test probe is inserted multiple times at the stationary contact point. 4. The method of claim 1, wherein insertion of the test probe into the bone creates one or more microscopic fractures in the bone. 5. The method of claim 1, wherein the evaluation of one or more parameters is performed by a computer. 6. A diagnostic instrument for assessing bone, comprising: a reference probe comprising a distal tubular section that is tapered asymmetrically;a test probe disposed within the reference probe and configured for movement relative to the reference probe, wherein the reference probe and test probe are configured for insertion through skin so as to place a distal tip of the reference probe in stationary contact with a surface of exposed bone and wherein the test probe is configured for cyclical movement with respect to the reference probe, the cyclical movement comprising repeated insertion and withdrawal of the test probe into the same location of the bone;a displacement sensor configured to measure displacement of the test probe into the bone; anda force sensor configured to measure the force to insert and withdraw the test probe. 7. The diagnostic instrument according to claim 6, wherein the test probe is tapered asymmetrically. 8. The diagnostic instrument according to claim 6, wherein the test probe comprises a rod having a conical tip. 9. The diagnostic instrument according to claim 6, further comprising a guide operatively connected to the reference probe, the reference probe removably mounted to the guide. 10. The diagnostic instrument according to claim 9, wherein the guide comprises a threaded neck disposed at a distal end thereof, the reference probe comprising an internally threaded opening configured to mate with the threaded neck. 11. The diagnostic instrument according to claim 6, further comprising a force generator operatively connected to the test probe. 12. The diagnostic instru... | The Regents of the University of California | A61B005/103 A61B005/117 G01N003/48 G01N003/42 G01N003/32 G01N003/44
US07869858 | B2 | 2011-01-11 | Patient table system and apparatus | 2007-04-26 | A patient table is disclosed. The patient table includes a cradle for linearly receiving and transferring a board comprising a top surface defined by a single plane for carrying a patient, a cradle supporting member for slidably receiving the cradle, a locker within the cradle for positioning the board and patient on the table, and for preventing the board from sliding off the cradle, and a stopper within the cradle for preventing the cradle form sliding off the cradle supporting member. | 1. A patient table comprising: a cradle for linearly receiving and transferring a board comprising a top surface defined by a single plane for carrying a patient;a cradle supporting member for slidably receiving the cradle;a locker extending through the cradle for positioning the board and patient on the cradle, the locker extending into and out of engagement with the board to prevent the board from sliding off the cradle; and,a stopper within the cradle for preventing the cradle from sliding off the cradle supporting member. 2. The patient table of claim 1 wherein the board comprises a frame and an insert defining the single plane. 3. The patient table of claim 1 wherein the board further comprises a set of accessory interfaces disposed along edges of the board. 4. The patient table of claim 3, further comprising: a patient restraint feature disposed upon the top surface of the board via at least one accessory interface of the set of accessory interfaces. 5. The patient table of claim 1, wherein the board is a one-piece board. 6. The patient table of claim 1 wherein the locker includes a locker manual release mechanism. 7. The patient table of claim 1 wherein the locker is automatically disengaged with the board when the patient table is docked with a second table capable of linearly receiving the board. 8. The patient table of claim 1 wherein the stopper is disengaged from the cradle supporting member when the patient table abuts a complementary support system capable of receiving the cradle. 9. The patient table of claim 1 wherein the stopper is disengaged from the cradle supporting member when the board is fully supported upon the cradle. 10. The patient table of claim 1 further comprising a radio frequency (RF) coil for use within an imaging process, a coil cable extending from the RF coil, a cable attachment box attached to the patient table for receiving the coil cable, and a system cable extending from the cable attachment box for attachment with an imaging system. 11. The patient table of claim 10, wherein the RF coil is disposed within the board. 12. The patient table of claim 10 wherein the RF coil is positioned by a hinged coil support attachable to the board, the coil disposed between the hinged coil support and the patient. 13. The patient table of claim 10 wherein the RF coil includes two lateral flex coils, and a coil cable extending from each of the lateral flex coils to the cable attachment box. 14. The patient table of claim 1 further comprising a docking assembly attached to a first end of the patient table, the docking assembly configured for attaching the first end of the patient table to a second table. 15. The patient table of claim 1 wherein the cradle and the transfer board are formed of an X-ray and magnetic resonance compatible fiber composite. | General Electric Company | A61B005/05 A47B071/00 A47B013/00 G01N023/04
US07873411 | B2 | 2011-01-18 | Interface device, interface method and control training device by the use of the interface device | 2005-08-30 | An interface device in accordance with the present claimed invention is an interface device 4 that intervenes between a measuring means 2 that measures brain activities of a subject H and an object 31 to be controlled, and that comprises a reference signal storing part D1 that stores a brain activity signal measured at a time when the subject H conducts predetermined thinking different from direct thinking in order to control the object 31 to be controlled in a desired mode as a reference signal, a discriminating part 41 that compares the brain activity signal measured by the measuring means 2 with the reference signal and discriminates whether or not the thinking of the subject H is the predetermined thinking, and a controlling part 42 that controls the object 31 to be controlled in the desired mode when the discriminating part 41 discriminates that the thinking of the subject H is the predetermined thinking. | 1. An interface device intervening between a measuring unit that measures brain activities of a subject and an object to be controlled, and comprising: a reference signal storing part that stores a brain activity signal measured at a time when the subject conducts predetermined thinking which causes movement of the object and is different from direct thinking in order to control the object to be controlled in a desired mode as a reference signal, wherein the predetermined thinking is multiple kinds of thinking about movements of body sites, the predetermined thinking is conducted in regions of the subject's brain corresponding to areas in a Penfield's brain functional map representative of portions of the brain that are separate from each other and correspond to the body sites,a discriminating part that compares the brain activity signal measured by the measuring unit with the reference signal and discriminates whether or not the thinking of the subject is the predetermined thinking,a controlling part that controls the object to be controlled in the desired mode at a time when the discriminating part discriminates that the thinking of the subject is the predetermined thinking, andan image display unit to provide the subject with predetermined information that is arranged on the image display unit, wherein the controlling part displays an indication on a screen to aid the subject in conducting the predetermined thinking when the object is being controlled and a control content related to the predetermined thinking, wherein the indication is an image showing movements of the body sites. 2. The interface device described in claim 1, wherein the predetermined thinking is six or more kinds of thinking wherein each region of a brain that becomes active at a time when the subject conducts each thinking differs mutually. 3. The interface device described in claim 1, wherein the controlling part further displays a controlling content of the desired mode on the screen. 4. The interface device described in claim 1, further comprising an image display unit to provide the subject with information, wherein the object to be controlled is a moving body displayed on a screen of the image display unit, and wherein the controlling part controls the moving body. 5. The interface device described in claim 4, wherein the moving body moves on a two-dimensional surface displayed on the screen of the image display unit. 6. The interface device described in claim 4, wherein the moving body moves in three-dimensional space displayed on the screen of the image display unit. 7. The interface device described in claim 4, wherein behavior of the moving body from a starting point to a destination point is set to be a parameter and an evaluating part that evaluates the parameter is provided. 8. The interface device described in claim 7, wherein the parameter is time until the moving body reaches the destination point from the starting point. 9. The interface device descr... | National Institute of Information and Communications Technology | A61B005/04
US07894874 | B2 | 2011-02-22 | Method and apparatus for enhancing the detecting and tracking of moving objects using ultrasound | 2006-05-08 | An ultrasonic pulse echo apparatus detects an object that is moving with respect to stationary objects. An ultrasonic transducer transmits a series of ultrasound pulses in a direction that intersects a path of the moving object. An ultrasound receiver receives a series of lines of echoes from objects in the field of view of the ultrasonic signal. Each echo line corresponds to one of the ultrasonic pulses. A signal processor processes the echo lines from the moving object. The echo lines are time shifted by different amounts and combined at different time shifts to produce different composite lines. The composite line having an optimal signal-to-noise ratio is selected. Other signal processing enhancements are performed. | 1. An ultrasonic pulse echo apparatus for detecting an object that is moving with respect to a stationary object, comprising: an ultrasound transducer for transmitting a series of ultrasound pulses in a direction that intersects a path of the moving object;an ultrasound receiver for receiving multiple lines of echoes from the objects in the field of view of the ultrasonic signal, each echo line corresponding to one of the ultrasonic pulses; anda signal processor for determining an envelope of each of the echo lines, time shifting the echo envelope lines different amounts, and preferentially enhancing the echoes from moving objects by combining the echo envelope lines at different time shifts and selecting one of the combined echo envelope lines with an optimal signal-to-noise ratio. 2. The apparatus in claim 1, wherein the time shifting aligns the echoes in the echo envelope lines, and wherein the signal processor is configured to determine an amplitude of each echo envelope at a particular range from the ultrasound transducer for each of the different time shifts. 3. The apparatus in claim 2, wherein the signal processor is configured to: determine a median value or an average value of the envelope amplitudes at each time shift,select the time shift where the envelope amplitudes have a minimum variance from the median value or average value, andcombine the echo envelope lines at the selected time shift to produce the combined echo envelope signal with the optimal signal-to-noise ratio for the particular range. 4. The apparatus in claim 2, wherein the signal processor is configured to: determine a median value of the echo envelope amplitudes at each time shift,select the time shift where the echo envelope amplitudes have a maximum median value, andcombine the echo envelope lines at the selected time shift to produce the combined echo envelope signal with the optimal signal-to-noise ratio for the particular range. 5. The apparatus in claim 1, wherein the signal processor is configured to sample each echo envelope signal at a reduced rate relative to higher-frequency sampling in the received echo lines required for accurate echo representation. 6. The apparatus in claim 5, wherein the reduced rate ensures that at least three sample points are contained within an envelope formed from a one-cycle ultrasonic echo. 7. The apparatus in claim 1, wherein the processor is configured to determine an envelope of each of the echo lines using a Hilbert transform and determine an envelope of each of the echo lines and remove a DC bias using a negative Hilbert transform. 8. The apparatus in claim 1, wherein the processor is configured to remove echoes associated with stationary objects that may be present in each of the echo lines. 9. The apparatus in claim 1, wherein the processor is configured to filter the combined echo envelope signal to remove one or more false peaks and to detect a peak corresponding to the moving object echo in each filtered echo ... | Luna Innovations Incorporated | A61B005/05 A61B008/08 A61B008/12 A61B008/14
US07869863 | B2 | 2011-01-11 | Apparatus and method for non-invasive, passive fetal heart monitoring | 2008-01-10 | The present invention provides a method and apparatus for monitoring the fetal condition in an objective, quantitative, and non-invasive manner by enabling continuous and highly specific data collection during pre-natal development and the intra-partum period. In an embodiment, monitoring the fetal condition comprises monitoring a fetal heart rate, a fetal heart rate variability, and/or the fetal heart vector orientation of a fetus in a pregnant mother. Automated signal processing techniques are applied to a multiplicity of transducers to select an optimum signal vector to monitor at least one of a fetal heart rate, a fetal heart rate variability, and/or the fetal heart vector orientation of a fetus in a pregnant mother. | 1. An automated method of monitoring the fetal electrocardiographic waveform of a fetus in a pregnant mother to extract a high quality fetal signal from an abundance of signal channels, said method comprising: a) measuring, in real-time, biopotential waveforms indicative of either one of a maternal heart or a maternal heart beat and a fetal heart beat;b) isolating a single biopotential waveform from among said measured biopotential waveforms for use as a reference fetal waveform;c) enhancing the isolated reference fetal waveform to remove correlated noise;d) monitoring the enhanced isolated reference fetal waveform in real-time via at least one fetal monitoring algorithm;e) monitoring an output quality metric via at least one quality metric monitoring algorithm to determine if it is within an acceptable range; andf) dynamically adjusting at least one processing parameter of the at least one fetal monitoring algorithm to maintain the output quality metric within said acceptable range. 2. A method according to claim 1, wherein said measuring step further comprises: acquiring said biopotential waveforms from at least one sensor located external to the mother;forming a plurality of differential electrocardiogram (ECG) bipolar vectors from the acquired biopotential waveforms. 3. A method according to claim 1, wherein said plurality of biopotential waveform signals are acquired as unipolar signals. 4. A method according to claim 1, wherein said isolating step further comprises: extracting a high quality fetal signal from an abundance of abdominal channels for use as said reference fetal waveform. 5. A method according to claim 1, wherein said enhancing step further comprises: linearly combining said measured biopotential waveforms to maximize the reference fetal waveform and reduce uncorrelated noise. 6. A method according to claim 1, wherein said monitoring step (d) further comprises: monitoring the enhanced reference fetal waveform over a short term re-evaluation interval on the order of substantially 10 to 30 seconds to determine whether the fetal heart rate is present and is within physiologically acceptable bounds. 7. A method according to claim 1, wherein said step (e) of monitoring said output quality metric further comprises: monitoring the enhanced reference fetal waveform over a long term re-evaluation interval on the order of substantially 1 to 2 minutes to determine whether a fetal heart vector has moved within a monitoring montage array. 8. A method according to claim 1, wherein said monitoring of said output quality metric further comprises: utilizing a quality score associated with the quality of the fetal waveform, based on one or more metrics, anddynamically adjusting appropriate parameters when said quality score falls below a predetermined threshold to ensure that the quality of the fetal waveform is optimized. 9. A method according to claim 4, wherein said quality metric is selected from the group consisting of: (a) outliers... | The Johns Hopkins University | A61B005/04
US07895061 | B2 | 2011-02-22 | Auctioning provider prices | 2008-03-31 | A computer-implemented method includes receiving, at a computer-based system for connecting consumers with providers in real time over the Internet, requested compensation amounts from a plurality of providers. The method also includes displaying a price associated with consulting with a particular provider to a consumer on a user interface, the price being based at least in part on the requested compensation amount received from the provider. | 1. A computer-implemented method, comprising: receiving, at a computer-based system for connecting consumers with medical service providers in real time over the Internet, requested compensation amounts from a plurality of medical service providers,receiving, at the computer-based system, a request from a consumer to consult with a medical service provider having a profile that satisfies at least some attributes in a set of attributes provided by the consumer in the request with one or more of the attributes being associated with medical qualifications;identifying, by the computer-based system, a plurality of medical service providers satisfying at least some of the attributes in the request from the consumer;determining prices associated with consulting with at least some of identified medical service providers with the price being based at least in part on the requested compensation amount received from the medical service provider;generating a user interface that, when rendered on a device, displays at least two of the identified medical service providers and the determined prices associated with consulting with the identified medical service providers to the consumer;receiving, from the consumer, a selection of one of the identified medical service providers from the at least two of the identified medical service providers displayed on the user interface; andproviding a communication channel to establish a communication between the consumer and the selected one of the medical service providers. 2. The method of claim 1, wherein: receiving the requested compensation amount from the plurality of medical services medical service providers comprises:receiving a first requested compensation amount from a first medical service provider; andreceiving a second requested compensation amount from a second medical service provider, the second requested compensation amount being different than the first requested compensation amount; andgenerating the user interface comprises:generating the user interface to display an identification of the first medical service provider and a price based on the received first requested compensation amount from a first medical service provider and an identification of the second medical service provider and a price based on the received second requested compensation amount from a second medical service provider. 3. The method of claim 2, further comprising: receiving, from the consumer, a selection of one of the first medical service provider or the second medical service provider; andproviding a communication channel to establish a communication between the consumer and the selected one of the first medical service provider or the second medical service provider. 4. The method of claim 3, further comprising: providing to a non-selected one of the first medical service provider the second medical service provider an indication of the requested compensation of the selected medical service provider. 5. The method of cl... | American Well Corporation | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07867174 | B2 | 2011-01-11 | Systems and methods for obtaining analytes from a body | 2008-04-01 | A method may include collecting at least one analyte from within a body, ejecting the collected at least one analyte from the body through at least one dermal layer of the body, and receiving the ejected at least one analyte outside the body. A system may include a means for collecting at least one analyte from within a body, a means for ejecting the collected at least one analyte from the body through at least one dermal layer of the body, and a means for receiving the ejected at least one analyte outside the body. | 1. A system, comprising: means for collecting at least one analyte from within a body;means for ejecting said at least one analyte from said body through at least one dermal layer of said body, wherein said means for ejecting includes means for ejecting said at least one analyte from said body through at least one dermal layer of said body under high pressure; andmeans for receiving said at least one analyte outside said body. 2. The system of claim 1, further comprising: means for transferring said at least one analyte from a sampling device within said body to an ejector device within said body. 3. The system of claim 2, wherein said sampling device is located distal to said ejector device and said at least one analyte is transported to said ejector device. 4. The system of claim 1, further comprising: means for transmitting a location signal enabling an ejector device to be located. 5. The system of claim 4, wherein said location signal comprises at least one of an electrical current, an electrical field, a magnetic flux, an optical signal, a radio frequency identification, an ultrasound, a vibration, an electromagnetic force, a force, or a pressure. 6. The system of claim 1, further comprising: means for sorting said at least one analyte. 7. The system of claim 1, further comprising: means for concentrating said at least one analyte. 8. The system of claim 1, further comprising: means for diluting said at least one analyte. 9. The system of claim 1, wherein said means for collecting at least one analyte from within a body comprises: means for collecting at least a portion of at least one of cells, proteins, bacteria, blood, a blood component, molecules, viruses, viral particles, pathogens, parasites, malarial parasites, oglionucleotides, lymph, a lymph component, or cerebral spinal fluid. 10. The system of claim 9, wherein said blood component comprises at least one of plasma or serum. 11. The system of claim 1, wherein said means for receiving said at least one analyte outside said body comprises: means for receiving said at least one analyte outside said body in a receiving device. 12. A system, comprising: means for collecting at least one analyte from within a body;means for ejecting said at least one analyte from said body through at least one dermal layer of said body;means for receiving said at least one analyte outside said body; andmeans for analyzing said at least one analyte within said body. 13. A system, comprising: means for collecting at least one analyte from within a body;means for ejecting said at least one analyte from said body through at least one dermal layer of said body;means for receiving said at least one analyte outside said body; andmeans for aligning said receiving device with an ejector device. 14. The system of claim 13, wherein said means for aligning said receiving device with an ejector device comprises: means for locating a fiducial. 15. The system of claim 14, wherein said body includes said f... | The Invention Science Fund I | A61B005/00
US08292832 | B2 | 2012-10-23 | Event-based health activity tracking with icon-based user interface | 2009-07-27 | A health activity plan may be prescribed for an individual comprising a set of health activity instances (e.g., physical therapy exercises, an exercise regimen, or a wound management program.) The performed health activity instances of the individual may be tracked by a device, and may be presented to a user of the device in correlation with the health activity plan (e.g., to demonstrate compliance of the individual with the health activity plan.) The health activity information may be presented using a set of icons that depict the health activity information, which may improve an easier understanding and a more consistent communication of health activity information among the individual and healthcare providers. | 1. A method of tracking health activities prescribed for an individual on a computer accessible to a user and having a processor, a memory, at least one input component, at least one display, and an icon store configured to store health activity icons depicting health activity types, the method comprising: executing on the processor instructions configured to: store in the memory a health activity plan comprising at least one prescribed health activity instance of a prescribed health activity type;upon receiving from at least one input component an at least partial recording of an activity of a person: store the at least partial recording of the activity of the person,using the at least partial recording of the activity of the person, identify a performed health activity instance of a performed health activity type relating to the individual, andstore in the memory a performed health activity record representing the performed health activity instance of the performed health activity type relating to the individual; andupon receiving from the user a request to present the performed health activity instances, present on the display component, for respective performed health activity records: an icon selected from the icon store depicting the performed health activity type,an at least partial recording of at least one activity of a person associated with a health activity instance, andan indicator of at least one prescribed health activity instance correlating to the performed health activity instance. 2. The method of claim 1: the instructions further configured to receive the health activity plan by displaying for the user icons selected from the icon store and depicting health activity types that may be prescribed for the individual; andstoring the health activity plan comprising: upon receiving a selection by the user of at least one icon, storing in the memory a prescribed health activity instance of the health activity type represented by the selected icon. 3. The method of claim 1, the input component configured to receive the notification of the performed health activity instance from the user. 4. The method of claim 3, wherein the identifying the performed health activity instances using the at least partial recording comprises: displaying for the user the at least partial recording of the activity of the person;displaying for the user at least one icon selected from the icon store and depicting the health activities that may be performed by the individual, andreceiving from the user a selection of at least one icon. 5. The method of claim 1, the performed health activity record further comprising at least one performed health activity instance descriptor selected from a set of health activity instance descriptors comprising: the performed health activity type;the time of the performed health activity instance;a proficiency of the performed health activity instance;an annotation received from the user and describing the preformed health... |  | A61B005/11
US08409113 | B2 | 2013-04-02 | Body fluid sampling unit | 2010-10-21 | A body fluid sampling unit includes a puncture device including a needle having a sharp needle point and a drive mechanism for operating the needle to puncture a living body surface with the needle point, a tip including an introducing section for introducing a body fluid flowing out from a puncture portion of the living body surface punctured with the needle point, a test paper for detecting a predetermined component in the body fluid, and a tip body mounted to a component measuring device for measuring the quantity and/or property of the predetermined component detected by the test paper; and a case including a mechanism containing section for containing the puncture mechanism so that the needle point of the puncture mechanism can be moved by the drive means, and a sampling implement containing section for containing the body fluid sampling implement so that the tip can be taken out. | 1. A method of obtaining a body fluid sample from a living body and using the body fluid sample comprising: removing a body fluid sampling implement from an interior of a container, the body fluid sampling implement comprising a reagent containing part comprising a reagent adapted to interact with a predetermined component in the body fluid, the interior of the container also containing a needle possessing a sharp needle point at a distal end of the needle;mounting the body fluid sampling implement on a component measuring device, the component measuring device being adapted to measure at least one of a quantity and a property of the predetermined component of the body fluid obtained by the body fluid sampling implement;the body fluid sampling implement being removed from the interior of the container either before mounting the body fluid sampling implement on the component measuring device or after mounting the body fluid sampling implement on the component measuring device;puncturing a surface of the living body with the needle by moving the needle relative to the container to produce a body fluid sample; andintroducing to the reagent containing part of the body fluid sampling implement the body fluid sample obtained by puncturing the surface of the living body with the needle so that the reagent interacts with the predetermined component in the body fluid sample. 2. The method as set forth in claim 1, wherein the body fluid sampling implement is removed from the interior of the container before mounting the body fluid sampling implement on the component measuring device, the body fluid sampling implement being mounted on the component measuring device after removing the body fluid sampling implement from the interior of the container. 3. The method as set forth in claim 1, wherein the body fluid sampling implement is mounted on the component measuring device as the body fluid sampling implement is in the interior of the container, the body fluid sampling implement being removed from the interior of the container after the body fluid sampling implement is mounted on the component measuring device. 4. The method as set forth in claim 3, wherein the body fluid sampling implement is mounted on the component measuring device by pushing the component measuring device towards the body fluid sampling implement, the pushing of the component measuring device causing a spring in the interior of the container to accumulate a spring bias which is applied to the needle during the puncturing of the surface of the living body. 5. The method as set forth in claim 1, wherein the interior of the container in which is located the body fluid sampling implement is covered by a removable membrane, and wherein the membrane is removed to expose the interior of the container before mounting the body fluid sampling implement on the component measuring device. 6. A method of obtaining a body fluid sample from a living body and using the body fluid sample comprising:... | Terumo Kabushiki Kaisha | A61B005/00 A61F002/02 B65D081/00
US08452374 | B2 | 2013-05-28 | Apparatus and method for substantially immobilizing a breast for medical imaging procedure | 2010-08-06 | Apparatus and methods are described for substantially immobilizing a breast for use in medical imaging. In an exemplary embodiment, an apparatus for substantially immobilizing a breast is provided, including a first support member connected to a first and second base, the bases capable of engagement with the torso of a patient, the first support member defining an inner area substantially covered by a first membrane. A corresponding second support member, defining a second inner area substantially covered by a second membrane, is engageable to the first support member such that the inner areas are substantially aligned to substantially immobilize a breast, the membranes defining a pocket to receive a breast. | 1. An apparatus for substantially immobilizing a breast of a patient relative to a torso of the patient, the apparatus comprising: a first support member having a first and a second base each capable of engagement with the torso, the first support member defining a first inner area and when the each base is engaged with the torso the first support member extending outwardly from the first base in a path over the breast to the second base;a corresponding second support member removably engageable to the first support member the second support member defining a second inner area and when engaged to the first support member the first and second inner areas in substantial alignment;a first and a second deformable membrane, the first deformable membrane removably connected to the first support member substantially covering the first inner area and the second deformable membrane removably connected to the second support member substantially covering the second inner area, the first and second membranes defining a pocket configured to receive the breast, to substantially immobilize the breast relative to the torso of the patient. 2. The apparatus of claim 1, further comprising a locking mechanism capable of locking the second support member to the first support member, when each base is engaged with the torso. 3. The apparatus of claim 2, wherein the first and the second membranes are non-magnetic and acoustically permeable. 4. The apparatus of claim 3, wherein each of the first and second membranes has a surface being an adhesive configured for adherence to at least one of the breast and the adhesive surface of the other of the membranes. 5. The apparatus of claim 4, wherein the first and second base are flexibly configured to conform to the torso of the patient. 6. The apparatus of claim 5, wherein the second base is hingedly connected to the first support member capable of engagement with a side of the torso of the patient. 7. The apparatus of claim 5, further comprising an extended membrane portion connected to at least one of the first and the second deformable membrane, the extended membrane portion having an adhesive surface capable of removable adherence to the torso. 8. The apparatus of claim 5, further comprising fidicual markers engaged to at least one of the first and second support members. 9. The apparatus of claim 5, further comprising a chest support member connected to the first and second bases capable of engagement with the torso. 10. An apparatus for substantially immobilizing a breast of a patient relative to a torso of the patient, the apparatus comprising: a first support member having a first and a second base each capable of engagement with the torso, the first support member defining a first inner area and when the each base is engaged with the torso the first support member extending outwardly from the first base in a path over the breast to the second base;a corresponding second support member removably hinged to the... | Hologic, Inc. | A61B005/055 H05G001/00
US08645164 | B2 | 2014-02-04 | Medical information visualization assistant system and method | 2010-05-27 | Methods, systems, devices and/or apparatus related to displaying and visualizing medical information for facilitating effective analysis by medical professionals. Example systems for displaying medical information related to a human subject may include a computer system, server system(s), data monitoring device(s), electronic medical record/interoperability tool, medical information database(s) and a medical information visualization tool having an interface, software and management database(s). Such system may simultaneously output at least two medical parameters (relating to near real-time medical information and historical medical information) to a display device using the medical information visualization tool. | 1. A system for displaying medical information related to a human subject, comprising: a computer system;one or more server systems in communication with the computer system;one or more data monitoring devices providing at least near real-time medical information to the computer system and/or the one or more server systems;one or more electronic medical records in communication with the one or more server systems, each electronic medical record being configured to store historical medical information associated with the human subject;a medical information interoperability tool in communication with the one or more electronic medical records, the medical information interoperability tool configured to translate and transmit real-time and/or historical medical information among the one or more data monitoring devices, one or more electronic medical records and the one or more server systems; anda medical information visualization tool stored on the one or more server systems and accessible by the computer system, the medical information visualization tool being configured to display at least near real-time medical information from the data monitoring devices and historical medical information from at least one of the electronic medical records and/or the medical information interoperability tool, the at least near real-time and historical medical information including medical parameters related to at least one of clinical data and human subject-generated data;wherein the medical information visualization tool includes a touch-screen operated scrubber tool that is configured to select a particular time on a timeline and to display the historical medical information on a graph having a dynamically adjustable time period, the displayed historical medical information including numeric values that are representative of the medical parameters, wherein a particular time on the timeline is selected by touching the scrubber tool on the touch-screen and sliding the scrubber tool along the timeline to a corresponding location representative of the particular time being selected;wherein the medical information visualization tool is configured to simultaneously display at least two parameters of the historical medical information and the at least near real-time medical information, the historical medical information being displayed on the graph having a dynamically adjustable time period; andwherein the medical information visualization tool is configured to allow a healthcare clinician to select and view a range of physiological data parameters for a patient over a select time period, the range of physiological data parameters being selectable and viewable from at least one customizable perspective selected from the group consisting of a clinician specific perspective, a patient specific perspective and a medical condition specific perspective. 2. The system of claim 1, wherein the medical information visualization tool is configured to be dynamically adjusta... | Indiana University Research and Technology Corporation | G06F019/00
US08649854 | B2 | 2014-02-11 | Method and device for the determination of murmur frequency band | 2009-02-06 | The present invention is related to a method for the determination of frequency band characteristics of a heart disease. A first set of phonocardiograms are recorded from a first set of reference healthy patients, and a second set of phonocardiograms from a second set of patients suffering of a heart disease. Spectral energies of all possible frequency bands are then calculated. These spectral energies are then compared in order to determine an optimized frequency band that gives rise to the maximal distinction between spectral energies of the phonocardiograms from first and second set of phonocardiograms. | 1. A method for a determination of frequency band characteristic of a heart disease comprising the steps of: using an electronic stethoscope, collecting and/or recording a first set of phonocardiograms from a first set of reference healthy patients, and a second set of phonocardiograms from a second set of patients suffering of said heart disease;defining a frequency sampling rate and a maximum cutoff frequency, producing a finite set of sampled frequencies;calculating spectral energies at all sampled frequencies, said spectral energies being calculated on the entire phonocardiogram;calculating spectral energies of all possible frequency bands by summing spectral energies at sample frequencies belonging to said frequency bands;comparing said spectral energies of all possible frequency bands; anddetermining a frequency band giving rise to a maximal distinction between spectral energies of the said phonocardiograms from the first and the second set of phonocardiograms. 2. A method for a detection of heart diseases comprising the steps of: determining frequency band giving rise to a maximal distinction between spectral energies of the said phonocardiograms from the first and the second set of phonocardiograms by using the method of claim 1; andcomparing said determined frequency band energy of phonocardiograms of a patient to corresponding frequency band energy of the first and the second set of phonocardiograms to define if said patient suffers of the heart disease. 3. The method according to claim 1 wherein the heart disease is selected from the group consisting of ventricular septal defect, patent ductus arteriosus, pulmonary valve stenosis, tetralogy of fallot, aortic stenosis, atrial septal defect, coarctation of aorta, transposition of great arteries, atrioventricular septal defect, truncus arteriosus, tricuspid atresia, anomalous pulmonary venous connection or a mixture thereof. 4. The method according to claim 1 wherein the maximum cutoff frequency is about 500 Hz. 5. The method according to claim 2 wherein the determined frequency band is comprised of between 88 Hz and 96 Hz and the heart disease is ventricular septal defect. 6. The method according to claim 2 wherein the determined frequency band is comprised of between 50 Hz and 60 Hz and the heart disease is tetralogy of fallot. 7. The method according to claim 2 wherein the determined frequency band is comprised of between 15 Hz and 121 Hz and the heart disease is ventricular patent ductus arteriosus. 8. The method according to claim 2 wherein the determined frequency band is comprised of between 17 Hz and 129 Hz and the heart disease is aortic stenosis. 9. The method according to claim 2 wherein the determined frequency band is comprised of between 16 Hz and 32 Hz and the heart disease is pulmonary valve stenosis. 10. The method according to claim 1 wherein the spectral energies are calculated only in the systolic segment of the phonocardiogram. 11. The method according to cl... | Capis SPRL | A61B007/04
US08672855 | B2 | 2014-03-18 | Methods and systems that monitor for an impending myocardial infarction | 2009-07-08 | Implantable systems, and methods for use therewith, are provided for monitoring for an impending myocardial infarction. A signal indicative of changes in arterial blood volume is obtained. Such a signal can be a photoplethysmography signal or an impedance plethysmography signal. For each of a plurality of periods of time, a metric indicative of the areas under the curve of the signal or number of inflections in the signal is determined. An impending myocardial infarction is monitored for based on changes in the metric indicative of the area under the curve of the signal or number of inflections in the signal, and an alert and/or therapy is triggered in response to an impending myocardial infarction being predicted. | 1. For use with an implanted system, a method for monitoring for an impending myocardial infarction (MI), the method comprising: (a) using an implanted photoplethysmography sensor remote from the patient's heart to produce a photoplethysmography signal that is indicative of changes in arterial volume;(b) for each of a plurality of periods of time, determining a metric indicative of the area under the curve of the photoplethysmography signal for the period of time;(c) monitoring for an impending MI based on changes in the metric indicative of the area under the curve of the photoplethysmography signal; and(d) triggering an alert and/or therapy in response to an impending MI being detected at step (c); wherein one or more of steps (b), (c) and (d) is/are performed using a processor. 2. The method of claim 1, further comprising determining a baseline of the metric indicative of area under the curve; wherein step (c) includes distinguishing between an impending MI and a transient myocardial ischemic event based on a length of time that the metric indicative of the area under the curve falls below the baseline by at least a specified threshold. 3. The method of claim 1, further comprising determining a baseline of the metric indicative of area under the curve; wherein step (c) comprises:(c.1) comparing the determined metrics indicative of the area under the curve to the baseline; and(c.2) predicting an impending MI if the determined metric indicative of the area under the curve falls below the baseline by at least a specified threshold for at least a specified length of time. 4. The method of claim 3, wherein step (c.2) includes distinguishing between an impending MI and a transient myocardial ischemic event based on a length of time that the metric indicative of the area under the curve falls below the baseline by at least the specified threshold. 5. The method of claim 1, wherein step (c) comprises predicting an impending MI if a decrease beyond a specified threshold is detected, in the metric indicative of the area under the curve of the photoplethysmography signal, from one of the periods of time to the immediately following one of the periods of time. 6. The method of claim 1, wherein step (c) comprises predicting an impending MI if the metric indicative of the area under the curve of the photoplethysmography signal decreases by at least a specified amount for at least a specified period of time. 7. The method of claim 1, wherein the metric indicative of area under the curve is selected from the group consisting of: an integral of at least one cycle of the photoplethysmography signal;a width from the start to the end of a cycle of the photoplethysmography signal; anda full width at half maximum (FWHM) of a cycle of the photoplethysmography signal. 8. The method of claim 1, wherein step (c) includes: (c.1) detecting changes in the area under the curve of the photoplethysmography signal that are indicative of changes in vascular stiffness; a... | Pacesetter, Inc. | A61B005/0295
US08784310 | B2 | 2014-07-22 | Vascular pressure sensor with electrocardiogram electrodes | 2010-05-24 | Devices, systems, and methods for sensing cardiac electrical activity and blood pressure within a blood vessel are disclosed. A vascular sensor includes a fixation element configured to secure the vascular sensor within a blood vessel, a sensor module housing coupled to the fixation element, a pressure sensor disposed within an interior portion of the housing and configured to sense blood pressure within the vessel, and a number of electrocardiogram sensing electrodes exposed to tissue and blood within the blood vessel and configured to sense cardiac electrical data within the blood vessel. A controller module processes the sensed cardiac electrical data and blood pressure data and stores the data within a memory unit and/or transmits the data to a communicating device in wireless communication with the vascular sensor. | 1. A leadless implantable vascular sensor, comprising: at least one fixation element configured to secure the vascular sensor within a blood vessel;a sensor module housing coupled to the fixation element and including: a sensor housing including a plurality of exterior surfaces defining an interior portion of the sensor housing;a power supply housing coupled to the sensor housing and containing an energy storage device;a pressure sensor disposed within the interior portion of the sensor housing, the pressure sensor configured to sense blood pressure data within the blood vessel;a first electrocardiogram sensing electrode disposed on an exterior portion of the sensor housing;a second electrocardiogram sensing electrode disposed on an exterior portion of the power supply housing;an insulative break disposed on at least a portion of at least one of the plurality of exterior surfaces of the sensor housing and electrically isolating the first electrocardiogram sensing electrode from the second electrocardiogram sensing electrode,the sensor module housing further comprising a pair of parallel and opposed ends, the pair of parallel and opposed ends comprising a first end disposed on the sensor housing and a second end disposed on the power supply housing, and a plurality of sides disposed between the first end and the second end and including the plurality of exterior surfaces of the sensor housing, wherein: the first electrocardiogram sensing electrode is disposed on the first end; andthe second electrocardiogram sensing electrode is disposed on the second end;the first and second electrocardiogram sensing electrodes configured to sense cardiac electrical data within the blood vessel; andcircuitry means for processing the cardiac electrical data and blood pressure data sensed by the electrocardiogram sensing electrodes and pressure sensor to correlate the timing of the blood pressure data with a timing of a cardiac cycle. 2. The vascular sensor of claim 1, wherein: the first electrocardiogram sensing electrode is electrically connected to the at least one fixation element; andthe second electrocardiogram sensing electrode is coupled to the sensor housing. 3. The vascular sensor of claim 2, further comprising: a tether wire extending proximally from the at least one fixation element; anda conductive ring disposed at a proximal end of the tether wire, wherein the conductive ring is coupled to the power supply housing, thereby electrically coupling the first electrocardiogram sensing electrode to the at least one fixation element. 4. The vascular sensor of claim 1, wherein the at least one fixation element includes a first fixation element and a second fixation element, the first fixation element electrically coupled to the first electrocardiogram sensing electrode, the second fixation element electrically coupled to the second electrocardiogram sensing electrode. 5. The vascular sensor of claim 1, further including a communication module configured t... | Cardiac Pacemakers, Inc. | A61B005/02 A61B005/0215 A61B005/00 A61B005/042
US09289154 | B2 | 2016-03-22 | Techniques for temperature measurement and corrections in long-term magnetic resonance thermometry | 2009-08-19 | Techniques for temperature measurement and correction in long-term MR thermometry utilize a known temperature distribution in an MR imaging area as a baseline for absolute temperature measurement. Phase shifts that arise from magnetic field drifts are detected in one or more portions of the MR imaging area, facilitating correction of temperature measurements in an area of interest. | 1. A method of performing proton resonance frequency (PRF) based magnetic resonance (MR) temperature measurement, the method comprising the steps of: using an MRI unit, acquiring a first phase image of an MR imaging area having a known temperature distribution, the MR imaging area comprising an area of interest;using the MRI unit, acquiring a second phase image of the MR imaging area subsequent to the acquisition of the first phase image;without modifying the first image, determining, from differences between the second phase image and the first phase image, one or more first corrections in one or more portions of the MR imaging area that have experienced a known or clinically insignificant change in temperature since the acquisition of the first phase image;determining one or more second corrections in at least the area of interest by extrapolating the one or more first corrections;assigning a first set of one or more absolute temperatures to the area of interest based at least in part on (i) the known temperature distribution and (ii) the one or more second corrections,detecting a movement and/or deformation of the area of interest;determining a new temperature distribution of at least the area of interest by registering the known temperature distribution to the new temperature distribution based at least in part on the detected movement and/or deformation;acquiring a third phase image of the MR imaging area subsequent to the detected movement and/or deformation;acquiring a fourth phase image of the MR imaging area subsequent to the acquisition of the third phase image; andassigning a second set of one or more absolute temperatures to the area of interest based at least in part on (i) the new temperature distribution and (ii) a difference between the fourth phase image and the third phase image. 2. The method of claim 1, further comprising: establishing the known temperature distribution in the MR imaging area. 3. The method of claim 2, further comprising: establishing a uniform temperature in the MR imaging area. 4. The method of claim 1, wherein the MR imaging area is centered on the area of interest. 5. The method of claim 1, wherein (i) the area of interest comprises a treatment area of a patient undergoing an MR-guided focused ultrasound treatment, and (ii) the MR imaging area comprises at least a portion of the patient's body. 6. The method of claim 5, wherein the treatment area comprises the patient's prostate area. 7. The method of claim 6, further comprising: establishing a uniform temperature in the MR imaging area by circulating heated water in a balloon pressed against the patient's rectal wall. 8. The method of claim 7, wherein the uniform temperature is substantially the same as the patient's body temperature. 9. The method of claim 1, further comprising: determining the second correction(s) with a two-dimensional (2-D) polynomial fit based on the first correction(s). 10. The method of claim 9, wherein the two-dimensional... | InSightec Ltd. | A61B005/055 G01R033/48 A61N007/00 G01R033/565
US09265435 | B2 | 2016-02-23 | Multi-electrode sensing patch for long-term physiological monitoring with swappable electronics, radio and battery, and methods of use | 2008-10-21 | Provided herein is an electrode device comprising an interface comprising at least one magnetic contact. The interface can be adaptable to be in communication with an electronic member, wherein the magnetic contact on the interface is adaptable to align and position the electronic member with respect to the interface. The interface can be further adaptable to remain affixed to a patient while an electronic member is removed and/or inserted from the interface. Further provided herein are methods of using the electrode device and kits. | 1. An electrode device comprising: a pouch formed of a flexible substrate having a first part and a second part that are inseparably joined together along a periphery, the pouch having an interior and comprising at least one electrode contact area on a surface of the first part and at least one contact patch on an interior surface of the second part, the at least one electrode contact area connected to the at least one contact patch via a conductive trace traversing a distance along the flexible substrate over both the first part of the pouch and the second part of the pouch, the at least one contact patch overlying at least one magnetic contact, the pouch adaptable to be in communication with an electronic member through the at least one contact patch and accept the electronic member within its interior between the first part and the second part of the flexible substrate without disturbing the pouch, the at least one magnetic contact adaptable to slidably align and position the electronic member within the interior of the pouch. 2. The device of claim 1 wherein the pouch is further adaptable to remain affixed to a patient while the electronic member is removed and/or inserted from the interior of the pouch. 3. The device of claim 1 wherein the electronic member comprises a corresponding magnetic contact and the device further comprises the electronic member within the interior of the pouch such that the at least one magnetic contact of the pouch aligns with the corresponding magnetic contact of the electronic member. 4. The device of claim 1 wherein the at least one magnetic contact of the pouch is a ferromagnetic material and the corresponding magnetic contact of the electronic member is a magnet. 5. The device of claim 3 wherein the at least one magnetic contact of the pouch adaptably aligns with the corresponding magnetic contact of the electronic member to cause the contact patch of the pouch overlying the at least one magnetic contact to directly contact a contact located on the electronic member, such that the contact patch of the pouch and the contact located on the electronic member are sandwiched between the at least one magnetic contact of the pouch and the corresponding magnetic contact of the electronic member. 6. The device of claim 1 wherein the device is adaptable to be used with a replaceable electronic member. 7. The device of claim 1 further comprising a ferromagnetic metal adaptable to facilitate a magnetic connection between the pouch and the electronic member. 8. The device of claim 7 wherein the pouch is adaptable to maintain the connection between the electronic member and the pouch. 9. The device of claim 1 further comprising an additional support structure adaptable to maintain the position of the electronics component. 10. The device of claim 1 wherein the at least one magnetic contact is adaptable to provide a low engagement force between the pouch and the electronic member. 11. An electrode device comprising... | HMICRO, INC. | A61B005/04 A61B005/0408 A61B005/0478 A61B005/0492
US20110004086 | A1 | 2011-01-06 | METHOD OF ANALYZING BIOLOGICAL COMPONENT, BIOLOGICAL COMPONENT ANALYZER, REACTION CARTRIDGE OF BIOLOGICAL COMPONENT ANALYZER AND EXTRACTION CARTRIDGE OF BIOLOGICAL COMPONENT ANALYZER | 2010-09-17 | This method of analyzing a biological component includes steps of arranging an extraction medium holding device on the skin of a subject, extracting an extract from the subject, attaching a reaction cartridge configured to be detachable on an extraction medium holding part of the extraction medium holding device and to be capable of transferring an extraction medium to a reaction part to the extraction medium holding part of the extraction medium holding device arranged on the skin, and transferring the extraction medium held in the extraction medium holding part to the reaction part. | 1. A method of analyzing a biological component comprising steps of: arranging an extraction medium holding device comprising an extraction medium holding part configured to be capable of holding an extraction medium for holding an extract extracted from a subject on the skin of the subject;extracting said extract from the subject to said extraction medium held in said extraction medium holding part;attaching a reaction cartridge comprising a reaction part allowing said extract to cause a prescribed reaction, configured to be detachable on said extraction medium holding part of said extraction medium holding device and to be capable of transferring said extraction medium held in said extraction medium holding part to said reaction part to said extraction medium holding part of said extraction medium holding device arranged on said skin;transferring said extraction medium held in said extraction medium holding part to said reaction part;detecting said extract from said prescribed reaction caused in said reaction part; andanalyzing a result of detecting said extract. 2. The method of analyzing a biological component according to claim 1, wherein said step of transferring said extraction medium to said reaction part includes a step of transferring said extraction medium through a transfer passage provided in said reaction cartridge for transferring said extraction medium to said reaction part. 3. The method of analyzing a biological component according to claim 2, wherein said step of transferring said extraction medium to said reaction part includes a step of transferring said extraction medium held in said extraction medium holding part to said reaction part through said transfer passage by capillarity. 4. A biological component analyzer, comprising: an extraction medium holding device having an extraction medium holding part configured to be capable of holding an extraction medium for holding an extract extracted from a subject;a reaction cartridge including a reaction part allowing said extract to cause a prescribed reaction and a transfer passage for transferring said extraction medium held in said extraction medium holding part to said reaction part, configured to be detachable on said extraction medium holding part of said extraction medium holding device;a detecting part detecting said extract from said prescribed reaction caused in said reaction part; andan analysis part analyzing a result of detecting said extract. 5. The biological component analyzer according to claim 4, wherein said transfer passage is configured to transfer said extraction medium held in said extraction medium holding part to said reaction part by capillarity. 6. The biological component analyzer according to claim 4, wherein said transfer passage is configured to come into contact with said extraction medium held in said extraction medium holding part when said reaction cartridge is attached to said extraction medium holding part of said extraction medium holding... | - | A61B005/00
US20110046495 | A1 | 2011-02-24 | DEVICE FOR MEASURING THE SIZE OF AN INTRACARDIAC OPENING | 2010-08-11 | A device (1) for measuring the diameter of an intracardiac opening (2), having a hollow body (3) made from a metal-wire mesh is provided that has twisted individual bars that form the jacket of the discontinuous wall (5) of the hollow body (3) and form longitudinal stabilization between two elastically deformable metal-wire meshes (6) forming the end sides of the hollow body (3) and are deformable in a position of use by an edge (17) of the opening (2) in a contact region into the interior of the hollow body (3). Thus, with the device (1), the size of the opening (2) can be determined in that, in an X-ray or ultrasound image, the deformation of the individual bars (4) is comparable through the opening (2) with markings (19, 20) arranged on the hollow body (3) and/or on the device (1) adjacent to the hollow body (3). | 1. Device (1) for measuring the size of an intracardiac opening (2), comprising a hollow body (3) that has an elastic wall (5) and is insertable into the opening (2) and can expand perpendicular to an insertion direction and having at least two transparent markings (19, 20) arranged at a predetermined distance to each other and the distance can be compared with a molding of the hollow body (3) caused by an edge (17) of the opening (2) to be measured, the hollow body (3) is permeable to liquid and is constructed as a metal cage, the hollow body (3) is discontinuous in a middle region between end sides (7, 8) thereof and the discontinuity is spanned by individual bars (4) that are each formed by at least two twisted wires (23). 2. The device according to claim 1, wherein the individual bars (4) are arranged parallel to each other and to a longitudinal middle axis of the hollow body (3). 3. The device (1) according to claim 1, wherein the hollow body (3) is made from at least one of an elastic material or a memory metal. 4. The device (1) according to claim 1, wherein the hollow body (3) has, in a measurement region, several of the individual bars (4) describe a cylindrical lateral surface and are sufficiently stable for stabilizing the hollow body (3) and are soft and flexible in the contact region (18) with the edge (17) of the opening (2). 5. The device (1) according to claim 1, wherein, in a position of use, the individual bars (4) are adapted to extend through the opening (2) and contact the edge (17) of the opening (2). 6. The device (1) according to claim 1, wherein the hollow body (3) has wire sections (23) that are twisted with each other and are in touching contact, in a position of use, with the edge (17) of the opening (2). 7. The device according to claim 1, wherein the hollow body (3) comprises a metal-wire mesh (6) that can be deformed elastically in a radial direction against a restoring force from a maximum, expanded position on each of its end sides (7, 8) arranged on both sides of the opening (2) in a position of use. 8. The device according to claim 1, wherein ends (7, 8) of the hollow body (3) are connected by a connection element (9) deformable elastically against a tensile stress and that extends centrally through the hollow body (3). 9. The device according to claim 1, wherein an insertion catheter is provided with a sleeve-shaped receptacle (12) for the hollow body (3) in a radially folded-together or compressed form. 10. The device according to claim 9, wherein the hollow body (3) is mounted on a push-and-pull element (11) and can be moved with the push-and-pull element (11) into the receptacle (12) and out from the receptacle (12). 11. The device according to claim 10, wherein the receptacle (12) is constructed on a shaft (14) of the insertion catheter holding the hollow body (3) in a position of use. 12. The device according to claim 1, wherein the markings (19, 20) are arranged adjacent to the hollow body (3) ... | - | A61B005/02
US07887549 | B2 | 2011-02-15 | Catheter system | 2010-02-03 | A robotic catheter system includes a housing and a drive assembly. The drive assembly is supported by the housing and configured to impart movement to a catheter device and includes a first channel configured to receive the catheter device. The robotic catheter system includes at least one magnet positioned adjacent to the first channel to allow the magnet to interact with the catheter device to assist in the retention of the catheter device within the first channel. | 1. A robotic catheter system, the system comprising: a housing;a drive assembly supported by the housing and configured to impart movement to a catheter device, the drive assembly including: an engagement structure having a pair of opposing engagement surfaces configured to engage opposing surfaces of the catheter device;a top deck coupled to the housing and positioned above the engagement structure;a first channel formed in the top deck configured to receive the catheter device, wherein the first channel comprises a bottom wall and a channel opening opposite the bottom wall, wherein the engagement structure engages the catheter device within the first channel; andat least one magnet positioned proximate the bottom wall of the first channel, wherein the magnet interacts with the catheter device to assist in loading the catheter device into the channel through the channel opening by attracting the catheter device in a direction from the channel opening toward the bottom wall and to assist in the retention of the catheter device within the first channel. 2. The robotic catheter system of claim 1, further comprising: a first cover moveable from a closed position overlying the first channel to an opened position exposing the first channel; anda restraint structure that acts to restrain movement of the catheter device within the first channel when the cover is moved to the closed position. 3. The robotic catheter system of claim 2, wherein the restraint structure comprises a tab extending from a lower surface of the first cover, and further wherein at least a portion of the tab is positioned within at least a portion of the first channel when the first cover is in the closed position. 4. The robotic catheter system of claim 3, wherein the force exerted by the at least one magnet on the catheter device is sufficient to assist in the retention of the catheter device within the first channel while permitting the catheter device to be moved by the drive assembly. 5. The robotic catheter system of claim 4, wherein the tab assists in holding the catheter device within the first channel after the cover is moved to the closed position. 6. The robotic catheter system of claim 5, further comprising a plurality of upwardly extending channel tabs positioned along the first channel to aid in alignment of the catheter device with the first channel during loading of the catheter device, wherein the tab is positioned between two of the plurality of channel tabs when the cover is in the closed position. 7. The robotic catheter system of claim 3, further comprising a plurality of magnets positioned beneath the first channel, wherein the plurality of magnets are spaced at even intervals along the length of the first channel. 8. The robotic catheter system of claim 3, wherein the catheter device is a guide wire, wherein the drive assembly includes a guide wire axial drive mechanism to releasably engage and drive the guide wire along the longitudinal axis of the gu... | Corindus Inc. | A61F011/00 A61B005/00
US07881763 | B2 | 2011-02-01 | Optimized sensor geometry for an implantable glucose sensor | 2006-05-02 | An implantable sensor for use in measuring a concentration of an analyte such as glucose in a bodily fluid, including a body with a sensing region adapted for transport of analytes between the sensor and the bodily fluid, wherein the sensing region is located on a curved portion of the body such that when a foreign body capsule forms around the sensor, a contractile force is exerted by the foreign body capsule toward the sensing region. The body is partially or entirely curved, partially or entirely covered with an anchoring material for supporting tissue ingrowth, and designed for subcutaneous tissue implantation. The geometric design, including curvature, shape, and other factors minimize chronic inflammatory response at the sensing region and contribute to improved performance of the sensor in vivo. | 1. An implantable sensor for use in measuring a concentration of an analyte in a bodily fluid, the sensor comprising: a body comprising an electrochemical sensing region adapted for transport of an analyte between the electrochemical sensing region and a bodily fluid, wherein the electrochemical sensing region comprises one or more electrodes, wherein the body comprises a first major surface and a second major surface, wherein the one or more electrodes are situated approximately at a center of the first major surface, wherein the electrochemical sensing region is located on a curved portion of the first major surface configured such that when a foreign body capsule forms around the sensor, a contractile force is exerted by the foreign body capsule toward the sensing region, and wherein the sensor is a subcutaneous sensor configured for implantation in a soft tissue of a host. 2. The sensor of claim 1, wherein the electrochemical sensing region comprises a plurality of electrodes. 3. The sensor of claim 1, wherein the second major surface is flat. 4. The sensor of claim 1, wherein the second major surface has a curvature. 5. The sensor of claim 1, wherein the first major surface, when viewed from a direction perpendicular to a center of the first major surface, has a substantially oval profile. 6. The sensor of claim 1, wherein the first major surface, when viewed from a direction perpendicular to a center of the first major surface, has a substantially circular profile. 7. The sensor of claim 1, wherein at least one of the first major surface or the second major surface has a non-electrode surface area comprising an anchoring material for supporting tissue ingrowth. 8. The sensor of claim 1, further comprising a mechanical anchoring mechanism formed on the body. 9. The sensor of claim 1, wherein the curved portion of the first major surface has a plurality of radii of curvature. 10. The sensor of claim 1, wherein the curved portion of the first major surface has a radius of curvature of from about 0.5 mm to about 10 cm. 11. The sensor of claim 1, wherein the first major surface and the second major surface together account for at least about 40% of the surface area of the body. 12. The sensor of claim 1, wherein edges of the first major surface are rounded and transition smoothly away from the first major surface. 13. The sensor of claim 1, wherein the body defines a surface area, and wherein from 10% to 100% of the surface area is convexly curved. 14. The sensor of claim 1, wherein the body comprises plastic. 15. The sensor of claim 14, wherein the plastic is selected from the group consisting of thermoplastic and thermoset. 16. The sensor of claim 15, wherein the thermoset is selected from the group consisting of epoxy, silicone, and polyurethane. 17. The sensor of claim 1, wherein the body comprises a material selected from the group consisting of metal, ceramic, and glass. 18. The sensor of claim 1, further comprising a poro... | DexCom, Inc. | A61B005/1473
US07892188 | B2 | 2011-02-22 | Method and apparatus for characterization of clot formation | 2004-10-22 | Methods, apparatus and systems for characterizing changes in at least one physical property of soft tissue. A series of acoustic pulses is generated and directed into the soft tissue such that at least one of the pulses is of sufficiently high intensity to induce physical displacement of the tissue. Waves reflected off the tissue, or a flexible member that moves with the tissue, are received and measured to estimate at least one characteristic of the physical displacement induced thereby. Repetition of the generating, receiving and estimating steps provides characterization of the at least one physical property over time. Methods, apparatus and systems for characterizing at least one physical property of blood, by generating a series of acoustic pulses and directing the series of pulses into the blood such that at least one of the pulses is of sufficiently high intensity to induce physical displacement of the blood. Acoustic pulses and/or optical waves reflected from the blood, or a flexible member in contact with the blood that moves with the blood, are received and measured to estimate at least one characteristic of the physical displacement induced thereby. | 1. A method of characterizing at least one physical property of a biological material, said method comprising: generating a series of acoustic pulses and directing said series of pulses into the material such that at least one of the pulses is of sufficiently high intensity to induce physical displacement of the material, and wherein each of said acoustic pulses is of finite duration and of brief duration relative to a duration of said generating and said directing;controlling a temperature of the material;measuring a displacement, either directly or indirectly, of the material resulting from said induced physical displacement thereof, wherein the temperature of the material is controlled over a duration of said measuring a displacement; andestimating at least one characteristic of the physical displacement based on said measuring. 2. The method of claim 1, wherein said measuring step comprises receiving at least two of said acoustic pulses reflected from the material and wherein said estimating comprises establishing a baseline based on at least one of said acoustic pulses received, and calculating an estimate of said at least one characteristic based on comparison of at least another of said acoustic pulses received to said baseline. 3. The method of claim 1, wherein said measuring step comprises receiving optical reflections from the material as the material is being physically displaced and said estimating is based on the optical reflections received. 4. The method of claim 1, wherein said measuring step comprises receiving at least two of said acoustic pulses reflected from a container or membrane in contact with the material that moves with the material. 5. The method of claim 1, wherein the material comprises soft tissue. 6. The method of claim 1, wherein the material comprises blood. 7. The method of claim 6, wherein said controlling the temperature of the material comprises placing the blood in contact with a second material, said second material having a controlled temperature. 8. The method of claim 7, wherein said controlling comprises controlling the temperature of the blood without placing the blood or material in a bath. 9. The method of claim 1, wherein the material comprises a first material, and wherein said controlling the temperature of the material comprises placing a second material, the temperature of which is controlled, in contact with said first material. 10. The method of claim 9, wherein said controlling the temperature of the material comprises controlling the temperature of the second material to maintain the first material at a controlled temperature. 11. The method of claim 9, wherein said controlling the temperature comprises controlling the temperature of the first material without placing the first material or second material in a bath. 12. The method of claim 1, further comprising timing said directing of said series of acoustic pulses into the material, and timing receipt of echoes of said pulses us... | Hemosonics, LLC | A61B005/103 A61B005/117 A61B005/00 A61B008/00
US07869851 | B2 | 2011-01-11 | System and method for determining insulin bolus quantities | 2004-12-23 | A method of computing an insulin bolus quantity when a measured glucose value in a current time interval exceeds a target glucose value for the current time interval and when a time duration of glucose lowering action of the bolus to be administered spans the current and a number of the next adjacent time intervals, may comprise determining a corresponding number of percentages of insulin action of the recommended bolus quantity that will each be used lowering the glucose level during a respective one of the number of time intervals, and computing the recommended bolus quantity as a function of the measured glucose value, the target blood glucose level for a last one of the number of next adjacent time intervals, the insulin sensitivities for the current and each of the number of next adjacent time intervals, and the corresponding number of percentages. | 1. A method of computing a recommended insulin bolus quantity when a measured glucose value in a current time interval of a number of adjacent time intervals, each of the number of adjacent time intervals defining an associated target glucose value and an insulin sensitivity value, exceeds the target glucose value for the current time interval and when a time duration of glucose lowering action of the recommended insulin bolus quantity to be administered extends from the current time interval into a number of next adjacent time intervals, the method comprising the steps of: determining with a processor a plurality of percentages each corresponding to a percentage of insulin action of the recommended bolus quantity that will be used lowering the glucose level during a respective one of the current time interval and each of the number of next adjacent time intervals, andcomputing with the processor the recommended bolus quantity as a function of the measured glucose value, the target blood glucose level for a last one of the number of next adjacent time intervals, the insulin sensitivities for the current and each of the number of next adjacent time intervals, and the plurality of percentages. 2. The method of claim 1 wherein the number of next adjacent time intervals is one, and wherein the step of determining with a processor a plurality of percentages includes:determining with the processor a first percentage corresponding to a percentage of insulin action of the recommended bolus quantity that will be used lowering the glucose level during the current time interval,determining with the processor a second percentage corresponding to a percentage of insulin action of the recommended bolus quantity that will be used lowering the glucose level during the next adjacent time interval,and wherein the computing step includes computing the recommended bolus quantity as a function of the measured glucose value, the target blood glucose level for the next adjacent time interval, the insulin sensitivities for the current and next adjacent time intervals, and the first and second percentages. 3. The method of claim 2 wherein the step of determining a first percentage includes determining the first percentage as a function of at least a current time in the current time interval at which the measured bolus value was measured, a beginning time of the next adjacent time interval and the time duration of glucose lowering action of the subsequently administered recommended insulin bolus quantity. 4. The method of claim 3 further including the step of computing with the processor a current interval bolus activity time as a difference between the beginning time of the next adjacent time interval and the current time, and wherein the step of determining a first percentage includes determining the first percentage as a function of the current interval bolus activity time and the time duration of glucose lowering action of the subsequently administered recommended ... | Roche Diagnostics Operations, Inc. Roche Diagnostics International AG | A61B005/00
US07891106 | B2 | 2011-02-22 | Growth chart device | 2010-01-18 | A growth chart device features presenting a wall facing side and an opposing chart side with a measuring scale extending along the chart side along a lengthwise dimension of the chart body. A marking area is defined on the chart side of the chart body proximate the measuring scale to receive markings at different positions along the measuring scale at different times to track a height of an individual. At least one door is pivotally mounted on the chart body and movable between open and closed positions in which at least a portion of the chart side of the chart body is respectively uncovered and covered to give selective control over access to said portion. | 1. A growth chart device comprising: a chart body presenting a wall side intended to face a wall or other vertical structure to which the growth chart is to be mounted and a chart side opposite said wall side;a measuring scale defined on the chart side of the chart body to extend along a lengthwise dimension of the chart body to be vertically oriented when mounted to the wall or other vertical structure;a marking area defined on the chart side of the chart body proximate the measuring scale to receive markings at different positions along the measuring scale at different times to track a height of an individual; andat least one door pivotally mounted on the chart body and movable between open and closed positions in which at least a portion of the chart side of the chart body is respectively uncovered and covered to give selective control over access to said portion. 2. The growth chart according to claim 1 wherein said portion of the chart side of the chart body selectively enclosable by the at least one door includes at least part of the measuring scale. 3. The growth chart according to claim 1 comprising a sliding indicator slidably mounted to the chart body for movement along the lengthwise dimension thereof to a position indicating the height of the individual along the measuring scale. 4. The growth chart according to claim 3 wherein the at least one door comprises a pair of doors mounted on the chart body to close over respective portions of the chart side of the chart body on opposite sides of the sliding indicator. 5. The growth chart according to claim 3 comprising a rail structure along which the sliding indicator is slidable and which projects from the chart side of the chart body past the sliding indicator to form a stop against which each door abuts in the closed position. 6. The growth chart according to claim 5 wherein the rail structure comprises a pair of opposing rails each having a track formed therein into which features of the sliding indicator project from between the opposing rails. 7. The growth chart according to claim 6 wherein at least one of the tracks comprises a through-slot in a respective one rail of the opposing rails. 8. The growth chart according to claim 3 wherein a projecting portion of the sliding indicator projects along the chart side of the chart body toward the measuring scale. 9. The growth chart according to claim 3 wherein a measuring arm of the sliding indicator is pivotable between storage and measuring positions extending along and projecting away from the chart side of the chart body respectively. 10. The growth chart according to claim 3 wherein the measuring scale and the marking area are defined on a same side of the sliding indicator. 11. The growth chart according to claim 3 wherein the sliding indicator is centrally located along a widthwise dimension of the chart body. 12. The growth chart device according to claim 1 comprising a display area defined on the chart side of the char... | - | A61B005/103 G01B003/04
US08308650 | B2 | 2012-11-13 | System and method for generating baseline data for automated management of cardiovascular pressure | 2010-11-03 | A system and method for generating baseline data for automated management of cardiovascular pressure is disclosed. Collected device measures are accumulated to record raw physiometry for a patient, wherein the patient is regularly monitored by an implantable medical device, beginning with an initial observation period. Derived device measures are generated to provide derivative physiometry determined at least in part from the collected device measures. A patient status indicator is determined by analyzing the collected and derived device measures to diagnose a pathophysiology indicative of an absence, onset, progression, regression, and status quo in cardiovascular pressure, wherein the collected and derived device measures and the patient status indicator originating from the initial observation period include baseline data. | 1. A programmer for generating baseline data for automated management of cardiovascular pressure, comprising: a memory to store device measures received from an implantable medical device, comprising: collected device measures to record raw physiometry for a patient regularly monitored by the implantable medical device, beginning with an initial observation period; andderived device measures to provide derivative physiometry determined at least in part from the collected device measures; anda processor to determine a wellness indicator through analysis of the collected and derived device measures to diagnose a patient status indicative of at least one of a change and a status quo in cardiovascular pressure,wherein the collected and derived device measures and the wellness indicator originating from the initial observation period comprise baseline data. 2. A programmer according to claim 1, wherein the analysis comprises one or more of: a monitor to accumulate collected device measures and to derive derived device measures for use in determining the wellness indicator relating to cardiovascular pressure;a patient well being analyzer to monitor patient well being in conjunction with at least one of potassium level, sodium level, hormonal level, and blood pressure; anda programming interface to apply programming data to effect therapy modification selected from the group comprising device parameters, drug infusion parameters, and neural stimulation parameters. 3. A programmer according to claim 1, further comprising: an evaluator to process physiometric data conjunction with monitoring of the patient, wherein the physiometric data is selected from the group comprising: analytic data recorded, collected, and stored for cardiovascular pressure evaluation;alert data generated in response to significant events, wherein the alert data is selected from the group comprising absolute measures, long term average measures, and short term average measures; andaction data executed in response to indications selected from the group comprising the significant events, ECG signals, activity level, temperature, pressure, and oxygen saturation. 4. A programmer according to claim 1, further comprising: physiometric data comprising one or more of the collected and derived device measures sensed through one or more of cardiovascular pressure, potassium level, sodium level, hormonal level, central nervous system injury, and central nervous system blood flow. 5. A programmer according to claim 1, further comprising: programming parameters, comprising at least one of: monitoring parameters to monitor the patient while asleep; andexclusion parameters to exclude physiometric data comprising one or more of the collected and derived device measures when the patient exhibits an increased respiratory rate due to exercise or activity level. 6. A programmer according to claim 1, further comprising: programmed trigger conditions with corresponding actions, wherein the actions ... | Cardiac Pacemakers, Inc. | A61B005/02
US07883688 | B2 | 2011-02-08 | Polycationic polyrotaxanes capable of forming complexes with nucleic acids | 2006-02-03 | A polycation capable of forming a complex with a nucleic acid for carriage thereof. The polycation comprising at least one cyclic compound having a cavity. A polymer backbone is threaded into the cavity of the cyclic compound. A pair of bulky moieties cap the terminals of the polymer backbone to inhibit dethreading of the cyclic compound from the polymer backbone. In use, the net positive charge of said polycation enables a complex to form with nucleic acid. | 1. A supramolecular polycation assembly capable of forming a complex with a nucleic acid for carriage thereof, the supramolecular polycation assembly comprising: at least one amino-substituted cyclodextrin having a cavity extending therethrough, wherein the amino-substituted cyclodextrin comprises one or a plurality of aminated D-glucose units of the formula:wherein: p is an integer from 0 to 8; andT is optional and is an alkyl selected from the group consisting of methyl, ethyl and propyl;a co-polymer backbone threaded into the cavity of the cyclodextrin; anda pair of bulky moieties capping the terminals of the co-polymer backbone to inhibit dethreading of the cyclodextrin from the co-polymer backbone,wherein the net positive charge of the supramolecular polycation assembly enables a complex to form with the nucleic acid. 2. The supramolecular polycation assembly of claim 1, wherein the bulky moieties are biocleavable. 3. The polycation of claim 1, wherein the co-polymer backbone has a shape selected from the group consisting of linear polymers, branched polymers, star polymers and combinations thereof. 4. The supramolecular polycation assembly of claim 1, wherein a polymer of the co-polymer backbone is selected from the group consisting of polyethylene glycol, polypropylene glycol, polybutylene glycol, polypentylene glycol, polyhexylene glycol, polymethyl vinyl ether, polyethyl vinyl ether, polyisoprene, polyisobutylene, and polybutadiene. 5. The supramolecular polycation assembly of claim 1, wherein the co-polymer backbone has a molecular weight in the range selected from the group consisting of 200 to 50000, 200 to 10000, 200 to 5000, and 200 to 2000. 6. The supramolecular polycation assembly of claim 1, wherein a plurality of cyclodextrin compounds are threaded on the co-polymer backbone. 7. The supramolecular polycation assembly of claim 1, wherein the bulky moieties are either one of a group having at least one benzene ring or a group having at least one tertiary butyl. 8. The supramolecular polycation assembly of claim 1, wherein the co-polymer backbone is a straight chain co-polymer and the bulky moieties are connected to the straight chain co-polymer backbone by a biocleavable linker. 9. The supramolecular polycation assembly of claim 8, wherein the biocleavable linker is selected from the group consisting of amides, amines, alkyls having from 1 to 8 carbon atoms, esters having from 1 to 8 carbon atoms, anhydrides and combinations thereof. 10. A method for introducing an exogenous nucleic acid molecule into a target cell comprising the steps of: (a) forming a complex between a polycation as claimed in claim 1 and a nucleic acid; and (b) introducing the formed complex into a target cell in vitro. 11. The supramolecular polycation assembly of claim 1, wherein the co-polymer backbone comprises polyethylene glycol and polypropylene glycol. 12. The supramolecular polycation assembly of claim 1, wherein p is 1 to 4. 13. The supram... | Agency for Science, Technology and Research National University of Singapore | A61B005/055 C12N005/00 C12N005/02 A61K048/00
US07869865 | B2 | 2011-01-11 | Current-based position sensing | 2005-01-07 | A method for position sensing includes inserting a probe comprising at least one electrode into a body of a subject, and passing electrical currents through the body between the at least one electrode and a plurality of locations on a surface of the body. Respective characteristics of the currents passing through the plurality of the locations are measured in order to determine position coordinates of the probe. | 1. A method for position sensing, comprising: inserting a probe comprising at least one electrode into a body of a subject;passing electrical currents through the body from the at least one electrode to a plurality of locations on a surface of the body and maintaining a constant voltage between each of the plurality of the locations and the at least one electrode;measuring respective characteristics of the electrical currents passing through the plurality of the locations, wherein the respective characteristics are indicative of respective electrical impedances between the plurality of the locations and the at least one electrode by measuring the electrical currents at the constant voltage; anddetermining position coordinates of the probe responsively to the measured characteristics. 2. The method according to claim 1, wherein the at least one electrode comprises a plurality of electrodes, and wherein passing the electrical currents comprises passing each of the electrical currents between one of the plurality of the electrodes and one of the plurality of the locations on the surface of the body. 3. The method according to claim 1, wherein passing the electrical currents comprises affixing conductive patches to the body at the plurality of locations, and passing the electrical currents through the conductive patches. 4. The method according to claim 1, wherein the plurality of the locations comprises at least three locations. 5. The method according to claim 4, wherein the plurality of the locations comprises exactly three locations. 6. The method according to claim 1, wherein inserting the probe comprises performing a diagnostic treatment on the subject using the probe. 7. The method according to claim 6, wherein the probe comprises a catheter, and wherein performing the diagnostic treatment comprises mapping a heart of the subject. 8. The method according to claim 7, wherein mapping the heart comprises sensing electrical potentials in tissue of the heart using the at least one electrode. 9. The method according to claim 1, wherein inserting the probe comprises performing a therapeutic treatment on the subject using the probe. 10. The method according to claim 9, wherein the probe comprises a catheter, and wherein performing the therapeutic treatment comprises ablating heart tissue using the catheter. 11. Apparatus for position sensing, comprising: a probe, comprising at least one probe electrode, which is adapted to be inserted into a body of a subject;a plurality of body surface electrodes, which are adapted to be fixed to a surface of the body at respective locations; anda controller, which is adapted to be coupled to the probe and to the body surface electrodes so as to pass electrical currents through the body from the at least one probe electrode to the plurality of body surface electrodes, and to determine position coordinates of the probe by measuring respective characteristics of the electrical currents passing through the plu... | Biosense Webster, Inc. | A61B005/04 A61B005/05 A61B008/00
US08858452 | B2 | 2014-10-14 | Check valve structure, diaphragm pump, and sphygmomanometer | 2009-04-10 | A check valve structure has a first space and a second space. The check valve structure is configured to allow a flow of a fluid from the first space to the second space and to inhibit flow of the fluid from the second space to the first space. The check valve structure has a partition wall arranged between the first space and the second space and provided with a communication hole providing communication between the first space and the second space, an elastic film body for covering the side of the second space of the communication hole so as to prevent a reverse flow of the fluid, an elastic member having a wall portion surrounding the communication hole for holding the elastic film body, and a nipping member for nipping the elastic member with the partition wall. | 1. A check valve structure, comprising: a first check valve;a first space and a second space, the first check valve being configured to allow a flow of a fluid from the first space to the second space and inhibit a flow of the fluid from the second space to the first space;a second check valve;another first space and another second space, the second check valve being configured to allow a flow of the fluid from the another first space to another second space and inhibit a flow of the fluid from the another second space to the another first space;a partition wall arranged between the first space and the second space and between the another first space and the another second space and provided with a communication hole providing communication between the first space and the second space and another communication hole providing communication between the another first space and the another second space;the first check valve comprising a first elastic film body that covers and directly contacts a side of the second space of the communication hole so as to prevent a reverse flow of the fluid when the first elastic film body is in a first closed position;the second check valve comprising a second elastic film body that covers and directly contacts a side of the another second space of the another communication hole so as to prevent a reverse flow of the fluid when the second elastic film body is in a second closed position;an elastic member having a first elastic wall portion surrounding the communication hole and a second elastic wall portion surrounding the another communication hole for holding the first elastic film body and the second elastic film body; anda nipping member for nipping the elastic member with the partition wall, wherein a surface of the elastic member is closely attached to a surface on the side of the second space of the partition wall,wherein a back surface on an opposite side of the surface is closely attached to the nipping member,wherein the nipping member has a first protruding portion and a second protruding portion that protrude toward the partition wall,wherein the first protruding portion is fitted into the second space surrounded by the first elastic wall portion so as to be in surface contact with the entire circumference of the first elastic wall portion to restrict positional displacement of the first elastic wall portion, so that deformation due to distortion of the first elastic film body and the first elastic wall portion is suppressed,wherein the first protruding portion comprises a first ventilation passage that communicates with another first ventilation passage formed inside the nipping member,wherein the second protruding portion is fitted into the another second space surrounded by the second elastic wall portion so as to be in surface contact with the entire circumference of the second elastic wall portion to restrict positional displacement of the second elastic wall portion, so that deformation due to disto... | OMRON HEALTHCARE Co., Ltd. | A61B005/0235 F16K015/00 F04B043/00 A61B005/021 F16K015/16 F16K015/14 F04B053/10
US09131896 | B2 | 2015-09-15 | In situ offset compensation for pressure sensors | 2010-08-10 | A pressure sensor having a substrate and a first, deformable membrane, partially supported by the substrate, which generates a first sensor reading when deformed by pressure. A second membrane is contiguous to the first membrane. When the second membrane is energized, it deforms the first membrane to alter the first sensor reading. | 1. A pressure sensor comprising: a substrate;a first, deformable membrane, partially supported by the substrate, which generates a first sensor reading when deformed by pressure; anda second deformable membrane contiguous to and touching the first membrane which, when energized, deforms and deforms the first membrane to alter the first sensor reading;wherein the first membrane carries a first electrode which forms capacitance with another electrode spaced from the first electrode to generate the first sensor reading. 2. The pressure sensor of claim 1 wherein the first membrane carries at least one strain gauge to generate the first sensor reading when deformed. 3. The pressure sensor of claim 2 wherein the strain gauge includes a piezoresistive element. 4. The pressure sensor of claim 1 wherein the second membrane includes piezoelectric material. 5. The pressure sensor of claim 4 wherein the piezoelectric material is deposited on the first membrane. 6. The pressure sensor of claim 1 wherein the second membrane is energizable by applying a voltage across it. 7. The pressure sensor of claim 1 wherein the substrate and first and second membranes are disposed within a housing that is implantable in a patient. 8. A system for measuring pressure within a patient, comprising: a housing implantable in a patient and containing a substrate;a first circuit including a first, deformable membrane, partially supported by the substrate, which generates a first sensor reading when deformed by pressure; anda second circuit including an energy source and a second deformable membrane, contiguous to and touching the first membrane which, when energized by the energy source, deforms and deforms the first membrane to alter the first sensor reading;wherein the first membrane carries a first electrode which forms capacitance with another electrode spaced from the first electrode to generate the first sensor reading. 9. The system of claim 8 wherein the first membrane carries at least one strain gauge to generate the first sensor reading when deformed. 10. The system of claim 8 wherein the second membrane includes piezoelectric material. 11. The system of claim 8 wherein the second circuit applies a voltage across the second membrane energize it. 12. A method of compensating a pressure sensor immersed in a fluid, comprising: obtaining a first sensor reading from a first, deformable membrane within the pressure sensor;energizing a second deformable membrane contiguous to and touching the first membrane to deform and deform the first membrane to alter the first sensor reading;determining a delta response from the first sensor reading and the altered sensor reading; andutilizing the delta response to determine the actual pressure of the fluid. 13. The method of claim 12 wherein energizing the second membrane includes applying a voltage across it. 14. The method of claim 12 wherein utilizing the delta response includes retrieving a pressure value corresponding t... | MEDOS INTERNATIONAL S.A.R.L. | A61B005/00 G01L019/02 G01L027/00 A61B005/03
US09636031 | B2 | 2017-05-02 | Stylets for use with apparatus for intravascular placement of a catheter | 2010-09-29 | An integrated catheter placement system for accurately placing a catheter within a patient's vasculature is disclosed. In one embodiment, the integrated system comprises a system console, a tip location sensor for temporary placement on the patient's chest, and an ultrasound probe. The tip location sensor senses a magnetic field of a stylet disposed in a lumen of the catheter when the catheter is disposed in the vasculature. The ultrasound probe ultrasonically images a portion of the vasculature prior to introduction of the catheter. ECG signal-based catheter tip guidance is included in the integrated system to enable guidance of the catheter tip to a desired position with respect to a node of the patient's heart. Stylets and catheters including various multiple bipolar and monopolar electrode configurations are also disclosed. | 1. A catheter placement system configured for positioning a catheter within a vasculature of a patient, comprising: a console having a display;a plurality of electrodes located at a distal end of the catheter, of each the plurality of electrodes configured for sensing separate ECG signals of the patient, and the plurality of electrodes being operably connected to the console; anda processor included with the console, the processor: (i) performing a differential comparison of a plurality of ECG signals sensed by a subset of the plurality of electrodes with a baseline ECG signal sensed by a baseline electrode of the plurality of electrodes to determine resultant ECG signals and a position of the distal end of the catheter based on at least one of the resultant ECG signals;(ii) identifying an amplitude of a P-wave from each of the resultant ECG signals; and(iii) generating an output for the display that plots the amplitudes of the P-wave from the resultant ECG signals according to a position of each of the plurality of electrodes on the catheter. 2. The catheter placement system as defined in claim 1, wherein the processor implements an algorithm for performing the differential comparison that provides the resultant ECG signals, the resultant ECG signals providing information relating to the position of the distal end of the catheter with respect to a node of the patient's heart producing the plurality of ECG signals. 3. The catheter placement system as defined in claim 2, wherein the resultant ECG signals include only differences between the plurality of ECG signals and the baseline ECG signal. 4. The catheter placement system as defined in claim 3, wherein the resultant ECG signals differences in amplitude of at least one of a P-wave and a QRS complex between each of the plurality of ECG signals and the baseline ECG signal. 5. The catheter placement system as defined in claim 1, wherein the plurality of electrodes is disposed on a stylet removably positionable within a lumen of the catheter and further includes a plurality of distal electrodes of a first polarity disposed proximate a distal end of the stylet, a counterpart electrode of a second polarity, and a ground electrode. 6. The catheter placement system as defined in claim 5, wherein the plurality of distal electrodes and the counterpart electrode cooperate to define a plurality of bipolar electrodes. 7. The catheter placement system as defined in claim 6, wherein six bipolar electrodes are included on the stylet, including six distal electrodes of a positive polarity and a proximally disposed counterpart electrode of a negative polarity. 8. The catheter placement system as defined in claim 5, wherein the stylet further includes at least one magnetic element for enabling tracking of the stylet by an external sensor. 9. The catheter placement system as defined in claim 1, wherein the plurality of electrodes include monopolar electrodes that are disposed on a stylet removably position... | C.R. Bard, Inc. | A61B005/042 A61B008/08 A61B034/20 A61B005/06 A61M025/01 A61B046/00 A61B090/40 A61B090/00
US20110021883 | A1 | 2011-01-27 | Remote health care system for enabling user to select a symptom diagnosis algorithm appropriate to his/her individual disease and method thereof | 2009-12-29 | This invention relates to a remote health care system and a method thereof for allowing medical experts specializing in various medical fields to upload various developed program files of symptom diagnosis algorithms to a database of a care server in the remote health care system for adding, revising or updating the program files stored therein, so that the care server can provide a variety of new and useful program files to a user and allow the user to select the program files more appropriate to his/her individual disease and to upload his/her physiological parameter data to the care server. Then, the care server can execute corresponding calculating and analyzing procedures to the uploaded physiological parameter data by using the symptom diagnosis algorithm in the selected program file, and transmit calculation and analysis results to the user or health care workers for precisely and objectively understand the user's individual physiological condition. | 1. A remote health care method for enabling a user to select a symptom diagnosis algorithm appropriate to his/her individual disease, which is applied to a remote health care system comprising a care server, at least one first user end and at least one second user end, wherein the care server is provided with a database which stores a plurality of program files and plural sets of file information, each said first user end and second user end link to the care server via a network, respectively, each said first user end uploads at least one program file to be stored in the database, and each said program file includes program codes of a symptom diagnosis algorithm, the method controlling the care server to execute steps comprising: receiving a selection command of selecting each said program file from the database by each said second user end;receiving at least one set of physiological parameter data uploaded by each said second user end;executing a calculation and analysis procedure to each said set of physiological parameter data according to the algorithm in each said selected program file through a calculation and analysis unit; andtransmitting a calculation and analysis result obtained by said calculation and analysis unit to the corresponding second user end. 2. The remote health care method according to claim 1, wherein the method further controls the care server to execute steps of: receiving each said program file uploaded by each said first user end; andstoring each said program file in the database. 3. The remote health care method according to claim 1, wherein the method further allows each said second user end to execute a step of: reading the program files in the database, and listing and showing file information of each said program file in series on a selectable field in an operational window of a display screen of each said second user end, so that the user is allowed to select the file information. 4. The remote health care method according to claim 2, wherein the method further allows each said second user end to execute a step of: reading the program files in the database, and listing and showing file information of each said program file in series on a selectable field in an operational window of a display screen of each said second user end, so that the user is allowed to select the file information. 5. The remote health care method according to claim 3, wherein the operation window further comprises a selected field, and the method further allows each said second user end to execute a step of: listing and showing the file information of each said selected program file on the selected field in series after the user selects each said program file for the user's confirmation. 6. The remote health care method according to claim 4, wherein the operation window further comprises a selected field, and the method further allows each said second user end to execute a step of: listing and showing the file information of each said ... | Aescu Technology Inc. | A61B005/00 G06Q050/00 G06Q030/00
US07879043 | B2 | 2011-02-01 | System and method for preventing intraoperative fracture in cementless hip arthroplasty | 2006-11-28 | A method and system for monitoring impaction of a femoral component of a hip prosthesis into a femur in which impaction data generated during the impaction of the femoral component into the femur is received from at least one measurement transducer attached to the femoral component and is normalized by a data acquisition and analysis device. An impaction monitoring metric is calculated based on the normalized impaction data, and femoral component fit and stability data is then generated and output to a user interface based on the impaction monitoring metric. | 1. A method for monitoring impaction of a femoral component of a hip prosthesis into a femur, comprising: receiving impaction data generated during the impaction of the femoral component into the femur from at least one accelerometer attached to the femoral component;normalizing the impaction data into a usable format as normalized impaction data;calculating a total norm of an acceleration signal including the normalized impaction data based on the following equation:NormTotal=∑i=1N⁢ai2 where N is a number of points in a time interval to be analyzed and a is an acceleration measurement taken at an ith point by the accelerometer;calculating an impaction monitoring metric Ej for a femoral component impaction hit as a fraction of the total norm of the acceleration signal during the time interval at a jth point based on the following equation:Ej=∑i=1j⁢ai2NormTotaldetermining a time t(j) corresponding to a fractional value Ej within a predetermined range indicative of the femoral component being fully seated for each of a plurality of femoral component impaction hits that is equal to a length of time required for Ej to reach a predetermined fractional value indicative of femoral component fit and stability after the impaction hit; andgenerating an output based on the time t(j) for the femoral component impaction hit, the output being indicative of the femoral component fit and stability. 2. The method of claim 1, wherein the length of time t(j) required for Ej to reach a predetermined fractional value indicative of femoral component fit and stability after the impaction hit corresponds to Ej=0.99. 3. The method of claim 1, wherein the receiving of impaction data comprises receiving output acceleration data generated during the impaction of the femoral component into the femur from a single accelerometer attached to the femoral component. 4. The method of claim 3, wherein the single accelerometer is oriented in an impaction direction of the femoral component along a relative z axis. 5. The method of claim 1, wherein the generating an output based on the time t(j) for the femoral component impaction hit comprises generating an output based on the time t(j) corresponding to Ej=0.99 for the femoral component impaction hit. 6. A system for monitoring femoral component fit and stability during impaction of a femoral component of a hip prosthesis into a femur, comprising: a testing component including at least one accelerometer that is attachable to the femoral component, the at least one accelerometer being operable to generate impaction data corresponding to the impaction of the femoral component into the femur;a data acquisition and analysis device in communication with the at least one accelerometer and configured to receive the impaction data, normalize the impaction data into a usable format as normalized impaction data and calculate an impaction monitoring metric by calculating (1) a total norm of an acceleration signal including the normalized i... | Robert Michael Meneghini Phillip John Cornwell Aaron Glen Rosenberg | A61B017/58 A61B017/60 A61B019/00 A61B005/103 A61B005/117 A61F002/00
US07890189 | B2 | 2011-02-15 | Mono-body defibrillaton probe | 2005-01-24 | A probe including at its distal extremity a tubular flexible sheath core supporting at least a winding forming a shock electrode and connected to a electrical conductor of connection extending in a internal lumen of the sheath core. The sheath core extends axially without a solution of continuity in the area supporting the winding. In particular, the sheath core comprises cavities to receive and hold conducting inserts, of homologous size with cavities formed locally close to the ends of the winding, the insert being connected to the interior side to the electrical conductor, and on the external side to the corresponding extremity of winding. A longitudinal slit connects the two cavities and allows, by elastic deformation of the sheath core, the introduction into the cavities and in the internal lumen of the unit formed by the final extremity of the electrical conductor beforehand equipped with its two inserts. | 1. A mono-body defibrillation probe, comprising a body having a distal extremity and a periphery area, a tubular flexible sheath core at said body distal extremity, and an internal lumen extending longitudinally along at least a length of said body, said body periphery area supporting at least one winding of wire forming a shock electrode for application of an energy of defibrillation or cardioversion, said winding of wire being electrically connected to an electrical connection conductor extending longitudinally in said internal lumen, wherein said sheath core extends axially without a solution of continuity in said periphery area supporting said winding and wherein said sheath core further comprises: a cavity in the vicinity of each extremity of the winding, and a narrower longitudinal slit connecting said cavities and radially extending from an external surface of said sheath core to said internal lumen, so as to allow, by elastic strain of the material of said sheath core on both sides of said slit, introduction into each of said cavities an insert made of a conducting material, of a size homologous with the size of each of said respective cavities, wherein each of said inserts has an interior side and exterior side, being electrically connected on said interior side with said electrical conductor and on said external side at the corresponding extremity of said winding. 2. The mono-body probe of claim 1, further comprising a junction ring for mechanical and electric connection of at least one of said inserts to said winding, said junction ring being a cylindrical conducting material ring, with an internal surface able to cooperate with a part turned towards an outside of said insert, and an external surface comprising a connection part able to cooperate with a part turned towards an interior of said extremity of winding. 3. The mono-body probe of claim 2, wherein said junction ring comprises, in the area of said internal surface able to cooperate with said insert, an assembly part able to allow a mechanical and electric solidarization of said ring to said insert. 4. The mono-body probe of claim 3, wherein said assembly part is a part comprising an opening to allow a solidarization of said ring to said insert by welding from the outside. 5. The mono-body probe of claim 3, wherein a diameter of said assembly part is greater than a diameter of said connection part, the difference of the diameters being approximately equal to double a thickness of said winding, so that an external surface of said ring is approximately level with the external surface of said winding. 6. The mono-body probe of claim 5, further comprising an external envelope made out of a flexible insulated material sheathing said sheath core over a length of said sheath core except in the area of said winding, a diameter of said external envelope being approximately equal to a diameter of said winding. 7. The mono-body probe of claim 6, wherein said junction ring also compri... | ELA Medical S.A.S. | A61N001/05 A61B005/0408
US20110034790 | A1 | 2011-02-10 | APPARATUS AND METHODS FOR CORONARY SINUS ACCESS | 2010-10-18 | Apparatus and methods for locating morphological features within a body cavity using a catheter including proximal and distal ends, a transparent balloon carried on the distal end, and an optical imaging assembly carried on the distal end for imaging through the balloon. The balloon includes a channel extending therethrough to a lumen extending through the catheter. A guidewire or other localization member is received in the lumen that is extendable through the channel. During use, the catheter is inserted into a right atrium of a heart, and the balloon is expanded and placed against the wall of the heart to locate the coronary sinus. Sufficient force is applied to clear blood between the surface and the wall and clear the field of view of the imaging assembly. The catheter is manipulated to locate the coronary sinus, whereupon the localization member is advanced into the coronary sinus. | 1. A method of assessing a tissue region of interest, comprising: inserting a distal end of a catheter into a blood-filled body lumen, the catheter having a barrier or membrane that is at least partially transparent and projecting from the catheter;expanding the barrier or membrane with a transparent fluid from an insertion configuration to a deployed configuration;contacting the tissue region with a distal surface of the barrier or membrane;imaging the tissue region through the transparent fluid with an imaging element positioned within the barrier or membrane proximal to the distal surface;sensing a physiological parameter of the tissue region via one or more electrodes extending along the barrier or membrane such that the electrodes contact the tissue region when the barrier or membrane is expanded; andrepositioning the barrier or membrane along the tissue region depending upon the physiological parameter sensed by the one or more electrodes while maintaining visualization of the tissue region. 2. The method of claim 1, wherein inserting comprises intravascularly advancing the distal end into a chamber of a heart. 3. The method of claim 1, wherein expanding comprises inflating the barrier or membrane with the transparent fluid. 4. The method of claim 1, wherein contacting further comprises advancing an instrument from the catheter. 5. The method of claim 4, wherein advancing an instrument from the catheter comprises advancing a guidewire from the catheter. 6. The method of claim 4, wherein advancing an instrument from the catheter comprises stabilizing the distal end of the catheter relative to the tissue region via the instrument. 7. The method of claim 1, further comprising delivering a fluid between the distal surface and the tissue region to at least partially clear blood such that imaging of the tissue region is enhanced. 8. The method of claim 1, further comprising infusing a therapeutic agent into the tissue region. 9. The method of claim 1, wherein sensing a physiological parameter of the tissue region comprises detecting an oxygen level of the tissue region. 10. The method of claim 1, wherein repositioning the barrier or membrane comprises manipulating the barrier or membrane to move along the tissue region while maintaining visualization through the barrier or membrane until a second location of the tissue region is located. 11. The method of claim 1, wherein repositioning the barrier or membrane comprises reconfiguring the barrier or membrane from its deployed configuration to its insertion configuration. 12. A method of assessing a tissue region of interest, comprising: inserting a distal end of a catheter into a blood-filled body lumen, the catheter having a barrier or membrane that is at least partially transparent and projecting from the catheter;expanding the barrier or membrane with a transparent fluid from an insertion configuration to a deployed configuration;imaging the tissue region contacted against the barrier... | The Board of Trustees of the Leland Stanford Jr. University | A61B005/04 A61M005/14 A61B005/145
US07881765 | B2 | 2011-02-01 | Interference-fit shroud-based cardiac electrodes | 2006-04-28 | Apparatus and method according to the disclosure relate to minimizing gaps between a substantially planar cardiac-sensing electrode and a shroud member utilizing a so-called interference-fit. For example, a relatively recessed area or aperture formed in an exemplary resin-based shroud member has slightly reduced dimensions relative to the electrode and requires compression forces during assembly (e.g., manually or in an automated process including a press, a tool or other means). The interference-fit promotes a very tight fit (or seal) between the metallic electrode and the resin-based shroud member and, importantly, minimizes gaps. Additionally, discrete interference structures promote fluid tight seals between the electrode and a recess or aperture adapted to receive the electrode. | 1. A subcutaneous cardiac activity sensing shroud, comprising: a shroud member for location adjacent at least a part of the peripheral portion of a housing for an implantable medical device (IMD), said shroud member including one of a recessed region and an aperture; anda substantially planar, plate-type electrode having a periphery that mechanically sealingly interlocks to one of the recessed region and the aperture,wherein at least a portion of an interior periphery of said recessed region or said aperture includes an interference structure configured to sealingly receive the electrode when said electrode is disposed therein. 2. A shroud according to claim 1, wherein said interference structure comprises a resilient material. 3. A shroud according to claim 1, further comprising a compressible member disposed intermediate the electrode and the shroud where the electrode and the shroud are abutting each other. 4. A shroud according to claim 3, wherein the compressible member comprises a resin-based member. 5. A shroud according to claim 3, wherein the compressible member comprises a biocompatible gasket and further comprising a medical grade adhesive coupled between one of: the biocompatible gasket and the electrode and the biocompatible gasket and the shroud. 6. A shroud according to claim 3, wherein the compressible member comprises a member having a C-shaped cross-sectional configuration. 7. A shroud according to claim 3, wherein the compressible member comprises a member having opposing C-shaped cross sectional configuration. 8. A shroud according to claim 1, further comprising an elongated conductor coupled to one of a major surface of the electrode and a peripheral portion of the major electrode. 9. A shroud according to claim 8, wherein the elongated conductor comprises an integrally formed structure with the electrode. 10. A shroud according to claim 1, wherein the electrode comprises one of a titanium material and a platinum material. 11. A shroud according to claim 1, further comprising medical grade adhesive disposed around between at least a part of the periphery of the electrode and the periphery of the recessed region or the aperture. 12. A shroud according to claim 1, wherein the electrode further includes a coating on at least a major surface thereof. 13. A shroud according to claim 12, wherein the coating comprises one of a nitride coating, a carbon black coating, a time-release coating. 14. A subcutaneous cardiac activity sensing shroud, comprising: a shroud member for location adjacent at least a part of the peripheral portion of a housing for an implantable medical device (IMD), said shroud member including one of a recessed region and an aperture;a substantially planar, plate-type electrode having a periphery that mechanically sealingly interlocks to one of the recessed region and the aperture,wherein at least a portion of an interior periphery of said recessed region or said aperture includes an interference str... | Medtronic, Inc. | A61B005/04
US07892627 | B2 | 2011-02-22 | Pattern transferable to skin for optical measurements during shaving | 2009-05-07 | A system for measuring 3D deformation and strain on the skin of a living person. A non contact optical 3D deformation measuring system is provided having a plurality of high resolution digital cameras for measuring 3D coordinates and 3D deformations operatively connected to a computer. A removable tattoo is also provided including a substrate and a pattern having a plurality of indicia distributed to form a pattern density of between about 40% and about 60%. The removable tattoo provides a reference marking for the non contact optical 3D deformation measuring system. | 1. A kit for measuring 3D deformation and strain on the skin of a living person comprising: a non-contact optical 3D deformation measuring system including a plurality of high resolution digital cameras for measuring 3D coordinates and 3D deformations operatively connected to a computer;a removable tattoo including a substrate and a pattern having a plurality of indicia distributed to form a pattern density of between about 40% and about 60%, wherein the removable tattoo provides a reference marking for the non-contact optical 3D deformation measuring system and;a shaving razor transmitting a force to the skin. 2. The kit of claim 1, wherein the indicia comprises dots. 3. The kit of claim 2, wherein the dots comprise two or more different sized dots. 4. The kit of claim 2, wherein the dots comprise three different sized dots. 5. The kit of claim 1, wherein the pattern density is about 50%. 6. The kit of claim 1, wherein the pattern density is about 42.5%. 7. The kit of claim 1, wherein the indicia is transferable to the skin by wetting with an alcohol. 8. The kit of claim 1, wherein the substrate is a paper. 9. The kit of claim 8, wherein the paper is chosen from a group of paper consisting of blotting paper and cigarette paper. 10. The kit of claim 1, wherein the indicia comprises a substantially water-insoluble ink or dye. 11. The kit of claim 10, wherein the indicia comprises an oil-based ink or dye. 12. The kit of claim 1, wherein the indicia is devoid of a cover layer formed above said indicia opposite said substrate. 13. The kit of claim 1, wherein the removable tattoo is devoid of an adhesive layer atop said indicia. | The Gillette Company | B41M005/00 B44C001/17 G03G007/00 A61B005/00 A61B005/103 A61B005/117
US20110046457 | A1 | 2011-02-24 | MULTILUMEN CATHETER | 2010-11-01 | A multilumen catheter having tubings extending into lumens within the catheter. The lumens may be used for blood, drugs or other medicants. The lumens may also be used for sensors. The junction element, external to the patient, connects the tubings to the lumens. The tubings, also external to the patient, connect to infusion members, to which one or more infusion systems may be connected to deliver blood, drugs and other medicants to the patient. A sensor having a sensing element may extend through the sensor lumen and be positioned internal to the patient for physiological parameter sensing. An external portion of the sensor may be connected to associated electronics to provide automatic monitoring of the physiological parameters and automatic delivery and control of the infusants. Also, a central line catheter for delivering fluids directly into a main artery or vein near the heart, which contains a first lumen to deliver a fluid through the central line catheter and a second lumen containing a sensor capable indicating a characteristic level in blood. The first lumen delivers the fluid downstream of the sensor to prevent any interference between the fluid delivery and the sensor readings. In other versions, the central line catheter can have additional lumens for additional purposes. In addition, the central line catheter can further include a flush sleeve to remove debris around the sensor. | 1.-38. (canceled) 39. A multilumen catheter comprising: a first tubing having a length dimension and a plurality of lumens;the plurality of lumens disposed within the first tubing and extending along the length dimension of the first tubing, the plurality of lumens including a sensor lumen and a first infusion lumen;a sensor lead extending within the sensor lumen along the length dimension of the first tubing to a sensor location;a sensor connected to the sensor lead and arranged at the sensor location;the first infusion lumen having an exit outlet on the first tubing at a location spaced apart along the length dimension of the first tubing from the sensor location. 40. A multilumen catheter of claim 39, wherein the sensor location is within the sensor lumen of the first tubing. 41. A multilumen catheter of claim 39, wherein: the first tubing has a distal end and a connection end;the sensor location is closer to the distal end of the first tubing than the exit outlet is to the distal end of the first tubing. 42. A multilumen catheter of claim 41, wherein the sensor lumen terminates before the distal end of the first tubing. 43. A multilumen catheter of claim 41, further comprising a junction element connected to the connection end of the first tubing, the junction element connected to at least one second tubing. 44. A multilumen catheter of claim 41, further comprising a junction element connected to the connection end of the first tubing, the junction element connected to a plurality of second tubings. 45. A multilumen catheter of claim 44, wherein the plurality of second tubings comprises a first infusion tubing and a second infusion tubing. 46. A multilumen catheter of claim 45, wherein: the first tubing includes a second infusion lumen; andthe first infusion tubing extends into the first infusion lumen and the second infusion tubing extends into the second infusion lumen. 47. A multilumen catheter of claim 39, wherein: the first tubing has a distal end and a connection end;the sensor location along the length dimension of the first tubing, is between the exit outlet and the distal end of the first tubing. 48. A multilumen catheter of claim 47, wherein the exit port is closer to the connection end than the sensor location is to the connection end. 49. A method of making a multilumen catheter, the method comprising: forming a first tubing having a length dimension and a plurality of lumens, the plurality of lumens disposed within the first tubing and extending along the length dimension of the first tubing, the plurality of lumens including a sensor lumen and a first infusion lumen;extending a sensor lead within the sensor lumen along the length dimension of the first tubing to a sensor location;connecting a sensor to the sensor lead and arranging the sensor at the sensor location; andforming an exit outlet of the first infusion lumen in the first tubing at a location spaced apart along the length dimension of the first tubing from t... | Medtronic MiniMed, Inc. | A61B005/145 B23P017/04
US20110028803 | A1 | 2011-02-03 | METHOD AND DEVICE FOR NON-INVASIVE DETERMINATION OF THE CONCENTRATION OF A SUBSTANCE IN A BODY FLUID | 2008-03-31 | The present invention relates generally to the field of measuring glucose levels in a body fluid of a subject, in particular to non-invasive measuring of blood glucose in blood of a subject. A device is disclosed comprising an electrically conducting probe comprising a plurality of electrodes adapted for measuring electrical impedance of the tissue of a subject, typically skin, a sensing device adapted for sensing at least one concentration of an ion in the tissue, typically acidity (pH), wherein said electrically conducting probe and said sensing device are adapted to obtain said electrical impedance of the tissue and said concentration of an ion in the tissue simultaneously. The device further comprises a blood glucose determining unit adapted for determining an estimate, typically a concentration, of blood glucose based on said electrical impedance of the tissue and said concentration of an ion in the tissue. Further, a device is disclosed comprising an electrically conducting probe comprising a plurality of electrodes adapted for measuring electrical impedance of the tissue of a subject, typically the skin, wherein at least one of said plurality of electrodes is adapted to, when being in voltage mode, sense a signal representative of a concentration of an ion in the tissue, typically acidity (pH), wherein said electrically conducting probe is adapted to simultaneously obtain said impedance of the tissue and said concentration of an ion in the tissue, the device also comprising a blood glucose determining unit adapted for determining an estimate, typically a concentration, of blood glucose based on said electrical impedance of the tissue and said concentration of an ion in the tissue. Furthermore, methods for non-invasive determination of an estimate of blood glucose in blood of a subject are disclosed. | 1.-44. (canceled) 45. A device for non-invasive determination of an estimate of glucose in the blood of a subject, comprising: an electrically conducting probe comprising a plurality of electrodes adapted to measure the impedance of the tissue of a subject,at least one sensing device adapted to sense at least one concentration acidity (pH) in the tissue of a subject,wherein said electrically conducting probe and said sensing device are adapted to substantially simultaneously measure the impedance of the tissue and sense a concentration of acidity (pH) in the tissue, respectively, anda blood glucose determining unit adapted to determine an estimate of blood glucose of said subject based on said tissue impedance and said acidity (pH) concentration. 46. A device for non-invasive determination of an estimate of blood glucose in the blood of a subject, comprising: an electrically conducting probe comprising a plurality of electrodes adapted to measure a value of the impedance of the tissue of a subject,wherein at least one of said plurality of electrodes is adapted for, when being in voltage mode, sense a signal representing a value of the concentration of acidity (pH) in the tissue of a subject, andwherein said electrically conducting probe is adapted to substantially simultaneously obtain said value of the impedance of the tissue of a subject and said value of the concentration of acidity (pH) in the tissue of a subject, anda blood glucose determining unit adapted to determine an estimate of blood glucose of said subject based on said tissue impedance and said acidity (pH) concentration. 47. The device according to claim 46, wherein at least one of said plurality of electrodes is made of an acidity-sensitive material. 48. The device according to claim 46 or 47, wherein said at least one of said plurality of electrodes is made of iridium, antimony, palladium, ruthenium, bismuth, zirkonium, oxides of iridium, antimony, palladium, ruthenium, bismuth, or zirkonium, or compositions of iridium, antimony, palladium, ruthenium, bismuth, or zirkonium, or oxides of iridium, antimony, palladium, ruthenium, bismuth, or zirkonium. 49. The device according to claim 45, wherein said electrically conducting probe comprises one pair of electrodes, wherein one electrode is a current injection electrode and the other is a voltage sensing electrode, or wherein said electrically conducting probe comprises two pairs of electrodes, each pair being a current injection electrode and a voltage sensing electrode. 50. The device according to claim 45 or 46, wherein said blood glucose determining unit is adapted to send an alert signal if said blood glucose estimate is below or above predetermined reference glucose levels. 51. The device according to claim 45 or 46, further comprising an insulin delivery device adapted to deliver an insulin dose based on said blood glucose estimate, and wherein said blood glucose determining unit is adapted to send a blood glucose estima... | - | A61B005/145
US20110028819 | A1 | 2011-02-03 | MEDICAL SENSOR DEVICE | 2010-07-22 | Electromyographic and mechanomyographic parameters of a patient are detected accurately with few artifacts in a medical sensor device (11) for a patient. The medical sensor device (11) includes an electrode (12) for detecting an electric voltage on a body surface of the patient, a holding element (10) with preferably at least one transmission means for transmitting or conducting signals or electric currents, at least one mechanical connection means (14) for the detachable mechanical connection of electrode (12) with holding element (10), and at least one electric connection means (15) for the detachable electric connection of electrode (12) with holding element (10). the holding element (10) comprising at least one sensor for detecting at least one medical parameter of the patient. | 1. A medical sensor device for a patient, the medical sensor device comprising: an electrode for detecting an electric voltage on a body surface of the patient;a holding element with a transmission means for transmitting or conducting signals or electric currents;a detachable mechanical connection means for the detachable mechanical connection of the electrode with the holding element;a detachable electric connection means for the detachable electric connection of the electrode with the holding element, the holding element comprising a sensor for detecting at least one medical parameter of the patient. 2. A medical sensor device in accordance with claim 1, wherein the sensor is a mechanomyographic sensor. 3. A medical sensor device in accordance with claim 1, wherein the sensor is an acceleration sensor and/or a microphone. 4. A medical sensor device in accordance with claim 1, wherein the detachable mechanical connection means has at least one elastic element for generating a prestress between the electrode and the holding element. 5. A medical sensor device in accordance with claim 1, wherein the detachable mechanical connection means provides a coefficient of friction between the holding element and the electrode that equals at least 0.05 to 0.25, including between a contact surface of the holding element and a contact surface of the electrode, at right angles to a contact surface of the electrode, and/or the prestressing force at the contact surface equals at least 0.5 N to 5 N. 6. A medical sensor device in accordance with claim 1, wherein the detachable mechanical connection provides static friction between the holding element and the electrode that equals 0.05 N to 10 N, including between the contact surface of holding element and the contact surface of the electrode, at right angles to contact surface of electrode. 7. A medical sensor device in accordance with claim 1, wherein the especially detachable mechanical connection means establishes at least one of a clip connection a locking connection, a frictionally engaged connection, a positive-locking connection and a nonpositive connection. 8. A medical sensor device in accordance with claim 1, wherein the detachable mechanical connection means comprises a mandrel formed at the electrode and a recess formed at the holding element, the mandrel being arranged in the recess. 9. A medical sensor device in accordance with claim 1, wherein the detachable mechanical connection means comprises an elastic element including a spring or an O-ring. 10. A medical sensor device in accordance with claim 1, wherein the electrode comprises a chamber and/or the contact surface including an adhesive surface portion. 11. A medical sensor device in accordance with claim 1, wherein the electrode is a disposable electrode and/or holding element that can be reused several times. 12. A medical sensor device in accordance with claim 1, wherein the transmission means for transmitting signals or electric c... | Drager Medical AG & Co. KG | A61B005/04
US20110004088 | A1 | 2011-01-06 | THE ECG SHIRT | 2009-05-13 | The electrocardiograph (ECG) Shirt is an extremely elastic non-electric and non-mechanical method that creates a more secure bond to the skin and more accurate and standardized ECG electrode lead placement for men, women, and children. The ECG Shirt reduces unexpected movement or dislocation. Contrasted with much prior art that analyzes errors after the fact the ECG Shirt eliminates the potential for errors to occur. The ECG Shirt is capable of using most ECG devices with different types of electrode lead connectors or pads. The ECG Shirt facilitates not only twelve lead resting ECG testing but also exercise testing, monitoring, fifteen lead, and two different types of eighteen lead ECG testing. The ECG Shirt is a tool for teaching expert ECG electrode lead placement. | 1. A method and apparatus that defines and secures Electrocardiogram (ECG) lead placement comprising; 1 An elastic method and apparatus comprising; a garment with or without sleeves; an elastic fabric; apertures, straps, and garters, holding firm contact of the ECG leads, garment, and ECG electrodes to the skin; labels, instructions, and color-coding for clear, easy, concise, and clinically diverse ECG electrode lead placement according to international standards;2 An embodiment of claim 1 except with a design of a brassiere; supporting, cupping, raising, and separating the breasts to facilitate correct ECG lead placement;3 An embodiment of claim 1 as a vest with an additional label on the abdomen clearly directing placement of the Embodiment of claim 5;4 An embodiment of claim 2 and an embodiment of claim 3 without sleeves;5 An embodiment of claim 1 comprising four distinct elastic garters for ECG lead placement of limbs, with apertures identical except for the color and labels, for use with a the embodiments of claim 2 and the embodiment of claim 4;6 The apertures of claim 1 and the embodiments of claim 2, claim 3, claim 4, and claim 5 wherein an electrode is able to be positioned in the precise anatomical location corresponding with accepted clinical protocols for ECG testing of the twelve lead, fifteen lead, and eighteen lead for testing and monitoring; resting or during exercise;7 A plurality of Labels and Instructions encompassing claim 1 and the embodiments of claim 2, claim 3, claim 4, claim 5, and claim 6 with bold visual cues and specific instructions according to the highest standards of clinical cardiovascular ECG testing placed in the precise location of ECG lead placement; facilitating and increasing the speed of ECG lead application to a patient' body; and providing additional advanced instruction for vital ECG tests using the right precordial leads and posterior leads; | - | A61B005/0408
US20110046455 | A1 | 2011-02-24 | METHODS AND DEVICES FOR EXAMINING A PARTICULAR TISSUE VOLUME IN A BODY, AND A METHOD AND A DEVICE FOR SEGMENTING THE PARTICULAR TISSUE VOLUME | 2010-08-17 | A tip of an elongate device is navigated into a particular tissue volume in order to examine the particular tissue volume in a body and part of the tissue volume is analyzed in real-time by way of a biosensor. In at least one embodiment, in the process, either the biosensor can be arranged on the tip of the device or the device includes a catheter, by which a substance to be analyzed is transported out of the tissue volume from the tip to the biosensor arranged at the proximal end of the catheter. | 1. A device for examining a tissue volume in a body, comprising: a guiding device; anda biosensor disposed on the guiding device, wherein the biosensor is guidable through the body into part of the tissue volume via the guiding device and wherein the biosensor is designed for real-time analysis of the part of the tissue volume. 2. The device as claimed in claim 1, wherein the guiding device comprises a catheter, a pump and a container, the catheter connecting a distal region, on which the biosensor is arranged, to a proximal region at a distance therefrom, and wherein the pump and container are connected to the catheter in the proximal region such that the pump can pump a substance situated in the container through the catheter to the distal region. 3. The device as claimed in claim 1, wherein a coil is arranged on the biosensor to improve imaging of the biosensor by way of a magnetic resonance scanner. 4. A device for examining a tissue volume in a body, comprising: a sample acquisition device;a catheter; anda biosensor, wherein the catheter connects the sample acquisition device to the biosensor such that a sample is transportable through the catheter from the sample acquisition device to the biosensor, wherein the sample acquisition device is guidable through the body into part of the tissue volume and wherein the biosensor is designed for real-time analysis of the sample. 5. The device as claimed in claim 4, further comprising a pump and a container, the pump and the container being connected to the catheter at a proximal end of the catheter on which the biosensor is arranged such that the pump can pump a substance situated in the container through the catheter to a distal end of the catheter. 6. The device as claimed in claim 4, wherein a coil is arranged on the sample acquisition device to improve imaging of the sample acquisition device by way of a magnetic resonance scanner. 7. The device as claimed in claim 2, the container is developed to hold particles with antibodies, oligonucleotides, enzymes or other specific binding molecules, wherein the device is developed such that the pump can be used to pump the particles through the catheter to the distal end of the catheter and wherein the biosensor is developed for real-time analysis by way of the particles. 8. The device as claimed in claim 1, wherein the biosensor comprises at least one of the following devices: a device for applying the quartz crystal microbalance technique,a detector for evanescent fields,an impedance sensor,a device for applying surface plasmon resonance technology,a device for applying calorimetry,a device for carrying out electrochemical detection,a device for carrying out fluorescence-optical analysis,a device for carrying out spectrophotometric analysis, anda device for carrying out enzymatic analysis. 9. The device as claimed in claim 1, wherein the biosensor comprises an arrangement of a plurality of biosensors in order to analyze a plurality of parameter... | - | A61B005/107 A61M025/00 A61B005/055 A61B005/145 A61B006/00 A61B005/053 A61B006/03
US20110009710 | A1 | 2011-01-13 | COMBINATION LEVEL ALARMS AND ALARM PERSISTENCE FOR PATIENT MONITORING | 2010-09-24 | Combination level alarms and alarm persistence for patient monitoring, as well as related methods and apparatus, are disclosed. For a combination level alarm, respective values for multiple physiological conditions are received. A determination is made as to whether the received values satisfy respective alarm criteria for the physiological conditions, and if so, an alarm is raised. An alarm persistence feature involves periodically receiving values of a physiological condition, and determining whether the received values have satisfied an alarm criterion for at least a predetermined portion of a predetermined period of time. If so, an alarm is maintained. Combination alarms and alarm persistence may be implemented individually or together. | 1. A method comprising: periodically receiving values of a physiological condition from a sensor;determining whether the received values have satisfied an alarm criterion for at least a predetermined portion of a predetermined period of time; andmaintaining an alarm where the received values have satisfied the alarm criterion for at least the predetermined portion of the predetermined period of time. 2. The method of claim 1, wherein the received values indicate values of the physiological condition at respective times, and wherein determining comprises determining whether at least a minimum number of the received values that indicate values of the physiological condition during a time period equal to the predetermined period of time immediately preceding a time of a most recently received value satisfied the alarm criterion. 3. The method of claim 1, further comprising: periodically receiving values of a further physiological condition from a further sensor,wherein determining further comprises determining whether the received values of the further physiological condition have satisfied a further alarm criterion for at least the predetermined portion of the predetermined period of time, andwherein maintaining comprises maintaining the alarm where both the received values of the physiological condition have satisfied the alarm criterion for at least the predetermined portion of the predetermined period of time and the received values of the further physiological condition have satisfied the further alarm criterion for at least the predetermined portion of the predetermined period of time. 4. The method of claim 1, wherein maintaining comprises maintaining an indication of the alarm in a GUI (Graphical User Interface). 5. A computer readable medium having computer readable instructions stored thereon that when executed implement the method according to claim 1. 6. Apparatus comprising: a sensor interface that enables the apparatus to periodically receive values of a physiological condition from a sensor; andan analysis module, operatively coupled to the sensor interface, that determines whether the received values have satisfied an alarm criterion for at least a predetermined portion of a predetermined period of time, and maintains an alarm where the received values have satisfied the alarm criterion for at least the predetermined portion of the predetermined period of time. 7. The apparatus of claim 6, wherein the received values indicate values of the physiological condition at respective times, and wherein the analysis module determines whether the received values have satisfied the alarm criterion for at least the predetermined portion of the predetermined period of time by determining whether at least a minimum number of the received values that indicate values of the physiological condition during a time period equal to the predetermined period of time immediately preceding a time of a most recently received value satisfied the alarm c... | BRYTECH INC. | A61B005/00
US20110004084 | A1 | 2011-01-06 | Method of Calibrating an Analyte-Measurement Device, and Associated Methods, Devices and Systems | 2010-09-17 | The invention relates to a method for calibrating an analyte-measurement device that is used to evaluate a concentration of analyte in bodily fluid at or from a measurement site in a body. The method involves measuring a concentration, or calibration concentration, of an analyte in blood from an “off-finger” calibration site, and calibrating the analyte-measurement device based on that calibration concentration. The invention also relates to a device, system, or kit for measuring a concentration of an analyte in a body, which employs a calibration device for adjusting analyte concentration measured in bodily fluid based on an analyte concentration measured in blood from an “off-finger” calibration site. | 1. An apparatus for use in calibrating a signal from a subcutaneous sensor, the apparatus comprising: one or more processors;a memory for storing instructions which, when executed by the one or more processors, causes the one or more processors to receive a calibration measurement from a calibration sensor placed in contact with capillary blood from an off-finger calibration site within a body, to evaluate whether the calibration measurement is within a predetermined range of analyte concentration, and to determine that the calibration measurement is suitable for use in converting an analyte signal from the subcutaneous sensor into an analyte concentration if the calibration measurement is within the predetermined range. 2. The apparatus of claim 1, wherein the analyte is glucose and the predetermined range is from about 60 mg/dL to about 350 mg/dL. 3. The apparatus of claim 1, wherein the analyte is glucose. 4. The apparatus of claim 1, wherein the calibration site is located in an arm of the body. 5. The apparatus of claim 1, wherein the calibration site is located in a leg of the body. 6. The apparatus of claim 1, wherein the subcutaneous sensor is located in an arm of the body. 7. The apparatus of claim 1, wherein the subcutaneous sensor is located in the abdomen of the body. 8. The apparatus of claim 1, wherein the subcutaneous sensor is located in one region of the body and the calibration site is located in another region of the body. 9. The apparatus of claim 1, wherein each of the subcutaneous sensor and the calibration site is located in substantially one region of the body. 10. The apparatus of claim 1, wherein the subcutaneous sensor is sufficient for electrochemically determining a concentration of an analyte. 11. The apparatus of claim 1, wherein the instructions for receiving the calibration measurement comprises instructions for determining the calibration measurement electrochemically. 12. The apparatus of claim 1, wherein the instructions for receiving the calibration measurement comprises instructions for determining the calibration measurement in less than or equal to about 1 μL of blood. 13. The apparatus of claim 1, wherein the instructions for receiving the calibration measurement comprises instructions for determining the calibration measurement in less than or equal to about 0.5 μL of blood. 14. The apparatus of claim 1, wherein the instructions for receiving the calibration measurement comprises instructions for determining the calibration measurement in less than or equal to about 0.2 μL of blood. 15. The apparatus of claim 1, wherein the instructions for receiving the calibration measurement comprises instructions for determining the calibration measurement within about five minutes after the subcutaneous sensor is inserted at a measurement site. 16. The apparatus of claim 1, wherein the instructions for receiving the calibration measurement comprises instructions for determining the calibration measurem... | Abbott Diabetes Care Inc. | A61B005/1473 A61B005/145
US20110004117 | A1 | 2011-01-06 | IMPLANT PARAMETER SELECTION BASED ON COMPRESSIVE FORCE | 2010-04-26 | Systems and techniques for determining in vivo mechanical load exerted on an implanted medical device (IMD) are described. A transfer function for determining a normal force exerted by a muscle based on an in-line force and muscle parameters is also described. | 1. A method comprising: determining a parameter of a muscle;determining a first force exerted by the muscle along a first direction;with a processor, determining a second force exerted by the muscle along a second direction substantially perpendicular to the first direction based on the parameter of the muscle and the first force; andselecting at least one implant parameter for an implantable medical device based on the second force. 2. The method of claim 1, wherein the first direction corresponds to a line of contraction of the muscle and the second direction is substantially perpendicular to the first direction. 3. The method of claim 1, wherein the first direction substantially extends between an origin of the muscle and a point of origin of the muscle. 4. The method of claim 1, wherein the second force comprises a compressive force exerted by the muscle. 5. The method of claim 1, wherein determining the first force comprises, with the processor, determining the first force based on an electromyogram signal of the muscle. 6. The method of claim 1, wherein the parameter of the muscle comprises at least one of a muscle volume, a muscle length, a muscle width, a muscle thickness, a physiologic cross-sectional area of the muscle, or a mass of the muscle. 7. The method of claim 1, wherein the parameter of the muscle comprises a geometric parameter. 8. The method of claim 1, wherein the implant parameter comprises at least one of an implant site within a patient for the implantable medical device or a design configuration of the implantable medical device. 9. The method of claim 8, wherein the design configuration comprises at least one of a housing thickness, shape, material or mass. 10. The method of claim 1, wherein determining the second force exerted by the muscle comprises determining the second force exerted by each of a plurality of muscles, and selecting at least one implant parameter for an implantable medical device based on the second force comprises: comparing the second force for each of the muscles; andselecting an implant site associated with the muscle that exerts a lowest relative second force. 11. A method comprising: determining a compressive force exerted by a muscle based on a transfer function that indicates a relationship between the compressive force and an in-line muscle force and at least one muscle parameter; andselecting at least one implant parameter for an implantable medical device based on the compressive force. 12. The method of claim 11, wherein the transfer function indicates that the compressive force substantially equals: 77.6+0.264*the in line force (pounds)−1.44*a length of the muscle (inches)−18.2*a width of the muscle (inches)−139*a thickness of the muscle (inches)+27.5*a physiologic cross-sectional area of the muscle (square inches)−0.139*a volume of the muscle (cubic centimeters)+0.110*a mass of the muscle (grams)+0.0326. 13. The method of claim 11, wherein the implant parameter comprises at l... | Medtronic, Inc. | A61B005/0488 A61B005/103
US07865374 | B2 | 2011-01-04 | Computerized system and method for verifying authority to modify clinical orders | 2006-02-22 | Computerized methods for verifying authority to modify clinical orders are provided. The methods may include receiving a clinical order having components associated with a plurality of phases; associating a plurality of the components to form a cross-phase group, the components including components associated with a first phase and a second phase, each of the components including a common attribute; receiving an indication to modify the common attribute of the component associated with the first phase; and determining whether a user from whom the indication is received has the authority to modify the common attribute. If it is determined that the user has the authority, the method may further include modifying all components that are associated in the group. If it is determined that the user does not have the authority, the method may further include not modifying the common attribute. Computerized systems for performing the disclosed methods are also provided. | 1. A method in a clinical computing environment for verifying authority to modify at least one clinical order, the method comprising: receiving, through a first computing process, the at least one clinical order, the at least one clinical order having components that are administered in two or more phases, wherein a phase is a sub-plan within a healthcare plan, wherein individual phases are separated from other phases by at least one day; and wherein a healthcare plan includes multiple orders for the treatment of a particular ailment in a patient;associating, through a second computing process, a plurality of the components to form a cross-phase group, the plurality of the components including at least a component associated with a first phase and a component associated with a second phase, wherein each of the components in the cross-phase group includes a common attribute, wherein the common attribute is a clinical detail related to administration of the at least one clinical order to the patient, and wherein the first phase and the second phase occur at least one day apart;receiving, through a third computing process, a user indication to modify the common attribute of the component associated with the first phase; anddetermining, through a fourth computing process, whether a user from whom the user indication is received has the authority to modify the common attribute of the component associated with the first phase; andmodifying, through a fifth computing process and in response to the user indication, the common attribute of all components in the cross-phase group in the same way, wherein the first, second, third, fourth, and fifth computing processes are performed on one or more computing devices. 2. The method of claim 1, further comprising displaying, through a sixth computing process, a message indicating that the common attribute of the component associated with the first phase will be modified in response to the indication received, wherein the sixth computing process is performed on the one or more computing devices. 3. The method of claim 1, wherein the first phase and the second phase are separated by a period of time greater than six days. 4. The method of claim 1, wherein the common attribute a route of administration for a medication. 5. The method of claim 3, wherein the common attribute is a medication dosage. 6. The method of claim 3, wherein the common attribute is an administration frequency for a medication. 7. A method in a clinical computing environment for verifying authority to modify at least one clinical order, the method comprising: receiving, through a first computing process, the at least one clinical order, the at least one clinical order having components associated with a plurality of phases, wherein a phase is one or more clinical orders occurring within a period of time;associating, through a second computing process, a plurality of the components to form a cross-phase group, the plurality of the compon... | Cerner Innovation, Inc. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07949399 | B2 | 2011-05-24 | Acquiring nerve activity from carotid body and/or sinus | 2010-10-29 | An exemplary includes acquiring an electroneurogram of the right carotid sinus nerve or the left carotid sinus nerve, analyzing the electroneurogram for at least one of chemosensory information and barosensory information and calling for one or more therapeutic actions based at least in part on the analyzing. Therapeutic actions may aim to treat conditions such as sleep apnea, an increase in metabolic demand, hypoglycemia, hypertension, renal failure, and congestive heart failure. Other exemplary methods, devices, systems, etc., are also disclosed. | 1. A method comprising: identifying one or more regions of neurons of the right carotid sinus nerve or the left carotid sinus nerve as being associated with at least one member of the group consisting of ventilation rate, blood carbon dioxide concentration, blood oxygen concentration and blood pH; selecting a first electrode configuration from an electrode array to sense a first type of nerve activity of the right carotid sinus nerve or the left carotid sinus nerve and selecting a second electrode configuration to sense a second type of nerve activity of the right carotid sinus nerve or the left carotid sinus nerve, wherein the first type of nerve activity is different from the second type of nerve activity; acquiring an electroneurogram of the one or more regions; storing the electroneurogram in an implantable device, wherein the implantable device comprises a memory, and wherein the electroneurogram is stored in the memory of the implantable device; communicating the stored electroneurogram to another device; and analyzing the electroneurogram to determine least one member of the group consisting of ventilation rate, blood carbon dioxide concentration, blood oxygen concentration and blood pH. 2. The method of claim 1 further comprising calling for one or more therapeutic actions based at least in part on the analyzing. 3. The method of claim 1 further comprising instructing the implantable device to call for one or more therapeutic actions based at least in part on the analyzing. 4. The method of claim 1 wherein the analyzing comprises comparing at least one of the chemosensory information and the barosenory information to historic information. 5. The method of claim 1 wherein the analyzing comprises use of at least one established relationship between carotid sinus nerve activity and blood pressure. 6. The method of claim 1 wherein the analyzing comprises use of at least one established relationship between carotid sinus nerve activity and blood chemistry. 7. The method of claim 1 further comprising making a diagnosis based at least in part on the analyzing. 8. The method of claim 7 wherein the diagnosis comprises a member selected from a group consisting of sleep apnea, an increase in metabolic demand, hypoglycemia and hypertension. 9. The method of claim 1, wherein the first electrode configuration senses nerve activity at a first location along the right carotid sinus nerve or the left carotid sinus nerve and the second electrode configuration senses nerve activity at a second location along the right carotid sinus nerve or the left carotid sinus nerve, wherein the first and second locations are different. 10. A method comprising: identifying one or more regions of neurons of the right carotid sinus nerve or the left carotid sinus nerve as being associated with at least one member of the group consisting of ventilation rate, blood carbon dioxide concentration, blood oxygen concentration and blood pH; selecting a first electrode con... | Pacesetter, Inc. | A61N001/08
US07892181 | B2 | 2011-02-22 | Apparatus for monitoring intra-abdominal pressure | 2005-09-01 | An improved apparatus for monitoring the intra-abdominal pressure of a hospitalized patient includes a urinary catheter connected to a urine valve having selectable communication positions between a discharge end of the urinary catheter and either a drain or a fluid source. Preferably, the urine valve has a housing adapted to resist patient discomfort from leg-valve contact. One operable protective housing may be embodied as a separate tray component. Plumbing structure desirably maintains fluid supply and drain conduits in a substantially parallel arrangement to assist routing those conduits between a patient's legs. When the urine valve is oriented for communication to the fluid source, an infusion pump may be used to introduce a known quantity of fluid through the urine valve and into the patient's bladder where the fluid's pressure can be measured. Desirably, a double check valve is included in a fluid supply path and arranged to permit repetitive operation of a syringe to inject a bolus of fluid into the patient's bladder. Subsequent to making a pressure measurement, the urine valve is returned to the bladder draining position. | 1. An apparatus for measuring hydraulic pressure in the bladder of a medical patient comprising: a catheter adapted for draining urine from said patient into a drain conduit;a container of fluid;a first fluid path between said container of fluid and a discharge end of said catheter;a pump disposed to effect a pressure on fluid in said first fluid path;a pressure transducer arranged to measure a pressure of said fluid in said first fluid path at a location downstream of said pump; anda drain valve defining, at most, two fluid flow paths there-through, each such flow path extending from a valve housing entrance to a valve housing exit and being configurable to permit uninterrupted flow of fluid the entire length there-between, the drain valve being operable to occlude said drain conduit while permitting fluid flow through said first fluid path, said valve being configured to remain in association with said drain conduit when said valve is in a configuration to permit fluid flow through said drain conduit, the valve comprising structure configured to cause routing of a length of the first path, on the container-side of the valve and in the vicinity of the valve, to be disposed substantially in parallel with both of: a first length of the drain conduit at a valve entrance, and;a second length of the drain conduit at a valve exit. 2. The apparatus of claim 1, said valve comprising: a multi-way valve comprising: a first flow portion disposed in said first fluid path and arranged to receive fluid from said container;a second flow portion disposed in a second fluid path configured as a drain for said catheter; anda third flow portion disposed in said first fluid path for fluid communication with a urine discharge end of said catheter;wherein:said multi-way valve is operable selectively to resist fluid flow between said third flow portion and said second flow portion; andsaid multi-way valve is configured to resist access of ambient air into the second flow portion during conventional operation of the valve. 3. The apparatus of claim 2, wherein: a urine draining lumen forming a flow path through said valve has a substantially uniform diameter in excess of about 3/16 inches (4.8 mm), and;a sealing element of said multi-way valve is structured to contain a dead volume of less than about 0.0006 cubic inches (0.01 ml). 4. The apparatus of claim 1, said drain valve further comprising: a proximal body portion comprising a single valve entrance port through which fluid may communicate on a path extending through said valve, a first alignment structure, and an axial connection structure;a distal body portion structured in harmony with said proximal body portion to permit relative motion there-between subsequent to assembly of said valve, said distal body portion comprising: a second alignment structure disposed for cooperation with said first alignment structure to indicate a valve fluid-flow configuration;a first aperture in fluid communication with a first d... | AbViser Medical LLC | A61B005/00
US08121972 | B2 | 2012-02-21 | Reduction of classification error rates and monitoring system using an artificial class | 2010-10-14 | Systems and methods for enhancing the accuracy of classifying a measurement by providing an artificial class. Seizure prediction systems may employ a classification system including an artificial class and a user interface for signaling uncertainty in classification when a measurement is classified in the artificial class. | 1. A seizure advisory system, comprising: a user interface, the user interface comprising: a first interface configured to provide an indication corresponding to a result of classification of data from a subject corresponding to the subject's susceptibility to a seizure;a second interface configured to provide an indication of uncertainty when a result of classification of data from a subject is indicative of uncertainty in the classification; anda third interface configured to receive an input from the subject. 2. The system of claim 1, wherein the input from the subject indicates that the subject has taken a medication. 3. The system of claim 1, wherein the input from the subject indicates that the subject has experienced a seizure. 4. The system of claim 1, further comprising: an implantable device configured to measure physiological signals from the subject;wherein the user interface is provided on a device external to the subject's body. 5. The system of claim 4, wherein the implantable device comprises a classification processor for generating the result of classification of data from the subject. 6. The system of claim 4, wherein the external device comprises a classification processor for generating the result of classification of data from the subject. 7. The system of claim 1, wherein the input from the subject comprises an audio input. 8. The system of claim 1, further comprising an implantable device configured to measure physiological signals from the subject. 9. The system of claim 8, wherein the implantable device is configured to measure brain activity signals. 10. The system of claim 8, wherein the implantable device is configured to measure blood oxygenation. 11. The system of claim 8, wherein the implantable device is configured to measure ECG/EKG. 12. The system of claim 1, further comprising an implantable device configured to measure physiological signals from the subject. 13. The system of claim 12, wherein the implantable device comprises subcutaneous electrodes. 14. The system of claim 12, wherein the implantable device comprises intracranial electrodes. 15. The system of claim 12, wherein the implantable device comprises a rechargeable battery. 16. A method of providing a seizure advisory to a subject, the method comprising: providing via a first interface an indication corresponding to a result of classification of data from the subject corresponding to the subject's susceptibility to a seizure;providing via a second interface an indication of uncertainty when a result of classification of data from a subject is indicative of uncertainty in the classification; andreceiving an input from the subject via a third interface. 17. The method of claim 16, wherein said receiving the input from the subject comprises receiving an input from the subject indicating that the subject has taken a medication. 18. The method of claim 16, wherein said receiving the input from the subject comprises receiving an input from... | NeuroVista Corporation | G06F009/44 G06N007/02 G06N007/06
US07867611 | B2 | 2011-01-11 | Biomedical sensors employing hydro-insensitive alternating current responsive composites | 2010-03-31 | A biomedical sensor is disclosed that includes a conductive material for coupling to monitoring equipment, and a composite. The composite includes a polymeric material and a polar material that is substantially dispersed within the polymeric material. The composite has a first side that is coupled to the conductive material and has a second side that is positionable with respect to a subject to be monitored. The polar material exhibits molecular compatibility with the polymeric material such that the polar material neither blooms to a surface of the polymeric material nor crystallizes within the polymeric material. | 1. A biomedical sensor comprising a conductive material for coupling to monitoring equipment; and a composite including a polymeric material and a polar material that is substantially dispersed within the polymeric material, said composite having a first side that is coupled to said conductive material and having a second side that is positionable with respect to a subject to be monitored, wherein said polar material exhibits molecular compatibility with the polymeric material such that the polar material neither blooms to a surface of the polymeric material nor crystallizes within the polymeric material. 2. The biomedical sensor as claimed in claim 1, wherein said polymeric material is an adhesive. 3. The biomedical sensor as claimed in claim 1, wherein said polymeric material is an acrylic adhesive. 4. The biomedical sensor as claimed in claim 1, wherein said polymeric material is a pressure sensitive adhesive. 5. The biomedical sensor as claimed in claim 1, wherein said polar material is a salt. 6. The biomedical sensor as claimed in claim 1, wherein said polar material is an organic salt. 7. The biomedical sensor as claimed in claim 1, wherein said composite includes a concentration of said polar material within said polymeric material of between about 5% and about 65% based on weight. 8. The biomedical sensor as claimed in claim 1, wherein said composite includes a concentration of said polar material within said polymeric material of between about 10% and about 35% based on weight. 9. The biomedical sensor as claimed in claim 1, wherein said composite has a DC resistance of at least about 100 k Ohms. 10. The biomedical sensor as claimed in claim 1, wherein said composite has a dielectric constant that changes in the presence of and changes with an alternating electric field. 11. A biomedical sensor comprising a conductive material including an access node for coupling to monitoring equipment; and a non-ionically conductive composite that includes a polymeric material and a polar material that is substantially dispersed within the polymeric material, said composite having a first side that is coupled to said conductive material, wherein said polar material exhibits molecular compatibility with the polymeric material such that the polar material neither blooms to a surface of the polymeric material nor crystallizes within the polymeric material such that said non-ionically conductive composite provides a non-ionically conductive coupling between a subject and said conductive material. 12. The biomedical sensor as claimed in claim 11, wherein said polymeric material is an acrylic adhesive. 13. The biomedical sensor as claimed in claim 11, wherein said polymeric material is a pressure sensitive adhesive. 14. The biomedical sensor as claimed in claim 11, wherein said polar material is an organic salt. 15. The biomedical sensor as claimed in claim 11, wherein said composite includes a concentration of said polar material within said... | FLEXcon Company, Inc. | A61B005/04 A61B005/0402 H01B001/12 C09J133/08 B32B027/18
US08145301 | B2 | 2012-03-27 | Blending cardiac rhythm detection processes | 2010-09-30 | Systems and methods are described for classifying a cardiac rhythm. A cardiac rhythm is classified using a classification process that includes a plurality of cardiac rhythm discriminators. Each rhythm discriminator provides an independent classification of the cardiac rhythm. The classification process is modified if the modification is likely to produce enhanced classification results. The rhythm is reclassified using the modified classification process. | 1. A method for classifying a tachyarrhythmia in a heart, the method comprising: sensing, at a patient device, one or more sequences of atrial and ventricular contractions associated with the tachyarrhythmia;selecting, at the patient device, one or more patterns that represent an onset pattern, the patterns including respective sequences of atrial and ventricular contractions; andobtaining, at the patient device, a first classification of the tachyarrhythmia based on whether a match exists between the one or more sequences of atrial and ventricular contractions and the one or more patterns. 2. The method in claim 1, comprising: detecting whether the tachyarrhythmia is sudden onset;determining whether an atrial rate is approximately equal to a ventricular rate; andif the tachyarrhythmia is sudden onset and the atrial rate is approximately equal to the ventricular rate, then classifying the tachyarrhythmia using whether the match exists. 3. The method of claim 1, wherein the obtaining the first classification comprises: identifying a sequence of atrial and ventricular contractions associated with the tachyarrhythmia; andclassifying the sequence as more likely indicative of ventricular tachyarrhythmia if two or more ventricular contractions occur between consecutive atrial contractions within the sequence. 4. The method of claim 1, wherein the obtaining the first classification comprises: identifying a sequence of atrial and ventricular contractions associated with the tachyarrhythmia; andclassifying the sequence as more likely indicative of supraventricular tachyarrhythmia if two or more atrial contractions occur between consecutive ventricular contractions within the sequence. 5. The method of claim 1, wherein the obtaining the first classification comprises: identifying a sequence of atrial and ventricular contractions associated with the tachyarrhythmia; andclassifying the sequence as more likely indicative of ventricular tachyarrhythmia if two or more consecutive ventricular contractions occur between consecutive atrial contractions without having two or more consecutive atrial contractions occur between consecutive ventricular contractions within the sequence. 6. The method of claim 1, comprising: establishing a matching threshold value using a result of the first classification; andperforming a comparison using the established matching threshold value to obtain a second classification of the tachyarrhythmia. 7. The method of claim 6, wherein the comparison is a morphology analysis. 8. The method of claim 6, comprising: determining whether a therapy is necessary;choosing a therapy using the second classification if a therapy is deemed necessary; anddelivering the therapy. 9. The method of claim 6, wherein the establishing the matching threshold comprises at least one of: raising or maintaining the matching threshold if the first classification is ventricular tachycardia; andlowering or maintaining the matching threshold if the first cl... | Cardiac Pacemakers, Inc. | A61N001/39 A61N001/362
US07865223 | B1 | 2011-01-04 | In vivo blood spectrometry | 2005-03-14 | A process and apparatus for determining the arterial and venous oxygenation of blood in vivo with improved precision. The optical properties of tissue are measured by determination of differential and total attenuations of light at a set of wavelengths. By choosing distinct wavelengths and using the measured attenuations, the influence of variables such as light scattering, absorption and other optical tissue properties is canceled out or minimized. | 1. An apparatus for determining blood oxygenation in a tissue, the apparatus comprising: a sensor interface having a first emitter and a second emitter configured to emit light into the tissue and having a first detector and a second detector configured to detect light from the tissue, wherein an emitter-detector distance is selected to provide a light path that travels through the tissue at a depth of penetration that includes a selected region of the tissue;a processor configured to:determine light attenuation LA(A1,wsj) of light emitted by the first emitter and received by the first detector;determine light attenuation LA(A4,wsj) of light emitted by the second emitter and received by the second detector;determine light attenuation LA(A2,wsj) of light emitted by the first emitter and received by the second detector;determine light attenuation LA(A3,wsj) of light emitted by the second emitter and received by the first detector; anddetermine an optical characteristic of said tissue based on a resulting light attenuation LAwsj at said wavelength wsj, the resulting light attenuation determined by summing the light attenuations of LA(A2,wsj) and LA(A3,wsj) and deducting the light attenuations LA(A1,wsj) and LA(A4,wsj), the optical characteristic representative of oxygenation and corresponding to oxygenation of venous blood in tissue of said selected; andan output device coupled to the processor and configured to receive the optical characteristic. 2. The apparatus of claim 1, wherein said sensor interface is configured to be coupled to the tissue by at least one of a fixation band, a sensor holder, medical glue, a finger glove, and a finger clip. 3. The apparatus of claim 2, further including a battery coupled to at least one of the sensor interface, the processor, and the output device. 4. The apparatus of claim 1, wherein said output device includes a wrist worn display. 5. The apparatus of claim 4, wherein at least one of the first emitter and the second emitter includes an LED. 6. The apparatus of claim 4, wherein at least one of the first detector and the second detector includes a photodiode detector. 7. The apparatus of claim 1 wherein the processor is configured to repeatedly determine light attenuations and repeatedly determine the optical characteristic. 8. The apparatus of claim 1 wherein the processor is configured to determine a weight for at least one of light attenuation LA(A1,wsj), light attenuation LA(A2,wsj), light attenuation LA(A3,wsj), and light attenuation LA(A4,wsj), wherein the weight is selected to reduce an error due to tissue inhomogeneity. 9. The apparatus of claim 1 wherein the processor is configured to determine light attenuation using a logarithm of light intensity. 10. Apparatus for measuring tissue oxygenation comprising: a sensor interface adapted to be coupled to a tissue site and including at least two emitters placed apart from each other on a surface of said interface and configured to emit light of at... | - | A61B005/1455
US07879058 | B2 | 2011-02-01 | Lancet device for forming incision | 2005-04-22 | An incision device capable of preventing inadvertent firing, and including a covered attached to the device. The incision device comprises a housing, and a lancet guide assembly having lancet guides and a lancet body. The lancet body further comprises a lancet blade. The lancet blade is covered by a removable protective cover having a stopper which prevents a trigger arm from being moved so as to fire the lancet blade. | 1. An incision device for forming an incision, the incision device comprising: (1) a housing including a first opening;(2) a lancet guide assembly including a lancet body and a lancet guide, the lancet body being mounted within the housing for pivotal movement and comprising a lancet blade covered by a removable protective cover, and the lancet guide being provided for guiding the lancet blade through an arcuate cutting stroke;(3) a trigger mechanism including a trigger arm mounted in the housing for pivotal movement between a ready-to-fire position and a firing position, the trigger arm having a trigger actuating element; and(4) a spring element arranged within the housing so as to be in an unbiased condition when the trigger arm is in the ready-to-fire position and so as to be biased by engaging a catch arranged in the housing as the trigger arm pivots towards the firing position,wherein the housing and lancet guide assembly are configured such that during the cutting stroke, the lancet blade extends through the first opening in the housing when the lancet body is pivoted between an unfired position and a fired position;wherein the trigger mechanism and the spring element are configured such that the spring element in the biased condition is released from the catch arranged in the housing when the trigger arm reaches the firing position;wherein the spring element is arranged such that when released from the catch the spring element exerts a force on the lancet body which pivots the lancet body from the unfired position to the fired position and thereby moves the lancet blade through the cutting stroke; andwherein the removable protective cover comprises a stopper which prevents the trigger arm from moving to reach the firing position when the protective cover covers the lancet blade by directly contacting the trigger actuating element. 2. The incision device according to claim 1 wherein each of the trigger mechanism and the lancet body are arranged independently in the housing. 3. The incision device according to claim 1 wherein each of the lancet body and the trigger arm pivots respectively. 4. The incision device according to claim 3 wherein the lancet body pivots about a first pivot axis at an end of the lancet body and a first bearing portion arranged in the housing which corresponds to the first pivot axis, and wherein the trigger arm pivots about a second pivot axis at an end of the trigger arm and a second bearing portion arranged in the housing which corresponds to the second pivot axis. 5. The incision device according to claim 1 wherein the trigger mechanism and the lancet body are integrally formed by a connecting portion located between them, and a notch is provided between the trigger mechanism and the connecting portion and a notch is provided between the connecting portion and the lancet body, the notches being configured to substantially allow pivot movements of the trigger mechanism and the lancet body. 6. The incision dev... | Asahi Polyslider Company, Limted | A61B017/14 A61B017/32 A61B005/00
US07878988 | B2 | 2011-02-01 | Method for measuring the strength of healing bone and related tissues | 2006-10-06 | A method of measuring healing strength of bone includes a bone fixation plate (14) implanted at the healing area between bone segments (12). The bone fixation plate (14) includes high aspect ratio material with electric current conduction properties responsive to strain. This method causes an induced, or directly applied, electric current to pass through the high aspect ratio material. Analysis of the change in this current by a series of increasing loads placed on the healing body part indicates what level of load produces strain on the bone fixation plate (14). As healing strength increases, evidence of strain on the bone fixation plate (14) occurs at a greater load. Physicians determine thereby the strength of bone healing and safe levels of activity for patients while bone healing progresses. | 1. A method of measuring bone healing strength on a bone area being studied, comprising: a) implanting a bone fixation device constituted at least in part of a material selected from the group consisting of carbon nanotubes and other high aspect ratio materials which alter the conduction of electric current when said bone fixation device is subjected to stress or strain; b) applying externally a testing cuff requiring no direct electric contacts to said bone fixation device, wherein said testing cuff provides a means to cause an electric current in said bone fixation device, and to measure said electric current; c) using an analyzer to study changes in said electric current resulting from the effects of stress and strain on the material of the bone fixation device; d) obtaining a series of test results using predetermined loads on the bone area; and e) determining bone healing strength. 2. The method of claim 1 wherein an anchoring mechanism attaching said bone fixation device to said bone area contains a nanotube coating whereby loosening from said bone area will increase resistance to said electric current compared to a baseline reading. 3. The method of claim 1 wherein said bone fixation device is selected from the group of fixation devices including plates, pins, wires, rods, clamps and prostheses. 4. The method of claim 1 wherein said testing cuff includes an active coil with a means to produce a primary induced electric current in said bone fixation device. 5. The method of claim 4 wherein said testing cuff includes a passive coil with a means to produce a secondary induced electric current corresponding to said primary induced electric current. 6. The method of claim 1 further comprising: a) measuring said electric current while said bone area is at total rest; b) measuring said electric current with a plurality of predetermined loads on said bone area; and c) determining bone healing strength by noting a minimum load producing strain on the bone fixation device. 7. The method of claim 1 wherein said analyzer sounds an alarm when test results indicate strain on said bone fixation device; a) said analyzer having the capability of displaying test results; b) said analyzer having the capability of storing said test results; and c) said analyzer having the capability of transmitting said test results. 8. The method of claim 6 wherein said predetermined loads include weight bearing by the patient. 9. The method of claim 1 wherein said predetermined loads include a continuous, gradual increase in effort until an alarm sounds, whereby the patient will know a safe limit of exertion. | - | A61B005/103 A61B017/80
US09131859 | B2 | 2015-09-15 | Blood pressure measurement apparatus, recording medium that records blood pressure derivation program, and blood pressure derivation method | 2009-03-13 | A blood pressure measurement apparatus includes a derivation control unit that performs derivation control to derive a blood pressure of a measured person. The derivation control unit includes an envelope extracting unit to extract an envelope of a volume pulse wave based on an arterial volume signal, a differential processing unit to differentiate the envelope with respect to a cuff pressure, a maximum value extracting unit to extract a maximum value of a differential value of the envelope, and a blood pressure decision unit that decides the cuff pressure used for the differential of the maximum value as a blood pressure. | 1. A blood pressure measurement apparatus, comprising: a cuff to be wound around a predetermined physical site of a measured person;a pressure detector to detect a cuff pressure representing a pressure inside the cuff;a volume detector to detect an arterial volume signal indicating a volume of an artery of the measured person;an adjustment unit to adjust the pressure inside the cuff;a drive control unit to control driving of the adjustment unit; anda derivation control unit to perform derivation control to derive a blood pressure of the measured person, based on the cuff pressure and the arterial volume signal, the derivation control unit including: a first extraction unit to extract an envelope of a volume pulse wave based on the arterial volume signal;a differential unit to differentiate the envelope with respect to the cuff pressure;a second extraction unit to extract a maximum value of a differential value of the envelope' anda decision unit to measure the blood pressure based on the cuff pressure obtained when the maximum value of the differential value is extracted,wherein the envelope includes a volume minimum point envelope in contact with arterial volume minimum points of respective pulse wave components included in the volume pulse wave, and a volume maximum point envelope in contact with arterial volume maximum points of the respective pulse wave components included in the volume pulse wave,wherein the differential unit differentiates the volume minimum point envelope, and the volume maximum point envelope,wherein the second extraction unit extracts a maximum value of a differential value of the volume minimum point envelope, and a maximum value of a differential value of the volume maximum point envelope,wherein the derivation control unit performs the derivation control when the cuff pressure is being decreased or increased at a constant rate by the drive control unit, andwherein the decision unit measures, as a diastolic blood pressure, the cuff pressure obtained when the maximum value of the differential value of the volume minimum point envelope is extracted by the second extraction unit, and as a systolic blood pressure, the cuff pressure obtained when the maximum value of a differential value of the volume maximum point envelope is extracted by the second extraction unit. 2. A blood pressure measurement apparatus, comprising: a cuff to be wound around a predetermined physical site of a measured person;a pressure detector to detect a cuff pressure representing a pressure inside the cuff;a volume detector to detect an arterial volume signal indicating a volume of an artery of the measured person;an adjustment unit to adjust the pressure inside the cuff;a drive control unit to control driving of the adjustment unit; anda derivation control unit to perform derivation control to derive a blood pressure of the measured person, based on the cuff pressure and the arterial volume signal, the derivation control unit including: a first ex... | OMRON HEALTHCARE Co., Ltd. | A61B005/02 A61B005/0225 A61B005/00 A61B005/026
US08623327 | B2 | 2014-01-07 | Imaging agents for use in magnetic resonance blood flow/perfusion imaging | 2007-06-19 | In one aspect, a method of imaging blood flow in a region of interest of a subject using magnetic resonance imaging (MRI) is provided. The method includes introducing an imaging agent into the subject, the imaging agent including a compound having at least one hyperpolarized nucleus having a T.sub.1 greater than 15 seconds and a water-octanol partition coefficient between −1 and 1, providing at least one excitation signal to the region of interest, the at least one excitation signal configured to invoke a nuclear magnetic resonance (NMR) effect at least in the introduced imaging agent, and detecting an NMR signal emitted by the region of interest in response to the at least one excitation signal. In another aspect, the imaging agent includes a carbon-13 enriched alcohol. | 1. A method of obtaining nuclear magnetic resonance information for use in imaging blood flow in a region of interest of a subject using magnetic resonance imaging (MRI), the method comprising: introducing an imaging agent into the subject, the imaging agent comprising a compound having at least one hyperpolarized nucleus having a T1 greater than 15 seconds and a log water-octanol partition coefficient between and including −1 and 1, wherein the compound is capable of substantial penetration of the blood-brain barrier, and wherein the compound is carbon-13 enriched tert-butyl alcohol (t-butanol) or carbon-13 enriched 2-methyl-2-butanol;providing at least one excitation signal to the region of interest, the at least one excitation signal configured to invoke a nuclear magnetic resonance (NMR) effect at least in the introduced imaging agent; anddetecting an NMR signal emitted by the region of interest in response to the at least one excitation signal. | Beth Israel Deaconess Medical Center, Inc. | A61B005/055
US09033901 | B2 | 2015-05-19 | System for determining individual user anthropometric characteristics related to mattress preference | 2010-09-20 | A measuring device to aid mattress selection has an upstanding frame with a plurality of vertically space-apart horizontal illuminated strips. A camera mounted in front of the frame takes front view and side view images of a person standing in front of the illuminated strips. These images are delivered to an associated computer having means to determine the body mass distribution and the spine curvature of the person based on the images of the strips taken by the camera and in particular based on those portions of the image of the illuminated strips blocked by the person standing in front of the illuminated strips. This information can then be used to aid in appropriate mattress selection for the person. | 1. A measuring device for determining individual user anthropometric characteristics related to mattress preference to aid mattress selection, including: a support frame,a plurality of illumination strips mounted spaced-apart on the support frame,a camera mounted in front of the frame having a lens directed at the illumination strips with said illumination strips in its view,a computer connected to the camera to receive images of the illumination strips from said camera,said computer having means to determine one or more anthropometric characteristics of a person standing between the camera and the illumination strips in a front view and in a side view based on images of the illumination strips taken by the camera when said illumination strips are illuminated,said computer being operable to determine said one or more anthropometric characteristics based on those portions of images of the illuminated strips blocked by the person standing between the camera and the illuminated strips in said front view and said side view and the distance between the illumination strips,said anthropometric characteristics including body mass distribution determined by delineation of a series of polygons characterising the body shape of said person. 2. The measuring device as claimed in claim 1 wherein said anthropometric characteristics include spine curvature of the person determined by calculating the radius of curvature by taking chords between three vertically adjacent points on the side view image of the illumination strips. 3. The measuring device as claimed in claim 1 wherein the illumination strips are operable to generate light of a particular colour and the camera is adapted to be preferentially sensitive to said colour. 4. The measuring device as claimed in claim 1 wherein the illumination strips are adapted to blink at a specified frequency and the computer is operable to compare an image received from the camera when the illumination strips are illuminated with an image received from the camera when the illumination strips are not illuminated in order to eliminate extraneous detail from the image resulting from ambient light. 5. The measuring device as claimed in claim 3 wherein the illumination strips are adapted to blink at a specified frequency and the computer is operable to compare an image received from the camera when the illumination strips are illuminated with an image received from the camera when the illumination strips are not illuminated in order to eliminate extraneous detail from the image resulting from ambient light. 6. The measuring device as claimed in claim 1 wherein the support frame is an upstanding frame and a plurality of horizontal illumination strips are mounted vertically spaced-apart on the support frame. 7. The measuring device as claimed in claim 1 wherein the camera is a personal computer web cam connected to a computer via USB. 8. The measuring device as claimed in claim 1 wherein the illumination strips comprise e... | - | A61B005/107 G01F017/00 G01B011/28 A61B005/00
US08636667 | B2 | 2014-01-28 | Systems and methods for processing physiological signals in wavelet space | 2009-07-06 | Methods and systems are disclosed for analyzing multiple scale bands in the scalogram of a physiological signal in order to obtain information about a physiological process. An analysis may be performed to identify multiple scale bands that are likely to contain the information sought. Each scale band may be assessed to determine a band quality, and multiple bands may be combined based on the band quality. Information about a physiological process may determined based on the combined band. In an embodiment, analyzing multiple scale bands in a scalogram arising from a wavelet transformation of a photoplethysmograph signal may yield clinically relevant information about, among other things, the blood oxygen saturation of a patient. | 1. A method for processing a physiological signal, comprising: receiving, from a sensor, an electronic signal representative of a physiological process;using processor equipment for: transforming the electronic signal into a transformed signal based at least in part on a continuous wavelet transformation,generating a scalogram based at least in part on the transformed signal,identifying, within the scalogram, a first scale band associated with the physiological process,identifying, directly within the scalogram, a plurality of scale bands related to the first scale band,determining a band quality associated with each identified scale band,combining the identified scale bands based at least in part on the associated band qualities, anddetermining information regarding the physiological process based at least in part on the combining; andoutputting the information to an output device. 2. The method of claim 1, wherein the first scale band comprises a region of the scalogram localized around a scale associated with a pulse rate. 3. The method of claim 2, wherein each of the plurality of scale bands comprises a region of the scalogram localized around a scale associated with a multiple of the pulse rate. 4. The method of claim 1, wherein determining a band quality associated with a scale band comprises comparing a total energy in the scale band with a total energy in a region of the scalogram. 5. The method of claim 1, wherein determining a band quality associated with a scale band comprises comparing a total energy in the scale band with a total energy in another scale band. 6. The method of claim 1, wherein determining a band quality associated with a scale band comprises evaluating the consistency of an amplitude of the scalogram within the scale band. 7. The method of claim 1, wherein combining the identified scale bands based at least in part on the associated band qualities comprises calculating a weighted combination of each of the scale bands, wherein the weight associated with a scale band is based at least in part on the associated band quality. 8. The method of claim 1, wherein combining the identified scale bands based at least in part on the associated band qualities comprises concatenating a portion of each of the scale bands, wherein the portion of a scale band is based at least in part on the associated band quality. 9. The method of claim 1, further comprising removing a portion of a scale band over a time interval based at least in part on a band quality over the time interval. 10. The method of claim 1, wherein transforming the electronic signal into a transformed signal comprises: calculating a first transformed signal from an electronic signal component associated with a first wavelength; andcalculating a second transformed signal from an electronic signal component associated with a second wavelength. 11. The method of claim 10, wherein determining a band quality associated with a scale band comprises correlating a sca... | Nellcor Puritan Bennett Ireland | A61B005/02 A61B005/024 A61B005/1455
US08992431 | B2 | 2015-03-31 | Method, apparatus and computer program for non-invasive blood pressure measurement | 2009-07-09 | A method, apparatus and computer program product are disclosed for non-invasively determining blood pressure related parameters of a subject. A cuff attached to a subject is inflated to a target pressure which is such that normal blood pressure oscillation of the subject appears in an output signal of a pressure sensor monitoring the pressure of the cuff. The output signal of the pressure sensor is acquired for at most a predetermined time period, while the cuff is maintained substantially at the target pressure, thereby to obtain blood pressure waveform data for the subject, and the blood pressure related parameters are derived from the blood pressure waveform data. | 1. A method for non-invasively determining blood pressure related parameters of a subject, the method comprising: inflating a cuff attached to a subject to an initial target pressure, wherein the initial target pressure is such that normal blood pressure oscillation of the subject appears in an output signal of a pressure sensor monitoring a pressure of the cuff;acquiring the output signal of the pressure sensor in a non-invasive blood pressure (NIBP) device while the cuff is maintained substantially at the initial target pressure, thereby to obtain blood pressure waveform data for the subject, wherein the acquiring continues for at most a predetermined time period;determining whether respiratory modulation is present in the output signal by tracking amplitude extremes of pressure pulses in the blood pressure waveform data;adjusting the target pressure if respiratory modulation is not detected in the output signal of the pressure sensor;reacquiring the output signal in the NIBP device while the cuff is maintained substantially at the adjusted target pressure to obtain the blood pressure waveform data for the subject, wherein the target pressure is readjusted and the output signal is reacquired at the adjusted target pressure until respiratory modulation is detected in the output signal, wherein the reacquiring continues for at most the predetermined time period; andderiving the blood pressure related parameters in the NIBP device from the blood pressure waveform data obtained at the target pressure at which respiratory modulation is present. 2. The method according to claim 1, further comprising starting a measurement timer in the NIBP device to measure the predetermined time period; andstopping the acquiring when the measurement timer expires. 3. The method according to claim 2, further comprising performing a data quality analysis in the NIBP device, wherein the performing comprises identifying high quality waveform data from the blood pressure waveform data, and wherein the deriving includes deriving the blood pressure related parameters from the high quality waveform data. 4. The method according to claim 3, further comprising monitoring the amount of high quality waveform data that has been identified; andstopping the acquiring when the monitoring indicates that the high quality waveform data identified fulfils predetermined criteria for the deriving. 5. The method according to claim 4, further comprising detecting respiratory cycles of the subject, wherein the performing of the data analysis comprises analysing the output signal of the pressure sensor during several consecutive respiratory cycles. 6. The method according to claim 4, wherein the stopping comprises stopping the acquiring when the monitoring indicates that the high quality waveform data identified fulfils predetermined criteria, in which the predetermined criteria requires that the high quality waveform data identified originates from a predetermined number of respiratory... | General Electric Company | A61B005/022 A61B005/021 A61B005/113 A61B005/00
US20110015514 | A1 | 2011-01-20 | MEDICAL IMAGING METHOD AND SYSTEM FOR PROVIDING A FINITE-ELEMENT MODEL | 2008-02-29 | A method comprising: -providing acquisition data (34,38) of a patient, comprising acquisition data of a bone structure of the patient in a given position, providing a knowledge base (22,43,47), comprising data related to such bone structures, -determining a patient-specific position-specific finite-element model of the bone structure comprising: patient-specific three-dimensional geometry, patient-specific mechanical characteristics of the bone structure, and patient-specific position-specific mechanical load applied to the bone structure in the position. | 1. A method comprising: providing acquisition data of a patient, comprising at least acquisition data of a bone structure of the patient in a given functional position,providing a knowledge base, comprising at least data related to structures of the same type as the bone structure of the patient,determining, at least from said acquisition data, a patient-specific three-dimensional geometry of the bone structure,determining, from said acquisition data and said knowledge base, patient-specific mechanical characteristics of the bone structure,determining, at least from the acquisition data, a patient-specific position-specific mechanical load applied to the bone structure in said functional given position,assembling the three-dimensional geometry, the mechanical characteristics and the mechanical load into a patient-specific position-specific finite-element model of the bone structure. 2. Method according to claim 1 wherein the knowledge database comprises at least data related to the three-dimensional geometry of structures of the same type as the structure of the patient, and wherein the patient-specific three-dimensional geometry is determined using the knowledge base. 3. Method according to claim 1 wherein the knowledge database comprises at least data related to the mechanical characteristics of structures of the same type as the structure of the patient. 4. Method according to claim 1 wherein the knowledge database comprises at least data related to mass and/or location of a centre of mass of body segments, and wherein the patient-specific position-specific mechanical load is determined using the knowledge base. 5. A method according to claim 4, wherein acquisition data further comprise acquisition data of at least a part of the external envelope of the patient, wherein the knowledge base further comprise data related to mass and/or location of a centre of mass correlated to external envelopes. 6. A method according to claim 1, wherein providing acquisition data further comprises providing the weight of the patient. 7. A method according to claim 1, wherein the mechanical load comprises the load of the mass of the patient situated above the bone structure in said functional position. 8. A method according to claim 7 wherein the mechanical load further comprises a muscular load applied to the bone structure. 9. A method according to claim 1 wherein acquisition data of the bone structure comprises at least one bi-dimensional image of the bone structure, and wherein the knowledge base comprises at least data related to the mechanical characteristics of bone structures of the same type as the bone structure of the patient. 10. A method according to claim 1 wherein acquisition data of the bone structure comprises at least one bi-dimensional image of the bone structure, and wherein the patient-specific three-dimensional geometry is obtained by fitting a generic three-dimensional mesh of the bone structure to the image, said 3D mesh comprisi... | ECOLE NATIONAL SUPERIEURE D'ARTS ET METIERS (ENSAM | A61B005/05 G06N005/00
US20110040205 | A1 | 2011-02-17 | Treatment for Alleviating Tinnitus and Hyperacusis with Auditory Stimulation by Compensating for Hearing Loss and Loss of Non-Linear Compressions | 2008-11-10 | An auditory stimulation treatment for alleviating tinnitus and hyperacusis by compensating for hearing loss and loss of non-linear compression. Natural auditory signals are delivered correcting for hearing loss and compressive non-linearity, both determined at separate frequency bands for the individual subject. This differs from existing treatment in that synthetic “masking” signals are not delivered, but rather the natural auditory input to the subject is modified. Consequently, the method is more akin to a conventional hearing aid. In contrast to conventional hearing aids, however, perception thresholds are specifically corrected and non-linear compression is matched to the specific hearing deficit of the user. | 1. A method for alleviating tinnitus and hyperacusis, comprising the steps of: determining hearing loss and compressive loss of a subject at separate frequency bands;modifying natural acoustic signals to compensate for compressive loss and loss of hearing sensitivity of the subject at said frequency bands; anddelivering this the modified signal as auditory stimulus to the subject. 2. The method of claim 1, wherein said determining step comprises measuring hearing sensitivity with high frequency resolution audiograms. 3. The method of claim 1, wherein said determining comprises measuring compressive hearing loss of the subject with high frequency resolution distortion product otoacoustic emissions (DPOAE). 4. The method of claim 1 further comprising: applying therapeutic auditory stimulation temporarily via a conventional audio device. 5. The method of claim 4, wherein the audio device is portable. 6. The method of claim 4, wherein audio device is a portable music player. 7. The method of claim 1, further comprising the step of: applying therapeutic stimulation chronically via a hearing-aid device. 8. The method of claim 4, further comprising the step of: synthesizing sounds to deliver sound energy to compensate for specific hearing loss of the subject. 9. The method of claim 8, further comprising the step of: synthesizing sounds to deliver sound energy to compensate for specific hearing loss of the subject. 10. The method of claim 1, further comprising the step of: implementing the method during limited times of the day with at least one of natural environmental sounds delivered with a corresponding hearing aid and by modifying sound output by a conventional personal electronic device. 11. The method of claim 10, wherein the limited times of the day is during nightly sleep. 12. The method of claim 1, wherein said step of determining hearing loss and compressive loss of a subject at separate frequency bands comprises: diagnosing the subject to determine optimal parameters for each frequency band for the subject. 13. The method of claim 12, wherein the diagnosis is performed via psychometric or physiological testing. 14. The method of claim 13, wherein the psychometric testing include at least one of hearing thresholds, loudness growth, two-tone suppression and distortion products. | The City University University of Ireland | A61B005/12
US07877273 | B2 | 2011-01-25 | System and method for evaluating and providing nutrigenomic data, information and advice | 2003-01-08 | A method for minimizing an individual's predisposition for developing a disease including providing personal information of an individual, querying a first database of disease information with the personal information to determine a disease for which the individual has a predisposition for developing, wherein the query produces disease information, querying one or more second databases with at least a portion of the disease information and/or personal information for obtaining second information for reducing the risk for the individual developing the disease and communicating a result of the query of the lifestyle database to the individual. | 1. A system for determining whether consumption of an item falls within a lifestyle for minimizing an individual's predisposition for developing a disease, the system comprising: inputting means for inputting a code of a product for consumption by an individual into a wireless communication device;first communicating means for communicating the code via the wireless communications device to a remote system;receiving means for receiving the code by the remote system;determining means for determining the product associated with the code;querying means for querying a database to determine whether the product or a closely related product is contained therein, wherein said database comprises data domains comprising a scientific domain, genetic domain, and transactional domain, wherein the transactional domain comprises the code for the product;analyzing means for interpreting the data in the data domains to determine whether the consumption of the product is related to minimizing a disease state; andsecond communicating means for communicating the analysis from the remote system to the wireless communication device. 2. A computer readable medium having computer instructions provided thereon for enabling a computer to perform a method for determining whether consumption of an item falls within a lifestyle for minimizing an individual's predisposition for developing a disease, the method comprising: inputting a code of a product for consumption by an individual into a wireless communication device;communicating the code via the wireless communications device to a remote system;receiving the code by the remote system;determining the product associated with the code;querying means for querying a database to determine whether the product or a closely related product is contained therein, wherein said database comprises data domains comprising a scientific domain, genetic domain, and transactional domain, wherein the transactional domain comprises the code for the product;analyzing means for interpreting the data in the data domains to determine whether the consumption of the product is related to minimizing a disease state; andcommunicating the analysis from the remote system to the wireless communication device. 3. An application program operational on a computer system for performing a method for determining whether consumption of an item falls within a lifestyle for minimizing an individual's predisposition for developing a disease, the method comprising: inputting a code of a product for consumption by an individual into a wireless communication device;communicating the code via the wireless communications device to a remote system;receiving the code by the remote system;determining the product associated with the code;querying means for querying a database to determine whether the product or a closely related product is contained therein, wherein said database comprises data domains comprising a scientific domain, genetic domain, and transactional d... | - | A61B005/00 G06F019/00
US08249715 | B2 | 2012-08-21 | Implantable medical device configured as a pedometer | 2010-10-26 | A system includes an implantable medical device. The implantable medical device includes a control circuit and a motion sensing device. The motion sensing device is coupled to the control circuit, and the motion sensing device is configured to transmit signals to the control circuit. The control circuit is configured to identify one or more steps of a patient using the motion sensing device signal. | 1. A system comprising: an implantable cardiac rhythm management device comprising: a control circuit; anda motion sensing device coupled to the control circuit, the motion sensing device configured to transmit a signal to the control circuit;wherein the control circuit is configured to: identify a candidate step of a patient using the motion sensing device signal; andconfirm that the candidate step is an actual individual step of the patient. 2. The system of claim 1, wherein the motion sensing device includes an accelerometer. 3. The system of claim 2, wherein the accelerometer comprises a three-axis accelerometer, and wherein the control circuit is configured to classify the actual individual step as one of a step up by the patient, a step down by the patient, and a step forward by the patient. 4. The system of claim 2, wherein the control circuit comprises a circuit to identify one or more actual individual steps of the patient by one or more of identifying a peak in an output of the accelerometer and by pattern matching an output of the accelerometer. 5. The system of claim 1, comprising an external device communicatively coupled to the implantable cardiac rhythm management device. 6. The system of claim 5, wherein the external device is a local external device configured to communicate with a remote external device. 7. The system of claim 5, wherein at least one of the implantable cardiac rhythm management device and the external device are configured to determine, using data from one or more actual individual steps of the patient, at least one of: a count of the one or more steps of the patient;a count of the one or more steps of the patient during a particular period of time;a physical activity category;a distance traveled by the patient during a particular period of time;an amount of time spent walking by the patient during a particular period of time;a caloric expenditure by the patient during a particular period of time;an amount of time between episodes of walking;a stride pattern of the patient;a measure of activity of the patient;a velocity of the patient;a length of a particular step; andan amount of time relating to a duration of the particular step. 8. The system of claim 1, comprising an electrical stimulation circuit coupled to the control circuit, the electrical stimulation circuit configured to deliver an electrical pulse, wherein delivery of the electrical pulse is based on data from one or more actual individual steps of the patient. 9. The system of claim 1, wherein the control circuit is configured to confirm that the candidate step is an actual individual step of the patient by identifying a first step followed by a second step within a particular period of time. 10. The system of claim 1, wherein the control circuit is configured to confirm that the candidate step is an actual individual step of the patient by identifying three consecutive steps. 11. A method comprising: receiving data from an implantable moti... | Cardiac Pacemakers, Inc. | A61B005/11
US07894875 | B2 | 2011-02-22 | Quantifying blood volume using magnetization transfer magnetic resonance imaging | 2005-06-01 | A magnetic resonance method for imaging blood volume in parenchyma via magnetic transfer (MT) includes: determining a MT effect of parenchyma; determining a MT effect of tissue; and quantifying the parenchymal blood volume using the difference between the MT effect of parenchyma and the MT effect of tissue. In one embodiment, the parenchymal blood volume is quantified through the following: MTRpar=MTRtissue(1−BV/Vpar), where MTRpar is the magnetization transfer ratio of parenchyma, MTRtissue is the magnetization transfer ratio of tissue, BV is the blood volume, and Vpar is a total parenchymal water volume. | 1. A magnetic resonance (MR) method for quantifying the blood volume in parenchyma via a magnetization transfer (MT) effect, comprising: determining a MT effect in parenchyma by operations including acquiring a MR signal of parenchyma in a presence of off-resonance radio frequency (RF) irradiation, acquiring a MR signal of parenchyma in the absence of off-resonance RF irradiation, and determining the MT of parenchyma based on the parenchymal MR signals;determining a MT effect in tissue;determining a MT effect in blood; andquantifying the parenchymal blood volume based on the MT effect of parenchyma, the MT effect of tissue, and the MT effect of blood. 2. The method of claim 1, wherein determining the MT of tissue includes: performing a blood-nulling MR pulse sequence;acquiring a MR signal of tissue in the presence of off-resonance RF irradiation;acquiring a MR signal of tissue in the absence of off-resonance RF irradiation; anddetermining the MT of tissue based on the MR tissue signals. 3. The method of claim 2, wherein performing a blood-nulling MR pulse sequence includes at least one of the following: applying a 180 degree inversion radio frequency pulse followed by an inversion recovery time that nulls a blood MR signal based on a correct T1 of blood water at a particular magnetic field strength;applying one of a strong coherent flow-based and an incoherent diffusion-based de-phasing gradient during evolution of the magnetization before acquisition;exciting and de-phasing water protons in the blood before flowing into a slice of interest; andnulling a MR signal of blood. 4. The method of claim 3, wherein nulling the MR signal of blood includes adding a relaxation agent to remove a water signal of blood. 5. The method of claim 1, wherein determining the MT of blood includes: acquiring a MR signal of blood in a presence of off-resonance RF irradiation;acquiring a MR signal of blood in the absence of off-resonance RF irradiation; anddetermining the MT of blood based on the blood MR signals. 6. The method of claim 5, wherein the blood MR signals are one of the following: isolated blood signals in vivo, andblood signals from a blood phantom maintained under physiological conditions. 7. The method of claim 1 further including: determining the MT effect without a perturbation;determining the MT effect with a perturbation; andcomputing the effect of a blood volume change using the difference between the MT effect without the perturbation and the MT effect with the perturbation. 8. The method of claim 7, further including introducing the perturbation through at least one of the following: CO2 breathing, breathhold, hyperventilation, brain activation, inhale chemicals, and ingested chemicals. 9. A magnetic resonance (MR) method for quantifying the blood volume in parenchyma via magnetization transfer ratio (MTR) comprising: determining a MTR in parenchyma by operations including acquiring a MR signal of parenchyma in a presence of off-resonance ... | - | A61B005/055
US08266349 | B2 | 2012-09-11 | Systems and methods for selecting parameters used in a portable patient monitor | 2009-08-21 | A patient monitoring system includes a patient monitor and a plurality of docking stations. Each docking station is associated with a set of patient monitor configurations. Interfacing the patient monitor with a particular docking station allows a user to select one or more of the set of patient monitor configurations associated with the particular docking station for use in the patient monitor. | 1. A patient monitor system, comprising: a plurality of docking stations, each docking station associated with a respective patient monitoring location;a portable patient monitor configured to be selectively coupled with and selectively decoupled from the respective docking stations as a user transports the portable patient monitor among the patient monitoring locations, the portable patient monitor comprising: a patient data acquisition unit for receiving patient data signals from one or more sensors attached to a patient; anda processor for processing the patient data signals, the processor configured to: detect a coupling of the portable patient monitor to a first docking station of the plurality of docking stations, the first docking station being associated with a first patient monitor location;in response to a predetermined condition, prompt the user to select one or more first default configuration settings associated with the first patient monitor location; andin response to a first user selection, load the selected one or more first default configuration settings into the portable patient monitor,wherein the predetermined condition comprises a determination by the processor that the first patient monitor location is a first non-transient hospital ward,wherein the processor is further configured to: detect a coupling of the portable patient monitor to a second docking station of the plurality of docking stations, the second docking station being associated with a second patient monitor location;determine that the second patient monitor location is a transient hospital ward where the patient is expected to remain for a relatively short period of time as compared to that spent in the first patient monitor location; andin response to the determination of the transient hospital ward, maintain a current set of configuration settings in disregard to one or more second default settings, if any, associated with the second patient monitor location, andwherein the processor is further configured to: detect a coupling of the portable patient monitor to a third docking station of the plurality of docking stations, the third docking station being associated with a third patient monitor location;determine that the third patient monitor location is a second non-transient hospital ward;in response to the determination that the third patient monitor location is the second non-transient hospital ward, prompt the user to select one or more third default configuration settings associated with the third patient monitor location; andin response to a second user selection, load the selected one or more third default configuration settings into the portable patient monitor. 2. The system of claim 1, wherein the predetermined condition further comprises a determination by the processor that a current set of configuration settings is different than the one or more first default configuration settings. 3. The system of claim 1, wherein the predetermined condition... | Mindray DS USA, Inc. | G06F003/00 G06F013/12 G06F013/00 H04L012/28 A61B005/00
US20110040165 | A1 | 2011-02-17 | INTEGRATED CARTRIDGE AND TUB ASSEMBLY | 2009-08-13 | An assembly for storing and dispensing a plurality of integrated lancet and test strips for use in a testing meter is provided. The assembly comprises a cartridge portion for housing a plurality of integrated lancet and test strips, and a tub portion coupled to the cartridge portion and providing a surface for advancing a single integrated lancet and test strip for use in the testing meter. A system is also provided which includes the assembly and a meter for determining an analyte concentration. | 1. An integrated cartridge and tub assembly for storing and advancing a plurality of integrated lancet and test strips for use in a testing meter, the assembly comprising: a cartridge portion for housing a plurality of integrated lancet and test strips; anda tub portion coupled to the cartridge portion, the tub portion providing a surface for advancing a single integrated lancet and test strip for use in the meter. 2. The integrated cartridge and tub assembly of claim 1, further comprising: a resilient biasing element coupled to the tub portion and cartridge portion, the resilient biasing element positioning the tub portion in a relative first position from the cartridge portion; anda sealing element positioned between the cartridge portion and the tub portion, wherein the resilient biasing element provides closing forces to the cartridge portion and tub portion, the closing forces providing an air-tight and moisture-tight environment for the plurality of integrated lancet and test strips. 3. The integrated cartridge and tub assembly of claim 2, wherein when a displacing force is applied to the assembly, the resilient biasing element is stressed and the tub portion is relatively displaced to a relative second position from the cartridge portion, wherein an opening is formed in the assembly when the tub portion is in the relative second position. 4. The integrated cartridge and tub assembly of claim 3, wherein the tub portion is irremovably coupled to the cartridge portion. 5. The integrated cartridge and tub assembly of claim 3, wherein the tub portion is removably coupled to the cartridge portion. 6. The integrated cartridge and tub assembly of claim 3, wherein the opening allows for the advancement of a single integrated lancet and test strip for use in the testing meter. 7. The integrated cartridge and tub assembly of claim 1, further comprising; a resilient biasing element coupled to the tub portion and cartridge portion, the resilient biasing element positioning the tub portion in a relative first position from the cartridge portion. 8. The integrated cartridge and tub assembly of claim 7, wherein when a displacing force is applied to the assembly, the resilient biasing element is stressed and the tub portion is relatively displaced to a relative second position from the cartridge portion, wherein an opening is formed in the assembly when the tub portion is in the relative second position. 9. The integrated cartridge and tub assembly of claim 8, wherein when the tub portion is in the relative second position, an opening is formed in the assembly. 10. The integrated cartridge and tub assembly of claim 9, wherein the opening is formed between the tub portion and the cartridge portion. 11. The integrated cartridge and tub assembly of claim 8, wherein the assembly comprises an obstructed opening when the tub portion is in the relative first position, the obstructed opening becoming unobstructed when the tub portion is in the relative s... | - | A61B005/00 B65D083/00
US20110004080 | A1 | 2011-01-06 | METHOD FOR NON-INVASIVE ANALYSIS OF A SUBSTANCE CONCENTRATION WITHIN A BODY | 2010-09-15 | Embodiments of the present methods measure a concentration of a substance, such as glucose, in a body. The described embodiments change the temperature of the surface of a body from a first temperature to a second temperature, then change the temperature of the surface of the body from the second temperature back to the first temperature, and measure a first amount of infrared (“IR”) radiation absorbed or emitted from the surface of the body in a first wavelength band, and a second amount of IR radiation absorbed or emitted from the surface of the body in a second wavelength band at predetermined time intervals during the time period that the temperature of the surface of the body changes from the second temperature back to the first temperature. The described embodiments also measure a temperature at the surface of the body and an ambient temperature. A normalized ratio parameter is calculated from the four measurements, and the concentration of the substance in the body is determined by correlating the normalized ration parameter with the body surface temperature and the ambient temperature. | 1. A method for non-invasive measurement of a concentration of a substance in a body, comprising the steps of: changing a temperature of a surface of a body from a first temperature to a second temperature;changing the temperature of the surface of the body from the second temperature to the first temperature;measuring a first amount of infrared radiation absorbed or emitted from a body in a first wavelength band at predetermined time intervals during the change of the surface of the body from the second temperature to the first temperature;measuring a second amount of infrared radiation absorbed or emitted from a body in a second wavelength band at predetermined time intervals during the change of the surface of the body from the second temperature to the first temperature;measuring a temperature at the surface of the body;measuring an ambient temperature;calculating a normalized ratio parameter based on the first amount, the second amount, the body surface temperature and the ambient temperature; anddetermining the concentration of the substance in the body by correlating the normalized ratio parameter with the body surface temperature and the ambient temperature. 2. The method of claim 1, wherein the first wavelength band is a wavelength band or bands in which the substance emits and absorbs infrared radiation, and the second wavelength band is the entire spectrum in which the body emits and absorbs infrared radiation in the far infrared spectrum. 3. The method of claim 1, wherein the first wavelength band is a wavelength band or bands in which the substance emits and absorbs infrared radiation, and the second wavelength band is the entire spectrum in which the substance has no or negligible emission and absorption of infrared radiation. 4. The method of claim 1, wherein the normalized ratio parameter is the average ratio selected from the group consisting of at least one of the ratio of the first amount normalized against a black body to the second amount normalized against a black body at each time interval, and a logarithm of the ratio of the first amount normalized against a black body to the second amount normalized against a black body at each time interval. 5. The method of claim 1, wherein the substance is glucose. 6. The method of claim 1, wherein the body is a human body. 7. The method of claim 1, wherein the concentration of the substance in the body is determined using an empirically deprived lookup table. 8. The method of claim 1 wherein the first wavelength range comprises about 8.5μ to about 10.0μ and the second wavelength range comprises about 10.5μ to about 15.0μ. 9. The method of claim 8 wherein the second wavelength range comprises about 7.0μ to about 15.0μ. 10. The method of claim 1 wherein the step of changing the temperature of the surface of the body from the second temperature to the first temperature is at a predetermined rate. 11. A method of noninvasively measuring a concentration of a substance in a body u... | GLUCOVISTA, LLC | A61B005/145
US20110009718 | A1 | 2011-01-13 | DETERMINATION OF PHYSIOLOGICAL PARAMETERS USING REPEATED BLOOD PRESSURE MEASUREMENTS | 2010-07-14 | A pulsewave detection unit is provided, at least a portion of the pulsewave detection unit being configured to be coupled to a portion of a subject's body, the pulsewave detection unit generating a signal that is responsive to arterial pressure of the portion of the subject's body. A pulsewave parameters determination unit receives respective first and second signals from the pulsewave detection unit, the signals being responsive to arterial pressure of the portion of the subject's body while the portion of the pulsewave detection unit that is coupled to the portion of the subject's body is at respective first and second heights, at respective times. An arterial parameters calculating unit determines an arterial property of the subject by processing the first and second signals, and generates an output in response to determining the arterial property. Other embodiments are also described. | 1. Apparatus, comprising: a pulsewave detection unit, at least a portion of which is configured to be coupled to a portion of a subject's body, the pulsewave detection unit being configured to generate a signal that is responsive to arterial pressure of the portion of the subject's body; anda control unit comprising:a pulsewave parameters determination unit configured to receive respective first and second signals from the pulsewave detection unit, the signals being responsive to arterial pressure of the portion of the subject's body while the portion of the pulsewave detection unit that is coupled to the portion of the subject's body is at respective first and second heights, at respective times; andan arterial parameters calculating unit configured to determine an arterial property of the subject by processing the first and second signals, and to generate an output in response to determining the arterial property. 2-4. (canceled) 5. The apparatus according to claim 1, wherein the pulsewave detection unit comprises a sensor selected from the group consisting of a photoplethysmographic sensor, a strain gauge plethysmograph, a cuff configured to be placed on the portion of the subject's body and to generate the signal responsively to changes in blood pressure of the subject, an intravascular pressure sensor, and a sensor configured to generate the signal responsively to measuring a spectral property of blood of the subject. 6-8. (canceled) 9. The apparatus according to claim 1, wherein the arterial parameters calculating unit is configured to facilitate the generation of the first and second signals while the portion of the pulsewave detection unit is at respective first and second heights by instructing the subject to move the portion of the subject's body to which the portion of the pulsewave detection unit is coupled, to respective first and second heights. 10-13. (canceled) 14. The apparatus according to claim 1, further comprising a height-related input unit coupled to the control unit, wherein the control unit is configured to receive height indications regarding the first and second heights from the subject via the height-related input unit, and the arterial parameters calculating unit is configured to determine the arterial property of the subject by processing the first and second signals and the height indications regarding the first and second heights. 15. The apparatus according to claim 1, further comprising a height-detecting unit configured to detect the first and second heights, wherein the control unit is configured to receive the height indications regarding the first and second heights from the height-detecting unit, and the arterial parameters calculating unit is configured to determine the arterial property of the subject by processing the first and second signals and the height indications regarding the first and second height. 16. The apparatus according to claim 15, wherein the height-detecting unit comprises a sens... | - | A61B005/021 A61B005/1455 A61B008/00
US07871382 | B2 | 2011-01-18 | Interface between a noninvasive blood pressure sensor and an invasive blood pressure monitor | 2006-11-16 | The preferred embodiment of the present invention comprises a single microprocessor-based interface that connects between a noninvasive blood pressure (NIBP) sensor and an invasive blood pressure (IBP) monitor or module. The interface effectively emulates an IBP transducer in such a way that the IBP monitor sees the interface as if it were a regular IBP transducer from a fluid-filled blood pressure monitoring line. It receives the signal from an NIBP sensor and determines the blood pressure corresponding to the signal. It accepts the excitation voltage provided by the IBP monitor. From the excitation voltage and a known transducer sensitivity which the IBP monitor is configured to work with, the interface emulates the IBP transducer output signal corresponding to the blood pressure. The interface also emulates the input and output impedances of the IBP transducer which the IBP monitor is configured to work with. Zeroing of the interface with the IBP monitor can be easily performed in a way that is similar to that for a fluid-filled system. A noninvasive system comprising a suitable NIBP sensor and this interface can be used as an alternative to the fluid-filled monitoring line. | 1. An interface configured to connect between a noninvasive blood pressure (NIBP) sensor and an invasive blood pressure (IBP) monitor comprising: a first input confiqured to receive a measurement signal indicative of the NIBP of a subject;a second input confiqured to receive a transducer excitation signal provided by the IBP monitor;at least one processor(s) confiqured to receive said measurement signal and said excitation signal and emulate an output signal indicative of the IBP of a subject according to predetermined instructions; andan output configured to provide said output signal in a form suitable for input to the IBP monitor;wherein said first input is connected to the NIBP sensor and said second input and said output are both connected to the IBP monitor; andwherein said second input and said output are configured to connect to the IBP monitor by means of a detachable or configurable connection between said interface and an IBP interface cable for the IBP monitor. 2. An interface as claimed in claim 1 wherein said connection includes an adapter with a first connector configured to connect to said interface and a second connector configured to connect to a transducer end of the IBP interface cable. 3. An interface as claimed in claim 2 wherein said adapter includes a cable between said first and second connectors. 4. An interface as claimed in claim 2 wherein said adapter does not include a cable between said first and second connectors. 5. An interface configured to connect between a noninvasive blood pressure (NIBP) sensor and an invasive blood pressure (IBP) monitor comprising: a first input configured to receive a measurement signal indicative of the NIBP of a subject;a second input configured to receive a transducer excitation signal provided by the IBP monitor;at least one processor(s) configured to receive said measurement signal and said excitation signal and emulate an output signal indicative of the IBP of a subject according to predetermined instructions; andan output configured to provide said output signal in a form suitable for input to the IBP monitor;wherein said processor(s) is configured such that a user can select the sensitivity of said output signal in relation to said measurement signal from a predetermined range of choices. 6. An interface as claimed in claim 5 wherein said sensitivity can be selected or specified by a user. 7. An interface as claimed in claim 6 wherein said range includes 5 μV/V/mmHg and 40 μV/V/mmHg. 8. An interface configured to connect between a noninvasive blood pressure (NIBP) sensor and an invasive blood pressure (IBP) monitor comprising: a first input configured to receive a measurement signal indicative of the NIBP of a subject;a second input configured to receive a transducer excitation signal provided by the IBP monitor;at least one processor(s) configured to receive said measurement signal and said excitation signal and emulate an output signal indicative of the IBP of a subject a... | Healthstats International Pte Ltd | A61B005/02
US07894884 | B2 | 2011-02-22 | System and method for ischemia classification with implantable medical device | 2007-01-31 | An implantable medical device monitors ST segment data collected from EGM. ST trends are established and monitored over time. The IMD is able to discern whether the data indicate supply ischemia, demand ischemia, or other physiological causes. The IMD is then able to provide appropriate information and alerts. | 1. A method comprising: obtaining an electrogram representing cardiac data from an implantable medical device for a plurality of cardiac cycles;determining an ST segment data point for each of the cardiac cycles;initiating a timer interval when one of the ST segment data point values exceeds a threshold value;determining a maximum ST segment data point value during the timer interval;determining an amount of time from initiation of the timer interval until the maximum ST segment data point value occurs; anddetermining whether the ST segment data points indicate ischemia based on the amount of time. 2. The method of claim 1, further comprising determining whether the ST segment data points indicate demand ischemia. 3. The method of claim 1, wherein determining an ST segment data point further comprises: identifying a peak of an R wave;adding a predetermined time interval to the identified peak; andidentifying a recorded value at a time identified by the predetermined time interval after the peak of the R wave as the ST segment data point. 4. The method of claim 3, further comprising: determining a PR isoelectric baseline value; andsubtracting the PR isoelectric baseline value from the recorded value to identify the ST segment data point. 5. The method of claim 1, further comprising: determining whether the ST segment data points indicate supply ischemia. 6. The method of claim 5, further comprising: determining that the ST segment data points do not indicate supply ischemia if the amount of time is less than a first short value or greater than a first long value. 7. The method of claim 5, further comprising: determining that the ST segment data points do not indicate supply ischemia if less than a predetermined percentage of the ST segment data points occurring during the timer interval exceed the threshold value. 8. The method of claim 5, further comprising: determining that the ST segment data points do not indicate supply ischemia if an ST segment data point occurring at a termination of the timer interval is less than a predetermined percentage of the maximum ST segment data point value. 9. The method of claim 5, further comprising: calculating a slope of the ST segment data points prior to the maximum ST segment data point; anddetermining that the ST segment data points do not indicate supply ischemia if the slope is less than a first slope value. 10. The method of claim 5, further comprising evaluating a secondary indicator to confirm the determination of whether the ST segment data points indicate supply ischemia. 11. The method of claim 1, further comprising: monitoring the ST segment data point values; andprocessing the values using a predetermined filter bank. 12. The method of claim 1, further comprising: monitoring the ST segment data point values; andprocessing the values by comparing to a template. 13. An apparatus comprising: means for obtaining an electrogram representing cardiac data from an implantable medical device ... | Medtronic, Inc. | A61B005/02
US08428675 | B2 | 2013-04-23 | Nanofiber adhesives used in medical devices | 2009-08-19 | Nanofiber adhesives and their uses with device, such as medical devices, are described. In one embodiment, a nanofiber adhesive layer may be disposed on a surface of a medical device, such as a backing layer of a sensor, for adhesion to a substrate. The nanofiber adhesive layer may allow durable adhesion to the substrate. Other described features may include materials and methods to determine the attachment of the medical device to a patient by determining the adhesive state of the adhesive layer. | 1. A device, comprising: a backing layer;an adhesive layer comprising a plurality of nanofibers attached to the backing layer to couple the device to a patient tissue; andtwo or more electrical contacts positioned within or beneath the adhesive layer, wherein the two or more electrical contacts are capable of conducting a current through the adhesive layer or through the patient tissue to provide a signal related to an adhesive state of the adhesive layer. 2. The device of claim 1, wherein the plurality of nanofibers of the adhesive layer undergo a conformational change when adhered to the substrate such that a characteristic of the current flowing between the two or more electrical contacts is altered. 3. The device of claim 1, wherein the electrical contacts are positioned within the adhesive layer. 4. The device of claim 1, wherein the device comprises a pulse oximetry sensor, electroencephalogram lead, or an electrocardiogram lead. 5. A medical device, comprising: a backing material;a sensor or lead attached to the backing material, wherein the sensor or lead is capable of generating physiological data monitored by a patient monitor;an adhesive layer comprising a plurality of nanofibers configured to adhere the medical device to a patient tissue, wherein the adhesive layer is attached to the backing material; andtwo or more electrical contacts configured to conduct a current through the adhesive layer or through the patient tissue to provide a signal related to an adhesive state of the adhesive layer. 6. The medical device of claim 5, wherein the plurality of nanofibers are aligned substantially perpendicular to the backing layer. 7. The medical device of claim 5, wherein the adhesive layer is attached to the backing material using a glue or a polymeric bonding agent. 8. The medical device of claim 5, wherein the plurality of nanofibers comprise a polymer material. 9. The medical device of claim 5, wherein the plurality of nanofibers are comprised of any one of or a combination of single, double, or multiwalled carbon nanotubes. 10. The medical device of claim 5, wherein the medical device comprises a pulse oximetry sensor, an electroencephalogram lead, or an electrocardiogram lead. 11. The medical device of claim 5, wherein the plurality of nanofibers comprise one or more of carbon, boron, silicon, boron carbide, silicon carbide, boron-silicon, cellulose or its synthetic analogs, collagen or its synthetic analogs, or keratin or its synthetic analogs. 12. A method of manufacturing a medical device, comprising: applying a sensor or lead on a backing material;applying an adhesive layer on the backing material, wherein the adhesive layer comprises a plurality of nanofibers configured to adhere the medical device to a patient tissue; andapplying electrical contacts between the backing material and the adhesive layer, or within the adhesive layer, to provide a signal related to an adhesive state of the adhesive layer. 13. The method as... | Covidien LP | A61B005/1455
US20110015515 | A1 | 2011-01-20 | Methods For Physiological Monitoring, Training, Exercise And Regulation | 2010-09-23 | Computer executable software and device for guiding brain activity training comprising: logic which takes data corresponding to activity measurements of one or more internal voxels of a brain and determines one or more members of the group consisting of: a) what next stimulus to communicate to the subject, b) what next behavior to instruct the subject to perform, c) when a subject is to be exposed to a next stimulus, d) when the subject is to perform a next behavior, e) one or more activity metrics computed from the measured activity, f) a spatial pattern computed from the measured activity, g) a location of a region of interest computed from the measured activity, h) performance targets that a subject is to achieve computed from the measured activity, i) a performance measure of a subject's success computed from the measured activity, j) a subject's position relative to an activity measurement instrument; and logic for communicating information based on the determinations to the subject in substantially real time relative to when the activity is measured. | 1. A method comprising: taking brain activity measurements from one or more regions internal to a brain of a subject;using a computer executable logic for determining information based upon said brain activity measurements; andusing a computer executable logic for communicating said information to said subject in substantially real time. 2. The method of claim 1 wherein said one or more regions internal to the brain comprise one or more voxels. 3. The method of claim 1 wherein said one or more regions internal to the brain comprise regions of interest. 4. The method of claim 1 wherein said taking brain activity measurements are made using an fMRI apparatus. 5. The method of claim 1 further comprising pre-training said subject. 6. The method of claim 1 further training said subject to modulate brain activity without taking brain activity measurements. 7. The method of claim 1 wherein said taking brain activity measurements are made using an apparatus capable of taking measurements from one or more internal voxels without substantial contamination of the measurements by activity from regions intervening between the internal voxels being measured and where the measurement apparatus collects the data. 8. The method of claim 1 wherein said information comprises an activity metric. 9. The method of claim 1 wherein said activity measurements are recorded. 10. The method of claim 1 wherein the region of interest has a primary function of releasing a neuromodulatory substance. 11. The method of claim 1 wherein said one or more regions internal to said brain comprise a scan volume. 12. The method of claim 1 wherein said one or more regions internal to said brain comprise a spatial pattern. 13. A method of treating a condition in a subject, wherein said condition is associated with activity of one or more regions internal to the brain of said subject comprising: training said subject to modulate said activity using fMRI. 14. The method of claim 13 wherein said training comprises training said subject to increase the level of said activity or decrease the level of said activity. 15. The method of claim 13 wherein said training comprises training said subject to change the activity from a first activity metric to a second activity metric. 16. The method of claim 13 wherein said one or more regions have a primary function of releasing a neuromodulatory substance. 17. The method of claim 13 further comprising communicating information based upon activity measurements made by said fMRI to said subject in substantially real time. 20. The method of claim 13 wherein said one or more regions has abnormally increased or decreased activity due to an injury or disease. | R. Christopher DeCharms | A61B005/055
US20110046459 | A1 | 2011-02-24 | Non-Invasive Patient Monitoring Using Near Infrared Spectrophotometry | 2010-06-29 | Methods, systems, and related computer program products for non-invasive spectrophotometric monitoring of an optical property of a medium are described. Respective light portions are propagated through respective at least partially non-overlapping subregions of the medium during each of a calibration interval and a subsequent monitoring interval. Detections of the light portions acquired during the calibration interval are processed to compute at least one algorithm compensation that causes a first result related to the optical property for the first subregion to be substantially equal to a second result related to the optical property for the second subregion. Subsequently, detections of respective light portions acquired during the monitoring interval are processed in conjunction with the at least one algorithm compensation to compute a monitoring result for the optical property. Advantages can include a reduction in the number of sources/detectors required and/or an obviation of symmetry requirements in source/detector layout. | 1. A method for non-invasive spectrophotometric monitoring of an optical property of a medium, comprising: securing each of a first optical source, a second optical source, a first optical detector, and a second optical detector to a surface of the medium;during each of a calibration interval and a monitoring interval, said monitoring interval being subsequent to said calibration interval, propagating a first portion of light from the first optical source through the medium to the first optical detector, a second portion of light from the second optical source through the medium to the first optical detector, a third portion of light from the first optical source through the medium to the second optical detector, and a fourth portion of light from the second optical through the medium to the second optical detector;processing detections of said first, second, third, and fourth light portions acquired during said calibration time interval to compute at least one algorithm compensation that causes a first result related to said optical property based on said first and second detected light portions to be substantially equal to a second result related to said optical property based on said third and fourth detected light portions; andprocessing detections of said first, second, third, and fourth light portions acquired during said monitoring interval in conjunction with said at least one algorithm compensation to compute a monitoring result for the optical property of the medium. 2. The method of claim 1, said first and second light portions corresponding to a first subregion of the medium, said third and fourth light portions corresponding to a second subregion of the medium that is at least partially non-overlapping with said first subregion, wherein said computing said at least one algorithm compensation comprises: (i) computing at least one error factor associated with at least one non-ideality of said optical sources and/or detectors to which a difference in said first and second results would be attributable if the optical property was known to be spatially homogenous throughout said first and second subregions during said calibration interval; and(ii) determining at least one compensation factor associated with said at least one error factor that causes said first and second results to be substantially equal for said calibration interval. 3. The method of claim 2, wherein said at least one compensation factor remains fixed throughout said monitoring interval. 4. The method of claim 2, wherein said at least one non-ideality is associated with one or more of intensity of the optical sources, sensitivity of the optical detectors, coupling efficiency of light from the optical sources into the medium, and coupling efficiency of light from the medium to said optical detectors. 5. The method of claim 4, wherein said first and second optical sources modulate said light portions according to a phase modulated spectrophotometry (PMS) modulation pat... | O2 MEDTECH, INC. | A61B005/00
US20110009730 | A1 | 2011-01-13 | MRI SYSTEM APPARATUS TO PRODUCE ELECTRIC CURRENT | 2008-11-27 | A magnetic resonance imaging system (10) includes a main field magnet (12) which generates a B 0 magnetic field through an examination region (14). One or more electric power generators (30) are disposed in the B 0 magnetic field. Each of the electric power generators includes at least one winding or coil (34) which is rotatably mounted for movement relative to the B0 magnetic field. A mechanical mechanism (60) such as vanes or propellers (62), a turbine (74) in conjunction with a fluid pump (72), or a motor (78) in conjunction with a drive shaft (76) and gear box (74), drive the at least one winding (34) to move in such a manner that it interacts with the B 0 magnetic field to generate an electric current. | 1. A magnetic resonance system comprising: a magnet which generates a B0 magnetic field;a system for inducing resonance and receiving induced resonance signals;an electric power generator disposed in the B0 magnetic field, the electric power generator including:at least one winding,a mechanism for mounting the winding for movement relative to the B0 magnetic field,a mechanical mechanism which drives the at least one winding to move such that it interacts with the B0 magnetic field to generate a current. 2. The system according to claim 1, wherein the mechanical mechanism includes fan or impeller blades. 3. The system according to claim 1, wherein the mechanical mechanism includes a fluid-driven turbine. 4. The system according to claim 1, wherein the drive mechanism includes a mechanical drive. 5. The system according to claim 1, wherein the system for inducing resonance and receiving induced resonance signals includes a local receive coil, the electric power generator being mounted to at least one of the local receive coil or a patient support adjacent to the local receive coil. 6. The system according to claim 5, wherein the local receive coil includes at least one of: an amplifier connected with the electric current generator to receive operating power therefrom; and/oran analog-to-digital converter connected with the electric current generator to receive operating power therefrom. 7. The system according to claim 1, further including an electronic device configured for selective deployment into the B0 field, the current generator being coupled with the electronic device to provide operating power thereto. 8. The system according to claim 7, wherein the electronic device includes at least one of a physiological monitor, a patient entertainment device, a music player, a video player, an fMRI probe, or a patient table position monitor. 9. The system according to claim 1, further including: a subject support for moving a subject into and out of an examination region, the electric current generator being mounted in the subject support. 10. The system according to claim 1, further including: a fluid cooling system which circulates cooling fluid to heat-generating components and wherein the current generator drive mechanism includes a fluid-driven turbine connected with the fluid cooling system. 11. The system according to claim 1, wherein the magnetic resonance system is a magnetic resonance imaging system which generates magnetic resonance images of a portion of a subject in an examination region and wherein the current generator is disposed outside of the examination region and in the B0 magnetic field. 12. An electric generator for use in a magnetic resonance system which generates a B0 magnetic field, the electric generator comprising: at least one winding;a mechanism for mounting the winding for movement relative to the B0 magnetic field; anda mechanical mechanism which drives the at least one winding to move such that it interacts ... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/055 G01R033/36
US20110046510 | A1 | 2011-02-24 | Glucose Monitoring Device And Method | 2010-08-23 | A method for determining the quantity of an analyte in a fluid is described along with various components of an apparatus designed to carry out the method. The method involves habituating a patient's eye to one or more colors, measuring the recovery time, and correlating the recovery time to the quantity of the analyte. A reproducible, objective, non-attentiveness-dependant test for assessing analyte levels is further disclosed. To this end, a device that measures saccadic eye movements may be used to assess the return of vision to the prehabituated state. The method and apparatus are particularly suited for noninvasively measuring blood glucose levels. | 1. (canceled) 2. A method of monitoring glucose levels in a patient, the method comprising: displaying a first image on a screen of a display module, wherein the first image is displayed for a first period of time to allow a patient's vision to reach a habituated state while viewing the first image;terminating the first image;monitoring saccadic eye movements of the patient using a sensor after the first image is terminated to determine when the patient's vision returns from the habituated state to a non-habituated state;determining a second period of time from when the patient's vision returns to the non-habituated state; andcorrelating the second period of time with a blood glucose level of the patient using a processor. 3. The method of claim 2 further comprising displaying a second image after the first image is terminated. 4. The method of claim 3 wherein the second image further comprises a white background. 5. The method of claim 2 wherein the first image further comprises a pattern having one or more colors. 6. The method of claim 2 wherein the first image further comprises an object having one or more colors. 7. The method of claim 3 wherein the second image further comprises a pattern having one or more colors. 8. The method of claim 3 wherein the second image further comprises an object having one or more colors. 9. The method of claim 3 wherein at least a portion of the second image moves in one or more directions on the screen. 10. An apparatus for monitoring glucose levels in a patient, the apparatus comprising: a processor configured to produce a first image on the screen of a display module, wherein the first image is displayed for a first period of time to allow the patient's vision to reach a habituated state while viewing the first image;a sensor configured to monitor saccadic eye movements of the patient after the processor terminates the first image, wherein the sensor provides a sensor signal to the processor to determine when the patient's vision returns from the habituated state to a non-habituated state;a circuit configured to track a second period of time from when the patient's vision returns to the non-habituated state and store the second period of time in a memory,wherein the processor analyzes the second period of time to compute a blood glucose level of the patient. 11. The apparatus of claim 10, wherein the processor displays a second image after the first image is terminated. 12. The apparatus of claim 11, wherein the second image further comprises a white background. 13. The apparatus of claim 10, wherein the first image further comprises a pattern having one or more colors. 14. The apparatus of claim 10, wherein the first image further comprises an object having one or more colors. 15. The apparatus of claim 11, wherein the second image further comprises a pattern having one or more colors. 16. The apparatus of claim 11, wherein the second image further comprises an object having one or more colo... | - | A61B005/00
US08961459 | B2 | 2015-02-24 | Fingertip positioned artery stabilizer | 2010-09-26 | An artery stabilizer device is provided for stabilizing an artery while a technician inserts a needle into the stabilized artery. Two stabilizer fingers are provided to hold an artery in place while a needle is passed between two extensions and further inserted into the artery beyond the stabilizer fingers; the device can be removed from the worksite while the tip of the needle remains lodged within the artery. A platform is disposed above the stabilizer fingers and extensions create ample distance between the stabilizer fingers and the platform for simplified palpation and improved visibility and safety. Each embodiment can create an augmented pulse at a targeted needle insertion site, making it easier for a technician to palpate the patient's pulse there. Arterial catheter insertions can be facilitated using this invention. | 1. A device for stabilizing a peripheral artery comprising two solid stabilizer fingers disposed parallel to each other, shaped to be pressed down over each side of a targeted peripheral artery, and spaced to straddle said artery; a fingerhold platform comprising an outspread top surface that is disposed horizontally over said stabilizer fingers; andsolid elongated extensions that connect said stabilizer fingers to said fingerhold platform whereby said elongated extensions distance said stabilizer fingers from said platform vertically to provide enough space to pass a needle between said stabilizer fingers under said platform and to keep a technician's extremities distanced from said needle during an arterial access procedure. 2. A fingertip positioned artery stabilizer device comprising two solid stabilizer fingers disposed parallel to each other, adapted to be pressed onto each side of a targeted peripheral artery, and spaced to straddle said artery; a solid fingerhold that is adapted to provide stable support against at least one fingertip of a human hand and is spread out horizontally above said stabilizer fingers whereby a technician can press down on said fingerhold to lodge said stabilizer fingers lengthwise over each side of a targeted peripheral artery; andsolid elongated extensions connecting said stabilizer fingers to said fingerhold and distancing them vertically whereby enough space is provided for a needle to be passed between said stabilizer fingers beneath said platform and whereby a technician's extremities are distanced from said needle during an arterial access procedure. 3. A fingertip positioned artery stabilizer device comprising two solid stabilizer fingers disposed parallel to each other and spaced to straddle a peripheral artery; a solid fingerhold platform with an outspread surface that is adapted to provide stable support against a fingertip of a human hand, is spread out horizontally above said stabilizer fingers, and allows stable fingertip control of the placement of said stabilizer fingers during use, whereby said fingerhold platform can be manually pressed down directly over said stabilizer fingers to stably lodge said stabilizer fingers lengthwise over each side of a targeted peripheral artery;a first solid elongated extension attached to said fingerhold platform; andsecond and third solid elongated extensions connecting said stabilizer fingers to said first solid extension and vertically distancing said stabilizer fingers from said platform whereby enough space is provided for a needle to be passed between said stabilizer fingers underneath said platform and a technician's extremities are distanced from said needle during an arterial access procedure. 4. A fingertip positioned artery stabilizer device comprising two solid stabilizer fingers positioned parallel to each other, adapted to be pressed onto each side of a targeted peripheral artery, and spaced to straddle said artery; a solid fingerhold that is adap... | - | A61M005/00 A61B017/12 A61B017/00
US07894886 | B2 | 2011-02-22 | Method and apparatus for discriminating cardiac signals in a medical device based on wavelet decomposition analysis | 2007-05-04 | A method and device for detecting cardiac signals in a medical device that includes decomposing sensed cardiac signals using a wavelet function to form a corresponding wavelet transform, generating a first wavelet representation corresponding to the wavelet transform that is responsive to RR intervals of the sensed cardiac signals, generating a second wavelet representation that is not responsive to RR intervals associated with the sensed cardiac signals, and determining a device failure in response to the first wavelet representation and the second wavelet representation. The method and device may also include decomposing sensed cardiac signals using a wavelet function to form a corresponding wavelet transform, generating a wavelet representation that is not responsive to RR intervals of the sensed cardiac signals, determining RR intervals associated with the sensed cardiac signals, and determining a device failure in response to the first wavelet representation and the determined RR intervals. | 1. A medical device, comprising: means for sensing cardiac signals;a cardiac lead, the means for sensing cardiac signals being positioned, at least in part, with respect to the cardiac lead;means for applying a wavelet function to the sensed cardiac signals to generate a wavelet transform from the sensed cardiac signals using dyadic wavelet decomposition;means for generating a first wavelet representation from the wavelet transform, the first wavelet representation being responsive to highpass frequency content of the sensed cardiac signals;means for generating a second wavelet representation from the wavelet transform, the second wavelet representation being responsive to lowpass frequency content of the sensed cardiac signals; andmeans for determining a lead failure of the cardiac lead in response to the first wavelet representation and the second wavelet representation. 2. The device of claim 1, wherein the wavelet transform is a Haar wavelet transform and the second wavelet representation corresponds to an approximation sequence. 3. The device of claim 1, wherein the wavelet transform is a Haar wavelet transform and the first wavelet representation corresponds to a detail sequence. 4. The device of claim 1, wherein the sensed cardiac signals are decomposed at a plurality of scales and further comprising: means for determining approximation coefficients in response to the plurality of scales; andmeans for reconstructing the second wavelet representation using predetermined approximation coefficients of the determined approximation coefficients. 5. The device of claim 4, further comprising: means for determining detail coefficients in response to the plurality of scales; andmeans for reconstructing the first wavelet representation using predetermined detail coefficients of the determined detail coefficients. 6. The device of claim 5, wherein the predetermined approximation coefficients correspond to a coarsest scale of the plurality of scales and the detail coefficients correspond to a finest scale of the plurality of scales. 7. The device of claim 6, wherein the approximation coefficients are determined at a fourth scale of the plurality of scales and the detail coefficients are determined at a first scale of the plurality of scales. 8. The device of claim 5, wherein the detail coefficients are determined at a first scale of the plurality of scales to form a first detail sequence and a second scale of the plurality of scales to form a second detail sequence, and wherein the first wavelet representation is reconstructed in response to a sum of the first detail sequence and the second detail sequence. 9. A medical device, comprising: means for sensing cardiac signals;a cardiac lead, the means for sensing being positioned, at least in part, with respect to the cardiac lead;means for applying a wavelet function to the sensed cardiac signals to generate a wavelet transform from the sensed cardiac signals using dyadic wavelet decomposition;m... | Medtronic, Inc. | A61B005/0428
US20110035232 | A1 | 2011-02-10 | METHODS OF TREATING HEPATIC ENCEPHALOPATHY | 2010-06-02 | Treatment of hepatic encephalopathy using gastrointestinal specific antibiotics is disclosed. One example of a gastrointestinal specific antibiotic is rifaximin. The instant application also provides methods for determining if a subject has a neurological condition or hepatic encephalopathy by determining the critical flicker frequency and/or the venous ammonia level of the subject at two or more time points. The further provided are methods for treating these subjects. | 1. A method of decreasing blood ammonia levels in a subject comprising: administering to the subject an effective amount of rifaximin, thereby reducing ammonia blood levels. 2-4. (canceled) 5. A method of decreasing fatigue in a subject having hepatic encephalopathy comprising: administering to the subject an effective amount of rifaximin, thereby reducing fatigue in a subject. 6. A method for treating a subject having HE, comprising: administering a GI specific antibiotic with lactulose, thereby treating a subject having HE. 7. (canceled) 8. (canceled) 9. The method of claim 6, further comprising advising a health care worker to administer the GI specific antibiotic prior to, with, or after administration of lactulose. 10-14. (canceled) 15. The method of claim 6 wherein the GI specific antibiotic is rifaximin. 16. A method of diagnosing hepatic encephalopathy (HE) in a subject comprising: determining the critical flicker frequency (CFF) of a subject at two or more time points, wherein a decrease in the CFF is indicative that the subject has hepatic encephalopathy. 17. The method of claim 16, wherein the CFF comprises the CFF time weighted average. 18. The method of claim 17, wherein the CFF time weighted average comprises less than about 24 Hz. 19. A method of determining a subject's risk of an HE breakthrough event, comprising: determining the critical flicker frequency (CFF) of a subject at two or more time points, wherein a decrease in the CFF is indicative that the subject has an increased risk of an HE breakthrough event. 20. The method of claim 19, wherein the CFF comprises the CFF time weighted average. 21. The method of claim 20, wherein the CFF time weighted average comprises less than about 24 Hz. 22. The method of claim 19, wherein a average CFF time weighted average 10 Hz is indicative that the subject has the greatest risk of a HE breakthrough event. 23. A method of determining the prognosis of a subject having HE, comprising: determining the critical flicker frequency (CFF) of a subject at two or more time points, wherein a decrease in the CFF is indicative that the subject has a poor prognosis. 24. The method of claim 23, wherein the CFF is the CFF time weighted average. 25. The method of claim 24, wherein the CFF time weighted average comprises less than about 24 Hz. 26. The method of claim 24, wherein the CFF time weighted average comprises less than about 20 Hz. 27. A method of treating or preventing an HE event, comprising: determining the critical flicker frequency (CFF) of a subject at two or more time points,administering to a subject having a decrease in the CFF average between the time points an effective amount of a GI specific antibiotic,thereby treating or preventing an HE event. 28. The method of claim 27, wherein the CFF comprises the CFF time weighted average. 29. The method of claim 28, wherein the CFF time weighted average comprises less than about 24 Hz. 30. The method of claim 28, wherein ... | SALIX PHARMACEUTICALS, LTD. | G06Q050/00 A61K031/437 A61P001/16 A61P025/00 A61P031/00 A61B005/00
US07865004 | B2 | 2011-01-04 | System, method, and program for medical image interpretation support | 2007-10-18 | The number of similar case images to be displayed is changed according to the number of times and progress of imaging diagnosis. First examination judgment means is added to a medical image interpretation support system comprising storage means for storing image data sets obtained by imaging subjects, similar case image search means for extracting similar case image data sets having a characteristic similar to an interpretation target image from the image data sets, and display control means for controlling display of the interpretation target image and the similar case images. The display control means controls the display so as to display a larger number of the similar case images together with the interpretation target image in the case where the interpretation target image has been judged to be a first examination image than in the otherwise case. | 1. A medical image interpretation support system comprising: storage means for storing image data sets obtained by imaging patients;similar case search means for extracting similar case image data sets representing similar case images having a characteristic similar to an interpretation target image or interpretation target images from the image data sets stored in the storage means; anddisplay control means for controlling display of the interpretation target image or images and the similar case images on a monitor or monitors, the medical image interpretation support system further comprising:first examination judgment means for judging whether the interpretation target image or images is/are a first examination image/first examination images, whereinthe display control means controls the display in such a manner that the number of the similar case images displayed together with the interpretation target image or images in the case where the first examination judgment means has judged that the interpretation target image or images is/are a first examination image or first examination images is larger than the number of the similar case images displayed together with the interpretation target image or images in the case where the first examination judgment means has judged that the interpretation target image or images is/are not a first examination image or first examination images. 2. The medical image interpretation support system according to claim 1, further comprising medical information search means for extracting medical information, wherein the display control means controls the display to further display the medical information extracted by the medical information search means in the case where the first examination judgment means has judged that the interpretation target image or images is/are a first examination image or first examination images. 3. The medical image interpretation support system according to claim 1, further comprising past image search means for extracting at least one past image data set representing a past image or past images obtained in the past regarding the same patient as the interpretation target image or images from the image data sets stored in the storage means, wherein the display control means controls the display to further display the past image or images extracted by the past image search means in the case where the first examination judgment means has judged that the interpretation target image or images is/are not a first examination image or first examination images. 4. The medical image interpretation support system according to claim 3 wherein the first examination judgment means judges that the interpretation target image or images is/are a first examination image or first examination images in the case where the past image search means cannot extract the past image data set or sets. 5. The medical image interpretation support system according to claim 3, wherein the past image search mean... | Fujifilm Corporation | G06K009/00 A61B005/00
US07895060 | B1 | 2011-02-22 | Systems and methods for administration of prescription drug benefits | 2006-02-03 | The invention relates to the administration of prescription drug formulary information. A list of drugs that a particular prescription drug plan will pay for in whole or in part is called a formulary. The invention involves the use of information and customizable rules associated with formularies, combined with information from patient medical records, to generate custom, dynamic formularies applicable to specific patients, groups of patients, or both. Depending on the choice of rules and the other information, this may help payers control prescription drug costs by encouraging the use of less-expensive drugs when medically appropriate, but without impairing the freedom of prescribers to prescribe specific drugs according to their professional judgments. Some embodiments may also be adapted to generate and store data about the use and functioning of the embodiments or aspects of them and to generate reports containing some or all such data in response to queries. | 1. A computer system for processing information related to formularies, comprising: at least one processor;at least one interface, operatively coupled to the at least one processor; andat least one memory, operatively coupled to the at least one processor and encoded with instructions that, when executed by at least one of the processors, cause the computer system at least to receive via the at least one interface medical information comprising at least one medical record of a patient from a first electronic information source;receive via the at least one interface, from a second electronic information source, formulary information that is associated with the patient and comprises at least one rule and a first formulary, the first formulary comprising information about payment by a payer for one or more drugs; andgenerate a second formulary, comprising information about payment by the payer for one or more drugs, by applying the at least one rule to the at least one medical record and the first formulary. 2. The computer system of claim 1, wherein the instructions comprise instructions that, when executed by at least one of the processors, cause the computer system at least to transmit information from the second formulary via the at least one interface. 3. The computer system of claim 2, wherein the transmitted information includes first information comprising identification of at least one drug that is on the second formulary and second information comprising identification of at least one drug that is not on the second formulary. 4. The computer system of claim 3, wherein the second information comprises identification of at least one reason for exclusion of at least one of the drugs that is not on the second formulary. 5. The computer system of claim 3, wherein the second information comprises identification of at least one condition such that if the at least one condition is fulfilled, then at least one of the drugs that is not on the second formulary before fulfillment of the condition will be on the second formulary after fulfillment of the condition. 6. The computer system of claim 5, wherein at least one of the conditions includes subjecting the patient to a medical test. 7. The computer system of claim 1, wherein the instructions comprise instructions that, when executed by at least one of the processors, cause the computer system at least to: receive first prescription information via the at least one interface, the first prescription information comprising identification of a drug that has been prescribed for a patient; andin response to receiving the first prescription information, transmit second prescription information via the at least one interface, the second prescription information comprising indication of whether the drug is on the second formulary;wherein generating the second formulary occurs in response to receiving the first prescription information. 8. The computer system of claim 7, wherein the second prescriptio... | Quest Diagnostics Investments, Inc. | A61B005/00 G06F019/00
US07890163 | B2 | 2011-02-15 | Method and apparatus for detecting fibrillation using cardiac local impedance | 2006-10-19 | A cardiac rhythm management (CRM) system detects tachyarrhythmia using cardiac local impedance indicative of cardiac local wall motion. A cardiac local impedance signal indicative of an impedance of a cardiac region is sensed by using a pair of bipolar electrodes placed in that cardiac region. Tachyarrhythmia such as VF is detected by analyzing one or more cardiac local impedance signals sensed in one or more cardiac regions. | 1. A cardiac rhythm management system for treating a heart, the system comprising: an implantable medical device configured to be coupled to an implantable lead including a pair of impedance sensing electrodes configured to be placed in the heart to sense a cardiac local impedance signal indicative of a cardiac local wall motion, the implantable medical device including: an impedance sensing circuit adapted to sense the cardiac local impedance signal using the pair of impedance sensing electrodes; andan impedance-based tachyarrhythmia detector adapted to detect a predetermined-type tachyarrhythmia using the cardiac local impedance signal, the impedance-based tachyarrhythmia detector including an impedance-based ventricular fibrillation (VF) detector including: a detection zone generator adapted to produce a VF detection zone specified by one or more threshold amplitudes; anda comparator including a signal input to receive the cardiac local impedance signal, one or more threshold inputs to receive the one or more threshold amplitudes, and an output to indicate a VF detection when the amplitude of the cardiac local impedance signal falls into the VF detection zone. 2. The system of claim 1, wherein the implantable medical device further comprises: a cardiac sensing circuit to sense one or more electrograms; andan electrogram-based tachyarrhythmia detector adapted to detect the predetermined-type tachyarrhythmia using the one or more electrograms. 3. The system of claim 1, wherein the impedance sensing circuit comprises: a current source circuit adapted to deliver a current through the pair of impedance sensing electrodes;a voltage sensing circuit adapted to sense a voltage across the pair of impedance sensing electrodes; andan impedance detector adapted to produce the cardiac local impedance signal using the sensed voltage. 4. The system of claim 3, wherein the impedance detector is adapted to produce the cardiac local impedance signal as a ratio of the sensed voltage to the delivered current. 5. The system of claim 1, wherein the detection zone generator is adapted to adjust the VF detection zone based on a trend of the cardiac local impedance signal. 6. The system of claim 1, further comprising the first implantable lead, wherein the first implantable lead includes a proximal end configured to be coupled to the implantable medical device, a distal end configured to be placed in the heart and including the pair of impedance sensing electrodes, and an elongate lead body coupled between the proximal end and the distal end. 7. The system of claim 6, wherein a distance between two impedance sensing electrodes of the first pair of impedance sensing electrodes is in a range of approximately 2 millimeters to 40 millimeters. 8. The system of claim 7, wherein the distance between the two impedance sensing electrodes is within approximately 20 millimeters. 9. A cardiac rhythm management system for treating a heart, the system comprising: an implanta... | Cardiac Pacemakers, Inc. | A61B005/053
US07873397 | B2 | 2011-01-18 | Spectroscopic optical system | 2007-06-19 | Innovative techniques that result in a better signal-to-noise ratio for spectrographic analysis of substances in a target than conventional techniques. In these techniques, light illuminates a target with at least some of the light penetrating the target. At least a portion of the light that penetrates the target is collected from a region on the target's surface that is not directly illuminated. Preferably, at least a majority of the collected light is light that penetrates the target. Also preferably, the light that illuminates the target is in a pattern that partially but not completely surrounds the region from which the portion of the light that penetrates the target is collected. A spectrum of at least a portion of the collected light is analyzed. | 1. A method of performing Raman spectroscopy, comprising the steps of: illuminating a target with light from a light source, with at least some of the light penetrating the target;collecting, with collection optics, at least a portion of the light that penetrates the target from a region on the target's surface not directly illuminated by the step of illuminating, wherein at least a majority of the light collected by the step of collecting is light that penetrates the target;analyzing, with Raman spectroscopic instrumentation, a spectrum of at least a portion of the light collected by the step of collecting; andwherein the light that illuminates the target is in a pattern that partially but not completely surrounds the region from which the portion of the light that penetrates the target is collected. 2. A method as in claim 1, wherein the light that illuminates the target is in a crescent-shaped, semi-circle-shaped, hyperbolic-shaped, or parabolic-shaped pattern. 3. A method as in claim 2, wherein the crescent-shaped, semi-circle-shaped, hyperbolic-shaped, or parabolic-shaped pattern is formed using only refractive, transmissive, or refractive and transmissive optics. 4. A method as in claim 2, wherein the crescent-shaped, semi-circle-shaped, hyperbolic-shaped, or parabolic-shaped pattern is formed using one of more of refractive, transmissive, masking, or diffractive optics. 5. A method as in claim 1, wherein the step of collecting the light is performed using collection optics directly in contact with the target. 6. A method as in claim 1, wherein the light that illuminates the target comprises coherent monochromatic light. 7. A method as in claim 6, wherein the coherent monochromatic light is from a laser or laser diode. 8. A method as in claim 1, wherein the light comprises infrared light. 9. A method as in claim 8, wherein the infrared light has a wavelength in the range of 785 nanometers to 850 nanometers. 10. A method as in claim 1, wherein the target comprises a person's skin. 11. A method as in claim 10, wherein the spectrum of light is measured to determine a glucose content of one or more of tissue, interstitial fluid, and blood in the person's skin. 12. A method of performing Raman spectroscopy, comprising the steps of: illuminating a target with light from a light source, with at least some of the light penetrating the target;collecting, with collection optics, at least a portion of the light that penetrates the target from a region on the target's surface not directly illuminated by the step of illuminating;analyzing, with Raman spectroscopic equipment, a spectrum of at least a portion of the light collected by the step of collecting; andwherein the light that illuminates the target is in a pattern that partially but not completely surrounds the region from which the portion of the light that penetrates the target is collected. 13. A method as in claim 12, wherein the light that illuminates the target is in a crescent-s... | - | A61B005/1455
US20110009758 | A1 | 2011-01-13 | SYSTEM AND METHOD FOR HEART MONITORING | 2009-11-05 | Disclosed are systems and methods for monitoring a heart. According to one embodiment, the system includes a registering unit positioned external to the patient's body. The registering unit comprises a first controller configured to register an electrical signal from the heart. The system includes a second controller in operable communication with the first controller. The second controller is configured to receive the data from the first controller corresponding to the registered electrical signal and to compare the registered electrical signal to a baseline electrical signal to determine whether the heart is functioning properly. | 1. A system for monitoring a patient's heart, comprising: a registering unit located external to the patient's body, wherein the registering unit comprises a first controller in electrical communication with the patient's heart, and wherein the first controller is configured to register from the patient's heart: (i) a baseline electrical signal; and (ii) at least one registered electrical signal after registration of the registered baseline electrical signal; anda second controller in operable communication with the first controller of the registering unit, wherein the second controller receives the registered baseline electrical signal and the at least one registered electrical signal from the first controller and generates: (i) a baseline template corresponding to the baseline electrical signal; and (ii) at least one second template corresponding to the at least one registered electrical signal, and wherein the second controller is configured to compare the at least one second template to the baseline template to determine whether the patient's heart is functioning properly. 2. A system according to claim 1, further comprising a plurality of electrodes coupled to the registering unit and positioned external to the patient's body, wherein the plurality of electrodes are configured to be in electrical communication with the patient's heart and communicate the baseline electrical signal and the at least one registered electrical signal to the first controller. 3. A system according to claim 1, wherein the registering unit further comprises a transmitter in operable communication with the second controller and configured to transmit the baseline electrical signal and the at least one registered electrical signal to the second controller. 4. A system according to claim 1 wherein the registering unit is in at least one of electrical or optical communication with the second controller. 5. A system according to claim 1, wherein the second controller is located remotely from the registering unit. 6. A system according to claim 1, wherein the second controller comprises a data repository for storing data corresponding to the baseline electrical signal and the at least one registered electrical signal. 7. A system according to claim 1, further comprising a relay unit in communication with the first controller of the registering unit, wherein the relay unit is configured to receive data from the first controller corresponding to the baseline electrical signal and the at least one registered electrical signal. 8. A system according to claim 7, wherein the relay unit is further configured to transmit instructions to the registering unit and to transmit the baseline electrical signal and the at least one registered electrical signal to the second controller. 9. A system according to claim 1, wherein the second controller measures the area between the baseline template and the at least one second template to determine whether the patient's heart is fun... | LifeScience Solutions LLC | A61B005/0402
US07874985 | B2 | 2011-01-25 | Method, system, and computer program product for the evaluation of glycemic control in diabetes from self-monitoring data | 2005-12-19 | A method, system, and computer program product related to the diagnosis of diabetes, and is directed to predicting the long-term risk of hyperglycemia, and the long-term and short-term risks of severe hypoglycemia in diabetics, based on blood glucose readings collected by a self-monitoring blood glucose device. The method, system, and computer program product pertain directly to the enhancement of existing home blood glucose monitoring devices, by introducing an intelligent data interpretation component capable of predicting both HbA1c and periods of increased risk of hypoglycemia, and to the enhancement of emerging continuous monitoring devices by the same features. With these predictions the diabetic can take steps to prevent the adverse consequences associated with hyperglycemia and hypoglycemia. | 1. A computerized method for evaluating hemoglobin A1c (HbA1c) of a patient based on blood glucose (BG) data collected over a predetermined duration, said method comprising: computing on a computer weighted deviation of the BG of said patient toward high blood glucose (WR) from a predefined normal level;computing on a computer mean blood glucose (Dr); andestimating on a computer HbA1c using a predetermined mathematical formula based on said computed WR and Dr;defining predetermined categories for the estimate of HbA1c, each of said HbA1c estimate categories representing a range of values for estimated HbA1c; andassigning said estimated HbA1c to at least one of said HbA1c estimate categories. 2. The method of claim 1, wherein: said computed WR is mathematically defined from a series of BG readings x1, x2, . . . xn taken at time points t1, t2, . . . tn as:WR=1n⁢∑i=1n⁢⁢wr⁢⁢(xi,1)where: wr(BG,b)=10ƒ(BG)b if ƒ(BG)>0 and 0 otherwise,where ƒ is a symmetrization function,b=1, representing a weighting parameter, andsaid computed Dr is mathematically defined as: Dr=average of sk+1−sk,where: sk=10S(k+t1)2 for k=0, 1, . . . , tn−t1,S(t1)=ƒ(xj), for j=1, . . . , n, where xj is a BG reading of said patient at time j. 3. The method of claim 1, wherein said estimating HbA1c from said BG monitoring data comprises determining Estimated HbA1c according to: Estimated HbA1c=0.9008(WR)−0.8207(DR)+6.7489. 4. The method of claim 1, wherein said HbA1c estimate categories are defined as follows: classified category 1, wherein said estimated HbA1c is less than about 7.8;classified category 2, wherein said estimated HbA1c is between about 7.8 and about 8.5;classified category 3, wherein said estimated HbA1c is between about 8.5 and about 9.0;classified category 4, wherein said estimated HbA1c is between about 9.0 and about 9.6;classified category 5, wherein said estimated HbA1c is between about 9.6 and about 10.3;classified category 6, wherein said estimated HbA1c is between about 10.3 and about 11.0; andclassified category 7, wherein said estimated HbA1c is above about 11.0. 5. The method of claim 4, further comprising: defining predicted confidence intervals for corresponding said HbA1c estimate categories, wherein said predicted confidence intervals are defined as follows: said classified category 1 corresponds with a predicted HbA1c less than about 8.0;said classified category 2 corresponds with a predicted HbA1c between about 8.0 and about 8.5;said classified category 3 corresponds with a predicted HbA1c between about 8.5 and about 9.0;said classified category 4 corresponds with a predicted HbA1c between about 9.0 and about 9.5;said classified category 5 corresponds with a predicted HbA1c between about 9.5 and about 10.1;said classified category 6 corresponds with a predicted HbA1c between about 10.1 and about 11.0; andsaid classified category 7 corresponds with a predicted HbA1c above about 11.0. 6. The method of claim 1, further comprising: defining predicted c... | University of Virginia Patent Foundation | A61B005/00 A61B005/05
US07867171 | B2 | 2011-01-11 | Management of gastro-intestinal disorders | 2004-02-20 | Methods of determining gastro-intestinal conditions by performing a succession of breath tests or others tests, particularly for determining the gastric emptying or gastric accomodation condition of a subject. Methods of performing successive gastric emptying tests with different test meals are also presented, enabling the gastric accommodation to be determined. The effects of different test means on the results is presented. Novel substrates for use in such tests are suggested, including the use of micro-encapsulation. Breath tests for the detection of bacterial overgrowth, lactose intolerance and combinations thereof are presented. | 1. A method of determining at least a first and a second gastro-intestinal condition in a subject, the method comprising: administering to said subject a test meal comprising a labeled marker whose by-products are absorbed and exhaled in breaths of said subject after exit from the stomach of said subject;performing on said subject a breath test indicative of at least one of the first and the second gastro-intestinal conditions; andautomatically determining the first and the second gastro-intestinal conditions in said subject based on the same meal. 2. A method according to claim 1 and wherein at least one of said first and said second gastro-intestinal conditions comprises at least one of dyspepsia and irritable bowel syndrome. 3. A method according to claim 1 and wherein at least one of said first and said second gastro-intestinal conditions comprises at least one of a gastric emptying disorder, a gastric accommodation disorder, and a Helicobacterpylori infection. 4. A method according to claim 1 and wherein at least one of said first and said second gastro-intestinal conditions comprises at least one of a sugar malabsorption disorder, a bacterial overgrowth, a gastric emptying disorder, and an orececal transit time disorder. 5. A method according to claim 4 and wherein said sugar malabsorption disorder is at least one of lactose intolerance, fructose, intolerance, sucroseintolerance and maltose intolerance. 6. A method according to claim 1, wherein a type of the breath test indicative of the second gastro-intestinal condition is selected based on a result of the breath test indicative of the first gastro-intestinal condition. 7. A method according to claim 1, wherein said first gastro-intestinal condition comprises a gastric emptying disorder and said second gastro-intestinal condition comprises a gastric accommodation disorder. 8. A method according to claim 7, wherein said test meal comprises a liquid meal having a predetermined volume, at least one average gastric emptying characteristic of said meal for a large plurality of normal subjects being known, wherein determining said first gastro-intestinal conditions comprises determining at least one emptying characteristic of said meal from the stomach of the subject,and wherein determining said second gastro-intestinal conditions comprises determining the gastric accommodation of the subject according to the deviation between said at least one emptying characteristic of said meal from the stomach of said subject and said at least one average emptying characteristic of said meal for a large plurality of normal subjects. 9. A method according to claim 8, and wherein said predetermined volume is sufficient to cause gastric distension in said subject. 10. A method according to claim 8, and wherein said predetermined volume is at least 750 milliliters of liquid. 11. A method according to claim 8, and wherein said liquid meal has a gastric retention characteristic arising from at least one... | Exalenz Bioscience Ltd. | A61B005/00
US07869859 | B2 | 2011-01-11 | X-ray CT apparatus having a display to display information inputted from a remote operator console | 2003-03-10 | A medical apparatus including a main unit including an X-ray CT imaging unit configured to image a patient and a display unit provided on the X-ray CT imaging unit and configured to display a name of the patient to be imaged, and a console unit remote from the main unit and including an input unit configured to input condition information of the main unit and individual information including the name of the patient. The console also includes a control unit configured to extract at least one item including the name of the patient in the individual information from the inputted condition information and the inputted individual information. The display unit is configured to display the extracted item including the name of the patient on a screen. | 1. A medical system, comprising: a main unit including a gantry in which an X-ray tube and an X-ray detector are provided;a display unit provided on the gantry and configured to display a name of a patient to be imaged;an input device provided on the gantry adjacent to the display unit and configured to receive input from the patient that confirms whether the name displayed on the display unit is accurate or not, and configured to transmit confirmation information indicating that the name displayed on the display unit is accurate when receiving the input from the patient; anda console unit remote from the main unit and including a processor configured to control imaging by the main unit and configured to postpone the imaging by the main unit until the console unit receives the confirmation information transmitted, in response to the input from the patient, from the input device provided on the gantry indicating that the name displayed on the display unit is accurate. 2. The medical system according to claim 1, wherein the console unit is configured to extract a name of the patient from the individual information received by the input device. 3. The medical system according to claim 1, wherein the input unit is provided in a different room from a room where the main unit is provided. 4. The medical system according to claim 1, wherein the display unit is provided on the gantry in such a position that the patient to be imaged can see the display unit. 5. The medical system according to claim 1, wherein the display unit is arranged in a position on or near a peripheral portion around an aperture of the main unit. 6. The medical system according to claim 1, wherein the display unit is arranged in a same room as the main unit. 7. The medical apparatus according to claim 1, wherein the display unit includes a color LCD unit or an LED unit. 8. The medical apparatus according to claim 1, wherein the input unit receives determination information to determine received condition information and received individual information; andthe display unit starts to display an item extracted from the received individual information or the received condition information when the determination information is received. 9. The medical apparatus according to claim 1, further comprising a voice generation unit configured to generate voice information of an item extracted from information received by the input device. 10. The medical apparatus according to claim 1, wherein the display unit ceases to display an item extracted from information received by the input device when treating or imaging of the patient is completed. 11. The medical apparatus according to claim 1, wherein the display unit displays situation information of treating or imaging of the patient when the treating or the imaging of the patient starts. 12. The medical apparatus according to claim 1, further comprising a memory unit configured to store an item extracted from information received by ... | Kabushiki Kaisha Toshiba | A61B005/05
US07871376 | B2 | 2011-01-18 | System and method for monitoring a physiological condition | 2007-05-14 | Method for treating a medical condition in a human patient comprising choosing a psychological strategy for treating the medical condition, encoding electronic instructions for an interactive video game in such a way that the interactive video game implements the psychological strategy, loading the electronic instructions into a microprocessor-based unit equipped with a display for displaying the interactive video game and with an patient input device for receiving responses to the interactive video game from the human patient, and instructing the human patient how and when to use the microprocessor-based unit to play the interactive video game. The interactive video game contains instructions for a scoring procedure for quantitatively analyzing the medical condition of the human patient, and/or counseling instructions or self-care instructions. The video game can be used in conjunction with a physical parameter measuring device connected to the microprocessor-based unit. | 1. A diabetes management system, comprising: (a) a blood glucose meter measuring blood glucose level data; and(b) a portable microprocessor-based unit in communication with the blood glucose meter such as to be capable of accessing said blood glucose level data, wherein said portable microprocessor-based unit comprises a multi-line display screen configured to present to a user (i) graphs of said blood glucose level data and (ii) a game-like presentation;(c) wherein input data corresponding to said blood glucose level data is used in a set of program instructions running on the portable microprocessor-based unit; and(d) wherein said set of program instructions include instructions to (i) send a signal to inject insulin based upon said blood glucose level data, (ii) present educational information to said user of said system regarding a health condition of said user, (iii) present to said user said blood glucose level data, (iv) receive from said user responses confirming said injection of said insulin, and (v) positively reinforce compliance with a diabetes management plan by providing bonuses in said game-like presentation on said multi-line display screen in response to (a) said blood glucose level data being within an optimal range and (b) said confirmation of said injection of said insulin. 2. The system of claim 1, wherein said set of program instructions include (i) instructions comprising a treatment plan, and (ii) instructions to send a signal to advance between treatment steps of said treatment plan. 3. The system of claim 1, wherein said insulin is injected by said user. 4. The system of claim 3, wherein said user uses a syringe to inject said insulin. 5. The system of claim 1, said set of program instructions including instructions to display an indicator to check a blood glucose level on said multi-line display screen of the portable microprocessor-based unit, wherein said multi-line display screen is configured to present a plurality of lines of alphanumeric characters. 6. The system of claim 1, said set of program instructions including instructions to display an indicator to select an insulin plan on said multi-line display screen of the portable microprocessor-based unit, wherein said multi-line display screen is configured to present a plurality of lines of alphanumeric characters. 7. The system of claim 1, said set of program instructions including instructions to display an indicator to get a menu of foods to eat on said multi-line display screen of the portable microprocessor-based unit. 8. The system of claim 1, said portable microprocessor-based unit being connectable with a remote communication unit. 9. The system of claim 1, said portable microprocessor-based unit being connectable with a computer in a hospital, wherein said computer is remotely located from said portable microprocessor-based unit. 10. The system of claim 9, said portable microprocessor-based unit being connectable with said computer via a telepho... | Health Hero Network, Inc. | A61B005/00
US08506496 | B2 | 2013-08-13 | Method and apparatus to determine the end of the systolic part of a pressure curve | 2008-11-24 | The invention relates to a method and an apparatus for determining the systolic phase interval (SP) of an arterial pressure curve with a starting point (t0) of the systolic phase interval (SP) and an end point (tN) of the systolic phase interval (SP) wherein a QT interval is defined as the start of the Q wave and the end of the T wave in the hearts electrical cycle and wherein the determination of the end point (tN) of the systolic phase interval (SP) is restricted to the measured points of arterial pressure which fulfill the condition that the difference in time between the end point (tN) to be determined of the systolic phase interval (SP) and the starting point (t0) of the systolic phase interval (SP) is smaller than the QT interval. | 1. A method for computing a systolic phase interval of an arterial pressure curve, comprising: determining, by a processor, a time interval between a starting point (t0) of the systolic phase interval and a time instant (tk), wherein tk is when a mean arterial pressure is reached for a first time following the starting point (t0);calculating, by the processor, a product of a factor and a QT interval, the QT interval being the time between a start of a Q wave and an end of a T wave in an electrical cycle of a heart; andsumming, by the processor, the time interval and the product to obtain an end point (tN) of the systolic phase interval. 2. The method of claim 1, wherein the factor is a value between 0.25 and 0.8. 3. The method of claim 1, wherein the factor is a square root of (SAP-MAP)/(SAP-P 0), wherein SAP is a maximum arterial pressure, MAP is the mean arterial pressure, and P0 is a base line pressure. 4. An apparatus comprising: an input port to receive signals representing an arterial pressure;a processor coupled to the input port for computing a systolic phase interval of an arterial pressure curve, the processor adapted to: determine a time interval between a starting point (t0) of the systolic phase interval and a time instant (tk), wherein tk is when a mean arterial pressure is reached for a first time following the starting instant point (t0);calculate a product of a factor and a QT interval, the QT interval being the time between a start of a Q wave and an end of a T wave in an electrical cycle of a heart; andsum the time interval and the product to obtain an end point (tN) of the systolic phase interval. 5. The apparatus of claim 4, wherein the factor is a value between 0.25 and 0.8. 6. The apparatus of claim 4, wherein the factor is a square root of (SAP-MAP )/(SAP-P0), wherein SAP is a maximum arterial pressure, MAP is the mean arterial pressure, and P0 is a base line pressure. | Edwards Lifesciences IPRM AG | A61B005/02 A61B005/04
US10542962 | B2 | 2020-01-28 | Adaptive radiotherapy treatment using ultrasound | 2010-07-07 | Radiation treatment is delivered to a patient by positioning the patient such that a radiation beam is delivered to a lesion within the patient along a beam-delivery path while securing a diagnostic imaging device about the patient such that the diagnostic imaging device does not intersect the beam-delivery path. Radiation therapy is simultaneously delivered along the beam-delivery path while diagnostic images are obtained using the imaging device. | 1. A method of monitoring a target tissue in a patient during a radiation treatment, comprising: positioning an ultrasound imaging device such that the ultrasound imaging device is in contact with the patient's perineum;using the ultrasound imaging device to scan, transperineally, a target tissue in a genitourinary region of the patient to obtain, prior to a radiation fraction, a baseline image of the target tissue and surrounding tissue;delivering the radiation fraction along at least one beam-delivery path to the target tissue using a radiation therapy device, wherein the at least one beam-delivery path is directed to the patient's anatomy in a non-transperineal direction;simultaneously delivering the radiation fraction while obtaining a plurality of ultrasound images using the ultrasound imaging device, wherein the ultrasound imaging device is positioned to obtain, from at least one transperineal direction, the plurality of ultrasound images, such that the ultrasound imaging device does not intersect the at least one beam-delivery path while obtaining the plurality of ultrasound images;comparing the one or more of the plurality of ultrasound images with the baseline image; anddetermining intrafractional movement of at least one of the target tissue or the surrounding tissue based on the compared one or more ultrasound images as the radiation fraction is delivered to the patient by the radiation therapy device. 2. The method of claim 1, wherein obtaining the plurality of ultrasound images comprises obtaining at least one two-dimensional ultrasound image. 3. The method of claim 1, wherein obtaining the plurality of ultrasound images comprises obtaining at least a three-dimensional ultrasound image. 4. The method of claim 3, wherein the ultrasound imaging device further comprises one or more tracking devices to facilitate real-time positional tracking of the ultrasound imaging device. 5. The method of claim 4, further comprising combining a plurality of temporally-displaced three-dimensional ultrasound images to create a four-dimensional ultrasound image. 6. The method of claim 1, wherein the ultrasound imaging device is movable within a predetermined envelope to obtain the plurality of ultrasound images from different respective locations. 7. The method of claim 6, wherein the ultrasound imaging device is attached to a brace, connected to a stationary structure, for holding the ultrasound imaging device at each of the different locations. 8. The method of claim 7, further comprising adjusting a position of the brace in response to a change in one or more radiation treatment parameters. 9. The method of claim 8, wherein the one or more radiation treatment parameters comprises at least one of a beam angle, a beam isocenter, a couch position, a beam aperture, a beam delivery path, a focal point, an energy level, or a duration of a radiation fraction. 10. The method of claim 7, further comprising rotating the brace such that the ultrasound ... | Elekta, LTD | A61B008/00
US07880717 | B2 | 2011-02-01 | Method, apparatus, and article for force feedback based on tension control and tracking through cables | 2004-03-26 | A haptic device for human/computer interface includes a user interface tool coupled via cables to first, second, third, and fourth cable control units, each positioned at a vertex of a tetrahedron. Each of the cable control units includes a spool and an encoder configured to provide a signal corresponding to rotation of the respective spool. The cables are wound onto the spool of a respective one of the cable control units. The encoders provide signals corresponding to rotation of the respective spools to track the length of each cable. As the cables wind onto the spools, variations in spool diameter are compensated for. The absolute length of each cable is determined during initialization by retracting each cable In turn to a zero length position. A sensor array coupled to the tool detects rotation around one or more axes. | 1. A haptic interface device to provide haptic interaction to a user manipulating a tool, the haptic interface device comprising: an attachment point;a plurality of cables coupled to the attachment point, the plurality of cables comprising a first, a second, a third, and a fourth length of cable extending from the attachment point;a first tool translation effecter device having a first cable guide, a second tool translation effecter device having a second cable guide, a third tool translation effecter device having a third cable guide, and a fourth tool translation effecter device having a fourth cable guide, the first, the second, the third and the fourth cable guides positioned, relative to each other, such that each of the first, the second, the third, and the fourth cable guides occupies a vertex of a tetrahedron, each tool translation effecter device having coupled thereto, a respective one of the first, the second, the third, and the fourth lengths of cable such that, as the attachment point moves, each of the first, the second, the third, and the fourth lengths of cables is retracted or paid out accordingly by the respective tool translation effecter device, each tool translation effecter device including controlling means for selectively varying an active tension on the respective length of cable;metering means for metering each of the first, the second, the third, and the fourth length of cables as they are retracted and paid out; andestablishing means for establishing a respective distance between each of the first, the second, the third, and the fourth cable guides and the attachment point, the establishing means including a memory configured to receive and store, prior to a complete shutdown of the haptic interface device, the respective distances between each of the first, the second, the third, and the fourth tool translation effecter devices and the attachment point, and to provide the stored distances during a startup procedure. 2. The haptic interface device of claim 1 wherein: the controlling means of each of the first, the second, the third, and the fourth tool translation effecter devices includes a spool and a motor coupled to rotatably drive the spool, the motor and spool selectively operable to wind and unwind the respective length of cable; andthe metering means includes: counting means for counting fractions of rotations of the spool of each of the first, the second, the third, and the fourth tool translation effecter devices. 3. The haptic interface device of claim 1 wherein the establishing means includes a controller configured to direct the first tool translation effecter device to retract, during an initialization procedure, the first length of cable until the attachment point is at a selected position relative to the first tool translation effecter device. 4. The haptic interface device of claim 1 wherein the controlling means of each of the first, the second, the third, and the fourth tool translation effecter d... | Mimic Technologies, Inc. | G05B011/01 G05B015/00 G09G005/00 G06F019/00 A61B005/103 G06G007/48
US07873402 | B2 | 2011-01-18 | System and method for radar-assisted catheter guidance and control | 2007-10-09 | A Catheter Guidance Control and Imaging (CGCI) system whereby a magnetic tip attached to a surgical tool is detected, displayed and influenced positionally so as to allow diagnostic and therapeutic procedures to be performed is described. The tools that can be so equipped include catheters, guidewires, and secondary tools such as lasers and balloons. The magnetic tip performs two functions. First, it allows the position and orientation of the tip to be determined by using a radar system such as, for example, a radar range finder or radar imaging system. Incorporating the radar system allows the CGCI apparatus to detect accurately the position, orientation and rotation of the surgical tool embedded in a patient during surgery. In one embodiment, the image generated by the radar is displayed with the operating room imagery equipment such as, for example, X-ray, Fluoroscopy, Ultrasound, MRI, CAT-Scan, PET-Scan, etc. In one embodiment, the image is synchronized with the aid of fiduciary markers located by a 6-Degrees of Freedom (6-DOF) sensor. The CGCI apparatus combined with the radar and the 6-DOF sensor allows the tool tip to be pulled, pushed, turned, and forcefully held in the desired position by applying an appropriate magnetic field external to the patient's body. A virtual representation of the magnetic tip serves as an operator control. This control possesses a one-to-one positional relationship with the magnetic tip inside the patient's body. Additionally, this control provides tactile feedback to the operator's hands in the appropriate axis or axes if the magnetic tip encounters an obstacle. The output of this control combined with the magnetic tip position and orientation feedback allows a servo system to control the external magnetic field. | 1. A method for controlling movement of a tool having a distal end to be inserted in a body, comprising; applying a force to said distal end of a tool inserted in a body by generating an external magnetic field;regulating said force to move said distal end in a desired direction within said body; andlocating said distal end in said body by radar by processing radar data to identify contrast between said distal end and body tissues near said distal end, said contrast resulting from the dielectric constants of said distal end and said body tissues wherein said processing of said radar data filters out the background clutter associated with said body tissues based on said dielectric constants to determine the position coordinates of said distal end. 2. The method of claim 1, further comprising changing a visual representation of said distal end in substantially real time as said distal end moves through the body. 3. The method of claim 1, further comprising controlling one or more electromagnets to produce said external magnetic field. 4. The method of claim 1, further comprising locating a plurality of fiduciary markers and synchronizing said markers with positions on a real-time image of at least a portion of the body. 5. The method of claim 1, further comprising determining a current position of said distal end in comparison to a desired location. 6. The method of claim 1, wherein determining said current position of said tool distal end comprises: inputting a dynamic cardio-position via said controller; andcalculating said current position as a function of said cardio-position. 7. The method of claim 1, further comprising computing a position error of said distal end. 8. The method of claim 7, further comprising altering at least one of a duty cycle and a polarity of modulation inputs to at least one of said X-axis controller, said Y-axis controller, and said Z-axis controller when said position error is greater than a specified minimum value. 9. The method of claim 7, further comprising producing a tactile feedback if said position error exceeds a predetermined amount along at least one axis. 10. The method of claim 7, wherein said system controller causes said tool distal end to move so that its position corresponds to position data from a Virtual Tip. | Magnetecs, Inc. | A61B005/05
US20110021896 | A1 | 2011-01-27 | Biosensor Membranes Composed of Polymers Containing Heterocyclic Nitrogens | 2010-09-30 | Novel membranes comprising various polymers containing heterocyclic nitrogen groups are described. These membranes are usefully employed in electrochemical sensors, such as amperometric biosensors. More particularly, these membranes effectively regulate a flux of analyte to a measurement electrode in an electrochemical sensor, thereby improving the functioning of the electrochemical sensor over a significant range of analyte concentrations. Electrochemical sensors equipped with such membranes are also described. | 1. (canceled) 2. A sensor for continuous measurement of an analyte concentration, the sensor comprising: an electrode and a membrane located over the electrode, the membrane comprising: a first domain comprising a polymer having a surface active group, the first domain configured to control a flux of an analyte therethrough and to substantially reduce or block a flux of an endogenous interferent therethrough; anda second domain comprising an enzyme. 3. The sensor of claim 2, wherein the first domain is configured to substantially reduce or block a flux of an exogenous interferent therethrough. 4. The sensor of claim 2, wherein the polymer comprises a blend of a base polymer and a hydrophilic polymer. 5. The sensor of claim 4, wherein the hydrophilic polymer is a polymer selected from the group consisting of polyvinyl acetate, poly(ethylene glycol), polyacrylamide, acetates, polyethylene oxide, poly ethyl acrylate, and polyvinylpyrrolidone. 6. A device for continuous measurement of an analyte concentration, the device comprising: a sensing mechanism configured to generate a signal associated with a concentration of an analyte in a host; anda sensing membrane located over the sensing mechanism, the sensing membrane comprising a bioprotective domain comprising a polymer comprising a surface-active group incorporated therein. 7. The device of claim 6, wherein the surface-active group is covalently bonded to the polymer. 8. The device of claim 6, wherein the surface-active group comprises a surface-active end group. 9. The device of claim 6, wherein the polymer is a polymer selected from the group consisting of polyvinyl acetate, poly(ethylene glycol), polyacrylamide, acetates, polyethylene oxide, poly ethyl acrylate, and polyvinylpyrrolidone. 10. The device of claim 6, wherein the analyte is glucose. 11. The device of claim 6, wherein the sensing mechanism comprises an electrode. 12. The device of claim 6, wherein the bioprotective domain is configured to control a flux of the analyte therethrough. 13. The device of claim 6, wherein the sensing membrane further comprises a resistance domain configured to control a flux of the analyte therethrough. 14. The device of claim 6, wherein the sensing membrane further comprises an enzyme domain comprising a catalyst. 15. The device of claim 14, wherein the catalyst is incorporated into the bioprotective domain. 16. The device of claim 6, wherein the polymer comprises at least one polymer selected from the group consisting of epoxies, polyolefins, polysiloxanes, polyethers, acrylics, polyesters, carbonates, and polyurethanes. 17. The device of claim 6, wherein the surface-active group comprises silicone. 18. The device of claim 6, wherein the bioprotective domain is configured to control a flux of an analyte therethrough and to substantially reduce or block a flux of an endogenous interferent therethrough. 19. A device for continuous measurement of glucose concentration, the device comprisin... | Abbott Diabetes Care Inc. | A61B005/1468 G01N027/40 C12Q001/00 G01N027/327 C12M001/34
US20110040194 | A1 | 2011-02-17 | METHOD AND SYSTEM FOR DETERMINING CARDIAC PARAMETERS | 2009-08-12 | The present invention explains a method and system for providing all the relevant cardiac parameters in real time and in a fast manner, which are required for the cardiac analysis. The invention involves determining a first area and a second area, the first area under a waveform representing an aortic pressure, and extends between an onset of the systole and the end of the systole and the second area extends between a waveform representing a left ventricle pressure and a waveform adapted to represent the left atrium pressure of the same systole. The method further involves presenting a functional relationship between the first area and the second area on the one hand side and the cardiac parameters on the other hand side and then finally determining the cardiac parameters based on the first area and the second area by using the representation of the functional relationship. | 1. A method for determining a cardiac parameter, comprising: determining a first area under a waveform representing an aortic pressure, wherein the first area extends between an onset of a systole and an end of the systole;providing a representation of a functional relationship between the first area and the cardiac parameter; anddetermining the cardiac parameter based on the first area by using the representation of the functional relationship. 2. The method according to claim 1, further comprising: determining a second area, wherein aid second area extends between a waveform representing a left ventricle pressure and a waveform adapted to represent a left atrium pressure in the systole;providing the representation of the functional relationship between the first area and the second area and the cardiac parameter; anddetermining the cardiac parameter based on the first area and the second area by using the representation of the functional relationship,wherein the first area represents a stroke volume (SV) and the second area represents a left ventricle systolic volume (LVSV). 3. The method according to claim 1, wherein the representation of functional relationship is a database containing a plurality of entries for the first area and for the cardiac parameter. 4. The method according to claim 1, wherein the representation of functional relationship is a function which maps the first area to the cardiac parameter. 5. The method according to claim 1, wherein the representation of functional relationship is established by a series of measurements of the first area and the cardiac parameter independently. 6. The method according to claim 1, further comprises finding an additional cardiac parameter, left ventricle end systolic volume (ESV), calculated by the equation, ESV=LVSV−SVwherein, LVSV is left ventricle systolic volume and SV is stroke volume. 7. The method according to claim 1, further comprises finding an additional cardiac parameter, left ventricle end diastolic volume (EDV), calculated by the equation, EDV=SV+ESVwherein, SV is stroke volume and ESV is left ventricle end systolic volume. 8. The method according to claim 1, further comprises finding an additional cardiac parameter, cardiac output (CO), calculated by the equation, CO=SV*HRwherein SV is stroke volume (SV) X heart rate (HR). 9. The method according to claim 1, further comprises finding an additional cardiac parameter, ejection fraction, calculated by the equation,EF=[(EDV-ESV)EDV]*100wherein EDV is left ventricle end diastolic volume and ESV is left ventricle end systolic volume. 10. A method for determining a cardiac parameter, comprising: determining a second area, wherein the second area extends between a waveform representing a left ventricle pressure and a waveform adapted to represent a left atrium pressure in a systole;providing a representation of a functional relationship between the second area and the cardiac parameter; anddetermining the cardiac parameter b... | - | A61B005/021
US20110028806 | A1 | 2011-02-03 | REFLECTANCE CALIBRATION OF FLUORESCENCE-BASED GLUCOSE MEASUREMENTS | 2009-07-29 | A noninvasive or minimally invasive procedure and system for measuring blood glucose levels is disclosed. A set of photodiodes detects the fluorescence and reflectance of light energy emitted from one or more emitters, such as LEDs, into a patient's skin. In an embodiment, small molecule metabolite reporters (SMMRs) that bind to glucose are introduced to the measurement area to provide more easily detected fluorescence. | 1. A method of measuring a glucose level comprising: emitting a first excitation wavelength;measuring a first fluorescence intensity;measuring a first reflectance intensity;calibrating said first fluorescence intensity measurement with said first reflectance intensity measurement; andcalculating a glucose level with said calibrated first fluorescence intensity measurement. 2. The method of claim 1, wherein said first reflectance intensity is approximately equal to said excitation wavelength. 3. The method of claim 1, wherein said first excitation wavelength is emitted from an LED at between 320-390 nm. 4. The method of claim 1, wherein said first fluorescence intensity is measured with a band-pass filter at one or both of: approximately 420 nm or approximately 440 nm. 5. The method of claim 1, wherein said first reflectance intensity is measured with a short-pass filter at approximately 400 nm. 6. The method of claim 1, wherein said first fluorescence intensity is measured from the fluorescence of a Small Molecule Metabolite Reporter (SMMR) that binds to glucose molecules. 7. The method of claim 6, wherein said SMMR is adapted to be injected into the skin with a micro-needle. 8. The method of claim 1, wherein said method of measuring a glucose level additionally comprises: emitting a second excitation wavelength;measuring a second fluorescence intensity;measuring a second reflectance intensity;correcting a measurement error in said first glucose level with one or more of said second fluorescence intensity and said second reflectance intensity. 9. The method of claim 8, wherein said second fluorescence intensity is measured at a skin location untreated with a Small Molecule Metabolite Reporter (SMMR). 10. An apparatus for measuring a glucose level comprising: an excitation module that emits an excitation signal adapted to activate a fluorophore in skin;a reflectance measurement module that measures the intensity of one or more wavelengths reflected by said skin from said excitation module;a fluorescence measurement module that measures the intensity of one or more wavelengths emitted by said fluorophore;a glucose calculation module that accepts reflectance intensity data from said reflectance measurement module and fluorescence intensity data from said fluorescence measurement module, wherein said glucose calculation module calculates an in vivo glucose level from said reflectance intensity data and said fluorescence intensity data;a system controller that accepts said in vivo glucose level from said glucose calculation module; anda display adapted to convey said glucose level to a user. 11. The apparatus of claim 10, wherein said excitation module comprises an LED emitting an excitation wavelength between 320-390 nm. 12. The apparatus of claim 10, wherein said reflectance measurement module comprises a short-pass filter at approximately 400 nm. 13. The apparatus of claim 10, wherein said fluorescence measurement module comprises a ban... | - | A61B005/1455
US08055334 | B2 | 2011-11-08 | Evaluation of gastrointestinal function using portable electroviscerography systems and methods of using the same | 2009-12-10 | Portable electroviscerography systems are provided. Aspects of the systems include a body-associated receiver and an ingestible identifier unit, both of which are configured to detect visceral electrical signals. The system communicates with an extra-corporeal data processor configured to receive data from the body-associated receiver and generate and electroviscerogram from the received data. Also provided are methods of producing electroviscerograms using systems of the present invention. | 1. A system for detection of myoelectric activity, the system comprising: a receiver unit capable of detecting voltage potential changes, wherein the receiver unit comprises a transmission module for sending encoded information to a remote location and at least one sensor electrode to gather information from a proximal environment;an ingestible identifier in communication with the receiver unit, wherein the identifier comprises: a housing including a first electrode and a second electrode secured thereto;a control unit electrically coupled to the first and second electrodes, wherein the control unit changes the conductance between the first and second electrodes; anda detection unit electrically coupled to the first and second electrodes, wherein the detection unit reads a surrounding environment to collect information,wherein electrochemical reactions between the first electrode and the surrounding environment and the second electrode and the surrounding environment respectively produce a voltage potential that powers the ingestible identifier such that the control unit switches the ingestible identifier between broadcast mode and detection mode; and a second ingestible identifier wherein the second ingestible identifier communicates with the ingestible identifier such that when the second ingestible identifier is in broadcast mode, the ingestible identifier is in detection mode thereby allowing the ingestible identifiers to communicate. 2. The system of claim 1, further comprising a second receiver unit that includes at least three electrodes and wherein each of the electrodes of the second receiver unit transmits a control signal that is detected by each of the ingestible identifiers to create a detection array within the subject wherein the location of each ingestible identifier can be determined. 3. The system of claim 1, wherein the ingestible identifier measures acidic conditions and the second ingestible identifier measures alkaline conditions. 4. The system of claim 1, wherein the system further comprises an extra-corporeal data processor configured to receive data from the receiver unit and generate an electroviscerogram from the received data. 5. The system of claim 1, wherein the ingestible identifier includes a disc secured to the housing, wherein the disk is non-conducting and wherein during the detection mode the ingestible identifier records data associated with myoelectric activity and during the broadcast phase the ingestible identifier produces a unique current signature including the data. 6. The system of claim 1, wherein during the detection mode the ingestible identifier detects and records data associated an implantable device and during the broadcast phase the ingestible identifier produces a unique current signature encoded with the data. 7. The system of claim 1, wherein the information encoded in the unique current signature produced by the ingestible identifier is identification information to indicate activatio... | Proteus Biomedical, Inc. | A61B005/04
US07887494 | B2 | 2011-02-15 | Fluid sample transport devices and methods | 2005-09-30 | Arrangements are provided including a fluid transport tube, or a needle, having a first end and a second end opposite the first end, and a lumen having an inner diameter. At least one fluid transport enhancing projection is disposed in the lumen and extends from the second end toward the first end. A discreet, wearable blood glucose monitor including such arrangements is also described. | 1. An arrangement comprising: a needle comprising a first end and a second end opposite the first end, and a lumen having an inner diameter and a longitudinal length; andat least one fluid transport enhancing projection comprising an outer peripheral surface and extending toward the first end;wherein the at least one projection extends from the base and is constructed and arranged to enhance fluid transport by enhancing capillary forces drawing fluid along the outer peripheral surface of the at least one projection. 2. The arrangement of claim 1, wherein the first end of the needle is constructed for piercing the skin. 3. The arrangement of claim 1, wherein the needle is formed from a metal, and the at least one projection is formed from a plastic or a metal. 4. The arrangement of claim 1, further comprising a base having a bore disposed therein extending from a first surface of the base through a second surface of the base, and at least the second end of the needle being received by the bore. 5. The arrangement of claim 4, wherein the bore extends beyond the second end of the needle. 6. The arrangement of claim 5, wherein the portion of the bore extending beyond the second end of the needle has an inner diameter that is substantially the same as the inner diameter of the lumen. 7. The arrangement of claim 4, wherein the second end of the needle is disposed such that it is substantially coplanar with the second surface. 8. The arrangement of claim 1, wherein the projection terminates before extending past the second end of the needle and into the lumen, or extends into the lumen and up to 100% of the longitudinal length of the needle. 9. The arrangement of claim 1, wherein the at least one projection comprises an integral extension of the base. 10. The arrangement of claim 1, wherein the at least one projection comprises at least one wire. 11. The arrangement of claim 1, further comprising an analyte quantification member in fluid communication with the bore. 12. The arrangement of claim 11, wherein the quantification member is in direct fluid communication with the bore. 13. The arrangement of claim 11, wherein the quantification member comprises an assay pad containing a chemical reagent chosen to react with a predetermined analyte. 14. The arrangement of claim 12, further comprising a cover overlying the quantification member. 15. The arrangement of claim 14, wherein the cover is constructed to permit optical communication with the quantification member. 16. The arrangement of claim 8, wherein the projection extends past the second end of the needle and into the lumen. 17. An integrated analyte detector device comprising the arrangement of claim 1. 18. The device of claim 17, wherein the device is configured to be wearable or handheld. 19. The device of claim 18, wherein the device is constructed to perform at least one blood glucose concentration measurement. 20. The device of claim 19, wherein the device comprises a dispos... | Intuity Medical, Inc. | A61B005/00 A61B005/05
US07890345 | B2 | 2011-02-15 | Establishment of a telephone based engagement | 2008-04-18 | In general, in one aspect, a computer-implemented method for establishing multiple communication channels between a service provider and a consumer includes establishing a text based communication between the consumer and the service provider over the Internet. The method also includes receiving, at a central system, a first telephone number provided by the consumer via a first user interface and a second telephone number provided by the service provider via a second user interface. The method also includes calling, from the central system, the first telephone number and the second telephone number and bridging the calls between the central system and the first telephone number and the central system and the second telephone number to establish a communication channel between the service provider and the consumer. | 1. A computer-implemented method, the method comprising: establishing, by a computer, a first communication channel as a real-time, text-based communication between a consumer and a service provider; and subsequent to establishing the first channel,establishing a second, different voice-based communication channel between the service provider and the consumer, by: receiving, at a central computer based system, a first telephone number provided by the consumer via a first user interface and a second telephone number provided by the service provider via a second user interface;calling, from the central computer based system, the first telephone number to establish a call between the central system and the first telephone number;calling, from the central computer based system, the second telephone number to establish a call between the central system and the second telephone number; andbridging the calls between the central system and the first telephone number and the central system and the second telephone number to establish the second, different voice-based communication channel between the service provider and the consumer. 2. The method of claim 1, wherein bridging the calls comprises establishing the communication second channel between the service provider and the consumer without disclosing the first telephone number to the service provider and without disclosing the second telephone number to the consumer. 3. The method of claim 1, further comprising: providing a first authentication number to the consumer via the first user interface;providing a second authentication number to the service provider via the second user interface; andreceiving authentication numbers from the consumer and the service provider via the telephone and validating the first and second authentication numbers, prior to bridging the calls. 4. The method of claim 1, further comprising: enabling the consumer and the service provider to provide an availability indication; andinitiating a telephone conference between the consumer and the service provider when the availability indications for the consumer and the service provider indicate that the consumer and the service provider are both available. 5. The method of claim 1, further comprising setting a time limit for the telephone call. 6. The method of claim 1, further comprising receiving a request from one of the consumer and the service provider to establish a telephone based communication. 7. The method of claim 1, wherein the first real-time communication channel comprises an instant messaging communication channel. 8. The method of claim 5, further comprising providing an indication to at least one of the consumer and the service provider at a time near the time limit for the telephone call. 9. The method of claim 5, further comprising allowing the consumer and the service provider to extend the length of the telephone call via the first and second user interfaces. 10. A computer program product residing ... | American Well Corporation | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07947378 | B2 | 2011-05-24 | Medical assembly suitable for long-term implantation and method for fabricating the same | 2010-10-18 | A process for fabricating a medical assembly having a medical device at least a portion of which is formed of inorganic material is provided. The medical assembly is suitable for substantially long-term implantation in a host animal. The process includes modifying a surface of the medical device to form a hydrophilic adhesion-promoting surface. The hydrophilic adhesion-promoting surface is coated with an alginate solution comprising alginate and the alginate is reacted with alkaline earth metal cations. | 1. A medical assembly suitable for substantially long-term implantation in a host animal, the medical assembly comprising: a medical device at least a portion of which is formed of inorganic material;a hydrophilic adhesion-promoting surface comprising polyvinyl pyrrolidone formed on said medical device; anda gelled alginate coating overlying said hydrophilic adhesion-promoting surface, wherein the medical assembly relates to a sensor. 2. The medical assembly of claim 1, said medical device having an oxidized surface underlying said hydrophilic adhesion-promoting surface. 3. The medical assembly of claim 1, said hydrophilic adhesion-promoting surface subjected to plasma bombardment. 4. The medical assembly of claim 1, said gelled alginate coating comprising an additional biocompatible material subjected to plasma bombardment. 5. The medical assembly of claim 1, said gelled alginate formed from a sodium alginate solution. 6. The medical assembly of claim 1, said gelled alginate formed by applying an alginate solution to said hydrophilic adhesion-promoting surface and contacting said alginate solution with a gelling solution comprising alkaline earth metal cations. 7. The medical assembly of claim 6, said alkaline earth metal cations comprising at least one of calcium and barium ions. 8. The medical assembly of claim 6, said gelled alginate formed by applying said alginate solution to said hydrophilic adhesion-promoting surface, contacting said alginate solution with a gelling solution comprising calcium chloride and then contacting said alginate solution with a gelling solution comprising barium chloride. 9. The medical assembly of claim 1 wherein the sensor being one of a pressure sensor, pH sensor, temperature sensor or flow sensor. 10. A medical assembly suitable for substantially long-term implantation in a host animal, the medical assembly comprising: a medical device at least a portion of which is formed of inorganic material;a hydrophilic adhesion-promoting surface comprising polyvinyl pyrrolidone formed on said medical device; anda gelled alginate coating overlying said hydrophilic adhesion-promoting surface, wherein the medical assembly relates to a biosensor, an actuator, a catheter, a stent, a fluid flow control valve, a microelectromechanical system, and a nanoelectromechanical system. 11. The medical assembly of claim 10 wherein the fluid flow control valve being a cerebral spinal fluid flow control valve. | Medtronic, Inc. | B32B027/30
US07991457 | B2 | 2011-08-02 | Discrimination of supraventricular tachycardia and ventricular tachycardia events | 2010-10-11 | A method and system for discrimination of supraventricular tachycardia and ventricular tachycardia events. Morphological features points are extracted from normal sinus rhythm (NSR) complexes and used to generate a NSR template. A numerical convolution is performed using the NSR template and the feature points for each sensed NSR to give a NSR filter output. Using a plurality of NSR complexes, a median NSR filter output template is determined, where the median NSR filter output template has a median value for each value in the NSR filter output. The median NSR filter output template is then used during a tachycardia event to distinguish tachycardia events as either ventricular tachycardia events or supraventricular tachycardia events. | 1. A method, comprising: storing a normal sinus rhythm (NSR) output template;sensing cardiac signals including tachycardia complexes;extracting tachycardia feature points from the tachycardia complexes;generating a tachycardia complex output using the NSR output template and the tachycardia feature points for each of the tachycardia complexes;comparing the tachycardia complex output to the NSR output template; andclassifying the tachycardia complexes as either ventricular tachycardia (VT) complexes or supraventricular tachycardia (SVT) complexes using an outcome of the comparison of the tachycardia complex output to the NSR output template. 2. The method of claim 1, further comprising: determining whether a number of the tachycardia complexes classified as the VT complexes exceeds a predetermined threshold; anddeclaring a VT event when the number of the tachycardia complexes classified as the VT complexes exceeds the predetermined threshold. 3. The method of claim 2, further comprising delivering a ventricular tachycardia therapy in response to the VT event being declared. 4. The method of claim 1, further comprising: sensing cardiac signals including NSR complexes; anddetermining the NSR output template using the NSR complexes. 5. The method of claim 4, wherein determining the NSR output template using the NSR complexes comprises: locating NSR feature points on the NSR complexes;generating a NSR template using the NSR feature points;creating a filter impulse response (h(t)) using the NSR template;performing a numerical convolution of the NSR template and the NSR feature points to create a NSR filter output; andcreating the NSR output template using the NSR filter output. 6. The method of claim 5, wherein generating the NSR template comprises generating a NSR template including an average value for each of the NSR feature points. 7. The method of claim 5, further comprising receiving cardiac rate signals and detecting a tachycardia event using the cardiac rate signals, wherein the tachycardia complexes are associated with the tachycardia event. 8. The method of claim 5, further comprising delivering appropriate cardioversion and defibrillation therapy for the classified tachycardia complexes. 9. The method of claim 8, wherein generating the tachycardia complex output comprises performing a numerical convolution of the NSR template with tachycardia feature points. 10. The method of claim 5, further comprising detecting the tachycardia complexes and the NSR complexes, wherein the tachycardia complexes and the NSR complexes are each a QRS-complex. 11. The method of claim 1, wherein sensing the cardiac signals including the tachycardia complexes comprises sensing cardiac signals including the tachycardia complexes and NSR complexes, and further comprising: receiving cardiac signals including NSR complexes;extracting a plurality of NSR feature points from each of the NSR complexes;generating an NSR template using the plurality of NSR feature... | Cardiac Pacemakers, Inc. | A61B005/0464
US07881764 | B1 | 2011-02-01 | Physiological recording device | 2008-12-22 | The present invention relates to a dry physiological recording electrode that can be used without skin preparation or the use of electrolytic gels. The dry physiological recording electrode comprising a substrate having an upper and a lower surface, and at least one penetrator(s) protruding from the upper surface of the substrate. The penetrator(s) is capable of piercing through the stratum corneum or outer layer of the skin, and transmitting an electric potential from the lower layers of the epidermis through the penetrator(s) which can be measured, or detecting agents from the lower layers of the epidmermis primarily the stratum germinativum layer. At least one epidermis stop may be provided resulting in the formation of detritus troughs interposed between adjacent penetrator(s) and epidermis stops. The present invention also includes a method of sensing biopotentials in the skin. | 1. A method of acquiring physiological signals from a subject having skin comprising forming at least one dry electrode from a negative image having desired surface features, where said surface features include at least one penetrator sized and shaped for piercing the stratum corneum or outer layer of the epidermis and accessing the lower layers of the epidermis,packaging the at least one dry electrode with a portable physiological recording device wherein the at least one dry electrode is capable of transmitting a signal to the recording device,placing the at least one dry electrode in contact with the skin of the subject, andacquiring a physiological signal from the subject. 2. The method of claim 1, wherein the at least one penetrator has an aspect ratio of less than about 1.5. 3. The method of claim 1, wherein the physiological signal acquired is ECG. 4. The method of claim 1, wherein at least two penetrators are formed on the surface of the dry electrode, and the at least two penetrators are individually addressable. 5. The method of claim 1, further including the step of applying a film or coating to the surface of the dry electrode. 6. A method of acquiring physiological signals from a subject having skin comprising forming a physiological electrode with at least one surface feature from an image having desired surface features, where the at least one surface feature is sized and shaped for accessing the lower layers of the epidermis,packaging the physiological electrode with a portable physiological recording device wherein the physiological electrode is capable of transmitting a signal to the portable physiological recording device,placing the physiological electrode in contact with the skin of the subject, and acquiring a physiological signal from the subject. 7. The method of claim 6, wherein the at least one surface feature has an aspect ratio of less than about 1.5. 8. The method of claim 6, wherein the physiological signal acquired is an ECG signal. 9. The method of claim 6, wherein at least two surface features are formed on the surface of the physiological electrode, and the at least two surface features are individually addressable. 10. The method of claim 6, further including the step of applying a film or coating to the surface of the dry electrode. 11. A method of acquiring physiological signals from a subject having skin comprising forming at least one physiological electrode with at least one surface feature from an image having desired surface features, where the at least one surface feature is sized and shaped for accessing the lower layers of the epidermis,packaging the at least one physiological electrode to provide a low impedance path from the electrode's surface to the package's output connector,placing the at least one physiological electrode in contact with the skin of the subject, andacquiring a physiological signal from the subject. 12. The method of claim 11, wherein the physiological signal acquired i... | Orbital Reseach Inc. | A61B005/04 H01R043/00
US07869866 | B2 | 2011-01-11 | Device for the monitoring of physiologic variables through measurement of body electrical impedance | 2005-09-06 | A device for the measurement of body impedance includes a current injecting device for injecting a test current through a first set of electrodes to the body of a patient, and a voltage reading device for reading voltage variations between two electrodes belonging to a second set of electrodes using a coherent demodulator. | 1. A device for the measurement of body impedance comprising: means for injecting a test current through a first set of electrodes to the patient's body, including: at least one oscillator and at least one voltage—current converter and means to read voltage variations between two electrodes belonging to a second set of electrodes including: a demodulator; at least one filter; at least one differential amplifier with two outputs having opposite phase and at least one low pass filter, wherein said at least one oscillator is a sinewave oscillator and wherein said demodulator is a coherent demodulator further including at least one comparator and at least one switch, and wherein the output of said sinewave oscillator is connected to the input of said voltage-current converter, the output of said voltage-current converter is connected to said first set of electrodes, said second set of electrodes is connected to the input of said at least one filter, the output of said at least one filter is connected to the input of said at least one differential amplifier, the first output of said at least one differential amplifier is connected to a first input of said at least one switch, the second output of said at least one differential amplifier is connected to a second input of said at least one switch and to the input of said at least one comparator, the output of said at least one comparator is connected to the control input of said at least one switch, the output of said at least one switch is connected to the input of said at least one low pass filter. 2. A device for the measurement of body impedance comprising: means for injecting a test current through a first set of electrodes to the patient's body, including: at least one oscillator and at least one voltage—current converter and means to read voltage variations between two electrodes belonging to a second set of electrodes including: a demodulator; at least one filter; at least one differential amplifier with two outputs having opposite phase and at least one low pass filter, wherein said at least one oscillator is a sinewave oscillator and wherein said demodulator is a coherent demodulator further including at least one comparator and at least one switch, and wherein the output of said sinewave oscillator is connected to the input of said voltage-current converter and to the input of said at least one comparator, the output of said voltage-current converter is connected to said first set of electrodes, said second set of electrodes is connected to the input of said at least one filter, the output of said at least one filter is connected to the input of said at least one differential amplifier, the first output of said at least one differential amplifier is connected to a first input of said at least one switch, the second output of said at least one differential amplifier is connected to a second input of said at least one switch, the output of said at least one comparator is connected to the contr... | Smartex S.r.l. | A61B005/00
US09138536 | B2 | 2015-09-22 | Apparatus and a method for monitoring a vascular access | 2008-04-01 | An apparatus for monitoring a vascular access includes a dialysis circuit having a venous needle and a blood pump of a peristaltic type which generates oscillatory perturbations in the circuit. An optical sensor, arranged on a patient's body by the venous needle, perceives the perturbations transmitted through the needle and sends a pulsating signal to an analysis device. The absence of a pulsating signal indicates a detachment of the needle. | 1. An apparatus for monitoring a vascular access, comprising: a fluid transport system for moving a fluid;at least one needle connected to the fluid transport system and to a vascular access of an individual, the needle being configured to supply the fluid to the vascular access, to remove the fluid from the vascular access, or to both supply the fluid to the vascular access and to remove the fluid from the vascular access;a perturbation generator configured to generate variations of at least one parameter in the fluid transport system, to generate variations of at least one parameter in the needle, or to generate variations of at least one parameter in both the fluid transport system and in the needle;a sensor configured to be applied directly to the individuals body independent of the needle and the fluid transport system,wherein the sensor comprises at least an emitter of light radiation and a receiver of light radiation configured to emit a signal indicating the variations of at least one parameter generated by the perturbation generator and transmitted through the needle, the emitter of light radiations and the receiver of light radiations configured to be operatively associated to at least a tissue of the individual's body; andan analysis device configured to recognize a detachment of the at least one needle, on a basis of said signal. 2. The apparatus of claim 1, wherein the perturbation generator is operatively directly associated to the fluid transport system. 3. The apparatus of claim 1, wherein the perturbation generator comprises at least an actuator arranged in the fluid transport system. 4. The apparatus of claim 1, wherein the fluid transport system comprises at least an actuator, the perturbation generator comprising the actuator. 5. The apparatus of claim 1, wherein the sensor comprises at least an active part configured to assume at least an operative configuration in which the sensor interacts with the individual's body and is not constrained rigidly to the needle. 6. The apparatus of claim 1, wherein: the at least one needle has a proximal end configured to penetrate a vascular access at a point of penetration and a distal end connected to the fluid transport system;the sensor comprises at least an active part configured to assume at least an operative configuration in which the sensor interacts with the individual's body such that, when the proximal end is in a normal operative configuration in which the proximal end penetrates into the vascular access, said point of penetration is arranged in an intermediate zone between the distal end and the active part. 7. The apparatus of claim 1, wherein the sensor comprises at least an active part configured to assume at least an operative configuration in which the active part interacts contactingy or nearly contactingly with a surface of the individual's body. 8. The apparatus of claim 1, wherein the sensor comprises a physiological sensor for detecting at least a physiologica... | Gambro Lundia AB | A61M001/00 A61M031/00 A61B005/00 A61M005/168 A61M037/00 A61M001/36
US07890348 | B2 | 2011-02-15 | Text-based generic script processing for dynamic configuration of distributed systems | 2002-05-20 | The present technique provides a system and method for dynamic configuration of medical diagnostic systems using distributable multi-component configuration files having extractable component-specific configuration data. The component-specific configuration data is extractable and processable at each component receiving a broadcast of the distributable multi-component configuration file. If a configuration change is desired in the system or in a particular component, then the change is made via the distributable multi-component configuration file. For example, the foregoing distribution and extraction techniques may be performed during operation of the distributed medical diagnostic system in response to a global or application change, such as different medical diagnostic applications. Accordingly, the present technique provides a flexible and architecture-independent mechanism for configuring components of a distributed medical diagnostic system. | 1. A medical diagnostic system, comprising: a plurality of medical diagnostic components communicatively coupled via communications circuitry; anda dynamic configuration system for the plurality of medical diagnostic components, comprising: a configuration data distributor of multi-component configuration data to the plurality of medical diagnostic components, wherein the multi-component configuration data includes a file having application-specific information for at least two medical diagnostic components of the plurality of medical diagnostic components;a plurality of multi-component configuration data receivers, wherein each of the at least two medical diagnostic components of the plurality of medical diagnostic components includes a respective multi-component configuration data receiver of the plurality of multi-component configuration data receivers to allow each of the at least two medical diagnostic components to receive the file having application-specific information for all of the at least two medical diagnostic components;a plurality of component-specific data extractors of the multi-component configuration data, wherein each of the at least two medical diagnostic components of the plurality of medical diagnostic components includes a respective component-specific data extractor of the plurality of component-specific data extractors to enable each of the at least two medical diagnostic components to identify a relevant portion of the multi-component configuration data specific to the respective medical diagnostic component and to extract the relevant portion for use with the respective medical diagnostic component;a plurality of configuration data processors, wherein each of the at least two medical diagnostic components of the plurality of medical diagnostic components includes a respective configuration data processor of the plurality of configuration data processors to facilitate configuration of each of the at least two medical diagnostic components based on the relevant portion of the multi-component configuration data extracted by the component-specific data extractor of the respective medical diagnostic component. 2. The medical diagnostic system of claim 1, wherein the plurality of medical diagnostic components comprise imaging components. 3. The medical diagnostic system of claim 2, wherein the imaging components comprise magnetic resonance imaging components. 4. The medical diagnostic system of claim 2, wherein the imaging components comprise computed tomography components. 5. The medical diagnostic system of claim 2, wherein the imaging components comprise ultrasound components. 6. The medical diagnostic system of claim 2, wherein the imaging components comprise x-ray components. 7. The medical diagnostic system of claim 1, wherein the dynamic configuration system is operable at runtime of the medical diagnostic system. 8. The medical diagnostic system of claim 1, wherein the dynamic configuration system is architect... | GE Medical Systems Global Technology Company, LLC | A61B005/00
US10123713 | B2 | 2018-11-13 | Patient data sensor device | 2009-02-24 | A patient data sensor device has an implantable patient data sensor and an implantable data transmission mechanism, which is in signal connection with the patient data sensor. The data transmission mechanism has an antenna for telemetric data transmission to an external data collection unit. In addition, the data transmission mechanism has a carrier plate for electronic components. The latter are covered on the carrier plate, at least to one side, by a cover. The antenna is configured in an edge region of the carrier plate and running around the latter and around the electronic components. The result is a patient data sensor device which is more tolerable for the patient. | 1. A brain parameter sensor device (1; 16) with an implantable brain parameter sensor (2),with an implantable data transmission mechanism (4), which is in signal connection with the brain parameter sensor (2),wherein the data transmission mechanism (4) has an antenna (11) for telemetric data transmission to an external data collection unit,a carrier plate (5, 7) for electronic components (9, 10),wherein the electronic components (9, 10) on the carrier plate (5, 7) are covered at least to one side by a cover in the form of a membrane,wherein the entirety of the cover (14) is made from a flexible material, and wherein an edge region of the cover (14) is connected to the carrier plate (5, 7),wherein the antenna (11) is arranged in an edge region (8) of the carrier plate (5, 7) and is configured running around the latter and around the electronic components (9, 10), andwherein the antenna is configured as a part of a flexible antenna layer (11), and said flexible antenna layer is connected in a planar manner to the carrier plate (5, 7) and surrounds the carrier plate (5, 7) at the edge. 2. A sensor device according to claim 1, wherein the electronic components (9, 10) on the carrier plate (5, 7) are covered toward both sides by a respective cover (14; 24) formed convexly at least in portions. 3. A sensor device according to claim 2, wherein the cover (24) is produced from one of ceramic and titanium. 4. A sensor device according to claim 1, wherein the antenna (11) is circular. 5. A sensor device according to claim 1 wherein the carrier plate (5, 7) is configured in two layers with two carrier layers, between which the antenna layer (11) is arranged. 6. A sensor device according to claim 1, wherein the antenna (11) is configured as a printed circuit. 7. A sensor device according to claim 1, wherein the carrier plate (5, 7) is covered, where it would be accessible from the outside, by at least one of a coating and an embedding (23) made of a biocompatible material. 8. A sensor device according to claim 7, wherein the coating or the embedding (23) is made by one of the group of silicone, silicone rubber, parylene and polyurethane. 9. A sensor device according to claim 1, wherein the sensor (2) is disposed outside of the cover (14). 10. A sensor device according to claim 1, wherein the cover (14) is connected to a flexible carrier layer (5) of the carrier plate (5, 7). 11. A brain parameter sensor device (1; 16) with an implantable brain parameter sensor (2),with an implantable data transmission mechanism (4), which is in signal connection with the brain parameter (2),wherein the data transmission mechanism (4) has an antenna (11) for telemetric data transmission to an external data collection unit,a carrier plate (5, 7) for electronic components (9, 10),wherein the electronic components (9, 10) on the carrier plate (5, 7) are covered at least to one side by a cover in the form of a membrane,wherein the entirety of the cover (14) is made from ... | RAUMEDIC AG | A61B005/00 A61B005/03 H01Q001/22 H01Q001/27 H05K001/16
US07874983 | B2 | 2011-01-25 | Determination of emotional and physiological states of a recipient of a communication | 2003-01-27 | A communication device capable of determining an emotional or physiological state of a user of the device includes an input device that provides a signal indicative of one of an emotional and a physiological state of a user of the device. A controller receives the signal from the input device and interprets the signal into state data identifying one of an emotional and a physiological state of a user of the device. The device makes the state data available to be downloaded such that a device sending a communication can appropriately adapt its communication method and transport mode of user interface to suit the recipient's condition. | 1. A communication system comprising: a first communication device including: an input device operable to receive at least one signal indicative of an emotional or physiological state of a user of the first communication device,a first processing unit coupled to the input device, the first processing unit operable to receive the at least one signal from the input device, interpret the at least one signal into state data indicating an emotional or a physiological state of the user of the first communication device, and at least temporarily store the state data, anda first transmitter coupled to the first processing unit and operable to transmit the state data; anda second communication device including: a receiver operable to receive the state data from the first communication device;a second processing unit coupled to the receiver and selectively supporting a plurality of communication modes for the second communication device, the second processing unit being operable to automatically select a communication mode of the plurality of communication modes with which to send a communication to the first communication device based on the received state data so as to take into account the emotional or physiological state of the user of the first communication device, the second processing unit being further operable to compare the selected communication mode with a current communication mode of the second communication device and, when the selected communication mode is different than the current communication mode, automatically change the current communication mode of the second communication device to the selected communication mode, anda second transmitter coupled to the second processing unit and operable to transmit the communication to the first communication device in accordance with the selected communication mode. 2. The system of claim 1, wherein the first processing unit is operable to interpolate stored state data to refine a determination of the emotional or a physiological state of the user of the first communication device. 3. The system of claim 1, wherein the second communication device includes a learning mode in which the second communication device is operable to adapt a communication to the first communication device to match previous communications with the first communication device taking into account the emotional or physiological state of the user of the first communication device. 4. The system of claim 1, wherein the second communication device adapts information by changing an interface mode of the second communication device. 5. The system of claim 1, wherein the input device monitors one or more of the group consisting of a biological process of the user of the first communication device, an environment of the user of the first communication device, and a context of the user of the first communication device. 6. The system of claim 1, further comprising: a memory coupled to the first processing unit, the memory havi... | Motorola Mobility, Inc. | A61B005/00
US20110028852 | A1 | 2011-02-03 | Implantable Pressure Sensor with Membrane Bridge | 2009-07-30 | An implantable pressure sensor having improved bend error performance is provided having a capsule housing a pressure sensing device, wherein the capsule includes an opening that allows the pressure sensing device to obtain pressure measurements from an environment surrounding the capsule. A rigid bridge is attached to the capsule so as to extend across the opening of the capsule. The bridge includes at least one opening that exposes the pressure sensing device to the surrounding environment. The rigidity of the bridge functions to prevent deformations from bending forces that are exerted on the sensor and also functions to divert loads from such bending forces away from the pressure sensing device to instead travel through the bridge, which in turn resists deformation. A fill material may be situated between the bridge and capsule so as to prevent bodily fluid or tissue from building up in the sensor under the bridge. | 1. An implantable pressure sensor, comprising: a capsule housing a pressure sensing device, wherein the capsule includes an opening that allows the pressure sensing device to obtain pressure measurements from an environment surrounding the capsule; anda rigid bridge member is attached to the capsule so as to extend across the opening of the capsule. 2. The implantable pressure sensor of claim 1, wherein the bridge member includes at least one opening that allows the pressure sensing device to obtain pressure measurements from an environment surrounding the capsule. 3. The implantable pressure sensor of claim 1, further comprising at least one gap positioned between the bridge member and the capsule that allows the pressure sensing device to obtain pressure measurements from an environment surrounding the capsule. 4. The implantable pressure sensor of claim 1, wherein the bridge member is formed of a material comprising at least one of titanium, stainless steel, MP35N alloy, platinum or another bio-compatible metal. 5. The implantable pressure sensor of claim 2, further comprising a soft fill material positioned at least in the opening in the bridge member. 6. The implantable pressure sensor of claim 5, wherein the soft fill material is further positioned to occupy an area above the pressure sensing device between the bridge member and the capsule. 7. The implantable pressure sensor of claim 5, wherein the soft fill material comprises a silicone material. 8. The implantable pressure sensor of claim 1, wherein the bridge member and opening in the capsule are configured so as to distribute bending forces exerted on the implantable pressure sensor through the bridge member and away from the pressure sensing device. 9. A membrane bridge for reducing bend error in an implantable pressure sensor, comprising: a rigid bridge member attachable to a capsule housing a pressure sensing device including a membrane that is responsive to pressure measurements from an environment surrounding the capsule, wherein the capsule includes an opening adjacent to the membrane that allows the pressure sensing device to obtain pressure measurements from the environment surrounding the capsule,wherein the rigid bridge member is formed to extend across the opening of the capsule over the membrane. 10. The membrane bridge of claim 9, wherein the bridge member includes at least one opening that exposes the membrane of the pressure sensing device to the environment surrounding the capsule for obtaining pressure measurements. 11. The membrane bridge of claim 10, wherein the bridge member is shaped so as to further form at least one gap positioned between the bridge member and the capsule that further exposes the membrane of the pressure sensing device to the environment surrounding the capsule for obtaining pressure measurements. 12. The membrane bridge of claim 9, wherein the bridge member is formed of a material comprising at least one of titanium, stainless steel, ... | - | A61B005/02 B23K026/00
US20110046466 | A1 | 2011-02-24 | Analyte Sensors Including Nanomaterials and Methods of Using Same | 2009-08-19 | Generally, embodiments of the invention relate to analyte determining methods and devices (e.g., electrochemical analyte monitoring systems) that have improved uniformity of distribution of one or more components, improved stability, and improved response time of the sensor by inclusion of a nanomaterial, such as inert inorganic nanomaterials, where the components are disposed proximate to a working electrode of in vivo and/or in vitro analyte sensors, e.g., continuous and/or automatic in vivo monitoring using analyte sensors and/or test strips. Also provided are systems and methods of using the, for example electrochemical, analyte sensors in analyte monitoring. | 1. An analyte sensor comprising: a working electrode comprising a sensing layer disposed proximate thereto;a counter and/or reference electrode; anda nanomaterial disposed proximate to the sensing layer. 2. The analyte sensor of claim 1, wherein at least a portion of the electrochemical analyte sensor is adapted to be subcutaneously positioned in a subject. 3. The analyte sensor of claim 1, further comprising a membrane disposed over the sensing layer. 4. The analyte sensor of claim 1, wherein the nanomaterial comprises a nanopowder, a nanoparticle, a nanotube, a nanofilament, or a fullerene. 5. The analyte sensor of claim 1, wherein the nanomaterial is an inorganic nanomaterial. 6. The analyte sensor of claim 5, wherein the inorganic nanomaterial comprises a metal nanomaterial, metal-alloy nanomaterial, ceramic nanomaterial, or dielectric nanomaterial. 7. The analyte sensor of claim 1, wherein the nanomaterial is configured to provide electrical conductivity. 8. The analyte sensor of claim 1, wherein the nanomaterial is a substantially inert nanomaterial. 9. The analyte sensor of claim 1, wherein the sensing layer comprises the nanomaterial. 10. The analyte sensor of claim 1, wherein the nanomaterial comprises carbon nanopowder. 11. The analyte sensor of claim 1, wherein the nanomaterial comprises aluminum, carbon, cobalt, copper, copper-zinc alloy, diamond, gold, iron, iron-nickel alloy, molybdenum, magnesium, nickel, palladium, platinum, silver, silver-copper alloy, tantalum, tin, indium doped tin oxide, titanium, titanium nitride, tungsten, zinc, calcium oxide, hydroxyapatite, indium, silica, silicon, silicon dioxide, silicon nitride, silicon carbide, cellulose, or clay. 12. The analyte sensor of claim 1, wherein the nanomaterial comprises a polymer, wherein the polymer is not covalently conjugated to components of the analyte sensor. 13. The analyte sensor of claim 12, wherein the polymer is polyethylene, polymethylene, polypropylene, or polystyrene. 14. The analyte sensor of claim 9, wherein the nanomaterial comprises from about 0.5% to about 60% by weight of the total formulation of the sensing layer. 15. The analyte sensor of claim 9, wherein the nanomaterial comprises from about 1% to about 10% by weight of the total formulation of the sensing layer. 16. The analyte sensor of claim 1, wherein the nanomaterial comprises a nanopowder. 17. The analyte sensor of claim 16, wherein the nanopowder comprises particles having a diameter from about 1 nm to about 300 nm. 18. The analyte sensor of claim 1, wherein the sensing layer comprises a glucose-responsive enzyme. 19. The analyte sensor of claim 1, wherein the sensing layer comprises a redox mediator. 20. The analyte sensor of claim 1, wherein the sensing layer comprises a crosslinked analyte responsive enzyme and redox polymer. 21. The analyte sensor of claim 19, wherein the redox mediator comprises a ruthenium-containing complex or an osmium-containing complex. 22. The a... | - | A61B005/1468 G01N027/26
US20110040176 | A1 | 2011-02-17 | METHOD AND DEVICE FOR NEAR-FIELD DUAL-WAVE MODALITY IMAGING | 2009-02-18 | A method of near-field imaging of a region of interest includes emitting an input of near-field electromagnetic radiofrequency energy from a near-field source device into the region of interest, detecting a mechanical wave response generated in the region of interest in response to the near-field electromagnetic energy input using a detector device, and providing image data representing an image of the region of interest based on the mechanical wave response. | 1-40. (canceled) 41. A method of near-field imaging of a region of interest, comprising: emitting an input of near-field electromagnetic radiofrequency energy from a near-field source device into the region of interest;detecting a mechanical wave response generated in the region of interest in response to the near-field electromagnetic energy input using a detector device; andproviding image data representing an image of the region of interest based on the mechanical wave response. 42. The method of claim 41, wherein the region of interest is placed within the near-field region of the near-field source device at a distance smaller than λ from the near-field source device, wherein λ is a wavelength of the electromagnetic energy absorbed in the near-field. 43. The method of claim 41, wherein the near-field electromagnetic energy is emitted as a pulse modulated signal. 44. The method of claim 43, further comprising performing at least one of amplitude, frequency, phase or digital modulation of the pulse modulated signal. 45. The method of claim 41, wherein the near-field energy is emitted as a modulated intensity signal. 46. The method of claim 45, further comprising performing at least one of amplitude, frequency, phase or digital modulation of the intensity modulated signal. 47. The method of claim 41, further comprising preventing emission of far-field radiation into a surrounding region. 48. The method of claim 41, wherein the mechanical wave response comprises at least one of an acoustic response, a pressure response, a shear wave response, a shock wave response, an interference response or thermal waves response. 49. The method of claim 41, wherein the mechanical wave response is detected by the detector device including an acoustic detector or an optical or interferometric detector. 50. The method of claim 49, wherein the detector device comprises at least one detector element or a plurality of detector elements arranged in an array format. 51. The method of claim 41, further comprising: reconstructing the image of the region of interest from the image data including utilizing information on the structure of the near-field, whereinthe image describes at least one of one-dimensional properties including depth profiling properties, two-dimensional properties including surface or cross-sectional properties, and three-dimensional properties including volumetric properties of the region of interest. 52. The method of claim 51, further comprising displaying the image of the region of interest. 53. The method of claim 41, further comprising adjusting a relative position or orientation of at least one of the object, the near-field source device and the detector device relative to each other before or during data acquisition. 54. The method of claim 53, wherein the adjusting comprises at least one of a rotation and a translation of at least one of the object, the near-field source device and the detector device. 55. The method of claim ... | Helmholtz Zentrum Muenchen Deutsches Forschungszentrum fur Gesundheit und Umwelt (GmbH) | A61B005/05 G06K009/00 G01S013/00
US20110004114 | A1 | 2011-01-06 | Cognitive information encoding conversion telecommunication system and conferencing method | 2009-07-06 | This invention relates to a scanless system for- and method of teleportation of cognitive perceptions as events in series in conjunction with- or without radio-telephone frequency applications. In the invention, an event includes the focal content of bound percepts in a source-subject. The invention describes the process whereby the quantum nature of the information-bearing content of those percepts, whether founded on physical or mental (imaginary) stimuli can be transceived through the invention as concept modules whose origins lie within the data-free elements of sensory organ electromagnetic radiation frequency wave responses to the source-subject's individually-directed and subjectively-generated emotions. At the core, the assembly identifies a method for capturing the engagement between photon-carrying information and transmissions to the receiving-subject or -devices using spintronic-based evaluation for a plurality of inputs and a plurality of outputs with a radio-frequency telephone-enabled assembly, governed by quantum key distribution protocols and time-dependent delay processors. | 1. a method of teleportation of the quantum nature of cognitive information content encoded on the electromagnetic radiation waves present in-, on- or near—as well as in direct association to primary sensory perception in a subject-source to a receiver-source and/or peripheral computation device and/or other receiving device. 2. a method of scanless information conferencing of a plurality of cognitive information content encoded on the electromagnetic radiation waves present in-, on- or near—as well as in direct association to primary sensory perception. 3. a method of decoding a plurality of output relationships between subject-source emotion and perception-activation values 4. a method for incorporating radio-frequency with the assembly-method for real-time communication of visual-/acoustic-sensory cognitive perceptions. 5. The method in claim #1, #2, #3, and #4a) a scanless mirroring component capable of detecting and transmitting concentrated localizations of information encoded electromagnetic activity within and running into primary sensory organs for perception-based cognitive functioning with radio-telephone frequency applicationsb) a scanless mirroring component capable of detecting and transmitting concentrated localizations of information encoded electromagnetic activity within and running into primary sensory organs for perception-based cognitive functioning without radio-telephone frequency applicationsc) a method of delineating the tangible relationship between neuron firing and projective geometries ideally suited for cognitive relationships according to complex Hilbert, Minkowski or other quantum mechanical space-time models in order to transfer arrays of perceptual-information encodings from the brain processes 6. As in claim 5, an assembly comprising Detector Conference System and/or Converter Assembly for Source-subject source 7. As in claim 5, an assembly comprising Convertor Assembly and Detector Conference System for Receiver-subject source 8. As in claim 1 through 6, an assembly comprising both Detector Conference System and Converter Assembly for both Source-subject and Receiver subject source 9. As in claim 1 through 8, any arrangement of Detector Conference system 10. As in claim 1 through 8, any arrangement of Converter Assembly | Farida Hanna Schenck (Campbell), | A61B005/0476
US08874191 | B2 | 2014-10-28 | Method and magnetic resonance device for graphically assisting navigation of an instrument | 2010-07-09 | In a method and magnetic resonance device for graphically assisting the navigation of an instrument (in particular a medical instrument such as a needle and/or a guide sleeve guiding a needle) within a procedure region of a target subject (in particular a patient), coordinates of a target position and/or target region to be addressed that are obtained from a high-contrast image data set acquired with the magnetic resonance device. Low-contrast image data showing the instrument are acquired continuously and/or intermittently during the navigation and/or the position and orientation of the instrument is determined (in particular measured) with a position determination device. The low-contrast image later are brought into registration with the magnetic resonance device continuously and/or intermittently during the navigation. A representation showing at least the relative position and orientation of the instrument relative to the target position and/or target region is generated and displayed under consideration of the coordinates and the low-contrast image data and/or the position and orientation of the instrument. | 1. A method for graphically assisting navigation of a medical instrument in a magnetic resonance device within a medical procedure region in a patient in the magnetic resonance device, comprising: operating the magnetic resonance device with the patient therein to obtain a high-contrast image data set representing said procedure region, said high-contrast image data set being acquired in a magnetic resonance device coordinate system associated with the magnetic resonance device;supplying said high-contrast image data set to a computerized processor and, in said computerized processor, automatically determining, from said high-contrast image data set, coordinates of a target, in said magnetic resonance device coordinate system, to be reached by the instrument in said procedure region, said coordinates of said target being selected from the group consisting of a target position and a target region;obtaining instrument-representing data representing said instrument during movement of said instrument in said procedure region toward said target by operating said magnetic resonance device to intermittently or continuously acquire low-contrast image data representing said instrument during movement of said instrument in said procedure region toward said target;supplying said low contrast data representing said instrument to said computerized processor and, in said processor, automatically generating a low-contrast image therefrom in which said instrument, and said target at said coordinates of said target, are visualized; andat a display device in communication with said computerized processor, visually displaying said low-contrast image to show said instrument relative to said coordinates of said target. 2. A method as claimed in claim 1 comprising representing said target at said display with a symbolic marking indicating an orientation of said target. 3. A method as claimed in claim 2 comprising presenting a representation at said display comprising a slice image of said low-contrast data, and entering said symbolic marking into said slice image when reconstructing said slice image. 4. A method as claimed in claim 2 comprising employing, as said symbolic marking, an indicator representing information identifying a distance of the target position perpendicular to the slice plane. 5. A method as claimed in claim 4 wherein said indicator has a shape, a color, and a size, and comprising configuring at least one of said shape and said color of said indicator to indicate a position of the target above or below the plane of the slice image, and indicating the distance of the target from the plane of the slice by said size of said indicator. 6. A method as claimed in claim 1 comprising forming said low-contrast image data from two slice images. 7. A method as claimed in claim 6 wherein said two slice images are substantially perpendicular to each other. 8. A method as claimed in claim 1 comprising incorporating at least one additional subject or area ... | Siemens Aktiengesellschaft | A61B005/05 A61B005/06 A61B005/055
US08391966 | B2 | 2013-03-05 | Sensory-evoked potential (SEP) classification/detection in the time domain | 2010-09-20 | Techniques are disclosed for sensory-evoked potential (SEPs, e.g., visual-evoked potentials) signal detection/classification by synchronizing EEG to the repeated presentation of sensory stimuli in the time domain. In some embodiments, a system receives a plurality of EEG signal samples, generates a stimulus-locked EEG and determines whether the plurality of EEG signal samples are evoked in response to a pattern of stimulus. In some embodiments, no prior knowledge about the update pattern (such as the flashing frequency of a visual stimulus) of the stimulus and no prior knowledge about an individual user's EEG pattern are required. | 1. A method for sensory-evoked potential signal classification, comprising: generating a pattern of stimulus characterized by a frequency that updates randomly;determining in real-time the frequency of the pattern of stimulus, the determined frequency of the pattern of stimulus being the current frequency of the pattern of stimulus;receiving a plurality of electroencephalography signal samples from a current user evoked in response to the pattern of stimulus;synchronizing the plurality of electroencephalography signal samples in time domain without conversion to a frequency domain to the pattern of stimulus in time domain;generating a stimulus-locked electroencephalography signal sample using a processor; andperforming a determination using a time domain classifier in real-time of whether the plurality of electroencephalography signal samples are evoked in response to the pattern of stimulus using the processor,wherein the determination of whether the plurality of electroencephalography signal samples are evoked in response to the pattern of stimulus being performed dynamically or adaptively in real-time in a manner such that the current user's electroencephalography pattern determined in real-time is used, the current user's electroencephalography response pattern determined before the current user's current electroencephalography response pattern is not used, the frequency of the pattern of stimulus determined in real-time is used, and a frequency of the pattern of stimulus determined before the current determined frequency of the pattern of stimulus is not used in the determination of whether the plurality of electroencephalography signal samples are evoked in response to the pattern of stimulus; andwherein the sensory-evoked potential signal classification is computed by calculating a mutual relationship between pairs of segments for the stimulus-locked electroencephalography signal sample, determining a statistical measure of the mutual relationship for each stimulus, and determining if a stimulus is being attended to by the current user according to the statistical measure of the mutual relationship for each stimulus user the time domain classifier. 2. The method recited in claim 1, wherein sensory-evoked potential signal classification is performed by synchronizing electroencephalography signal samples to repeated presentation of sensory stimuli in the time domain. 3. The method recited in claim 1, wherein sensory-evoked potential signal classification is performed by synchronizing electroencephalography signal samples to repeated presentation of sensory stimuli in the time domain, wherein the sensory-evoked potential signal classification includes visually evoked potentials. 4. The method recited in claim 1, further comprising: detecting a plurality of electroencephalography signals. 5. The method recited in claim 1, further comprising: recording the plurality of electroencephalography signal samples. 6. The method recited in claim 1... | NeuroSky, Inc. | A61B005/04 A61B005/05
US08306611 | B2 | 2012-11-06 | Method and system for use in monitoring left ventricular dysfunction | 2010-08-12 | A method and system are presented for use in determining a patient's heart condition. First and second data are provided, where the first data is indicative of the patient's heart rate, and the second data is indicative of bioimpedance peak value during a cardiac cycle. A relation between a certain predetermined value and a product of the first and said second data is determined, where such relation is indicative of the patient's left ventricular condition, enabling to identify a left ventricular dysfunction. | 1. A method for use in determining a patient's heart condition, the method comprising: providing first data indicative of the patient's heart rate and second data indicative of bioimpedance peak value during a cardiac cycle, wherein said first and second data are characterized by at least one of the following:(a) said first data comprises a product of the patient's heart rate value and a coefficient kHR which is specific for a patient and which is a correction coefficient for correction of the heart rate; and(b) said second data comprises a product of a normalized systolic impedance variation, ΔR/R, and a systolic peak time, α;determining a relation between a product of said first and said second data and a certain predetermined value, said relation being indicative of the patient's left ventricular condition enabling to identify left ventricular dysfunction. 2. The method according to claim 1, wherein said relation indicative of the patient's left ventricular dysfunction is determined such that said product of said first and said second data being less than said certain predetermined value corresponds to a condition of the patient's left ventricular dysfunction. 3. The method according to claim 1, wherein said first data indicative of the patient's heart rate comprises a patient's heart rate value. 4. The method according to claim 1, wherein the correction coefficient kHR is equal to 1 when the measured patient's heart rate HRmeas within a certain range of normal values between a bottom limit BL and a top limit TL, is equal to BL/HRmeas when the measured patient's heart rate HRmeas is less than the bottom limit BL of the normal range, and is equal to TL/HRmeas when the heart rate HRmeas higher than the top limit TL of the normal range. 5. The method according to claim 4, wherein the correction coefficient kHR is equal to 1 when the measured patient's heart rate HRmeas is within a range of 60-90, and is equal to 60/HRmeas when the measured heart rate HRmeas is less than 60, and is equal to 90/HRmeas when the measured heart rate HRmeas higher than 90. 6. The method according to claim 5, wherein said second data indicative of electrical bioimpedance changes during a cardiac cycle comprises a product of a normalized systolic peak value impedance, ΔR/R, and a systolic peak time, α. 7. The method according to claim 6, wherein said certain predetermined value is equal to 10. 8. The method according to claim 4, wherein said second data indicative of electrical bioimpedance changes during a cardiac cycle comprises a product of a normalized systolic impedance variation, ΔR/R, and a systolic peak time, α. 9. A system for use in determining a patient's heart condition, the system comprising: a data input utility for receiving first data indicative of the patient's heart rate and receiving second data indicative of bioimpedance variations during a cardiac cycle, wherein said first and second data are characterized by at least one of the following:(a)... | N.I Medical Ltd. | A61B005/04
US07890155 | B2 | 2011-02-15 | Feature emphasis and contextual cutaways for image visualization | 2007-09-28 | A method for visualizing one or more objects of interest within a volume of imaging data includes the steps of receiving a volume of imaging data and displaying the one or more objects of interest within the volume by adaptively changing the amount of information displayed from the volume based on (1) a viewing direction of the volume, (2) a position of the one or more objects of interest within the volume, and (3) importance values of materials within the volume. | 1. A method for visualizing one or more objects of interest within a volume of imaging data, the method steps comprising: receiving, by a processor, the volume of imaging data; anddisplaying the one or more objects of interest within the volume by adaptively changing by the processor an amount of information displayed from the volume based on a viewing direction of the volume, a position of the one or more objects of interest within the volume, and importance values of materials within the volume, wherein the displaying comprises: selecting a viewing angle that is relative to the viewing direction;determining regions of the volume that fall within the viewing angle by determining a first region having a first angle away from the viewing direction, the first region having a bottom that is a first distance away from the object of interest and determining a second region having a second angle away from the viewing direction, the second region having a bottom that is a second distance away from the object of interest, wherein the first angle is less than the second angle and the first distance is less than the second distance;determining occlusion values for each point of the regions based on the degree that each point occludes the object of interest; andmodifying an opacity value for each point based on the occlusion value of the point and the importance material that the point corresponds to. 2. The method of claim 1, further comprising using the display of the one or more objects of interest within the volume to aid a surgical procedure. 3. The method of claim 2, wherein a first one of the objects of interest is a needle and the surgical procedure comprises guiding the needle into a second one of the objects of interest. 4. The method of claim 1, wherein the receiving the volume of imaging data comprises: receiving two dimensional (2D) imaging data; andregistering the volume of imaging data with the 2D imaging data to generate the one or more objects of interest within the volume of imaging data. 5. The method of claim 1, wherein the volume of imaging data comprises data from a three dimensional (3D) computed tomography (CT) scan. 6. The method of claim 4, wherein the 2D imaging data comprises data from a tracked ultrasound image. 7. The method of claim 1, further comprising shading the materials within the volume based on the corresponding importance values. 8. The method of claim 1, wherein the materials represent different types of anatomical tissue structures within the volume. 9. A method for visualizing an object of interest, the method steps comprising: selecting, by a processor, a viewing direction with respect to an object of interest in a volume of image data, and first and second angles with respect to the viewing direction, the second angle being less than the first angle;determining, by the processor, a first region within the volume defined by the first angle, a second region within the volume defined by the second angle, and... | Siemens Medical Solutions USA, Inc. | A61B005/05
US07883470 | B2 | 2011-02-08 | Apparatus and method for outputting heart sounds | 2005-01-18 | An apparatus for outputting heart sounds includes an implantable system and an external system. The implantable system includes a sensor for generating sensed signals representing detected heart sounds, an interface circuit and a control circuit for receiving the sensed signals, generating data representing the heart sounds therefrom, and transmitting the data to the external system via the interface circuit. The external system includes an interface circuit for communicating with the implantable system, and a control circuit for receiving the data representing the heart sounds and for generating control signals that cause an output device to generate outputs representing the sounds. The implantable system may also include a sensor(s) for detecting cardiac electrical signals. In this case, outputs representing the cardiac electrical signals are also output. | 1. An external system for communicating with an implantable system sensing an accelerometer signal indicative of heart sounds and cardiac electrical signals indicative electrical cardiac events, the external system comprising: an external device communicatively coupled to the implantable system, the external device including: an interface circuit configured to communicate with the implantable system;a control circuit, coupled to the interface circuit, configured to receive data representing the heart sounds and the cardiac electrical signals telemetered from the implantable system, detect the heart sounds and the electrical cardiac events from the received data, and calculate timing differences each between a heart sound of the detected heart sounds and an electrical cardiac event of the detected electrical cardiac events; andan output device coupled to the control circuit, the output device including a screen display including: a plurality of horizontal traces including simultaneously displayed one or more accelerometer traces related to the heart sounds and one or more cardiac electrical signal traces each showing one of the cardiac electrical signals;visual indicia indicative of the heart sounds, the visual indicia displayed in association with locations of the heart sounds on at least one of the displayed one or more accelerometer traces; andthe calculated timing differences. 2. The external system of claim 1, wherein the screen display further includes a pair of horizontally movable vertical lines for comparing timing between any traces of the plurality of horizontal traces. 3. The external system of claim 1, further comprising one or more input devices coupled to the interface circuit, the one or more input devices adapted to allow for vertically moving a trace of the plurality of horizontal traces to superimpose that trace over another trace of the plurality of horizontal traces. 4. The external system of claim 1, further comprising a further output device and a transmission media coupled between the interface circuit and the further output device, the further output device adapted to display the data representing the heart sounds and the cardiac electrical signals. 5. The external system of claim 1, wherein the one or more accelerometer traces comprise a processed accelerometer trace showing the accelerometer signal after being processed by the external system. 6. The external system of claim 5, wherein the one or more cardiac electrical signal traces comprise one or more electrogram (EGM) traces. 7. The external system of claim 6, wherein the visual indicia are indicative of the heart sounds displayed in association with locations of the heart sounds on the processed accelerometer trace. 8. The external system of claim 7, wherein the screen display further includes visual indicia indicative of electrical cardiac events displayed in association with locations of electrical cardiac events on the displayed one or more EGM traces. 9.... | Cardiac Pacemakers, Inc. | A61B005/02
US07873409 | B2 | 2011-01-18 | Invasive cardiology digital signal amplifier and acquisition device | 2006-08-25 | The digital signal amplifier and acquisition system includes at least one catheter input module (CIM) configured to receive, route and digitize incoming analog cardiac signals from a number of catheters, as well as outgoing stimulator pulses. A plurality of CIMs are mechanically stacked and electrically daisy-chained and coupled with a mounting. The mounting platter is clamped to the bedrail and provides a single digital output cable to the base. The system also includes an acquisition device and the aforementioned base, which is configured to collect, filter and distribute the acquired data. Lastly an analog output module receives filtered digital signals from the base and is configured to reconstruct analog representations of such filtered digital signals for outside devices. | 1. A data acquisition system in communication to receive a set of analog cardiac data from a catheter of a patient situated on a support structure, the system comprising: at least one catheter input module configured to receive the set of analog data from a patient and to convert the set of analog cardiac data to a set of digital data, wherein the catheter input module includes a plurality of connectors configured to receive analog data, and further wherein the catheter input module is configured to be physically and electrically coupled with additional catheter input modules;a mount platter removably coupled to a patient bed in close proximity to the patient, and further coupled in communication to receive the set of digital data from the at least one catheter input module, wherein the mount platter is removably coupled to the at least one catheter input module and supporting the at least one catheter input module; anda base unit coupled in communication to receive the set of digital data from the mount platter and further configured to filter and distribute the set of digital data, wherein the base unit is physically coupled to the patient bed, wherein the mount platter includes an emergency stimulator connector coupled with a STIM input and configured to provide a pulse to the catheter from the STIM input, wherein the pulse from the emergency stimulator connector is provided in the event of total system failure. 2. The system as claimed in claim 1, further comprising an analog output module electrically coupled with the base unit and configured to receive a filtered set of digital data from the base unit, and further configured to convert the filtered set of digital data to a reconstructed analog signal. 3. The system as claimed in claim 1, further comprising an acquisition device coupled with the base unit, the acquisition device configured to control the collection of a plurality of physiological parameters from a patient. 4. The system as claimed in claim 1, wherein the mount platter includes an auxiliary reference input configured to collect a non-catheter patient input. 5. The system as claimed in claim 1, wherein the at least one catheter input module includes a plurality of inputs. 6. The system as claimed in claim 5, wherein the mount platter includes a test signal generator configured to test the functionality of any of the plurality of inputs of the catheter input module. 7. The system as claimed in claim 5, further comprising a plurality of catheter input modules, wherein each of the plurality of catheter input modules is electrically and mechanically coupled to each other and with the mount platter. 8. A data acquisition system in communication to receive a set of analog cardiac data from a catheter of a patient situated on a support structure, the system comprising: at least one catheter input module configured to receive the set of analog data from a patient and to convert the set of analog cardiac data to a set of digital... | General Electric Company | A61B005/02
US07881792 | B1 | 2011-02-01 | Methods and systems for detecting the presence of T-wave alternans | 2005-09-16 | Embodiments of the present invention relate to implantable systems, and method for use therein, that can detect T-wave alternans. In accordance with specific embodiments of the present invention, intrinsic premature contractions of the ventricles are detected, and at least one metric of T-waves is measured in a specified number of beats that follow each detected intrinsic premature contraction of the ventricles. A determination of whether T-wave alternans are present is made based on the measured T-wave metrics. In alternative embodiments, rather than waiting for intrinsic premature contractions of the ventricles, premature contractions of the ventricles are caused on demand by inducing premature atrial contractions. In still other embodiments, a patient's vagus nerve is stimulated to simulate premature contractions of the ventricles. This abstract is not intended to be a complete description of, or limit the scope of, the invention. | 1. In an implantable system, a method for detecting T-wave alternans, comprising: (a) stimulating a patient's vagus nerve to thereby simulate at least one premature contraction of the ventricles;(b) measuring at least one metric of T-waves in a specified number of beats that follow each simulated premature contraction of the ventricles that is caused by stimulating the patients vagus nerve; and(c) determining, based on the measured T-wave metrics, whether T-wave alternans are present. 2. The method of claim 1, wherein step (a) includes stimulating a patient's vagus nerve to thereby simulate a plurality of premature contractions of the ventricles. 3. The method of claim 2, wherein steps (a) and (b) are repeated a plurality of times and wherein the patient's heart is monitored for premature ventricular contractions to ensure that there are at least 3 beats in which a premature contraction of the ventricles does not occur between each stimulating of the vagus nerve at step (a). 4. The method of claim 1, wherein step (b) comprises measuring at least one metric of T-waves in M beats that follow each simulated premature contraction of the ventricles, where M is an integer between 2 and 10, inclusive. 5. The method of claim 1, wherein step (b) includes measuring at least one of the following T-wave metrics: T-wave amplitude;T-wave width;T-wave slope;T-wave area;T-wave morphology;QT interval; andevoked QT interval. 6. An implantable system for detecting T-wave alternans, comprising: means for stimulating a patient's vagus nerve to thereby simulate premature contractions of the ventricles;means for measuring at least one metric of T-waves in a specified number of beats that follow the simulated premature contractions of the ventricles that are caused by stimulating the patients vagus nerve; anda TWA detector to determine, based on the measured T-wave metrics, whether T-wave alternans are present. 7. The system of claim 6, further including means for monitoring for premature ventricular contractions and means for ensuring that there are at least 3 beats in which a premature contraction of the ventricles does not occur between each stimulating of the vagus nerve. 8. The system of claim 6, wherein at least one metric of T-waves is measured in M beats that follow each simulated premature contraction of the ventricles, where M is an integer between 2 and 10, inclusive. 9. The system of claim 6, wherein said means for measuring measures at least one of the following T-wave metrics: T-wave amplitude;T-wave width;T-wave slope;T-wave area;T-wave morphology;QT interval; andevoked QT interval. | Pacesetter, Inc. | A61N001/362 A61B005/0452
US07881781 | B2 | 2011-02-01 | Thoracic impedance detection with blood resistivity compensation | 2009-12-01 | This document discusses, among other things, a cardiac rhythm management device or other implantable medical device that uses thoracic impedance to determine how much fluid is present in the thorax, such as for detecting or predicting congestive heart failure, pulmonary edema, pleural effusion, hypotension, or the like. The thoracic fluid amount determined from the thoracic impedance is compensated for changes in blood resistivity, which may result from changes in hematocrit level or other factors. The blood-resistivity-compensated thoracic fluid amount can be stored in the device or transmitted to an external device for storage or display. The blood-resistivity-compensated thoracic fluid amount can also be used to adjust a cardiac pacing, cardiac resynchronization, or other cardiac rhythm management or other therapy to the patient. This document also discusses applications of the devices and methods for predicting or indicating anemia. | 1. A method comprising: obtaining a first blood resistivity quantity during a first time using at least one implanted intravascular or intracardiac electrode in a subject;obtaining a second blood resistivity quantity during a second time using the at least one implanted intravascular or intracardiac electrode;comparing, using a controller, the second blood resistivity quantity to the first blood resistivity quantity; anddeclaring an indication or prediction of anemia to be present in the patient when the second blood resistivity quantity is less than the first blood resistivity quantity by at least a first amount. 2. The method of claim 1, wherein obtaining the first blood resistivity quantity and obtaining the second blood resistivity quantity are performed in the same blood vessel or heart chamber. 3. The method of claim 1, wherein obtaining the first blood resistivity quantity and obtaining the second blood resistivity quantity are performed under a like condition with respect to at least one of a cardiac cycle, a respiratory cycle, a posture, and a circadian cycle. 4. The method of claim 1, comprising communicating an indication of anemia predicted or anemia present to an external device. 5. The method of claim 1, wherein declaring an indication or prediction of anemia includes: declaring an indication of anemia to be present in the patient when the second blood resistivity quantity is less than the first blood resistivity quantity by at least a second amount; anddeclaring a prediction of anemia to be present in the patient when the second blood resistivity quantity is less than the first blood resistivity quantity by at least the first amount and by less than the second amount. 6. The method of claim 1, comprising communicating information regarding the indication or prediction of anemia. 7. The method of claim 1, comprising displaying information regarding the indication or prediction of anemia. 8. The method of claim 7, wherein displaying the information includes displaying the information using an external device. 9. The method of claim 1, comprising providing an anemia warning with the indication or prediction of anemia. 10. The method of claim 1, wherein at least one of: obtaining the first blood resistivity quantity includes detecting a first blood resistivity measurement during the first time; orobtaining the second blood resistivity quantity includes detecting a second blood resistivity measurement during the second time. 11. The method of claim 1, wherein at least one of: obtaining the first blood resistivity quantity includes combining two or more first blood resistivity measurements acquired during the first time; orobtaining the second blood resistivity quantity includes combining two or more second blood resistivity measurements acquired during the second time. 12. The method of claim 1, wherein at least one of: obtaining the first blood resistivity quantity includes averaging two or more first blood resistivity measu... | Cardiac Pacemakers, Inc. | A61B005/00
US07874992 | B2 | 2011-01-25 | Method for continuous baroreflex sensitivity measurement | 2006-01-31 | A medical device and method for determining baroreflex sensitivity (BRS) based on one or more respiration cycles. The BRS determination may be performed continuously based on measurements of heart rate, blood pressure, and respiration cycles. | 1. A method for calculating a continuous baroreflex sensitivity (BRS) value, comprising: detecting respiration cycles of a patient;measuring heart rate of the patient over a first set of respiration cycles;measuring a pressure metric of the patient over a second set of respiration cycles; andcalculating a BRS value based on the heart rate and pressure measurements, wherein the calculating the BRS value is based on heart rate and pressure measurements taken over different respiration cycles. 2. A method according to claim 1, wherein the calculating the BRS value is performed for each respiration cycle. 3. A method according to claim 1, wherein the respiration cycles are detected from the pressure measurement. 4. A method according to claim 1, wherein the measuring the pressure metric comprises measuring a right ventricular systolic pressure metric. 5. A method according to claim 1, wherein the calculating the BRS value includes determining a maximum change in the pressure metric over each respiration cycle. 6. A method according to claim 1, wherein calculating the BRS value includes determining a maximum change in an R-R interval over each respiration cycle. 7. A method according to claim 1, wherein the heart rate is measured from an ECG signal. 8. An implantable medical device, comprising: one or more sensors for sensing patient heart rate, blood pressure, and respiration cycles; anda housing supporting a processor and adapted for implantation in a patient, the processor coupled to the one or more sensors, the processor for calculating continuous baroreflex sensitivity (BRS) for each detected respiration cycle based on the sensed heart rate, blood pressure and respiration cycles, wherein:the processor compares the BRS calculation against a zone of predetermined BRS thresholds, andthe processor changes at least one zone threshold if the BRS calculation intermittently produces values outside the zone. 9. A medical device according to claim 8, wherein the respiration cycles and blood pressure are sensed by the same one of the one or more sensors. 10. A medical device according to claim 8, wherein the processor calculates a trendline of BRS values obtained over multiple respiration cycles. 11. A medical device according to claim 8, wherein the processor initiates an alarm if the BRS calculation is outside the zone of predetermined BRS thresholds. 12. A medical device according to claim 8, wherein the processor initiates a treatment if the BRS calculation is outside of the zone of predetermined thresholds. 13. A computer-readable medium programmed with instructions for baroreflex sensitivity (BRS) measurement, the medium comprising instructions for causing a programmable processor to: detect respiration cycles of a patient;measure heart rate of the patient over a first set of respiration cycles;measure a pressure metric of the patient over a second set of respiration cycles; andcalculate a BRS value based on the heart rate and pressure mea... | Medtronic, Inc. | A61B005/02
US08483793 | B2 | 2013-07-09 | Dual electrode system for a continuous analyte sensor | 2010-10-29 | Disclosed herein are systems and methods for a continuous analyte sensor, such as a continuous glucose sensor. One such system utilizes first and second working electrodes to measure analyte or non-analyte related signal, both of which electrode include an interference domain. | 1. A method for providing a substantially noise-free glucose signal for a glucose sensor implanted in a host, the method comprising: providing a sensor, the sensor comprising a first working electrode disposed beneath an active enzymatic portion of a membrane system, and a second working electrode disposed beneath an inactive-enzymatic or a non-enzymatic portion of the membrane system, wherein the membrane system is configured to substantially reduce one or more interfering species from reaching the first working electrode and the second working electrode;generating a first signal associated with the first working electrode;generating a second signal associated with the second working electrode;processing the first signal and the second signal to produce a glucose signal that is substantially without signal contribution from interfering species; andperforming, using electronic circuitrty, a self-diagnosis by measuring an auxiliary signal to determine a change in sensitivity of the sensor using one or more electrodes of the sensor, independent of a blood glucose reference value. 2. The method of claim 1, wherein determining a change in sensitivity comprises measuring a non-glucose analyte concentration. 3. The method of claim 1, further comprising requesting an external reference value in response to the determined sensitivity change. 4. The method of claim 1, further comprising calibrating the glucose signal at a frequency responsive to the determined sensitivity change. 5. The method of claim 1, further comprising calibrating the glucose signal based on the determined change in sensitivity. 6. The method of claim 1, further comprising determining a sensor stability based on the self-diagnosis. 7. The method of claim 6, further comprising controlling a display of the glucose signal based on the determined sensor stability. 8. The method of claim 1, further comprising implanting the sensor in a host. 9. A continuous glucose monitoring system configured for measuring glucose concentration in a host, the system comprising: a sensor comprising a first working electrode disposed beneath an active enzymatic portion of a membrane system and configured to generate a first signal, and a second working electrode disposed beneath an inactive-enzymatic or a non-enzymatic portion of the membrane system and configured to generate a second signal, wherein the membrane system is configured to substantially reduce interfering species from reaching the first working electrode and second working electrode; andelectronics configured to process the first signal and the second signal to produce a glucose signal that is substantially without signal contribution from interfering species, wherein the electronics are configured to perform a self-diagnosis based on a measurement from an auxiliary signal indicative of a change in sensitivity of the sensor, and wherein the measurement is obtained from one or more electrodes of the sensor independent of a blood glucos... | DexCom, Inc. | A61B005/05
US10299686 | B2 | 2019-05-28 | Method apparatus and system for analyzing images | 2008-12-28 | A method of identifying a thermally distinguishable region in a living body is disclosed. The method comprises determining transient thermal history of a surface of the living body, and estimating a location or presence of the thermally distinguishable region in the living body based on the transient thermal history. | 1. A method of identifying an internal thermally distinguishable region in a living body comprising: operating a thermospatial imaging system for imaging the living body and generating and storing in a memory a first thermospatial image describing a section of the living body when the living body is at a first posture, and a second thermospatial image describing said section of the living body when the living body is at a second posture, wherein each thermospatial image is defined over a three-dimensional (3D) spatial image of said body and has thermal data associated with a surface of the 3D spatial image;retrieving said thermospatial images from said memory;using a data processor of a computing platform for calculating a deformation function which maps a first 3D shape of said section of the living body, corresponding to said first posture, to a second 3D shape of said section of the living body, corresponding to said second posture, to determine correspondence between sets of picture-elements belonging to said first and said second thermospatial images;determining, by said computing platform, transient thermal variation for each picture-element in at least one set and storing said transient thermal variation in the memory; andestimating, by said computing platform, a location or presence of the internal thermally distinguishable region in the living body based on said transient thermal variations and generating, by said computing platform, an output indicative of the internal thermally distinguishable region. 2. The method of claim 1, further comprising estimating, by said computing platform, a size of said internal thermally distinguishable region, based on said transient thermal variation. 3. The method of claim 1, further comprising comparing, by said computing platform, said transient thermal variation to a reference transient thermal history, so as to determine said presence of said internal thermally distinguishable region. 4. The method of claim 1, further comprising heating or cooling a surface of said section of the living body between said first and said second thermospatial images. 5. The method of claim 1, wherein said estimating said location of said internal thermally distinguishable region comprises determining, by said computing platform, amount of temperature change for each point of a plurality of points over a surface of said section of the living body. 6. The method of claim 5, wherein said estimating said location of said internal thermally distinguishable region comprises determining, by said computing platform, a temperature change rate for at least one point of said plurality of points. 7. The method of claim 6, wherein said estimating said location of said internal thermally distinguishable region comprises comparing, by said computing platform, a temperature change rate of one point of said plurality of points to a temperature change rate of another point of said plurality of points. 8. The method of claim 5, w... | Real Imaging Ltd. | A61B001/00 A61B005/00 A61B005/01
US09475050 | B2 | 2016-10-25 | Container containing reference solution | 2010-06-21 | The present disclosure provides a container containing reference solution for clinical chemistry, whereby the container includes a closure part-covered outlet at which reference solution can be removed by puncturing with a puncturing element. The disclosure further provides a wick that is soaked with the reference solution and arranged in the container to extend from the outlet and such that a majority of its length extends into an interior space of the container. | 1. A container containing reference solution for clinical chemistry, including: an outlet for extracting reference solution by puncturing with a puncturing element of a puncturing device;a closure part covering the outlet;a wick having a length, the wick being positioned in the container and extending from the outlet such that a majority of the length extends into an interior space of the container to contact reference solution contained in the interior space; anda rigid molded wall that forms a stop that resists pressure exerted on the container by the puncturing device during extraction of a sample of the reference solution;wherein the rigid molded wall forms a channel extending between the interior space and the outlet, the wick being secured to the channel. 2. A container according to claim 1, wherein the wall surrounds the wick and carries the closure part. 3. A container according to claim 1, wherein the wall is configured to withstand a press-against force of at least 3 N when the puncturing device engages the container during extraction of the sample without reference solution being forced from the outlet. 4. A container according to claim 1, wherein the closure part is connected to the container by a pre-determined breakage zone. 5. A container according to claim 1, further including a pressure compensation channel that is closed by the closure part and extends into the interior space adjacent to the wick. 6. A container according to claim 1, wherein the wick includes a fiber bundle. 7. A container according to claim 1, wherein a majority of the length of the wick touches against the channel. 8. A container according to claim 1, wherein the closure part is removable and the wick is positioned in the container to project from the container, when the closure part is removed, by a distance that is smaller than a width of the wick. 9. A container according to claim 1, wherein the wick includes fibers that are distributed over a cross-section of the wick. 10. A container according to claim 1, wherein the wick immerses into the reference solution, when the outlet faces upwardly. 11. A container according to claim 1, wherein one end of the wick is positioned at the outlet. 12. A container containing reference solution for clinical chemistry, including: a wick disposed within an interior space of the container in contact with the reference solution contained in the interior space;a container wall configured to be pierced by a puncturing device upon extraction of the reference solution; anda ring-shaped disc inside the container that surrounds the wick and supports the container wall thereby forming a stop that resists pressure exerted on the container wall by the puncturing device during extraction of the reference solution. 13. A container containing reference solution for clinical chemistry, including: an outlet for extracting reference solution by puncturing with a puncturing element of a puncturing device;a closure part covering ... | Roche Diabetes Care, Inc. | B01L003/00
US20110021907 | A1 | 2011-01-27 | BIOMEDICAL IMAGING APPARATUS AND BIOMEDICAL TOMOGRAPHIC IMAGE GENERATION METHOD | 2010-07-22 | A biomedical imaging apparatus according to the present invention includes: an ultrasound generating section configured to output ultrasound to a predetermined region in an object under examination; an illuminating light generating section configured to emit illuminating light to the predetermined region upon which the ultrasound is incident; a phase component detecting section configured to time-resolve return light of the illuminating light emitted to the predetermined region, from the first time point to the Nth time point, and thereby detect the first to the Nth phase components of the return light corresponding to the first time point to the Nth time point; and a computing section configured to perform a process for subtracting a sum of the first to the (N−1)th phase components from the Nth phase component based on the phase components detected by the phase component detecting section. | 1. A biomedical imaging apparatus comprising: an ultrasound generating section configured to output ultrasound to a predetermined region in an object under examination;an illuminating light generating section configured to emit illuminating light to the predetermined region upon which the ultrasound is incident;a phase component detecting section configured to time-resolve return light of the illuminating light emitted to the predetermined region, from the first time point to the Nth time point, and thereby detect the first to the Nth phase components of the return light corresponding to the first time point to the Nth time point; anda computing section configured to perform a process for subtracting a sum of the first to the (N−1)th phase components from the Nth phase component based on the phase components detected by the phase component detecting section. 2. The biomedical imaging apparatus according to claim 1, wherein the computing section generates a tomographic image of the predetermined region using process results of the process as a pixel component. 3. The biomedical imaging apparatus according to claim 1, wherein the illuminating light is coherent light. 4. A biomedical tomographic image generation method comprising the steps of: outputting ultrasound to a predetermined region in an object under examination;emitting illuminating light to the predetermined region upon which the ultrasound is incident;time-resolving return light of the illuminating light emitted to the predetermined region, from the first time point to the Nth time point, and thereby detecting the first to the Nth phase components of the return light corresponding to the first time point to the Nth time point;performing a process for subtracting a sum of the first to the (N−1)th phase components from the Nth phase component based on the phase components detected by the phase component detecting section; andgenerating a tomographic image of the predetermined region using process results of the process as a pixel component. 5. The biomedical tomographic image generation method according to claim 4, wherein the illuminating light is coherent light. | OLYMPUS MEDICAL SYSTEMS CORP. | A61B008/14 A61B005/05
US07869861 | B2 | 2011-01-11 | Flexible tracking article and method of using the same | 2002-10-25 | A method of registering an article having a surface to previously created scan data of the article includes the steps of providing a flexible substrate having multiple tracking points attached to the substrate, applying the substrate to the article to be registered, creating a model of the surface of the article from a location of each tracking point; and registering the model with the previously created scan data. A method of creating a model of a surface having a shape and an article to enable the creation of a model of a surface are also disclosed. | 1. A method of registering an article having a surface to previously created scan data of the article, the method comprising the steps of: providing a flexible substrate having multiple tracking points attached to the substrate;applying the flexible substrate to the surface of the article to be registered;creating a model of the surface of the article from a location of each tracking point; andregistering the location of each tracking point in the model with the previously created scan data of the article by matching the locations of the tracking points in the model to the scan data of the article within a predetermined error level. 2. The method of claim 1 wherein the previously created scan data is pre-surgical scan data. 3. The method of claim 1 wherein there are at least 10 tracking points. 4. The method of claim 1 wherein the method is used to track surgical implements using a computer assisted surgical navigation system. 5. The method of claim 1 wherein the tracking points are optical. 6. The method of claim 5 wherein the tracking points are light emitting diodes. 7. The method of claim 1 wherein the substrate is adhesively applied to the article. 8. The method of claim 1 wherein the article is a portion of human anatomy. 9. The method of claim 8 wherein the human anatomy is forehead and nose. 10. The method of claim 1 wherein there are at least 25 tracking points. 11. The method of claim 1 including the additional step of trimming the flexible substrate to fit the article without deactivating any tracking points remaining attached to the flexible substrate. 12. A method of creating a model of a surface of a patient, wherein the surface has a shape, the method comprising the steps of: providing a flexible substrate having multiple tracking points attached to the substrate;attaching the flexible substrate to the surface;determining a location for each tracking point using a surgical navigation system; andcalculating the shape of the surface from the locations of the tracking points relative to each other. 13. The method of claim 12 wherein there are at least 10 tracking points. 14. The method of claim 12 wherein the method is used to track surgical implements using a computer assisted surgical navigation system. 15. The method of claim 12 wherein the tracking points are optical. 16. The method of claim 15 wherein the tracking points are light emitting diodes. 17. The method of claim 12 wherein the substrate is adhesively applied to the article. 18. The method of claim 12 wherein the article is a portion of human anatomy. 19. The method of claim 18 wherein the human anatomy is a forehead and nose. 20. The method of claim 12 wherein there are at least 25 tracking points. 21. The method of claim 12 including the additional step of trimming the flexible substrate to fit the surface without deactivating any tracking points remaining attached to the flexible substrate. | Howmedica Leibinger Inc. | A61B005/00
US07885832 | B2 | 2011-02-08 | Insurance policy and method for providing an insurance policy having dormancy features | 2006-10-02 | An insurance policy having a dormancy provision is provided in which the dormancy provision enables a policyholder to activate and deactivate benefits provided under the insurance policy. Also provided is a method for enabling a holder of an insurance policy with a dormancy provision to activate and deactivate benefits of an insurance policy, and includes purchasing an insurance policy having a dormancy provision and a set of associated benefits, paying premiums to maintain the set of benefits in an active state, deactivating the set of benefits; paying dormancy premiums to maintain the set of deactivated benefits in a deactivated state, and reactivating the set of deactivated benefits. | 1. A computer-implemented system for generating insurance policy data associated with an insurance policy providing insurance coverage for one or more entities, comprising: a computer processor programmed to: receive underwriting data relating to an entity to be insured;generate insurance policy data based upon the underwriting data, wherein the insurance policy data defines terms associated with an insurance policy for providing insurance coverage to the entity to be insured, wherein the insurance policy data defines an active coverage state and coverage and benefits to be provided to the entity during the active coverage state, and wherein the insurance policy data further defines a dormant state during which no coverage or benefits are to be provided to the entity, terms under which the dormant state can be commenced, and terms under which the active coverage state can be commenced to terminate the dormant state;determine an active coverage premium that the entity is required to pay during the active coverage state to maintain active coverage under the insurance policy using the underwriting data;determine a first dormancy premium that the entity is required to pay during the active coverage state to maintain the option of exercising the dormancy provision in the insurance policy using the underwriting data; anddetermine a second dormancy premium that the entity is required to pay during the dormant state to maintain the option of recommencing the active coverage period using the underwriting data; andan electronic database for storing the insurance policy data, the active coverage premium, and the first and second dormancy premium. 2. The system according to claim 1, wherein the active coverage state precedes the dormant state. 3. The system according to claim 1, wherein the dormant state precedes the active coverage state. 4. The system according to claim 1, wherein the first dormancy premium is less than the second dormancy premium. 5. The system according to claim 1, wherein the insurance policy data further defines one or more fees associated with commencement or termination of the active coverage state. 6. The system according to claim 1, wherein the underwriting data includes age and gender of the entity, health rating class of the entity, geographic area of the entity, and the provider network, plan and optional benefits of the entity. 7. The system according to claim 1, wherein the insurance policy is a health insurance policy. 8. The system according to claim 1, wherein the insurance policy is a life insurance policy. 9. The system according to claim 1, wherein the insurance policy is a property insurance policy. 10. A computer-implemented method for generating insurance policy data associated with an insurance policy providing insurance coverage for one or more entities, comprising: receiving underwriting data relating to an entity to be insured at a computer processor;using the computer processor to generate insurance poli... | Golden Rule Insurance Company | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07862522 | B1 | 2011-01-04 | Sensor glove | 2006-08-08 | A Sensor glove is an instrument, in the same way that a piano is an instrument. The Sensor glove is purchased and the owner then learns to use it and the sensors therein to develop the performance of a sound, character or model, or to “conduct” or “play” the sensor movements and experience the response of the model in real time. Several performers with a number of the Sensor glove can control over 200 motors or hundreds of control points on a character at one time, manipulate audio data, or computer images etc. The characteristics of the sensors on the gloves can be adjusted to the needs and demands of the wearer or the many performer and wearers in those venues where multiple performers are used with many gloves to control or produce the motions and sounds of many distinct characters in a screen play. | 1. A sensor glove worn on the fingers and hand of a performer, for providing analog signals representing the motion of one or more fingers in the glove, the sensor glove comprising: a palmer region, andat least one finger region, the finger region being worn on a finger, the finger region being attached to the palmer region, the finger region and palmer region each havinga corresponding palmer surface anda dorsal surface,a rigid reference base, the rigid reference base having a top surface and an underside, the rigid reference base underside being coupled to the dorsal surface of the palmer region,a mechanical link and means for coupling motion between the dorsal surface of the finger region and the rigid reference base to a control input for a pot to change the resistance of the pot in response to a bending motion of the finger in the finger region with respect to the rigid reference base,a mechanical link and means for translating a first and second degree of freedom motion of the thumb with respect to the rigid reference base into a thumb control signal, anda rigid flap, the rigid flap having an upper and lower surface, a first force sensitive resistor being coupled to the rigid flap upper surface, the rigid flap being pivotally coupled to the rigid reference base on a rotational axis transverse to the extended direction of a finger, the force sensitive resistor being positioned on the rigid flap upper surface at a location under the finger, the rigid flap lower side being supported or raised by motion of the thumb to permit the finger to provide a force to the first force sensitive resistor to change the resistance of the first force sensitive resistor in response to an increase in pressure between the thumb and the finger with the rigid flap and force sensitive resistor therebetween. 2. The sensor glove of claim 1 further comprising: an index finger region for receiving the index finger, the sensor glove havinga vertical motion sensing pot having a case coupled to the top of the rigid reference base and an input axis transverse to the longitudinal axis, the input axis of the vertical sensing pot being above the top surface of the rigid reference base, the vertical sensing pot input axis having a shaft with a bell crank there on, a link coupling the distal surface of the finger region of the index finger with an arm of the bell crank to provide rotational motion in a vertical plane to the input axis shaft in response to vertical motion of the top surface of the index finger region, anda horizontal motion sensing pot having a case coupled to a vertical flange, the vertical flange being integrally coupled to the top of the rigid reference base, the horizontal motion sensing pot having an input axis normal to the plane of the palmer region,the input axis of the horizontal sensing pot having a shaft with a bell crank there on, a link coupling the distal surface of the finger region of the index finger with an arm of the horizontal sensing pot be... | David Barclay | A61B005/103 A61B005/117 A61H001/00 A61H001/02 A61H005/00 B25J001/00 B25J003/00 B25J015/00 B25J019/00 B25J003/04 B25J013/08
US07890342 | B1 | 2011-02-15 | Method and system for tracking and monitoring patient compliance with medical device usage prescription | 2003-08-21 | A computer-implemented method for tracking and monitoring a patient's compliance with a medical device usage prescription for a medical device is disclosed. This method and system includes the steps of: (a) providing a minimum medical device usage compliance value for the medical device for a predetermined compliance period; (b) providing the quantity of compliance periods in a measurement cycle; and (c) determining a compliance period value as the number of compliance periods in the measurement cycle in which an actual medical device usage value is at least equal to the minimum medical device usage compliance value. In another embodiment, the method also includes: providing a weighting factor, and applying the weighting factor to at least one actual medical device session usage value. | 1. A method of measuring and utilizing efficiency of resources utilized by health care providers in the care of medical conditions, comprising: (a) obtaining at a computing device a minimum medical device usage compliance for a medical device for a predetermined compliance;(b) obtaining at the computing device a quantity of the compliance periods in a measurement cycle;(c) obtaining at the computing device measurements of actual medical device usage of the medical device during the measurement cycle;(d) weighting, within the computing device, the measurements of medical device usage for individual usage sessions during the measurement cycle according to a predetermined weighting scheme;(e) determining, within the computing device, whether a compliance period value for the measurement cycle, is at least equal to the minimum medical device usage compliance value, wherein the compliance period value for the measurement cycle is determined on the computing device as the number of compliance periods in the measurement cycle for which the weighted measurements of medical device usage indicate that the actual medical device usage has satisfied the compliance value, and wherein the weighted measurements of medical device usage indicate actual medical device usage has satisfied the compliance value, for a given compliance period responsive to the sum of the weighted measurements for the usage sessions that occurred during the given compliance period being greater than or equal to the compliance value;(f) obtaining, at the computing device, a medical device usage prescription value; and(g) comparing via the computing device, the compliance period value with the medical device usage prescription value. 2. The method of claim 1, wherein the compliance period value is determined as a percentage of compliance periods in the measurement cycle in which the weighted measurements of medical device usage indicate actual medical device usage has satisfied the compliance value. 3. The method of claim 1, further comprising: (f) obtaining, at the computing device, a measurement cycle compliance value;(g) comparing, in the computing device, the compliance period value with the measurement cycle compliance value; and(h) providing a compliance indicator from the computing device based upon results of the comparison. 4. The method of claim 3, wherein the compliance indicator indicates (1) a compliance signal responsive to the compliance period value being at least equal to the measurement cycle compliance value, and (2) a non-compliance signal responsive to the compliance period value being less than the measurement cycle compliance value. 5. The method of claim 3, wherein the compliance indicator indicates a compliance signal, a non-compliance signal, a marginal compliance signal, a warning signal, a trend signal, a compliance value signal, or any combination thereof. 6. The method of claim 3, further comprising (i) creating a report based upon the compliance indicat... | RIC Investments, LLC | G06Q100/00 G06Q050/00 A61B005/00 G06F019/00
US08337423 | B2 | 2012-12-25 | Blood glucose sensor | 2009-07-14 | The invention is directed to an apparatus and method for blood sample acquisition and testing. The apparatus incorporates a test strip having a blood sample volume detection area in which a blood sample is accumulated and a minimum blood sample volume is detected. Once a minimum blood sample volume is detected, the test strip is moved with respect to the blood sample extraction site so that a bending surface of the test strip is presented to the blood sample, utilizing the surface tension of the blood sample droplet to facilitate movement of the blood sample to a measurement site on the strip where blood glucose in the sample is measured. | 1. A blood glucose sensor for measuring blood glucose in a blood sample withdrawn from a blood sample extraction site on a subject's body, comprising: a test strip;conductive contacts positioned on the test strip and defining a blood sample volume detection area;said test strip having a blood transport channel having a mouth at one end in fluid communication with the blood sample volume detection area; anda measurement site at an opposite end of the blood transport channel;the test strip being bent at a plurality of inflections and moved between a first inflection at which the blood sample volume detection area is opposite the blood sample extraction site and a second inflection at which the mouth of the blood transport channel is opposite the blood sample extraction site,wherein a blood sample bridging the contacts defining the blood sample volume detection area permits electrical communication between the contacts to detect a blood sample volume; andwherein the detected blood sample volume is sufficiently large that the blood sample moves through the blood transport channel to contact the measurement site when the mouth of the blood transport channel is moved to said blood sample extraction site. 2. The blood glucose sensor of claim 1, further comprising a bending surface supporting the test strip when bent at an inflection opposite the blood sample extraction site. 3. The blood glucose sensor of claim 2, wherein the bending surface is a roller and the test strip is movable over the roller between the first and second inflections. 4. The blood glucose sensor of claim 3, wherein the test strip arranged on the roller is continuous, having a plurality of functional areas, each functional area on the test strip comprising a volume detection area, blood transport channel and measurement site, and each functional area is separated by a neutral portion, the functional areas being arranged on the test strip for successive blood glucose measurements. 5. The blood glucose sensor of claim 4, further comprising a housing and a lancet arranged in the housing to pass through the test strip at the blood volume detection area and to withdraw a blood sample from a subcutaneous space in a subject's body beneath the blood sample extraction site. 6. The blood glucose sensor of claim 5, further comprising a first seal isolating the functional areas of the test strip that have been used for blood glucose measurement from an environment outside of the housing, and a second seal isolating the unused areas of test strip from the environment. 7. The blood glucose sensor of claim 1, further comprising a cartridge containing a plurality of test strips, and a guide member for bending each individual test strip at a plurality of inflections opposite a blood sample extraction site on a patient's body. 8. The blood glucose sensor of claim 1, wherein the blood sample volume detection area is sized so that a blood sample having a volume of about 0.2 μL to about 1.0 μL ca... | Becton, Dickinson and Company | A61B005/00 B65D081/00
US08221337 | B2 | 2012-07-17 | System and method for monitoring bladder distention in a variety of patients having differing anatomical proportions | 2009-07-02 | A system and method are provided for monitoring bladder distention. Independently controllable acoustic transmitters are coupled to the surface of a user's abdomen region in a spaced-apart fashion. Each transmitter can introduce a broadband acoustic signal into a portion of the abdomen region. Acoustic receivers are also coupled to the surface of the abdomen region to detect acoustic reflections generated when the acoustic signals are introduced into the abdomen region. A controller is positioned on the user and is correlated to an orientation of the user. The controller stores a profile indicative of anatomical features of the user. The controller initiates periodic activations of selected ones of the transmitters based upon the profile and the orientation of the user. The controller uses a weighted combination of all of the acoustic reflections to determine bladder distention of the user. The weighted combination is defined by at least the profile and orientation of the user. | 1. A system for monitoring bladder distention, comprising: a plurality, of independently controllable acoustic transmitters adapted to be coupled to the surface of a user's abdomen region in a spaced-apart fashion, wherein a first portion of said transmitters is arranged on one side of the abdomen region relative to a vertical centerline thereof extending along the user's torso, a second portion of said transmitters is arranged on an opposing side of the abdomen region in an approximate mirror image fashion with respect to said first portion, and a third portion of said transmitters is arranged on an upper region of the abdomen region, each of said transmitters including an active element made from a piezocomposite material for broadband acoustic transmission in a range from approximately 500 KHz to approximately 2 MHz, said piezocomposite material being semi-spherical in shape with a convex portion thereof being adapted to face the surface of the abdomen region when said transmitters are coupled thereto wherein each of said transmitters introduces a broadband acoustic signal into a portion of the abdomen region;a plurality of acoustic receivers adapted to be coupled to the surface of the abdomen region, each of said receivers positioned for detecting acoustic reflections generated when each said acoustic signal is introduced into the abdomen region; anda controller coupled to said transmitters and said receivers, said controller adapted to be positioned on the user and correlated to an orientation of the user, said controller storing a profile indicative of anatomical features of the user, said controller initiating periodic activations of selected ones of said transmitters from said first portion and said second portion based upon said profile and said orientation of the user wherein each said acoustic signal resulting from said periodic activations generates one of said acoustic reflections associated therewith for detection by one of said receivers,said controller initiating selected activation of said third portion of said transmitters wherein each said acoustic signal resulting from said selected activation generates one of said acoustic reflections associated therewith for detection by at least one of said receivers, andsaid controller using a weighted combination of all of said acoustic reflections to determine bladder distention of the user, wherein said weighted combination is defined by at least said profile and said orientation of the user, said controller generating an alarm signal when said bladder distention indicates that a bladder voiding is needed. 2. A system as in claim 1, wherein each of said transmitters from said first portion and said second portion is combined with one of said receivers to form a composite trancducer. 3. A system as in claim 1, wherein said controller includes a device for continually sensing orientation of said controller, and wherein orientation of said controller is indicative of said orientation of ... | ODU Research Foundation | A61B005/103 A61B008/00 G08B023/00 G01S003/80
US08556812 | B2 | 2013-10-15 | Blood testing device | 2009-03-10 | A blood testing device in which a blood sensor does not contaminate a piercing section when discarded. The blood testing device (11) is provided with the piercing section (13) having an upper holder (13a) and a lower holder (13b), and also with a connector (18). In piercing operation, the connector (18), the upper holder (13a), and the lower holder (13b) sandwich and hold the sensor (24). To discharge the sensor, the lower holder (13b) is pivoted so as to be separated from the upper holder (13a), which creates a space between the lower holder (13b) and the upper holder (13a) to expose the sensor (24). | 1. A blood test apparatus comprising: a first holding section composed of a first holder and a second holder that are configured to sandwich a blood sensor therebetween;a second holding section that is configured to hold the blood sensor sandwiched between the first holder and the second holder; anda puncturing section that punctures skin from a first holder side of the blood sensor sandwiched between the first holder and the second holder, wherein;the second holding section is rotatable around a rotation axis while holding the blood sensor,the second holder is rotatable around the rotation axis so as to define a space between the first holder and the second holder, andthe second holding section is pivotably mounted on a rotating shaft that is coaxial with the second holder. 2. The blood test apparatus according to claim 1, wherein the puncturing section is configured to penetrate the blood sensor to puncture the skin. 3. The blood test apparatus according to claim 1, wherein the first holder is fixed to a housing, andthe second holder faces the first holder and is pivotably mounted on a rotating shaft of the housing. 4. The blood test apparatus according to claim 1, wherein the first holder is fixed to a housing, andthe second holder faces the first holder and is pivotably mounted on a rotating shaft of the first holder. 5. The blood test apparatus according to claim 1, wherein the second holding section is pivotably mounted on a rotating shaft of a housing. 6. The blood test apparatus according to claim 1, wherein the second holding section is pivotably mounted on a rotating shaft of the first holder. 7. The blood test apparatus according to claim 1, wherein the second holding section has a sliding mechanism that is configured to release a holding of the blood sensor when the second holder rotates. 8. The blood test apparatus according to claim 7, wherein the sliding mechanism includes an inclined part having a tapered surface to push the blood sensor. 9. The blood test apparatus according to claim 1, wherein the second holding section is pivotably mounted on a rotating shaft of a housing; andwhen the second holding section rotates by a predetermined angle, the second holder rotates by an angle greater than the predetermined angle of the second holding section. 10. The blood test apparatus according to claim 1, further comprising: a sensor cartridge in which blood sensors are stacked and stored; anda conveying section that conveys the blood sensors stacked and stored in the sensor cartridge to the puncturing section. 11. The blood test apparatus according to claim 1, further comprising: a pushing section provided in the second holder and configured to push the blood sensor so as to release a holding of the blood sensor; andan ejecting button which is configured to operate the pushing section. 12. The blood test apparatus according to claim 11, wherein the pushing section includes a lifting section that is configured to lift the blood... | Panasonic Corporation | A61B005/00
US08655437 | B2 | 2014-02-18 | Analysis of the mirror neuron system for evaluation of stimulus | 2009-08-21 | The human mirror neuron system includes neurons that fire both when an individual performs an action and when the individual observes the action being performed by another. Neuro-response data involving the mirror neuron system is collected as a subject is exposed to stimulus material. The stimulus material may include individuals performing actions such as making a purchase, accepting an offer, participating in an activity, etc. Neuro-response data involving the mirror neuron system of the subject is analyzed to determine the propensity of the subject to act. | 1. A method, comprising: analyzing, using a processor, neuro-response data from a subject exposed to first stimulus material, wherein the neuro-response data includes data obtained from areas of a brain of the subject associated with a mirror neuron system, and the neuro-response data includes gamma band data and mu band data;identifying at least one of (1) an increase in mirror neuron activity based on an increase in gamma band data and a decrease in mu band data or (2) a decrease in mirror neuron activity based on a decrease in gamma band data and an increase in mu band data; andgenerating, using the processor, a propensity to act index indicating at least one of (a) that the subject is likely to take an action based on the increase in the mirror neuron activity or (b) that the subject is not likely to take the action based on the decrease in the mirror neuron activity. 2. The method of claim 1, wherein the gamma band data comprises gamma rhythms. 3. The method of claim 1, wherein the neuro-response data includes beta band data and the propensity to act index is further based on a ratio of the gamma band data to the beta band data, and the beta band data comprises beta rhythms. 4. The method of claim 1, wherein the mu band data comprises mu rhythms. 5. The method of claim 1, wherein the propensity to act is further based on a ratio of the gamma band data to the mu band data. 6. The method of claim 1, wherein the neuron-response data is obtained using electroencephalography. 7. The method of claim 1, wherein the neuro-response data is obtained using magnetoencephalography. 8. The method of claim 1 further comprising generating a baseline by analyzing second neuro-response data gathered from the subject while exposed to second stimulus material before the subject is exposed to the first stimulus material. 9. The method of claim 1, wherein the propensity to act index is calibrated based on a demographic group. 10. The method of claim 9, wherein the propensity to act index is calibrated based on a gender. 11. The method of claim 1 further comprising generating a propensity to reach index based on the neuro-response data. 12. The method of claim 1 further comprising generating a learning index based on the neuro-response data. 13. The method of claim 1 further comprising generating an empathy index based on the neuro-response data. 14. A system, comprising: a data collector to obtain neuro-response data from a subject exposed to first stimulus material, wherein the neuro-response data includes data obtained from areas of a brain of the subject associated with a mirror neuron system, the neuro-response data including gamma band data and mu band data; anda data analyzer to: identify at least one of (1) an increase in mirror neuron activity based on (a) an increase in gamma band data and (b) a decrease in mu band data or (2) a decrease in mirror neuron activity based on (a) a decrease in gamma band data and (b) an increase in mu band data;... | The Nielsen Company (US), LLC | A61B005/04 A61B005/05
US20110004085 | A1 | 2011-01-06 | PERFORMANCE REPORTS ASSOCIATED WITH CONTINUOUS SENSOR DATA FROM MULTIPLE ANALYSIS TIME PERIODS | 2010-04-29 | A method of providing performance reports to persons associated with the monitoring of an analyte level, such as blood glucose concentrations in a host, is provided. The performance reports can indicate a measure of an analyte variability of the host over a first analysis time period compared to a measure of analyte variability of the host over one or more other analysis time periods. The analyte level of the host can be received in each of a plurality of sensor readings from a continuous analyte sensor, such as multiple sensor readings per hour, so that the analyte variability that is used in determining the performance report of the host can be based on not just a few metered readings taken by the host, but a plurality of intermittent readings that are taken by a continuous analyte sensor, for example. Various types of performance reports can be generated, such as reports that comprise charts, graphs, tables, and/or other visual indicia of a host's performance during one analysis period versus another analysis period. | 1. A method of analyzing continuous analyte sensor data associated with a host, wherein the continuous analyte sensor data is representative of glucose levels of the host, the method comprising: accessing a plurality of blood glucose readings associated with a host during two or more analysis time periods, wherein each blood glucose reading is indicative of a blood glucose level of the host at a respective time; andgenerating a performance report indicating a relationship between the plurality of blood glucose readings of the two or more analysis time periods, wherein the performance report is indicative of glucose control by the host over the two or more analysis time periods. 2. The method of claim 1, wherein the performance report further indicates a relationship between the plurality of blood glucose readings of the two or more analysis time periods and target sensor data. 3. The method of claim 1, further comprising: initiating transmission of appropriate clinical feedback to the host in response to determining that the relationship between the plurality of blood glucose readings meets criteria indicative of providing the host with clinical feedback. 4. The method of claim 1, wherein the plurality of blood glucose readings comprises more than 72 blood glucose readings for each 24 hour period of an at least one week period. 5. The method of claim 1, further comprising transmitting the performance report to a caretaker computing device. 6. A method of analyzing continuous analyte sensor data associated with a host, wherein the continuous analyte sensor data is representative of glucose levels of the host, the method comprising: receiving from a user an indication of a first starting date of a first analysis time period;determining a first ending date of the first analysis time period by adding a predetermined time period to the first starting date;determining a second starting date of a second analysis time period;determining a second ending date of the second analysis time period by adding the predetermined time period to the second starting date;accessing a plurality of blood glucose levels associated with a host during the first analysis time period and the second analysis time period, wherein each blood glucose level is indicative of a blood glucose level of the host at respective times;determining one or more first statistical values associated with the first analysis time period based on each of the blood glucose levels of the host received during the first analysis time period;determining one or more second statistical values associated with the second analysis time period based on each of the blood glucose levels of the host received during the second analysis time period; anddetermining a comparison indicator indicating a relationship between the first one or more statistical values and the second one or more statistical values. 7. The method of claim 6, wherein the second starting data of the second analysis time period is a da... | DexCom, Inc. | A61B005/145
US07885697 | B2 | 2011-02-08 | Transcutaneous analyte sensor | 2005-03-10 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. An analyte measuring system, the system comprising: a mounting unit adapted for mounting on a skin of a host;an electronics unit detachably connectable to the mounting unit;a sensor for measuring a level of the analyte, the sensor comprising an in vivo portion for insertion through the skin at an exit site of the host and an ex vivo portion connected to the mounting unit, wherein the sensor is operably connected to the electronics unit when the electronics unit is detachably connected to the mounting unit; anda communication station for receiving and processing analyte measurements received from the electronics unit, wherein the mounting unit comprises an opening configured to releasably receive the electronics unit therein, and wherein the electronics unit and the mounting unit each comprise at least one electrical contact, wherein each of the at least one electrical contacts is protected from moisture by an elastomeric sealing member that fits within the mounting unit, wherein the elastomeric sealing member forms an interference fit with the electronics unit. 2. The system of claim 1, wherein the communication station comprises a hand-held receiver. 3. The system of claim 1, wherein the electronics unit body is designed to be used with a plurality of sensors. 4. A transcutaneous analyte sensor system comprising: an electrode configured to measure a level of glucose in a host;a membrane disposed over the electrode and configured to limit a transport of glucose to the electrode; andsensor electronics operably connected to the electrode and configured to measure a current flow produced by the electrode, wherein the sensor electronics are configured to measure the current flow in at least a picoAmp range, and wherein the system is configured to have a glucose sensitivity of from about 1 pA/mg/dL to about 9 pA/mg/dL, and wherein the electrode comprises an exposed electroactive working electrode with a surface area between about 0.000084 cm2 and about 0.016 cm2. 5. The transcutaneous analyte sensor system of claim 4, wherein the sensor electronics is configured to directly measure the level of current flow produced by the electrode. 6. The transcutaneous analyte sensor of claim 4, further comprising an analog-to-digital converter configured to translate the current flow measurement to a digital signal. 7. The system of claim 4, wherein the system is configured to have a glucose sensitivity of from about 3.5 pA/mg/dL to about 7.5 pA/mg/dL. 8. The transcutaneous analyte sensor of claim 4, wherein the membrane comprises a resistance domain configured with a permeability ratio of at least about 50:1 analyte to co-analyte concentration. 9. The transcutaneous analyte sensor of claim 8, wherein the permeability ratio is at least about 200:1. 10. A sensor system for measuring an analyte concentration in a host, the sensor system comprising: a sensor comprising at least one electrode;sensor electronics configured to measure the analyte concentrati... | DexCom, Inc. | A61B005/00 A61B005/05
US08613707 | B2 | 2013-12-24 | System and method for monitoring cardiac output | 2008-11-14 | A method for monitoring cardiac output (pulmonary blood flow) of a subject, the method including: measuring a first net pulmonary uptake or elimination of a breathed gas species by the subject and a first partial pressure of the gas species at a first time, and at a second time later than the first time, determining a first pulmonary blood flow of the patient at the first time, and determining a pulmonary blood flow of the subject at the second time on the basis of the first pulmonary blood flow, the first net pulmonary uptake or elimination, the second net pulmonary uptake or elimination, the first partial pressure, and the second partial pressure. | 1. A method for monitoring cardiac output (pulmonary blood flow) of a subject, the method including: (i) making a first measurement of a gas species breathed by said subject at a first time to determine a first measurement value;(ii) determining a first pulmonary blood flow of said subject at said first time;(iii) making a second measurement of said gas species breathed by said subject at a second time later than said first time to determine a second measurement value; and(iv) determining a pulmonary blood flow of said subject at said second time on the basis of a product of said first pulmonary blood flow and a square of a ratio of the first measurement value and the second measurement value;wherein:(a) said first measurement value represents a net rate of pulmonary uptake or elimination of said gas species by said subject at said first time, and said second measurement value represents a net rate of pulmonary uptake or elimination of said gas species by said subject at said second time; or(b) said first measurement value represents a difference in partial pressures of said gas species between inspired and end-tidal gas at said first time, and said second measurement value represents a difference in partial pressures of said gas species between inspired and end-tidal gas at said second time; or(c) said first measurement value represents a partial pressure in end-tidal gas of said gas species at said first time, and said second measurement value represents a partial pressure in end-tidal gas of said gas species at said second time. 2. A method as claimed in claim 1, including: determining a first rate of change of partial pressure of said gas species in lungs of said subject at said first time; anddetermining a second rate of change of partial pressure of said gas species in lungs of said subject at said second time;wherein said pulmonary blood flow of said subject at said second time is determined on the basis of said first and second rates of change of partial pressure of said gas species in lungs of said subject at said first and second times. 3. A method according to claim 2, wherein {dot over (Q)}tk, the pulmonary blood flow ({dot over (Q)}t) for a breath k ({dot over (Q)}tk), is determined according to:Q.⁢tk=Q.⁢ti·(V.GkV.Gi)2where {dot over (V)}Gi and {dot over (V)}Gk are net rates of pulmonary uptake or elimination of a breathed gas species G for breath k and for an earlier breath i, respectively, and {dot over (Q)}ti is a pulmonary blood flow of said subject for said breath i. 4. A method according to claim 3, wherein the net rate of pulmonary uptake or elimination of said gas species G for each breath is approximated by a difference between an inspired partial pressure of said gas species G (PIG) and its partial pressure in end-tidal gas (PE′G) for a corresponding breath, so that said pulmonary blood flow {dot over (Q)}tk for a breath k of said subject is determined from {dot over (Q)}ti, a pulmonary blood flow of said subject for an ... | Austin Health | A61B005/02
US08560035 | B2 | 2013-10-15 | Physiological sensor with a tail | 2010-08-20 | An exemplary sensor includes a sensor pad defining a plurality of openings and a circuit board. The circuit board is at least partially disposed in the sensor pad and has a light source configured to generate near-infrared light and a light detector configured to receive near-infrared light. The light source and the light detector are each aligned with one of the openings of the sensor pad so that near-infrared light generated by the light source can travel through a part of a patient's body to the light detector and the near-infrared light received by the light detector indicates oxygen saturation of the part of the patient's body through which the light travelled. The circuit board further includes an integrally formed tail at least partially disposed in the sensor pad. Additionally, the sensor pad includes an overlay disposed on the tail. | 1. A sensor comprising: a sensor pad having a first layer and a second layer, the second layer defining a plurality of openings; anda circuit board at least partially disposed in the sensor pad and having a light source configured to generate near-infrared light and a light detector configured to receive near-infrared light, wherein the light source and the light detector are each aligned with one of the openings of the sensor pad so that near-infrared light generated by the light source can travel through a part of a patient's body to the light detector and the near-infrared light received by the light detector indicates oxygen saturation of the part of the patient's body through which the light travelled, andwherein the circuit board has an integrally formed tail at least partially disposed in the sensor pad, and wherein the sensor pad includes a first overlay disposed on one side of the tail and integrally formed with the first layer and a second overlay disposed on another side of the tail and integrally formed with the second layer, wherein the first and second overlays extend along the entire length of the tail. 2. A sensor as set forth in claim 1, wherein the tail is substantially flat. 3. A sensor as set forth in claim 2, wherein the first overlay and second overlay are substantially flat. 4. A sensor as set forth in claim 1, wherein the tail has a length that is at least three-fourths of a circumference of a region of the patient's body in which the sensor is applied. 5. A sensor as set forth in claim 1, wherein the tail is flexible. 6. A sensor as set forth in claim 1, wherein the circuit board includes a flexible printed circuit board. 7. A sensor as set forth in claim 1, further comprising a cable configured to carry control signals from a controller to the light source and signals representative of oxygen saturation from the light detector to the controller. 8. A sensor as set forth in claim 7, wherein the tail, first overlay, and second overlay, when combined, have a thickness that is less than a thickness of the cable. 9. A sensor as set forth in claim 7, further comprising a connector disposed on the circuit board and configured to electrically connect the cable to the circuit board. 10. A sensor as set forth in claim 9, wherein the connector is disposed at an end of the tail. 11. A sensor as set forth in claim 1, further comprising a spacer disposed in the sensor pad and on the circuit board. 12. A sensor as set forth in claim 1, wherein the first layer is spaced from the second layer. 13. A sensor as set forth in claim 1, further comprising an adhesive disposed on the sensor pad. 14. A sensor as set forth in claim 1, wherein the sensor pad is at least partially formed from a flexible material. 15. A sensor as set forth in claim 1, wherein the sensor pad is at least partially formed from a light-blocking material. 16. A sensor comprising: a first layer formed from a flexible material and defining a plurality of ope... | Covidien LP | A61B005/00
US07869849 | B2 | 2011-01-11 | Opaque, electrically nonconductive region on a medical sensor | 2006-09-26 | A medical sensor may be adapted to prevent unwanted light and electrical interference from corrupting physiological measurements. Sensors are provided with features that reduce the amount of outside light or shunted light that impinge the detecting elements of the sensor. The sensor is adapted to reduce crosstalk between electrical signals, increasing the accuracy of measurements. The sensor is also adapted to reduce the effect of outside light or shunted light on pulse oximetry measurements. | 1. A sensor comprising: a sensor body supporting respective electrical connectors connected to an emitter and a detector, wherein the emitter and the detector are disposed on a tissue-contacting surface of the sensor body; andat least one opaque member disposed on the sensor body and spanning only a portion of the sensor body between the emitter and the detector, the opaque member comprising a substantially electrically nonconductive material, and wherein the opaque member does not surround the emitter and detector. 2. The sensor, as set forth in claim 1, wherein the sensor comprises at least one of a pulse oximetry sensor or a sensor for measuring a water fraction. 3. The sensor, as set forth in claim 1, wherein the at least one opaque member comprises a polyester-based material. 4. The sensor, as set forth in claim 1, wherein the at least one opaque member comprises a silicone-based material. 5. The sensor, as set forth in claim 1, wherein the at least one opaque member is conformable to the tissue. 6. The sensor, as set forth in claim 1, wherein the emitter comprises at least one light emitting diode and the detector comprises at least one photodetector. 7. The sensor, as set forth in claim 1, wherein the at least one opaque member is substantially dark in color. 8. The sensor, as set forth in claim 1, wherein the at least one opaque member is substantially light in color. 9. A pulse oximetry system comprising: a pulse oximetry monitor; anda pulse oximetry sensor adapted to be operatively coupled to the monitor, the sensor comprising:a sensor body; andat least one opaque member disposed on the sensor body and covering at least a portion of the sensor body supporting respective electrical leads coupled to an emitter and a detector, the opaque member comprising a substantially electrically nonconductive material, and wherein the opaque member does not surround the emitter and detector. 10. The system, as set forth in claim 9, wherein the pulse oximetry sensor transmits an electrical signal, wherein the electrical signal comprises at least one measurement. 11. The system, as set forth in claim 9, wherein the at least one opaque member is disposed on at least 90% the sensor body. 12. The system, as set forth in claim 9, wherein the at least one opaque member is adhesively disposed on the sensor body. 13. The system, as set forth in claim 9, wherein the at least one opaque member comprises a polyester-based material. 14. The system, as set forth in claim 9, wherein the at least one opaque member comprises a silicone-based material. 15. The system, as set forth in claim 9, wherein the at least one opaque member comprises an adhesive material adapted to contact the tissue. 16. The system, as set forth in claim 9, wherein the at least one opaque member is conformable to patient tissue. 17. The system, as set forth in claim 9, wherein the emitter comprises at least one light emitting diode and the detector comprises at least one photode... | Nellcor Puritan Bennett LLC | A61B005/1455
US07875037 | B2 | 2011-01-25 | Infant umbilical cord cardiac monitoring system and method | 2007-07-23 | An infant umbilical cord monitoring system having a hinged housing adapted to receive both an umbilical cord clamping means and a means for measuring physiological data of the infant. A method of using an infant's umbilical cord for obtaining physiological data includes applying an umbilical cord clamping means to the infant's umbilical cord, positioning a hinged housing adapted to receive both the umbilical cord clamping means and a means for measuring physiological data of the infant against the infant and detecting the physiological data of the infant. | 1. An infant umbilical cord monitoring system utilized at birth comprising: a hinged housing having a central opening configured to surround an umbilical cord, the hinged housing having a first side and a second side opposite to the first side that is configured to contact the infant, the first side of the hinged housing is adapted to receive an umbilical cord clamping means such that the umbilical cord clamping means is configured to engage the umbilical cord as the umbilical cord passes through the central opening, and a means for measuring immediately and in real time physiological data of the infant continuously from the moment of birth which passes through an orifice on another side of the hinged housing wherein the means for measuring immediately and in real time physiological data is configured to be in direct contact with the infant and/or the infant's umbilical cord. 2. An infant umbilical cord monitoring system as claimed in claim 1 wherein the first side of the hinged housing further comprises at least one slot adapted to receive the umbilical cord clamping means. 3. An infant umbilical cord monitoring system as claimed in claim 2 wherein the umbilical cord clamping means is a clamp having a hinged first end and a second end having a locking mechanism. 4. An infant umbilical cord monitoring system as claimed in claim 3 wherein the clamp further comprises at least one slot member adapted to engage the slot in the hinged housing to secure the clamp to the hinged housing. 5. An infant umbilical cord monitoring system as claimed in claim 1 wherein the hinged housing is a clamp having a first portion and a second portion able to move between an open position and a closed position. 6. An infant umbilical cord monitoring system as claimed in claim 5 wherein the first portion and the second portion are joined together by a hinge. 7. An infant umbilical cord monitoring system as claimed in claim 5 wherein the first and second portions are secured together by a locking mechanism. 8. An infant umbilical cord monitoring system as claimed in claim 5 wherein the orifice is positioned in the hinged housing at an angle and is adapted to receive the means for measuring physiological data of the infant. 9. An infant umbilical cord monitoring system as claimed in claim 1 wherein the central opening has a textured surface configured to provide traction against the infant umbilical cord. 10. An infant umbilical cord monitoring system as claimed in claim 1 wherein the means for measuring physiological data of the infant is an ultrasound unit. 11. An infant umbilical cord monitoring system as claimed in claim 1 wherein the means for measuring physiological data of the infant is an electric device, a mechanical device or an electrical/mechanical device. 12. An infant umbilical cord monitoring system as claimed in claim 11 wherein the electrical device is an electrocardiogram. 13. An infant umbilical cord monitoring system as claimed in claim 11 wh... | Jengstar-MD Medical Technologies Ltd. | A61B017/42 A61B005/00 A61B008/14
US07887492 | B1 | 2011-02-15 | Monitoring device, method and system | 2006-06-24 | A monitoring device (20) and method (200) for monitoring the health of a user is disclosed herein. The monitoring device (20) is preferably an article (25), an optical sensor (30), a circuitry assembly (35) a display member (40) and a control component (43). The monitoring device (20) preferably displays the following information about the user: pulse rate; calories expended by the user of a pre-set time period; target zones of activity; time; and distance traveled. | 1. A monitoring device for monitoring the health of a user, the monitoring device comprising: an article configured to be worn on a user's hand, the article comprising a thumb portion and a main body portion having a palm portion and a back portion, the article having a mass ranging from once ounce to five ounces;an optical sensor to measure blood flow through an artery of the user and to generate a digital signal corresponding to the flow of blood through the artery of the user, the optical sensor disposed on the thumb portion of the article, wherein the optical sensor is a light-to-frequency photodetector capable of transmitting a digital signal and at least one light emitting diode capable of radiating light ranging from 600 nanometers to 1100 nanometers;a microprocessor configured to process the digital signal from the optical sensor to generate a real-time heart rate value for the user from average pulse values filtered by a first heart rate value based on an average pulse value of a secondary set of time periods subtracted from an average pulse value of a primary set of time periods microprocessor disposed on the article;means for visually displaying the real-time heart rate value of the user, the visually displaying means attached to an exterior surface of the back portion of the article, the visually displaying means comprising a light emitting diode array having a plurality of columns and a plurality of rows, each of the plurality of columns capable of separate illumination to provide the appearance of a moving display; andmeans for controlling the input information and the output of information displayed on the visually displaying means, the controlling means extending from the exterior surface of the article. 2. The monitoring device according to claim 1 wherein the controlling means is a plurality of control buttons to input information and control the output of information on the visually displaying means. 3. The monitoring device according to claim 1 wherein the main body portion further comprises an attachment means and the thumb portion further comprises an attachment means. 4. The monitoring device according to claim 1 wherein the visually displaying means is capable of displaying a real-time pulse waveform of the user. 5. A monitoring device for monitoring the health of a user, the monitoring device comprising: an article to be worn on the user's hand, the article comprising a main body portion and a thumb portion, the main body portion comprising a palm portion, a back portion and an attachment means, the thumb portion comprising an annular portion for placement around a portion of a thumb of the user's and an attachment means for securing the thumb portion to the user's thumb, the article having a mass ranging from one ounce to five ounces;an optical sensor disposed on an interior surface of the thumb portion of the article, the optical sensor generating a digital signal corresponding to a flow of blood through an artery o... | Impact Sports Technologies, Inc. | A61B005/02
US07885706 | B2 | 2011-02-08 | System and device for seizure detection | 2006-09-20 | A device comprises a head mounting arrangement sized and shaped to be worn on a user's head and a plurality of electrodes disposed on the arrangement so that, when the arrangement is worn on the user's head, the electrodes contact target portions of a scalp to detect electrical activity of a brain of the user in combination with an image capture device disposed on the arrangement so that, when the arrangement is worn on the user's head, a field of view of the image capture device includes a portion of an anatomy of the user and a processing unit generating EEG data from the electrical activity, wherein, when the EEG data is indicative of an epileptic event, the processing unit activates the image capture device to capture video data of the user and may store the EEG and/or the video data with transmission of warning signals to one or more remote displaying and/or computing arrangements. | 1. A device, comprising: a head mounting arrangement sized and shaped to be worn on a user's head;a plurality of electrodes disposed on the arrangement so that, when the arrangement is worn on the user's head, the electrodes contact target portions of a scalp to detect electrical activity of a brain of the user;an image capture device disposed on the arrangement so that, when the arrangement is worn on the user's head, a field of view of the image capture device includes a portion of an anatomy of the user; anda processing unit generating EEG data from the electrical activity,wherein, when the EEG data is indicative of an epileptic event, the processing unit activates the image capture device to capture video data of the user. 2. The device according to claim 1, wherein the arrangement is one of a headband and a cap. 3. The device according to claim 1, wherein the electrodes include at least four electrodes disposed in a predetermined configuration on the arrangement. 4. The device according to claim 1, wherein each of the electrodes is contained in an electrode unit, the electrode unit further comprising an operational amplifier coupled directly to the corresponding electrode. 5. The device according to claim 4, wherein the electrode unit further comprises: an adhesive delivery arrangement securing the electrode unit to the scalp; anda conductive paste delivery arrangement applying a conductive paste to electrically couple the electrode to the scalp. 6. The device according to claim 1, wherein the image capture device is one of a video camera and a digital camera. 7. The device according to claim 1, wherein the field of view of the image capture device includes one of trunk, hands and feet of the user. 8. The device according to claim 1, further comprising: a memory,wherein the processing unit writes at least a portion of one of the EEG data and the video data to the memory. 9. The device according to claim 8, wherein the memory is a removable memory device. 10. The device according to claim 8, wherein the memory includes reference data corresponding to EEG data obtained from one of the subject and a group of subjects. 11. The device according to claim 10, wherein the processing unit compares the EEG data to the reference data to determine if the user is experiencing the epileptic event. 12. The device according to claim 11, wherein the epileptic event is a seizure. 13. The device according to claim 11, wherein the epileptic event is brain activity preceding a seizure. 14. The device according to claim 1, further comprising: an indicator activated upon detection of the epileptic event. 15. The device according to claim 14, wherein the indicator is one of a light-emitting diode and a speaker. 16. The device according to claim 1, further comprising: a wireless transceiver transmitting a warning signal upon detection of the epileptic event to one of a remote display device and a remote computing device. 17. The device according to c... | New York University | A61B005/0476
US08521256 | B2 | 2013-08-27 | Retrospective motion compensation in magnetic resonance imaging | 2009-04-14 | A magnetic resonance imaging apparatus comprises a data acquisition system with motion compensation means for acquiring an image data set in respect of a moving object within an imaging region, the image data set comprising a plurality of acquisitions each in respect of one or more slice planes. Further there are means for averaging, on a slice-by-slice basis, said plurality of acquisitions to generate a single, average image set of said moving object. Also means are provided for calculating the actual scan location of all of said slice planes from all of said acquisitions, means for averaging a selected set of slice planes to generate a single, average image set of a static object within said imaging region. This averaging results in a high image quality in respect of the moving object (12) because for the static object, slice locations change for each acquisition. | 1. A magnetic resonance imaging apparatus, comprising: a data acquisition system including means for motion compensation, the data acquisition system being configured to acquire an image data set of a moving object within an imaging region, said image data set comprising a plurality of acquisitions, each acquisition comprising one or more slice planes located at a same relative scan location of said moving object, respectively, throughout said plurality of acquisitions;means for averaging respective one or more slice planes from said plurality of acquisitions to generate a single, average image set of said moving object;means for calculating an actual scan location of all of said one or more slice planes from all of said acquisitions with respect to a static object within the imaging region; andmeans for averaging a selected set of slice planes, based on said calculation, to generate a single, average image set of the static object. 2. The apparatus according to claim 1, wherein said means for motion compensation comprises navigator echo means. 3. The apparatus according to claim 1, further comprising storage means for storing said image data set and object displacement measured by said means for motion compensation. 4. A magnetic resonance imaging method, the method comprising: acquiring a motion compensated image data set of a moving object within an imaging region, said motion compensated image data set comprising a plurality of acquisitions and each acquisition comprising one or more slice planes located at a same relative scan location of said moving object, respectively, throughout said plurality of acquisitions;averaging respective one or more slice planes from said plurality of acquisitions to generate a single, averaged image set of said moving object;calculating an actual scan location of all of said one or more slice planes from all of said acquisitions with respect to a static object within the imaging region; andaveraging a selected set of slice planes, based on said calculation, to generate a single, averaged image set of the static object. 5. A non-transitory computer readable medium configured to perform the magnetic resonance imaging method of claim 4. | Koninklijke Philips N.V. | A61B005/05
US08886291 | B2 | 2014-11-11 | Systems and methods for combined ultrasound and electrical impedance imaging | 2009-01-07 | A dual imaging probe 300 for obtaining both ultrasound and electrical impedance data is disclosed along with methods of using the dual imaging probe 300 to interrogate tissue. An electrical impedance imaging overlay 330 is adapted to be positioned on a transducer window 304 of an ultrasound probe 320, and may be integrally formed as part of the ultrasound probe 320 or as a modular adapter for coupling with, and optionally uncoupling from, an ultrasound probe 320 to form the dual imaging probe 300. A method (FIG. 6) of reconstructing composite images using both ultrasound and electrical impedance data is described. Applications for medical diagnosis are described. A particular use for prostate imaging is described. | 1. A method for imaging tissue using a dual imaging probe, comprising: transmitting acoustic signals from an ultrasound transducer through at least some metal electrodes of an array of electrodes into a tissue;receiving reflected acoustic signals through at least one electrode of the electrodes;transducing the reflected acoustic signals into transduced electrical signals;sending the transduced electrical signals as ultrasound data to an electronic system;transmitting electrical current through a first subset of the electrodes of an electrical impedance imaging (EII) overlay into a tissue, the EII overlay comprising a plurality of electrodes;measuring parameters selected from the group consisting of electrical current and electrical voltage at each electrode of a second subset of the electrodes of the EII overlay;sending the measured electrical parameters as EII data to the electronic system; andanalyzing and co-registering the ultrasound data and the EII data to create at least one tomographic image of the tissuewherein the tomographic images are constructed by:extracting regions from the ultrasound data, the ultrasound data obtained simultaneously with the EII data and comprising images;constructing a finite element model of the tissue, the finite element model having impedance parameters at each node of the model;constraining the model with the regions extracted from the ultrasound data; anddetermining impedance at nodes of the model such that simulation of the model matches measured EII data. 2. The method of claim 1, further comprising displaying the at least one tomographic image on an output device. 3. The method of claim 1, further comprising evaluating the one or more tomographic images to make a diagnosis of damage or disease to the tissue. 4. The method of claim 1, wherein the acoustic signals have a frequency between 6 MHz and 30 MHz. 5. The method of claim 4, wherein the acoustic signals have a frequency between 6 MHz and 15 MHz. 6. The method of claim 4, wherein the electrical current is less than 10 mA. 7. The method of claim 6, wherein the one or more tomographic images are three-dimensional. 8. The method of claim 1, wherein a false-color image is constructed having edges extracted from the ultrasound data and colors representing impedance at nodes of the model corresponding to the regions extracted from the ultrasound data. 9. The method of claim 1, wherein at least one electrode of the EII overlay is located over an ultrasound imaging window of the probe. 10. A method for imaging tissue using a dual imaging probe, comprising: transmitting acoustic signals from a transducer of the probe through at least one metal electrode of a group of electrical impedance imaging (EII) electrodes of the probe;receiving reflected acoustic signals through at least one electrode of a group of electrical impedance imaging (EII) electrodes of the probe;transducing the reflected acoustic signals into transduced electrical signals;sending the... | The Trustees of Dartmouth College | A61B005/05 A61B008/08 A61B008/12 A61B005/053 A61B008/00
US08417355 | B2 | 2013-04-09 | Cardiac force sensor and methods of use | 2010-11-05 | A method and apparatus permit sensing one or more forces exerted by one or more portions of a heart. A force transducer and displacement sensor are disclosed. A movement of one or more portions of a heart can be translated into one or more signals indicative of force. These signals can be used to provide information such as to diagnose or treat one or more conditions. | 1. A method comprising: sensing one or more forces exerted by at least a portion of a heart by detecting a compression or expansion of an implanted intracardiac element;obtaining sensed force signal responsive to the one or more forces by measuring a displacement of the compression or expansion of the implanted intracardiac element, wherein the implanted intracardiac element includes a force sensor including: a displacement sensor configured to sense a displacement of the portion of the heart during a heart movement, the displacement sensor including a first portion movable with respect to a second portion along an axial direction and in response to the displacement of the portion of the heart axial movement of the first portion with respect to the second portion being representative of the displacement of the portion of the heart, the displacement sensor being configured to provide an output displacement signal indicative of the displacement sensed; anda contractile structure separate from the displacement sensor, the contractile structure being configured to axially expand or axially contract an axial distance representative of the axial movement of the first portion of the displacement sensor with respect to the second portion, wherein the force sensor is configured to provide the sensed force signal indicative of a force with which at least the portion of the heart moves during the heart movement, the force sensor configured to determine the force as a function of the distance that the contractile structure axially expands or axially contracts; andcommunicating sensed force signal to an electronics unit of an implantable medical device. 2. The method of claim 1, wherein sensing the one or more forces exerted by at least the portion of the heart by detecting the compression or expansion of the implanted intracardiac element includes using the force sensor of the implanted intracardiac element, the force sensor associated with a distal portion of a fully implantable intravascular cardiac lead of the implanted intracardiac element. 3. The method of claim 1, comprising determining a value of the one or more forces exerted by at least the portion of the heart by using a force-displacement relationship of the implanted intracardiac element and the displacement of the compression or expansion. 4. The method of claim 3, comprising determining a pressure exerted on the implanted intracardiac element by at least the portion of the heart using the sensed one or more forces exerted by at least the portion of the heart. 5. The method of claim 1, comprising selecting at least one of a pacing voltage and a pacing location based on the sensed one or more forces exerted by at least the portion of the heart. 6. The method of claim 1, comprising determining a paradoxical heart movement using the sensed one or more forces exerted by at least the portion of the heart. 7. The method of claim 1, comprising: storing an indication of the electrical signal over ... | Cardiac Pacemakers, Inc. | A61N001/05
US09247894 | B2 | 2016-02-02 | Apparatus and method for white-matter-enhancement processing | 2009-12-24 | The present invention provides an apparatus for white-matter-enhancement processing, and a method and a program for white-matter-enhancement processing that makes it possible to analyze distinguishing characteristics of the brain of a living body, such as degree of growth and development, left/right brain dominance, areas of strength and weakness, and personality traits. The apparatus for white-matter-enhancement processing of the present invention has a memory part for inputting and storing contrast images of a brain of a living body; a white-matter-enhanced image-creating part for creating white-matter-enhanced images, based on stored contrast images; a display processing part for displaying the white-matter-enhanced images thus created; and a branching analysis part for analyzing distinguishing characteristics of the brain of a living body, based on any one of or a combination of the characteristics of branches in the white-matter-enhanced images that are created: their number, thickness, length, location, intensity, shape and/or size. | 1. A white-matter-enhancement apparatus comprising: a white-matter-enhanced image-creating computer processor that creates a white-matter-enhanced image based on a contrast image of a brain of a living body received from a magnetic resonance imaging (MRI) apparatus, wherein white matter is enhanced by adjusting the contrast image in such a way that the white matter stands out, a cortex does not stand out, and branches of the white matter are visible;a display device that displays the white-matter-enhanced image; anda branching analysis computer processor that quantitatively measures at least an intensity of selected demarcating regions of the branches of the white matter in the white-matter-enhanced image and analyzes distinguishing characteristics of the brain, based on at least the intensity of the selected demarcating regions, wherein the selected demarcating regions are assigned as regions where specified functions and roles are performed in the brain of the living body;wherein the branching analysis computer processor quantitatively measures the intensity of the selected demarcating regions by counting one of a first number of black voxels and a second number of white voxels in the selected demarcating regions in a left brain and a right brain and analyzes at least left/right brain dominance as a distinguishing characteristic by calculating a Laterality Index LI according to a formula LI=(R−L)/(R+L),wherein L=the first number of black voxels or the second number of white voxels counted in the selected demarcating regions in the left brain, and R=the first number of black voxels or the second number of white voxels counted in the selected demarcating regions in the right brain, andwherein if LI is negative, left brain dominance is strong for the selected demarcating regions, and if LI is positive, right brain dominance is strong for the selected demarcating regions. 2. The white-matter-enhancement apparatus of claim 1, wherein the contrast image is a T1-weighted image, a proton density-weighted image, a T2-weighted image, a FLAIR (Fluid Attenuated Inversion Recovery) image or a diffusion-weighted image, or an image that is a combination thereof. 3. The white-matter-enhancement apparatus of claim 1, wherein the white-matter-enhanced image is created by adjusting a window width and a window level. 4. The white-matter-enhancement apparatus of claim 1, wherein addresses for demarcating regions of the brain are displayed together with the white-matter-enhanced image. 5. The white-matter-enhancement apparatus of claim 1, wherein the branching analysis computer processor further creates and analyzes a plurality of types of classifications based on measurements of at least intensities of the selected demarcating regions of white-matter-enhanced images of brains of living bodies. 6. The white-matter-enhancement apparatus of claim 1, wherein the distinguishing characteristics include a degree of growth and development, left/right brain dominance, ... |  | A61B005/00 A61B005/055 A61B005/107 G01R033/56 G06T005/00
US07885702 | B2 | 2011-02-08 | Segmentation of the airway tree using hyperpolarized noble gases and diffusion weighted magnetic resonance imaging | 2006-04-19 | Diagnostic MR images of the lung are segmented to distinguish the airway tree from lung parenchyma. Diffusion weighted images of the lungs are acquired using hyperpolarized gas MRI and a segmentation image is produced from the apparent diffusion coefficient (ADC) image calculated from the diffusion weighted images. Voxels in the diagnostic MR images having an ADC above a selected amount are segmented as being in the airway tree. | 1. A method for segmenting a diagnostic magnetic resonance image of the lung, the steps comprising: a) filling the lung with a NMR responsive gas;b) acquiring diffusion weighted image (DWI) data from the lung with an MRI system operating at the larmor frequency of the NMR responsive gas;c) reconstructing an apparent diffusion coefficient (ADC) image of the lungs from the acquired DWI data;d) producing a segmentation image from the ADC image in which voxels therein having an ADC value above a selected value are identified as being in the airway tree; ande) using the segmentation image to segment the diagnostic magnetic resonance image. 2. The method as recited in claim 1 in which the NMR responsive gas is a noble gas. 3. The method as recited in claim 2 in which the noble gas is helium. 4. The method as recited in claim 1 which further includes: f) registering the segmentation image with the diagnostic magnetic resonance image before performing step e). 5. The method as recited in claim 1 in which step d) includes growing regions in the segmentation image having corresponding voxels in the ADC image with ADC values above the selected value. 6. The method as recited in claim 1 in which the diffusion weighted image data is acquired in step b) using a two-dimensional pulse sequence that is repeated to acquire data for a series of two-dimensional slices. 7. The method as recited in claim 6 in which step c) includes reconstructing a series of two-dimensional slice images, and calculating the apparent diffusion coefficient (ADC) at each voxel in each of the series of slice images. 8. The method as recited in claim 7 in which step d) includes; d)i) locating a voxel in one of the two-dimensional ADC images having a value that is above a first ADC threshold value;d)ii) growing a region in a corresponding segmentation image that indicates airway tree by locating contiguous voxels in the two-dimensional ADC image that have values above a second ADC threshold value. 9. The method as recited in claim 8 in which the NMR responsive gas is helium and the first ADC threshold value is 0.4 cm2/s and the second ADC threshold value is 0.3 cm2/s. 10. The method as recited in claim 8 in which steps d)i) and d)ii) are repeated for each of the other two-dimensional ADC images having a voxel value above the first ADC threshold value. 11. The method as recited in claim 1 in which the segmentation image indicates at each of its voxels whether the voxel location contains airway or lung tissue. 12. The method as recited in claim 11 in which step e) is performed by multiplying voxel values in the segmentation image by corresponding voxel values in the diagnostic magnetic resonance image. 13. The method as recited in claim 1 in which step b) includes: b)i) acquiring a first image data set using a pulse sequence that includes a diffusion weighted gradient; andb)ii) acquiring a second image data set using a pulse sequence with substantially no diffusion weighting gradi... | Wisconsin Alumni Research Foundation | A61B005/05
US07867477 | B2 | 2011-01-11 | Methods for in vivo evaluation of pulmonary physiology and/or function using NMR signals of polarized | 2007-01-16 | In certain embodiments, methods of the present invention obtain dynamic data sets of an NMR spectroscopy signal of polarized 129Xe in a selected structure, environment, or system. The signal data can be used to evaluate: (a) the physiology of a membrane or tissue; (b) the operational condition or function of a body system or portion thereof (when at rest or under stimulation); and/or (c) the efficacy of a therapeutic treatment used to treat a diagnosed disorder, disease, or condition. Thus, the present invention provides methods for screening and/or diagnosing a respiratory, cardiopulmonary disorder or disease such as chronic heart failure, and/or methods for monitoring the efficacy of therapeutics administered to subject to treat the disorder or disease. | 1. An in vivo method for evaluating whether a subject has a respiratory disorder, or a cardiopulmonary disorder, comprising: delivering polarized 129Xe in vivo to a subject such that the polarized 129Xe travels across the alveolar-capillary membrane to be taken up in the blood across the membrane, the polarized gas in the blood having a polarized 129Xe NMR chemical shift signal frequency;destroying the polarization of the polarized 129Xe in the blood and the membrane;obtaining an NMR spectroscopic signal of the polarized gas in the subject over time at the blood and/or tissue resonant frequency to generate at least one dynamic data set to generate at least one NMR 129Xe uptake curve at least one chemical shift frequency of interest of signal strength values over time; andevaluating the NMR derived dynamic data set to assess whether the subject has a respiratory or cardiopulmonary disorder. 2. A method according to claim 1, wherein the subject is evaluated for chronic heart failure. 3. A method according to claim 1, further comprising: calculating a time constant associated with the time it takes the polarized gas to travel across the membrane and then enter the blood after said destroying step; anddetermining a diffusion thickness based on data provided by said obtaining and calculating steps. 4. A method according to claim 3, wherein the step of obtaining comprises obtaining a plurality of signal data points over a time which is greater than about twice the time constant. 5. A method according to claim 3, wherein said obtaining step is carried out when the subject is at rest and then repeated while the subject is exposed to conditions of actual or chemically induced exercise, and wherein said method further comprises comparing the time constants associate therewith to thereby assess the function of the alveolar-capillary membrane. 6. A method according to claim 4, further comprising calculating asymptotes and intercepts of the NMR 129Xe signal uptake curves from the NMR data and determining the membrane thickness, wherein the membranes have a thickness in the range of between about 1 micron to about 100 microns. 7. A method according to claim 1, wherein the obtaining step is carried out a plurality of times, including at least once while the subject is at rest and at least once when the subject is under or just after actual or simulated physical activity when the heart rate is elevated. 8. A method according to claim 1, wherein the obtaining step is performed after a therapeutic agent is administered to the subject to evaluate the efficacy in treating the disorder or to evaluate its impact on the thickness of the alveolar-capillary membrane. 9. A method according to claim 8, wherein the obtaining step is carried out both before and after the administration of the therapeutic to the subject. 10. A method according to claim 1, wherein the evaluating step comprises evaluating at least one of thickness of the alveolar-capillary membrane, pe... | Medi-Physics Inc. | A61B005/055
US07878984 | B2 | 2011-02-01 | Medical device for navigation through anatomy and method of making same | 2003-07-25 | Medical devices for navigation through anatomy, including guidewires, which may have a core wire, a slotted tubular member, or both. Embodiments may have coils, including non-circular cross-section edge-wound marker coils, extended coil tips, and soldered or glued mesial joint coils. Core wires may have a step, ridge, or taper at the joints to the tubular member, and may be flattened at the distal tip. Radiopaque material may be located inside the tubular member, and the distal tip may be heat treated to make it shapeable. Additional tubular members or coils may be used concentrically or in line and may end in depth radunequal in depth opacity, reduce friction, or reduce material or manufacturing cod in groups, such as groups of three, and may be equally spaced around the axis or offset to provide a steerable or compressible tip. | 1. A guidewire, comprising: an elongate core wire having a distal portion;a metallic tubular member disposed over the distal portion, the tubular member having a proximal end and a distal end;wherein the tubular member has a plurality of slots defined therein;an edge-wound coil disposed between the core wire and the tubular member, the coil having a distal end that is disposed adjacent the distal end of the tubular member and a proximal end disposed distally of the proximal end of the tubular member;wherein the edge-wound coil is formed from a wire having a greater dimension in the radial direction than in the axial direction; anda material disposed between the core wire and the tubular member, wherein the material fills or partially fills at least some of the slots. 2. The guidewire of claim 1, wherein the core wire includes stainless steel. 3. The guidewire of claim 1, wherein the tubular member includes nickel-titanium alloy. 4. The guidewire of claim 1, wherein the tubular member includes a super-elastic nickel-titanium alloy. 5. The guidewire of claim 1, wherein the edge-wound coil includes a radiopaque material. 6. The guidewire of claim 1, wherein the edge-wound coil is attached to the core wire, the tubular member, or both with an adhesive at the proximal end of the edge-wound coil. 7. The guidewire of claim 1, wherein the edge-wound coil is attached to the core wire, the tubular member, or both with an adhesive at the distal end of the edge-wound coil. 8. The guidewire of claim 1, wherein the edge-wound coil is configured to increase the radiopacity of the guidewire. 9. The guidewire of claim 1, wherein the edge-wound coil is configured to have a minimal impact on the flexibility of the guidewire. 10. The guidewire of claim 1, wherein the edge-wound coil is configured to improve torsional stiffness, torsional strength, or both of the guidewire. 11. The guidewire of claim 1, wherein a space is defined within the tubular member between the proximal end of the edge-wound coil and the proximal end of the tubular member, and wherein the space is substantially free of any other structures of the guidewire. 12. A guidewire, comprising: an elongate core wire having a distal portion;a metallic tubular member disposed over the distal portion, the tubular member having a proximal end and a distal end;wherein a plurality of slots are formed in the tubular member;a radiopaque coil disposed between the core wire and the tubular member, the coil having a distal end that is disposed at the distal end of the tubular member and a proximal end disposed distally of the proximal end of the tubular member;wherein the guidewire is free of another coil between the core wire and the tubular member other than the radiopaque coil; andwherein the radiopaque coil is an edge-wound coil that is configured to increase the radiopacity of the guidewire while have a minimal impact on the flexibility by being formed from a wire having a greater dimension in the ... | Boston Scientific Scimed, Inc. | A61B005/00 A61M025/00
US07887491 | B2 | 2011-02-15 | Impedance based device for non-invasive measurement of blood pressure and ankle-brachial index | 2007-10-19 | A device for non-invasive measurement of blood pressure includes a blood pressure cuff, a plethysmographic electrode for acquiring an impedance plethysmogram distal to the cuff and a processing device to inflate and deflate the cuff generate the impedance plethysmogram and to determine the systolic and diastolic blood pressures. The device can be used to measure systolic or diastolic blood pressure or both. It can also be used to take ankle-brachial measurements. An autocorrelation technique can be used to correct noise. | 1. A method for non-invasively measuring blood pressure of a patient, the method comprising: (a) attaching a pressure cuff to an extremity of the patient;(b) attaching a plethysmographic device to the extremity at a distal location relative to the pressure cuff:(c) inflating and deflating the cuff;(d) taking a series of waveforms representing blood pulsations by using the plethysmographic device;(e) deriving, using a processor, a mathematical parameter from a combination of successive ones of the waveforms taken in step (d) and measuring the blood pressure from the mathematical parameter; and(f) outputting the blood pressure measured in step (e);wherein step (e) comprises multiplying corresponding points in groups of the waveforms to produce product waveforms and measuring the blood pressure from the product waveforms. 2. The method of claim 1, wherein the groups of the waveforms are groups of even numbers of the waveforms. 3. The method of claim 2, wherein the groups of the waveforms are groups of at least four of the waveforms. 4. A method for non-invasively measuring blood pressure of a patient, the method comprising: (a) attaching a pressure cuff to an extremity of the patient;(b) attaching a plethysmographic device to the extremity at a distal location relative to the pressure cuff:(c) inflating and deflating the cuff;(d) taking a series of waveforms representing blood pulsations by using the plethysmographic device;(e) deriving, using a processor, a mathematical parameter from a combination of successive ones of the waveforms taken in step (d) and measuring the blood pressure from the mathematical parameter; and(f) outputting the blood pressure measured in step (e);wherein the blood pressure comprises systolic blood pressure; andwherein step (e) is performed twice, once with a first deflation rate and once with a second deflation rate which is slower than the first deflation rate. 5. A method for non-invasively measuring blood pressure of a patient, the method comprising: (a) attaching a pressure cuff to an extremity of the patient;(b) attaching a plethysmographic device to the extremity at a distal location relative to the pressure cuff:(c) inflating and deflating the cuff;(d) taking a series of waveforms representing blood pulsations by using the plethysmographic device;(e) deriving, using a processor, a mathematical parameter from a combination of successive ones of the waveforms taken in step (d) and measuring the blood pressure from the mathematical parameter; and(f) outputting the blood pressure measured in step (e);wherein the blood pressure comprises diastolic blood pressure; andwherein step (e) comprises:(i) performing a point-by-point multiplication on the waveforms to obtain multiplied waveforms: and(ii) determining a maximum of the multiplied waveforms. 6. A method for non-invasively measuring blood pressure of a patient, the method comprising: (a) attaching a pressure cuff to an extremity of the patient;(b) attaching a ple... | Smithmarks, Inc. | A61B005/02
US07874030 | B2 | 2011-01-25 | Device for transferring a cradle for use with a medical imaging equipment | 2007-08-17 | Embodiments of the present invention may provide a cradle transferring device, which moves and withdraws a cradle to and from an examination table of an MRI apparatus in a PET-MRI hybrid system. The cradle transferring device includes: a cradle, on which a subject lies; a shuttle table supporting the cradle and moving and withdrawing the cradle to and from the examination table of the MRI apparatus; and a cradle receiving member provided on the examination table of the MRI apparatus. The shuttle table is provided with a transfer member for holding and transferring the cradle and a transfer member driving means for reciprocating the transfer member along the shuttle table. The cradle transferring device is further provided with a cradle locking means, which locks the cradle to the examination table after the cradle is moved to the examination table. The cradle transferring device moves the cradle to the examination table of the MRI apparatus and accurately withdraws the cradle from the examination table of the MRI apparatus after the examination in the MRI apparatus. | 1. A device for transferring a cradle for use with a medical imaging equipment, comprising: a cradle having a first guide member lengthwisely disposed on a lower side thereof, wherein a subject lies on an upper side thereof;a shuttle table including: a transfer frame for supporting the lower side of the cradle; a transfer member for transferring the cradle, wherein the transfer member is mounted to the transfer frame so as to be movable along the transfer frame; and a transfer member driving means for reciprocating the transfer member between a first position where the transfer member is situated at one end of the transfer frame and a second position where the transfer member is situated at the other end of the transfer frame, wherein the shuttle table is positioned so that one end of an examination table of the medical imaging equipment and the other end of the transfer frame are opposed to each other; anda cradle receiving member located on the examination table of the medical imaging equipment for supporting the cradle transferred from the shuttle table, wherein the cradle receiving member has a second guide member lengthwisely disposed on an upper side thereof, and wherein the second guide member is configured to contact the first guide member, wherein the device further includes a cradle locking means, and wherein the cradle locking means includes: a first locking member movably mounted to the cradle; and a second locking member fixed to the cradle receiving member, wherein the first locking member is engaged and released to and from the second locking member. 2. The device of claim 1, wherein the first locking member comprises a shaft formed with a male thread at one end thereof and the second locking member comprises a nut formed with a female thread therein. 3. The device of claim 2, wherein the cradle locking means further includes: a second motor mounted to the shuttle table and producing normal and reverse rotation; and a transmission member for transmitting the rotation of the second motor to the first locking member, wherein the transmission member comprises a first transmission member disposed on the shuttle table in close proximity of the one end of the transfer frame and a second transmission member disposed on the lower side of the cradle as joined together with the first locking member, andwherein when the transfer member is situated at the second position, the first transmission member and the second transmission member are connected to each other and the rotation of the second motor is transmitted to the rotation of the first locking member. 4. The device of claim 3, wherein the device further comprises signal generating means for generating a first signal for stopping the second motor and a second signal for actuating the second motor. 5. The device of claim 4, wherein the signal generating means includes a switch block disposed at the first locking member and a limit switch disposed at the one end of the transfer frame, ... | Gachon University of Medicine & Science Industry—Academic Cooperation Foundation | A61B006/04 A61B005/055 A61G007/08
US07875261 | B2 | 2011-01-25 | Fluorinated resorufin compounds and their application | 2008-01-16 | The invention provides novel fluorinated resorufin compounds that are of use in a variety of assay formats. Also provided are methods of using the compounds and kits that include a compound of the invention and instructions detailing the use of the compound in one or more assay formats. | 1. A method for determining the presence or absence of peroxide in a sample, wherein the method comprises: a) contacting the sample with a fluorogenic compound to prepare a labeled sample, wherein the fluorogenic compound has the formula;wherein A is OR8 or NR9R10;E is OR8 or NR9R10 wherein R8 is hydrogen, substituted alkyl, unsubstituted alkyl, substituted heteroalkyl, unsubstituted heteroalkyl, substituted heterocycloalkyl, unsubstituted heterocycloalkyl, substituted carboxyalkyl, unsubstituted carboxyalkyl, substituted sulfoalkyl, unsubstituted sulfoalkyl, substituted acyl, unsubstituted acyl, substituted haloalkyl, unsubstituted haloalkyl, substituted alkoxy, unsubstituted alkoxy, or a reactive group;R9 is hydrogen, substituted alkyl, unsubstituted alkyl, substituted heteroalkyl, unsubstituted heteroalkyl, substituted heterocycloalkyl, unsubstituted heterocycloalkyl, substituted carboxyalkyl, unsubstituted carboxyalkyl, substituted sulfoalkyl, unsubstituted sulfoalkyl, substituted acyl, unsubstituted acyl, substituted haloalkyl, unsubstituted haloalkyl, substituted alkoxy, unsubstituted alkoxy, or a reactive group;R10 is hydrogen, substituted alkyl, unsubstituted alkyl, substituted heteroalkyl, unsubstituted heteroalkyl, substituted heterocycloalkyl, unsubstituted heterocycloalkyl, substituted carboxyalkyl, unsubstituted carboxyalkyl, substituted sulfoalkyl, unsubstituted sulfoalkyl, substituted acyl, unsubstituted acyl, substituted haloalkyl, unsubstituted haloalkyl, substituted alkoxy, unsubstituted alkoxy, or a reactive group; ora member independently selected from R9 in combination with R10;R9 in combination with R3;R9 in combination with R6;R10 in combination with R4; andR10 combination with R5 together with the atoms to which they are joined, form a ring which is a 5-, 6- or 7-membered heterocycloalkyl, a substituted 5-, 6- or 7-membered heterocycloalkyl, a 5-, 6- or 7-membered heteroaryl, or a substituted 5-, 6- or 7-membered heteroaryl;R1 is substituted alkyl, unsubstituted alkyl, substituted heteroalkyl, unsubstituted heteroalkyl, substituted aryl, unsubstituted aryl, substituted heteroaryl, unsubstituted heteroaryl, substituted heterocycloalkyl, unsubstituted heterocycloalkyl, OR8 or NR9R10;X1 is oxygen or sulfur;X is oxygen or sulfur;R2 is hydrogen, halogen, substituted alkyl, unsubstituted alkyl, substituted alkoxy, unsubstituted alkoxy, substituted alkylthio, unsubstituted alkylthio, substituted aryl, unsubstituted aryl, substituted heteroaryl, unsubstituted heteroaryl, sulfo, nitro, carboxyl, hydroxyl, or a reactive group;R3 is hydrogen, halogen, substituted alkyl, unsubstituted alkyl, substituted alkoxy, unsubstituted alkoxy, substituted alkylthio, unsubstituted alkylthio, substituted aryl, unsubstituted aryl, substituted heteroaryl, unsubstituted heteroaryl, sulfo, nitro, carboxyl, hydroxyl, or a reactive group;R4 is hydrogen, halogen, substituted alkyl, unsubstituted alkyl, substituted alkoxy, unsubstituted alkoxy, substitute... | Life Technologies Corporation | A61B010/00 A61B005/00
US20110040204 | A1 | 2011-02-17 | NEUROLOGICAL STIMULATION AND ANALYSIS | 2007-10-19 | In one embodiment, a method for neurological interrogation and response analysis is provided. The method includes inducing, at the device, a non-optical excitation for a user. The device comprises a wearable device being worn by the user such that the non-optical excitation can be sensed. A response of a voluntary conscious movement is detected from the user. An outcome is analyzed based on the response and the non-optical excitation. An action can be performed in response to the analysis. | 1. A method for neurological interactive analysis, the method comprising: inducing, at the device, a non-optical excitation for a user, wherein the device comprises a wearable device being worn by the user such that the non-optical excitation can be sensed;detecting a response of a movement from the user;analyzing an outcome based on the response and the non-optical excitation; andperforming an action in response to the analysis. 2. The method of claim 1, wherein the non-optical excitation comprises a minimally obtrusive neurological test that includes a haptic excitation that requires the response from the user. 3. The method of claim 1, wherein the outcome from the user comprises voluntary reaction from the user. 4. The method of claim 1, wherein performing the action comprises triggering a notice that the outcome is not passed; andsending the notice to a computing device configured to process failures of the test. 5. The method of claim 1, wherein performing the action comprises storing information for later analysis. 6. The method of claim 1, wherein analyzing the outcome comprises comparing the outcome to a threshold to determine if the outcome passes a test. 7. The method of claim 6, further comprising triggering an alarm based on the determination if the outcome does not pass the test. 8. The method of claim 1, wherein the non-optical excitation comprises an excitation of the user's peripheral nervous system, wherein the response comprises a response from the user's central nervous system. 9. The method of claim 1, wherein the response does not require a selection of an input on the device. 10. The method of claim 1, wherein analyzing the outcome comprises analyzing response time or a quality of the response of the movement from the user. 11. An apparatus configured to perform a neurological interactive analysis, the apparatus comprising: one or more processors; andlogic encoded in one or more tangible media for execution by the one or more processors and when executed operable to:induce a non-optical excitation for a user, wherein the apparatus is being worn by the user such that the non-optical excitation can be sensed;detect a response of a movement from the user;analyze an outcome based on the response and the non-optical excitation; andperform an action in response to the analysis. 12. The apparatus of claim 11, wherein the non-optical excitation comprises a minimally obtrusive neurological test that includes a haptic excitation that requires the response from the user. 13. The apparatus of claim 11, wherein the outcome from the user comprises voluntary reaction from the user. 14. The apparatus of claim 11, wherein the logic operable to perform the action comprises logic further operable to: trigger a notice that the outcome is not passed; andsend the notice to a computing device configured to process failures of the test. 15. The apparatus of claim 11, wherein the logic operable to perform the action comprises logic fur... | The Regents of The University of California | A61B005/16 A61B005/11
US08568328 | B2 | 2013-10-29 | Electronic blood pressure gauge for measuring blood pressure by using volume compensation method | 2009-03-13 | An electronic sphygmomanometer, which measures blood pressure in accordance with the volume compensation method, detects a cuff pressure inside a cuff attached to a measurement site of the blood pressure. An arterial volume detection circuit detects an arterial volume signal of the measurement site. A drive control unit, after setting the cuff pressure to an initial cuff pressure, servo-controls a cuff pressure adjustment unit so that a volume of an artery becomes constant, based on the detected arterial volume signal. While the servo control is being performed, an arterial volume change amount is detected based on the detected arterial volume signal. In the servo control, when it is detected that the change amount in the arterial volume is minimal, a volume change elimination rate calculating unit calculates a volume change elimination rate (an amplitude of the arterial volume signal during the control/an amplitude of the arterial volume signal before the control) to output the same as a barometer of arteriosclerosis. | 1. An electronic sphygmomanometer to measure blood pressure in accordance with the volume compensation method, the electronic sphygmomanometer comprising: a cuff adapted for attaching to a measurement site of the blood pressure, wherein the measurement site is selected from the group consisting of: a site from a wrist to a base of an arm, and a site from an ankle to a base of a leg;a pressure detector to detect a cuff pressure representing a pressure inside the cuff;a volume detector provided in the cuff and intended to detect an arterial volume signal indicating a volume of an artery of the measurement site;a cuff pressure adjustment unit to adjust the cuff pressure by increasing and decreasing the pressure; anda control unit comprising: a servo control unit to servo-control the cuff pressure adjustment unit so that the volume of the artery becomes constant, based on the detected arterial volume signal after setting the cuff pressure to the initial cuff pressure;a volume change detector to detect an amount of change in the volume of the artery prior to constant volume, based on the detected arterial volume signal, while the control by the servo control unit is being performed,a first control unit comprising: a control target value detector to detect a control target value by controlling the cuff pressure adjustment unit until the cuff pressure becomes a specific pressure at which the volume change detector detects that the amount of volume change in the volume of the artery is maximum, wherein the control target value is a maximum value of the amount of change in the volume of the artery; anda cuff pressure setting unit to set the cuff pressure to an initial cuff pressure representing the specific pressure value;an amplitude ratio detector to detect a ratio between an amplitude of the arterial volume signal detected when the cuff pressure is set to the initial cuff pressure representing the specific cuff pressure, that is, when the volume change detector detects that the amount of change in the volume of the artery is maximum, and an amplitude of the arterial volume signal detected when the volume change detector detects that the amount of change in the volume of the artery is minimal; andan output unit to output the ratio of the amplitude detected by the amplitude ratio detector as a barometer of a degree of arteriosclerosis. 2. The electronic sphygmomanometer according to claim 1, wherein the servo control unit determines the arterial volume when the maximum amplitude of the arterial volume signal is detected to be a target value of the servo control, andbased on a difference between the arterial volume indicated by the detected arterial volume signal and the target value, a servo gain is adjusted so that the amount of change in the volume of the artery detected by the volume change detector becomes minimal, by which the cuff pressure adjustment unit is servo-controlled by the servo control unit. 3. The electronic sphygmomanometer according ... | Omron Healthcare Co., Ltd. | A61B005/02
US20110004113 | A1 | 2011-01-06 | METHOD FOR CONTROLLING AND/OR REGULATING A TRAINING AND/OR REHABILITATION UNIT | 2008-10-17 | The invention relates to a method for controlling and/or regulating a training and/or rehabilitation unit, wherein a) a sensor unit is used in the flow of inspiration and expiration air of a person or an animal using the training and/or rehabilitation unit, b) physiological parameters of ventilation and/or gas exchange of the person or the animal are determined using the respiratory gas composition and/or breath volume measured using the sensor unit, c) one or more maximum performance variables are determined on the basis of the determined parameters under submaximal loading, using a regression function and/or by limit loading to the maximum performance capability, and d) a resistance or brake arrangement of the training and/or rehabilitation unit is controlled and/or regulated as a function of at least one of the determined maximum performance variables. | 1. A method for controlling and/or regulating a training and/or rehabilitation unit, wherein a) a sensor unit is arranged in the flow of inspired and expired air of a person or an animal which uses the training and/or rehabilitation unit,b) the respiratory gas composition and/or the breath volume measured by the sensor unit is/are used to determine physiological parameters of ventilation and/or gas exchange of the person or animal,c) one or more maximum performance characteristics is/are determined on the basis of the parameters which have been determined during sub-maximal exercise, with the aid of a regression function, and/orduring exhaustive exercise until the performance maximum is reached, andd) a resistance or brake arrangement of the training and/or rehabilitation unit is controlled and/or regulated as a function of at least one of the maximum performance characteristic(s) which have been determined. 2. A method according to claim 1, characterized in that O2 uptake (Vo2), CO2 output (Vco2), and/or parameters derived therefrom, namely the respiratory anaerobic threshold (AT), the respiratory quotient (RQ), and/or the oxygen pulse (O2Puls), are determined as gas exchange parameters. 3. A method according to claim 1, characterized in that the tidal volume (VT), the respiratory frequency (fR), and the minute ventilation (VE), and/or the ventilatory equivalent ratio for oxygen (VE/Vo2) which is derived therefrom are determined as ventilation parameters. 4. A method according to claim 1, characterized in that the maximal oxygen uptake (Vo2max) is determined as the maximum performance characteristic. 5. A method according to claim 1, characterized in that the resistance or brake arrangement of the training and/or rehabilitation unit is controlled and/or regulated in such a manner that the O2 uptake (Vo2) of the person or animal is adjusted to a predefinable partial value of the maximal oxygen uptake (Vo2max). 6. A method according to claim 1, characterized in that the resistance or brake arrangement of the training and/or rehabilitation unit is controlled and/or regulated in such a manner that the O2 uptake (Vo2) is maintained at a constant value of between 10% and 100% of the maximal oxygen uptake (Vo2max) during exercise. 7. A method according to claim 1, characterized in that the sensor unit determines the oxygen concentration and/or determines the carbon dioxide concentration with the aid of one or more liquid electrolyte sensor(s). 8. A method according to claim 1, characterized in that the sensor unit determines the oxygen concentration with the aid of a heatable electrochemical solid electrolyte sensor, and/or determines the carbon dioxide concentration with the aid of another heatable electrochemical solid electrolyte sensor,and the heating power of heating elements of the sensors is controlled as a function of the breath volume of the person with the aid of a micro-controller in a sensor control unit in order to maintain constant... | - | A61B005/083
US20110034823 | A1 | 2011-02-10 | Lumen diameter and stent apposition sensing | 2010-09-13 | A stent balloon is provided with two conductive rings, created by a thin metallized coating deposited directly on the balloon, adjacent to the ends of the stent. The impedance between those rings and the body of the patient is measured at different AC frequencies. As the balloon approaches the vessel wall the impedance increases rapidly. Once the balloon forms full contact with vessel wall the impedance increases slowly. The changing impedance provides a guide for optimal apposition of the stent.;The same conductive rings can also detect stent slippage and stent position relative to the balloon. With the addition of an extra conductive pad and wire, stent spring-back can be measured and corrected for. | 1. A method of expanding a stent inside a body lumen by using a stent balloon having at least one electrode outside the area covered by the stent, the method comprising of monitoring the electrical impedance between said electrode and the body, and determining the apposition of said stent inside said lumen by said impedance. 2. A method of expanding a stent inside a body lumen by using a stent balloon having at least one electrode outside the area covered by the stent, the method comprising of monitoring the electrical impedance between said electrode and the body, and determining the apposition of said stent inside said lumen by the rate of change of said impedance. 3. A method of positioning a stent balloon inside a body lumen relative to an existing stent, comprising: adding at least one electrode to the balloon outside the area covered by the stent, and monitoring the electrical impedance between said electrode and the body. 4. A method as in claim 1 wherein each electrodes is made of a metallic ring deposited on the balloon and attached to a thin insulated electrical wire. 5. A method as in claim 2 wherein each electrodes is made of a metallic ring deposited on the balloon and attached to a thin insulated electrical wire. 6. A method as in claim 3 wherein each electrodes is made of a metallic ring deposited on the balloon and attached to a thin insulated electrical wire. 7. A method as in claim 1 wherein the electrical impedance is measured at multiple frequencies. 8. A method as in claim 3 wherein the electrical impedance is measured at multiple frequencies. 9. A method as in claim 1 wherein said balloon has two electrodes placed adjacent to the ends of said stent. 10. A method as in claim 2 wherein said balloon has two electrodes placed adjacent to the ends of said stent. 11. A method as in claim 3 wherein said balloon has two electrodes placed adjacent to the ends of said stent. 12. A method as in claim 1 wherein the electrical impedance is measured at a frequency of between 10 KHz and 10 MHz. 13. A method as in claim 1 wherein said balloon is removed after expanding the stent and a second balloon is inserted, said second balloon having at least one electrode used to locate said second balloon relative to stent and for stent apposition, based on monitoring electrical impedance between said electrode and body. 14. A method as in claim 3 wherein said stent balloon is used to accurately place a plurality of stents relative to each other. 15. A method as in claim 1 wherein said balloon is also used to measure stent spring-back. 16. A method as in claim 1 used for the automatic deployment of stents. 17. A method as in claim 2 used for the automatic deployment of stents. 18. A method as in claim 2 wherein a rapid increase in electrical impedance indicates correct stent apposition. | - | A61B005/05 A61F002/82
US07878975 | B2 | 2011-02-01 | Metabolic monitoring, a method and apparatus for indicating a health-related condition of a subject | 2006-08-25 | An apparatus for indicating a health-related condition of a subject has an input interface for receiving a sequence of samples of a first biological quantity derived by a first measurement method, the first measurement method being an invasive measurement and having a first impact on the subject, and for receiving a sequence of samples of a second biological quantity derived by a second measurement method, the second measurement method being a non-invasive measurement and having a second impact on the subject, wherein the first biological quantity gives a more accurate indication of the health-related condition of the subject than the second biological quantity, wherein the first biological quantity and the second biological quantity have a correlation to the health-related condition of the subject, and wherein the second impact is smaller than the first impact; a predictor for providing, for a certain time, for which no sample for the first biological quantity exists, an estimated value of the first biological quantity using samples for the first biological quantity and, as far as available, samples for the second quantity; and an output interface for outputting the estimated value or data derived from the estimated value so that an indication for the health-related condition of the subject is obtained. | 1. An apparatus for indicating a health-related condition of a subject, wherein the health-related condition is a diabetes-related or glucose-related or insulin-related metabolic condition, comprising: an input interface for receiving a raw sequence of samples of a first biological quantity derived by a first measurement method, the first measurement method being an invasive measurement and having a first impact on the subject, and for receiving a raw sequence of samples of a second biological quantity derived by a second measurement method, the second being a non-invasive measurement and having a second impact on the subject, the biological quantities having a useful variation and a non-useful variation;wherein the first biological quantity gives a more accurate indication of the health-related condition of the subject than the second biological quantity, wherein the first biological quantity and the second biological quantity have a correlation to the health-related condition of the subject, and wherein the second impact is smaller than the first impact;a predictor for providing, for a certain time, for which no sample for the first biological quantity exists, an estimated value of the first biological quantity as a predicted sample using samples for the second biological quantity and, as far as available, samples for the first quantity;a filter for filtering a sequence having samples of the first biological quantity and at least one predicted sample, the filtered sequence having a useful variation and a reduced non-useful variation compared to the sequence before filtering, andan output interface for outputting at least an increase indication, a decrease indication or a remain unchanged indication as a trend of the data, the trend being representative to a useful variation of the health-related condition of the subject. 2. The apparatus in accordance with claim 1, in which the first measurement method is a blood or plasma glucose measurement, andin which the second measurement method is a heart rate measurement, a blood pressure measurement or a method for obtaining a product of heart rate and blood pressure. 3. The apparatus in accordance with claim 1, in which the input interface is arranged to receive, as the sequence of samples of the second biological quantity, a first sub-sequence of samples and a second sequence of samples, and wherein the apparatus further includes a combiner for combining the first sub-sequence of samples and the second subsequence of samples to obtain the sequence of samples of the second biological quantity, the combiner being arranged to perform a sample-wise multiplication. 4. The apparatus in accordance with claim 1, in which the first biological quantity is a blood glucose level, a blood lipid level or a blood insulin level of the subject. 5. The apparatus in accordance with claim 1, in which the first biological quantity is a blood glucose level. 6. The apparatus in accordance with claim 1, in which the pr... | Diabetes Tools Sweden AB | A61B005/00
US07894891 | B2 | 2011-02-22 | Diffusion-based magnetic resonance methods for characterizing bone structure | 2006-06-29 | A method of in vitro or in vivo nuclear magnetic resonance and/or magnetic resonance imaging, to determine bone properties by measuring the effects of molecular diffusion inside the bone specimen to derive parameters that are related to the structure of the trabecular bones. The method is a non-invasive probe that provides topological information on trabecular bone without requiring a full high-resolution image of its structure, and is compatible with clinical use. | 1. A method of strength estimation for a trabecular bone by identifying structural parameters and properties of the trabecular bone, the method comprising: a) using a magnetic resonance technique that provides signals for measuring molecular diffusion over two or more diffusion time periods of different duration by monitoring characteristics of a net molecular displacement inside the trabecular bone so as to provide measured data;b) measuring the molecular diffusion at the two or more diffusion time periods and storing the measured data in a processor;c) using the stored data to derive at least one parameter that is related to a structure of the trabecular bone; andd) estimating the strength of the trabecular bone from the derived at least one parameter by comparing the derived at least one parameter with data relating to known properties of trabecular bones. 2. The method according to claim 1, further comprising calculating a bone volume fraction of the trabecular bone and wherein estimating the strength of the trabecular bone includes estimating the strength of the trabecular bone based on one of the at least one parameter and the calculated bone volume fraction, other prerecorded data or some combination thereof. 3. The method according to claim 1, further comprising calculating a bone mineral density of the trabecular bone wherein estimating the strength of the trabecular bone includes estimating the strength of the trabecular bone based on one of the at least one parameter and the calculated bone mineral density, other prerecorded data or some combination thereof. 4. The method according to claim 1, wherein deriving the at least one parameter includes deriving a surface-to-volume ratio of the trabecular bone. 5. The method according to claim 1, wherein measuring includes performing a decay from diffusion in an internal field (DDIF) measurement on the trabecular bone to obtain measured DDIF decay data. 6. The method according to claim 5, further comprising performing a Laplace inversion on the measured DDIF decay data to provide a DDIF spectrum. 7. The method according to claim 1, wherein measuring includes performing a pulsed-field gradient measurement on the trabecular bone to obtain measured apparent diffusion coefficient (ADC) data. 8. The method according to claim 7, further comprising performing a fit on the measured ADC data to obtain the projected surface to volume ratio. 9. The method according to claim 1, wherein measuring includes performing a pulsed-field gradient measurement on the trabecular bone using field gradients along different directions. 10. The method according to claim 9, further comprising performing a fit on the measured ADC data to obtain the surface to volume tensor. 11. The method according to claim 1, wherein the data is from the group consisting of one of trend data determined from repeating steps (a) to (d) and data relating known properties of one or more other similar trabecular bones or trend data ... | Schlumberger Technology Corporation | A61B005/04
US07870011 | B2 | 2011-01-11 | Automated processing of electronic medical data for insurance and disability determinations | 2010-04-26 | Methods and systems are described for automated processing of medical data for insurance and disability determinations. Based on medical conditions claimed by a claimant, medical evidence queries are automatically generated to provide instructions to medical providers for conducting physical exams and laboratory tests and for retrieving medical records. After medical evidence is collected according to the queries, the medical evidence and related rating codes and decisions are displayed to rating personnel in a user-friendly format to assist in making a rating decision. | 1. A computer-implemented method of assisting medical evidence collection for adjudication of a medical disability request, the method comprising: receiving from a claimant at least one claim of a medical condition;selecting, with a processor, from a predetermined set of medical evidence queries, a plurality of claim-type specific medical evidence queries based on and limited by the at least one claim and based on and limited by a disability rating rules collection, the set containing a greater number of queries than the number of queries in the plurality of queries;displaying the generated medical evidence queries in a data collection protocol;receiving medical evidence data responsive to the generated medical evidence queries from an electronic data storage;wherein the generating the plurality of medical evidence queries comprises excluding queries, from the predetermined set of medical evidence queries, that are, based on the disability rating rules collection, not specific to the claim or claim type. 2. The method of claim 1, wherein the medical evidence data comprises responses to the generated medical evidence queries. 3. The method of claim 1, wherein the electronic data storage comprises electronic medical records. 4. The method of claim 1, wherein the receiving medical evidence data comprises receiving medical evidence data through a local area network. 5. The method of claim 1, wherein the receiving medical evidence data comprises receiving medical evidence data over the Internet. 6. The method of claim 1, wherein the receiving medical evidence data comprises receiving medical evidence data over an encrypted Internet connection. 7. The method of claim 1, wherein the receiving medical evidence data further comprises receiving medical evidence data collected by a medical provider. 8. The method of claim 7, wherein the receiving medical evidence data collected by a medical provider further comprises reviewing and extracting data from medical records. 9. The method of claim 7, wherein the receiving medical evidence data collected by a medical provider further comprises taking a medical history of the claimant. 10. The method of claim 9, wherein the medical history comprises at least one of a family history and a social history. 11. The method of claim 7, wherein the receiving medical evidence data collected by a medical provider further comprises receiving data from a physical exam. 12. The method of claim 7, wherein the receiving medical evidence data collected by a medical provider further comprises receiving data representative of laboratory tests. 13. The method of claim 7, wherein the receiving medical evidence data collected by a medical provider further comprises receiving data from radiological study results. 14. The method of claim 7, wherein the receiving medical evidence data collected by a medical provider further comprises receiving data from nuclear medicine laboratory results. 15. The method of claim 1, further ... | QTC Management, Inc. | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07883477 | B2 | 2011-02-08 | Human body backbone measuring/displaying system | 2005-10-26 | A human spinal column measuring and displaying system includes a spinal column measuring apparatus capable of measuring bending or a circumflex (twist) angle of the spinal column of the human body as a detaching amount measuring data from a reference value, an input apparatus for inputting gender, height data of the measurement subject, an image processing apparatus for inputting the detaching amount measuring data from the spinal column measuring apparatus and data from the input apparatus and generating a three-dimensional spinal column image of the measurement subject based on the detaching amount measuring data and the gender, height data of the measurement subject, and a display apparatus for displaying an image data from the image processing apparatus. The system can calculate a measured data of three-dimensional coordinates and a circumflex angle of the spinal column by scanning a predetermined probe along the spinal column, and simulate a shape of the spinal column from the measured data to be able to display a three-dimensional pseudonymous spinal column image by utilizing computer graphics. | 1. A human spinal column measuring and displaying system, comprising: a spinal column measuring apparatus including a measurement bed on which a measurement subject lies in a lying state,a base provided on a side of one side face of the measurement bed,a measuring direction support arm fixed to the base and configured to be movable in a measuring direction constituting a longitudinal or X axis direction of the measurement bed,a parallel support arm fixed to the measuring direction support arm and configured to be movable in a width or Y axis direction orthogonal to the measuring direction,a vertical support arm fixed to the parallel support arm and configured to be movable in a thickness or Z axis direction, anda probe pivotably fixed to a front end of the vertical support arm to detect amounts of detaching from reference positions in the X direction, the Y direction, the Z direction, and a twist angle θ direction centering on the X axis direction by pinching the probe, which is provided at a front end of a scanning arm configured to move in the X axis direction, the Y axis direction, the Z axis direction, and the twist angle θ direction of the spinal column of the measurement subject on the measurement bed, between a second finger and a third finger of a measuring person and three-dimensionally moving front ends of the second and third fingers in a state of moving the front ends of the second and third fingers from a position of the first cervical vertebra or a position of the first thoracic vertebra to a position of the fifth lumber vertebra of the spinal column of the measurement subject to output as detaching amount measuring data;an input apparatus to input a gender and height data of the measurement subject; andan image processing apparatus configured to input the detaching amount measuring data from the spinal column measuring apparatus and the gender and height data of the measurement subject from the input apparatus, and to generate a three-dimensional spinal column image of the measurement subject, based on the detaching amount measuring data and the gender and height data of the measurement subject, to provide the spinal column image to a display apparatus,wherein the image processing apparatus includes converting means for converting the detaching amount measuring data in the X axis direction, the Y axis direction, the Z axis direction, and the θ direction outputted by the spinal column measuring apparatus to predetermined provided data to be stored in a converted data storing unit;a basic diagram database that stores an average size by the gender and height of the measurement subject and a basic shape of the measurement subject with regard to the respective vertebrae constituting the spinal column of a human body;data selecting means for selecting to input the respective vertebrae in correspondence with the gender and height data from the basic diagram database in accordance with the gender and height data of the measurement subject ... | Nihon University MASARU SEIKI Corporation | A61B005/103 A61B005/117 G01C009/00 G01C017/00 G01C019/00
US08755854 | B2 | 2014-06-17 | Methods and apparatus for producing and using lightly filtered photoplethysmograph signals | 2009-07-31 | One or more physiological conditions of a patient can be observed by obtaining a photoplethysmograph (“PPG”) signal from the patient and by only lightly filtering that signal. The light filtering of the PPG may be such as to only remove (for example) high frequency noise from that signal, while leaving in the signal most or all frequency components that are due to physiological events in the patient. In this way, such physiological events can be observed via a visual display of the lightly filtered PPG signal and/or via other signal processing of the lightly filtered PPG signal to automatically extract certain physiological parameters or characteristics from that signal. | 1. Apparatus for monitoring a physiological condition of a patient comprising: photoplethysmograph (“PPG”) circuitry configured to obtain a PPG signal from the patient;filter circuitry configured to lightly filter the PPG signal to produce a lightly filtered PPG signal; andoutput circuitry configured to provide a visible display of a trace of the lightly filtered PPG signal, thereby making it possible for a clinician user of the apparatus to observe at least some patient physiological condition information contained in the lightly filtered PPG signal. 2. The apparatus defined in claim 1 wherein the filter circuitry comprises: a low-pass filter configured to pass components of the PPG signal having frequencies less than a threshold frequency and for blocking components of the PPG signal having frequencies greater than the threshold frequency. 3. The apparatus defined in claim 2 wherein the threshold frequency is approximately 4 Hz. 4. The apparatus defined in claim 1 wherein the filter circuitry comprises: a band-pass filter configured to pass components of the PPG signal having frequencies in a range between upper and lower threshold frequencies and for blocking components of the PPG signal having frequencies that are not in said range. 5. The apparatus defined in claim 4 wherein the upper threshold frequency is approximately 4 Hz. 6. The apparatus defined in claim 4 wherein the lower threshold frequency is approximately 0.005 Hz. 7. The apparatus defined in claim 1 wherein the filter circuitry has a frequency threshold at which the filter circuitry switches from passing frequency components of the PPG signal to blocking frequency components of the PPG signal, and wherein the frequency threshold is controllably variable. 8. The apparatus defined in claim 7 further comprising: processing circuitry configured to analyze the lightly filtered PPG signal to produce a control signal for controlling the frequency threshold. 9. The apparatus defined in claim 1 wherein the output circuitry comprises: processing circuitry configured to analyze the lightly filtered PPG signal to produce an output signal indicative of a physiological parameter extracted from the patient physiological condition information contained in the lightly filtered PPG signal. 10. The apparatus defined in claim 1 further comprising: polysomnogram (“PSG”) circuitry configured to obtain a PSG signal from the patient; and wherein the output circuitry additionally makes it possible for the clinician user to observe at least some additional patient physiological condition information contained in the PSG signal. 11. The apparatus defined in claim 10 wherein the PSG circuitry comprises: SpO2 circuitry configured to measure saturation of the patient's hemoglobin with oxygen. 12. The apparatus defined in claim 10 wherein the PSG circuitry comprises: breathing airflow monitoring circuitry. 13. The apparatus defined in claim 1 further comprising: circuitry configured to extract from... | Nellcor Puritan Bennett Ireland | A61B005/1455 A61B005/145 A61B005/00
US07877126 | B2 | 2011-01-25 | Hat-based oximeter sensor | 2006-07-26 | A method for use and an improved oximeter sensor substrate that is conforming to the shape of the patient's forehead. In one embodiment, the present invention is an oximeter sensor, having a substrate with a shape similar to a shape of at least a portion of a patient's forehead and including a section adapted to substantially fit over a portion of a forehead of a patient; an emitter disposed on the substrate at a position located on the section; and a detector disposed on the substrate at a distance from the emitter. In one embodiment, the substrate includes a hat that holds the emitter and the detector in a spaced-part manner against the patient's forehead. | 1. A sensor comprising: a stocking cap;a substrate disposed on the stocking cap;at least one sensing component disposed on the substrate; andan electrical connector coupled to the at least one sensing component, wherein the electrical connector extends from the substrate through an open portion of the stocking cap, wherein the open portion of the stocking cap is adapted to guide the electrical connector away from a patient's torso and is configured to be proximate to a top of the stocking cap when the stocking cap is applied to the patient. 2. The sensor of claim 1, wherein the substrate is conformable to a forehead-contacting surface of the stocking cap. 3. The sensor of claim 1, wherein the substrate comprises an adhesive layer adapted to attach the substrate to the stocking cap. 4. The sensor of claim 1, wherein the stocking cap comprises a neonatal stocking cap. 5. The sensor of claim 1, wherein the at least one sensing component is adapted to be attached to a forehead-contacting surface of the stocking cap. 6. The sensor of claim 1, wherein the electrical connector is secured by a portion of the stocking cap. 7. The sensor of claim 1, wherein the at least a portion of the electrical connector comprises a cable. 8. The sensor of claim 1, wherein the at least one sensing component comprises at least one of an emitter or a detector. 9. A method of manufacturing a cap-based sensor comprising: providing a stocking cap;providing a substrate disposed on the stocking cap;providing at least one sensing component disposed on the substrate; andproviding an electrical connector coupled to the at least one sensing component, wherein the electrical connector extends from the substrate through an open portion of the stocking cap, wherein the open portion of the stocking cap is adapted to guide the electrical connector away from a patient's torso and is configured to be proximate to a top of the stocking cap when the stocking cap is applied to the patient. 10. The method of claim 9, wherein the substrate is conformable to a forehead-contacting surface of the stocking cap. 11. The method of claim 9, comprising providing an adhesive layer adapted to attach the substrate to the stocking cap. 12. The method of claim 9, wherein the stocking cap comprises a neonatal stocking cap. 13. The method of claim 9, wherein the at least one sensing component is adapted to be attached to a forehead-contacting surface of the stocking cap. 14. The method of claim 9, comprising providing an opening adapted to accommodate the at least one sensing component. 15. The method of claim 14, wherein the opening comprises a hole or a pocket. 16. The method of claim 9, wherein the electrical connector is secured by a portion of the stocking cap. 17. The method of claim 9, wherein the at least a portion of the electrical connector comprises a cable. 18. The method of claim 9, wherein the sensing component comprises at least one of an emitter or a detector. 19. A method ... | Nellcor Puritan Bennett LLC | A61B005/1455
US08639319 | B2 | 2014-01-28 | Watertight ECG monitor and user interface | 2009-03-04 | An ECG monitoring system for ambulatory patients includes a small multi-electrode patch that adhesively attaches to the chest of a patient. A reusable battery-powered ECG monitor clips onto the patch and receives patient electrical signals from the electrodes of the patch. A processor continuously processes received ECG signals and stores the signals in memory in the monitor. Processed ECG signals and cardiac event information are sent wirelessly to a cellphone handset for relay to a monitoring center. The ECG monitor is contained in a watertight sealed case with only electrical contacts on the outside of the case. The electrical contacts electrically couple the ECG monitor to the electrodes of the patch during patient monitoring and to a charger during recharge of the battery. | 1. A cardiac monitoring system for ambulatory patients including an ECG monitor comprising: a nonconductive watertight case;a rechargeable battery located in the case;an ECG signal processor located in the case and having an input for receiving ECG signals from a patient;a memory, located in the case, for storing processed ECG signal information;a wireless transceiver, located in the case and coupled to the ECG signal processor, and operable to wirelessly transmit ECG signal information to a receiver;a user interface, located on the outside of the case and electrically coupled to one or more components inside the case, the user interface comprising only a plurality of electrical contacts; anda power management circuit operable to sense both of a connection of the ECG monitor to an external charging dock andthat the ECG monitor is electrically coupled to receive patient ECG signals,wherein the power management circuit is further operable to switch to a low power consumption state when the monitor is not coupled to receive patient ECG signals and not connected for charging the rechargeable battery, andfurther wherein the power management circuit switches to a high power state responsive to the sensed connection for charging. 2. The cardiac monitoring system of claim 1, further comprising battery charging circuitry located in the case and coupled to the rechargeable battery, wherein the electrical contacts are coupled to the battery charging circuitry and the input of the ECG signal processor. 3. The cardiac monitoring system of claim 2, wherein the charging dock comprises electrical contacts for supplying charging current, wherein the battery charging circuitry contacts of the ECG monitor are coupled to the electrical contacts of the charging dock during recharging of the rechargeable battery. 4. The cardiac monitoring system of claim 1, further comprising an electrode patch having an electrode and an adhesive surface for attachment to the skin of a patient, wherein a one or more of the electrical contacts of the ECG monitor are electrically coupled to the electrode patch during cardiac monitoring of a patient. 5. The cardiac monitoring system of claim 4, wherein the ECG signal processor further includes a plurality of inputs coupled to electrical contacts of the user interface; wherein the electrode patch further includes a plurality of electrodes and a plurality of contacts electrically coupled to the electrodes,wherein the plurality of contacts engage the ECG signal processor input electrical contacts during cardiac monitoring of a patient. 6. The cardiac monitoring system of claim 1, wherein the ECG monitor is always powered to at least the low power consumption state when the battery is charged. 7. The cardiac monitoring system of claim 6, wherein the ECG monitor is powered to a high powered state which is higher than the low power consumption state during operation of the wireless transceiver and during patient ECG monitoring. 8. The c... | Koninklijke Philips N.V. | A61B005/0402
US09186092 | B2 | 2015-11-17 | System, garment and method | 2008-08-22 | The present invention relates to a system and garment that incorporates sensors that can be used for measuring or monitoring pressure or forces in feet, the stumps of limbs of an amputee that are fitted with prosthetic devices, or any other parts of the body that are subject to forces such as the buttock while seated or when external pressure inducing devices are employed, such as for example, pressure bandages. The invention also relates to a method to monitor or diagnose any foot or limb related activity for recreational, sporting, military or medical reasons and is particularly aimed at the treatment of neuropathic or other degenerating conditions. | 1. A system that is wearable by a person to monitor or sense pressure, the system including: a) a flexible substrate in the form of a garment that can be worn by or fitted directly or indirectly to surfaces of the person subject to loads or pressure; andb) at least one pressure sensor fixed to the substrate or movably connected to the substrate, with leads extending, from each pressure sensor, the leads terminating, in contact surfaces located on an external surface of the flexible substrate and spaced a distance from an edge boundary of the flexible substrate. 2. The system according to claim 1, wherein the at least one pressure sensor has a low profile that is equal to or less than 5.0 mm in thickness. 3. The system according to claim 2, wherein the at least one pressure sensor has a low profile that is equal to or less than 0.5 mm in thickness. 4. The system according to claim 1, wherein the at least one pressure sensor is held in position on the substrate using either one or a combination of adhesive materials, hook and loop type fasteners, clasps or other mechanical fasteners that have a low profile. 5. The system according to claim 1, wherein the at least one pressure sensor comprises a material selected from the group consisting of: conductive yarns or threads and metal coated threads or yarns, and wherein the at least one sensor comprises substantially inert electrodes to measure a change in electrical potential difference across the sensor. 6. The system according to claim 1, wherein the at least one pressure sensor comprises a material selected from the group consisting of: conductive yarns or threads and conductive polymers that are spaced apart by a spacing in the range of 2 to 15 mm in a direction transverse to the direction in which the substrate compresses. 7. The system according to claim 1, wherein the at least one pressure sensor has an area substantially transverse to a direction of pressure application that is equal to or less than 400 mm2. 8. The system according to claim 1, wherein the at least one pressure sensor is used for monitoring biomechanical movement of the person's foot and the at least one pressure sensor is located at a heel, metatarsal or arch region of the person's foot. 9. The system according to claim 1, wherein the substrate is a garment in the form of one of socks, stockings, underpants, long johns, a singlet, leggings, inserts, inner socks, inner garments, under garments, tubular sleeves, or garments used to cover a stump or terminated limb of an amputee. 10. The system according to claim 1, wherein the at least one pressure sensor can be used to monitor the level of pressure applied to the person by an external source, and the flexible substrate is in the form of a compression garment. 11. The system according to claim 1, additionally comprising electrical circuitry that receives signals from the contact surfaces located on the substrate that communicates with the at least one pressure sensor, w... | Sensoria, Inc. | A61B005/103 A61B005/01 A61B005/00
US08585607 | B2 | 2013-11-19 | Monitoring, predicting and treating clinical episodes | 2010-11-03 | Apparatus and methods are provided for use with a subject who is undergoing respiration. A motion sensor senses motion of a subject. A breathing pattern analysis unit analyzes components of the sensed motion that result from the subject's respiration. The breathing pattern analysis unit includes double-movement-respiration-cycle-pattern-identification functionality that designates respiration cycles as being double-movement-respiration-cycles (DMRC's) by determining that the cycles define two subcycles. Double-movement-respiration-cycle-event-identification functionality of the breathing pattern analysis unit identifies a DMRC event by detecting that the subject has undergone a plurality of DMRC's. An output is generated that is indicative of the subject having used accessory muscles in breathing, in response to identification of the double-movement-respiration-cycle event. Other embodiments are also described. | 1. Apparatus for use with a subject who is undergoing respiration, the apparatus comprising: a motion sensor configured to sense motion of a subject;a control unit comprising: a breathing pattern analysis unit that is configured to analyze components of the sensed motion that result from the subject's respiration, the breathing pattern analysis unit comprising: double-movement-respiration-cycle-pattern-identification functionality that is configured to designate respiration cycles as being double-movement-respiration-cycles by determining that the cycles define two subcycles; anddouble-movement-respiration-cycle-event-identification functionality configured to identify a double-movement-respiration-cycle event by detecting that the subject has undergone a plurality of double-movement-respiration-cycles; andoutput-generation functionality configured to generate an output that is indicative of the subject having used accessory muscles in breathing, in response to identification of the double-movement-respiration-cycle event; anda user interface configured to display the output,wherein the double-movement-respiration-cycle-pattern-identification functionality is configured to distinguish between a first cycle that defines two subcycles, and a second cycle that does not define two subcycles, even if the first and second cycles have substantially similar amplitudes to one another. 2. The apparatus according to claim 1, wherein the sensor is configured to sense motion of the subject without contacting the subject or clothes the subject is wearing. 3. The apparatus according to claim 1, wherein the sensor is configured to sense motion of the subject without contacting or viewing the subject or clothes the subject is wearing. 4. The apparatus according to claim 1, wherein the double-movement-respiration-cycle-pattern-identification functionality is configured to determine that a cycle defines two subcycles at least partially in response to a slope of a respiration motion signal of the subject. 5. The apparatus according to claim 1, wherein the control unit is configured to predict an onset of a clinical condition of the subject in response to the double-movement-respiration-cycle-event-identification functionality identifying the double-movement-respiration-cycle event. 6. The apparatus according to claim 5, wherein the control unit is configured to predict the onset of the clinical condition by predicting the onset of a clinical condition selected from the group consisting of: elevated heart rate and elevated respiration rate. 7. The apparatus according to claim 5, wherein the control unit is configured to predict the onset of the clinical condition at least 30 minutes before the onset of the clinical condition. 8. Apparatus for use with a subject who is undergoing respiration, the apparatus comprising: a motion sensor configured to sense motion of a subject;a control unit comprising: a breathing pattern analysis unit that is configured to analyz... | EarlySense Ltd. | A61B005/08
US10471027 | B2 | 2019-11-12 | Pharmacology of visual cycle modulators | 2010-07-02 | Described herein are methods for the treatment of an ophthalmic disease or disorder comprising the administration of non-retinoid visual cycle modulators. Ophthalmic diseases and disorders which can be treated according to the methods provided herein include age-related macular degeneration (AMD), Stargardt's disease, glaucoma, and other disorders that adversely affect the retina. | 1. A method of treating Stargardt's macular dystrophy, comprising administering daily a single oral dose of an amount from about 5 mg to about 10 mg of a non-retinoid compound (R)-3-amino-1-(3-(cyclohexylmethoxy)phenyl)propan-1-ol having the formula: 2. The method of claim 1, wherein the normalized electroretinogram response is from about 5% to about 15%, from about 15% to about 25%, from about 25% to about 35%, or from about 35% to about 50%. 3. The method of claim 1, wherein the normalized electroretinogram response is determined after about 12 hours to about 16 hours, after about 16 hours to about 20 hours, after about 20 hours to about 24 hours, after about 24 hours to about 30 hours, after about 30 hours to about 36 hours, after about 36 hours to about 42 hours, post administration of said non-retinoid compound. 4. The method of claim 1, wherein administration of a single dose of the non-retinoid compound results in a greater normalized electroretinogram response on day 1 than on day 2. 5. The method of claim 1, wherein administration of the non-retinoid compound results in a greater therapeutic response on day 2 than on day 1 after the administration of a single dose. 6. The method of claim 5, wherein the therapeutic response is determined by electroretinography. 7. The method of claim 1, wherein administration of the non-retinoid compound results in a normalized electroretinogram response of less than about 50% for a time period of about 4 hours to about 10 hours, about 10 hours to about 16 hours, about 16 hours to about 24 hours, or about 24 hours to about 36 hours after the plasma concentration of said non-retinoid compound has declined to 0.3 Cmax. 8. The method of claim 1, wherein non-retinoid compound is administered in the morning, is administered upon waking from sleep, or is administered upon waking from sleep in the morning. 9. The method of claim 1, wherein the amount administered is about 5 mg, about 7 mg, or about 10 mg of the non-retinoid compound. 10. The method of claim 1, wherein the cone response amplitude of the electroretinogram remains within about 10%, within about 20%, or within about 30% of the pre-treatment amplitude. 11. The method of claim 1, wherein the dosage of non-retinoid compound is adjusted to provide no noticeable deficiency in night vision. 12. The method of claim 1, wherein the method results in no detectable effect on sensitivity in photopic conditions. 13. The method of claim 1, wherein the non-retinoid compound is an in the form of a controlled-release solid dosage form. 14. The method of claim 13, wherein administration of the controlled-release solid dosage form results in a plasma Tmax 12 hours post-administration. 15. The method of claim 1, wherein non-retinoid compound is administered at a time other than the morning. 16. The method of claim 15, wherein non-retinoid compound is administered during lunch or subsequent to lunch. 17. The method of claim 1, wherein the administration ... | ACUCELA, INC. | A61K031/135 A61K031/137 A61K031/166 A61K009/00 A61K009/50 A61K031/19 A61K031/195
US11169010 | B2 | 2021-11-09 | Method for the calibration of an implantable sensor | 2009-07-27 | An implant includes a processor, RF communication circuitry, optical communication circuitry, a power source and a memory, all of which being hermetically sealed within a housing having a transparent window. Sensor readings are transmitted by RF using the RF communication circuitry to a remote reader after receiving interrogation signals from the reader. During calibration of the sensor, corrective coefficients are calculated by comparing actual sensor pressure readings with known pressure readings. The corrective coefficients are transmitted to the memory of the control circuitry using optical communication wherein modulated light is transmitted through the transparent window of the housing to the photo-detector. | 1. A method for calibrating a sensor of an implant before implantation into a body and using the sensor and the implant in the body, wherein said sensor is connected to control circuitry which includes a processor, a radio frequency (RF) communication circuitry, optical communication circuitry including a light-receiving device, a power source and memory, all of which being hermetically sealed within a housing, the sensor is disposed on a base plate, said method comprising the steps of: creating a bore through the base plate;placing the sensor over the bore so that the sensor is exposed to an external atmosphere while being hermetically sealed within the housing;placing said implant into a controlled environment outside of the body;creating an RF communication link between a remote controller and said RF communication circuitry of said control circuitry, whereby the control circuitry converts RF energy from the remote controller to help power at least the processor and RF communication circuitry;adjusting at least one aspect of said controlled environment to a first known value;allowing said sensor to measure said at least one aspect of said controlled environment and generate a first sensor reading;transmitting said first sensor reading from said implant to said remote controller using said RF communication link;comparing said received first sensor reading with said first known value of said at least one aspect of said controlled environment to create a first corrective coefficient;creating an optical communication link between said remote controller and said control circuitry of said implant before implantation into the body;transmitting said first corrective coefficient to said memory of said control circuitry using said optical communication link;allowing said control circuitry of said implant to use said first corrective coefficient to adjust said first sensor reading to align with said first known value for future readings;implanting said sensor and said implant within said body; andtransmitting, after calibration and implantation, RF energy from said remote controller to said implant in said body for activating said sensor, powering at least said processor and said RF communication circuitry of said sensor to read output of said sensor and use the first corrective coefficient to calculate calibrated sensor data, and extracting calibrated sensor data in real time from said sensor implanted in said body. 2. The method for calibrating according to claim 1, wherein said at least one aspect is selected from the list of pressure, temperature, humidity, acidity, oxygen, glucose and flow rate. 3. The method for calibrating according to claim 1, wherein said housing includes a transparent region through which said optical communication may occur. 4. The method for calibrating according to claim 3, wherein said base plate is made from glass. 5. The method for calibrating according to claim 3, wherein said base plate is made from a glass that is ... | INTEGRA LIFESCIENCES SWITZERLAND SÀRL | A61B005/1495 G01D011/24 G01D018/00 A61B005/0215 G01L019/08 A61B005/00 G01L019/14 G01D005/353
US20110004317 | A1 | 2011-01-06 | ORTHOPAEDIC IMPLANTS | 2008-12-17 | A method for making an orthopaedic implant, the method comprising: characterizing at least a portion of a bone corresponding to the bone to be replaced, said corresponding bone being on the contralateral side of the patient; providing a model of the orthopaedic implant based on a mirror image of the contralateral bone; and forming the orthopaedic implant based on the model. | 1-64. (canceled) 65. An implant for replacing a bone in a wrist, hand, ankle or foot of a patient, said implant having a body which is sized and shaped to anatomically replicate the bone being replaced and which is based on a mirror image of a corresponding bone on a contralateral side of the patient. 66. An implant according to claim 65, wherein the implant body includes at least one attachment element for attaching the implant to soft or hard tissue at a site of implantation in use, the attachment element being an opening formed through the implant body for receiving a fixing element. 67. An implant according to claim 65, wherein the implant has an articulating surface which has a topography corresponding to that of a surface against which the implant will articulate or to a surface of a corresponding bone on a contralateral side of the patient. 68. An implant according to claim 65, wherein the bone being replaced is a scaphoid or a lunate bone of a wrist, the implant being surgically insertable at an implantation site adjacent a radius bone, the implant body having an articulating surface for articulation against at least a portion of an articulating surface of the radius bone. 69. An implant according to claim 65, further comprising a component for replacing at least a portion of a surface of a bone articulating against the implant in use, the component comprising a component body and a component surface, the component surface corresponding substantially in topography to the portion of the surface of the articulating bone which the component is replacing, to a corresponding contralateral bone, or to an articulating surface of the implant. 70. An implant according to claim 69, wherein the component body includes at least one component attachment element for attaching the component to the articulating bone, wherein the at least one component attachment element includes at least one opening formed through the component body for receiving a fixing element. 71. An implant according to claim 69, wherein the component is sized and shaped to replace at least a portion of a radius bone of a wrist of the patient. 72. A method for making an implant for replacing a bone in a wrist, hand, ankle or foot of a patient, the method comprising: characterizing at least a portion of a bone corresponding to the bone to be replaced, said corresponding bone being on the contralateral side of the patient;providing a three-dimensional digital model of the implant based on a mirror image of the contralateral bone; andforming the implant based on the three-dimensional digital model. 73. A method according to claim 72, wherein the characterizing of the contralateral bone comprises imaging the portion of the contralateral bone corresponding to the bone to be replaced using Magnetic Resonance Imaging or Computed Tomography. 74. A method according to claim 72, further comprising characterizing a surface topography of an articulating bone intended to articulate wit... | - | A61F002/42 A61B005/05 A61B017/16
US07871385 | B2 | 2011-01-18 | Urine sample collection device | 2009-04-14 | A urine sample collection device comprising a urine receptor having an outlet aperture, a generally elongate tubular member extending from said receptor outlet aperture to an open end, the tubular member having an opening (21) formed in the side thereof, a coupling means (5′) for releasably mounting a urine collection container, the coupling means having a passage extending therethrough which meets said opening whereby urine can flow from the tubular member into a mounted container, and a flow director (24) located at or adjacent said opening and formed to direct urine past the opening. | 1. A urine sample collection device, comprising a urine receptor into which a user urinates;wherein the urine receptor contains a single outlet aperture flaring out to a rim defining a perimeter of an inlet area into which the user urinates;a hollow tubular member having a generally elongate form defining a closed tubular chamber, an inlet end containing an inlet opening communicating with said chamber, and an outlet end opposite said inlet end and containing an outlet opening communicating with said chamber;wherein the tubular member includes an inner surface and a sampling outlet fixture formed in a side wall thereof intermediate the inlet end and the outlet end;wherein the inlet end of the tubular member is connected with the single outlet aperture of the urine receptor so that all urine from the user flows out of the urine receptor and into the tubular member chamber;wherein the sampling outlet fixture defines a passage extending through the side wall of the tubular member and communicating with said chamber;wherein the fixture has an external portion which extends externally out from the tubular member and the external portion contains an external end opening of the passage;wherein the fixture external portion includes a coupling for releasably mounting an open topped urine sample collection container thereto for collecting urine flowing through the passage and out of the external end opening;wherein the fixture has an internal portion which extends internally into the tubular member chamber through the side wall of the tubular member and the internal portion contains an internal end opening of the passage; anda flow diverter formed on the internal portion of the fixture, said flow diverter comprising a surface inclined relative to said inner surface of the closed tubular member to direct smaller urine flows from the urine receptor along the tubular member inner surface past the internal end opening while allowing larger urine flows from the urine receptor within the tubular member to pass through the internal end opening into the passage. 2. A device according to claim 1, wherein said inclined surface extends towards a longitudinal axis of the tubular member. 3. A device according to claim 1, wherein said inclined surface is located upstream of the internal end opening. 4. A device according to claim 3, wherein said inclined surface comprises a wall which extends partially across the tubular member inner surface to an extent corresponding to an upstream edge of said internal end opening. 5. A device according to claim 3, wherein said flow diverter further comprises a projection located downstream of the internal end opening. 6. A device according to claim 1, wherein said flow diverter further comprises a projection shaped to define a channel on one or both sides of the internal end opening. 7. A device according to claim 1, wherein said flow diverter comprises a projection which extends into the tubular member by an amount correspondi... | - | A61B005/00 B65D081/00
US07894676 | B2 | 2011-02-22 | Diagnosis support apparatus and control method therefor | 2010-03-26 | A diagnosis support apparatus includes a storage unit which stores pieces of image feature information and finding sentences as interpretation results in correspondence with each other. This apparatus acquires image feature information of a target area designated on an image to be interpreted, searches the storage unit for image feature information similar to the acquired image feature information, acquires a finding sentence stored in correspondence with the retrieved image feature information from the storage unit, and creates a finding sentence concerning interpretation of the designated target area by changing a description of the acquired finding sentence based on image feature information of the designated target area. | 1. A diagnosis support apparatus comprising: a storage unit configured to store image feature information and a finding sentence as an interpretation result in correspondence with each other;an acquisition unit configured to acquire image feature information of a target area designated on an image to be interpreted;a search unit configured to search said storage unit for image feature information similar to the image feature information acquired by said acquisition unit and acquire a finding sentence stored in correspondence with the retrieved image feature information from said storage unit; anda creation unit configured to created a finding sentence concerning interpretation of the designated target area by changing, based on image feature information of the target area acquired by said acquisition unit, a description of the finding sentence acquired by said search unit,wherein said creation unit comprises:an extraction unit configured to extract a description corresponding to each feature item of image feature information by performing syntax analysis of the finding sentence acquired by said search unit; anda changing unit configured to change the description extracted by said extraction unit, based on a content of each feature item of the image feature information of the target area. 2. The diagnosis support apparatus according to claim 1, wherein said search unit calculates a similarity between the image feature information of the target area and image feature information stored in said storage unit, and acquires a finding sentence stored in correspondence with image feature information corresponding to a calculated similarity exceeding a threshold. 3. The diagnosis support apparatus according to claim 2, further comprising a designation unit configured to designate a similarity calculation method in said search unit. 4. The diagnosis support apparatus according to claim 1, further comprising a registration unit configured to receive a sentence as an interpretation result to be registered and image feature information and register the sentence and the image feature information in said storage unit in association with each other. 5. The diagnosis support apparatus according to claim 1, wherein said search unit calculates a similarity between the image feature information of the target area and image feature information stored in said storage means, and acquires a plurality of finding sentences stored in correspondence with image feature information corresponding to a calculated similarity exceeding a threshold, and said creation unit creates a plurality of candidate finding sentences of a finding sentence concerning the target area by using the plurality of finding sentences acquired by said search unit. 6. The diagnosis support apparatus according to claim 5, further comprising a removal unit configured to remove candidate finding sentences having the same content from the plurality of candidate finding sentences. 7. The diagnosis supp... | Canon Kabushiki Kaisha | G06K009/62 A61B005/00 G06F017/30
US08998885 | B2 | 2015-04-07 | System and method for intraparenchymal drug infusion | 2010-09-14 | A system for performing an intraparenchymal drug infusion including a pump device, a delivery tube, a sensor, and a processor. The tube is fluidly coupled to the pump device, establishing an infusate pathway from the pump to an infusate exit port of the delivery tube. The sensor is positioned to sense a parameter indicative of pressure in the infusate pathway. Finally, the processor is programmed to generate information indicative of infusate delivery effectiveness of a drug infusion procedure based upon information from the sensor. In some embodiments, the processor generates a net infusion pressure profile, such as a pressure-time curve, and prompts display of the pressure profile to a neurosurgeon for subsequent evaluation of infusate delivery effectiveness. | 1. A system for use with an intraparenchymal drug infusion procedure in which a drug is delivered from a source to an intracerebral target site via an infusate pathway, the system comprising: a sensor positioned to sense a parameter indicative of pressure in the infusate pathway;a display device; anda processor programmed to: generate a pressure profile for a current intraparenchymal drug infusion procedure infusing drug through the infusate pathway to the intracerebral target site based upon information from the sensor,generate comparative information relating to the current intraparenchymal drug infusion procedure,prompt the display device to simultaneously display the pressure profile and the comparative information, the simultaneous display being indicative of infusate delivery effectiveness of the current intraparenchymal drug infusion procedure, including indicative of whether a therapeutically effective volume of distribution of the infusate remains at the intracerebral target site. 2. The system of claim 1, further comprising a conduit fluidly forming a sensor pathway that connects the sensor to the infusate pathway. 3. The system of claim 2, wherein the conduit branches from a delivery tube otherwise defining the infusate pathway such that the sensor pathway is fluidly in-line with the infusate pathway. 4. The system of claim 1, further comprising a conduit fluidly connected to the sensor and terminating in a sensor exit port, wherein the sensor exit port is located in close proximity to an exit port of the infusate pathway. 5. The system of claim 4, wherein the conduit is connected to a delivery tube otherwise defining the infusate pathway. 6. The system of claim 1, wherein the processor is programmed to generate the pressure profile on a real-time basis. 7. The system of claim 1, wherein the pressure profile is a pressure-time curve. 8. The system of claim 1, wherein the processor is programmed to prompt the display device to display the pressure profile in graph form. 9. The system of claim 1, wherein the processor is programmed to determine baseline pressure information of an intended target site of the current intraparenchymal drug infusion procedure, the comparative information including the baseline pressure information. 10. The system of claim 9, wherein the baseline pressure information is based upon information from the sensor, and is a sum of an interstitial fluid pressure and a hydrostatic pressure difference between a height of the sensor and a height of the infusate exit port. 11. The system of claim 9, wherein the processor is programmed to determine baseline pressure information on a real-time basis. 12. The system of claim 9, wherein the processor is further programmed to formulate a net infusion pressure profile as a function of fluid pressure, as indicated by the sensor, and the baseline pressure information. 13. The system of claim 1, wherein the system further includes a memory storing a library of refere... | Medtronic, Inc. | A61K009/22 A61M005/145 A61M005/142 A61M005/168
US20110015513 | A1 | 2011-01-20 | TRANSPORTABLE DIGITAL AUTONOMOUS UNIT TO PERFORM ANTHROPOMETRIC MEASUREMENTS, OF TACTILE MANUAL LOCALIZATION, GEOMETRICALLY BASED ON THE PROJECTION OF POINTS IN A CARTESIAN (VIRTUAL) PLANE, INCORPORATING TECHNOLOGY OF SENSORS AND SOFTWARE | 2006-10-18 | A systematic procedure and device able to measure the body dimensions of persons, both young and adult, relating to anthropometry, and more concrete methods for anthropometric studies. The device includes an autonomous, transportable digital system to carry out the anthropometric measurements, which permits and facilitates the location and collection of the anthropometric points that are extracted of the individual (8) to be measured. Specifically, the device includes a posture correction element (3), a Cartesian framework (4), data transmitters (4.3 and 4.4), operating software (5), and an information processing component that is preferably a notebook type of computer (5.1). | 1-30. (canceled) 31. Autonomous digital transportable system for the performance of anthropometric measurements, which makes it easier to localize and collect the anthropometric points that are extracted from the person to be measured; storage process, digitalized and automatized processing and post-processing, which in at its phase during the measurement (FIG. 15.I) characterized because such system is formed by, a postural Rectifier 3) that receives the person to be measured (8), a Cartesian Frame (4) that is manipulated by the measurer (9),a computer application program (5), notebook or similar (5.1), wired or wireless data transmitters (6 and 7); mentioned Cartesian Frame (4), specifically the Cartesian manipulators (4.4 and 43) emit a bundle of light rays (433) that projects itself towards the measurable person (8), destined to refer to the localization of the required anthropometric point in a orthogonal and oblique way, this information processed by the sensors housed in the mentioned Cartesian manipulators (4.4 and 4.3) are pre-visualized in real time by means of a pivoting display (4.3.9), these anthropometric dimensions are digitally stored in the processor housed in vertical drained greater body (4.2.1.1), transmitted through wired or wireless data transmitters (6) to the computer application program (5) installed in a notebook or similar (5.1), performed during or subsequently to the measurement, avoiding besides improving the data visualization measurement's operating time. 32. System according to claim 1, characterized because the postural Rectifier (3) is composed of an upper support (3.1), a lower support (3.2) and a column (3.3), the latter, which joins and articulates them, being in turn, formed by Armrests (3.4) and Seat (3.5); this assembly allows to receive the postures of the person to be measured (8) and isolate him/her from the physical place where the measurement is intended to be performed. 33. System according to claim 2, characterized because the postural Rectifier (3), besides when the measured user (8) finds himself/herself over lower Support (3.2) collects the person's weight in order to transmit it with wire data transmitters (7) to the computer application program (5) installed in notebook or similar (5.1); avoids the utilization of auxiliary elements delaying the measurement process. 34. System according to Claim, characterized because the Upper Support (3.1), is a limited hinged positional receptor-rotational support. 35. System according to claim 2, characterized because the Lower Support (3.2), is a container-support, limited four-positional divergent receptor-rotational. 36. Device according to claim 2, characterized because the Column (3.3), is a pivoting support of two hinged vertebras, which form by means of its plates and planes. 37. Device according to claim 2, characterized because the Arm Support (3.4), is a support for upper parts of both children and adults, adjustable, extensible-rotating mul... | - | A61B005/107 A61B006/00
US20110015495 | A1 | 2011-01-20 | METHOD AND SYSTEM FOR MANAGING A USER'S SLEEP | 2010-07-16 | A sleep management method and system for improving the quality of sleep of a user which monitors one or more objective parameters relevant to sleep quality of the user when in bed and receives from the user in waking hours via a portable device such as a mobile phone feedback from objective test data on cognitive and/or psychomotor performance. | 1. A method of managing the sleep of a user, the method comprising: (i) monitoring, using at least one sensor, one or more objective parameters relevant to sleep quality of the user when in bed, said parameters being selected from physiological parameters and environmental parameters and including at least movement of the user and/or electrical signals indicative of brain activity of the user;(ii) collecting, using a sensor unit, signal data from said at least one sensor and communicating the signal data to a processing means;(iii) collecting, using a portable user interaction device, objective test data from the user when awake, the objective test data being indicative of cognitive and/or psychomotor performance and communicating the objective test data to said processing means; and(iv) processing data from said sensor unit and said user interaction device to generate at least a combined sleep indicator metric, which takes both sensor unit data and cognitive and/or psychomotor data into consideration, or a cognitive and/or psychomotor performance metric together with a sleep metric or a sleep quality metric; and(v) receiving at the portable user interaction device information relating to the user's sleep behaviour for display to the user. 2. A method according to claim 1 wherein said at least one sensor do not include any sensor attached to the user. 3. A method as claimed in claim 1, further comprising at least one of: a) supplementing said objective test data with subjective feedback from the user on sleep-related parameters inputted via said user interaction device;b) repeating steps (i) to (iv) at a subsequent time thereby to generate a combined sleep indicator metric for the subsequent time, or a cognitive and/or psychomotor performance metric together with a sleep metric or sleep quality metric for the subsequent time;c) sampling, within the sensor unit, signal data from each sensor at a frequency in the range of 0.1 Hz to 100 Hz before storage to a memory. 4. A method as claimed in claim 1 further comprising displaying information to the user by the user interaction device, the information including one or more recommendations for affecting the behaviour of the user in order to improve subsequent sleep quality. 5. A method as claimed in claim 4 wherein said one or more recommendations are selected from a group of recommendations including behavioural programs and/or actions by the user. 6. A method as claimed in claim 5, further comprising prompting the user, via said portable user interaction device, to implement a behavioural change. 7. A method as claimed in claim 5 wherein said processing means monitors efficacy of and/or compliance with any behavioural program or action presented to and chosen by the user and provides (i) one or more of warning, guidance, advice and message(s) of encouragement to the user to aid efficacy of and/or compliance with any selected program or action and/or (ii) one or more updated recommendations for... | SHARP KABUSHIKI KAISHA | A61B005/00
US20110021940 | A1 | 2011-01-27 | INCENTIVE AUDIO FOR PULMONARY FUNCTION DIAGNOSTICS | 2009-07-21 | A system and method for obtaining measurements of a person's pulmonary diagnostic is described. One embodiment includes a method of providing audible signals related to measurement of pulmonary function, the method comprising receiving a pulmonary input through a receiver, transferring the pulmonary input to a sensor, generating an input signal, using the sensor, based on the pulmonary input, and generating a responsive audible signal through a speaker, wherein the responsive audible signal is based on the input signal. The pulmonary input may comprise, for example, an inhalation, exhalation, gas content reading, or other pulmonary input. The responsive audible signal may comprise a verbal audible signal, a tone, or some combination of both. The responsive audible signal could be for instruction and/or encouragement. The method could further include generating audible instruction signals that are not necessarily related to the pulmonary input. | 1. A method of providing audible signals related to measurement of pulmonary function, the method comprising: receiving a pulmonary input through a receiver;transferring the pulmonary input to a sensor;generating an input signal, using the sensor, based on the pulmonary input; andgenerating a responsive audible signal through a speaker, wherein the responsive audible signal is based on the input signal. 2. The method of claim 1, wherein the pulmonary input comprises an exhalation. 3. The method of claim 1, wherein the responsive audible signal comprises a verbal responsive audible signal. 4. The method of claim 1, wherein generating the responsive audible signal through the speaker comprises: determining a value of a characteristic of the pulmonary input using a processor, wherein the processor determines the value of the characteristic of the pulmonary input based on the input signal; andgenerating the responsive audible signal through the speaker based on the value of the characteristic of the pulmonary input. 5. The method of claim 4, wherein determining the value of the characteristic of the pulmonary input using the processor comprises: determining a flow rate of the pulmonary input using the processor, wherein the flow rate is based on the input signal. 6. The method of claim 1, wherein generating the responsive audible signal through the speaker comprises: determining a first input level using the input signal at a first point in time;determining a second input level using the input signal at a second point in time; andgenerating the responsive audible signal through the speaker based on the difference between the first input level and the second input level. 7. The method of claim 1, wherein the responsive audible signal comprises a motivational responsive audible signal. 8. The method of claim 1, further comprising: generating an audible instruction signal through the speaker. 9. The method of claim 1, further comprising: receiving a pulmonary test type selection; andgenerating an audible instruction signal through the speaker based on the pulmonary test type selection. 10. A method of providing audible feedback while measuring pulmonary function, the method comprising: accepting an input from a pulmonary system of a person;analyzing the input, within a device, to measure a value of the input; andgenerating, using the device, an audible feedback to the person in response to the value of the input. 11. The method of claim 10, wherein analyzing the input to measure the value of the input comprises: analyzing the input within the device to measure the value of the input two or more times in succession. 12. The method of claim 11, wherein generating the audible feedback to the person in response to the value of the input comprises: generating, using the device, the audible feedback to the person in response to each of the measured values of the input. 13. The method of claim 10, wherein the input comprises an inhalation. 14. The... | - | A61B005/08 G09B023/28
US20110034818 | A1 | 2011-02-10 | METHOD AND SYSTEM FOR QUANTITATION OF RESPIRATORY TRACT SOUNDS | 2009-04-07 | Provided is a system and method for analyzing respiratory tract sounds. Sound transducers are fixed on the skin over the thorax that generates signals indicative of pressure waves at the location of the transducer. Processing of the signals involves performing an event search in the signals and determining event parameters for events detected in the search. | 1.-42. (canceled) 43. A system for making a differential diagnosis between two respiratory tract disorders, comprising: (a) an integer N of transducers, each transducer configured to be fixed on a surface of the individual over the thorax, the ith transducer being fixed at a location xi and generating a signal Z(xi,t) indicative of pressure waves at the location xi; i=1 to N at times t during a predetermined time interval;(b) a processor configured to: receive the signals Z(xi,t) and to process the signals, wherein the processing comprises performing at least one event search; determine values of one or more initial event parameters; determine values of the one or more final event parameters andcompare the values of the initial event parameters to the final event parameters, and to make differential diagnosis based on the comparison. 44. The system according to claim 43, wherein an event search is performed on one or more of the signals Z(xi,t). 45. The system according to claim 43, wherein an event search is performed on one or more signals P(xi,t) wherein the signals P(xi,t) are obtained after performing one or more procedures on one or more of the signals Z(xi,t) selected from filtering, denoising, smoothing, envelope extraction, applying a mathematical transformation. 46. The system according to claim 43, wherein the transducers are divided into one or more subsets and the processing comprises, for each of one or more of the subsets, calculating a representative signal from one or more of the signals Z(xi,t) or P(xi,t) obtained from transducers in the subset and performing one or more event searches on one or more of the representative signals. 47. The system according to claim 43, wherein one or more of the events are selected from an entire breathing cycle, an inspiratory phase of a breathing cycle, and an expiratory phase of a breathing cycle. 48. The system according to claim 47, wherein the representative signal of a transducer subset is a summation or an average signal of the signals obtained by the transducers in the subset. 49. The system according to claim 43, wherein the event search comprises performing any one or more of a peak search, an autocorrelation, a cross correlation with a predetermined function, and a Fourier transform. 50. The system according to claim 43, wherein one or more of the event parameters are selected from the group consisting of a time that an event occurred, a duration of an event, a magnitude of an event, a height of a peak associated with the event, the width of a peak associated with the event in a signal at half peak height, a half time to rise of a peak associated with the event in a signal, a half time to fall of a peak, an area under a peak; a maximum of the signal during the event, a ratio of a maximum during an inspiratory phase to a maximum during an expiratory phase, a ratio of a duration of an inspiratory phase to a duration of an expiratory phase, and a morphology of a signal during the... | DEEPBREEZE LTD. | A61B005/08
US07871597 | B2 | 2011-01-18 | Polyol and polyether iron oxide complexes as pharmacological and/or MRI contrast agents | 2003-04-09 | Pharmacological compositions, and methods for administration, of the type employing an iron oxide complex with a polyol or polyether. The methods of administration may comprise parenteral administration of an effective dose of the complex formulated in a biocompatible liquid delivered at a rate of from about 1 mL/sec to less than 1 mL/min and wherein upon administration the complex provides minimal detectable free iron in a subject, and minimal incidence of anaphylaxis. The pharmacological compositions are of the type employing a polyol or polyether iron oxide complex, which, upon parenteral administration to a subject, are substantially immunosilent, provide minimal anaphylaxis and minimal free iron, and undergo minimal dissolution in vivo. | 1. A unit dosage of a pharmacological composition comprising a terminally sterilizable autoclavable reduced carboxymethylated dextran iron oxide complex composition which is stable at 121 C and is prepared, in a total volume of biocompatible liquid from about 1 mL to about 15 mL and for a total single dose from about 50 mg to about 600 mg, the pharmacological composition capable of being parenterally administered to a subject at a rate substantially greater than 1 mL/min and wherein the unit dosage provides upon administration: minimal detectable free iron in the subject; andminimal incidence of anaphylaxis. 2. A unit dosage of a pharmacological composition according to claim 1, wherein the unit dosage provides minimal dissolution upon administration in a human subject. 3. A unit dosage of a pharmacological composition according to claim 1, wherein the unit dosage provides minimal incidence of anaphylaxis upon administration in a human subject. 4. A unit dosage of a pharmacological composition according to claim 3, wherein the unit dosage is substantially immunosilent upon administration in a human subject. 5. A method for administering a unit dosage of a pharmaceutical composition wherein the method comprises: administering parenterally an effective dose of unit dosage of a pharmaceutical composition as claimed in claim 1, the complex formulated in a biocompatible liquid so that upon administration the complex provides minimal detectable free iron in a subject, and minimal incidence of anaphylaxis; andeffecting such administration at a rate substantially greater than 1 mL/min. 6. A method for administering a unit dosage of a pharmaceutical composition according to claim 5, further comprising effecting administration at a rate of about 1 mL/sec. 7. A method for administering a unit dosage of a pharmaceutical composition according to claim 5, wherein the complex is formulated in a biocompatible liquid so that it provides minimal dissolution of the complex upon administration in a human subject. 8. A method for administering a unit dosage of a pharmaceutical composition according to claim 7, wherein the complex is formulated in a biocompatible liquid so that it is substantially immunosilent upon administration in a human subject. | AMAG Pharmaceuticals, Inc. | A61B005/055 A01N043/04 A61K031/715
US20110009715 | A1 | 2011-01-13 | INGESTIBLE EVENT MARKER DATA FRAMEWORK | 2009-07-02 | The ingestible event marker data framework provides a uniform, comprehensive framework to enable various functions and utilities related to ingestible event marker data (IEM data). The functions and utilities include data and/or information having an aspect of data derived from, collected by, aggregated by, or otherwise associated with, an ingestion event. | 1. A system comprising: ingestible event marker data;a hub to receive the ingestible event marker data; andat least one ingestible event marker data system to receive, directly or indirectly, the ingestible event marker data from the hub. 2. The system of claim 1, wherein the ingestible event marker data comprises information about an ingestion event. 3. The system of claim 1, wherein the ingestible event marker data comprises information about a response to an ingestion event. 4. The system of claim 1, wherein the ingestible event marker data is generated by an ingestible event marker device. 5. The system of claim 1, wherein the ingestible event marker device comprises an ingestible event marker and a receiver. 6. The system of claim 1, wherein the receiver comprises a personal signal receiver. 7. The system of claim 1, wherein the ingestible event marker data is associated with a unique identifier of an individual. 8. The system of claim 1, wherein the hub comprises at least one element selected from a group consisting essentially of a base station, a personal communication device, and a mobile telephone. 9. The system of claim 8, wherein the base station comprises at least one element selected from a group consisting essentially of a desktop computer, a laptop computer, and an intelligent device/appliance. 10. The system of claim 8, wherein the personal communication device comprises a mobile computer. 11. The system of claim 1, wherein the at least one ingestible event marker data system comprises at least one element selected from a group consisting essentially of: a feedback loop system, a decision support system, an auto refill system, patient tools, a behavioral medicine system, an incentive system, personalized commercial products/services, an auto billing system, a tracking system, an interdiction system, a subscription system, an ingestible event marker data collection, an approval system, a forecasting system, a financial system, an ingestible event marker data phone system, and a social network system. 12. The system of claim 1, wherein the at least one ingestible event marker data system comprises at least one element selected from a group consisting essentially of a computer, a receiver, a transmitter, an application, a data storage medium, a processor, a memory component, a personal communication device, software, and a mobile telephone. 13. The system of claim 1, further comprising: an ingestible event marker data source to at least one of generate at least a portion of the ingestible event marker data and transmit at least a portion of the ingestible event marker data. 14. The system of claim 13, wherein the ingestible event marker data source is selected from a group consisting essentially of an ingestible event marker data source device, a product, an event, a patient specific parameter, an ingestible event marker data algorithm, and a storage repository. 15. The system of claim 1, further comprising at least on... | - | A61B005/07
US07874993 | B2 | 2011-01-25 | System and method for diagnosing and monitoring congestive heart failure | 2005-06-10 | A system for diagnosing and monitoring congestive heart failure for automated remote patient care is presented. A database stores a plurality of monitoring sets relating to patient information recorded on a substantially continuous basis. A server retrieving and processing the monitoring sets includes a comparison module determining patient status changes by comparing at least one recorded measure from one of the monitoring sets to at least one other recorded measure from another of the monitoring sets with both recorded measures relating to a type of patient information, and an analysis module testing each patient status change for one of an absence, an onset, a progression, a regression, and a status quo of congestive heart failure against a predetermined indicator threshold corresponding to a type of patient information as the recorded measures. The indicator threshold corresponds to a quantifiable physiological measure of a pathophysiology indicative of congestive heart failure. | 1. A non-transitory computer-readable storage medium storing code for remote automated congestive heart failure patient management, comprising: code for receiving information to identify a patient enrolled in automated patient care, the information comprising at least one of treatment profile and patient medical history;code for receiving collected device measures recorded by a medical device as raw physiometry for the patient;code for deriving device measures as derivative physiometry based on the collected device measures;code for comparing at least one of the derived device measures to at least one other of the derived device measurescode for determining a pathophysiology indicative of at least one of a change and a status quo of congestive heart failure;code for receiving hysteresis parameters as a temporally-defined threshold for changes in the pathophysiology; andcode for adapting the hysteresis parameters depending upon patient medical history and peer-group guidelines. 2. A non-transitory computer-readable storage medium according to claim 1, further comprising: code for receiving evaluation parameters for the pathophysiology in terms of at least one of a cardiovascular and cardiopulmonary disease; andcode for determining a wellness indicator analyzing one or more of the derived device measures against the evaluation parameters. 3. A non-transitory computer-readable storage medium according to claim 2, further comprising: code for determining a patient comprising code for evaluating the wellness indicator to identify at least one of a change and a status quo of congestive heart failure. 4. A non-transitory computer-readable storage medium according to claim 2, wherein the code for determining the wellness indicator further comprises code for analyzing one or more of the derived device measures against the hysteresis parameters. 5. A non-transitory computer-readable storage medium according to claim 1, further comprising: code for receiving quality of life measures as qualified voice feedback data, wherein the voice feedback data is electronically recorded by the patient contemporaneous to recordation of the collected device measures. 6. A non-transitory computer-readable storage medium according to claim 1, wherein one or more of the derived device measures are selected from the group comprising self referencing, peer group, and general population physiometric measures. 7. A non-transitory computer-readable storage medium according to claim 1, further comprising: code for generating reprogramming parameters for the medical device; andcode for applying the programming parameters before the medical device has recorded the collected device measures. 8. A non-transitory computer-readable storage medium according to claim 1, wherein the medical device comprises one of an implantable medical device and an external medical device. 9. A non-transitory computer-readable storage medium according to claim 1, further comprising: code for rece... | Cardiac Pacemakers, Inc. | A61B005/04
US07862507 | B2 | 2011-01-04 | Confocal Raman Spectroscopy for dermatological studies | 2006-08-28 | Use of Confocal Raman Spectroscopy (CRS) for dermatological studies, including a method for determining the thickness of the Stratum Corneum (SC) on a test area of the skin, and to a method for quantifying the effectiveness of a skin care composition. The methods of the invention can be carried in vitro (either artificial skin or a sample of skin) or in vivo (directly on the human skin of a person). | 1. A method for determining the effectiveness of a skin care composition, comprising the steps of: (a) determining, using Confocal Raman Spectroscopy, a pre-application thickness of the Stratum Corneum within a test area;(b) applying a skin care composition to the test area;(c) determining, using Confocal Raman Spectroscopy, a post-application thickness of the Stratum Corneum within the test area; and(d) calculating, using a computer, the effectiveness of the skin care composition; wherein the thickness of the Stratum Corneum before and after the application of the skin care composition are input as parameters when calculating the effectiveness of the skin care composition. 2. A method for determining the effectiveness of a skin care composition, comprising the steps of: (i) selecting a Raman-active substance linked to the effectiveness of the skin care composition to be determined;measuring the concentration profile of said Raman-active substance as a function of depth within a test area of skin using Confocal Raman Spectroscopy;(iii) determining the pre-application thickness of the Stratum Corneum within said test area;(iv) applying the skin care composition to said test area;(v) measuring the concentration profile of said Raman-active substance as a function of depth within said test area using Confocal Raman Spectroscopy;(vi) determining the post-application thickness of the Stratum Corneum within said test area; and(vii) calculating, using a computer the effectiveness of the skin care composition as a function of both: (a) the concentration profile of the Raman-active substance before and after the application of the skin care composition; and(b) the thickness of the Stratum Corneum before and after the application of the skin care composition. 3. A method according to claim 2, wherein the thickness of the Stratum Corneum is determined by a method comprising the steps of: (a) measuring the concentration profile of a Raman-active material as a function of depth within the test area using Confocal Raman Spectroscopy; then(b) processing the data obtained to determine the thickness of the Stratum Corneum. 4. A method according to claim 3, wherein the data is processed using the Weibull algorithm. 5. A method according to claim 2, wherein the measurement is made in vivo. 6. A method according to claim 2, wherein the Raman-active material is selected from the group consisting of water, NMF, vitamins and combinations thereof. 7. A method according to claim 6, wherein the Raman-active material is water. 8. A method according to claim 2, wherein the skin care composition comprises at least one moisturizing agent. 9. A method according to claim 2, wherein the effect to be determined is skin moisturization. 10. A method according to claim 2, wherein the Raman-active substance linked to the effectiveness of the skin care composition is water. 11. A method according to claim 2 wherein the effectiveness of the skin care composition is calculated... | The Procter & Gamble Company | A61B005/00 A61B006/00 A61B005/103
US07889905 | B2 | 2011-02-15 | Fast 3D-2D image registration method with application to continuously guided endoscopy | 2006-05-19 | A novel framework for fast and continuous registration between two imaging modalities is disclosed. The approach makes it possible to completely determine the rigid transformation between multiple sources at real-time or near real-time frame-rates in order to localize the cameras and register the two sources. A disclosed example includes computing or capturing a set of reference images within a known environment, complete with corresponding depth maps and image gradients. The collection of these images and depth maps constitutes the reference source. The second source is a real-time or near-real time source which may include a live video feed. Given one frame from this video feed, and starting from an initial guess of viewpoint, the real-time video frame is warped to the nearest viewing site of the reference source. An image difference is computed between the warped video frame and the reference image. The viewpoint is updated via a Gauss-Newton parameter update and certain of the steps are repeated for each frame until the viewpoint converges or the next video frame becomes available. The final viewpoint gives an estimate of the relative rotation and translation between the camera at that particular video frame and the reference source. The invention has far-reaching applications, particularly in the field of assisted endoscopy, including bronchoscopy and colonoscopy. Other applications include aerial and ground-based navigation. | 1. A method of registering video frames of a body lumen received from an endoscope inserted into the body lumen to previously acquired images of the body lumen, the method comprising the steps of: computing a plurality of reference images of a body lumen, each reference image including a viewpoint with a depth map and image gradients;receiving video frames from the endoscope wherein said receiving step comprises receiving a live bronchoscopic video frame associated with an airway; andproviding a processor to register each video frame received from the endoscope to at least one of the reference images, said processor being programmed to perform the steps of: a) warping the video frame to the viewpoint of one reference image of the plurality of reference images or warping one reference image of the plurality of reference images at a viewpoint to the video frame;b) automatically computing an image difference between (i) the warped video frame and the reference image, or (ii) the warped reference image and the video frame;c) updating a viewpoint of the video frame using an optimization technique; andd) repeating steps a) through c) for each video frame until the viewpoint of the video frame converges on a final viewpoint or a next video frame becomes available. 2. The method of claim 1, wherein at least one reference image is updated before, during, or after registration. 3. The method of claim 1, including the step of guessing an initial viewpoint of the video frame. 4. The method of claim 1, wherein the final viewpoint gives an estimate of the relative rotation and translation between the video frame of the video feed and the reference image. 5. The method of claim 1, further including the step of using a resolution pyramid wherein all images, depth maps, and gradients are down-sampled. 6. The method of claim 1, further including the step of using a weighted normalized cross-correlation objective function. 7. The method of claim 1, wherein the optimization technique is based on a Gauss-Newton parameter update. 8. The method of claim 1 wherein the computed set of reference images of a body lumen comprises a 3-D model of a bronchial tree. 9. The method of claim 1 wherein the body lumen is an airway. 10. A method of registering video frames of a body lumen received from an endoscope inserted into the body lumen to previously acquired images of the body lumen, the method comprising the steps of: computing a plurality of reference images of a body lumen, each reference image including a viewpoint with a depth map and image gradients;receiving at least one video frame from the endoscope; andproviding a processor to register each video frame received from the endoscope to at least one of the reference images using of the plurality, said processor being programmed to perform the steps of: a) warping the video frame to the viewpoint of one reference image of the plurality of reference images or warping one reference image of the plurality of refere... | The Penn State Research Foundation | G06K009/00 G06K009/32 A61B005/05
US20110028827 | A1 | 2011-02-03 | SPATIOTEMPORAL PATTERN CLASSIFICATION OF BRAIN STATES | 2009-07-28 | A multivariate, pattern-based system and method for recognition of brain states and generation of feedback considering both the spatial and temporal pattern of network of brain activity is described. The system can be applied for enhancing a desired function or behavior, or to alleviate a behavioral or neurological problem. The system can also be used to investigate the spatiotemporal evolution of a network of brain activity corresponding to a brain function, by changing the activity as an independent variable and studying its effect on behavior. | 1. An apparatus comprising: a real-time spatiotemporal pattern classifier configured to classify spatial and/or temporal spectroscopic data of brain states in real time, which includes a feature extraction module anda pattern classification module; anda read-out having spatial and/or temporal components. 2. The apparatus of claim 1, wherein the real-time spatiotemporal pattern classifier is configured to classify both spatial and temporal spectroscopic data and the read-out has both spatial and temporal components. 3. The apparatus of claim 1, wherein the feature extraction module comprises a data driven algorithm or a model driven algorithm. 4. The apparatus of claim 3, wherein the data driven algorithm comprises a principal component analysis (PCA) model. 5. The apparatus of claim 3, wherein the data driven algorithm comprises an independent component analysis (ICA) model. 6. The apparatus of claim 3, wherein the data driven algorithm comprises a general linear model (GLM) and a Granger causality model (GCM). 7. The apparatus of claim 1, wherein the pattern classification module comprises a Hidden Markov Model (HMM) or a Support vector machine (SVM). 8. The apparatus of claim 1, wherein the real-time spatiotemporal pattern classifier is configured to classify data in about 0.5-3.0 seconds. 9. A system comprising: a spectroscope;a real-time spatiotemporal pattern classifier configured to classify spatial or temporal data in real time, which includes a feature extraction module anda pattern classification module; anda read-out having spatial and/or temporal components. 10. The system of claim 9, wherein the spetroscope is a functional MRI device (fMRI). 11. The system of claim 9, wherein the spectroscope is a near infrared spectroscope (NIRS). 12. The system of claim 10, wherein the fMRI is configured to utilize an echo planer imaging sequence. 13. A method comprising: obtaining a spectroscopic signal from a subject brain, the signal having spatial or temporal data;extracting one or more features from the signal in real time;classifying a pattern of a brain state in real time; andgenerating a read-out comprising said spatial and/or temporal data. 14. The method of claim 13, wherein two or more brain states are classified in real time. 15. The method of claim 13, wherein the read-out is associated with a specific physical, cognitive, or emotional task of the subject. 16. The method of claim 13, further comprising providing feedback to the subject. 17. The method of claim 13, further comprising recognizing a brain state in the subject brain. 18. The method of claim 13, further comprising recognizing, diagnosing, or treating a disease or condition in the subject. 19. The method of claim 18, wherein the condition is the utterance of a false statement. 20. The method of claim 18, wherein the condition is reduced limb function. | - | A61B005/055 G06K009/00
US20110021894 | A1 | 2011-01-27 | GLUCOSE SENSOR EMPLOYING SEMICONDUCTOR NANOELECTRONIC DEVICE | 2010-09-30 | A glucose sensor employs a programmable glucose sensor array of a relatively large number of nanoelectronic devices (e.g. semiconductor field-effect devices) having control surfaces functionalized with a glucose-reactive substance and generating sensing signals indicative of sensed glucose level of a bodily fluid. The devices are divided into sub-sets sequentially enabled over successive intervals to achieve overall sensor lifetime many times longer than the lifetime of any single device in operation. | 1. A glucose sensor, comprising: a nanoelectronic device having a control surface functionalized with a glucose-reactive substance; anda fluid interface structure configured to allow contact between the control surface and a bodily fluid. 2. A glucose sensor according to claim 1, comprising an array of the nanoelectronic devices having respective control surfaces also functionalized with the glucose-reactive substance, and wherein the fluid interface structure is configured to allow contact between the control surfaces and the bodily fluid. 3. A glucose sensor according to claim 1, wherein the nanoelectronic device is configured such that chemical interaction between the glucose-reactive substance and glucose in the bodily fluid affects electrical conduction characteristics of the nanoelectronic device. 4. A glucose sensor according to claim 1, wherein the nanoelectronic device is configured such that chemical interaction between the glucose-reactive substance and glucose in the bodily fluid effects capacitive or other parametric changes of the nanoelectronic device. 5. A glucose sensor according to claim 1, wherein the nanoelectronic device has a sensing element critical dimension less than 100 nm. 6. A glucose sensor according to claim 1, wherein the nanoelectronic device has a sensing element critical dimension less than 500 nm. 7. A glucose sensor, comprising: an array of nanoelectronic devices having respective control surfaces functionalized with a glucose-reactive substance which chemically interacts with glucose to affect electrical conduction characteristics of the nanoelectronic devices, the array being configured to allow for intimate contact between the control surfaces and a glucose-carrying bodily fluid, the array of nanoelectronic devices being logically organized into a plurality of individually operable subsets of the nanoelectronic devices, each subset being operable for only a limited period before operational degradation due to interaction between the bodily fluid and operating nanoelectronic sensors of the subset;device selection circuitry operative in response to control inputs to enable electrical sensing operation of a selected one of the subsets of the nanoelectronic devices to generate respective sensing output signals while simultaneously disabling such electrical sensing operation of remaining ones of the subsets of the nanoelectronic devices; andcontrol circuitry operative to generate the control signals so as to serially enable electrical operation of successive ones of the subsets of the nanoelectronic devices over an extended period generally equal to the product of the limited period and the number of the subsets of the nanoelectronic devices. 8. A glucose sensor according to claim 7 wherein the array is configured for implantation into a body tissue to provide for the intimate contact between the control surfaces and the glucose-carrying bodily fluid. 9. A glucose sensor according to claim 7 wherein the n... | TRUSTEES OF BOSTON UNIVERSITY | A61B005/1468 G01N027/26
US07865375 | B2 | 2011-01-04 | System and method for multidimensional extension of database information using inferred groupings | 2006-01-03 | A system and method for receiving medical or other database information and pregrouping and extending that data include a data enhancement layer configured to generate additional stored dimensions capturing the data and relevant attributes. Data sources such as hospitals, laboratories and others may therefore communicate their clinical data to a central warehousing facility which may assemble and extend the resulting aggregated data for data mining purposes. Varying source format and content may be conditioned and conformed to a consistent physical or logical structure. The source data may be extended and recombined into additional related dimensions, pre-associating meaningful attributes for faster querying and storage. The attributes, data and other pieces of information may likewise in embodiments be subjected to an inference analysis to determine whether previously unidentified or unexploited relationships may exist within the universe of source data, for instance using correlation, inference or other analytic techniques. Newly detected, identified or inferred data groupings, which may for instance reveal hidden trends or patterns residing in the data, may then be added back to the enhanced data groupings. Users running analytics against the resulting medical or other datamarts may therefore access a richer set of related information, more powerful sets of predictive models as well as have their queries and other operations run more efficiently. | 1. A system for generating an inferred data grouping, comprising: an input interface receiving clinically relevant source data from a first data source and clinically relevant source data from a second data source, wherein each of the first data source and the second data source comprises a healthcare provider, a hospital, an insurer, a laboratory, or a medical facility;an inference engine, the inference engine communicating with the input interface to access the clinically relevant source data,automatically detecting an inferred clinical relationship between at least one attribute or data in the clinically relevant source data from the first data source and at least one attribute or data in the clinically relevant source data from the second data source being separate from the first data source comprising the healthcare provider, the hospital, the insurer, the laboratory, or the medical facility, wherein the inferred clinical relationship is automatically detected using an inference measure to detect and substantiate the inferred clinical relationship,generating an inferred data grouping in accordance with the detected inferred clinical relationship, andcombining the inferred data grouping with an existing clinical data grouping to generate an inferentially enhanced data grouping that is used to analyze past healthcare-related trends and model future healthcare-related events to project a healthcare-related cost, outcome, trend, or interest; anda query engine, the query engine receiving a user query,in response to receiving the user query, utilizing the inferentially enhanced data grouping to analyze one or more healthcare-related trends, andproviding a user with output based on the one or more healthcare-related trends. 2. A system according to claim 1, wherein the clinically relevant source data from the first data source comprises data generated by at least one of a health care provider, a hospital, an insurer and a laboratory. 3. A system according to claim 1, wherein the inference measure comprises a correlation measure between the at least one attribute or data in the clinically relevant source data from the first data source and at least one attribute or data in the clinically relevant source data from the second data source. 4. A system according to claim 1, wherein the inferred data grouping comprises a dimensional or multidimensional data grouping. 5. A system according to claim 4, wherein the dimensional or multidimensional data grouping is combined with the existing clinical data grouping to generate an inferentially enhanced dimensional or multidimensional data grouping. 6. A system according to claim 1, wherein the inference engine stores the inferred data grouping to a datamart. 7. A system according to claim 6, wherein the datamart is configured to receive queries. 8. A system according to claim 7, wherein the queries comprise standard query language queries. 9. A system according to claim 7, wherein the queries comprise ... | Cerner Innovation, Inc. | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07862515 | B2 | 2011-01-04 | Apparatus for detecting sleep apnea using electrocardiogram signals | 2006-03-24 | An apparatus for diagnosing sleep apnea only uses an electrocardiogram signal and a computer processor and associated computer code that are configured to analyze the electrocardiogram signal and classify each time period of the electrocardiogram signal as either apneic or normal. A diagnostic measure of sleep apnea for the human patient is provided based on classification results obtained by combining a results for various electrocardiogram signal time periods. A computer-readable medium has computer code that causes the computer to analyze an electrocardiogram signal relating to the human patient and classify each time period in a set of time periods as either apneic or normal. A diagnostic measure of sleep apnea is provided based on classification results obtained by combining a plurality of results from the set of time periods using either time or frequency domain processing, or both. | 1. An article of manufacture comprising a memory containing computer software program instructions therein which, when executed by a processor, cause the processor to: analyze an electrocardiogram signal obtained from a human patient;partition the electrocardiogram signal into a plurality of time periods;derive an estimated respiration signal from the electrocardiogram signal;classify each time period of the plurality of time periods of the electrocardiogram signal as either apneic or normal by analyzing both the electrocardiogram signal and the estimated respiration signal; andprovide a diagnostic measure of sleep apnea for the human patient based on a combined classification result obtained by combining a plurality of the classifications from the plurality of time periods. 2. The article of manufacture of claim 1, wherein the computer software program instructions causes the processor to determine a plurality of R-R intervals in the electrocardiogram signal. 3. The article of manufacture of claim 1, wherein the computer software program instructions causes the processor to calculate a power spectral density of an RR interval. 4. The article of manufacture of claim 1, wherein the computer software program instructions causes the processor to determine the estimated respiration signal derived from the electrocardiogram signal, wherein the electrocardiogram signal is provided as a single lead electrocardiogram signal. 5. The article of manufacture of claim 4, wherein the computer software program instructions causes the processor to calculate a power spectral density of the estimated respiration signal. 6. The article of manufacture of claim 1, wherein the computer software program instructions causes the processor to use time domain processing to provide the diagnostic measure of sleep apnea. 7. The article of manufacture of claim 1, wherein the computer software program instructions causes the processor to use frequency domain processing to provide the diagnostic measure of sleep apnea. 8. The article of manufacture of claim 1, wherein the computer software program instructions causes the processor to use both time domain and frequency domain processing to provide the diagnostic measure of sleep apnea. 9. The article of manufacture of claim 1, wherein the computer software program instructions causes the processor to calculate, for each time period in the plurality of times periods, an associated probability that each of the plurality of time periods is apneic. | Biancamed Limited | A61B005/0402
US07862521 | B1 | 2011-01-04 | Uterine contraction sensing system and method | 2005-05-06 | The present invention provides a system for enabling the sensing of changes in pressure in response to uterine muscle tone changes relating to contractions, and for enabling the evaluation of contractions based thereon. The system enables a portion thereof to effectively and efficiently project at least partially below the non-compressed abdominal surface, and provides electrical isolation of the system elements. It enables ease of use and enhanced patient comfort. It also enables maximized sensitivity and minimized system loading retention pressure. | 1. A system for sensing the frequency and duration of uterine contractions during labor and estimating the relative intensity thereof, able to sense changes in pressure in response to uterine muscle tone changes relating to contractions, and to enable a fetal monitor to be connected thereto for evaluation of contractions, comprising: a force collector, for collecting force generated in response to uterine muscle tone changes relating to contractions, positionable against the exterior abdominal wall proximate the uterine muscle, and supportable so as to project into the abdominal wall to reside at least partially below the non-compressed surface of the abdomen;a pressure sensor, for sensing changes in pressure in the force collector responsive to uterine muscle tone changes, positionable adjacent to the force collector which is positionable against the exterior abdominal wall proximate the uterine muscle, which pressure sensor is supportable so as to project into the abdominal wall to reside at least partially below the non-compressed surface of the abdomen, and is connectable to a fetal monitor for evaluation of contractions during labor, which includes a supporting member, which includes a plurality of contact pads, and the system further comprises a cable for connecting to a fetal monitor at one end thereof, a plurality of connectors extending from the opposite end of the cable and secured to the plurality of contact pads, and a plurality of wires extending from the sensing member to the contact pads, and a sensing member for sensing changes in responsive to uterine muscle changes relating to contractions, supportable in the supporting member so as to project into the abdominal wall to reside at least partially below the non-compressed surface of the abdomen, and connectable to a fetal monitor for evaluation of contractions during labor; anda supporting element, for supporting the force collector so as to be positionable against the exterior abdominal wall proximate the uterine muscle and for supporting the pressure sensor adjacent to the force collector, and for supporting the force collector and the pressure sensor such that the force collector and the pressure sensor projects into the abdominal wall to reside at least partially below the non-compressed surface of the abdomen. 2. A system as in claim 1, wherein the plurality of connectors each comprise solid strand telephone cable. 3. A system as in claim 1, wherein the plurality of connectors in the cable are able to be soldered directly to the contact pads. 4. A system for sensing the frequency and duration of uterine contractions during labor and estimating the relative intensity thereof, able to sense changes in pressure in response to uterine muscle tone changes relating to contractions, and to enable a fetal monitor to be connected thereto for evaluation of contractions, comprising: a force collector, for collecting force generated in response to uterine muscle tone changes relating ... | Parker-Hannifin Corporation | A61B005/103 A61B005/117
US07878980 | B2 | 2011-02-01 | Gas flow diverter for respiratory monitoring device | 2004-06-08 | A gas flow diverter for increasing the turbulence of an incoming respiratory gas flow through a respiratory gas sensor is provided. The diverter is positioned at a patient or inlet end of a tube section of the sensor in order to divert the respiratory gas flow from the patient generally equally across the cross-sectional area of the tube section. The diverter can be formed integrally as a part of the sensor, or can be releasably attached to the sensor or to an adapter that is releasably connected to the inlet end of the sensor for use in situations in which the volume of the incoming respiratory gas flow from the patient is greatly reduced due to the size of the patient, i.e., when the patient is an infant. | 1. A gas flow diverter in combination with a respiratory measuring device, the combination comprising: a respiratory sensor having a tube section with an inlet end and an outlet end and adapted for communicating a respiration flow through an interior passage;a number of apertures formed through the tube section near a restriction, each aperture being fluidly connected to the interior passage; anda diverter comprising: a body defining a generally blunt end and having a generally conical surface extending from the blunt end, the blunt end being engageable with the tube section of the respiratory sensor such that the conical surface extends away from the blunt end and terminates nearer the restriction and any of the apertures of the tube section than the blunt end, the entire body being longitudinally offset, in a common direction aligned with a direction of the respiration flow, from the restriction and any of the apertures such that the diverter increases turbulence of the respiration flow before the respiration flow enters the restriction to generate a pressure differential on generally opposite sides of the restriction relative to the respiration flow. 2. The combination of claim 1 wherein the generally blunt end includes a peripheral ridge engageable with the tube section. 3. The combination of claim 1 wherein the generally blunt end includes at least one port adapted to direct an incoming gas flow around the diverter. 4. The combination of claim 3 wherein the generally blunt end includes a number of protrusions disposed on opposite sides of the at least one port. 5. The combination of claim 3 wherein the generally blunt end includes at least four ports. 6. The combination of claim 1 wherein the body is formed from a non-porous material. 7. The combination of claim 6 wherein the body is formed from a material selected from the group consisting of: a metal, a plastic, and a rubber. 8. The combination of claim 1 wherein the body has a generally conical concave surface. | Treymed, Inc. | A61B005/08
US07876942 | B2 | 2011-01-25 | System and method for optical position measurement and guidance of a rigid or semi-flexible tool to a target | 2007-03-29 | A system for measuring the position of a hand-held tool relative to a body in six degrees of freedom employs a camera attached via a mechanical linkage to the tool so that the camera moves together with the proximal portion of the tool and has a field of view including the distal end of the tool. A processing system processes images from the camera to determine a position of at least part of the tool. By employing optically identifiable fiducial reference points defined on the external surface of the body, a projection of the tool tip position onto a plane containing the target can be derived and displayed together with the target position, thereby facilitating guiding the tool to the target. | 1. A system for measuring the position of a hand-held tool relative to a body in at least five degrees of freedom, the system comprising: (a) a rigid or semi-flexible tool having a distal end for insertion into the body, and a proximal portion for manual manipulation outside the body;(b) a marker arrangement comprising a patch configured to be applied to an external surface of the body at a point of insertion of the tool into the body, said patch carrying a plurality of fiducial points;(c) a camera for generating images;(d) a mechanical linkage mounting said camera on said proximal portion of said tool such that: (i) said camera moves together with said proximal portion of said tool, and(ii) said camera is directed with a field of view including at least part of said patch when said patch is deployed on an external surface of the body and said distal portion of said tool in inserted into the body; and(e) a processing system in data communication with said camera and configured to process images from said camera to determine a position of at least part of said tool relative to said fiducial points. 2. The system of claim 1, wherein said plurality of fiducial points includes at least one set of four fiducial points, and wherein said patch is configured to maintain said set of four fiducial points substantially in a common plane. 3. The system of claim 1, wherein said patch includes a plurality of markers configured to be readily detected by a non-optical imaging system. 4. The system of claim 3, wherein said markers are coincident with said fiducial points on said patch. 5. The system of claim 1, wherein said patch is configured to delineate a point of penetration of the distal end of the tool into the body. 6. The system of claim 5, wherein said plurality of fiducial points include a first set of fiducial points and a second set of fiducial points optically distinguishable from said first set of fiducial points, said first set of fiducial points being more closely spaced than said second set of fiducial points. 7. The system of claim 1, wherein said processing system is further configured to derive a current tip position of the distal end of the tool, said deriving including calculating an estimation of flexion of the tool, and employing said estimation of flexion to determine said current tip position. 8. The system of claim 1, wherein said camera and at least part of said processing system are implemented on a common processor chip. 9. A method for guiding a distal end of a rigid or semi-flexible tool to a target within a body, the tool having a proximal portion for manual manipulation from outside the body, the method comprising the steps of: (a) determining a spatial relation between a plurality of optically identifiable fiducial reference points defined on an external surface of the body and the target;(b) providing a camera mechanically attached to the proximal portion of the tool; and(c) during insertion of the tool into the body: (... | Activiews Ltd. | G06K009/00 A61B005/05
US07892240 | B2 | 2011-02-22 | Femoral head center locating apparatus and method | 2007-09-13 | A femoral head center locating mechanism and method for using same. The femoral head center locating mechanism may include a frame and a plurality of lasers which emit laser beams which intersect to identify a reference point relative to a center of a femoral head. | 1. A method for identifying a reference point relative to a femoral head center of a femur having a femoral head, comprising the steps of: moving the femur along an arc of a circle to a first position;moving the femur along the arc of the circle to a second position;moving the femur along the arc of the circle to a third position;projecting a first light beam in a general direction toward the femoral head from a first midpoint between the first position and the second position;projecting a second light beam in a general direction toward the femoral head from a second midpoint between the second position and the third position; andidentifying the reference point relative to the femoral head center by determining an intersection of the first light beam and the second light beam. 2. The method of claim 1, wherein the first position and the second position define a first chord of the circle, said first projecting step comprising projecting the first light beam in a direction substantially perpendicular to the first chord. 3. The method of claim 2, wherein said first projecting step further comprises projecting the first light beam in a direction which substantially bisects the first chord. 4. The method of claim 1, wherein the second position and the third position define a second chord of the circle, said second projecting step comprising projecting the second light beam in a direction substantially perpendicular to the second chord. 5. The method of claim 4, wherein said second projecting step further comprises projecting the second light beam in a direction which substantially bisects the second chord. 6. The method of claim 1, further comprising the step of providing a frame having a first post corresponding to the first position, a second post corresponding to the second position, and a third post corresponding to the third position, wherein the first, second, and third posts are arranged parallel to each other, wherein the first position, the second position, and the third position lie in a plane extending perpendicular to each of the first, second, and third posts. 7. The method of claim 1, further comprising the step of providing a line extending through the reference point and extending perpendicular to the first light beam and the second light beam, the line extending through the femoral head center. | Zimmer, Inc. | A61F002/46 A61B017/74 A61B005/103
US09006660 | B2 | 2015-04-14 | Scanning terahertz probe | 2008-08-28 | A THz radiation probe (1) for examining an object (8), the probe comprising a first portion configured to be inserted into an opening of said object in a first direction (5), said probe further comprising at least one THz emitter (15), directing means (7) for directing THz radiation emitted from said emitter to said object via an aperture (2) located at said first portion and subsequently from said object to at least one THz detector (17) and, means for scanning said emitted THz radiation across said object in a scan direction, said scan direction having a component in said first direction (5). | 1. A THz radiation probe for examining an object, the probe comprising a first portion configured to be inserted into an opening of said object in a first direction, said probe further comprising at least one THz emitter,a directing unit for directing THz radiation emitted from said emitter to said object via an aperture located at said first portion and subsequently from said object to at least one THz detector and,a scanning unit for scanning said emitted THz radiation across said object in a scan direction, said scan direction having a component in said first direction,wherein the directing unit comprises an elongate solid dielectric member, the elongate direction of the elongate solid member being parallel to the first direction, the elongate solid dielectric member comprising a plane surface reflective to THz radiation angled to the first direction, the refractive index of the solid dielectric member and the plane surface being configured such that THz radiation emitted by the emitter undergoes reflection at the plane surface and is directed towards the aperture, the scanning of the THz beam across the plane surface causing the scanning of the THz beam across the aperture. 2. A THz radiation probe according to claim 1, wherein a part of said THz radiation probe comprising said first portion or a part thereof is configured to be detachable. 3. A THz radiation probe according to claim 1, wherein THz radiation that is reflected from said object is directed by said directing unit to said THz detector. 4. A THz radiation probe according to claim 1 wherein said emitted THz radiation is focused at said aperture. 5. A THz radiation probe according to claim 1 wherein said scanning unit comprises a pair of Risley prisms. 6. A THz probe according to claim 1 wherein said THz emitters are photoconductive antennas that comprise at least one dipole antenna having a resonant frequency in the THz frequency regime. 7. A THz radiation probe according to claim 1, wherein said planar surface is oriented such that the normal to said surface is at an angle of approximately 45° to said first direction. 8. A THz radiation probe according to claim 1 further comprising a sheath that substantially encloses said first portion. 9. A THz radiation probe according to claim 8, wherein said sheath comprises a material that is substantially transmissive to THz radiation. 10. A THz radiation probe according to claim 8, wherein said sheath comprises a region of material that is substantially transmissive to THz radiation, that is located adjacent to said aperture such that THz radiation may pass through said region. 11. A THz radiation probe according to claim 8, wherein said sheath or a part thereof is disposable. 12. A THz radiation probe according to claim 1, wherein THz radiation is emitted from said THz emitter is pulsed radiation. 13. A THz radiation probe according to claim 12, further comprising a time resolved detection unit. 14. A THz probe according to c... | TeraView Limited | G01J005/02 A61B005/00 G01N021/3581 A61B001/00
US09179847 | B2 | 2015-11-10 | System and method to provide career counseling and management using biofeedback | 2009-07-16 | An apparatus and method for measuring a person's biometric data as well as associated data and for using that data to determine the person's talents and well-being state, as well as predicting an optimal career path for the person. Biometric data is measured using a sensor, a memory configured to store the biometric signals, a database configured to store and retrieve profiles, and a processor configured to compare biometric data as well as associated data with anonymous profiles stored in the database and create a profile for the person. | 1. A system for creating biometric profiles comprising: a sensor configured to measure a person's biometric signals;memory configured to store the biometric signals;a database configured to store and retrieve biometric profiles; anda processor configured to process the biometric signals to produce a life fire graph, comprising a heat map that correlates an intensity of change in the biometric signals at different frequencies at different points in time, configured to compare the life fire graph to a plurality of biometric profiles stored in the database, and configured to evaluate at least one personality characteristic of the person based on correlations with the plurality of stored profiles. 2. The system of claim 1, wherein the measured biometric data includes Heart Rate Variability data. 3. The system of claim 1, wherein the measured biometric data includes skin conductivity. 4. The system of claim 1, wherein the measured biometric data includes brainwaves. 5. The system of claim 1, wherein the database includes a knowledge management system where at least one of competency, associated well-being state, and an associated performance rating of a plurality of users are stored. 6. The system of claim 1, wherein the processor includes an alert mechanism which is triggered when a mismatch occurs between a given person's biometric data and a job role. 7. The system of claim 1, wherein the sensor, memory, database, and processor are contained within a single, self-contained device that is small enough to be worn under a user's clothing. 8. A method for determining an optimal job for a given user comprising: creating a biometric profile for a given user, including the user's physical and mental characteristic traits, by processing biometric signals to produce a life fire graph, comprising a heat map that correlates an intensity of change in the biometric signals at different frequencies at different points in time;correlating the user's biometric profile with a plurality of known biometric profiles using a processor to determine job roles with a high probability of good performance and health for the user; andmatching the given user to a job role based upon the characteristic traits of the given user's biometric profile. 9. The method of claim 8, wherein the biometric data includes Heart Rate Variability information. 10. The method of claim 8, wherein the characteristic traits include at least one competency. 11. The method of claim 8, wherein the characteristic traits include the given user's well-being state. 12. The method of claim 11, wherein the information to determine the characteristic traits and the matching job role are retrieved from a knowledge management system where the characteristic traits and the associated performance rating of a significant number of users are stored. 13. The method of claim 12 wherein a competency includes information about personal characteristics such as motives, traits, self-image, social role, and ... | INTERNATIONAL BUSINESS MACHINES CORPORATION | A61B005/00 A61B005/024 A61B005/16 G06Q030/02
US09237868 | B2 | 2016-01-19 | Fitness test | 2009-05-04 | This document describes a fitness test application which allows the user to exercise freely (running and cycling outdoors, bicycle ergometer, rowing ergometer, treadmill) and which provides an estimate of user's fitness during and/or after the exercise. The analysis can be performed either as real-time or as post-analysis. | 1. A method for evaluating cardiorespiratory fitness of a user during an exercise, said evaluating being performed by a device including a processor, memory and software stored therein, and a user interface, the device being one of a heart rate monitor, a mobile phone, a PDA device, a wristop computer, and a personal computer, the method comprising the steps of: selecting exercise type and inputting user's physiological background parameters using the user interface of said device,during at least one exercise session, continuously collecting heart beat data comprising heart beat time series by an equipment suitable for obtaining heart beat data, and continuously collecting performance data by one of a satellite navigation system (GPS, Galileo), an accelerometer, bicycle ergometer, rowing ergometer, treadmill, and wristop computer,segmenting said collected heart beat data to data segments with coherent intensity by grouping continuous heart beat data having values that are within specified ranges,calculating reliability of the data segments based on a duration and variance of each data segment, and calculating weighting coefficients for the data segments based on their calculated reliability relative to a threshold reliability value, andobtaining an estimate of user's cardiorespiratory fitness level based on a linear relationship between heart rate and a performance output, wherein the heart rate is obtained from the heart beat data segments, the performance output is calculated from the performance data, and the calculated weighting coefficients are used in determining the linear relationship between the heart rate and the performance output. 2. A method as claimed in claim 1, wherein the heart beat data is a heart beat derivable parameter. 3. A method as claimed in claim 1, wherein the heart beat data can be inputted by the user. 4. A method as claimed in claim 1, wherein the heart beat data is rating of perceived exertion. 5. A method as claimed in claim 1, wherein the heart beat data is blood lactate level. 6. A method as claimed in claim 1, wherein the heart beat data is respiration rate. 7. A method as claimed in claim 1, wherein the heart beat data is ventilation. 8. A method as claimed in claim 1, wherein the data segments each have a duration of 20 s-10 min. | Firstbeat Technologies OY | A61B005/02 A61B005/22 A61B005/11
US07879314 | B2 | 2011-02-01 | Optical imaging of rheumatoid arthritis | 2006-06-28 | The present invention concerns the use of indocarbocyanine dyes, in particular indotricarbocyanine dyes for the diagnosis of inflammatory diseases, in particular rheumatoid arthritis, methods of diagnosing inflammatory diseases and an apparatus useful to carry out the method of diagnosing. | 1. A method for the diagnosis of an inflammatory disease of a joint comprising (i) administering to a subject a diagnostically effective amount of (a) an indocarbocyanine dye of formula (I)wherein C stands for a radical (II) or (III)wherein the position that is labelled with the star means the point of linkage with radical A and wherein A stands for a group (IV), (V), (VI), (VII), or (VIII)whereinR1 and R2, independently of one another, stand for a C1-C4-sulfoalkyl chain or a saturated or unsaturated, branched or straight-chain C1-C50-alkyl chain, which optionally is substituted by 0 to 15 oxygen atoms and/or by 0 to 3 carbonyl groups and/or with 0 to 5 hydroxyl groups or optionally interrupted by 0 to 15 oxygen atoms and/or by 0 to 3 carbonyl groups and/or can be substituted with 0 to 5 hydroxyl groups;R3 stands for a straight-chain carbohydrate chain with up to 20 carbon residues, which is substituted with one or more —OH, —COOH, —SO3 groups and/or optionally interrupted one or more times by —O—, —S—, —CO—, —CS—, —CONH, —NHCO—, NHCSNH—, —SO2—, —PO431 —, -aryl- and/or —NH— group;R4 stands for the group —COOE1, —CONE1E2, —NHCOE1, —NHCONHE1, —NE1E2, —OE1, —OSO3E1, —SO3E1, —SO2NHE1 or -E1, whereinE1 and E2, independently of one another, stand for a hydrogen atom, a C1-C4-sulfoalkyl chain, a saturated or unsaturated, branched or straight-chain C1-C50-alkyl chain, which optionally is interrupted by 0 to 15 oxygen atoms and/or by 0 to 3 carbonyl groups and/or is substituted with 0 to 5 hydroxyl groups;R5 stands for a hydrogen atom, or a fluorine, chlorine, bromine or iodine atom, methyl, ethyl, propyl or iso-propyl;b means the number 2 or 3; and X and Y, independently of one another, stand for O, S, ═C(CH3)2 or—(CH═CH)—,or a salt or solvate of thereof or(b) an indocarbocyanine dye which is Cy7, indocyanine green (ICG), an analog of ICG, or indotricarbocyanine (ITC),wherein no targeting moiety is attached to the dye (a) or (b),(ii) exposing the joint to be imaged to light of a wavelength between 700 nm and 900 nm and(iii) measuring fluorescence at multiple time points within 500 seconds after administration of said dye,wherein the distribution of said indocarbocyanine dye in the joint regions suspected of having inflammation is indicative of the presence of inflammatory joint disease. 2. A method according to claim 1, wherein the inflammatory disease is arthritis. 3. A method according to claim 2, wherein the arthritis is selected from the group consisting of rheumatoid arthritis, osteoarthritis, psoriatic arthritis, traumatic arthritis, bacterial arthritis, post-infectious arthritis, lyme disease (Borreliose), ankylosing spondyliitis and rubella arthritis. 4. A method according to claim 1, wherein the indocarbocyanine dye is administered in an amount of 10 mg/kg body weight or less. 5. A method of diagnosing an inflammatory joint disease comprising: (a) administering to a subject a diagnostically effective amount of an indocarbocyanine dye, by bol... | Bayer Schering Pharma AG | A61B005/00 A61B008/00 A61B010/00 A61K049/00
US08303500 | B2 | 2012-11-06 | Prescription zero: a non-pharmaceutical prescription device for prescribing, administering, monitoring, measuring and motivating a therapeutic lifestyle regimen for prevention and treatment of chronic diseases | 2009-08-21 | The invention discloses a wearable/handheld personal communication device with hardware and software sensor modules that sense and analyze all caregiver prescribed/monitored user-lifestyle activities, and deploys such analysis in improving user's overall health in terms of reduced risks for all-cause morbidities/mortalities and eventually a life without drugs. Termed Rx Zero, such method may be prescribed not just for maintaining a healthy lifestyle, but for treatment of chronic diseases with intent to wean the patients to minimal or zero pharmacological intervention, or in combination with medications to improve prognosis of the disease under treatment.;The benefits of the Rx Zero method of the present invention extend not only to the individual and the community through high quality healthcare at lower cost, but payers by reducing the loss ratio on account of reduced cost of medical claims, and to the caregivers in terms of an effective tool that disseminates, implements and redefines “Primary Health Care” and “Prevention” at levels beyond the terms' currently understood scope that has transformed healthcare to sickcare. | 1. A non-pharmaceutical method and prescription device for achieving a goal of zero drug intervention (Rx Zero) in treatment and prevention of chronic diseases by means of prescribing, administering, measuring, quantifying, monitoring, motivating and implementing a lifestyle behavior modification approach involving therapeutic physical activity and related dietary and behavioral lifestyle changes resulting in consequent health benefits in terms of risk reduction for future morbidities and mortalities, prognosis for current morbidities, reduction, weaning out, or elimination of chronic pharmaceutical intervention, and rewards accruing for achieving a prescribed lifestyle goal, comprising of: a) either a wearable or a handheld communication device or a personal user device, which may be a mobile wearable computer, a PDA or a mobile phone that is capable of performing, recording and displaying data regarding user's free-living lifestyle usage of functions pertaining to telecommunication, surfing Internet, playing music, power management, recording physical movements, and recording periods of inactivity that occurs during user's sleep or during a period the device is connected to external power source for charging;b) a sensor module comprising of the following two sensor sub-modules: i. a first sensor sub-module wherein the said sensor is a physical activity sensing MEMS (Micro Electro Mechanical System) single, double or triple axis accelerometer hardware sensor and a device activity sensing software sensor application either integrated within a hand held communication device or a mobile wearable computer or a PDA or a mobile phone, or resident in a disparate wireless device attached to user's body, or to user's apparel, or to a personal communication device carrying case, or to an exercise apparatus such as a bicycle or a treadmill, in each of which case the first sensor sub-module communicates wirelessly with a mobile wearable computer, or a handheld communication device, or a PDA or a mobile phone using close range radiofrequency transmission, and,ii. a second sensor sub-module wherein the said sensor is a software client server application that monitors user's usage data regarding different functions of a mobile wearable computer, or a handheld communication device, or a PDA or a mobile phone, that is analyzed to provide the user's free-living lifestyle that includes sleep patterns, physical activity or inactivity patterns, socializing patterns, mood patterns, and also to provide greater accuracy of accelerometer recordings by switching the accelerometer to sleep mode when the device is connected to the external power source for charging or when the user goes to sleep.c) a distributed sensor data analysis module, which is a client-server application that has a client-resident local analysis sub-module and a server resident global analysis sub-module operating in tandem to analyze free-living sensing data generated by the sensor module, and outpu... |  | A61B005/00 A61B005/103 A61B005/117 G06Q010/00 G06Q040/00 G06Q050/00
US20110009761 | A1 | 2011-01-13 | Method and system for real-time profiling respiratory motion with variable-horizon prediction capability | 2010-07-09 | A method of profiling respiratory motion is provided that includes estimating a temporal respiratory pattern, using a warping function to map the temporal respiratory pattern to a corresponding phase value, using a baseline drift function to determine drift in the temporal respiratory pattern, and noise filtering the temporal respiratory pattern, where variations in a respiratory motion are provided. The warping function includes combining an elliptical shape prior for providing iso-phase events in real-time, where the elliptical shape prior is in an augmented state space and Poincare´ sectioned. Parameters of the ellipse are estimated, where a projection of a center of the ellipse onto an observed respiratory position provides a real-time estimation for baseline drift of the respiratory motion. | 1. A method of profiling respiratory motion, comprising: a. estimating a temporal respiratory pattern;b. using a warping function to map said temporal respiratory pattern to a corresponding phase value;c. using a baseline drift function to determine drift in said temporal respiratory pattern; andd. noise filtering said temporal respiratory pattern, wherein variations in a respiratory motion are provided. 2. The method of claim 1, wherein said estimate of said temporal respiratory pattern is obtained by applying an inverse-warping function to respiratory observations and performing a minimum mean-squared error (MMSE) estimation to results of said inverse-warping. 3. The method of claim 1, wherein said warping function comprises combining an elliptical shape prior for providing iso-phase events in real-time, wherein said elliptical shape prior is in an augmented state space and Poincare' sectioned. 4. The method of claim 3, wherein parameters of said ellipse are estimated, wherein a projection of a center of said ellipse onto an observed respiratory position provides a real-time estimation for baseline drift of said respiratory motion. 5. The method of claim 1, wherein an observed breathing trajectory comprises a periodic fundamental pattern, wherein said periodic fundamental pattern is frequency modulated by said phase warping function. 6. The method of claim 1, wherein an observed breathing trajectory comprises a periodic fundamental pattern, wherein said periodic fundamental pattern is displacement modulated by said baseline drift function. 7. The method of claim 1, wherein variations in a respiratory baseline, a respiratory frequency and a fundamental respiratory pattern are decoupled. 8. The method of claim 1, wherein respiratory baseline, a respiratory frequency and a fundamental respiratory pattern are filtered to achieve denoising, wherein said denoising comprises a smoothness assumption. 9. The method of claim 1, wherein respiratory baseline, a respiratory frequency and a respiratory fundamental pattern are assembled back into an original signal space to obtain respiratory prediction results. 10. The method of claim 1, wherein variation in said respiratory motion is modeled as the composite result of baseline drift in said respiratory pattern, frequency variation in said respiratory pattern, fundamental change in said respiratory pattern and additive random observation noise in said respiratory pattern. 11. The method of claim 1, wherein phase estimation comprises augmentation of respiratory pattern observations, ellipse fitting said augmented respiratory pattern observations, Poincare' sectioning said augmented respiratory pattern observations, mapping said Poincare' intersections back to a time-displacement relation as iso-phase points, and i) providing linear interpolation, ii) linear extrapolation or i) and ii) to generate a continuous phase estimation. 12. The method of claim 11, wherein an elliptical shape is combined prio... | - | A61B005/08
US20110028831 | A1 | 2011-02-03 | PERMANENTLY VISIBLE IMPLANTABLE FIDUCIARY TISSUE MARKER | 2009-07-30 | A fiduciary marker includes a hollow cylindrical body comprised of a bio-compatible polymer and including a first end, a second end, an inner surface extending between the first end to the second end and an outer surface extending between the first end to the second end. A plurality of apertures are formed in the body and extend from the inner surface to the outer surface. | 1. A fiduciary marker comprising: a hollow cylindrical body comprised of a bio-compatible polymer and including: a first end;a second end;an inner surface extending between the first end to the second end;an outer surface extending between the first end to the second end; anda plurality of apertures formed in the body and extending from the inner surface to the outer surface. 2. The fiduciary marker of claim 1, wherein the plurality of apertures are formed in a plurality of rows about the hollow cylindrical body and extending between the first end to the second end of the body. 3. The fiduciary marker of claim 2, wherein each row comprises at least two apertures. 4. The fiduciary marker of claim 1, wherein the first and the second ends of the hollow cylindrical body comprise an outer diameter defined by the outer surface and an inner diameter defined by the inner surface such that the first end and the second end include openings to an interior of the hollow cylindrical body. 5. The fiduciary marker of claim 1, wherein the first and the second ends of the hollow cylindrical body comprise an outer diameter defined by the outer surface such that the first and the second end enclose an interior of the hollow cylindrical body. 6. The fiduciary marker of claim 1, further comprising a bead enclosed in the cylindrical hollow body. 7. The fiduciary marker of claim 6, wherein the bead is a spherical bead comprising a diameter that is greater than a diameter of the plurality of apertures formed in the body. 8. The fiduciary marker of claim 6, wherein the bead comprises a glass bead. 9. The fiduciary marker of claim 6, wherein the bead comprises a metallic bead. 10. The fiduciary marker of claim 1, further comprising a metal bar extending at least partially within the interior of the hollow cylindrical body along the inner surface at least between the first and the second end. 11. A fiduciary marker comprising: a cylindrical body comprised of a bio-compatible polymer compounded with a radiopaque material and having a first end, a second end and an outer surface. 12. The fiduciary marker of claim 11, wherein the radiopaque material comprises a metal oxide. 13. The fiduciary marker of claim 11, wherein the cylindrical body comprises a hollow cylindrical body having a first end, a second end, an inner surface extending between the first end to the second end, an outer surface extending between the first end to the second end, and a plurality of apertures formed in the body and extending from the inner surface to the outer surface, wherein the metal oxide comprises barium sulfate. 14. The fiduciary marker of claim 13, wherein the first and the second ends of the hollow cylindrical body comprise an outer diameter defined by the outer surface and an inner diameter defined by the inner surface such that the first end and the second end include openings to an interior of the hollow cylindrical body. 15. The fiduciary marker of claim 13, wherein the fi... | - | A61B005/05
US07894872 | B2 | 2011-02-22 | Computer assisted orthopaedic surgery system with light source and associated method | 2005-12-26 | A computer assisted orthopaedic surgery system comprises an image sensor and a light source. The image sensor is configured to sense an image of an area associated with an orthopaedic surgical site and has a field of view. The light source projects light distinct from any ambient lighting onto the area in a pattern visually indicative of at least a portion of the field of view. An associated method is disclosed. | 1. A method of operating a computer assisted orthopaedic surgery system, the method comprising: detecting an unlocking of an image sensor for movement relative to a base to which the image sensor is secured, andprojecting light from a light source distinct from any ambient lighting onto an area associated with an orthopaedic surgical site in a pattern visually indicative of a field of view of the image sensor configured to sense an image of the area, wherein the pattern is coterminous with the field of view of the image sensor, wherein projecting light comprises operating the light source in response to detection of the unlocking of the image sensor. 2. The method of claim 1, wherein projecting light comprises projecting the light onto the area in a linear pattern. 3. The method of claim 2, wherein projecting light comprises (i) emitting light from a light generator and (ii) modifying the emitted light into the linear pattern with an optical device. 4. The method of claim 1, wherein projecting light comprises operating a laser or a laser diode. 5. The method of claim 1, further comprising moving the image sensor, wherein projecting light comprises generating the light during movement of the image sensor. 6. A computer assisted orthopaedic surgery system, comprising: a base,an image sensor moveably coupled to the base, the image sensor being configured to sense an image of an area associated with an orthopaedic surgical site, the image sensor having a field of view,a user control configured to unlock the image sensor for movement relative to the basea light source to project light distinct from any ambient lighting onto the area in a pattern coterminous with the field of view, anda controller electrically coupled to the image sensor and the light source, the controller comprising (i) a processor, and (ii) a memory device electrically coupled to the processor, the memory device having stored therein a plurality of instructions which, when executed by the processor, cause the processor to:detect an unlocking of the image sensor for movement relative to the base, andoperate the light source to project light distinct from any ambient lighting onto the area in the pattern so as to visually indicate the field of view of the image sensor in response to detection of the unlocking of the image sensor. 7. The computer assisted orthopaedic surgery system of claim 6, wherein: the image sensor comprises a stereoscopic camera head, andthe plurality of instructions, when executed by the processor, cause the processor to operate the light source so as to generate a visual indication of the field of view of the stereoscopic camera head. 8. The computer assisted orthopaedic surgery system of claim 6, further comprising a second user control electrically coupled to the controller, wherein the plurality of instructions, when executed by the processor, cause the processor to operate the light source in response to operation of the second user control. 9. The c... | DePuy Orthopaedics, Inc | A61B005/00
US08334693 | B2 | 2012-12-18 | NMR device for detection of analytes | 2010-07-27 | This invention relates generally to detection devices having one or more small wells each surrounded by, or in close proximity to, an NMR micro coil, each well containing a liquid sample with magnetic nanoparticles that self-assemble or disperse in the presence of a target analyte, thereby altering the measured NMR properties of the liquid sample. The device may be used, for example, as a portable unit for point of care diagnosis and/or field use, or the device may be implanted for continuous or intermittent monitoring of one or more biological species of interest in a patient. | 1. A device for the magnetic resonance detection of one or more analytes in a liquid sample, the device comprising a support defining a well holding a liquid sample comprising magnetic particles and the one or more analytes and having an RF coil disposed about the well, the RF coil configured to detect an echo response produced by exposing the liquid sample to a bias magnetic field created using one or more magnets and an RF excitation, wherein the magnetic particles have binding moieties on their surfaces, the binding moieties operative to alter an aggregation of the magnetic particles in the presence of the one or more analytes, and wherein at least one of the binding moieties is operative to bind to at least one of the following: (i) the one or more analytes; (ii) another of the binding moieties; and (iii) an aggregation-inducing molecule in the liquid sample, thereby producing an aggregate of multiply-linked magnetic particles as a function of the presence or concentration of the one or more analytes in the liquid sample. 2. The device of claim 1, wherein the well and the RF coil are configured to provide a filling factor of at least about 0.1. 3. The device of claim 1, wherein the well has a volume of less than about 300 μL. 4. The device of claim 1, wherein the well and the RF coil are configured such that the volume circumscribed by the RF coil is less than about 300 μL. 5. The device of claim 1, wherein the RF coil has a characteristic dimension from about 10 μm to about 1000 μm. 6. The device of claim 1, further comprising a tube for holding the liquid sample, the tube having a varying cross section. 7. The device of claim 1, wherein an aggregate of multiply-linked magnetic particles is disaggregated as a function of the presence or concentration of the one or more analytes in the liquid sample. 8. The device of claim 1, further comprising a plurality of compartments for holding one or more populations of magnetic particles, wherein magnetic particles having different binding moieties are disposed in different compartments for detection of the one or more analytes, the different binding moieties promoting aggregation or disaggregation of the magnetic particles in the presence of the one or more analytes. 9. The device of claim 8, wherein magnetic particles having different binding moieties are disposed in different compartments for detection of multiple analytes. 10. The device of claim 8, wherein magnetic particles having the same binding moieties are disposed in different compartments for replicate measurements. 11. The device of claim 8, wherein magnetic particles having the same binding moieties are disposed in different compartments for detection of varying analyte concentrations in the liquid sample. 12. The device of claim 1, wherein the bias magnetic field has a strength of from about 0.45 Tesla to about 7 Tesla. 13. The device of claim 1, wherein the device is implantable. 14. The device of claim 1, wherein the devi... | T2 Biosystems, Inc. | G01V003/00
US07894887 | B2 | 2011-02-22 | Biosignal measuring equipment | 2008-03-26 | A signal-to-noise ratio and measurement precision is increased in electrode units disposed on the left and right sides of a steering wheel. A plurality of electrode units is disposed in the left and right handholds of a steering wheel. The contact impedances of all the electrode units are measured. A pair of left and right electrode units to be used to measure an electrocardiographic signal is designated from among the electrode units whose measured contact impedances are less than or equal to a first threshold. The results of measurement of an electrocardiographic signal by the designated electrode units are added in order to minimize noise. An electrode unit with high contact impedance is used to measure induction noise and remove the induction noise component from the electrocardiographic signal measurement result. | 1. Biosignal measuring equipment comprising: a plurality of electrode units disposed in left and right positions of a steering wheel; anda measuring means for measuring a biosignal associated with a subject using designated ones of the plurality of electrode units, the subject capable of holding the steering wheel using the left and right positions,wherein the measuring means measures a plurality of biosignals using the designated ones of the plurality of electrode units and determines final information on the basis of the plurality of measured biosignals, whereineach of the plurality of electrode units includes a first electrode and a second electrode; andthe measuring means includes: a differential amplification means for producing a biosignal by amplifying a difference between two input signals;an impedance measurement means for measuring an impedance between two electrodes;a switching circuit unit that connects the first and second electrodes of the each of the each of the plurality of electrode units to one of the differential amplification means, the impedance measurement means, and a reference potential; andan arithmetic control means for measuring a plurality of biosignals from the subject by connecting the electrodes of each of the electrode units to the one of the differential amplification means, impedance measurement means, and the reference potential via the switching circuit unit, and for deducing a final complex. 2. The biosignal measuring equipment according to claim 1, wherein the plurality of electrode units are disposed in the left and right positions such that a left palm and a right palm of the subject simultaneously come into contact with multiple ones of the plurality of electrode units when the subject holds the steering wheel. 3. The biosignal measuring equipment according to claim 2, wherein left ones and right ones of the plurality of electrode units are disposed within respective left and right circular areas having a diameter is 15 cm or less, around the left and right positions respectively. 4. The biosignal measuring equipment according to claim 2, wherein: the plurality of electrode units disposed near left and right joint portions where a ring portion and left and right spoke portions of the steering wheel are joined, the plurality of electrode units disposed in the circumferential direction of the ring portion; andat least one of the plurality of electrodes is formed to protrude from the ring portion toward the spoke portion. 5. The biosignal measuring equipment according to claim 1, wherein the measuring means is mounted in the steering wheel. 6. The biosignal measuring equipment according to claim 1, wherein the measuring means includes a plurality of differential amplification means, a number of the differential amplification means corresponding to a number of the plurality of electrode units disposed in each of the handholds. 7. The biosignal measuring equipment according to claim 6, wherein the measuri... | Denso Corporation | A61B005/04
US09028406 | B2 | 2015-05-12 | Oral-esophageal-gastric device to diagnose reflux and/or emesis | 2010-09-09 | A system and method diagnoses esophageal reflux in a patient. A catheter includes a first pressure sensor configured to measure intra-abdominal pressure when the catheter is inserted into the stomach of a patient. A second pressure sensor is located proximal to the distal tip and configured to measure intra-thoracic pressure. A pH indicator is carried by the catheter body and configured to indicate that reflux has occurred in response to an involuntary reflex cough test. A processing device is configured to receive data regarding the first and second pressure sensors and regarding the pH indicator and process that data and assess if reflux is present during the involuntary epoch. | 1. A system for diagnosing esophageal reflux in a patient, comprising: a catheter comprising a first pressure sensor configured to measure intra-abdominal pressure when the catheter is inserted into the stomach of a patient and a second pressure sensor located proximal to the distal tip and configured to measure intra-thoracic pressure and a pH indicator carried by the catheter body and configured to indicate that reflux has occurred in response to an involuntary reflex cough event;a nebulizer containing an agent that induces the involuntary cough event within the patient;at least one electromyogram (EMG) pad configured to be attached to the lumbar region of the patient's back and obtain an EMG from involuntary cough activated paraspinal muscles; anda processing device configured to receive data regarding the first and second pressure sensors and the pH indicator and the EMG pad and process said data and assess if reflux is present during the involuntary event. 2. The system according to claim 1, wherein said processing device comprises a portable handheld device comprising; a housing configured for handheld use;at least one interface carried by the housing and configured to receive the pressure measurements obtained during the involuntary cough reflex event; anda processor carried by the housing and connected to the interface and configured to receive the pressure measurements obtained during the involuntary cough reflex event. 3. The system according to claim 1, wherein said catheter comprises an air charged catheter. 4. A method for diagnosing esophageal reflux in a patient, comprising: inserting a nasogastric/orogastric (Ng/Og) catheter into the stomach of a patient;inducing an involuntary reflex cough epoch within the patient;measuring intra-abdominal and intra-thoracic pressure during the involuntary reflex cough epoch;determining a pressure gradient over the lower esophageal sphincter (LES) during the involuntary reflex cough epoch;obtaining an electromyogram (EMG) from involuntary cough activated paraspinal muscles during the involuntary reflex cough epoch;assessing if reflux is present during the involuntary epoch; andassessing, by a processing device, the severity of the reflux compared to the response of the involuntary cough test and the magnitude of the involuntary cough epoch to stratify reflux severity and determine a course of treatment. 5. The method according to claim 4, and further comprising placing the catheter by measuring from the lips or nares to the temporomandibular joint to a predetermined distance sub-xyphoid in an adult. 6. The method according to claim 4, wherein said Ng/Og catheter comprises a catheter body having a distal tip and a first pressure sensor located at the distal tip and configured to measure intra-abdominal pressure when the catheter is inserted into the stomach of a patient, and a second pressure sensor located proximal to the distal tip and configured to measure intra-thoracic pressure. 7. The ... | Pneumoflex Systems, LLC | G06F019/00 A61B005/00 A61B005/20 A61B005/0488 A61B005/145
US08597184 | B2 | 2013-12-03 | Techniques for evaluating urinary stress incontinence and use of involuntary reflex cough as a medical diagnostic tool | 2010-09-09 | A system permits diagnosis of a patient for a physiological abnormality while protecting their airway. An esophageal airway protection device comprises an elongate device body having a distal end for insertion into the stomach through the esophagus and a proximal end. The device includes a main lumen extending the length of the device and an inflatable esophageal cuff carried by the device body mid-esophagus. Emesis and/or reflux is blocked from passing out of the stomach past the esophageal cuff positioned mid-esophagus when it is inflated to protect a patient's airway during an involuntary cough event. At least one electromyogram (EMG) pad is configured to obtain an EMG from an involuntary cough activated paraspinal muscles. A processing device is configured to receive the EMG and process the EMG to determine a physiological abnormality. | 1. A system of diagnosing a patient for a physiological abnormality while protecting their airway, comprising: an esophageal airway protection device that comprises an elongate device body having a distal end for insertion into the stomach through the esophagus and a proximal end, and having a main lumen extending the length of the device and an inflatable esophageal cuff carried by the device body mid-esophagus and an inflation lumen formed within the device body and connecting the inflatable esophageal cuff through which the esophageal cuff is inflated and deflated, wherein upon inflation of said esophageal cuff, emesis and/or reflux is blocked from passing out of the stomach past the esophageal cuff positioned mid-esophagus to protect a patient's airway during an involuntary cough event;a nebulizer lumen extending along the device body and comprising a port through which medication is delivered for administrating the involuntary cough event, wherein said esophageal cuff protects the airway during the involuntary cough event;at least one electromyogram (EMG) pad configured to obtain an EMG from involuntary cough activated paraspinal muscles; anda processing device configured to receive the EMG and process the EMG to determine a physiological abnormality. 2. The system according to claim 1, and further comprising a nebulizer venturi connecting the nebulizer lumen and configured to deliver nebulized medication around the device body. 3. The system according to claim 1, and further comprising at least one pressure sensor located on the device body and configured to measure intra-abdominal pressure (IAP), wherein said processing device receives the IAP and correlates the IAP and EMG and involuntary reflex cough epoch. 4. The system according to claim 3, wherein said at least one pressure sensor comprises a pressure transducer and a transducer lead connecting the pressure transducer. 5. The system according to claim 3, and further comprising at least one pH sensor carried by the device body, and said processing device receives pH measurements and processes same with EMG and IAP. 6. The system according to claim 3, wherein the said processing device comprises a portable handheld device comprising, a housing configured for handheld use; at least one interface carried by the housing and configured to receive the pressure measurements obtained during the involuntary cough reflex test; and a processor carried by the housing and connected to the interface and configured to receive the IAP and EMG. 7. The system according to claim 1, and further comprising a sump port at the distal end and sump lumen formed the length of the device body and configured for venting gas and preventing adherence of the device against the gastric wall. 8. The system according to claim 1, wherein the esophageal cuff is radio-opaque to aid in placing the esophageal cuff at a desired location within the esophagus. 9. A method of diagnosing a patient for a physiological abn... | Pneumoflex Systems, LLC | A61B005/00
US08391941 | B2 | 2013-03-05 | System and method for memory switching for multiple configuration medical sensor | 2009-07-17 | A system is provided that includes a monitor, a multiple configuration sensor having a first configuration and a second configuration, and a sensor connector. The sensor connector includes a first memory device and a second memory device, such that the first memory device is accessible by the monitor in the first sensor configuration and the second memory device is accessible by the monitor in the second sensor configuration. Another system is provided that includes a sensor adaptor having a processing circuit that reads a memory of a multiple configuration sensor and provides a first set of calibration data of a first sensor configuration and a second set of calibration data for a second sensor configuration. | 1. A system, comprising: a monitor;a sensor configured to be coupled to the monitor, the sensor comprising: a sensor body being configurable between a first configuration adapted to be applied to a first tissue site and a second configuration adapted to be applied to a second tissue site, wherein the sensor is configured to measure a physiological characteristic and wherein the second configuration results from separating a removable portion from the sensor body; anda component coupled to the sensor, comprising: a first memory device storing a first set of calibration data;a second memory device storing a second set of calibration data;wherein only the first memory device is accessible by the monitor in the first configuration and only the second memory device is accessible by the monitor in the second configuration. 2. The system of claim 1, wherein the monitor comprises a pulse oximetry monitor. 3. The system of claim 1, wherein the component comprises a sensor connector or a sensor cable. 4. The system of claim 1, wherein the sensor comprises an indicator circuit configured to indicate the first configuration or the second configuration. 5. The system of claim 4, wherein the indicator circuit comprises a conductive trace, wherein the conductive trace is intact in the first configuration and broken in the second configuration. 6. The system of claim 4, wherein the component comprises a switch configured to switch between the first memory device and the second memory device based on the indicator circuit. 7. The system of claim 4, wherein the component comprises a resistor, capacitor, inductor, diode, or other electronic component or combination thereof configured to block access to the second memory device in the first configuration and to permit access to the second memory device in the second configuration. 8. The system of claim 4, wherein the second memory device is connected to a ground connection in the first configuration and the second memory device is isolated from the ground connection in the second configuration. 9. The system of claim 1, wherein the component comprises a processing circuit configured to provide the first set of calibration data to the monitor in the first configuration and provide the second set of calibration data to the monitor in the second configuration. 10. A medical sensor, comprising: a sensor body;a circuit operably coupled to the sensor body comprising: a first memory device; anda second memory device, wherein the circuit is configured to make the first memory device electrically accessible only when the sensor body is in a first configuration and configured to make the second memory device electrically accessible only when the sensor body is in a second configuration, wherein the second configuration results from separating a removable portion from the sensor body. 11. The medical sensor of claim 10, wherein the circuit comprises a switch configured to make the first memory device accessible only... | Covidien LP | A61B005/1455
US08718957 | B2 | 2014-05-06 | Method and devices for monitoring flow circuits | 2009-04-16 | A device is provided for monitoring the integrity of a flow circuit in fluid communication with a receptacle. The flow circuit includes a pumping device for transferring fluid through the flow circuit. The device operates according to a monitoring method in which a pressure signal is received from a pressure sensor, the pressure signal being indicative of fluid pressure in the receptacle or the flow circuit. The pressure signal is then processed for detection of a beating signal. The beating signal manifests itself as an amplitude modulation of the pressure signal and is formed by interference between pressure waves generated by a pulse generator associated with the receptacle and pressure waves generated by the pumping device. The integrity of the flow circuit is determined based at least partly on the presence or absence of the beating signal. The device and the flow circuit may be part of an apparatus for extracorporeal blood treatment, and the method may be implemented as a computer program product. | 1. A method for monitoring the integrity of a flow circuit in fluid communication with a receptacle, said flow circuit comprising a pumping device configured to transfer fluid through the flow circuit, said method comprising: receiving a pressure signal from a pressure sensor, said pressure signal being indicative of fluid pressure in the receptacle or the flow circuit,analyzing the pressure signal for a beating signal formed by interference between pressure waves generated by a pulse generator associated with the receptacle and pressure waves generated by the pumping device, anddetermining the integrity of the flow circuit based at least partly on the presence or absence of the beating signal. 2. The method of claim 1, wherein the beating signal is detected by analyzing the pressure signal in a time domain. 3. The method of claim 1, further comprising obtaining one or more specific frequencies related to a pumping frequency of the pumping device, and wherein said analyzing comprises creating at least one filtered pressure signal in which all but one of said specific frequencies are removed. 4. The method of claim 3, wherein said at least one specific frequency comprises one or more of: half the pumping frequency, the pumping frequency, and harmonics of the pumping frequency. 5. The method of claim 3, wherein said analyzing comprises determining an envelope of the filtered pressure signal. 6. The method of claim 5, wherein determining the envelope comprises extracting an array of temporally sequential peak values from the filtered pressure signal. 7. The method of claim 5, wherein said analyzing further comprises calculating at least one of a sum of derivatives and a variance based on the envelope. 8. The method of claim 3, further comprising: matching at least part of the filtered pressure signal against one or more predetermined signal patterns so as to detect the beating signal. 9. The method of claim 3, wherein said obtaining comprises at least one of analyzing the pressure signal in a frequency domain to identify said one or more specific frequencies, deriving a frequency measurement signal from the pumping device, and deriving a set value of a controller configured to control the pumping frequency of the pumping device. 10. The method of claim 1, further comprising: processing the pressure signal for detection of a signal component generated by the pulse generator, wherein said determining the integrity of the flow circuit is also based on the presence or absence of said signal component. 11. The method of claim 10, wherein said analyzing of the pressure signal for detection of a beating signal is conditioned upon absence of said signal component in the pressure signal. 12. The method of claim 1, further comprising: causing, in the absence of a beating signal, a predetermined change in the pumping frequency of the pumping device. 13. The method of claim 1, further comprising: causing the pumping device to be temporarily inactiva... | Gambro Lundia AB | G01F017/00
US08430547 | B2 | 2013-04-30 | Compact motion-simulating device | 2009-08-03 | A reaction-testing system is provided to simulate motion to an individual. Generally, the reaction-testing system includes a motion-simulating device, a reflective device, a control unit, and an input device that is activated by an individual In exemplary embodiments, the motion-simulating device includes a beam that accommodates a first and second plurality of light sources capable of, at least, alternating between an active and idle condition. These light sources are arranged amongst occlusion features in a mounting pattern such that light emitted from the first plurality is directly viewable by an individual and light emitted from the second plurality is indirectly viewed by the individual through the reflecting surface. The control unit is configured to sequentially activate the light sources sequentially to generate pulses of light on that, when perceived by the individual, appear as a light moving along a vector. | 1. A motion-simulating device that simulates motion to an individual viewing the motion-simulating device from a first position relative to the motion-simulating device, the motion-simulating device comprising: at least one reflecting surface viewable from the first position;a first plurality of light sources directly viewable from the first position;a second plurality of light sources longitudinally offset from the first plurality of light sources, the second plurality of light sources indirectly viewable from the first position using the at least one reflecting surface; anda beam that holds the first plurality of light sources and the second plurality of light sources between the at least one reflecting surface and the first position,wherein the light sources in the first plurality of light sources and the second plurality of light sources activate in a temporal pattern to create the appearance of movement of a light from the first position. 2. The motion-simulating device of claim 1, wherein the first plurality of light sources are not indirectly viewable from the first position using the at least one reflecting surface. 3. The motion-simulating device of claim 1, wherein the second plurality of light sources are not directly viewable from the first position. 4. The motion-simulating device of claim 1, wherein only one light source of the first plurality of light sources and the second plurality of light sources is activated at any given time. 5. The motion-simulating device of claim 4, wherein: the first plurality of light sources are not indirectly viewable from the first position using the at least one reflecting surface; andthe second plurality of light sources are not directly viewable from the first position. 6. The motion-simulating device of claim 5, wherein the at least one reflecting surface comprises a mirror. 7. The motion-simulating device of claim 5, wherein: the first plurality of light sources comprise display devices that display indicia having at least one visual characteristic; andthe second plurality of light sources comprise display devices that display indicia having at least one visual characteristic. 8. The motion-simulating device of claim 5, wherein the beam extends longitudinally between the at least one reflecting surface and the first position. 9. The motion-simulating device of claim 5, wherein the beam extends laterally between the at least one reflecting surface and the first position. 10. The motion-simulating device of claim 1, further comprising: an input device to be engaged by the individual at the first position to indicate when the individual believes that light at a predetermined position is activated. 11. The motion-simulating device of claim 10, further comprising: a secondary input device that receives an input from the individual indicating whether the individual believes they engaged the input device before, after, or while light at a predetermined position was activated. 12. The motion-s... | Nike, Inc. | F21V099/00
US08655441 | B2 | 2014-02-18 | Methods and apparatus for monitoring patients and delivering therapeutic stimuli | 2010-07-13 | In an exemplary implementation of this invention, a user wears comfortable biosensors. These sensors gather physiological data from the user and transmit this data to another radio-enabled device. This other device may be a mobile phone. The data is further transmitted, via this other radio-enabled device, to one or more networks (such as wireless networks or the Internet). A processor analyzes this transmitted data, in real time, to recognize patterns in the data that indicate the need for therapeutic intervention. Upon recognition of such a pattern, the processor outputs instructions for a transducer to deliver therapeutic stimuli. These instructions are transmitted, over one or more wired or wireless networks, to a transducer which delivers the therapeutic stimuli to the user. | 1. A system comprising, in combination: (a) at least one wearable biosensor configured (i) for measuring physiological data and (ii) for wirelessly transmitting the data to a hub for direct or indirect transmission to a web server, which server is included in a set of one or more servers;(b) the one or more servers;wherein (i) the at least one biosensor includes a wireless module,(ii) the one or more servers are configured (A) to use a pattern recognition algorithm to recognize patterns in the data, and(B) based on a pattern in the data recognized by the algorithm, to output instructions for one or more transducers to produce psychologically therapeutic stimuli,(iii) the psychologically therapeutic stimuli comprise auditory, visual or audiovisual stimuli, and(iv) the one or more servers are further configured to perform machine learning to repeatedly update classifiers in a pattern recognition algorithm, which machine learning includes using data from the biosensor for training. 2. The system of claim 1, wherein the hub comprises a mobile computing device. 3. The system of claim 2, wherein the one or more transducers are included in the mobile computing device. 4. The system of claim 1, wherein said pattern is indicative of a drug craving or panic state. 5. The system of claim 1, wherein the psychologically therapeutic stimuli comprise a verbal message. 6. The system of claim 1, wherein the psychologically therapeutic stimuli comprise audio or visual content. 7. The system of claim 1, wherein the psychologically therapeutic stimuli comprise an SMS message or MMS message. 8. The system of claim 1, wherein the one or more transducers are wearable. 9. The system of claim 1, further comprising the one or more transducers. 10. The system of claim 1, wherein the pattern recognition algorithm comprises a Dynamic Bayesian Network. 11. The system of claim 1, wherein the hub is wearable and is configured to store the data, after wireless receiving it and before wirelessly transmitting it. 12. The system of claim 1, wherein said system is adapted to repeatedly: accept updated physiological data,use at least some of said updated physiological data as training data for machine learning,based on said machine learning, revise at least one pattern recognition algorithm, anduse said revised pattern recognition algorithm in a processor housed onboard a wearable or handheld device. 13. The system of claim 1, wherein the one or more servers are further adapted for analyzing said physiological data regarding a user, accessing audio or visual content stored in machine-readable form in memory, and outputting instructions for at least one said transducer to deliver said content to said user in a form perceptible to said user. 14. The system of claim 1, wherein said system is adapted for performing said pattern recognition in real time. 15. The system of claim 1, wherein said system is further adapted for soliciting at least one live human communication and... | Massachusetts Institute of Technology | A61B005/0482 A61B005/04 A61B005/0404 A61B005/021
US10709330 | B2 | 2020-07-14 | Ambulatory medical telemetry device having an audio indicator | 2005-11-04 | An ambulatory medical telemetry device (10) is provided. This device includes at least one sensor (18) for detecting at least one physiological parameter of a patient and a housing that is securable to the patient. A circuit (50) is located in the housing for receiving and processing a signal representative of the physiological parameter from the sensor to generate recordable physiological data. An audible indicator (16) is located in the housing and operationally coupled to the circuit for generating an audio signal indicating at least one operational state of the device. | 1. An ambulatory medical telemetry device, comprising: at least one sensor configured for detecting at least one physiological parameter of a patient;a housing securable to the patient;a circuit located in said housing and configured for receiving and processing a signal representative of the physiological parameter from the at least one sensor to generate recordable physiological data;at least one processor located in said housing and connected to the circuit, said at least one processor programmed to perform a self-test of an overall operation and each of a plurality of operational states of the ambulatory medical telemetry device, the plurality of operational states including at least three operational states; andan audible indicator located in said housing and operationally coupled to said circuit and said at least one processor to generate an audio signal indicating at least one operational state of the device including a successful completion of the self-test indicating the operational state is operating correctly for each of the plurality of operational states of the ambulatory medical telemetry device;wherein said audible indicator is configured to generate a plurality of different tones for indicating each of the overall operation and each operational state of the ambulatory medical telemetry device, each different tone corresponding to a different operational state of the device, the plurality of different tones further including a tone indicating the successful completion of the self-test for each of the plurality of operational states of the ambulatory medial telemetry device. 2. The device of claim 1, wherein said audio signal generated by the audible indicator further indicates physiological-related information. 3. The device of claim 1, further comprising: a wireless receiver operationally coupled to the circuit and the audible indicator and located in said housing for receiving a wireless paging signal such that the audible indicator, in response to the wireless paging signal, generates an audio paging signal. 4. The device of claim 3, wherein said audible indicator generates a plurality of different tones indicative of each different operational state and the audio paging signal. 5. The device of claim 1, wherein one of the operational states of the device indicates a successful completion of a test result to obtain the physiological data. 6. The device of claim 1, wherein each of said plurality of tones differs in at least one characteristic selected from the group consisting of frequency, duration, intervals between tones, intensity, and tunes. 7. The device of claim 1, wherein at least one of the tones represents positive feedback, and at least one of the tones represents negative feedback in response to the self-test for the overall operation and each operational state of the device. 8. The device of claim 1, wherein the physiological parameter detected by the sensor includes at least one of: electrocardiogram/respirat... | KONINKLIJKE PHILIPS N.V. | A61B005/00 G06F019/00 G08B021/02 G08B021/18 A61B005/02 G16H040/63 G16H040/40 G08B021/00
US07881760 | B2 | 2011-02-01 | Measuring structure for magneto encephalographic equipment with a superconducting magnetic-shield | 2007-06-20 | A magneto-encephalographic equipment superconducting magnetic-shield comprising a vacuum-tight body comprising an outer enclosure wall, a first inner enclosure wall inserted in the outer enclosure wall to define a upper closed space, and a second inner enclosure wall to define a lower open space. The first and second inner enclosure walls are arranged with the bottom of the first inner enclosure wall facing the ceiling of the second inner enclosure wall. A first enclosure of HTSC and a second enclosure of high-permeability material are concentrically arranged in the annular vacuum space defined between the first and second inner enclosure walls and the outer enclosure wall. A head-accommodating area is delimited by the hollow partition between, the bottom of the first inner enclosure wall and the ceiling of the second inner enclosure wall both facing each other, and a plurality of SQUID sensors are arranged in the upper closed space, encircling the head-accommodating area. | 1. A measuring structure for a magneto encephalographic equipment in a superconducting magnetic-shield comprising: an vacuum-tight structure comprising: an outer enclosure wall;a first inner enclosure wall inserted in the outer enclosure wall to define an upper closed space; anda second inner enclosure wall to define a lower open space, the first and second inner enclosure walls being arranged with the bottom of the first inner enclosure wall facing the ceiling of the second inner enclosure wall, thereby forming the vacuum-tight space by the first inner enclosure wall, the second inner enclosure wall and the outer enclosure wall;a first enclosure of high critical temperature superconductor;a second enclosure of high permeability material, the first and second enclosures being arranged in the annular vacuum space defined between the first and second inner enclosure walls and the outer enclosure wall; anda plurality of SQUID sensors arranged in the upper closed space, encircling the head-accommodating area, which is delimited by the hollow partition between the bottom of the first inner enclosure wall and the ceiling of the second inner enclosure wall; the SQUID sensors being adapted to be cooled by the coolant in the upper closed space whereas the first inner enclosure of high critical temperature superconductor being adapted to be cooled by the coolant in the annular vacuum space. 2. A measuring structure for a magneto encephalographic equipment in a superconducting magnetic-shield according to claim 1 wherein the cooling medium contained in the upper closed space is liquid helium whereas the cooling medium to cool the first enclosure of high critical temperature superconductor is cooled helium gas. 3. A measuring structure for a magneto encephalographic equipment in a superconducting magnetic-shield according to claim 1 wherein it further comprises an adiabatic expansion compartment, the cooling medium contained in the upper closed space being liquid helium whereas the cooling medium fed to the adiabatic expansion compartment being cooled helium gas, which is converted into the liquid helium by adiabatic expansion, and is fed to the upper closed space to make up for evaporation, whereas the cooling medium to cool the first enclosure of high critical temperature superconductor being cooled helium gas. 4. A measuring structure for a magneto encephalographic equipment in a superconducting magnetic-shield according to claim 1 wherein it further comprises an adiabatic expansion compartment, the cooling medium fed to the adiabatic expansion compartment being cooled helium gas, which is converted into the liquid helium by adiabatic expansion to indirectly cool the SQUID magnetic sensors via an associated thermal conduction body, the cooling medium to cool the first enclosure of high critical temperature superconductor being cooled helium gas. | National Institute of Information and Communications Technology | G01R033/035 A61B005/05
US07892182 | B2 | 2011-02-22 | Modulation and analysis of cerebral perfusion in epilepsy and other neurological disorders | 2007-10-26 | A system including an implantable neurostimulator device capable of modulating cerebral blood flow to treat epilepsy and other neurological disorders. In one embodiment, the system is capable of modulating cerebral blood flow (also referred to as cerebral perfusion) in response to measurements and other observed conditions. Perfusion may be increased or decreased by systems and methods according to the invention as clinically required. | 1. An implantable device for identifying changes in cerebral perfusion in a human patient comprising: a detection subsystem configured to be coupled to at least one sensor that generates a signal representative of the patient's cerebral perfusion over time, wherein the detection subsystem is operative to receive and process the signal to:identify a baseline cerebral perfusion level;identify if the cerebral perfusion is trending upwardly or downwardly from the baseline cerebral perfusion level, and, if trending occurs, adjust the baseline cerebral perfusion level to reflect the trend; andidentify at least one condition in the patient if the cerebral perfusion meets or exceeds a predetermined value or range of values, where at least one of the predetermined values or at least one of the predetermined range of values is based on the baseline cerebral perfusion level and the adjusted baseline cerebral perfusion level. 2. The implantable device of claim 1 wherein the at least one condition is an epileptic seizure, onset of an epileptic seizure, or precursor to the onset of an epileptic seizure. 3. The implantable device of claim 1 wherein the at least one condition is representative of a symptom of the neurological disorder. 4. The implantable device of claim 1, further comprising a therapy subsystem configured to be coupled to at least one therapy output, wherein the therapy subsystem is operative to selectively initiate delivery of a therapy to the therapy output. 5. The implantable medical device of claim 4 wherein the detection subsystem is programmed to cause the therapy subsystem to initiate application of the therapy in response to the at least one condition. 6. The implantable medical device of claim 4, wherein the therapy subsystem is adapted to modulate cerebral perfusion. 7. The implantable device of claim 1 wherein the signal representative of the patient's cerebral perfusion corresponds to one of the following: an optical oximetry measurement, an electromagnetic field measurement, a thermographic measurement, an ultrasonic transducer, and an electrochemical oxygen concentration measurement. 8. The implantable device of claim 1 wherein the signal representative of the patient's cerebral perfusion is an impedance plethysmography measurement. 9. The implantable device of claim 1, wherein the signal representative of the patient's cerebral perfusion is electrographic signal. 10. The implantable device of claim 4 wherein the therapy comprises electrical stimulation. 11. A system for identifying changes in cerebral perfusion in a human patient comprising: at least one sensor, wherein the sensor is configured to generate a signal representative of the patient's cerebral perfusion over time; anda detection subsystem coupled to the at least one sensor, wherein the detection subsystem is operative to receive and process the signal to:identify a baseline cerebral perfusion level;identify if the cerebral perfusion is trending upwardly or do... | NeuroPace, Inc. | A61B005/02
US20110028812 | A1 | 2011-02-03 | METHOD FOR SPECTROPHOTOMETRIC BLOOD OXYGENATION MONITORING | 2010-08-06 | A method and apparatus for non-invasively determining a blood oxygenation level within a subject's tissue is provided. The method includes the steps of: a) providing a spectrophotometric sensor operable to transmit light into the subject's tissue, and to sense the light; b) inputting into the sensor at least one of the subject's age, weight, brain development, and head size; c) spectrophotometrically sensing the subject's tissue along a plurality of wavelengths using the sensor, and producing signal data from sensing the subject's tissue; and d) processing the signal data utilizing the at least one of the subject's age, weight, brain development, and head size, to determine the blood oxygen saturation level within the subject's tissue using a difference in attenuation between the wavelengths. | 1. A method for non-invasively determining a blood oxygenation level within a subject's tissue, comprising the steps of: providing a spectrophotometric sensor operable to transmit light into the subject's tissue, and to sense the light;inputting into the sensor at least one of the subject's age, weight, brain development, and head size;spectrophotometrically sensing the subject's tissue along a plurality of wavelengths using the sensor, and producing signal data from sensing the subject's tissue; andprocessing the signal data utilizing the at least one of the subject's age, weight, brain development, and head size, to determine the blood oxygen saturation level within the subject's tissue using a difference in attenuation between the wavelengths. 2. The method of claim 1, wherein the sensor includes a processor that is adapted to include one or more calibration constants that relate to subject age, weight, brain development, and head size. 3. The method of claim 2, wherein the processor is adapted to utilize one or more of a graph, a database structure, and a mathematical relationship to relate the one or more calibration constants to subject age, weight, brain development, and head size. 4. The method of claim 3, wherein the one or more of a graph, a database structure, and a mathematical relationship are based on empirically collected data. 5. An apparatus for non-invasively determining a blood oxygenation level within a subject's tissue, comprising: a sensor having one or more transducer portions and a processor portion;wherein each of the one or more transducer portions includes at least one light source and at least one light detector, and the light source is operable to transmit light along a plurality of wavelengths into the subject's tissue, and the light detector is operable to detect light along the wavelengths traveling through the subject's tissue, and each of the transducer portions is operable to produce signal data representative of the light sensed within the subject's tissue; andwherein the processor portion is operably connected to the one or more transducer portions, and is adapted to receive input of at least one of the subject's age, weight, brain development, and head size, and the processor portion is adapted to process the signal data utilizing at least one of the subject's age, weight, brain development, and head size, to determine the blood oxygen saturation level within the subject's tissue using a difference in attenuation between the wavelengths. 6. The apparatus of claim 5, wherein the processor portion is adapted to include one or more calibration constants that relate to a subject age, weight, brain development, and head size. 7. The apparatus of claim 6, wherein the processor is adapted to utilize one or more of a graph, a database structure, and a mathematical relationship to relate the one or more calibration constants to a subject age, weight, brain development, and head size. 8. The apparatus of claim 7... | CAS MEDICAL SYSTEMS, INC. | A61B005/1455
US20110046474 | A1 | 2011-02-24 | METHOD TO CORRECT INTERFERENCE SIGNALS IN CARDIAC NOISE-TRIGGERING OF MEDICAL DATA ACQUISITION, AND MEDICAL SYSTEM OPERATING ACCORDING TO THE METHOD | 2010-08-19 | In a method and medical data acquisition system for the correction of interference signals (in particular those caused by gradient coils) in a magnetic resonance apparatus given the determination of a point in time in the cardiac cycle of a patient by the acquisition of cardiac noises with a sound sensor arranged on the patient, a calculated and/or pre-measured interference signal that describes the interference noise of the gradient coils is obtained and is subtracted from a raw signal measured by the sound sensor. | 1. A method for correcting cardiac noise signals acquired in a medical data acquisition system, comprising the steps of: acquiring a raw signal from a sound sensor arranged on the patient that contains a signal component representing cardiac noises in the patient, in an environment in a medical data acquisition system in which the patient is located in which interference noise exists;prior to acquiring medical data from the patient in said medical data acquisition system, predetermining, in a procedure selected from the group consisting of a calculation procedure and a measurement procedure, an interference signal that describes said interference noise, and subtracting said interference signal from said raw signal to generate a corrected raw signal that is substantially free of said interference noise; andderiving a trigger signal from the corrected raw signal and triggering acquisition of medical data from the patient in said medical data acquisition system using said trigger signal. 2. A method as claimed in claim 1 comprising evaluating said corrected raw signal to identify a point in time in the cardiac cycle of the patient, and deriving said trigger signal to cause triggering of said acquisition of medical data at said point in time. 3. A method as claimed in claim 2 comprising determining said point in time by analyzing the corrected raw signal to identify an occurrence of the R-spike therein. 4. A method as claimed in claim 1 wherein said medical data acquisition system is a magnetic resonance apparatus comprising gradient coils, and wherein said interference noise is generated by operating said gradient coils with respective gradient currents supplied thereto, and comprising predetermining said interference signal by a calculation with a linear dependence on the respective gradient currents. 5. A method as claimed in claim 4 comprising supplying respective current pulses to said gradient coils and, with said sound sensor, detecting respective pulse responses of said interference noise resulting from the respective operation of said gradient coils with said current pulses, and calculating said interference signal by convolving the respective pulse responses with the respective current pulses. 6. A method as claimed in claim 1 comprising implementing a calibration measurement prior to said medical data acquisition to obtain calibration measurement results, and predetermining said interference signal from said calibration measurement results. 7. A method as claimed in claim 5 comprising, in said calibration measurement, positioning said patient in said data acquisition system, and positioning said sound sensor on said patient, at respective positions corresponding to positions that will be used to acquire said medical data from the patient, and operating said medical data acquisition system in said calibration procedure in a same manner of operation that will be used for acquiring said medical data from the patient. 8. A method as cla... | - | A61B005/055
US20110046494 | A1 | 2011-02-24 | Blood Pressure Cuff and Connector Incorporating an Electronic Component | 2009-08-19 | A blood pressure cuff is equipped with an electronic component providing encoding of cuff properties. The connectors and hose connecting the cuff to a blood pressure measurement instrument are provided with conductors and electrical coupling, allowing the measuring instrument to access the electronic component encoding cuff properties. The arrangement of the cuff, hose, and connectors makes simultaneous pneumatic and electrical connection when the cuff is attached to the hose. | 1. A blood pressure measurement device comprising: a blood pressure cuff detachably connected to a blood pressure monitoring instrument by means of a hose assembly and connectors, in which the blood pressure cuff contains an electronic component capable of exchanging information with the blood pressure monitoring instrument, the hose assembly is provided with electrical conductors in addition to one or more pneumatic lumens, and at least the mating connectors between the cuff and hose assembly are provided with electrical coupling in addition to pneumatic coupling, such that simultaneous electrical and pneumatic connection is established between the cuff and instrument when the cuff is connected to the hose. 2. The device of claim 1, in which the electronic component includes encoding of properties of the cuff. 3. The device of claim 1, in which the electronic component includes a sensor. 4. The device of claim 1, in which the electronic component is an impedance network. 5. The device of claim 4, in which the impedance network is a resistor. 6. The device of claim 4, in which the impedance network contains non-linear elements. 7. The device of claim 1, in which the electronic component contains a read-only memory device. 8. The device of claim 1, in which the electronic component contains a re-writable memory device. 9. The device of claim 8, in which the memory device is used to store the number of uses of the cuff. 10. The device of claim 8, in which the memory device is used to store patient data. 11. The device of claim 1, in which the blood pressure cuff electronic component includes encoding of a unique identifier or serial number. 12. The device of claim 11, in which the blood pressure cuff and its associated unique identifier are used as a patient identifier. 13. A blood pressure hose assembly having integrated electrical conductors in addition to one or more pneumatic lumens. 14. The blood pressure hose assembly of claim 13, in which the electrical conductors are located within a pneumatic lumen of the hose. 15. The blood pressure hose assembly of claim 13, in which the electrical conductors are located in a lumen not used for pneumatic purposes. 16. The blood pressure hose assembly of claim 13, in which the electrical conductors are imbedded in the wall of the hose. 17. The blood pressure hose assembly of claim 13, in which a common outer jacket integrates the electrical conductors with the pneumatic portion of the hose. 18. A blood pressure cuff connector integrating electrical connection as well as pneumatic connection, such that the pneumatic and electrical circuits are simultaneously engaged when the connector is mated. 19. The connector of claim 18, in which the electrical connection is made by one or more electrical contacts located adjacent to the pneumatic connection. 20. The connector of claim 18, in which one or more electrical contacts are arranged as annular rings surrounding the pneumatic connection. 21... | Mindray DS USA, Inc. | A61B005/022
US20110040201 | A1 | 2011-02-17 | Periodic Disordered Breathing Detection | 2010-10-25 | Systems and methods are directed to evaluating breathing disorders, such as periodic disordered breathing. A signal representative of patient respiration is developed, typically patient-internally. An envelope of the signal is provided. Periodicity of the envelope is detected, and presence and severity of periodic disordered breathing is determined based on the periodicity of the envelope. | 1. A system, comprising: an implantable housing;detection circuitry provided in the housing and configured to detect a signal representative of patient respiration; anda processor configured to detect periodicity of an envelope of the signal and determine presence of periodic disordered breathing based on the periodicity of the envelope by computing a parameter characterizing envelope cycle length for multiple envelope cycles occurring within a time window, the parameter comprising one or both of a standard deviation (KSD) of duration for n envelope cycles and the predetermined duration divided by a mean (KMEAN) indicative of average envelope cycle length, wherein n is a number. 2. The system of claim 1, wherein the processor is configured to detect regularity of envelope periodicity. 3. The system of claim 1, wherein the processor is configured to determine presence of periodic disordered breathing by determining regularity of envelope periodicity and a duration of envelope cycles occurring within the time window of predetermined duration. 4. The system of claim 1, wherein the processor is configured to determine a duration of envelope periodicity. 5. The system of claim 1, wherein the processor is configured to determine envelope periodicity if the standard deviation (KSD) is less than a first threshold and n is greater than a second threshold. 6. The system of claim 1, wherein the processor is configured to determine presence of periodic disordered breathing if the standard deviation (KSD) is less than a first threshold, n is greater than a second threshold, and KMEAN falls within a third threshold range. 7. The system of claim 1, wherein the processor is configured to compute an estimated apnea-hypopnea index based on the periodicity of the envelope. 8. The system of claim 1, wherein the processor is configured to compute an estimated apnea-hypopnea index based on a number of periodic disordered breathing events occurring within each of a plurality of the time windows, a total duration of the plurality of time windows defining a duration of patient sleep. 9. The system of claim 1, wherein the processor is configured to determine severity of the periodic disordered breathing. 10. The system of claim 1, wherein the processor is configured to determine severity of the periodic disordered breathing at least in part by distinguishing between central sleep apnea and obstructive sleep apnea. 11. The system of claim 1, wherein the processor is configured to determine severity of a patient's heart failure condition by determining presence of the periodic disordered breathing in each of a plurality of physiological states. 12. The system of claim 1, wherein the processor is configured to determine severity of the periodic disordered breathing at least in part by determining a frequency of the periodicity and a stability of the periodicity. 13. The system of claim 1, wherein the processor is configured to determine severity of the periodic d... | - | A61B005/08
US20110046517 | A1 | 2011-02-24 | METHOD AND SYSTEM FOR MONITORING SKELETAL DEFECTS | 2010-05-17 | The present invention relates to a handheld, point-of-care device that uses non-hazardous low level ultrasound to quantitatively detect, monitor and supply, in real time, information on the status of bone fractures from inception to full healing. The device includes a plurality of transducers housed in the handheld unit or which can be extended from the handheld unit to be placed at or near the site of the bone fracture. A plurality of acoustic parameters of the transducers are selected to provide optimal detection and monitoring of fractures in bone. The information can be displayed as a numerical readout indicating the severity of the break and provides ability for indicating small stress and fatigue fractures. The portable device provides rapid and inexpensive detection and diagnosis of musculoskeletal problems using low level ultrasound which can measure bone density, determine fracture status and monitor healing rate. | 1. A method of measuring, in vivo, mechanical strength and structural integrity of bone structure in a body having soft tissue and skin by the following procedures: (a) mounting three ultrasound transducers over the skin and propagating a signal from one of the transducers, said transducer being focused and angled such that the signal will propagate through the soft tissue and along the bone;(b) establishing a fixed reference signal for the ultrasound transducers;(c) receiving the propagated ultrasound signal at another one of said transducers;(d) sorting the received ultrasound signal according to a relative propagation delay time by correlating the received ultrasound signal with the fixed reference signal;(e) determining at least one acoustical parameter from the ultrasound signal propagated along the bone and the surrounding soft tissue based upon the propagated ultrasound sorting of (d); and(f) relating the at least one acoustical parameter to the mechanical strength and the structural integrity of the bone. 2. The method according to claim 1 further comprising the step of: establishing an additional reference signal by monitoring the ultrasound signal propagated through a normal part of the bone. 3. The method of claim 1 wherein the acoustical parameters are selected from the group of flight time, pulse duration, rise time, maximum amplitude, counts and energy. 4. The method of claim 1 wherein the acoustical parameters are selected from a velocity, a propagation energy and a degree dispersion. 5. A method as claimed in claim 1 wherein step (d) comprises the substep of sorting the received propagated ultrasound signal propagated along the bone and the received propagated ultrasound signal along the soft tissue. 6. The method of claim 1 further comprising the steps of: storing the mechanical strength and structural integrity of the bone;repeating steps (a) through (f) over time; anddisplaying the mechanical strength and structural integrity of the bone over time. 7. The method of claim 1 wherein the angle of at least one of said transducers can vary as a function of soft tissue thickness. 8. The method of claim 1 wherein the transducers have a resonance frequency greater than 1 MHz. 9. A method of measuring mechanical strength and structural integrity of bone in a body having bone and soft tissue said method comprising: (a) placing a focused and angled ultrasound transmitter and a first ultrasound receiver over the body with a first distance from the ultrasound transmitter therebetween;(b) placing a second ultrasound receiver over the body at a second distance less than the first distance from the ultrasound transmitter;(c) transmitting low level ultrasound into the body so as to propagate the transmitted ultrasound through the soft tissue and along the bone;(d) generating a fixed reference signal by joining the ultrasound transmitted signal with one of the ultrasound receivers;(e) receiving ultrasound propagated along the bone at the... | - | A61B005/103
US07894876 | B2 | 2011-02-22 | Combined MR-optical coil for prostate, cervix and rectum cancer imaging diagnostics | 2005-02-01 | We present, in exemplary embodiments of the present invention, a system combining anatomical imaging technologies (e.g., MR) with optical technologies. The system can be used for a variety of applications, including, but not limited to, (1) cancer diagnosis and staging; (2) image guidance; and (3) radiation therapy planning. Image guidance may include guiding a biopsy. For example, a prostatectomy potentially has severe side effects, such as impotence and incontinence. Thus, a histologically-confirmed diagnosis, such as one provided from a biopsy, may prevent unnecessary prostatectomy. Image guidance may also include guiding minimal invasive therapy, such as brachytherapy focused ultrasound. The present invention may be used to plan radiation therapy, for example, by detecting, and thus sparing, healthy tissue from radiation exposure. | 1. An apparatus for providing anatomical and molecular diagnostic imaging, comprising: a magnetic resonance (MR) coil assembly comprising MR coil wires and a translucent support between the MR coil wires for inserting into a body cavity; andan optical probe disposed within the MR coil assembly at a distal end of the apparatus adjacent the MR coil wires and moveable along an axis of the MR coil assembly relative to the MR coil,wherein the optical probe is configured to provide light to the body cavity and to detect fluorescence in the body cavity in response to the provided light, andwherein the optical probe is connected within the apparatus to enable the optical probe to be moved closer to a part of the MR coil in a first direction along the axis and moved further away from a part of the MR coil in a second and opposing direction along the same axis. 2. The apparatus of claim 1, wherein the optical probe detects near infra-red fluorescence. 3. The apparatus of claim 1, wherein the optical probe rotates independent of the MR coil. 4. The apparatus of claim 1, wherein movement of the optical probe is reproducible. 5. The apparatus of claim 1, further comprising a handheld component operatively connected to the MR coil. 6. The apparatus of claim 1, wherein the MR coil comprises a flexible coil. 7. The apparatus of claim 1, further comprising a sheath encompassing the MR coil. 8. The apparatus of claim 7, wherein the sheath comprises at least one light-transparent balloon for inflation after insertion into the body cavity to fix the MR coil in the body cavity. 9. The apparatus of claim 8, wherein the light-transparent balloon has a toric configuration. 10. The apparatus of claim 1, wherein the MR coil is operatively connected to a magnetic resonance amplifier. 11. The apparatus of claim 1, further comprising a first wire for receiving data from the MR coil and a second wire for receiving data from the optical probe. 12. The apparatus of claim 1, further comprising a modulation component for distinguishing between data from the MR coil and data from the optical probe. 13. The apparatus of claim 1, wherein the optical probe comprises a light wire for projecting the light and an optical head for detecting the fluorescence. 14. The apparatus of claim 13, wherein the light wire comprises at least one fiber, and wherein the at least one fiber feeds one of a circular diffuser and discrete microscopic lenses assembled in a circle. 15. The apparatus of claim 13, further comprising a rotating minor operatively connected to one end of the light wire, the rotating minor for circularly projecting the light in the body cavity. 16. The apparatus of claim 1, wherein the optical probe comprises a digital mirror array. 17. The apparatus of claim 1, further comprising a light source. 18. The apparatus of claim 17, wherein the light source comprises one of a broadband light source, a light-emitting diode (LED), and a laser. 19. The apparatus of claim ... | Siemens Medical Solutions, Inc. | A61B005/05
US08452367 | B2 | 2013-05-28 | Forehead sensor placement | 2010-07-26 | Forehead oximetry sensor devices and methods for determining physiological parameters using forehead oximetry sensors. One method includes placing an oximetry sensor on the forehead of a patient, such that the sensor is placed on the lower forehead region, above the eyebrow with the sensor optics placed lateral of the iris and proximal the temple; and operating the pulse oximeter to obtain the physiological parameter. In one aspect, the method also includes providing and placing a headband over the oximetry sensor, or alternately, the sensor is a headband-integrated sensor. The headband has an elastic segment sized to fit around the patient's head. The headband also includes a non-elastic segment that is smaller than and attached with the elastic segment. The non-elastic segment is sized to span a portion of the elastic segment when the elastic segment is stretched. In addition, the non-elastic segment is larger than the portion of the elastic segment it spans when the elastic segment is not stretched. When the headband or the headband-integrated sensor is sufficiently tight, it delivers a pressure in the range higher than the venous pressure and lower than the capillary pressure to the forehead of the patient. | 1. A method of determining a physiological parameter using a pulse oximeter, comprising: placing an oximetry sensor comprising an arcuate substrate having an emitter and a detector on a lower forehead region of a patient, such that the emitter and the detector are placed about a line centered over the iris such that the emitter or the detector is on one side of the iris and that the other of the emitter or the detector is on the other side of the iris; andoperating the pulse oximeter to obtain the physiological parameter. 2. The method of claim 1, wherein the placing comprises placing the sensor such that an axis connecting the emitter and the detector lies within 8 mm of the top of the patient's eyebrow. 3. The method of claim 1, wherein the placing comprises placing the sensor such that an axis connecting the emitter and the detector lies within 5 mm of the top of the patient's eyebrow. 4. The method of claim 1, wherein the substrate is resilient and has a shape conformable to the lower forehead of the patient. 5. The method of claim 1, wherein the emitter and detector are disposed on the substrate such that an axis connecting the emitter and the detector does not bisect the substrate. 6. The method of claim 1, comprising providing a headband comprising an elastic segment sized to fit around the patient's head and a non-elastic segment being smaller than and attached with the elastic segment, the non-elastic segment sized to span a portion of the elastic segment when the elastic segment is stretched, the non-elastic segment being larger than the portion of the elastic segment it spans when the elastic segment is not stretched; and placing the headband over the oximetry sensor. 7. The method of claim 6, wherein the non-elastic segment is attached with the elastic segment in such a manner that the non-elastic segment projects out from the surface of the elastic portion when the headband is not sufficiently tight, thus creating a loop which provides a visual indication that the headband needs re-tightening. 8. The method of claim 1, wherein the physiological parameter is an oxygen saturation value. 9. The method of claim 1, wherein the physiological parameter is a pulse rate. 10. A system comprising: a monitor configured to determine a physiological parameter based at least in part upon a signal related to the physiological parameter; anda sensor coupled to the monitor and capable of generating the signal related to the physiological parameter, the sensor comprising: a substrate being adapted to directly contact a portion of a forehead of a patient, wherein the substrate is arcuate when placed flat upon a planar surface in the absence of any restraining force;an emitter and a detector disposed on the substrate; anda headband configured to hold the substrate against the patient's forehead, the headband sized to fit around the patient's head, wherein the headband comprises: an elastic segment sized to fit around a patient's head; anda non-e... | Covidien LP | A61B005/1455
US20110028864 | A1 | 2011-02-03 | Uterine sound device | 2009-07-31 | A uterine sound device includes a first component having a handle and a stem, and a slidable member positioned around the stem. An inner wall of the slidable member fits around the stem such that the slidable member remains at a fixed position on the stem unless a predetermined amount of force is applied to the slidable member longitudinally along the stem. The stem includes graduated markings to indicate units of measure that are preferably formed of a non-leachable, biocompatible ink printed on the stem. In use, the sound stem is inserted past the cervix with the slidable member remaining on the outside of the cervix with insertion pressure causing the slidable member to slide along the stem toward the handle. The position of the slidable member according to the indicia indicates a depth into the uterus that the stem has reached. | 1. A uterine sound device for measuring the depth of a uterine cavity, the sound device comprising: a first component having a handle and a stem; anda slidable member positioned around the stem,wherein the stem includes graduated markings to indicate a unit of measure, the graduated markings printed on the stem, and wherein an interior wall of the slidable member fits around the stem such that the slidable member remains at a fixed position on the stem unless a predetermined amount of force is applied to the slidable member longitudinally along the stem. 2. The uterine sound device of claim 1, wherein the slidable member is a circular ring. 3. The uterine sound device of claim 2, wherein the circular ring is formed of a non-toxic plastic material. 4. The uterine sound device of claim 2, wherein the circular ring is an O-ring. 5. The uterine sound device of claim 1, wherein the stem is a semi-flexible stem having a predetermined flexibility such that the stem is rigid enough to penetrate a cervix yet flexible enough to bend about a contoured region of a uterus, whereby a course of the uterus may be measured. 6. The uterine sound device of claim 1, wherein the stem has a rounded tip. 7. The uterine sound device of claim 1, wherein the stem has a blunt tip. 8. The uterine sound device of claim 1, wherein the handle and the stem are formed of a material capable of undergoing a standard autoclaving sterilization process. 9. The uterine sound device of claim 8, wherein the slidable member is also formed of a material capable of undergoing a standard autoclaving sterilization process. 10. The uterine sound device of claim 1, wherein the first component is a single integral piece with a handle portion and a stem portion. 11. The uterine sound device of claim 1, wherein the stem has a first diameter and the slidable member has an inner diameter smaller than the first diameter when the slidable member is removed from the stem. 12. The uterine sound device of claim 11, wherein the slidable member is formed of a stretchable material. 13. The uterine sound device of claim 1, wherein the handle is structured to define a recessed portion for receiving a thumb of a user to provide an improved grip. 14. The uterine sound device of claim 1, wherein the stem is curved a predetermined amount in a region of the stem opposite the handle. 15. The uterine sound device of claim 1, wherein the graduated markings are formed of a non-leachable, biocompatible ink printed on the stem. 16. A uterine sound device for measuring a depth of a uterine cavity, the sound device comprising: a first component having a handle and a stem; anda circular ring positioned around the stem,wherein the stem includes graduated markings to indicate a unit of measure, the graduated markings printed on the stem, wherein the stem has a rounded tip, and wherein an interior wall of the circular ring fits around the stem such that the ring remains at a fixed position on the stem unless a... | - | A61B005/107
US08414750 | B2 | 2013-04-09 | Subcutaneous glucose electrode | 2010-09-29 | A small diameter flexible electrode designed for subcutaneous in vivo amperometric monitoring of glucose is described. The electrode is designed to allow “one-point” in vivo calibration, i.e., to have zero output current at zero glucose concentration, even in the presence of other electroreactive species of serum or blood. The electrode is preferably layered, with the layers serially deposited within a recess upon the tip of a polyamide insulated gold wire. A first glucose concentration-to-current transducing layer can be overcoated with an electrically insulating and glucose flux limiting layer (second layer) on which, optionally, an immobilized interference-eliminating horseradish peroxidase based film is deposited. An outer layer is preferably biocompatible. | 1. A glucose sensor comprising: a working electrode; anda multilayer membrane complex comprising: an analyte responsive sensing layer disposed on top and in electrical contact with the working electrode, the analyte responsive sensing layer comprising a glucose-specific enzyme, a mediator and a first polymer selected from the group consisting of poly(vinylimidazole) and poly(vinyl pyridine), wherein at least one of the glucose-specific enzyme and the mediator is chemically bonded to the first polymer; anda diffusion-limiting layer comprising a crosslinked second polymer disposed on top and in contact with the analyte responsive sensing layer. 2. The glucose sensor of claim 1, wherein at least one of the glucose-specific enzyme and the mediator is covalently bonded to the first polymer. 3. The glucose sensor of claim 2, wherein at least one of the glucose-specific enzyme and the mediator forms one or more of carbon-carbon, carbon-nitrogen, or metal-carbon covalent bonds with the first polymer. 4. The glucose sensor of claim 2, wherein at least one of the glucose-specific enzyme and the mediator is covalently bonded to the first polymer through a linker moiety. 5. The glucose sensor of claim 2, wherein the mediator is covalently bonded to the first polymer through one or more ligands of the mediator. 6. The glucose sensor of claim 5, wherein the one or more ligands comprises a substituted or unsubstituted heterocyclic nitrogen-containing moiety. 7. The glucose sensor of claim 5, wherein one or more ligands of the mediator are substituents of the first polymer backbone. 8. The glucose sensor of claim 2, wherein the glucose-specific enzyme is covalently bonded to the first polymer through one or more side chains of the enzyme. 9. The glucose sensor of claim 8, wherein the one or more side chains of the glucose-specific enzyme is selected from the group consisting of substituted or unsubstituted amines, alcohols, thiols, phenols, imidazoles, indols, and carboxylic acids. 10. The glucose sensor of claim 9, wherein one or more side chains of the glucose-specific enzyme are substituents of the first polymer backbone. 11. The glucose sensor of claim 1, wherein the first polymer is crosslinked. 12. The glucose sensor of claim 11, wherein at least one of the glucose-specific enzyme and the mediator is crosslinked with the first polymer. 13. The glucose sensor of claim 12, wherein the mediator is crosslinked to the first polymer. 14. The glucose sensor of claim 11, wherein at least one of the glucose-specific enzyme and the mediator is crosslinked with the first polymer through one or more crosslinking agents. 15. The glucose sensor of claim 11, wherein one or more side chains of the glucose-specific enzyme are crosslinked to the first polymer. 16. The glucose sensor of claim 1, wherein the first polymer forms a coordination complex with the mediator. 17. The glucose sensor of claim 16, wherein one or more nitrogen-containing moieties of the ... | Abbott Diabetes Care Inc. | C12M001/00 A61B005/00 C12Q001/00
US09155482 | B2 | 2015-10-13 | Apparatus and method for indicating cardiac output | 2009-03-04 | An apparatus for indicating cardiac output comprises means for monitoring a patient's transthoracic impedance and generating a corresponding impedance signal, and signal processing means for (a) deriving a signal S1 which is a measure of the average amplitude of the impedance signal, (b) filtering the impedance signal at a plurality of different wavelengths within a predetermined frequency band, (c) for each filter deriving a signal S2 which is a measure of the average amplitude of the respective filter output, (d) calculating the ratio of the maximum one of the signals S2 derived from step (c) to the signal S1 derived from step (a), and (e) using the ratio from step (d) in a decision tree to provide a signal indicating cardiac output or not. | 1. An external defibrillator, comprising: a processor;a frequency generator;a pair of electrodes; anda non-transitory computer-readable storage medium having instructions stored which, when executed by the processor, cause the processor to perform operations comprising: signaling the frequency generator to generate an impedance test signal;when the pair of electrodes are in contact with a patient, applying the impedance test signal to the patient via the pair of electrodes:receiving an impedance measurement in response to the impedance test signal; repeating the impedance test signal with the patient to acquire a plurality of impedance measurements;averaging the plurality of impedance measurements over a defined period of time, to yield an average impedance measurements;filtering each of the plurality of impedance measurements at a plurality of different spaced frequencies within a predetermined band, to yield a plurality of filtered impedance measurements;averaging each of the plurality of filtered impedance measurements over a defined period of time, to yield a plurality of average filtered impedance measurements;identifying a maximum average filtered impedance measurement from the plurality of average filtered impedance measurements;calculating a ratio of the maximum average filtered impedance measurement to the average impedance measurement;comparing the calculated ratio with a predetermined threshold value to provide a signal indicating the presence or absence of a cardiac output; and providing the signal indicating the presence or absence of the cardiac output to a user. 2. The external defibrillator of claim 1, wherein the threshold value is in a range inclusively comprising 0.25 to 0.75; andwhen the calculated ratio exceeds the threshold value, a first signal is provided indicating the presence of a cardiac output; andwhen the calculated ratio is below the threshold value, a second signal is provided indicating the absence of a cardiac output. 3. The external defibrillator of claim 2, the computer-readable storage medium having additional instructions stored which, when executed by the processor, result in operations, comprising: when the calculated ratio is below the first threshold value,comparing the calculated ratio with a second threshold value that is lower than the first threshold value; andwhen the calculated ratio is less than the second threshold value, providing a distinct signal indicating the absence of a cardiac output. 4. The external defibrillator of claim 3, wherein the second threshold value is in a range inclusively comprising 0.25 to 0.03. 5. The external defibrillator of claim 1, wherein the defined period of time is in a range inclusively comprising 1-12 seconds. 6. The external defibrillator of claim 1, wherein the predetermined frequency band has a bandwidth in a range inclusively comprising 1 to 5 Hz. 7. The external defibrillator of claim 6, wherein the plurality of different spaced frequencies number at le... | Heartsine Technologies Limited | A61B005/05 A61B005/029 A61B005/053 A61N001/39
US09623266 | B2 | 2017-04-18 | Estimation of alignment parameters in magnetic-resonance-guided ultrasound focusing | 2009-08-04 | A magnetic-resonance-guided focused ultrasound system may be calibrated by generating ultrasound foci using ultrasound transducers, establishing coordinates of the foci and of magnetic-resonance trackers associated with the transducers, and determining a geometric relationship between the trackers and the transducers. | 1. A method for calibrating a magnetic-resonance-guided focused-ultrasound system, comprising: (a) providing a phased array of ultrasound transducers, the array having at least one magnetic resonance (MR) tracker associated therewith at a fixed relative position with respect thereto;(b) establishing MR coordinates of the MR trackers;(c) for each of a plurality of sonication schemes, (i) creating an ultrasound focus, and (ii) establishing parameters indicative of MR coordinates of the ultrasound focus, wherein each sonication scheme is associated with a scheme-specific set of production errors; and,(d) based at least in part on the parameters, determining a geometric relationship between the ultrasound transducers and the MR trackers. 2. The method of claim 1 wherein the MR trackers and the array are part of a single rigid structure. 3. The method of claim 1 wherein the MR trackers comprise at least one of a micro-coil or an MR imaging marker. 4. The method of claim 1 wherein the ultrasound focus is created in a phantom. 5. The method of claim 1 wherein establishing the MR coordinates of the ultrasound focus comprises at least one of thermal MR imaging or acoustic-radiation-force MR imaging. 6. The method of claim 1 wherein the parameters comprise the MR coordinates of the ultrasound focus. 7. The method of claim 1 wherein the parameters comprise MR coordinates of a projection of the ultrasound focus. 8. The method of claim 1 wherein the determination of the geometric relationship is further based on the sonication schemes. 9. The method of claim 1 wherein the geometric relationship comprises coordinates of the MR trackers in a transducer coordinate system. 10. The method of claim 9 wherein determining the geometric relationship comprises determining a transformation between the MR coordinate system and the transducer coordinate system based, at least in part, on the parameters and the sonication schemes. 11. The method of claim 10 wherein determining the geometric relationship further comprises using the transformation to determine the coordinates of the MR trackers in the transducer coordinate system. 12. The method of claim 1 wherein the geometric relationship comprises production errors indicative of a deviation of the fixed relative positions between the MR trackers and the transducers from nominal relative positions. 13. The method of claim 12 wherein determining the geometric relationship comprises determining a transformation between the parameters and the production errors based, at least in part, on the sonication scheme. 14. The method of claim 13 wherein determining the geometric relationship further comprises applying a linear estimation method to the transformation and the parameters, thereby determining the production errors. 15. The method of claim 14 wherein the linear estimation method comprises a least square method. 16. A method for operating a magnetic-resonance-guided focused-ultrasound system, comprising: (a) ... | INSIGHTEC LTD. | A61B005/055 A61N007/02 G01R033/48 A61B008/00 A61B017/00 A61N007/00 A61B090/00
US09918635 | B2 | 2018-03-20 | Systems and methods for optimizing insulin dosage | 2010-06-18 | Embodiments of a testing method suitable for diabetic persons to optimize their administered insulin dosage comprise collecting one or more sampling sets of biomarker data, wherein each sampling set comprises a sufficient plurality of non-adverse sampling instances and wherein each sampling instance comprises an acceptable biomarker reading at a single point in time recorded upon compliance with adherence criteria, determining a biomarker sampling parameter from each sampling set, comparing the biomarker sampling parameter to a target biomarker range, calculating an insulin adjustment parameter associated with the biomarker sampling parameter if the biomarker sampling parameter falls outside the target biomarker range, adjusting the insulin dosage by the insulin adjustment parameter if the biomarker sampling parameter falls outside the target biomarker range and if the insulin dosage does not exceed maximum dosage, and exiting the testing method if the adjusted insulin dosage is optimized. The insulin dosage is optimized when one or more biomarker sampling parameters fall within a target biomarker range. | 1. A testing method suitable for a diabetic person to optimize an administered insulin dosage comprising: providing a structured collection procedure from a server of a health care provider to a collection device configured to guide the diabetic person through the structured collection procedure and optimize the administered insulin dosage, wherein the collection device comprises a meter configured to measure one or more selected biomarkers, a processor disposed inside the meter and coupled to a memory, wherein the memory stores the structured collection procedure provided from the server, and software having instructions that when executed by the processor causes the processor to instruct the diabetic person to collect one or more sampling sets of biomarker data in accordance with the structured collection procedure, wherein the server is a central repository for a plurality of the structured collection procedures;executing the software on the collection device and according to instructions and timing provided by the structured collection procedure:collecting one or more sampling sets of biomarker data, wherein each of the one or more sampling sets comprises a plurality of sampling instances recorded over a collection period and each sampling instance of the plurality of sampling instances comprises an acceptable biomarker reading recorded upon compliance with one or more acceptance criterion that is applied to the one or more sampling sets of biomarker data, wherein the one or more acceptance criterion comprises one or more considerations selected from the group consisting of: a patient diet, fasting or eating regimen; a patient exercise regimen; a patient lifestyle; a patient sleep regimen; insulin dosage amounts; and timing of biomarker reading collection;wherein after each collection of the one or more sampling sets of biomarker data, said processor performs the processes of:determining a biomarker sampling parameter from the one or more sampling sets of biomarker data whereby only these biomarker data are considered which are in compliance with the one or more acceptance criterion;comparing the biomarker sampling parameter to a target biomarker range;calculating an insulin adjustment parameter associated with the biomarker sampling parameter in response to the biomarker sampling parameter falling outside the target biomarker range;adjusting an insulin dosage by an amount of the insulin adjustment parameter in response to the biomarker sampling parameter falling outside the target biomarker range and in response to the insulin dosage not exceeding maximum dosage;exiting the testing method in response to the adjusted insulin dosage being optimized as an optimized insulin dosage such that the optimized insulin dosage is administered as the administered insulin dosage, otherwise repeating with a next sampling set of the one or more sampling sets of biomarker data, the optimized insulin dosage being achieved when the one or more biomarker sampli... | Roche Diabetes Care, Inc. | A61B005/00 A61B005/145 G06F019/00 A61M031/00 A61B005/05 C12Q001/00
US20110040166 | A1 | 2011-02-17 | EXTRACTING KIT, EXTRACTING DEVICE, AND EXTRACTING METHOD | 2010-10-21 | An extracting kits for extracting analyte through the skin of a living body are disclosed that comprises an inserting device comprising needles for forming a plurality of extraction holes in the skin through the stratum corneum but not reaching the subcutaneous tissue and an extracting device for extracting the analyte into a liquid supplied into the extraction holes by supplying the liquid into the plurality of extraction holes. An extracting device and extracting method are also described. | 1-17. (canceled) 18. An extracting method for extracting an analyte through skin of a living body, the method comprising: forming a plurality of extraction holes in the skin through stratum corneum but not reaching subcutaneous tissue utilizing a plurality of needles configured to form the plurality of extraction holes;removing the plurality of needles from the skin;supplying a liquid to the formed plurality of extraction holes; andextracting the analyte contained within the supplied liquidgenerating a signal representative of the extracted analyte utilizing a sensor placed in contact with the supplied liquid. 19. The extracting method of claim 18, further comprising: transferring the analyte contained in the liquid supplied to the extraction holes in response to a transfer voltage applied to the liquid. 20. The extracting method of claim 19, wherein the voltage applied to the liquid in electrical communication with the living body is of a magnitude such that a current does not effectively flow to the stratum corneum of the skin. 21. The extracting method of claim 18, wherein each of the plurality of needles projects from a support by a length of between about 20 μm and about 1,000 μm. 22. The extracting method of claim 18, wherein extracting further comprises: arranging an extraction device adjacent to the skin of the living body, andcollecting the analyte transferred out of the extraction holes in the transferring by the extracting device. 23. The extracting method of claim 18, further comprising: forming a position marker on the skin of the living body to specify the position of at least one of the plurality of extraction holes. 24. The extracting method of claim 18, further comprising: checking a state of extraction hole formation after removing the inserting device from the skin;extracting, if the checked state of the extraction hole formation is sufficient, the analyte. 25. The extracting method of claim 24, wherein checking further comprises: applying an electric field to the living body, andelectrically checking the formation state of at least one of the plurality of extraction holes based on a current value or a voltage value associated with the applied electric field. 26. The extracting method of claim 25, wherein the formation state of the at least one of the plurality of extractions is determined to be inadequate when the current value is less than a predetermined value. 27. The extracting method of claim 25, wherein the formation state of the at least one of the plurality of extraction holes is determined to be inadequate when the voltage value exceeds a predetermined value. 28. The extracting method of claim 24, further comprising: outputting a message in response to a determination of inadequate extraction hole formation. 29. (canceled) 30. An extracting method for extracting analyte through skin of a living body, the method comprising: forming a plurality of extraction holes in the skin through stratum corneum but not... | Sysmex Corporation | A61B005/145
US08287454 | B2 | 2012-10-16 | Analyte monitoring device and methods of use | 2010-09-27 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. A glucose monitoring system comprising: an in vivo sensor for monitoring a glucose level of a user;a sensor control unit coupled to the sensor to output one or more signals associated with the monitored glucose level of the user; anda receiver/display unit in communication with the sensor control unit and comprising: a receiver to receive the outputted one or more signals from the sensor control unit;an input device to receive an input from the user, wherein the input comprises a drug delivery information, a food intake, and an exercise level;a processor to evaluate the one or more signals from the sensor control unit and the input from the user to produce a result comprising the glucose level in interstitial fluid of the user, the glucose level in blood of the user, a rate of change of the glucose level, an acceleration in the rate of change of the glucose level, a dose of a medication and an insulin sensitivity factor; anda display in communication with the processor,wherein the processor temporarily disables a one or more therapy related functions executable by the user if one or more of the glucose level, the rate of change of the glucose level and the acceleration in the rate of change of the glucose level meets or exceeds a threshold level. 2. The system of claim 1, wherein the result further comprises one or more of an insulin action curve, a carbohydrate action curve and a basal profile information. 3. The system of claim 1, wherein the medication comprises insulin. 4. The system of claim 1, wherein the display displays a warning indication to the user if one or more of the glucose level, the rate of change of the glucose level and the acceleration in the rate of change of the glucose level meets or exceeds a threshold level. 5. The system of claim 4, wherein the warning indication comprises one of more of: an indication of hypoglycemia; an indication of impending hypoglycemia; an indication of hyperglycemia; an indication of impending hyperglycemia; an indication of a malfunction; and an indication of an error condition. 6. The system of claim 1, wherein the processor tags the one or more signals with an event marker. 7. The system of claim 6, wherein the event marker comprises one or more of a hypoglycemic episode, a hyperglycemic episode, the food intake, the drug delivery information, and the exercise level. 8. The system of claim 1, wherein the system comprises an integrated unit including the sensor, the sensor control unit and the receiver/display unit. 9. The system of claim 1, wherein the sensor control unit and the receiver/display unit are separate units. 10. The system of claim 9, wherein the sensor control unit comprises a transmitter to transmit the one or more signals from the sensor control unit to the receiver/display unit. 11. The system of claim 1, wherein the receiver/display unit further comprises a data storage unit to store at least one of the one or more signals and the result. 12. The system of claim... | Abbott Diabetes Care Inc. | A61B005/00 A61M031/00 A61M001/00
US08078252 | B2 | 2011-12-13 | Intracranial neural interface system | 2010-04-22 | In some preferred embodiments, without limitation, the present invention comprises an implantable, intracranial neural interface node which is an integrated and minimally invasive platform system and supports cross-modal neural interfaces to the cerebrum and other associated structures in the central nervous system. The neural interfaces comprise electrical and chemical interfaces for neural recording, electrical stimulation, chemical delivery, chemical sensing, chemical sampling, cell delivery, genetic material delivery and/or other functions of interest. | 1. An implantable neural interface node, comprising: a base insertable into a mammalian cranium, wherein the base includes side walls that define a lumen;a plurality of neuroprobe assemblies coupled to the base; andan electronic component coupled to at least one of the neuroprobe assemblies;wherein the base further includes a shelf and the electronic component includes a circuit board on the shelf within the base, wherein the circuit board is substantially flat and circular and defines an open center. 2. The implantable neural interface node of claim 1, wherein the side walls of the base are adapted to extend towards a surface of the cranium. 3. The implantable neural interface node of claim 2, wherein the side walls of the base are adapted to extend towards a lower surface of the cranium. 4. The implantable neural interface node of claim 2, wherein the base includes a top surface adapted to extend towards a top surface of the cranium, further comprising a plug that forms a seal over the top surface of the base and around an edge of the lumen of the base. 5. The implantable neural interface node of claim 1, wherein the base includes a bottom surface adapted to extend part way through the cranium. 6. The implantable neural interface node of claim 1, wherein the base includes a top surface adapted to extend towards a top surface of the cranium. 7. The implantable neural interface node of claim 6, wherein the top surface is adapted to be flush with the top surface of the cranium. 8. The implantable neural interface node of claim 6, wherein the top surface includes a flange. 9. The implantable neural interface node of claim 6, further comprising a plug, coupled to the base, that forms a seal over the top surface of the base. 10. The implantable neural interface node of claim 9, wherein the plug includes a cover that forms a seal around an edge of the lumen of the base. 11. The implantable neural interface node of claim 1, wherein the base is sized to fit in a hole in the cranium. 12. The implantable neural interface node of claim 11, wherein the hole in the cranium has a diameter greater than 0.5 centimeters. 13. The implantable neural interface node of claim 1, wherein at least one neuroprobe assembly is positioned within the lumen of the base. 14. The implantable neural interface node of claim 13, wherein at least one neuroprobe assembly includes a microelectrode array. 15. The implantable neural interface node of claim 13, wherein at least one neuroprobe assembly includes a fluidic component. 16. The implantable neural interface node of claim 1, wherein the electronic component is stacked in the base. 17. The implantable neural interface node of claim 16, wherein the circuit board is stacked in the base. 18. The implantable neural interface node of claim 1, wherein the electronic component is coupled to the base. 19. The implantable neural interface node of claim 1, wherein the electronic component is permanently coupled to the b... | - | A61B005/04 A61N001/05
US08310231 | B2 | 2012-11-13 | NMR device for detection of analytes | 2010-07-27 | This invention relates generally to detection devices having one or more small wells each surrounded by, or in close proximity to, an NMR micro coil, each well containing a liquid sample with magnetic nanoparticles that self-assemble or disperse in the presence of a target analyte, thereby altering the measured NMR properties of the liquid sample. The device may be used, for example, as a portable unit for point of care diagnosis and/or field use, or the device may be implanted for continuous or intermittent monitoring of one or more biological species of interest in a patient. | 1. A device for the detection of one or more analytes in a liquid sample, the device comprising: (a) a reader configured to receive a signal from the liquid sample, the reader comprising (i) one or more permanent magnets defining a magnetic field, (ii) an RF excitation coil for transmitting an RF excitation to the liquid sample, (iii) an RF coil configured to detect an NMR relaxation response produced by exposing the liquid sample to the magnetic field and the RF excitation, (iv) an electrical element in communication with the RF coil, the electrical element configured to amplify, rectify, transmit, and/or digitize the signal corresponding to the NMR relaxation response, and (v) a display for indicating the presence or concentration of the one or more analytes; and (b) a support defining a well holding the liquid sample comprising magnetic particles and the one or more analytes and having the RF coil disposed about the well, the RF coil configured to detect an echo response produced by exposing the liquid sample to a bias magnetic field created using the one or more permanent magnets and the RF excitation, wherein the magnetic particles have binding moieties on their surfaces, the binding moieties operative to alter an aggregation of the magnetic particles in the presence of the one or more analytes. 2. The device of claim 1, wherein the well and the RF coil are configured to provide a filling factor of at least about 0.1. 3. The device of claim 1, wherein the well has a volume of less than about 300 μL. 4. The device of claim 1, wherein the well and the RF coil are configured such that the volume circumscribed by the RF coil is less than about 300 μL. 5. The device of claim 1, wherein the RF coil has a characteristic dimension from about 10 μm to about 1000 μm. 6. The device of claim 1, further comprising a tube for holding the liquid sample, the tube having a varying cross section. 7. The device of claim 1, wherein at least one of the binding moieties is operative to bind to at least one of the following: (i) the one or more analytes; (ii) another of the binding moieties; and (iii) an aggregation-inducing molecule in the liquid sample, thereby producing an aggregate of multiply-linked magnetic particles as a function of the presence or concentration of the one or more analytes in the liquid sample. 8. The device of claim 7, wherein an aggregate of multiply-linked magnetic particles is disaggregated as a function of the presence or concentration of the one or more analytes in the liquid sample. 9. The device of claim 1, further comprising a plurality of compartments for holding one or more populations of magnetic particles. 10. The device of claim 9, wherein magnetic particles having different binding moieties are disposed in different compartments for detection of the one or more analytes, the different binding moieties promoting aggregation or disaggregation of the magnetic particles in the presence of the one or more analytes. 11. T... | T2 Biosystems, Inc. | G01V003/00
US07887490 | B2 | 2011-02-15 | Method and device for removing respiratory artefacts from measured blood pressure data | 2006-09-28 | The invention relates to a method and a device for removing respiratory artefacts in invasive blood pressure measurements. The level of CO2 in the expired air is monitored during blood pressure measurement and used to approximate and remove the respiratory artifacts. In one embodiment, the respiratory frequency from the respiratory signal is extracted by identifying the end-tidal CO2 level for each breath. A model is derived for the respiratory artifacts using the respiratory frequency, and the model is subtracted from the measured blood pressure data. | 1. A method for removing artifacts caused by patient respiration in measured blood pressure data acquired invasively in the heart or an artery of the patient, comprising: acquiring the level of CO2 in the patient's expired air as a respiratory signal during the blood pressure measurement by a sensor;using the respiratory signal to approximate the artifacts caused by patient respiration by a data processing unit; andremoving the artifacts based on the respiratory signal by the data processing unit. 2. The method according to claim 1, further comprising after the CO2 acquiring step; generating a reference signal from the measured respiratory signal, wherein the reference signal has a predetermined function of the same length at the measured respiratory signal;deriving a model for the respiratory artifacts from the reference signal; andsubtracting the model from the measured blood pressure data. 3. The method according to claim 2, wherein the model deriving step further comprises, deriving a finite impulse response model and, adapting the model using an iterative algorithm. 4. The method according to claim 3, wherein the adapting step comprises, using a Recursive Mean Square Algorithm. 5. The method according to claim 4, wherein the adapting step comprises, adapting the finite impulse response model by choosing only blood pressure data sampled during the diastolic period. 6. The method according to claims 2, wherein the reference signal generating step comprises, analyzing the respiratory signal for generating the reference signal, and wherein the analyzing step comprises calculating the start, stop and maximum data points for each breath. 7. The method according to claim 1, further comprising a step of compensating for Respiratory Sinus Arrhythmia. 8. The method according to claim 7, the compensating step further comprising: recording an ECG signal during the blood pressure measuring step;calculating the instantaneous heart rate during the blood pressure measurement from the ECG signal;calculating a relative heart rate as the difference between the instantaneous heart rate and the mean heart rate;using the relative heart rate as input to an autoregressive model with exogenous input for modeling the systolic blood pressure of each heart beat;adapting the model to the variations in the measured systolic blood pressure data; andsubtracting the adapted autoregressive model from the measured systolic blood pressure data. 9. The method according to claim 8, wherein the model subtracting step comprises, subtracting the autoregressive model from the measured systolic blood pressure data in a manner so as to not alter the shape of the blood pressure waveform. | Siemens Aktiengesellscahft | A61B005/08
US07873403 | B2 | 2011-01-18 | Method and device for determining a three-dimensional form of a body from two-dimensional projection images | 2004-07-13 | A method for determining a three-dimensional form of a body from two-dimensional projection images includes producing at least one projection image of the body and producing corresponding model projection images from a generic model of the body. The body projection images and the model can be input into a computer-assisted processing unit and the generic model can be adjusted to the body by comparing image contents of the body projection images with the model projections images at a two-dimensional level only. The three-dimensional form of the body can be determined from the adjusted generic model. | 1. A method for determining a three-dimensional form of an anatomical body from two-dimensional projection images, said method comprising: producing at least one projection image of the body;producing corresponding model projection images from a generic model of the body;comparing the at least one body projection image and the model projection images from the generic body model;adjusting the generic model to the body based on the comparing of the at least one body projection image and the model projection images;determining the three-dimensional form of the body from the adjusted generic model;wherein adjusting the generic model to the body includes comparing image contents of the at least one body projection image with the model projection images, andwherein comparing the image contents of the at least one body projection image and the model projection images includes comparing (i) outer contours of the projection images and (ii) depth information in the projection images, wherein depth information includes features in the interior of the outer contours of the projection images. 2. The method as set forth in claim 1, wherein adjusting the generic model to the body includes adjusting the form and/or the position of the generic model. 3. The method as set forth in claim 1, wherein comparing image contents of the at least one body projection image with the model projection images is determined at only a two-dimensional level. 4. The method as set forth in claim 1, wherein the at least one body projection image is an x-ray image. 5. The method as set forth in claim 4, wherein the model projection images are digitally reconstructed x-ray images which originate from volume or surface data sets for the model. 6. The method as set forth in claim 1, wherein producing the model projection images includes determining projection parameters of the at least one body projection image while producing the at least one body projection image. 7. The method as set forth in claim 6, wherein the projection parameters are determined using a tracking system for the body and a recording device for the at least one body projection image. 8. The method as set forth in claim 1, wherein the generic model is automatically adjusted by way of statistical deformation. 9. The method as set forth in claim 1, wherein the generic model is adjusted at least partially interactively with the aid of user inputs, wherein a free-form deformation model is used. 10. The method as set forth in claim 1, wherein a generic surface data set model is used for producing corresponding model projection images, data from the generic surface data set model being used when adjusting the model. 11. The method as set forth in claim 1, wherein a generic volume data set model is used for producing corresponding model projection images, data from the generic volume data set model being used when adjusting the model. 12. The method as set forth in claim 1, wherein a generic volume data set model ... | BrainLAB AG | A61B005/05 G06K009/00 G06T017/00
US08317728 | B2 | 2012-11-27 | Method of using flocked swab for collecting biological specimens | 2010-10-13 | The present invention relates to a swab for collecting biological specimens of the type consisting of a rod terminating in a tip covered with fiber with hydrophilic properties to allow absorption of said specimens, wherein said fiber covers said tip in the form of a layer deposited by flocking. | 1. A method for collecting a biological specimen to be analyzed, the method being carried out by using a swab comprising a rod terminating in a tip, in which a layer of fibers is disposed by flocking to cover the tip, said layer of fibers having a thickness of 0.6 to 3 mm and a fiber count of 1.7 to 3.3 Dtex, the method comprising collecting a specimen to be analyzed by absorbing a quantity of 40 μl to 100 μl of specimen in said layer of fibers. 2. The method according to claim 1 in which said step of collecting a specimen to be analyzed is carried out by absorbing said quantity of 40 μl to 100 μl of specimen in said layer of fibers by capillarity effect. 3. The method according to claim 1 in which said specimen absorbed in said layer of fibers is a liquid specimen. 4. The method according to claim 1 wherein said specimen comprises a material selected from a group consisting of: bacteria, viruses, DNA, RNA, and a mixture of any of the preceding materials. 5. The method according to claim 1 wherein said quantity of specimen is absorbed in said layer of fibers by capillarity hydrophilic properties of the layer of fibers. 6. The method according to claim 1 in which said fibers are disposed perpendicularly to the surface of the tip by flocking. 7. The method according to claim 1 further comprising the step of storing and transporting said quantity of specimen in said layer of fibers. 8. The method according to claim 1 further comprising a step selected from the group consisting of: releasing a quantity of said specimen to be analyzed from said layer of fibers. 9. The method according to claim 1 further comprising the step of analyzing said quantity of liquid specimen after releasing said specimen from said layer of fibers. 10. The method according to claim 1 wherein said fibers define capillaries in said layer of fibers conferring hydrophilic properties by capillarity to the layer of fibers. 11. The method according to claim 1 in which said layer of fibers has capillarity hydrophilic properties and in which said step of absorbing a quantity of liquid specimen in said layer of fibers is carried out by capillarity effect. 12. The method according to claim 1 in which said layer of fibers is of uniform thickness. 13. The method according to claim 1 in which said layer of fibers is disposed on said tip in an ordered arrangement by flocking. 14. The method according to claim 1 further comprising: releasing a quantity of said specimen to be analyzed from said layer of fibers by swabbing. 15. The method according to claim 1 further comprising: releasing, from said layer of fibers, a quantity from 90% to 100% of said specimen absorbed. | Copan Italia S.p.A. | A61B005/00 B65D081/00 A61M035/00 C12M001/00 C12N003/00
US07885704 | B2 | 2011-02-08 | Intracavity probes and interfaces therefor for use in obtaining images and spectra of intracavity structures using high field magnetic resonance systems | 2005-11-15 | An intracavity probe for use with an MR system allows images or spectra of a region of interest within a cavity of a patient to be obtained. The probe includes a shaft, a balloon at one end thereof, and a coil loop within the balloon. The coil loop preferably includes two drive capacitors and a tuning capacitor, all of which in series. A junction node between the drive capacitors serves as a ground for electrically balancing the coil loop. Diametrically opposite the junction node, the tuning capacitor enables the coil loop to resonate at the operating frequency of the MR system. Across each drive capacitor is connected an output cable having an electrical length of SL+n(λ/4). The output cables terminate in a plug that is used to connect the coil loop to an interface device for the intracavity probe. | 1. An intracavity probe for use with a magnetic resonance (MR) system for obtaining images or spectra of a region of interest within a cavity of a patient, said intracavity probe comprising: (a) a first coil loop for receiving MR signals from the region of interest, said first coil loop having a plurality of capacitors therein, said plurality of capacitors including (i) a first drive capacitor and a second drive capacitor of approximately equal value serially connected within said first coil loop and at a junction node thereof forming a virtual ground for electrically balancing and impedance matching said first coil loop and (ii) a tuning capacitor serially connected within said first coil loop diametrically opposite said junction node of said drive capacitors and having a value selected to resonate said first coil loop at an operating frequency of said MR system;(b) a first output cable connected at one end thereof across said first drive capacitor; and(c) a second output cable connected at one end thereof across said second drive capacitor;wherein said first and said second output cables (i) each have an electrical length of SL+n(λ/4) wherein SL is a supplemental length whose reactance is of a same magnitude as a reactance of said drive capacitor corresponding thereto, n is an odd integer, and λ is a wavelength of the operating frequency of said MR system; and (ii) terminate in a plug therefor for connecting said first coil loop to an interface device for said intracavity probe. 2. The intracavity probe of claim 1 further comprising: (a) a flexible shaft having a tip at a distal end thereof, said tip being substantially more flexible than a remainder of said flexible shaft;(b) an inner balloon connected to said distal end of said flexible shaft and enclosing said tip thereof;(c) a non-stretchable material for securing said first coil loop to an anterior surface of said inner balloon; and(d) an outer balloon connected to said distal end of said shaft enclosing both said inner balloon and said first coil loop secured thereto, said outer balloon for use in positioning said inner balloon within the cavity of the patient;such that said non-stretchable material affects inflation of said inner balloon within said outer balloon to enable said first coil loop therein to be positioned approximate the region of interest for optimal reception of said MR signals therefrom. 3. The intracavity probe of claim 2 wherein said non-stretchable material focuses inflation of said inner balloon to force a posterior surface of said outer balloon against a wall of the cavity then forcing an anterior surface of said outer balloon against a correspondingly-shaped interior contour of the cavity thereby bringing said first coil loop approximate the region of interest for optimal reception of said MR signals therefrom. 4. The intracavity probe of claim 2 wherein said outer balloon has an anterior surface for conformably fitting to a correspondingly-shaped interior contou... | MEDRAD, Inc. | A61B005/055
US07883015 | B2 | 2011-02-08 | Host apparatus and method providing calibration and reagent information to a measurement apparatus which makes use of a consumable reagent in a measuring process | 2006-05-16 | Apparatuses and methods providing data such as, for example, calibration and reagent information (30) over a network (6) to a measurement apparatus (4) which makes use of a consumable reagent (18) in a measuring process are disclosed. The present invention permits a measurement apparatus (4) to download the necessary code assignment (24) and calibration data (32) of the consumable reagent (18) by means of an electronic communication protocol from a host (33). The host (33) can be a software application running on a host computer (34), another measurement apparatus (50, 58) with or without a code-key slot used to read a memory device providing the above mentioned information (30), a dedicated connectivity device, such as a docking station (44), a portable memory reader (48), a remote computer (54), and combinations thereof. | 1. A measurement system comprising: a network;a measurement apparatus provided with a first set of lot data files containing calibration data and lot identification codes each associated with respective ones of the calibration data, said measurement apparatus uses a consumable reagent in a measurement process, said consumable reagent having an associated lot data file containing calibration data and associated lot identification codes each associated with respective ones of the calibration data, wherein said measurement apparatus transmits over said network said associated lot data file of said consumable reagent when said associated lot identification codes fail to match any of said lot identification codes of the first set of lot data files provided on the measurement apparatus; anda host apparatus provided with a second set of lot data files containing calibration data and lot identification codes each associated with respective ones of the calibration data, said host apparatus receives said associated lot data file containing the associated lot identification codes each associated with respective calibration data of said consumable reagent transmitted from said measurement apparatus over said network, to subsequently identify whether the associated lot identification codes in the received lot data file match any of the lot identification codes of the second set of lot data files, and to distribute said lot data file of said consumable reagent containing the calibration data and the lot identification codes each associated with respective ones of the calibration data via the network when one of said associated lot identification codes matches one of the lot identification codes of the second set of lot data files. 2. The measurement system of claim 1 wherein said network is an in-building wireless network. 3. The measurement system of claim 1 wherein said network is a wired network. 4. The measurement system of claim 1 wherein said network is a peer-to-peer connection. 5. The measurement system of claim 1 wherein said network is selected from a group comprising an in-building wireless network, a wired network, a peer-to- peer connection, and combinations thereof. 6. The measurement system of claim 1 wherein said measurement apparatus is a handheld diagnostic device. 7. The measurement system of claim 1 wherein said measurement apparatus is a clinical chemistry analyzer. 8. The measurement system of claim 1 wherein said measurement apparatus comprises a lot identification code reader which is configured to read said associated lot data file of said consumable reagent. 9. The measurement system of claim 1 wherein said associated lot data file is provided in a machine-readable memory, and said measurement apparatus communicates with a reader for the machine-readable memory which reads said machine-readable memory in order to provide said associated lot data file to said measurement apparatus. 10. The measurement system of claim 1 wherein... | Roche Diagnostics Operations, Inc. | G06K007/00 G06K007/10 A61B005/00 G01N031/00
US07878978 | B2 | 2011-02-01 | Use of relaxin to increase arterial compliance | 2005-03-18 | The present invention provides methods for increasing arterial compliance. The methods generally involve administering to an individual in need thereof an effective amount of relaxin. The present invention further provides methods of increasing arterial compliance in individuals who have Type 1 or Type 2 diabetes. The present invention further provides methods of increasing arterial compliance in perimenopausal, menopausal, and post-menopausal women. The present invention further provides methods of increasing arterial compliance in individuals who have or who are at risk of developing age-associated arterial stiffness. | 1. A method for increasing arterial compliance in a subject, said method comprising: measuring global arterial compliance in said subject;determining that said global arterial compliance is diminished in said subject relative to global arterial compliance in a healthy subject; andadministering to said subject a pharmaceutical formulation comprising relaxin to increase arterial compliance in said subject. 2. The method of claim 1, wherein said global arterial compliance is measured from the diastolic decay of the aortic pressure waveform using the area method. 3. The method of claim 1, wherein said global arterial compliance is calculated as the stroke volume-to-pulse pressure ratio, and wherein said stroke volume is defined as the ratio of cardiac output to heart rate. 4. The method of any of claims 1-3, wherein said global arterial compliance in said subject is increased by at least 10% following administration of said pharmaceutical formulation to said subject. 5. The method of any of claims 1-3, wherein said global arterial compliance in said subject is increased by 15-20% following administration of said pharmaceutical formulation to said subject. 6. The method of any of claims 1-3, wherein said pharmaceutical formulation is administered to said subject at a predetermined rate so as to maintain a serum concentration of relaxin of from 0.5 to 80 ng/ml. 7. The method of any of claims 1-3, wherein said relaxin is recombinant human relaxin. 8. The method of claim 1, wherein said recombinant human relaxin is H2 relaxin. 9. The method of any of claims 1-3, wherein said pharmaceutical formulation is an injectable formulation. 10. The method of any of claims 1-3, wherein said pharmaceutical formulation is a sustained release formulation. 11. The method of any of claims 1-3, wherein said pharmaceutical formulation is delivered by continuous infusion. 12. The method of any of claims 1-3, wherein said subject is diagnosed with one or more ailments selected from the group consisting of: atherosclerosis, Type 1 diabetes, Type 2 diabetes, coronary artery disease, scleroderma, stroke, diastolic dysfunction, familial hypercholesterolemia, isolated systolic hypertension, primary hypertension, secondary hypertension, left ventricular hypertrophy, arterial stiffness associated with long-term tobacco smoking, arterial stiffness associated with obesity, arterial stiffness associated with age, systemic lupus erythematosus, preeclampsia, and hypercholesterolemia. 13. The method of any of claims 1-3, wherein said subject is a perimenopausal woman, a menopausal woman, or a post-menopausal woman. 14. A method for increasing arterial compliance in a subject, said method comprising: measuring local arterial compliance in said subject;determining that said local arterial compliance is diminished in said subject relative to global arterial compliance in a healthy subject; andadministering to said subject a pharmaceutical formulation comprising relaxin to incr... | University of Pittsburgh- Of The Commonwealth System of Higher Education | A61B005/00
US07871381 | B2 | 2011-01-18 | Noninvasive blood pressure monitoring system | 2006-08-10 | The preferred embodiment of the present invention comprises a single microprocessor-based interface that connects between a noninvasive blood pressure (NIBP) sensor and an invasive blood pressure (IBP) monitor or module. The interface effectively emulates an IBP transducer in such a way that the IBP monitor sees the interface as if it were a regular IBP transducer from a fluid-filled blood pressure monitoring line. It receives the signal from an NIBP sensor and determines the blood pressure corresponding to the signal. It accepts the excitation voltage provided by the IBP monitor. From the excitation voltage and a known transducer sensitivity which the IBP monitor is configured to work with, the interface emulates the IBP transducer output signal corresponding to the blood pressure. The interface also emulates the input and output impedances of the IBP transducer which the IBP monitor is configured to work with. Zeroing of the interface with the IBP monitor can be easily performed in a way that is similar to that for a fluid-filled system. A noninvasive system comprising a suitable NIBP sensor and this interface can be used as an alternative to the fluid-filled monitoring line. | 1. An interface configured to connect between a noninvasive blood pressure (NIBP) sensor and an invasive blood pressure (IBP) monitor comprising a first input configured to receive a measurement signal indicative of the NIBP of a subject;a second input configured to receive a transducer excitation signal provided by an IBP monitor; configured to determine the differential voltage of said excitation signal; and configured to determine the midpoint voltage of said excitation signal;at least one processor(s) configured to receive said measurement signal, said differential voltage and said midpoint voltage and emulate an output signal indicative of the IBP of a subject according to predetermined instructions; andan output configured to provide said output signal in a form suitable for input to an IBP monitor; and such that the midpoint of said output signal is substantially similar to that of said midpoint voltage. 2. An interface as claimed in claim 1 wherein said second input includes a voltage divider, a differential amplifier, or a combination of the two, configured to sense and condition said differential voltage. 3. An interface as claimed in claim 2 wherein said processor(s) is configured to receive both positive and negative ranges of said conditioned differential voltage through a circuit that includes an analog-to-digital converter (ADC). 4. An interface as claimed in claim 3 wherein said analog-to-digital converter (ADC) is a bipolar ADC. 5. An interface as claimed in claim 1 wherein said second input includes a voltage divider or a combination of a voltage divider and a differential amplifier, configured to sense and condition said midpoint voltage. 6. An interface as claimed in claim 5 wherein said processor(s) is configured to receive both positive and negative ranges of said conditioned midpoint voltage through a circuit that includes an analog-to-digital converter (ADC). 7. An interface as claimed in claim 6 wherein said analog-to-digital converter (ADC) is a bipolar ADC. 8. An interface as claimed in claim 1 wherein said processor(s) includes a bipolar digital-to-analog converter (DAC) to provide said midpoint voltage to said output. 9. An interface as claimed in claim 8 wherein said output includes a voltage divider or a combination of a voltage divider and a differential amplifier, configured to scale and condition the output signal of said digital-to-analog converter (DAC) such that the midpoint voltage of the scaled and conditioned signal is substantially similar to said midpoint voltage. 10. An interface as claimed in claim 5 wherein said output includes a circuit configured to scale and condition said output signal such that the differential voltage of said output signal is equal to the result of the following mathematical expression: VEXC×SENS×Pwherein VEXC is the root-mean-square (RMS) differential voltage across excitation terminals of the IBP monitor, SENS is the transducer sensitivity with which the IBP monitor is... | Healthstats International Pte Ltd | A61B005/02
US07861711 | B2 | 2011-01-04 | Method for judging the reverse connection of a flow sensor and a respiratory mechanics measuring module used therein | 2006-12-08 | A respiratory mechanics measuring module includes a flow sensor, a differential-pressure sensor and a sampled-signal processing circuit which are connected in series. The module further includes a flow sensor reverse connection judging module and an alarm. The flow sensor reverse connection judging module judges the reverse connection of the flow sensor. The alarm gives an alarm signal indicating the flow sensor reverse connection to the main controller when the flow sensor reverse connection judging module judges that the flow sensor is reversedly connected. | 1. A method for judging the reverse connection of a flow sensor in a respiratory mechanics measurement, comprising the following steps of: a. connecting the flow sensor to a respiratory mechanics measuring module;b. sampling consecutively airway pressures and flows of a respiratory gas flowing between a patient and an anaesthetic machine; andc. judging that the flow sensor is reversedly connected and sending an alarm signal indicating the reverse connection of the flow sensor to a main controller when any one of the following two instances appears:c1) when the airway pressure decreases to the lowest point, the flow has a negative value while the patient is inhaling; andc2) around the time when the pressure intensity decreases the fastest, the flow has a positive value. 2. The method for judging the reverse connection of a flow sensor according to claim 1, wherein the step of sampling consecutively the airway pressures and flows of the respiratory gas flowing between the patient and the anesthetic machine comprises the following steps of: collecting a time sequence of the pressure intensity x1, x2, x3 and x4 respectively corresponding to four consecutive sampling points of time within an arbitrary time interval and a flow value S corresponding to that time interval as the basis for judgment, and calculating three pressure difference variables using the equations d×1=x1−x2, d×2=x2−x3 and d×3=x3−x4; and wherein the two instances, the occurrence of anyone of which indicates the reverse connection of the flow sensor, are as follows:for step c1), x1, x2 and x3 are all smaller than zero and meanwhile x1 and x3 are equal to or greater than x2, and the flow value S is smaller than a predetermined first threshold value f1; andfor step c2), d×1, d×2 and d×3 are all greater than zero and meanwhile d×1 and d×3 are smaller than or equal to d×2, and the flow value S is greater than a predetermined second threshold value f2. 3. The method for judging the reverse connection of a flow sensor according to claim 2, wherein the first threshold value f1 is set at −2 liters/minute and the second threshold value f2 is set at 3 liters/minute, both of which are used for eliminating misjudgment caused by tiny and weak interference. 4. A respiratory mechanics measuring module, comprising a flow sensor, a differential pressure sensor and a sampled signal processing circuit which are connected in series, the sampled signal processing circuit transferring calculation results to a main controller, and further comprising a flow sensor reverse connection judging module and an alarm, wherein the flow sensor reverse connection judging module judges whether the flow sensor is reversedly connected based on output signals of the differential pressure sensor, and the alarm sends a prompting signal indicating the flow sensor reverse connection to the main controller when the flow sensor reverse connection judging module judges that the flow sensor is reversedly connected. 5. The res... | Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | A61B005/08 A61M016/00 A62B009/00 A62B027/00 G08B003/00 G08B005/00
US07890152 | B2 | 2011-02-15 | Handheld apparatus and method for the automated application of cosmetics and other substances | 2008-02-11 | An applicator head is provided for a reflectance modifying agent (RMA) applicator is moved across the skin by means of a floating ring having dispersed raised contact points to maintain a proper distance from the surface to be treated, reduce the influence of outside light during scanning, and limit smudging during deposition. During an application session, software on the computer uses a camera to sense aspects of color and texture on human features, calculates cosmetic enhancements, and uses the printer head to apply RMA precisely to the features to create those enhancements. Skin landmarks are used for registration. The head uses differential lighting by providing a sequence of directional lighting, with some exposures left dark to adjust for ambient light leakage. The exposures are co-synchronized in stacks, where each stack is a grouping of data about a particular instant of time during the scanning. | 1. A handheld reflectance modifying agent applicator, comprising a reflectance modifying agent applicator head comprising a case,an inkjet printer head,a camera,a cosmetic reservoir,a plurality of light sources,the applicator head in communication with a computing environment that processes a plurality of images of an area of skin captured using the camera, each image corresponding to a directional orientation of a light source of the plurality of light sources relative to the case, that calculates cosmetic enhancements to the area of skin and that determines a position of the applicator head relative to a skinmark in the area of skin based on the plurality of images to apply a reflectance modifying agent to the skinmark based on the calculated cosmetic enhancements, and a light source controller, such that the light source controller provides a lighting cycle comprising a plurality of lighting modes, the lighting modes comprising: a first mode where only light sources representing a first directional orientation are turned on,a second mode where only light sources representing a second directional orientation are turned on, anda third mode where only light sources representing a third directional orientation are turned on. 2. The handheld reflectance modifying agent applicator of claim 1, further comprising: a power and data cable from the applicator head to a computer of the computing environment; andcosmetic enhancement software that is processed by the computer. 3. The handheld reflectance modifying agent applicator of claim 1, further comprising a floating ring having a plurality of raised contact points that maintain a distance between the applicator head and the area of skin. 4. The handheld reflectance modifying agent applicator of claim 3, wherein the plurality of raised contact points comprise a plurality of pounce wheels. 5. The handheld reflectance modifying agent applicator of claim 4, wherein the plurality of pounce wheels have a plurality of points having a height of about 4 to 5 millimeters. 6. The handheld reflectance modifying agent applicator of claim 3, wherein the plurality of raised contact points comprise a comb-like walker. 7. The handheld reflectance modifying agent applicator of claim 1, further comprising a field lens assembly comprising: a telecentric field lens; andan LED assembly. 8. The handheld reflectance modifying agent applicator of claim 1, wherein the lighting modes further comprise a dark mode where no light sources are turned on. 9. The handheld reflectance modifying agent applicator of claim 3, wherein the computing environment processes the plurality of images, calculates the cosmetic enhancements and determines the position of the applicator head during an application session, the application session beginning when the applicator head touches the area of skin and ending then the applicator head is retracted from the area of skin. 10. A method of applying a reflectance modifying agent, the method ... | TCMS Transparent Beauty LLC | A61B005/00
US07890341 | B2 | 2011-02-15 | System and a method for providing integrated access management for peritoneal dialysis and hemodialysis | 2003-02-28 | A system and a method for tracking patients having renal disease, particularly Chronic Kidney Disease (CKD), are provided. The system and method track patients through a common system with multiple healthcare professionals, such as doctors, nurses and the like, having access to the system and updating the system based on information regarding the care, treatment, diagnosis, therapy, or the like provided to the patients. The system and method interface with remote machines and supply chains by a modem, via the internet or by wireless devices. The system and method provide a continuum of care for renal patients, i.e., CKD, peritoneal dialysis (PD), home hemodialyis (HHD), in-center hemodialysis and transplant within a single database associated with the system. The system and method track the access information of PD patients and/or HD patients. Specifically, the system and method track access placement, infections, non-infection related problems, medications and hospitalizations. The system and method may track PD and HD information in one comprehensive module. | 1. A method for organizing care for a patient undergoing a plurality of therapies for kidney disease using a system comprising at least one input device, at least one display device, at least one memory device and at least one processor, the method comprising: entering information about the patient using the at least one input device and storing the information about the patient in the at least one memory device;the processor organizing the information for presentation in a plurality of screens, wherein each screen presents information pertinent to a particular aspect of the care; andthe at least one display device displaying information about at least one of a renal peritoneal dialysis access management system and a renal hemodialysis access management system, wherein the information is presented in the plurality of screens and wherein access refers to a medical device providing access to the peritoneal cavity or the vascular system of the patient, the plurality of screens including at least: (i) an access summary screen including an implant date, an implant method, an implant location, an implant physician and a summary of patient problems sorted by infection problems and non-infection problems, and(ii) an access implant detail screen including: (a) at least one item of information selected from the group consisting of physician, date, location, facility, facility type, first use, graft type and graft configuration and (b) if the implant has been removed or abandoned, at least one item of information selected from the group consisting of: a physician that removed the implant, a date the implant was removed, a location the implant was removed and a reason the implant was removed. 2. The method of claim 1, further comprising using the information to track care for the patient and to help decide a subsequent treatment step for the patient. 3. The method of claim 1, further comprising presenting information in screens concerning a treatment history and a hospitalization summary. 4. The method of claim 1, further comprising presenting information in at least one screen concerning a clinical assessment. 5. The method of claim 1, further comprising presenting information concerning a dynamic assessment in at least one screen. 6. The method of claim 1, further comprising presenting information in at least one screen concerning infection problems and non-infection problems. 7. The method of claim 1, further comprising presenting information concerning prescriptions or medications for the patient in at least one screen. 8. The method of claim 1, wherein the information, includes a digital signature for a prescription by a prescribing physician. 9. The method of claim 1, further comprising adding information that is pertinent and hiding information that is not pertinent to the particular screen being viewed. 10. The method of claim 1, wherein the information may be entered and viewed at a remote location and at a central location. 11. The method... | Baxter International Inc. Baxter Healthcare S.A. | G06Q050/00 G06Q010/00 G06Q040/00 A61B005/00 G06F019/00
US08340747 | B2 | 2012-12-25 | Method and apparatus for trending a physiological cardiac parameter | 2010-09-21 | The present invention relates to an implantable cardioverter-defibrillator or pacemaker whose standard circuitry is used to trend a physiological cardiac parameter using intra-cardiac impedance measurements. The trend information may be used to predict the onset of a sudden cardiac death (SCD) event. By being able to predict the onset of an SCD event, patients and their physicians may be forewarned of a life-threatening event allowing them to respond accordingly. The trend information may also be used to predict the efficacy of cardiac-related medications, monitor progress of congestive heart failure, detect the occurrence of myocardial infarction, or simply track changes in sympathetic tone. | 1. A system for monitoring congestive heart failure, the system comprising: a measuring device configured to measure, using a lead, intra-cardiac impedance of a patient when the patient is typically sleeping;a derivation module configured to: derive a physiologic cardiac parameter from the measured impedance, the derived parameter including at least one of stroke volume, ejection fraction, and pre-ejection period; andderive sympathetic activity from the physiologic cardiac parameter;a trending module configured to trend the derived sympathetic activity over spaced time intervals to create trended sympathetic activity data; andan analyzing module configured to analyze the trended sympathetic activity data to monitor congestive heart failure status from the analysis. 2. The system of claim 1, wherein the analyzing module is configured to: compare the trended sympathetic activity data; anddetect a difference between the trended sympathetic activity data. 3. The system of claim 1, comprising a reporting module configured to report the trended sympathetic activity data to an outside source. 4. The system of claim 1, wherein the derivation module and the trending module are packaged with an implantable device. 5. The system of claim 1, comprising an external device configured to store the trended sympathetic activity data. 6. The system of claim 1, comprising a separately located device, wherein the system is configured to download the trended sympathetic activity data to the separately located device. 7. The system of claim 6, wherein the separately located device includes an external patient management system. 8. The system of claim 1, wherein the measuring device is configured to measure intra-cardiac impedance using an implantable cardiac rhythm management device. 9. The system of claim 1, wherein the measuring device is coupled to at least two electrodes configured to measure a voltage resulting from a current applied between the at least two electrodes, wherein the voltage is used to determine the intra-cardiac impedance. 10. A device for trending sympathetic activity, the device comprising: an impedance module configured to measure, using a lead, an intra-cardiac impedance of a patient at spaced time intervals when the patient is typically sleeping;a derivation module configured to: derive a physiologic cardiac parameter from the measured impedance, the derived parameter including at least one of stroke volume, ejection fraction, and pre-ejection period; andderive the sympathetic activity using the physiologic cardiac parameter;a trending module configured to generate trended sympathetic activity data using the derived sympathetic activity; andan analyzing module configured to analyze the trended sympathetic activity data to monitor congestive heart failure status from the analysis. 11. The device of claim 10, comprising a reporting module configured to report the trended sympathetic activity data to an outside source. 12. The device ... | Cardiac Pacemakers, Inc. | A61B005/053
US08386015 | B2 | 2013-02-26 | Integration of micro and macro information for biomedical imaging | 2009-10-23 | Macroscopic imaging data, such as from a CT, MR, PET, or SPECT scanner, is obtained. Microscopic imaging data of at least a portion of the same tissue is obtained. To align the microscopic imaging data with the macroscopic imaging data, intermediate data is also obtained. For example, photographic data is acquired at an intermediary stage of a process of preparing tissue for microscopic scan. The macroscopic and microscopic data are registered to the intermediary photographic data. Once registered to the intermediary data, the spatial relationship between the macroscopic and microscopic data is known and may be used for imaging or quantification. | 1. A method for biomedical imaging, the method comprising: obtaining macroscopic data representing a volume of tissue;obtaining photographic data representing cross-sections of the volume of the tissue;obtaining microscopic data representing the cross-sections of the volume of the tissue;registering, with a processor, the microscopic data to the photographic data with a first transform of coordinate systems for the microscopic data and the photographic data;registering, with the processor, the macroscopic data to the photographic data with a second transform of coordinate systems for the macroscopic data and the photographic data; andgenerating, on a display, an image as a function of the microscopic data, macroscopic data, or both microscopic and macroscopic data and as a function of the registering of the microscopic data to the photograph data and as a function of the registering of the macroscopic data to the photographic data;wherein obtaining the macroscopic data comprises obtaining in vivo imaging data of the tissue, wherein obtaining the photographic data comprises photographing the cross-sections of the tissue in vitro prior to slicing, and wherein obtaining the microscopic data comprises obtaining in vitro imaging data of slices after slicing. 2. The method of claim 1 wherein obtaining microscopic data comprises obtaining digital microscopy data from mounts of the tissue in slides, the mounts corresponding to the cross-sections of the volume. 3. The method of claim 1 wherein obtaining macroscopic data comprises obtaining computed tomography data, magnetic resonance data, positron emission tomography data, single photon emission tomography data, or combinations thereof. 4. The method of claim 1 wherein obtaining microscopic and macroscopic data comprises obtaining data with different imaging modalities, the imaging modality for the microscopic data having cellular, sub-cellular or molecular level resolution for the tissue and the imaging modality for the macroscopic data having a less resolution associated with tissue structure without cellular or more detailed structure. 5. The method of claim 1 further comprising: constructing a dataset representing the volume with the photographic data;surface segmenting the dataset; andsurface segmenting the macroscopic data;wherein registering the macroscopic data to the photographic data comprises aligning a first surface from the surface segmenting of the dataset with a second surface from the surface segmenting of the macroscopic data. 6. The method of claim 1 wherein generating the image comprises three-dimensionally rendering the image from the macroscopic and microscopic data as a function of a spatial mapping based on the registrations. 7. The method of claim 1 wherein generating the image comprises fusing the macroscopic data with the microscopic data into a dataset and rendering the image from the dataset. 8. The method of claim 1 wherein obtaining the photographic data comprises obt... | Siemens Aktiengesellschaft Beth Israel Deaconess Medical Center | A61B005/00
US08483809 | B2 | 2013-07-09 | Real-time electrocardiogram monitoring system and method, patch-type electrocardiograph, telecommunication apparatus | 2009-03-12 | The present invention relates to a real-time electrocardiogram monitoring system and method, a patch-type electrocardiograph and a telecommunication apparatus. The real-time electrocardiogram monitoring system includes: a patch-type electrocardiograph which is attached to a skin of a subject in the form of a patch, measures an electrocardiogram of the subject, and generates and transmits first electrocardiogram data of a first frame format for near-field transmission of the measured electrocardiogram signal; a relay-type communication device which receives the first electrocardiogram data from the patch-type electrocardiograph, extracts electrocardiogram data, and combines the extracted electrocardiogram with identification information of the subject to generate second electrocardiogram data of a second frame format for far-field transmission; a monitoring server which receives the second electrocardiogram data from the relay-type communication device, extracts the electrocardiogram data and the subject identification information, and classifies, stores and outputs in real time the electrocardiogram data based on the subject identification information; and a repository which classifies and stores the electrocardiogram data. The real-time electrocardiogram monitoring system further includes: a broadcasting server which relays the electrocardiogram stored by the monitoring server to a remote place; and a viewer which views the electrocardiogram data received from the broadcasting server. | 1. A real-time electrocardiogram monitoring system comprising: a patch-type electrocardiograph which is attached to a skin of a subject in the form of a patch-type form, configured to measure an electrocardiogram of the subject, and to generate and transmit first electrocardiogram data of a first frame format for near-field transmission of the measured electrocardiogram signal;a relay-type communication device configured to receive the first electrocardiogram data from the patch-type electrocardiograph, to extract electrocardiogram data, and combine the extracted electrocardiogram with identification information of the subject to generate second electrocardiogram data of a second frame format for far-field transmission;a monitoring server configured to receive the second electrocardiogram data from the relay-type communication device, to extract the electrocardiogram data and the subject identification information of the subject, to classify, to store and to output in real time the electrocardiogram data based on the subject identification information of the subject; anda repository configured to classify and store the electrocardiogram data, wherein in response to receipt of requests by viewers of the electrocardiogram data, an authenticator is configured to perform a multilevel authentication to determine accessible class of the electrocardiogram data, the multilevel authentication specifying a different accessible class of viewer which is classified based on each of the viewers and a data selector that is configured to transmit only the electrocardiogram data upon detection of a successful authentication corresponding to each level of the multilevel authentication based on each of the viewers. 2. The real-time electrocardiogram monitoring system of claim 1, further comprising: a broadcasting server configured to relay the electrocardiogram stored by the monitoring server to a remote place; and to display the electrocardiogram data received from the broadcasting server. 3. The real-time electrocardiogram monitoring system of claim 1, wherein the patch-type electrocardiograph comprises: an adhesive pad having three electrodes located on a single plane;an analog signal processor configured to perform amplification and noise cancellation of the electrocardiogram signal input to the three electrodes;a digital signal processor configured to convert an analog signal output from the analog signal processor into a digital signal and to process the electrocardiogram data into the first electrocardiogram data of the first frame format for near-field transmission; anda UHF band transmitter configured to transfer the first electrocardiogram data in the near-field. 4. The real-time electrocardiogram monitoring system of claim 3, wherein the analog signal processor comprises: an amplifier configured to receive and to amplify the electrocardiogram signal from the adhesive pad;a band pass filter configured to separate a signal of power supply noise and a si... | Keimyung University Industry Academic Cooperation Foundation | A61B005/04
US08494606 | B2 | 2013-07-23 | Photoplethysmography with controlled application of sensor pressure | 2009-08-19 | Systems, methods, and devices for obtaining consistently reproducible diagnostic measurements with a photoplethysmographic sensor are provided. In one embodiment, a method for obtaining such a diagnostic measurement includes applying a pressure between a photoplethysmographic sensor and a patient, increasing the pressure until the photoplethysmographic sensor outputs a plethysmographic waveform of minimal amplitude, decreasing the pressure by a predetermined fraction, and obtaining a diagnostic measurement using the photoplethysmographic sensor. The pressure may be applied using a pressure device that includes, for example, a clip, a wrap, an inflatable balloon or bladder, or an inflatable cuff or any combination thereof. | 1. A method comprising: applying a pressure between a photoplethysmographic sensor and a patient;increasing the pressure to a maximal pressure that occurs when the photoplethysmographic sensor outputs a plethysmographic waveform of minimal amplitude;decreasing the pressure by a predetermined fraction of the maximal pressure; andobtaining a diagnostic measurement using the photoplethysmographic sensor after decreasing the pressure by the predetermined fraction. 2. The method of claim 1, wherein the predetermined fraction is approximately one half. 3. The method of claim 2, wherein decreasing the pressure by approximately half comprises decreasing a pressure of between approximately 100 and 150 mm Hg to a pressure of between approximately 50 and 70 mm Hg. 4. The method of claim 1, wherein the pressure is applied using a pressure device comprising a clip, a wrap, an inflatable balloon or bladder, or an inflatable cuff or any combination thereof. 5. The method of claim 1, wherein the pressure is applied using a pressure device controlled by an electronic patient monitor. 6. The method of claim 1, wherein the pressure is increased until the photoplethysmographic sensor outputs a waveform of minimal amplitude at a predetermined wavelength. 7. The method of claim 6, wherein the predetermined wavelength is suitable for a pulse oximetry measurement or for a measurement of total hemoglobin. 8. The method of claim 1, wherein the obtained diagnostic measurement comprises a pulse oximetry measurement or a measurement of total hemoglobin, or a combination thereof. 9. A system comprising: a photoplethysmographic sensor configured to be applied to a tissue site on a patient and configured to provide a plethysmographic signal associated with the patient;a pressure device coupled to the photoplethysmographic sensor and configured to apply a pressure against the tissue site;a pressure sensor coupled to the photoplethysmographic sensor and configured to indicate the pressure applied by the pressure device; andan electronic patient monitor configured to receive and to monitor the plethysmographic signal, and to control the pressure device to increase the pressure applied to a maximal pressure that occurs when the plethysmographic signal reaches a minimal amplitude, and to control the pressure device to decrease the pressure applied by a predetermined fraction of the maximal pressure to enable a diagnostic measurement using the photoplethysmographic sensor. 10. The system of claim 9, wherein the pressure device comprises a clip, a wrap, an inflatable balloon or bladder, or an inflatable cuff, or any combination thereof. 11. The system of claim 9, wherein the pressure sensor comprises a piezo-resistive pressure sensor. 12. The system of claim 9, wherein the electronic patient monitor is configured to regularly determine the maximal pressure at a predetermined interval. 13. The system of claim 9, wherein the electronic patient monitor is configured to contro... | Covidien LP | A61B005/1455
US09636239 | B2 | 2017-05-02 | System and method for mapping activity in peripheral nerves | 2010-08-20 | Systems and methods are provided for controlling an entity in response to activity in a peripheral nerve comprising a plurality of fascicles. A multicontact electrode assembly is configured to record activity from the peripheral nerve. A processing component includes a sensor mapping component configured to quantify activity associated with a proper subset of the plurality of fascicles, an evaluation component configured to determine an adjustment of the status of the controlled entity from the quantified activity of the proper subset of the plurality of fascicles, and a controller configured to provide a control signal, representing the adjustment of the status of the controlled entity, to the controlled entity. | 1. A system comprising: a multicontact electrode assembly configured to record voltage activity from a peripheral nerve, wherein each contact of the multicontact electrode assembly is associated with a weight representing a presence of an ideal point source at a location associated with the respective contact, wherein a set of the weights are concentrating into a transformation matrix representing a spatial sensitivity of the peripheral nerve;a processing component comprising: a sensor mapping component configured to map the voltage activity detected at the multicontact electrode onto a spatial distribution representing the peripheral nerve based on the transformation matrix to determine a relative voltage activity of each ideal point source;an evaluation component configured to classify the relative voltage activity into an event class associated with a respective command for medical device based on at least one of the ideal point sources; anda controller component configured to provide a control signal to the medical device, wherein the control signal is determined based on the event class associated with the respective command for the medical device,wherein the medical device changes a function based by adjusting at least one physical parameter based on the control signal. 2. The system of claim 1, the sensor mapping component being configured to map the voltage activity at the peripheral nerve as a representation comprising a temporal component, such that at least one of the direction and the speed of the voltage activity is determined. 3. The system of claim 1, the sensor mapping component being configured to represent the activity measured at the multi-contact electrode assembly as a vector of voltages and multiply the vector of voltages by the transformation matrix. 4. The system of claim 3, the sensor mapping component being configured to construct a filter associated with at least one fascicle as a linear combination of a plurality of vectors from the transformation matrix. 5. The system of claim 1, the sensor mapping component being configured to quantify activity from at least two fascicles and determine a custom filter for a given fascicle from the quantified activity for the at least two fascicles. 6. The system of claim 1, the medical device comprising at least one of an external assistive device, and implanted medical device, and a prosthetic limb, the controller comprising an interface for providing the control signal to the medical device. 7. The system of claim 1, and the controller being configured to produce a stimulation pattern for the medical device, wherein the medical device comprises the multicontact electrode assembly. 8. The system of claim 7, wherein the evaluation component is configured to monitor reflex activity in the peripheral nerve in response to the stimulation pattern. 9. The system of claim 1, the controller being configured to produce a therapeutic stimulation pattern for a second multicontact elect... | Case Western Reserve University | A61B005/05 A61F002/70 A61F002/72 A61N001/05 A61B005/04 A61B005/00 A61N001/36
US09468734 | B2 | 2016-10-18 | Gas flow system, adaptor, and method | 2010-02-01 | An airway adaptor that measures the concentration of at least one of a specific component and the flow rate of a supply gas and a return gas. The airway adapter may include a tube having a flow path in which the supply gas and the return gas passes. The tube may include a first end having a first aperture for guiding the supply gas into the flow path, and second end having a second aperture for guiding the return gas. Further, the airway adapter may include a measurement device operable to measure at least one of the supply gas and the return gas in the air flow path and a partial shielding disposed at the first aperture and structured to shield a portion of the air flow path. The supply gas may be a respiratory gas and the return gas may be inhaled or exhaled breath of a person. | 1. An airway adaptor operable to measure a concentration of at least one of a specific component and a flow rate of a supply gas and a return gas, the airway adapter comprising: a tube having a flow path in which the supply gas and the return gas passes, the tube including only a single partial shielding, a first end having a first aperture for guiding the supply gas into the flow path, and a second end having a second aperture for guiding the return gas; anda measurement device operable to measure at least one of the supply gas and the return gas in the air flow path, whereinthe partial shielding is located only at an outermost end of the first end of the adapter, and structured to shield a portion of the air flow path, the partial shielding including two elongate structures that extend across an entire width of the flow path and intersect each other in the flow path. 2. The airway adapter according to claim 1, wherein the first end is configured to connect to a supply of a respiratory gas. 3. The airway adapter according to claim 1, wherein the second end is configured to conduct an inhaled or exhaled breath of a person. 4. The airway adapter according to claim 1, wherein the tube further includes an inner wall defining the flow path, and a pair of translucent windows, the translucent windows facing each other along a respective portion of the inner wall, and allowing detection light irradiated from an outside of the tube from the measurement device. 5. The airway adaptor according to claim 4, wherein the pair of translucent windows is provided along the inner wall located at a side of the flow path, and the partial shielding is structured to shield a portion of the flow path excluding at least one of a lower part and an upper part of the flow path. 6. The airway adaptor according to claim 1, wherein the partial shielding extends across and divides the first aperture into at least two apertures. 7. The airway adaptor according to claim 1, wherein the partial shielding is provided integrally with the tube. 8. The airway adaptor according to claim 1, wherein the measurement device further includes a respiratory concentration sensor structured to optically measure a concentration of the specific component. 9. The airway adaptor according to claim 1, wherein the measurement device further includes a respiratory flow rate sensor structured to measure the flow rate. 10. An airway adaptor operable to measure a concentration of at least one of a specific component and a flow rate of a supply gas and a return gas, the airway adapter comprising: a tube having a flow path in which the supply gas and the return gas passes, the tube including only a single partial shielding, a first end having a first aperture for guiding the supply gas into the flow path, and a second end having a second aperture for guiding the return gas; anda measurement device operable to measure at least one of the supply gas and the return gas in the air flow path, whereinth... | NIHON KOHDEN CORPORATION | A61B005/08 A61M016/08 A61B005/097 A61B005/083 A61B005/087 A61M016/10
US09955270 | B2 | 2018-04-24 | Bone conduction device fitting | 2009-03-31 | Methods and systems for fitting a bone conduction device are provided herein. These methods and systems comprise obtaining dynamic range parameters for a bone conduction device. These dynamic range parameters may include threshold and a maximum comfort levels for the bone conduction device. Once determined, the bone conduction device may use the dynamic range parameters in applying stimulation to a recipient. | 1. A method comprising: providing a test signal to a recipient of a bone conduction device using an acoustic fitting system;receiving a response from the recipient to the provided test signal;determining a dynamic range parameter based on the response;determining map data for the bone conduction device using the determined dynamic range parameter; andproviding the map data to the bone conduction device for use by the bone conduction device in providing stimulation to the recipient. 2. The method of claim 1, wherein the dynamic range parameter comprises: one or more of threshold level and comfort level. 3. The method of claim 2, wherein the bone conduction device is configured to provide stimulation over a plurality of frequencies, the method further comprising: determining a threshold level for at least one of the frequencies, wherein the threshold level corresponds to a minimum intensity of stimulation at which the recipient perceives sound in response to the provided test signal; andwherein determining map data comprises: determining the map data using the determined threshold level. 4. The method of claim 3, further comprising: determining a comfort level for the at least one of the frequencies; andwherein determining the map data comprises: determining the map data using the determined comfort level. 5. The method of claim 3, further comprising: determining a comfort level for the at least one of the frequencies,wherein determining the map data comprises: determining the map data using the determined comfort level. 6. The method of claim 1, further comprising: enabling the bone conduction device to clip or scale down a stimulation signal if the stimulation signal exceeds a threshold of the map data provided to the bone conduction device. 7. The method of claim 1, further comprising the action of: transforming the map data from a domain of the acoustic fitting system to a domain of the bone conduction device. 8. The method of claim 1, wherein: the response is a verbal response from the recipient. 9. The method of claim 1, wherein: wherein the action of providing the test signal to the recipient entails providing the bone conduction device instructions to produce a vibration, wherein the produced vibration corresponds to the provided test signal. 10. The method of claim 1, wherein: wherein the action of providing the test signal to the recipient entails providing the bone conduction device with an acoustic signal from a speaker, wherein the speaker is part of the acoustic fitting system, wherein the bone conduction device processes the acoustic signal to provide stimulation to the recipient, and wherein the provided stimulation corresponds to the provided test signal. 11. The method of claim 1, wherein: wherein the action of providing the test signal to the recipient entails providing the bone conduction device with an acoustic signal from a speaker located away from the recipient and away from the bone conduction device, wherein the ... | Cochlear Limited | A61N001/00 H04R025/00 A61M005/142
US09452252 | B2 | 2016-09-27 | Method and device for monitoring a vascular access and extracorporeal blood treatment device comprising a device for monitoring vascular access | 2009-03-24 | The invention relates to a method and a device for monitoring a vascular access during an extracorporeal blood treatment. The method and the device according to the invention are based on the monitoring of the difference between the venous pressure measured by a venous pressure sensor and the arterial pressure measured by an arterial pressure sensor (in the extracorporeal blood circuit. According to the method and the device according to the invention, a test function describing disturbances in the extracorporeal blood circuit is determined. Said test function is used to determine a noise-free differential pressure from the measured venous and arterial pressure, said differential pressure being evaluated in an arithmetic and evaluation unit to identify a defective vascular access. | 1. A method of monitoring a vascular access during extra-corporeal blood treatment in which blood flows from the vascular access via an arterial segment of an extra-corporeal blood circuit into a blood-treating element, and from the blood-treating element via a venous segment back into the vascular access, comprising the steps of: measuring a pressure in the venous and arterial segments of the extra-corporeal blood circuit in a first interval of time wherein there is a proper vascular access;determining a test function that defines interference affecting the measured arterial and venous pressures in the first interval of time, wherein the test function is an estimated function that defines interference affecting the arterial and venous pressures for a subsequent second interval of time;removing the interference from the individual arterial and venous pressures measured in the first interval of time;determining a first pressure differential that is freed from interference from the measured venous and arterial pressures in the first interval of time;measuring a pressure in the venous and arterial segments of the extra-corporeal blood circuit in the subsequent second interval of time;determining a second pressure differential from the measured venous and arterial pressures in the subsequent second interval of time;freeing the second pressure differential from interference by comparing the second pressure differential to the test function; andanalyzing the second pressure differential that has been freed from interference to detect a faulty vascular access. 2. The method according to claim 1, further comprising: determining the test function successively during the extra-corporeal blood treatment. 3. The method according to claim 2, further comprising: determining the difference between the second pressure differential and the test function; andconcluding that there is a faulty vascular access if the difference between the second pressure differential and the test function is more than a preset limiting value. 4. The method according to claim 3, wherein the test function is determined from a linear combination of trigonometrical functions. 5. The method according to claim 4, wherein the test function is defined by the following equation:P^VA⁡(t)=B0+∑k=1K⁢∑n=1N⁢[Ank⁢sin⁡(n⁢⁢ωk⁢t)+Bnk⁢cos⁡(n⁢⁢ωk⁢t)]where the coefficients A and B determine the contributions made by the higher harmonics of order n of the cyclic interference k of frequency ωk. 6. The method according to claim 1, further comprising: determining the difference between the second pressure differential and the test function; andconcluding that there is a faulty vascular access if the difference between the second pressure differential and the test function is more than a preset limiting value. 7. The method according to claim 1, wherein the test function is defined by the following equation:P^VA⁡(t)=B0+∑k=1K⁢∑n=1N⁢[Ank⁢sin⁡(n⁢⁢ωk⁢t)+Bnk⁢cos⁡(n⁢⁢ωk⁢t)]where the coefficients A and B deter... |  | A61M001/14 A61B005/021 A61M001/36
US07888125 | B2 | 2011-02-15 | Calibration of fluidic devices | 2006-03-24 | The present invention provides methods of calibrating a fluidic device useful for detecting an analyte of interest in a bodily fluid. The invention also provides methods for assessing the reliability of an assay for an analyte in a bodily fluid with the use of a fluidic device. Another aspect of the invention is a method for performing a trend analysis on the concentration of an analyte in a subject using a fluidic device. | 1. A method of improving the accuracy of calibrating a fluidic system, comprising: a) providing a system for detecting an analyte in a bodily fluid from a subject comprising a fluidic device for providing said bodily fluid, said fluidic device having a calibration assembly and a reader assembly for detecting the presence of said analyte;b) measuring one or more parameters that are fitted to a calibration curve associated with said fluidic device;c) comparing said one or more parameters with predetermined parameters associated with said fluidic device;d) adjusting a signal output from the fluidic device by multiplying a ratio of said predetermined parameters to said one or more parameters, wherein the ratio is selected from the group consisting of: a ratio of maximum signal level during factory calibration to maximum signal level measured within the assay, and a ratio of minimun signal level during factory calibration to minimum signal level measured within an assay. 2. The method of claim 1, wherein subsequent to said measuring, said one or more parameters are used to calculate an assay signal as a function of an analyte concentration. 3. The method of claim 1, wherein subsequent to said measuring, said one or more parameters are used to calculate an analyte concentration as a function of an assay signal. 4. The method of claim 1 wherein said one or more parameters is selected from the group consisting of: a maximum signal, a maximum signal, a minimum signal level half way between the maximum signal and the minimum signal, and a shape parameter. 5. The method of claim 1 wherein said predetermined parameters are parameters determined at the time the fluidic device is manufactured. 6. The method of claim 5 wherein said calibration curve is used to scale a signal that is indicative of said analyte concentration. | Theranos, Inc. | G01N031/00 A61B005/00
US08517953 | B2 | 2013-08-27 | Techniques for prediction and monitoring of coughing-manifested clinical episodes | 2010-07-23 | A method is provided for predicting an onset of a clinical episode, the method including sensing breathing of a subject, determining at least one breathing pattern of the subject responsively to the sensed breathing, comparing the breathing pattern with a baseline breathing pattern, and predicting the onset of the episode at least in part responsively to the comparison. Other embodiments are also described. | 1. Apparatus comprising: a sensor configured to sense a breathing-related sound of a subject selected from the group consisting of: coughing of a subject and wheezing of the subject, and to generate a signal responsively to the selected breathing-related sound;a breathing sensor configured to sense breathing of the subject and generate a breathing signal responsively thereto; anda control unit configured to: receive the signal generated responsively to the selected breathing-related sound;receive the breathing signal;in response to the received signals, determine a correlation between timing of a breathing cycle of the subject and timing of the breathing-related sound; andin response thereto, monitor a condition of the subject selected from the group consisting of: asthma, chronic obstructive pulmonary disease (COPD), congestive heart failure (CHF), cystic fibrosis (CF), and epilepsy. 2. The apparatus according to claim 1, wherein the control unit is configured to assess a progression of an episode of the condition. 3. The apparatus according to claim 1, wherein the sensor configured to sense the selected breathing-related sound is selected from the group consisting of: an acoustic sensor configured to sense the selected breathing-related sound and a mechanical vibration sensor configured to sense mechanical vibrations caused by the selected breathing-related sound. 4. The apparatus according to claim 1, wherein the sensor configured to sense the selected breathing-related sound is configured to sense the selected breathing-related sound without contacting the subject or clothes the subject is wearing. 5. The apparatus according to claim 1, wherein the sensor configured to sense the selected breathing-related sound is configured to sense the selected breathing-related sound without requiring subject compliance during the sensing. 6. A method comprising: sensing a breathing-related sound of a subject selected from the group consisting of: coughing of a subject and wheezing of the subject;sensing breathing of the subject with a sensor;determining a correlation between timing of a breathing cycle of the subject and timing of the breathing-related sound; andin response thereto, monitoring a condition of the subject selected from the group consisting of: asthma, chronic obstructive pulmonary disease (COPD), congestive heart failure (CHF), cystic fibrosis (CF), and epilepsy,the method being performed using apparatus including: a sensor configured to sense the selected breathing-related sound and to generate a signal responsively to the selected breathing-related sound;a breathing sensor configured to sense breathing of the subject and generate a breathing signal responsively thereto; anda control unit configured to: receive the signal generated responsively to the selected breathing-related sound;receive the breathing signal;in response to the received signals, determine the correlation between timing of the breathing cycle of the subject and timi... | EarlySense Ltd. | A61B005/08
US09011347 | B2 | 2015-04-21 | Methods and apparatus for determining breathing effort characteristics measures | 2009-07-02 | One or more respiratory characteristics of a patient are measured by coupling patient monitor apparatus (e.g., a photoplethysmograph (“PPG”)) to the patient in order to produce a patient monitor signal that includes signal indicia indicative of effort the patient is exerting to breathe. A breathing or respiratory effort signal for the patient is extracted from the patient monitor signal. A respiratory characteristic signal is extracted (at least in part) from the effort signal. This may be done, for example, on the basis of an amplitude feature of the effort signal and a relative time of occurrence of that amplitude feature. Alternatively, the respiratory characteristic signal may be based on a relationship between two amplitude features of the effort signal, with or without regard for specifics of the times of occurrence of those amplitude features. A breath air flow meter may also be coupled to the patient, if desired, in order to produce a flow signal. One or more of the respiratory characteristic measures may also be partly based on the flow signal. | 1. Apparatus for determining a respiratory characteristic of a patient comprising: apparatus for coupling to the patient to produce a patient monitor signal containing signal information indicative of effort the patient is exerting to breathe;first circuitry for extracting from the patient monitor signal a breathing effort signal indicative of the patient's exertion of effort in order to breathe; andsecond circuitry for: detecting first and second amplitude features of the effort signal;determining, using the effort signal, an elapsed time between the first amplitude feature of the effort signal and the second amplitude feature of the effort signal; andextracting a respiratory characteristic signal for the patient that is based on the elapsed time. 2. The apparatus defined in claim 1 wherein the elapsed time is from a minimum to a subsequent maximum in the amplitude of the effort signal. 3. The apparatus defined in claim 1 wherein elapsed time is from one maximum to a next subsequent maximum in the amplitude of the effort signal. 4. The apparatus defined in claim 1 a first precursor respiratory characteristic signal is indicative of the elapsed time, wherein the second circuitry is further adapted to extract from the effort signal a second precursor respiratory characteristic signal for the patient that is based on at least one other amplitude feature of the effort signal, and wherein the second circuitry is still further adapted to produce a further respiratory characteristic signal for the patient that is indicative of a ratio between the first and second precursor respiratory characteristic signals. 5. The apparatus defined in claim 4 wherein elapsed time is from a minimum to a subsequent maximum in the amplitude of the effort signal, and wherein the second precursor signal is indicative of an elapsed time from one maximum to a next subsequent maximum in the amplitude of the effort signal. 6. The apparatus defined in claim 4 wherein the second precursor signal is indicative of a magnitude of an amplitude feature of the effort signal. 7. The apparatus defined in claim 1 further comprising: a flow meter for coupling to the patient to produce a breathing air flow signal for the patient, wherein the flow signal is applied to the second circuitry, and wherein the respiratory characteristic signal is additionally based on at least one amplitude feature of the flow signal and a relative time of occurrence of that flow signal amplitude feature. 8. The apparatus defined in claim 7 wherein the amplitude feature of the flow signal indicates cessation of effective breathing by the patient, and wherein an amplitude feature of the effort signal indicates commencement of an increase in breathing effort by the patient. 9. The apparatus defined in claim 8 wherein the respiratory characteristic signal is additionally based on an elapsed time between the amplitude feature of the flow signal and the amplitude feature of the effort signal indicating commen... | Nellcor Puritan Bennett Ireland | A61B005/02 A61B005/08 A61B005/00 A61B005/1455 A61B005/087
US20110021899 | A1 | 2011-01-27 | CONDUCTIVE POLYMER COATINGS | 2010-07-07 | An electrically conductive coating composition that includes a polymeric mixture, an electrically conductive material dispersed within the polymeric mixture and, optionally, one or more bioactive agents is described. | 1. An electrically conductive coating composition comprising: (a) a polymeric mixture comprising a plurality of polymers, including a first polymer component selected from the group consisting of: poly(alkyl)(meth)acrylates and poly(aromatic(meth)acrylates); and a second polymer component selected from the group consisting of: ethylene copolymers with other alkylenes; polybutenes; aromatic group-containing copolymers; epichlorohydrin-containing polymers; poly(alkylene-co-alkyl(meth)acrylates); diolefin-derived, non-aromatic polymers and copolymers; poly(ethylene-co-vinyl acetate), and combinations and mixtures thereof; and(b) an electrically conductive material combined with the polymeric mixture. 2. The electrically conductive coating composition of claim 1, further comprising one or more bioactive agents dispersed within the polymeric mixture. 3. The electrically conductive coating composition of claim 1, wherein the electrically conductive material comprises a conductive polymeric material. 4. The electrically conductive coating composition of claim 3, wherein the conductive polymeric material is selected from the group consisting of poly(acetylene), poly(pyrrole), poly(thiophene), polyaniline, polythiophene, poly(p-phenylene sulfide), poly(para-phenylene vinylene), polyindole, polypyrene, polycarbazole, polyazulene, polyazepine, poly(fluorene), polynaphthalene, and combinations and mixtures thereof. 5. The electrically conductive coating composition of claim 3, wherein the conductive polymeric material is selected from the group consisting of polyaniline, polypyrrole, poly(3,4-ethylenedioxythiophene) (PEDOT), polyacetylene, Para-phenylene diamine (PPD), and combinations and mixtures thereof. 6. The electrically conductive coating composition of claim 1, wherein the electrically conductive material comprises biocompatible inorganic electrically conducting filler. 7. The electrically conductive coating composition of claim 6, wherein the inorganic electrically conducting filler is selected from the group consisting of: carbon nanotubes, metal micro/nano-particles, and carbon black. 8. The electrically conductive coating composition of claim 1 comprising at least about 0.5% by weight and up to about 50% by weight electrically conductive material. 9. The electrically conductive coating composition of claim 1, wherein the bioactive agent is selected from the group consisting of anti-inflammatory agents, anti-proliferative agents, antibiotics and antimicrobial agents. 10. A durable electrically conductive coating composition, comprising: (a) a non-toxic hydrophobic polymeric mixture;(b) an electrically conductive material dispersed throughout the polymeric mixture to form a polymeric coating having a resistivity and/or impedance sufficient to allow the electrical performance of the associated device; and(c) one or more bioactive agents combined with the polymeric mixture. 11. The medical device of claim 10, wherein the electrically conduc... | SurModics, Inc. | A61B005/0408 H01B001/16 H01B001/18 A61K047/02 A61K047/04 A61K047/34 A61K047/32 A61N001/04
US20110004096 | A1 | 2011-01-06 | ELECTROMAGNETIC CELLULAR TOMOGRAPHY | 2010-07-06 | The invention provides an electromagnetic cellular tomograph and methods of operating such a device. An array of structures is configured to apply probe signals to cells or tissues of interest that are held in a sample holder. The array also includes structures that can receive a response signal from the sample of interest. Data processing and control circuits are provided to manipulate and analyze the response and to allow the result to be recorded, transmitted to other data systems, or displayed to a user. | 1. An electromagnetic cellular tomograph, comprising: an electromagnetic structure configured to apply a probe signal to a sample of interest and to receive a response signal from said sample of interest, said electromagnetic structure having at least one input terminal connected to an electromagnetic signal application structure and having at least one output terminal connected to an electromagnetic signal sensing structure;a source circuit configured to provide an electromagnetic probe signal, said source circuit having at least one output terminal in electrical communication with said at least one input terminal of said electromagnetic structure configured to apply said probe signal to said at least one input terminal of said electromagnetic structure, and having at least one control input terminal configured to receive a control signal;a sensing circuit configured to sense said response signal from said sample of interest, said sensing circuit having at least one input terminal to receive said response signal from said at least one output terminal of said electromagnetic structure, said sensing circuit having at least one output terminal at which said sensed response signal is provided;a control circuit having at least one terminal in electrical communication with said at least one control input terminal of said source circuit, having at least one terminal in electrical communication with said at least one output terminal of said sensing circuit, and having at least one pair of input and output terminals for communication with an external electrical apparatus, said control circuit configured to control said source circuit, and said control circuit configured to communicate the sensed signals from the sensing circuit to said external electrical apparatus, and to accept control signals said external electrical apparatus; anda sample holder adjacent said electromagnetic structure, said sample holder configured to hold a sample of interest. 2. The electromagnetic cellular tomography of claim 1, wherein said electromagnetic structure is fabricated on a single chip. 3. The electromagnetic cellular tomography of claim 1, wherein said electromagnetic structure is configured as an N×M array, N and M being integers greater than or equal to 1. 4. The electromagnetic cellular tomography of claim 1, wherein said electromagnetic structure is configured to apply a probe signal comprising an electrical current to said sample of interest. 5. The electromagnetic cellular tomography of claim 1, wherein said electromagnetic structure is configured to apply a probe signal comprising an electrical voltage to said sample of interest. 6. The electromagnetic cellular tomography of claim 1, wherein said electromagnetic structure is configured to apply a probe signal comprising an electrical field to said sample of interest. 7. The electromagnetic cellular tomography of claim 1, wherein said electromagnetic structure is configured to apply a probe signal comprisi... | California Institute of Technology | A61B005/05
US20110046505 | A1 | 2011-02-24 | Impedance measurement process | 2008-08-07 | A method for use in performing impedance measurements on a subject. The method includes, in a processing system, determining at least one first impedance value, measured at a site using a first electrode configuration, determining at least one second impedance value, measured at the site using a second electrode configuration, and determining the presence, absence or degree of an anomaly using the first and second impedance values. | 1) A method for use in performing impedance measurements on a subject, the method including, in a processing system: a) determining at least one first impedance value, measured at a site using a first electrode configuration;b) determining at least one second impedance value, measured at the site using a second electrode configuration; and,c) determining the presence, absence or degree of an anomaly using the first and second impedance values. 2) A method according to claim 1, wherein the method includes using a tetrapolar electrode arrangement, the first and second electrode configurations using a different configuration of drive and measurement electrodes. 3) A method according to claim 2, wherein the method includes, in the processing system, determining an impedance value for each of four electrode configurations. 4) A′ method according to any one of claims 1 to 3, wherein the method uses apparatus including a signal generator, a sensor, a switching device and an electrode array having a number of electrodes, and wherein the method includes in the processing system, controlling the electrode configuration by: a) selectively interconnecting the signal generator and electrodes using the switching device; and,b) selectively interconnecting the sensor and electrodes using the switching device. 5) A method according to any one of claims 1 to 4, wherein the method includes, in the processing system: a) causing at least one drive signals to be applied to the subject;b) measuring at least one induced signal across the subject; and,c) determining at least one impedance value using an indication of the excitation signal and the induced signal. 6) A method according to any one of claims 1 to 5, wherein the method includes, in the processing system: a) determining impedance values at a number of different sites; and,b) determining an impedance map using the impedance values at each site. 7) A method according to claim 6, wherein the method includes, in the processing system: a) determining the presence of an anomaly at any one of the sites; and,b) determining the impedance map taking the anomaly into account. 8) A method according to claim 7, wherein the method includes, in the processing system, for a site having an anomaly, at least one of: a) excluding the site from the impedance map;b) modifying the impedance value determined for the site. 9) A method according to any one of claims 1 to 8, wherein the method includes, in the processing system: a) determining a difference between the first and second impedance values; and,b) determining the presence, absence or degree of an anomaly using the determined difference. 10) A method according to any one of claims 1 to 9, wherein the method includes, in the processing system: a) determining a difference between the first and second impedance values;b) comparing the difference to a reference; and,c) determining an anomaly depending on the result of the comparison. 11) A method according to any one o... | IMPEDIMED LIMITED | A61B005/053 G01R027/00 G06F019/00
US20110021891 | A1 | 2011-01-27 | NON-INVASIVE GLUCOSE MONITOR | 2010-07-22 | In a non-invasive glucose monitor is configured to, by using a heat conductive member having a body surface contacting part for contacting a surface of a human body at its one end, measures a first temperature adjacent to the body surface contacting part and a second temperature adjacent to the other end apart from the body surface. The computing device takes in data of the first and second temperature, environmental temperature, radiation heat from the body surface, reflection light caused by reflecting the light of the two different wavelength at the body surface contacting part. The computing device stores a relationship between parameters and blood glucose levels in advance, thereby converts the above-mentioned data to parameters, and calculates a blood glucose level by applying the parameters to the relationship. | 1. A non-invasive glucose monitor comprising: a heat flux measurement device including a heat conductive member having a body surface contacting part for contacting a surface of a human body at its one end, a first temperature sensor with which the heat conductive member is provided adjacent to the body surface contacting part thereof, and a second temperature sensor with which the heat conductive member is provided adjacent to the other end apart from the body surface thereof;an environmental temperature sensor that measures environmental temperature;a radiation heat sensor that measures radiation heat from the body surface;a light source that irradiates light having at least two different wavelengths toward the body surface contacting part;a photo sensor that senses reflection light caused by reflecting the light at the body surface contacting part;a computing device including a conversion part that converts respective outputs from the first temperature sensor, the second temperature sensor, the environmental temperature sensor, the radiation heat sensor and the photo sensor to respective parameters, and a processing part that stores a relationship between parameters and blood glucose levels and calculates a blood glucose level by applying the parameters converted from the respective outputs to the relationship, and a display part that displaying result outputted from the computing part. 2. The non-invasive glucose monitor according to claim 1, wherein the computing device is configured to, after starting temperature measurement at the first temperature sensor and the second temperature sensor in the heat flux measuring device, determines a regression line by making use of a least square method from temperature rise curves obtained at the first temperature sensor and the second temperature sensor with a predetermined time interval, stores relationships between the gradient of the regression line and equilibrium temperatures of the first temperature sensor and the second temperature sensor respectively in advance as database, and predict the equilibrium temperature for a subject to be monitored from the database. 3. The non-invasive glucose monitor according to claim 1 characterized in that as the body surface a finger is used. 4. The non-invasive glucose monitor according to claim 1 characterized in that the heat flux measuring device is cooled with a louver or a fan. 5. The non-invasive glucose monitor according to claim 1 characterized in that the database is prepared in a form of grouping respectively with respect to items of room temperature, sex, age and anamnesis. | HITACHI MEDIA ELECTRONICS CO., LTD. | A61B005/1455
US08892185 | B2 | 2014-11-18 | Apparatus, systems, and methods for localizing markers or tissue structures within a body | 2010-06-25 | Apparatus, systems, and methods are provided for localizing lesions within a patient's body, e.g., within a breast. The system may include one or more markers implantable within or around the target tissue region, and a probe for transmitting and receiving electromagnetic signals to detect the one or more markers. During use, the marker(s) are into a target tissue region, and the probe is placed against the patient's skin to detect and localize the marker(s). A tissue specimen, including the lesion and the marker(s), is then removed from the target tissue region based at least in part on the localization information from the probe. | 1. A system for localization of a target tissue region within a patient's body, comprising: a probe comprising a substantially flat patient contact surface and one or more antennas for transmitting electromagnetic signals into a patient's body and receiving reflected signals from the patient's body, through the patient contact surface, the probe analyzing the reflected signals to identify one or more objects within the patient's body; anda marker sized and configured for implantation directly in tissue within a patient's body independent of a delivery device used to deliver the marker; the marker comprising an active reflector to enhance identification of the marker by the probe. 2. The system of claim 1, wherein the marker comprises a plurality of beads carried by an elongate core member, the beads configured for enhancing reflection of signals from the probe to facilitate identification of the marker. 3. The marker of claim 2, wherein a predetermined surface finish is provided on one or more of the beads to modulate signals reflected from the one or more beads to enhance identification of the marker. 4. The system of claim 1, wherein the marker comprising an electronic circuit that modulates incident signals from the probe striking the marker such that reflected signals from the marker to enhance identification by the probe. 5. The system of claim 4, wherein the electronic circuit imposes a phase shift on the reflected signals. 6. The system of claim 1, wherein the marker is configured for enhancing a radar cross section (“RCS”) of the marker to enhance identification of the marker by the probe. 7. The system of claim 1, wherein the marker is biased to assume a nonlinear shape yet is resiliently flexible to allow the core member to be straightened to facilitate introduction of the marker into a target tissue region. 8. The system of claim 1, further comprising a delivery device for introducing the marker into a patient's body, the delivery device comprising a substantially rigid cannula having a sharpened tip for penetrating through tissue. 9. The system of claim 1, wherein the marker comprises an elongate core member, and an electronic circuit that modulates electromagnetic signals striking the marker such that reflected signals enhance identification of the marker. 10. The system of claim 9, wherein the core member is biased to a non-linear shape yet is resiliently flexible to allow the core member to be straightened to facilitate loading the marker into a delivery device. 11. The system of claim 9, wherein the electronic circuit comprises one or more of a diode and a FET transistor. 12. The system of claim 1, wherein the marker has a length no greater than four millimeters (4.0 mm). 13. The system of claim 1, wherein the marker is configured for implantation independent of any other devices or tethers. 14. The system of claim 1, wherein the marker includes first and second free and untethered ends. 15. A system for localization... | Cianna Medical, Inc. | A61B005/06 A61B005/05 A61B017/00 A61B019/00
US07885701 | B2 | 2011-02-08 | Registration pointer and method for registering a bone of a patient to a computer assisted orthopaedic surgery system | 2006-06-30 | A registration pointer includes a camera located in an elongated shaft having a distal end configured to be contacted to a bone of a patient to register the bone to a computer assisted orthopaedic surgery system. The camera may be a hemispherical camera and may include a panoramic camera and a wide-angle camera equipped with a fish-eye lens. The registration pointer is configured to transmit images received by the cameras to the computer assisted orthopaedic surgery system for display thereon. The computer assisted orthopaedic surgery system may be configured to generate an image, such as a hemispherical image, based on the images received from the registration pointer. | 1. A computer assisted orthopaedic surgery system comprising: a registration pointer including (i) a handle, (ii) an elongated shaft extending from handle, the elongated shaft having a distal end configured to be placed in contact with a bone of a patient, (iii) a panoramic camera positioned at the distal end, (iv) a wide-angle camera having a fish-eye lens positioned at the distal end, and (v) a sensor array to provide an indication of a location of the registration pointer;a display device;a processor communicatively coupled to the registration pointer and the display device; anda memory device electrically coupled to the processor, the memory device having stored therein a plurality of instructions, which when executed by the processor, cause the processor to: receive a first image from the panoramic camera and a second image from the wide-angle camera;generate a third image based on the first image and the second image; and display the third image on the display device. 2. The computer assisted orthopaedic surgery system of claim 1, wherein the third image is a hemispherical image. 3. The computer assisted orthopaedic surgery system of claim 1, wherein the panoramic camera and the wide-angle camera form a hemispherical camera. 4. The computer assisted orthopaedic surgery system of claim 1, wherein to receive the first image from the panoramic camera and the second image from the wide-angle camera comprises to wirelessly receive the first image from the panoramic camera and the second image from the wide-angle camera. 5. The computer assisted orthopaedic surgery system of claim 1, wherein the display device is a heads-up display device. 6. The computer assisted orthopaedic surgery system of claim 1, wherein to display the third image comprises to display the third image superimposed on a rendered image of a bone. 7. The computer assisted orthopaedic surgery system of claim 1, wherein the plurality of instructions further cause the processor to: receive a signal from the registration pointer; andstore the third image based on the signal. 8. The computer assisted orthopaedic surgery system of claim 1, wherein: the plurality of instructions further cause the processor to:determine a location of the registration pointer based on a location of the sensor array; and display indicia of the location of the registration pointer on the display device. 9. The computer assisted orthopaedic surgery system of claim 1, wherein the panoramic camera has a horizontal field of view of 360 degrees and a vertical field of view of 120 degrees. 10. The computer assisted orthopaedic surgery system of claim 1, wherein the wide-angle camera has a field of view of 60 degrees. 11. A method for displaying an image of a patient during the performance of an orthopaedic surgical procedure, the method comprising: receiving a first image from a panoramic camera positioned at a distal end of an elongated shaft of a registration pointer, the distal end being configured... | DePuy Products, Inc. | A61B005/05
US09042969 | B2 | 2015-05-26 | Heart rate correction system and methods for the detection of cardiac events | 2009-08-12 | A device for detecting a cardiac event is disclosed. Detection of an event is based on a test applied to a parameter whose value varies according to heart rate. Both the parameter value and heart rate (RR interval) are filtered with an exponential average filter. From these filtered values, the average change in the parameter and the RR interval are also computed with an exponential average filter. Before computing the average change in the parameter, large changes in the parameter over short times, which may be caused by body position shifts, are attenuated are removed, so that the average change represents an average of small/smooth changes in the parameter's value that are characteristic of acute ischemia, one of the cardiac events that may be detected. The test to detect the cardiac event depends on the heart rate, the difference between the parameter's value and its upper and lower normal values, and its average change over time, adjusted for heart rate changes. The upper and lower normal parameter values as a function of heart rate are determined from long term stored data of the filtered RR values and parameter values. Hysteeresis related data and transitory deviations from normal (e.g. vasospasm related data) are excluded from the computation of normal upper and lower parameter bounds. | 1. A cardiac event detection system for detecting a cardiac event in a patient comprising: (a) at least two sensors for sensing an analog signal from said patient's heart;(b) a device coupled to said sensors, said device having an analog-to-digital circuit system contained therein for digitizing said electrical signal to produce a digitized waveform of a heart rate of said patient over a plurality of heart beats, said heart rate defining a first parameter;(c) a processor electrically coupled to said analog-to-digital circuit system, said processor configured for: (i) computing a digitized heart rate time series data from said digitized waveform over said plurality of heart beats;(ii) low pass filtering and averaging of the digitized heart rate time series data, thereby producing a filtered and averaged heart rate time series signal;(iii) performing a test to detect the cardiac event, wherein the test is based on the filtered and averaged heart rate time series signal, said test comparing a second parameter value of at least one beat with a heart rate dependent threshold, wherein the threshold is selected as a function of according to the filtered and averaged heart rate time series signal. 2. The system of claim 1 wherein the heart rate time series data is filtered with a recursive filter. 3. The system of claim 1 wherein the recursive filter is an exponential average filter. 4. The system of claim 1 wherein the filter involves the computation of a moving average. 5. The system of claim 1 herein the processor and the device are disposed in a housing suitable for implantation into an animal. 6. The system of claim 1 wherein the processor is further configured to compute a change in heart rate value from the filtered heart rate time series signal, and wherein the test to detect the cardiac event is based on the change in heart rate value. 7. The system of claim 1 wherein the cardiac event is acute ischemia. 8. A cardiac event detection system for detecting a cardiac event in a patient comprising: (a) at least two sensors for sensing an analog signal from said patient's heart;(b) a device coupled to said sensors, said device having an analog-to-digital circuit system contained therein for digitizing said electrical signal to produce a digitized waveform of a heart rate of said patient over a plurality of heart beats, said heart rate defining a first parameter;(c) a processor electrically coupled to said analog-to-digital circuit system, said processor configured for: (i) computing both heart rate time series data and a second parameter value data from the digitized waveform over said plurality of heart beats;(ii) low pass filtering and averaging of the heart rate time series data and the second parameter value data, thereby a producing filtered and averaged heart rate time series data and filtered parameter value data;(iii) based on the filtered and averaged heart rate time series data and filtered parameter value data, generating upper and l... | Angel Medical Systems, Inc. | A61B005/04 A61B005/0452
US08378832 | B2 | 2013-02-19 | Breathing disorder treatment system and method | 2010-03-10 | The breathing disorder treatment system is a continuous monitoring and an actuated stimulation system for the treatment of breathing disorders, such as sleep apnea. The system includes computer readable memory in communication with a processor, and a threshold blood oxygen saturation (BOS) level is recorded therein. A BOS sensor is provided for continually measuring the blood oxygen saturation level in the patient. The BOS sensor is in communication with the processor, such that the measured BOS level is continually compared with the threshold BOS level. A timer is in communication with the processor, so that if a series of measured BOS levels are less than or equal to the threshold BOS level, a time of apnea or low BOS occurrence is recorded in the memory, and an alarm is actuated according to a calculated moving average period based upon the times of occurrence. | 1. A method of treating a breathing disorder, comprising the steps of: establishing a threshold blood oxygen saturation level;continually measuring blood oxygen saturation in a patient;continually comparing the measured blood oxygen saturation with the threshold blood oxygen saturation level;recording a time of low blood oxygen saturation occurrence in computer readable memory when the measured blood oxygen saturation is less than or equal to the threshold blood oxygen saturation level;calculating a time period between sequential times of low blood oxygen saturation occurrence;generating an alarm signal to actuate an alarm at regular intervals based upon the calculated time period to raise the patient's level of consciousness;generating the alarm signal when the measured blood oxygen saturation is less than or equal to the threshold blood oxygen saturation level; andestablishing a period of time of deactivation of the alarm signal. 2. The method of treating a breathing disorder as recited in claim 1, further comprising the steps of: establishing a present time;comparing the present time with the time of deactivation; andgenerating a deactivation signal to override the alarm signal and deactivate an alarm. 3. The method of treating a breathing disorder as recited in claim 1, further comprising the step of recording the measured blood oxygen saturation when the measured blood oxygen saturation level is less than or equal to the threshold blood oxygen saturation level. 4. A breathing disorder treatment system, comprising: a processor;computer readable memory in communication with the processor, a threshold blood oxygen saturation level being recorded in the computer readable memory;a blood oxygen saturation sensor for continually measuring a blood oxygen saturation level in a patient, the blood oxygen saturation sensor being in communication with the processor, the processor continually comparing signals from the sensor representing the measured blood oxygen saturation level with the threshold blood oxygen saturation level;a clock circuit in communication with the processor for keeping time, the processor being programmed to record the time of low blood oxygen saturation occurrence in the computer readable memory when the measured blood oxygen saturation level is less than or equal to the threshold blood oxygen saturation level, the processor being programmed to calculate a time period between sequential times of low blood oxygen saturation occurrence;an alarm in communication with the processor, the processor being programmed to generate an alarm signal actuating the alarm at regular intervals based upon the calculated time period in order to raise the patient's level of consciousness; andmeans for selectively and temporarily deactivating said alarm. 5. The breathing disorder treatment system as recited in claim 4, wherein said alarm is an implantable electrical stimulator. 6. The breathing disorder treatment system as recited in claim 4, wher... | - | G08G023/00
US20110021904 | A1 | 2011-01-27 | DYNAMIC PLANNING TOOL FOR USE IN CONTRAST-ENHANCED DYNAMIC SCAN IN MAGNETIC RESONANCE IMAGING | 2009-03-11 | The timing scheme of a dynamic contrast-enhanced MRI scan of the breast is critical for any subsequent analysis or CAD. This invention proposes a planning device configured to calculate the timing scheme of a dynamic MRI investigation, to be integrated in either the MRI scanner or a breast analysis or CAD software package. The planned scan may be transferred to the scanner by hand or by means of an ExamCard. The planning may also be used as the input for an analysis or CAD software pack-age. | 1. A device (40) for planning timing of a dynamic contrast-enhanced Magnetic Resonance Imaging scan, said device being configured to: receive (41) a user-defined parameter; andcalculate (42) a timing scheme for said dynamic contrast-enhanced Magnetic Resonance Imaging scan at least based on said user-defined parameter. 2. The device according to claim 1, wherein said calculation of said timing scheme is performed by calculating (43) a further parameter based on said user-defined parameter, and wherein said further parameter is required by said timing scheme. 3. The device according to claim 1, wherein the user-defined parameter comprises information regarding: an injection protocol comprising injection speed, contrast agent concentration (11), total amount of contrast agent, or injection delay;a scanning protocol comprising duration of an acquisition of the dynamic image dataset, acquisition delay, or sampling mode of k-space in time; ora patient characteristic comprising patient's weight or blood flow. 4. The device according to claim 2, wherein said further calculated parameter pertains to an injection protocol or scanning protocol. 5. The device according to claim 1, being comprised in an MRI analysis package, a CAD system, or an MRI scanner. 6. The device according to claim 5, wherein said timing scheme is utilized in subsequent image analysis. 7. The device according to claim 1, further configured to generate an ExamCard. 8. The device according to claim 1, further configured to control an injector. 9. The device according to claim 1, wherein the calculation of the timing scheme can use a parameter from a previously performed dynamic scan or a previously calculated timing scheme. 10. The device according to claim 1, further configured to visualize (43) said timing scheme on a display. 11. The device according to claim 1, further configured to simulate a dynamic scan based on said timing scheme. 12. The device according to claim 1, wherein said timing scheme comprises duration of an acquisition of a dynamic image dataset comprised in the dynamic scan, k-space ordering or cardiac flow. 13. The device according to claim 1, further configured to retrieve a parameter for the timing scheme from an external device. 14. The device according to claim 1, wherein said user-defined parameter comprises information regarding user preferences. 15. The device according to claim 14, configured to display user-defined parameters conflicting with each other in the calculated timing scheme. 16. A graphical user interface (30) for planning the timing of a dynamic contrast-enhanced Magnetic Resonance Imaging scan, said graphical user interface being configured to receive a user-defined parameter,calculate a timing scheme based on said user-defined parameter, andpresent said timing scheme to a user. 17. Use of the device according to claim 1 for the calculation of the timing scheme for a dynamic Magnetic Resonance Imaging scan of breast tissue for ... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/055
US08868151 | B2 | 2014-10-21 | Electrochemical impedance spectroscopy enabled continuous glucose monitoring sensor system | 2010-08-16 | The use of electrical impedance spectroscopy to adjust calibration settings in an in vivo monitoring system, such as an in vivo continuous glucose monitoring sensor. The adjustments can compensate for the condition of the sensor membrane in vivo. | 1. A method of testing the condition of a continuous analyte monitoring sensor in vivo comprising: obtaining a reference capacitance value for the continuous analyte sensor;inserting the continuous analyte sensor in vivo;performing electrochemical impedance spectroscopy on the in vivo sensor to obtain an in vivo capacitance value; andcomparing, by electrochemical impedance equipment, the in vivo capacitance value to the reference capacitance value to identify particular characteristics of the in vivo sensor, wherein the step of performing electrochemical impedance spectroscopy is conducted during in vivo insertion. 2. The method of claim 1, wherein the sensor further comprises an external counter electrode, wherein when the in vivo capacitance value is zero or intermittently zero during insertion, the sensor is not in sufficient contact with interstitial fluid or the external counter electrode is not in contact with a skin of a user. 3. The method of claim 2, further comprising warning the user to reinsert the sensor. 4. The method of claim 1, wherein when the in vivo capacitance value is greater than the reference capacitance value during insertion, the sensor coating is damaged, and a warning to the user to replace the sensor is provided. 5. The method of claim 1, wherein the step of obtaining the reference capacitance value is conducted after the step of inserting. 6. The method of claim 1, wherein the step of comparing is conducted by a Bode plot to identify phase shifts. 7. The method of claim 1, wherein the sensor further comprises a working electrode within a sensor membrane and a membrane layer overlies an outer surface of the sensor membrane. 8. The method of claim 7, wherein the comparing step identifies whether the sensor is in a proper position and whether the membrane layer is damaged. 9. The method of claim 1, further comprising determining whether the in vivo capacitance value is equal to, less than, or greater than the reference capacitance value during insertion to identify the particular characteristic of the sensor. 10. The method of claim 9, further comprising an external electrode, wherein when the in vivo capacitance value is zero or intermittently zero during insertion, the sensor is identified as being in an improper position. 11. The method of claim 9, wherein when the in vivo capacitance value is greater than the reference capacitance value during insertion, the sensor coating is identified as being damaged. 12. The method of claim 9, wherein when the in vivo capacitance value is less than the reference capacitance value during insertion, the sensor coating is identified as being intact. 13. The method of claim 1, wherein the in vivo capacitance value is a first in vivo capacitance value, the method further comprising: performing electrochemical impedance spectroscopy on the in vivo sensor to obtain a second in vivo capacitance value, andcomparing the second in vivo capacitance value to the first in vivo capa... | Bayer HealthCare LLC | A61B005/05 A61B005/145 A61B005/053
US09694358 | B2 | 2017-07-04 | Sampling and dispensing device | 2009-03-27 | The present invention relates to a device and method for easy collection, dilution, mixing and dispensing of a liquid specimen in a self-contained system. The present invention can be used in combination with numerous testing means for performing chemical, biochemical or biomedical qualitative or quantitative assays in the field of both clinical and hygiene testing. | 1. A device comprising: a container comprising an open end and a closed end;a first penetrable member disposed within said container and cooperating with said container to form a first sealed chamber, said first penetrable member is movable within said container and operable as a piston within said container;a sampler defining a passage for sampling a sample which is open at first and second ends, at least a portion of said passage extending from said first end comprising a capillary passage operable to draw said sample into said passage by capillary action, said sampler being movable relative to said container, said sampler comprising a penetrating means configured to penetrate said first penetrable member such that said passage is in communication with said first sealed chamber once said penetrable member has been penetrated. 2. The device according to claim 1 wherein said first penetrable member comprises a puncturable member and said penetrating means comprises a puncturing means configured to puncture said puncturable member. 3. The device according to claim 1 wherein said first penetrable member comprises a member having an aperture and said penetrating means comprises a plug, said plug being arranged to seal said aperture, wherein said plug is disposed at a distal end of said penetrating means such that moving said penetrating means into said penetrable member moves said plug into said chamber. 4. The device according to claim 1 wherein said first penetrable member comprises a member having an aperture extending entirely therethrough, the device further comprising a plug, said plug being arranged to seal said aperture, said penetrating means pushing said plug through said aperture and into said chamber. 5. The device according to claim 1 wherein said sampler is configured to slide inside said container. 6. The device according to claim 5 wherein said sampler comprises a protrusion configured to abut against said container and arranged to limit movement of said sampler into said container after moving said sampler into said container. 7. The device according to claim 5, wherein said sampler comprises a shaft or rod. 8. The device according to claim 1 wherein said sampler comprises an engagement means configured to engage with said penetrable member after said penetrating means penetrates said penetrable member. 9. The device according to claim 8 wherein said container is a cylinder in which said first penetrable member is slidable, said engagement means being configured to engage said first penetrable member to operate said first penetrable member as said piston. 10. The device according to claim 9 wherein said first penetrable member is operable to expel contents of said passage. 11. The device according to claim 10 wherein said first penetrable member is operable to expel said contents of said passage into said first sealed chamber. 12. The device according to claim 10 wherein said first penetrable member is operable to expel s... | ORION DIAGNOSTICA OY | A61M001/00 A61B005/00 B01L001/00 B01L003/02 A61B010/00 A61B005/15 A61J001/20 A61M005/28
US09697336 | B2 | 2017-07-04 | Electronically initiating an administration of a neuromodulation treatment regimen chosen in response to contactlessly acquired information | 2009-07-28 | In an embodiment, a method includes acquiring electronic data indicative of a property of a mammal without making physical contact with the mammal. The method also includes extracting digital information indicative of a physiological characteristic of the mammal from the electronic data indicative of the property of the mammal. The method further includes choosing from an electronically accessible treatment database a neuromodulation treatment regimen for administration to the mammal in response to the information indicative of a physiological characteristic of the mammal. The method also includes electronically initiating an administration of the chosen neuromodulation treatment regimen to a nervous system component of the mammal. | 1. A method comprising: acquiring electronic data indicative of a property of a mammal without making physical contact with the mammal;extracting digital information indicative of a physiological characteristic of the mammal from the electronic data indicative of the property of the mammal;choosing from an electronically accessible treatment database a neuromodulation treatment regimen for administration to a patient mammal based on the information indicative of a physiological characteristic of the mammal;determining that the mammal and the patient mammal have a substantial probability of being the same mammal; andelectronically initiating an administration of the chosen neuromodulation treatment regimen to a nervous system component of the mammal only if it is determined that the mammal and the patient mammal have the substantial probability of being the same mammal. 2. The method of claim 1, wherein acquiring electronic data indicative of a property of a mammal without making physical contact with the mammal includes: acquiring the electronic data indicative of the property of the mammal present in a space without making physical contact with the mammal; andwherein extracting digital information indicative of a physiological characteristic of the mammal from the electronic data indicative of the state of the mammal includesin an operation performed at a location remote from the space, extracting the digital information indicative of the physiological characteristic of the mammal from the electronic data indicative of a state of the mammal. 3. The method of claim 1, wherein acquiring electronic data indicative of a property of a mammal without making physical contact with the mammal includes: acquiring electronic data indicative of the property of the mammal present in a space without making physical contact with the mammal; andwherein choosing from an electronically readable treatment database a neuromodulation treatment regimen for administration to a patient mammal in response to the information indicative of a physiological characteristic of the mammal includes in an operation performed at a location remote from the space, choosing from the electronically readable treatment database the neuromodulation treatment regimen for administration to the patient mammal in response to the information indicative of the physiological characteristic of the mammal. 4. The method of claim 1, wherein acquiring electronic data indicative of a property of a mammal without making physical contact with the mammal includes: acquiring the electronic data indicative of the property of the mammal present in a space without making physical contact with the mammal;wherein extracting digital information indicative of a physiological characteristic of the mammal from the electronic data indicative of the property of the mammal includesin an operation performed at a location remote from the space, extracting the digital information indicative of the physiological ch... | GEARBOX, LLC | A61B005/024 G06F019/00 A61B005/05 A61B005/00
US10441344 | B2 | 2019-10-15 | Surgical apparatus comprising a nerve testing device | 2009-04-07 | A surgical apparatus for facilitating testing nerves during HF surgery. The surgical apparatus comprises a converter for converting a high-frequency treatment current into a nerve stimulating current. A controllable selector switch optionally feeds either the treatment current or the nerve stimulating current to the electrosurgical instrument. | 1. A surgical apparatus for facilitating testing nerves during high-frequency surgery, comprising: a high-frequency generator for generating a high-frequency treatment current,a converter for converting the high-frequency treatment current into a nerve stimulating current;connecting lines for feeding the high-frequency treatment current from the high-frequency generator to the converter; anda controllable selector switch for selectively feeding one of the high-frequency treatment current and the nerve stimulating current to a high-frequency electrosurgical instrument,wherein the converter comprises a rectifier circuit for generating a direct current as the nerve stimulating current. 2. The surgical apparatus of claim 1, wherein the converter comprises a pulse modulation unit for generating a temporally defined pulse or temporally defined pulse sequence of the nerve stimulating current. 3. The surgical apparatus of claim 1, wherein the converter comprises a constant current source for generating the nerve stimulating current, wherein the nerve stimulating current is constant. 4. The surgical apparatus of claim 3, wherein the constant current source is adjustable to predetermined current strengths. 5. The surgical apparatus of claim 1, wherein the converter comprises at least one driver device adapted to control output current parameters and voltage parameters of the high-frequency treatment current. 6. The surgical apparatus of claim 1, wherein the converter is arranged in a special housing or housing insert and configured such that said converter can be connected to the high-frequency electrosurgical instrument via a plug connection. 7. The surgical apparatus of claim 5, wherein the driver device is capable of adjusting the output current parameters and the voltage parameters of the high-frequency treatment current to match defined nerve stimulating current parameters. | ERBE ELEKTROMEDIZIN GMBH | A61B018/10 A61B018/12 A61B005/04 A61B005/00 A61B018/00 A61N001/36
US20110021892 | A1 | 2011-01-27 | SYSTEMS AND METHODS FOR RESPIRATION MONITORING | 2009-07-23 | According to embodiments, techniques for determining respiratory parameters are disclosed. More suitable probe locations for determining respiratory parameters, such as respiration rate and respiratory effort, may be identified. The most suitable probe location may be selected for probe placement. A scalogram may be generated from the detected signal at the more suitable location, resulting in an enhanced breathing band for determining respiratory parameters. Flexible probes that allow for a patient's natural movement due to respiration may also be used to enhance the breathing components of the detected signal. From the enhanced signal, more accurate and reliable respiratory parameters may be determined. | 1. A method for determining a respiratory parameter of a patient, comprising: receiving a first signal from a first probe at a first probe location;computing a first index for the first probe location, the first index based at least in part on the respiratory components in the first signal;receiving a second signal from a second probe at a second probe location, wherein the second probe location is different than the first probe location;computing a second index for the second probe location, the second index based at least in part on the respiratory components in the second signal;comparing the first index to the second index; andselecting, based at least in part on the comparing, one of the first probe location and the second probe location for determining at least one respiratory parameter. 2. The method of claim 1 further comprising determining the at least one respiratory parameter at the selected probe location. 3. The method of claim 1 wherein the at least one respiratory parameter is selected from the group consisting of respiration rate and respiratory effort. 4. The method of claim 1 wherein receiving a first signal from a first probe location comprises receiving a photoplethysmograph (PPG) signal from a pulse oximetry probe. 5. The method of claim 1 wherein computing a first index for the first probe location comprises: performing a continuous wavelet transform on the first signal to produce a first transformed signal; andgenerating a scalogram based at least in part on the first transformed signal. 6. The method of claim 5 further comprising: identifying a breathing band in the scalogram; anddetermining the energy associated with the breathing band. 7. The method of claim 6 wherein the first index is proportional to the determined energy associated with the breathing band. 8. The method of claim 6 further comprising: identifying a pulse band in the scalogram; anddetermining the energy associated with the pulse band. 9. The method of claim 8 wherein the first index is proportional to the ratio of the determined energy associated with the breathing band to the determined energy associated with the pulse band. 10. The method of claim 1 wherein the first probe location and the second probe location exhibit natural movement due to respiration of the patient. 11. A system for determining a respiratory parameter of a patient, comprising: a first probe capable of receiving a first signal from a first probe location;a second probe capable of receiving a second signal from a second probe location; anda processor capable of: computing a first index for the first probe location, the first index based at least in part on the respiratory components in the first signal;computing a second index for the second probe location, the second index based at least in part on the respiratory components in the second signal;comparing the first index to the second index; andselecting, based at least in part on the comparing, one of the first probe loc... | Nellcor Puritan Bennett Ireland | A61B005/1455 A61B005/08
US20110021943 | A1 | 2011-01-27 | NEURAL INTERFACE | 2009-01-16 | This invention relates to the neural interface comprising one or more electrode-coupled microchannels. Each electrode-coupled microchannel comprises a microchannel for accommodating regenerating nerve axons; and one or more electrodes exposed to the interior of said microchannel. The one or more electrodes may comprise (i) a stimulation array suitable for generating an extracellular stimulus current which induces an action potential in an axon in the microchannel; and/or (ii) a recording array which detects an extracellular signal in the microchannel indicative of an action potential in an axon in the microchannel. Interfaces as described may be useful in blocking pain, treating misrouted motor nerves; repair or treating nerve injury; or controlling or providing sensation and/or proprioception from a prosthesis. | 1. A neural interface comprising one or more electrode-coupled microchannels, each electrode-coupled microchannel comprising;a microchannel for accommodating regenerating nerve axons; and,one or more electrodes exposed to the interior of said microchannel.wherein the one or more electrodes comprise;(i) a stimulation array suitable for generating an extracellular stimulus current which induces an action potential in an axon in the microchannel; and/or(ii) a recording array for detecting an extracellular signal in the microchannel indicative of an action potential in an axon in the microchannel. 2. (canceled) 3. A neural interface according to claim 1 comprising two or more recording arrays which are axially spaced in the microchannel. 4. A neural interface according to claim 1 wherein the recording array comprises three electrodes connected in a tripole to an amplifier circuit. 5-6. (canceled) 7. A neural interface according to claim 1 wherein the recording array comprises a single electrode connected to a differential amplifier circuit. 8-20. (canceled) 21. A neural interface according to claim 1 wherein the stimulation array is suitable for generating an extracellular stimulus current which specifically induces an action potential in motor axons. 22. A neural interface according to claim 21 wherein the action potential induced in the axon is bi-directional. 23-25. (canceled) 26. A neural interface according to claim 21 wherein the action potential induced in the axon is uni-directional. 27. A neural interface according to claim 26 wherein the action potential is induced in the axon by a first stimulus current in a first region of the microchannel and a second stimulus current in a second region of the microchannel, wherein the first and second regions are longitudinally spaced in the microchannel and the first stimulus current is greater than the second stimulus current,such that the action potential is induced in the axon in the direction from the first region to the second region of the microchannel. 28-35. (canceled) 36. A neural interface according to claim 1 wherein the electrode coupled microchannel comprises a stimulation array and a recording array, said arrays being arranged in the microchannel such that an antidromic action potential is induced in the axon by the stimulation array in response to orthodromic action potentials detected in the axon by the recording array. 37-40. (canceled) 41. A neural interface according to claim 1 wherein the microchannels are formed from a biocompatible dielectric material. 42-47. (canceled) 48. A method of detecting an action potential in a nerve axon comprising; providing a neural interface according to claim 1 comprising an electrode-coupled microchannel containing a nerve axon in a conductive medium, and;detecting electrical signals in the medium which are indicative of an action potential in the nerve axon. 49-54. (canceled) 55. A method according to claim 48 comprising determin... | Cambridge Enterprise Limited | A61B005/04 A61N001/00
US07889900 | B2 | 2011-02-15 | Medical image viewing protocols | 2006-02-07 | A method, computer program and device for creating a viewing protocol for medical images is described. At least a first site of interest is identified in a medical imaging data set captured from the patient. Patient record data or computer assisted detection information can be used to identify the site of interest, which may be a potential lesion. A viewing protocol for displaying medical images to a user is planned. The viewing protocol includes a viewing path along which an image of the site of interest will be displayed. The viewing protocol also includes a trigger associated with the site of interest. When the trigger event is encountered the dynamic mode of image display is reconfigured to dynamically highlight the site of interest. The viewing protocol can then be used to control the display of images so as to provide, for example, a virtual endoscope. | 1. A computer implemented method for creating a viewing protocol for medical images relating to a patient, comprising: determining (2, 4) the existence of at least a first site of interest (86) in a medical imaging data set captured from the patient; andplanning (6, 52) a viewing protocol (88) for displaying medical images to a user, said viewing protocol including a viewing path which can display an image of a part of said first site of interest, wherein the viewing protocol (88) also includes at least a first trigger associated with the first site of interest (86) which can cause the dynamic mode of image display to be reconfigured (58) to dynamically highlight display (62) of the part of the first site of interest (86). 2. The method as claimed in claim 1, wherein the existence of a plurality of sites of interest is determined and the viewing protocol includes a plurality of triggers, each associated with a respective site of interest, and the dynamic mode of image display being reconfigured at least twice. 3. The method as claimed in claim 1, wherein determining the existence of the site of interest includes accessing an electronic patient record and using patient record data to determine the existence of the site of interest. 4. The method as claimed in claim 1, wherein determining the existence of the site of interest includes applying a computer assisted detection algorithm to the medical imaging data set to identify the site of interest. 5. The method of claim 1, wherein the site of interest corresponds to a potential or actual lesion or pathology. 6. The method of claim 1, wherein reconfiguring the dynamic mode of image display to highlight the site of interest, is selected from one, or a combination of: changing the time for which an image is displayed; changing the rate at which images are displayed; changing the magnification of at least a part of the image; changing the resolution of at least a part of the image; changing the direction of view of the image; and viewing the image from a plurality of different angles. 7. The method as claimed in claim 1, wherein the trigger makes the displayed site of interest an active entity with which a user can interact by entering a command. 8. The method as claimed in claim 1, wherein the trigger starts a macro which carries out a display routine for the site of interest. 9. The method of claim 1, further comprising: displaying a plurality of medical images according to the viewing protocol created. 10. A data processing device (22) for creating a viewing protocol for medical images relating to a patient, the device comprising: a memory (34) storing instructions storing a medical imaging data set captured from the patient;a data processor (30) designed to determine the existence of at least a first site of interest in a medical imaging data set,characterized in that the data processor comprises instructions to plan a dynamic image display protocol for displaying medical images to a user,... | Koninklijke Philips Electronics N.V. | G06K009/00 A61B005/05
US08764655 | B2 | 2014-07-01 | Remote medical diagnosis device including bio-mouse and bio-keyboard, and method using the same | 2009-03-02 | Disclosed herein are a remote medical diagnostic device that includes both a bio-mouse and a bio-keyboard, and a method using the same. | 1. A remote medical diagnostic device, comprising: a user computer connected to a wired or wireless internet to send and/or receive information regarding medical services;a bio-mouse connected to the user computer via a wired or wireless network, which includes at least one medical examination device selected from a group consisting of a temperature sensor, an image sensor, a ultrasound sensor, a stethoscope sensor, a blood pressure sensor and a wireless transmission/reception sensor to send/receive information of a capsule endoscope, as well as an optical mouse part;a bio-keyboard including a computer keyboard, a tray on which a bio-disc or optical disc may be horizontally loaded, a motor for rotating the loaded bio-disc or optical disc, and a bio-disc drive to control the same; anda central processing unit (CPU) providing a graphic user interface (GUI) to the user. 2. The diagnostic device according to claim 1, wherein the bio-disc comprises at least one device selected from a group consisting of a disease diagnostic analyzer, a nucleic acid hybrid analyzer, a biomaterial analyzer, a body constitutional analyzer for personalized medicine, a urine analyzer, a blood analyzer, an environmental pollution analyzer, a bio-chemical analyzer and an immunological analyzer. 3. The diagnostic device according to claim 1, wherein the bio-mouse includes at least one device selected from a group consisting of a selection switch to select a general mode or a medical examination mode, at least two light receiving lenses to switch light receiving characteristics according to the general mode and the medical examination mode, and a mouse wheel having a multi-colored LED to display the general mode and the medical examination mode. 4. The diagnostic device according to claim 1, wherein the bio-mouse further includes a position control device to control a position and a posture of the capsule endoscope inside the human body or an RF generation part to supply energy to the capsule endoscope. 5. The diagnostic device according to claim 4, wherein the position control device is a permanent magnet or electromagnet. 6. The diagnostic device according to claim 1, wherein the bio-mouse further includes a marker for tracking a position or an orientation of the bio-mouse. 7. The diagnostic device according to claim 1, wherein the bio-mouse further includes an RF ID reader to read an RF ID or contents of an RF IC contained in the bio-disc. 8. The diagnostic device according to claim 1, wherein the bio-mouse further includes a memory to store production information including version, production year and production ID of the bio-mouse as well as the data obtained from the medical examination device. 9. The diagnostic device according to claim 1, wherein the bio-keyboard further includes a memory to store product information including version, production year and production ID of the bio-keyboard as well as the data obtained from the bio-mouse. 10. The diagnostic devi... | Samsung Electronics Co., Ltd. | A61B005/00 G06F019/00 G01N021/00
US07877274 | B2 | 2011-01-25 | Messaging to remote patients in a networked health-monitoring system | 2003-09-16 | A networked health-monitoring system configured to collect and process patient health-related data. A plurality of remote patient sites each includes at least one display, a data management unit for collecting patient health-related data, a memory, and stored program instructions for generating health-monitoring related information on the display. A central server connects to the data management unit at each patient site. The system transmits messages for display at the patient sites and at least one of these messages is a message selected by a health care professional. | 1. A networked health-monitoring system, comprising: a plurality of remote patient sites, each site of said plurality of remote patient sites including a device having (i) at least one display, (ii) a data management unit configured to facilitate collection of health-related data of at least one patient, (iii) a memory circuit, and (iv) stored program instructions for collecting said health-related data and generating information in response to said health-related data; andat least one central server connectable for communication with the device at each remote patient site, wherein (A) said at least one central server is configured to (i) receive said health-related data from said device, (ii) analyze said health-related data, (iii) produce reports, including standardized reports, from said health related data, (iv) transmit said reports to a health care professional associated with said at least one patient, and (v) transmit at least one message for presentation on said at least one display and said at least one message is sent by said health care professional in response to at least one of said reports and (B) (i) the data management unit is physically separate from the at least one display, (ii) the memory circuit and the at least one display form a part of the device, (iii) the device comprises a handheld device, (iv) the handheld device is capable of displaying pictorial information related to said health-related data, (v) the handheld device is capable of displaying animated information related to said health-related data and (vi) the memory circuit comprises a program cartridge. 2. The system of claim 1, wherein the at least one message includes step-by-step instructions. 3. The system of claim 1, wherein the at least one message includes results of a test. 4. The system of claim 1, wherein the at least one message is a multi-line message. 5. The system of claim 1 wherein the health care professional generates the at least one message. 6. The system of claim 1, wherein the at least one message is educational or motivational. 7. The system of claim 1, wherein the at least one central server is configured to transmit the at least one message to a specific patient. 8. The system of claim 7, wherein the at least one central server is configured to transmit the at least one message automatically to the specific patient. 9. The system of claim 7, wherein the at least one central server enables the specific patient to choose when to receive the at least one message. 10. The system of claim 7, wherein the at least one message is stored before being transmitted to the specific patient. 11. The system of claim 1, wherein the device at each patient site is configured to monitor the health-related data of said at least one patient. 12. The system of claim 11, wherein the data management unit facilitates collection of said health-related data by receiving said health-related data from the device. 13. The system of claim 11, wherein the devi... | Health Hero Network, Inc. | A61B005/00 G06Q010/00 G06F019/00
US07885707 | B2 | 2011-02-08 | Method of scaling navigation signals to account for impedance drift in tissue | 2005-09-15 | A method for scaling the impedance measured during the course of an electrophysiology study accounts for impedance drifts. By scaling the impedance there is greater assurance that previously recorded positional information can be used to accurately relocate an electrode at a prior visited position. The scale factor may be based upon a mean value across several sensing electrodes. Alternatively, the scale factor may be calculated specifically with respect to an orientation of a dipole pair of driven electrodes. | 1. A method for scaling impedance measurements in an electrophysiology study, the method comprising providing an electrophysiology system capable of generating an electric field across a patient's body;driving a first dipole along a first axis to create an electric field across a patient's body;measuring at least one biologic impedance encountered by the electric field with respect to at least one surface sensor;calculating the mean of the absolute value of the at least one measured biologic impedance continuously as a function of time Pm(t);saving an initial calculated mean of the absolute value of the at least one measured biologic impedance as Pa; andmultiplying an impedance measurement between a measurement electrode and a reference electrode by the ratio of Pa/Pm(t). 2. The method of claim 1, wherein the at least one surface sensor comprises a plurality of surface sensors;the step of calculating further comprises calculating the mean of the absolute values of the at least one measured biologic impedance as a function of time at each of the plurality of surface sensors; andthe step of saving further comprises saving the initial calculated mean of the absolute values of the of the measured biologic impedance at each of the plurality of surface sensors. 3. The method of claim 1, wherein the step of driving further comprises driving a second dipole along a second axis across the patient's body;the step of calculating further comprises calculating means of the absolute values of the measured biologic impedance as a function of time at the at least one surface sensor with respect to both the first dipole and the second dipole separately as Pm(t,i);the step of saving further comprises saving an initial calculated mean of the absolute values of the measured biologic impedance for each of the first dipole and the second dipole separately as Pa(i); andthe step of multiplying further comprises multiplying an impedance measurement at the measurement electrode with respect to each of the first dipole and the second dipole by the respective ratio of Pa(i)/Pm(t,i). 4. The method of claim 3, wherein the at least one surface sensor comprises a plurality of surface sensors;the step of calculating further comprises calculating the mean of the absolute values of the measured biologic impedance as a function of time at each of the plurality of surface sensors; andthe step of saving further comprises saving the initial calculated mean of the absolute values of the of the measured biologic impedance at each of the plurality of surface sensors. 5. The method of claim 3, wherein the first axis and the second axis are generally orthogonal to each other. 6. The method of claim 3, wherein the step of driving further comprises driving a third dipole along a third axis across the patient's body;the step of calculating further comprises calculating means of the absolute values of the measured biologic impedance as a function of time at the at least one surface sensor... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/05
US07881949 | B2 | 2011-02-01 | Method for automated recording of patient actions | 2003-03-25 | A method is for automated recording of patient actions, for example in the course of telemedical care, in the course of clinical studies for the merit of a form of therapy or for trialing the effectiveness of drugs. The medical equipment or devices used in this context are provided with recording devices for identifying the patient so that the data obtained during use of the equipment or accessories can automatically be attributed to the patient in question. | 1. A method for recording of patient actions for tracking patient progress of participants in the course of clinical studies suitable for trialing the effectiveness of drugs, comprising: providing medical equipment able to obtain patient medical data from the patient at a treatment facility with a recording device for acquiring patient identification information from an electronic wirelessly communicating patient identification device carried by the patient, wherein said medical equipment is unaware of patient identification information prior to acquisition of said identification information;acquiring, from said patient identification device, data for identifying a patient and a patient location address;automatically associating data obtained during use of the medical equipment in a clinical study with the data identifying the patient and patient location address; andautomatically communicating said associated data via a communications network to an evaluation station for monitoring progress of the participants in the clinical study; andstoring said associated data in a storage device. 2. The method as claimed in claim 1, wherein the recording device for identifying the patient includes at least one of a chip card reader and a transponder and said patient identification device is individually assigned to a respective patient and is at least one of a counterpart chip card and a counterpart of the transponder. 3. The method as claimed in claim 2, wherein the storage device is located in at least one of the recording device for patient identification, in therapy equipment, in data collection equipment, in a technical accessory and in an individual chip card. 4. The method as claimed in claim 2, wherein the storage device is a data collection point residing in a computer connected to the recording device. 5. The method as claimed in claim 2, further comprising using an apparatus residing in a recording device in at least one of the medical equipment or device and the patient identification equipment for determining at least one of treatment facility location and patient location. 6. The method as claimed in claim 1, wherein said storage device is at least one of in the recording device for patient identification, in therapy equipment, in data collection equipment in a technical accessory and in an individual chip card. 7. The method as claimed in claim 6, wherein the data obtained and associated with the patient are stored in a computer connected to the recording device. 8. The method as claimed in claim 1, wherein said storage device is a data collection point located at the treatment facility. 9. The method as claimed in claim 1, wherein said evaluation station is a central evaluation station remotely located from the treatment facility and said storage device is a central computer networked to a plurality of different data collection points at a plurality of different treatment facilities. 10. The method as claimed in claim 9, further com... | Siemens Aktiengesellschaft | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07869854 | B2 | 2011-01-11 | Apparatus for magnetically deployable catheter with MOSFET sensor and method for mapping and ablation | 2006-02-23 | A mapping and ablation catheter is described. In one embodiment, the catheter includes a MOSFET sensor array that provides better fidelity of the signal measurements as well as data collection and reduces the error generated by spatial distribution of the isotropic and anisotropic wavefronts. In one embodiment, the system maps the change in potential in the vicinity of an activation wavefront. In one embodiment, the mapping system tracks the spread of excitation in the heart, with properties such as propagation velocity changes. In one embodiment, during measurement, the manifold carrying the sensor array expands from a closed position state to a deployable open state. Spatial variation of the electrical potential is captured by the system's ability to occupy the same three-dimensional coordinate set for repeated measurements of the desired site. In one embodiment, an interpolation algorithm tracks the electrogram data points to produce a map relative to the electrocardiogram data. | 1. A system for mapping endocardial tissue, said system, comprising: a catheter having a catheter distal end for insertion into a patient;a plurality of sensors associated with said distal end, wherein said plurality of sensors are configured to obtain measurements of electrical potential of endocardial tissue in contact with said sensors such that current flow in each sensor is proportional to said electrical potential of said endocardial tissue, wherein each of the plurality of sensors comprise: a housing in contact with said tissue;a transistor with a first P-N junction between a first gate junction and a source terminal and between said gate junction and a drain terminal, wherein said source terminal and drain terminal are connected to said housing in contact with said endocardial tissue such that said gate junction assumes said potential of said endocardial tissue;a computer accessible memory that stores said measurements of said electrical potential of said endocardial tissue obtained by said plurality of sensors over a time duration; andone or more computer processors that correlate said measurements of said electrical potential of said endocardial tissue with a location of said distal end of said catheter to create an endocardial map of said tissue. 2. The system of claim 1, said distal end of said catheter further comprising a magnet. 3. The system of claim 2, further comprising a deployment mechanism proximate to said catheter distal end and configured to increase a separation between a distal end of a first deployable member and a distal end of a second deployable member. 4. The system of claim 2, wherein said deployment mechanism further comprises a coil that is configured to increase said separation between said distal end of said first deployable member and said distal end of said second deployable member when current is provided to said coil. 5. The system of claim 1, wherein said said PN junction is further in communication with a second gate. 6. The system of claim 1, wherein P material of said first PN junction is electrically closer to said tissue than N material and said second electrical contact comprises a second PN junction oriented such that N material of said second PN junction is electrically closer to said tissue than P material. 7. The system of claim 1, further comprising a second PN junction, wherein an said first and second PN junctions are configured with opposite polarity. 8. The system of claim 1, further comprising an antenna in communication with at least one of said plurality of sensors. 9. The system of claim 1, wherein said catheter comprises a lumen. 10. The system of claim 1, further comprising an operator interface unit. 11. The system of claim 1, further comprising: a magnetic field source for generating a magnetic field, said magnetic field source comprising a first coil corresponding to a first magnetic pole and a second coil corresponding to a second magnetic pole, wherein said first magneti... | Magnetecs, Inc. | A61B005/04
US07890351 | B2 | 2011-02-15 | Managing utilization | 2008-03-31 | A computer-implemented method includes receiving a request from a consumer of medical services to engage with a provider of medical services in real time over the Internet using a computer-based brokerage system, accessing stored information related to past engagements of the consumer using the computer-based brokerage system, and determining allocation of a fee between the consumer and a third party for an engagement between the consumer and the provider based at least in part on the information related to the past engagements of the consumer. | 1. A computer-implemented method, comprising: receiving by a computer a request from a consumer of medical services to engage with a provider of medical services in real time over the Internet using a computer-based brokerage system;determining by the computer a fee for an engagement between the consumer and the provider;accessing stored information related to past engagements of the consumer using the computer-based brokerage system;accessing stored information related to a medical history of the consumer;accessing stored information related to insurance covering care for the consumer; anddetermining by the computer an allocation of the fee between the consumer and a third party for the engagement between the consumer and the provider based at least in part on the information related to the past engagements of the consumer, the information related to the medical history of the consumer, and the information related to the insurance covering the consumer;displaying a price associated with consulting with a particular service provider to the consumer on a user interface, wherein displaying the price comprises displaying a total cost, displaying an amount to be paid by the insurance company, and displaying an amount to be paid by the consumer. 2. The method of claim 1, wherein the information related to past engagements comprises a total number of engagements by the consumer during a period of time. 3. The method of claim 1, wherein the information related to past engagements comprises a total financial expenditure of an insurance company during a period of time. 4. The method of claim 1, wherein determining the allocation of the fee between the consumer and a third party comprises determining a percentage of a total cost to be paid by the consumer. 5. The method of claim 1, wherein determining the allocation of the fee between the consumer and a third party comprises allocating the entire fee to the consumer if an amount paid by the third party is greater than a threshold set by the third party. 6. The method of claim 1, wherein the third party comprises an insurance carrier. 7. The method of claim 1, further comprising: receiving a payment scale from the insurance company. 8. The method of claim 1, determining the distribution of cost between the consumer and the insurance company comprises: determining a total amount the insurance company has paid for a particular consumer using the system during a period of time prior to receiving the request; andcomparing the total amount to a maximum amount an insurance company will provide in co-payments;if the total amount exceeds the maximum amount the insurance company will provide in co-payments, providing an indication to the consumer that the consumer will be responsible for the entire cost. 9. The method of claim 1, further comprising: identifying one or more service providers satisfying at least some of attributes in the request from the consumer;sorting the identified one or more service prov... | American Well Corporation | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07867175 | B2 | 2011-01-11 | Body fluid sampling unit | 2007-09-27 | A body fluid sampling unit includes a puncture device including a needle having a sharp needle point and a drive mechanism for operating the needle to puncture a living body surface with the needle point, a tip including an introducing section for introducing a body fluid flowing out from a puncture portion of the living body surface punctured with the needle point, a test paper for detecting a predetermined component in the body fluid, and a tip body mounted to a component measuring device for measuring the quantity and/or property of the predetermined component detected by the test paper; and a case including a mechanism containing section for containing the puncture mechanism so that the needle point of the puncture mechanism can be moved by the drive means, and a sampling implement containing section for containing the body fluid sampling implement so that the tip can be taken out. | 1. A body fluid sampling unit comprising: a container having an interior;a needle positioned in the interior of the container, the needle possessing a sharp needle point at a distal end of the needle;at least one spring positioned in the interior of the container;the at least one spring being operatively engageable with the needle and being configured to accumulate a spring bias to be applied to the needle to move the needle in a puncturing direction to puncture a living body surface with the needle point;a body fluid sampling implement removably positioned in the interior of the container:the body fluid sampling implement comprising: an introducing section for conveying body fluid flowing out from a puncture portion of the living body surface punctured with the needle point; a reagent containing part positioned adjacent the introducing section to receive the body fluid conveyed by the introducing section, the reagent containing part comprising a reagent adapted to interact with a predetermined component in the body fluid; and a mounted section configured to be mounted to a component measuring device for measuring a quantity and/or property of the predetermined component; andwherein the container comprises an inner tube movably positioned within an outer tube, the body fluid sampling implement being positioned inside the inner tube, the outer tube comprising first and second openings, the inner tube comprising a pawl engaging the first opening, the body fluid sampling implement being configured to be mounted to the component measuring device while the body fluid sampling implement is contained in the sampling implement containing section by pushing the component measuring device towards the body fluid sampling implement to engage the body fluid sampling implement with the component measuring device, said pushing causing the inner tube to move within the outer tube to cause the pawl to be released from the first opening and to engage the second opening. 2. The body fluid sampling unit as set forth in claim 1, wherein the interior of the container comprises a sampling implement containing section in which is positioned the body fluid sampling implement, the sampling implement containing section being covered by a removable membrane member which seals the sampling implement containing section in an air-tight manner. 3. A body fluid sampling unit comprising: a puncture mechanism comprising a needle possessing a sharp needle point at a distal end of the needle, and drive means for operating the needle to puncture a living body surface with the needle point;a body fluid sampling implement comprising an introducing section for introducing a body fluid flowing out from a puncture portion of the living body surface punctured with the needle point, a detecting section for detecting a predetermined component in the body fluid introduced through the introducing section, and a mounted section configured to be mounted to a component measuring device for meas... | Terumo Kabushiki Kaisha | A61B005/00 A61F002/02 B65D081/00
US07890156 | B2 | 2011-02-15 | Medical image display method and apparatus | 2004-09-09 | A medical image display method of this invention includes an organ display step for displaying an image of a predetermined organ including a desired blood vessel in a medical image on a display unit; a blood vessel selection step for selecting the desired blood vessel from the image of the predetermined organ displayed by the organ display step and designating a specific position of the desired blood vessel; and an image display step for displaying a pseudo image in a pseudo imaging mode as if a contrast, agent were made to flow through the blood vessel selected in the blood vessel selection step. The pseudo image of the flow of the contrast agent in the blood vessel assists doctor in conducting diagnoses and therapies. | 1. A medical image display method comprising: an organ display step for displaying an image of a predetermined organ including a desired blood vessel in a medical image on a display unit;a blood vessel selection step for selecting the desired blood vessel from the image of the predetermined organ displayed by the organ display step and designating a specific position of the desired blood vessel; andan image display step for displaying a simulation image on which is observable a pseudo flow of a virtual contrast agent through the blood vessel selected in the blood vessel selection step, on the image of the predetermined organ. 2. The medical image display method according to claim 1, further comprising: a tissue border calculation step for calculating a tissue border of a related region relating to the desired blood vessel selected by the blood vessel selection step, whereinthe image display step in the pseudo imaging mode enables to display a pseudo image as if the contrast agent were made to flow from the specified position to the selected blood vessel and the related region of the blood vessel based on the tissue border calculated by the tissue border calculation step. 3. The medical image display method according to claim 1, wherein the image display step in the pseudo imaging mode-comprises sequentially displaying the pseudo flow of the contrast agent from the specified-position to the related region calculated by the tissue border calculation step. 4. The medical image display method according to claim 1, wherein the blood vessel selection step comprises setting a region of interest in extracting the image of the predetermined organ and the tissue border calculation step comprises calculating the tissue border based on the set region of interest. 5. The medical image display method according to claim 1, wherein the blood vessel selection step comprises arbitrarily setting the specified position in the desired blood vessel as a virtual contrast agent injection start position. 6. The medical image display method according to claim 5, further comprising a step for setting a termination position corresponding to the flow start position, wherein the image display step in the pseudo imaging mode comprises sequentially displaying a pseudo image of a flow of the contrast agent from the injection start position to the termination position. 7. The medical image display method according to claim 1, wherein the image display step in the pseudo imaging mode comprises sequentially displaying a pseudo image of a flow of the contrast agent to a substantially whole part of the related region of the selected blood vessel. 8. The medical image display method according to claim 1, wherein the image display step in the pseudo imaging mode comprises sequentially highlighting a locus of the flow of the contrast agent in displaying the pseudo image. 9. The medical image display method according to claim 8, wherein the highlighting comprises marking out the ... | Hitachi Medical Corporation | A61B005/05
US07869860 | B2 | 2011-01-11 | Method for enhancing the accuracy of PET/CT images | 2006-11-10 | A method is disclosed for obtaining linear attenuation coefficients for interpreting a PET scan of a region. The method is suitable for use when high molecular weight materials are present in the region, such as contrast agents or metal objects. The method includes obtaining first and second x-ray CT data sets of the region of interest at two different energies or voltage potentials and differencing corresponding CT numbers. The difference values are used to distinguish portions of the region that are bone from portions of the region that are contrast agent or other high molecular weight material. The obtained CT data set is then used to obtain an attenuation coefficient map of the region at the PET energy of 511 keV, for example, using a linear scaling factor suited to the particular identified material. Difference values at or near zero may be used to identify soft tissue portions of the region. | 1. A method for estimating linear attenuation coefficients at a particular photon energy for a three-dimensional region of a body comprising the steps of: scanning the three-dimensional region with an x-ray beam at a first kVp setting to generate a first CT image data set of values that indicate radiodensity at points in the three -dimensional region;scanning the three-dimensional region with an x-ray beam at a different kVp setting to generate a second CT image data set of values that indicate radiodensity at points in the three-dimensional region;calculating the differences in corresponding values in the first and second CT image data sets wherein the differences are calculated between points in the first and second CT image data sets that correspond to the same physical location in the three-dimensional region;estimating from the calculated differences portions of the three-dimensional region that are relatively low molecular weight material;estimating from the calculated differences portions of the three-dimensional region that are relatively high molecular weight material; andestimating linear attenuation coefficients for the three-dimensional region using one of the first and second CT image data sets and a first conversion factor for the relatively low molecular weight material and a second conversion factor for the relatively high molecular weight material, wherein the first conversion factor is different from the second conversion factor. 2. The method of claim 1, wherein the relatively low molecular weight material is bone. 3. The method of claim 1, wherein the relatively high molecular weight material is a contrast agent. 4. The method of claim 3, wherein the contrast agent is a solution containing one of barium and iodine. 5. The method of claim 1, wherein the first x-ray scan is taken at a tube voltage potential of between about 80 kVp and 100 kVp, inclusive. 6. The method of claim 1, wherein the second x-ray scan is taken at a tube voltage potential of between about 120 kVp and 140 kVp, inclusive. 7. The method of claim 1, wherein the attenuation coefficients are estimated for a photon energy of 511 keV. 8. The method of claim 1, further comprising estimating a difference threshold as a function of radiodensity, wherein the calculated differences above the threshold value are identified as relatively high molecular weight material. 9. The method of claim 8, wherein the difference threshold is a linear function of radiodensity. 10. The method of claim 1, further comprising using one of the first and second CT image data sets to identify portions of the three dimensional region that are a soft tissue. 11. A method for calculating linear attenuation coefficients in a three-dimensional region of a body comprising the steps of: using x-ray computed tomography at a first kVp setting to generate a first CT image data set containing radiodensity values of the three-dimensional region;using x-ray computed tomography at a different ... | University of Washington | A61B005/05 G01N023/04
US07869852 | B2 | 2011-01-11 | Diabetes management system | 2007-10-26 | A method of diabetes management, comprising (a) providing a remote communications unit in communication with a portable microprocessor-based unit; (b) transmitting blood glucose level data into the portable microprocessor-based unit or the remote unit, or both; (c) running a program of instructions on the portable microprocessor-based unit or on the remote unit, or both, using data corresponding to the blood glucose level data; and (d) providing a signal including instructions to inject insulin when the blood glucose level data indicates. | 1. A method of diabetes management, comprising: (a) providing a communications unit remotely located from and in communication with a portable microprocessor-based unit;(b) advising a user to measure a blood glucose level by transmitting blood glucose level data into the portable microprocessor-based unit or the communications unit, or both;(c) running a program of instructions on the portable microprocessor-based unit or on the communications unit, or both, using data corresponding to the blood glucose level data, wherein said portable microprocessor-based unit comprises a multi-line display configured to present to said user (i) graphs of said blood glucose level data and (ii) a game-like presentation; and(d) advising said user to inject insulin according to an insulin plan;(e) presenting to said user (a) said blood glucose level data and (b) information to educate said user regarding management of diabetes;(f) receiving from said user responses confirming said blood glucose level has been measured; and(g) positively reinforcing compliance with a diabetes management plan by providing bonuses in said game-like presentation on said multi-line display in response to said blood glucose level being within an optimal range. 2. The method of claim 1, further comprising signal coupling a blood glucose meter with said portable microprocessor-based unit from which said blood glucose data are downloaded. 3. The method of claim 1, wherein (i) said program of instructions comprises a diabetes management plan and (ii) further comprising communicating a signal to said portable microprocessor-based unit to advance between steps of said diabetes management plan. 4. The method of claim 1, further comprising checking said blood glucose level as a result of running the program of instructions. 5. The method of claim 1, further comprising selecting said insulin plan as a result of running the program of instructions. 6. The method of claim 1, further comprising getting a menu of foods to eat as a result of running the program of instructions. 7. The method of claim 1, wherein said program of instructions are configured such that said advising said user to inject said insulin occurs when said blood glucose levels do not remain in a predetermined range. 8. One or more processor-readable storage devices having program code embodied therein for programming one or more processors to perform a method of diabetes management using a system including a communications unit remotely located from and in communication with a hand-held microprocessor-based unit, wherein the method comprises: (a) advising a user to measure a glucose level;(b) providing blood glucose level data based on the measured glucose level into the hand-held microprocessor-based unit;(c) running a program of instructions on the hand-held microprocessor-based unit or on the communications unit, or both, using data corresponding to the blood glucose level data, wherein said hand-held microprocessor-bas... | Health Hero Network, Inc. | A61B005/00
US07865225 | B2 | 2011-01-04 | Method for implementing a medical procedure on a subject with different medical apparatuses that each process patient data | 2005-12-15 | In a method for implementation of a medical procedure on a patient with a first medical apparatus that processes patient data and a second medical apparatus that processes patient data, both apparatuses have interfaces that can be connected with one another for exchange of the patent data. Patient data of the patient is electronically stored in the first apparatus, the interfaces of both apparatuses are connected, the patient data is transferred from the first apparatus to the second apparatus, and the procedure is implemented with the second apparatus. | 1. A method for implementing a medical procedure on a subject comprising the steps of: designating and reserving a group of medical apparatuses for implementing an entirety of a medical procedure on a subject, said group including a first medical apparatus having a data interface;electronically storing patient data associated with the subject, needed for implementing said medical procedure, in said first medical apparatus and implementing an initial portion of said medical procedure on said subject with said group including said first medical apparatus;only after beginning implementation of said initial portion of said medical procedure, identifying a second medical apparatus that is not in said group and that performs a medical function not performed by any apparatus in said group, as being necessary for completion of said entirety of said medial procedure, said second medical apparatus also having a data interface, and connecting said data interface of said first medical apparatus with said data interface of said second medical apparatus, exclusively through a direct connection between said data interfaces that did not exist prior to beginning said implementation of said initial portion of said medical procedure, said second medical apparatus not having said patient data stored therein at a time said direct connection is made;after making said direct connection, transferring said patient data from said first medical apparatus to said second medical apparatus via the respective data interfaces thereof; andsubsequently proceeding with said medical procedure on said subject using said second medical apparatus. 2. A method as claimed in claim 1, comprising generating additional patient data with said second medical apparatus while proceeding with said medical procedure using said second medical apparatus, and exchanging said patient data and said additional patient data in both directions via said data interfaces between said first medical apparatus and said second medical apparatus. 3. A method as claimed in claim 2, comprising upon completion of said medical procedure at said second medical apparatus, electronically storing all patient data associated with said medical procedure at said first medical apparatus. 4. A method as claimed in claim 3, comprising operating said first medical apparatus as a client of a server of an information system selected from the group consisting of hospital information systems and radiology information systems, and wherein the step of electronically storing said patient data in said first medical apparatus comprises transferring said patient data from said server to said first medical apparatus, and operating said first medical apparatus as a sub-server of said information system and operating said second medical apparatus as a client of said sub-server after connecting the data interfaces of the first and second medical apparatuses. 5. A method as claimed in claim 4, comprising transferring said patient data f... | Siemens Aktiengesellschaft | A61B005/05 B23C001/12
US07865222 | B2 | 2011-01-04 | Method and apparatus for reducing coupling between signals in a measurement system | 2006-01-23 | A method and an apparatus for separating a composite signal into a plurality of signals is described. A signal processor receives a composite signal and separates a composite signal in to separate output signals. Feedback from one or more of the output signals is provided to a configuration module that configures the signal processor to improve a quality of the output signals. In one embodiment, the signal processor separates the composite signal by applying a first demodulation signal to the composite signal to generate a first output signal. In one embodiment, the signal processor also applies a second demodulation signal to the composite signal to generate a second output signal. In one embodiment, a phase and/or amplitude of the first demodulation signal and a phase and/or amplitude of the second demodulation signal are selected to reduce crosstalk. In one embodiment, the composite signal is obtained from a detector in a system for measuring one or more blood constituents. | 1. An apparatus for measuring one or more blood constituents in a subject, said apparatus comprising: a first signal source which applies a first input signal during a first time interval in a first modulation scheme;a second signal source which applies a second input signal in a second time interval in a second modulation scheme;a detector which detects a first parametric signal responsive to said first input signal passing through a portion of said subject having blood therein and which detects a second parametric signal responsive to said second input signal passing through said portion of said subject, said detector generating a detector output signal responsive to said first and second parametric signals; anda signal processor which receives said detector output signal, said signal processor demodulating said detector output signal by applying a first demodulation signal to a signal responsive to said detector output signal to generate a first demodulator output signal and applying a second demodulation signal to said signal responsive to said detector output signal to generate a second demodulator output signal, said first demodulation signal having at least one component comprising a frequency, a first amplitude, and a first phase, and said second demodulation signal having at least one component comprising a second frequency, a second phase, and a second amplitude, at least one of said first phase, said second phase, said first amplitude, and said second amplitude chosen to reduce crosstalk from said first parametric signal to said second demodulator output signal and to reduce crosstalk from said second parametric signal to said first demodulator output signal;wherein said signal processor is further configured to adjust at least one of the first and second modulation schemes and adjust a corresponding at least one of said first and second demodulation signals based on the output signals in order to further reduce crosstalk. 2. The apparatus of claim 1, wherein at least a portion of one of said first phase, said second phase, said first amplitude, and said second amplitude is determined by turning off one of said first and second signal sources and measuring the crosstalk between one of the parametric signals and the non-corresponding output signal. 3. The apparatus of claim 1, wherein said signal processor is further configured to adjust at least one of the first and second modulation schemes and adjust a corresponding at least one of said first and second demodulation signals based on the output signals in order to further reduce crosstalk continuously. 4. The apparatus of claim 1, wherein said signal processor is further configured to adjust at least one of the first and second modulation schemes and adjust a corresponding at least one of said first and second demodulation signals based on the output signals in order to further reduce crosstalk interspersed with measurements. 5. A method of demodulating a composite signal generate... | Masimo Laboratories | A61B005/1455 H04B015/00
US07865234 | B1 | 2011-01-04 | Quantitative method for the therapeutic treatment of sleep disorders | 2004-12-22 | The present invention relates to a method of analyzing a subject for excessive daytime sleepiness, and more particularly to a quick (short duration), quantitative method of sleep disorder analysis. The present invention additionally relates to a method, which can be used to quantitatively measure the treatment endpoints for the subject, i.e., appropriate levels of stimulants. Additionally, the present invention relates to a device for sleep disorder analysis. | 1. A method of therapeutically treating a subject for sleep disorders comprising the steps of: a) determining that a subject has maintained what the subject believes sleeping pattern;b) using a processor or computer to quantitatively analyze a subject's brain wave signals and using the quantitative analysis in estimating or determining whether the subject has a sleeping disorder;c) a physician, technician or veterinarian recommending a change in the physical sleeping conditions of the subject or giving or prescribing the subject medication to make an improvement to the subject's sleeping disorder based in part on the quantitative analysis;d) using a processor or computer to quantitatively analyze a second time the subject's brain wave signals to estimate or determine the extent of the improvement to the subject's sleeping disorder; ande) making an additional, change to the physical sleeping conditions of the subject or reducing or increasing the medication based on the extent of the improvement to the subject's sleeping disorder estimated or determined in previous step, or, alternatively, determining that no change to the physical sleeping conditions of the subject and no reduction or increase of the medication is necessary based on the extent of the improvement to the subject's sleeping disorder estimated of determined, in the previous step. 2. The method in claim 1, further including a step of providing one or more types of stimulus to the subject while the subject's brain wave signals are collected or measured. 3. The method in claim 2, wherein the subject is instructed to respond after recognizing one or more of the types of stimulus. 4. The method in claim 3, further including the step of measuring the subject's response to one or more types of stimulus and wherein the stimulus is provided on an intermittent basis. 5. The method in claim 1, wherein the subject brain wave signals are not reviewed by a human prior to the quantitative analysis. 6. The method in claim 1, wherein the subject is given medication and the amount of medication is decreased based on the extent of the improvement to the subject's sleeping disorder estimated of determined by quantitatively analyzing the subject's brain wave signals a second time. 7. A method of therapeutically treating a subject for sleep disorders comprising the steps of: a) using a processor or computer to quantitatively analyze a subject's brain wave signals and using the quantitative analysis in estimating or determining whether the subject has a sleeping disorder;b) making a change in the physical sleeping conditions of the subject to make an improvement to the subject's sleeping disorder based in part on the quantitative analysis; andc) using a processor or computer to quantitatively analyze a second time the subject's brain wave signals to estimate or determine the extent of the improvement to the subject's sleeping disorder; andd) making an additional change in the physical sleeping condit... | Cleveland Medical Devices Inc. | A61B005/04
US08078270 | B2 | 2011-12-13 | Heart failure monitor quicklook summary for patient management systems | 2010-10-04 | Continuous remote monitoring of patients based on data obtained from an implantable hemodynamic monitor provides an interactive patient management system. Using network systems, patients are remotely monitored to continuously diagnose and treat heart-failure conditions. A screen displayable summary provides continuous feedback and information to physicians, patients and authorized third parties. The quick look summary includes various sites and presentation tailored to match the patients' and physicians' needs. The quick look summary further includes intelligent features that understand and retain the user's interests, preferences and use patterns. Patients, physicians and other caregivers are seamlessly connected to monitor and serve the chronic needs of heart-failure patients in a reliable and economic manner. | 1. A heart failure monitor including a quick look summary implemented in a network having a web browser and portal interfaces to transfer and manage data from the heart failure monitor, the web browser and portal interfaces comprising: a patient portal;a physician portal; anda device manufacturer portal;said patient, physician and device manufacturer portals having shared databases and further including secure databases and encryption systems in data communications thereof. 2. The interface of claim 1 wherein said patient portal includes a quick look summary tailored to the patient. 3. The interface of claim 2 wherein said quick look summary included in said patient portal is customizable for various patient preferences. 4. The interface of claim 1 wherein said physician portal includes a quick look summary tailored to the physician. 5. The interface of claim 1 wherein said quick look summary is customizable for various physician preferences. 6. The interface of claim 1 wherein the data includes at least a one of the following: an internal cardiac pressure value, an oxygen saturation value, a pulmonary artery diastolic pressure, a pulmonary artery systolic pressure, a cardiac-blood temperature value, a right atrial pressure value. | Medtronic, inc. | A61B005/04
US09055888 | B2 | 2015-06-16 | Exhaled breath sampling with delivery of gas | 2010-01-04 | A system for sampling exhaled breath and for supply of a gas, the system comprising: a gas delivery cannula comprising at least one nasal prong for insertion into a nostril, the nasal prong comprising a distal end; an exhaled breath sampling cannula for insertion into the nostril, the exhaled breath sampling cannula comprising a distal end; and a connector for coupling the gas delivery cannula to the exhaled breath sampling cannula, such that the distal end of the exhaled breath sampling cannula is disposed deeper in the nostril than the distal end of the nasal prong, to reduce dilution of sampled exhaled breath by delivered gas. The connected is configured to facilitate adjustability of an insertion depth of the exhaled breath sampling cannula into said nostril. | 1. A method for measuring exhaled breath of a patient being supplied with an oxygen containing gas, the method comprising: determining a peak exhalation flow of the patient by measuring the patient's breath flow;providing the patient with a flow of an oxygen containing gas at a first flow rate;wherein the flow rate of the oxygen containing gas is reduced during the determined peak exhalation flow and maintained at the first flow rate when the patient's exhalation flow is not at its peak;sampling exhaled breaths for CO2 measurements during the determined peak exhalation flow; andestimating an end-tidal CO2 (EtCO2), by providing a higher weight to CO2 measurements of exhaled breaths which were performed during the peak in the patient's exhalation flow, than to CO2 measurements of exhaled breaths which were performed when the patient's exhalation flow was not at its peak. 2. The method according to claim 1, wherein the reducing of the flow of the gas comprises stopping the flow of the gas. 3. The method according to claim 1, wherein said peak exhalation flow is determined using a flow sensor. 4. The method according to claim 1, wherein estimating the EtCO2 is further based on a highest CO2 peak detected in at least one of a plurality of samplings. | ORIDION MEDICAL (1987) LTD. | A61B005/08 A61B005/097 A61B005/083 A61M016/06 A61M016/08 A61B005/087
US07876943 | B2 | 2011-01-25 | System and method for lesion detection using locally adjustable priors | 2008-09-30 | According to an aspect of the invention, a method for training a classifier for classifying candidate regions in computer aided diagnosis of digital medical images includes providing a training set of annotated images, each image including one or more candidate regions that have been identified as suspicious, deriving a set of descriptive feature vectors, where each candidate region is associated with a feature vector. A subset of the features are conditionally dependent, and the remaining features are conditionally independent. The conditionally independent features are used to train a naïve Bayes classifier that classifies the candidate regions as lesion or non-lesion. A joint probability distribution that models the conditionally dependent features, and a prior-odds probability ratio of a candidate region being associated with a lesion are determined from the training images. A new classifier is formed from the naïve Bayes classifier, the joint probability distribution, and the prior-odds probability ratio. | 1. A method for training a classifier for classifying candidate regions in computer aided diagnosis of digital medical images, said method comprising the steps of: providing a training set of images, each said image including one or more candidate regions that have been identified as suspicious by a candidate generation step of a computer aided diagnosis system, and wherein each said image has been manually annotated to identify lesions;deriving a set of descriptive feature vectors from a feature computation step of a computer aided diagnosis system, wherein each candidate region is associated with a feature vector, wherein a subset of said features are conditionally dependent, and the remaining features are conditionally independent;using said conditionally independent features to train a naïve Bayes classifier that classifies said candidate regions as lesion or non-lesion;determining a joint probability distribution from said training images that models the conditionally dependent features;determining from said training images a prior-odds probability ratio of a candidate region being associated with a lesion; andforming a new classifier from a product of said naïve Bayes classifier, said joint probability distribution for the conditionally dependent features, and said prior-odds probability ratio. 2. The method of claim 1, further comprising determining an operator threshold of said classifier from an ROC curve of predictions of said classifier. 3. The method of claim 1, wherein said joint probability distribution for the conditionally dependent features is a ratio of whether or not a candidate region is a lesion given said conditionally dependent features, and wherein determining said joint probability distribution comprises using a kernel-density estimation to estimate a distribution for said candidate region to be a lesion, and to estimate a distribution for said candidate region to be a non-lesion. 4. The method of claim 1, wherein said conditionally dependent features incorporate information regarding a spatial location of said candidate region in an object of interest. 5. The method of claim 4, wherein said object of interest is a colon, and said spatial location is a normalized distance from a rectum measured along a colon centerline. 6. The method of claim 4, wherein said object of interest is a pair of lungs, and said spatial location is specified in a lung coordinate system. 7. The method of claim 6, wherein said lung coordinate system is specified in terms of a distance from a lung center reference or hilum, a distance from a lung wall, a distance from a lung apex or basal point or carina, and where it is specified whether said spatial location is in a left lung or a right lung. 8. The method of claim 6, wherein said lung coordinate system is specified in terms of distances from key landmarks, vessels and airway tree structures in and around the lungs. 9. A method for training a classifier for classifying candidate regions i... | Siemens Medical Solutions USA, Inc. | G06K009/00 A61B005/00
US10987015 | B2 | 2021-04-27 | Dry electrodes for electroencephalography | 2009-08-24 | An electroencephalography (EEG) system includes a dry electrode design having a jagged, angular, comb, etc. shaped support housing. Each dry electrode housing includes multiple electrodes where each electrode has multiple contacts for scalp placement with minimal interference from hair. Signals from individual contacts may be disregarded and each housing may provide one or more aggregated signals for data analysis. Each electrode may be placed in close proximity with neighboring electrodes as no conductive gel is required and may be attached to the scalp using straps, elastic cap, spring-type materials, tape, etc. The dry electrode design effectively measures bio-signals including neurological activity. | 1. An apparatus comprising: a comb having a comb base, a first extension, and a second extension, the first extension and the second extension extending from the comb base, the comb base including a sensor to detect eye movement, the sensor disposed on an end of the comb base;a first electrode carried by the first extension, the first electrode including a first transceiver to receive a first control signal;a first filter communicatively coupled to the first transceiver to automatically dynamically modify a first frequency range of signals to be detected by the first electrode in response to the first control signal;a second electrode carried by the second extension, the second electrode including a second transceiver to receive a second control signal;a second filter communicatively coupled to the second transceiver to automatically dynamically modify a second frequency range of signals to be detected by the second electrode in response to the second control signal;a third electrode carried by the comb base, the third electrode including a third transceiver to receive a third control signal; anda third filter communicatively coupled to the third transceiver to automatically dynamically modify a third frequency range of signals to be detected by the third electrode in response to the third control signal;wherein at least one of the first control signal received by the first transceiver, the second control signal received by the second transceiver, or the third control signal received by the third transceiver is based on at least one of (a) a region of detection of the first electrode, (b) a region of detection of the second electrode, and (c) a region of detection of the third electrode. 2. The apparatus of claim 1, wherein the first extension and the second extension have a non-linear shape. 3. The apparatus of claim 1, wherein the third electrode has a plurality of contacts. 4. The apparatus of claim 3, wherein at least one of the contacts has a rectangular cross-section and a conical tip. 5. The apparatus of claim 1, wherein the first extension has a pointed tip, the first extension including a plurality of electrodes, the first electrode and the plurality of electrodes spaced from the pointed tip. 6. The apparatus of claim 1, wherein the first extension has a first radius of curvature and the second extension has a second radius of curvature substantially the same as the first radius of curvature. 7. The apparatus of claim 1, further including a third extension and a fourth electrode carried by the third extension. 8. The apparatus of claim 7, further including a fifth electrode carried by the first extension, a sixth electrode carried by the second extension, a seventh electrode carried by the third extension, and an eighth electrode carried by the comb base. 9. The apparatus of claim 7, wherein the first extension has a first radius of curvature, the second extension has a second radius of curvature substantially the same as the fir... | NIELSEN CONSUMER LLC | A61B005/0478 A61B005/00 A61B005/0484 A61B005/048
US09237864 | B2 | 2016-01-19 | Analyte sensors and methods of manufacturing same | 2010-07-01 | Analyte sensors and methods of manufacturing same are provided, including analyte sensors comprising multi-axis flexibility. For example, a multi-electrode sensor system 800 comprising two working electrodes and at least one reference/counter electrode is provided. The sensor system 800 comprises first and second elongated bodies E1, E2, each formed of a conductive core or of a core with a conductive layer deposited thereon, insulating layer 810 that separates the conductive layer 820 from the elongated body, a membrane layer deposited on top of the elongated bodies E1, E2, and working electrodes 802′, 802″ formed by removing portions of the conductive layer 820 and the insulating layer 810, thereby exposing electroactive surface of the elongated bodies E1, E2. | 1. A continuous analyte sensor configured for in vivo use, the continuous analyte sensor comprising: an elongated conductive body comprising a first electrode comprising an electroactive surface, wherein the elongated conductive body further comprises a second electrode; anda membrane covering at least a portion of the first electrode;wherein the continuous analyte sensor has a yield strength at least 60 ksi, wherein the continuous analyte sensor has a fatigue life of at least 20 cycles of flexing from about 45° to about −45° at a bend radius of about 0.031-inches. 2. The continuous analyte sensor of claim 1, wherein the elongated conductive body comprises an elongated core, an insulating layer covering at least a portion of the elongated core, a conductive layer comprising the second electrode and covering at least a portion of the insulating layer, and a membrane covering at least a portion the first electrode. 3. The continuous analyte sensor of claim 2, wherein the elongated core comprises an elongated body and a layer of conductive material covering at least a portion of the elongated body. 4. The continuous analyte sensor of claim 3, wherein the elongated body comprises at least one material selected from the group consisting of stainless steel, titanium, tantalum, and a polymer. 5. The continuous analyte sensor of claim 3, wherein the layer of conductive material covering at least a portion of the elongated body comprises a conductive material selected from the group consisting of platinum, platinum-iridium, gold, palladium, iridium, alloys thereof, graphite, carbon, and a conductive polymer. 6. The continuous analyte sensor of claim 2, wherein the conductive layer comprises a silver-containing material. 7. The continuous analyte sensor of claim 6, wherein the silver-containing material has a particle size associated with a maximum particle dimension that is less than about 100 microns. 8. The continuous analyte sensor of claim 7, wherein the silver-containing material has a particle shape that is substantially spherical. 9. The continuous analyte sensor of claim 2, wherein the insulating layer comprises at least one polymer selected from the group consisting of polyurethane and polyimide. 10. The continuous analyte sensor of claim 2, wherein a ratio of a thickness of the conductive layer to a thickness of insulating layer is from about 1:5 to about 1:1. 11. The continuous analyte sensor of claim 2 wherein the conductive material is a silver-containing material. 12. The continuous analyte sensor of claim 1, wherein the membrane comprises a polymer having a Shore hardness of from about 70 A to about 55 C. 13. The continuous analyte sensor of claim 1, having an ultimate tensile strength of from about 80 kPsi to about 500 kPsi. 14. The continuous analyte sensor of claim 1, having an ultimate tensile strength of from about 150 kPsi to about 280 kPsi. 15. The continuous analyte sensor of claim 1, having a fatigue life of at least ... | DexCom, Inc. | A61B005/00 A61B005/1473 A61B005/145 A61B005/1486 B05C003/10
US20110028850 | A1 | 2011-02-03 | Process for quantitative display of blood flow | 2009-07-28 | A method for the quantitative representation of the blood flow in a tissue or vascular region is based on the signal of a contrast agent injected into the blood. In the method, several individual images of the signal emitted by the tissue or vascular region are recorded at successive points in time and are stored. Based on the respective signal, a quantity characteristic for the blood flow and a quantity characteristic for the position of the blood vessels are determined for image areas of individual images. These quantities are represented superimposed for the respective image areas such that both the blood flow quantity and the position of the fine blood vessels become clearly visible in the representation and can be differentiated from the tissue. | 1. A method for the quantitative representation of the blood flow in a tissue or vascular region based on the signal of a contrast agent injected into the blood, said method comprising the steps of: recording and storing, at successive points in time, several individual images of the signal emitted by the tissue or vascular region,based on the respective signal, determining a quantity characteristic for the blood flow and a quantity characteristic for the position of the blood vessels for image areas of several individual images, andrepresenting these quantities superimposed for the respective image areas. 2. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 1, wherein the superimposition is carried out image point by image point. 3. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 1, wherein the superimposition is carried out via a weighted addition of quantities that are characteristic for the blood flow and for the position of the blood vessels. 4. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 1, wherein a multicolor representation is superimposed with a grayscale representation. 5. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 1, wherein the superimposition is carried out using a control quantity that is used to decide which quantity shall be employed for each image point of the superimposed representation. 6. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 5, wherein the control quantity is derived from the quantity that is characteristic for the position of the blood vessels. 7. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 5, wherein the quantity that is characteristic for the blood flow comprises a continuous data set and the quantity that is characteristic for the position of the blood vessels a binary data set. 8. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 7, wherein the quantity that is characteristic for the position of the blood vessel constitutes an edge image. 9. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 1, wherein a three-dimensional representation is selected for the superimposition of the quantities. 10. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth in claim 1, wherein the quantities to be superimposed are represented based on at least two quantities of the HSL color space (hue, saturation, luminance). 11. A method for the quantitative representation of the blood flow in a tissue or vascular region as set forth ... | - | A61B006/00 A61B005/02
US08100852 | B2 | 2012-01-24 | Methods and apparatuses for detecting medical device acceleration, temperature, and humidity conditions | 2010-10-16 | An ambulatory medical device for detecting acceleration, temperature, and/or humidity conditions in or around the medical device is provided. The medical device includes one or more acceleration, thermal, and/or humidity sensors which detect acceleration, temperature, and/or humidity conditions in or around the medical device. In response to detected conditions, the medical device may, among other things, alter the operation of the device, provide an alarm or warning to the user, or transmit data about the detected conditions to another device. | 1. An external infusion device for infusing fluid into a person from a reservoir, comprising: a reservoir;a housing adapted to be carried on an exterior of the person;a drive mechanism contained in the housing and operatively coupled to the reservoir to deliver fluid from the reservoir into the person;a processor contained in the housing;an indicator operatively coupled to the processor and adapted to provide at least one of a visual indication, an audible indication, or a tactile indication to indicate information about the infusion device to the person; andan acceleration sensor coupled to the processor and adapted to provide an acceleration output signal as a function of acceleration forces acting on the housing, wherein the acceleration sensor is an impact switch disposed within the housing, the impact switch comprising: an electrical input and an electrical output,an electrically conductive seismic mass having an anchored end and a free end, the mass being electrically coupled to the electrical output, and coupled to the housing at the anchored end and adapted for resilient deflection at the free end in a plurality of directions, anda plurality of electrically conductive contact surfaces fixedly mounted adjacent to the mass, each of the contact surfaces being electrically coupled to the electrical input,wherein the mass is adapted to deflect in one of the plurality of directions from a first position to a second position, touch one of the plurality of contact surfaces, and return from the second position to the first position such that a circuit between the electrical input and the electrical output is closed to provide the acceleration output signal when a predetermined acceleration force in one of the plurality of directions acts on the housing, andwherein the seismic mass comprises: an electrically conductive arm member having an anchored end and a free end, the arm member being electrically coupled to the electrical output and coupled to the housing at the anchored end and adapted for resilient deflection at the free end in the plurality of directions; andan electrically conductive impact head mounted on the free end of the arm member, the impact head being electrically coupled to the arm member,wherein the arm member is adapted to deflect in one of the plurality of directions from the first position to the second position, the impact head is adapted to touch one of the plurality of contact surfaces, and the arm member is adapted to return from the second position to the first position such that a circuit between the electrical input and the electrical output is closed to provide the acceleration output signal when a predetermined acceleration force in one of the plurality of directions acts on the housing, and wherein the arm member has a curved bend,wherein the processor is adapted to control the infusion device in accordance with the acceleration output signal. 2. The external infusion device of claim 1, wherein the plurality of elec... | Medtronic MiniMed, Inc. | A61M031/00 A61M037/00 A61M001/00
US08333705 | B2 | 2012-12-18 | System and method for navigating an ultrasound catheter to image a beating heart | 2010-09-09 | Catheter navigation is coupled with ultrasound imaging to yield a context map showing the location on a heart of the ultrasonically imaged frame. | 1. A system for navigating an ultrasound catheter to image a beating heart, comprising: a catheter carrying an ultrasonic transducer configured to generate a sound wave and sense an echo wave to obtain ultrasound data for a first portion of the interior surface of the heart and a sensor element configured to generate a signal indicative of a position of the first portion; and,an imaging system for simultaneously displaying an ultrasound image of the first portion and a three-dimensional context map of the heart and simultaneously highlighting a first surface on the context map corresponding to the first portion and a second surface on the context map corresponding to a second portion of the interior surface of the heart for which ultrasound data was previously captured by the ultrasonic transducer. 2. The system of claim 1 wherein the imaging system displays the ultrasound image in real time. 3. The system of claim 1 wherein the first and second portions partially overlap. 4. The system of claim 1 wherein the context map displays a section of the heart that includes the first portion and additional portions of the heart adjacent the first portion. 5. The system of claim 1, further comprising a robotic mechanism operatively coupled to the catheter to position and orient the catheter. 6. The system of claim 1 wherein the position of the first portion is determined responsive to a position of the sensor element and the ultrasound data. 7. A system for navigating an ultrasound catheter to image a beating heart, comprising: a catheter carrying a sensor element and an ultrasonic transducer configured to generate a sound wave and sense an echo wave to obtain ultrasound data for a first portion of the interior surface of the heart;a navigation system operatively coupled to the sensor element for determining a position of the first portion; and,an imaging system for simultaneously displaying an ultrasound image of the first portion and a three-dimensional context map of the heart and simultaneously highlighting a first surface on the context map corresponding to the first portion and a second surface on the context map corresponding to a second portion of the interior surface of the heart for which ultrasound data was previously captured by the ultrasonic transducer. 8. The system of claim 7, wherein said navigation system determines a position and an orientation of the sensor element. 9. The system of claim 7 wherein said sensor element is disposed on a longitudinal axis of the catheter and the catheter includes a plurality of additional sensor elements, each one of the additional sensor elements spaced radially from the axis and at approximately equal circumferential distances from circumferentially adjacent others of the additional sensor elements. 10. The system of claim 7 wherein said additional sensor elements are in a common plane. 11. The system of claim 7 wherein the imaging system displays the ultrasound image in real time. 12. The sy... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/05 A61B008/14
US07875260 | B2 | 2011-01-25 | Imaging agents for functional imaging of lymphatic structures | 2007-08-24 | Novel imaging agents targeted to a lymph vascular cell receptor and a hyaluranon cell receptor are disclosed. The disclosed imaging agents incorporate biological molecules such as hyaluronic acid which bind to the receptors. Lymph vascular cell receptor expression may be related to the beginnings of tumor formation. As such, embodiments of the imaging agents may be used to stain lymph structures for detailed imaging of lymph architecture as well as serving as potential markers for tumor angiogenesis, tumor metastases, etc. | 1. A method of imaging one or more lymph structures comprising lymph vessels and lymph nodes, the method comprising: a) administering to one or more lymph structures under a tissue surface at least one packet containing an imaging agent comprising a fluorescent, indol-containing dye, or a combination of indol-containing dyes, conjugated to one or more binding moieties that bind to one or more receptors selected from the group consisting of a lymph vascular cell receptor and a hyaluronan receptor;b) binding the binding moiety of said imaging agent to the one or more receptors on said one or more lymph structures to stain the one or more lymph structures;c) illuminating the tissue surface with an excitation light to excite the imaging agent;d) continuously, non-invasively detecting fluorescent emissions from the imaging agent to image at least said lymph vessels and lymph nodes;(e) capturing a plurality of images of said one or more lymph structures, based on said detected fluorescent images; and(f) tracking in vivo propulsion of said at least one packet through the one or more lymph structures, based on said captured images. 2. The method of claim 1 wherein (a) comprises injecting the imaging agent intradermally into an interstitial space for transport into the one or more lymph structures. 3. The method of claim 1 wherein (a) comprises using a catheter to inject the imaging agent directly into the one or more lymph structures. 4. The method of claim 1 wherein (a) comprises administering the imaging agent intravenously. 5. The method of claim 1 wherein the imaging agent comprises IR-783 dye. 6. The method of claim 1 wherein the one or more binding moieties comprises hyaluronic acid. 7. The method of claim 1 wherein the imaging agent has an excitation wavelength ranging from about 700 nm to about 900 nm. 8. The method of claim 1 wherein (d) comprises using an intensified charge-coupled camera. 9. The method of claim 1 wherein the one or more lymph structures is at least about 3 cm beneath the tissue surface. 10. The method of claim 1 wherein (c) comprises illuminating the tissue surface with an excitation light source selected from group consisting of laser diodes, semiconductor laser diodes, gas lasers, light emitting diodes, and combinations thereof. 11. The method of claim 6 wherein the hyaluronic acid is coupled to said indol-containing dye by an alkyl diamine linkage. 12. The method of claim 11 wherein said alkyl diamine linkage comprises from 2 to 10 carbon atoms. 13. The method of claim 6 wherein the hyaluronic acid has a molecular weight ranging from about 1,000 Da to about 25,000 Da. 14. The method of claim 13 wherein said hyaluronan receptor is HARE. 15. The method of claim 1 further comprising detecting an increase in fluorescent intensity to sense hyaluronidase activity. 16. The method of claim 1 wherein, in e), capturing said plurality of images comprises a camera integration time in the range of about 10 milliseconds to... | Baylor College of Medicine | A61B005/00 A61B010/00
US07881772 | B2 | 2011-02-01 | Electronic portal imaging for radiotherapy | 2002-03-15 | Some embodiments include transmission of a pulse to imaging elements in order to read an image frame from the imaging elements, prior to exposure of the imaging elements to radiation that is to be converted to and captured as image data, wherein the pulse is transmitted so that the image frame is read and the imaging elements become ready to capture image data associated with the radiation substantially immediately before the imaging elements are exposed to the radiation. Embodiments may also include transmission of a second pulse to the imaging elements in order to read a second image frame from the imaging elements, wherein the second pulse is transmitted at a predetermined time after the exposure of the imaging elements to the radiation is complete, wherein one of the imaging elements produces second radiation after the exposure of the imaging elements to the radiation is complete, and wherein the predetermined time is a time suitable to substantially capture image data associated with the second radiation. | 1. A method comprising: transmitting an electrical pulse while not transmitting radiation to imaging elements in order to initiate reading a first image frame from the imaging elements;reading the first image frame from the imaging elements;exposing the imaging elements to radiation emitted from a linear accelerator substantially immediately after the first image frame is read;detecting a linear accelerator pulse while the imaging elements are exposed to the radiation;in response to the linear accelerator pulse, transmitting a second electrical pulse while not transmitting radiation to the imaging elements in order to initiate reading a second image frame from the imaging elements;detecting a second linear accelerator pulse while the imaging elements are exposed to the radiation;determining that the second linear accelerator pulse was detected while the second image frame is being read from the imaging elements;detecting a third linear accelerator pulse while the imaging elements are exposed to the radiation, wherein no electrical pulse is transmitted to the imaging elements to initiate reading an image frame from the imaging elements between the detection of the second linear accelerator pulse and the detection of the third linear accelerator pulse;determining that the third linear accelerator pulse was not detected while the second image frame is being read from the imaging elements;in response to the determination that the third linear accelerator pulse was not detected while the second image frame is being read, transmitting a third electrical pulse to the imaging elements in order to initiate reading a third image frame from the imaging elements; anddisplaying the second image frame, and the third image frame to an operator. 2. The method according to claim 1, wherein the imaging elements comprise a plurality of charge-storing devices, andwherein the electrical pulse causes substantial discharge of the plurality of charge-storing devices. 3. The method according to claim 1, further comprising: determining that the exposure of the imaging elements to the radiation is complete; andtransmitting a fourth electrical pulse to the imaging elements in order to read a fourth image frame from the imaging elements after an immediately-previous image frame is read from the imaging elements. 4. A non-transitory computer-readable medium storing processor-executable process steps, the process steps comprising: a step to transmit an electrical pulse while not transmitting radiation to imaging elements in order to initiate reading a first image frame from the imaging elements;a step to read the first image frame from the imaging elements;a step to expose the imaging elements to radiation emitted from a linear accelerator substantially immediately after the first image frame is read;a step to detect a linear accelerator pulse while the imaging elements are exposed to the radiation;a step to transmit, in response to the linear accelerator pulse, a second el... | Siemens Medical Solutions USA, Inc. | A61B005/05
US07881771 | B2 | 2011-02-01 | Bone reposition device, method and system | 2005-08-03 | A bone reposition device includes first and second bone supports for supporting first and second portions of a fracture or osteotomize bone about a fracture or osteotomize site, a plurality of sequentially connected connection members and a plurality of joints each with at least a pair of adjacent parts for connecting the connection members therebetween and to the first and second bone supports. The plurality of joints includes at least six joints, each of which possesses one degree of freedom and allows controllable relative rotation or translation of said pair of adjacent parts about one of three axes respectively. | 1. A bone reposition device, comprising: first and second bone supports for supporting first and second portions of a fractured or osteotomized bone about a fracture or osteotomize site;first and second marker holders attachable to the first and second bone supports;the first and second marker holders respectively include first and second fiducial markers for defining relative orientations and positions of the first and second bone portions, respectively, andeach fiducial marker having a fixed position relative to its respective bone portion when the marker holders are attached to the bone supports, respectively,a series of connection members extending between the first and second bone supports; anda plurality of joints each having a single degree of freedom of movement and providing relative movement between adjacent connection members and between connection members adjacent the first and second bone supports,the plurality of joints including at least six joints and allowing controllable relative rotation or translation of the first bone support relative to the second bone support about six degrees of freedom,the first and second fiducial markers for determination of an adjustment factor being representative of a relative displacement required to move the first and second bone portions to a reduced arrangement;the first and second bone supports being disengageable from the bone portions so as to allow the joints to be moved and locked relative to each other and in accordance with the adjustment factor, andthe first and second bone supports being reattachable to the respective bone portions after the joints are moved and locked relative to each other in accordance with the adjustment factor so as to maintain the first and second bone portions at said reduced arrangement. 2. The bone reposition device of claim 1, wherein the six joints are about one of a set of three joints, at its neutral configuration, and the three joints are at least substantially perpendicular to each other. 3. The bone reposition device of claim 1, wherein the first and second bone supports each includes at least two elongate members for attaching to the respective bone portion. 4. The bone reposition device of claim 3, further comprising first and second offset locaters attachable to the first bone support and third and fourth offset locators attachable to the second bone support for reattaching the first and second bone supports to the elongate members after adjustment of the joints. 5. The bone reposition device of claim 4, wherein each offset locater is attachable to the respective elongate member. 6. The bone reposition device of claim 3, wherein each marker holder includes at least three non-collinear metal fiducial markers embedded on its external surface for defining a local coordinate system of its respective bone portion. 7. The bone reposition device of claim 1, further comprising at least six precision scales or position transducers attachable to the six jo... | The Hong Kong Polytechnic University | A61B005/05 A61B017/56 A61B017/00 A61B017/58
US08235895 | B2 | 2012-08-07 | Medical sensor kit for combination with a chair to enable measurement of diagnostic information | 2010-03-15 | A medical examination chair kit includes a group of sensors including electrodes in finger-tip sensors, various other types of sensors incorporated in covers attachable or placeable on a chair and load cells on an adjustable grid supporting the chair, all of which provide data which is partially processed in a transmission pocket mounted on the chair. | 1. A medical examination diagnostic kit for attachment to a chair, comprising: a chair, including: a generally horizontal seat, four legs, two arm rests and a back support;said kit including an independent load sensor device comprising an adjustable sensor support grid having four adjustably spaced load cell sensors for placement respectively of a single load cell under each of said legs, each of said load cell sensors comprising an element for detecting the load independently on each of the load cell sensors to provide a set of four independent data signals for medical diagnosis;means for collecting and comparing the four signals thereby providing at least one of the following: a diagnostic ECG;a diagnostic respiration rate;neurological disorders;neuromuscular pathology;stroke; andbalance. 2. The chair of claim 1 further including at least two arm pads incorporating one or more sensor elements, and one or more of the following supplemental sensor elements located in one of said sensor pads: blood pressure sensor;pulse oximetry sensor;sound sensor;skin resistance sensor;manual grip sensor;fluid flow rate sensor;glucometer; andfurther including a wireless sensor signal collection, storage and transmission device mounted on an attachment strap for attachment to the chair, said wireless sensor signal collection, storage and transmission device for collection of sensor signals from said sensor elements. 3. The kit of claim 1 wherein each load cell is mounted on a plate, said plates connected together to form a quadrilateral configuration, and further including an adjustable linkage between each pair of adjacent plates in said quadrilateral configuration of plates. 4. The kit of claim 3 wherein adjacent plates are connected by a link comprised of a telescoping rod in a tube, said links forming said quadrilateral grid. 5. The kit of claim 3 further including sensor pads for placement on the arm rests and back support, each sensor pad including independent sensor elements and a signal transmission device for transmission of a signal to said means for collecting and comparing. 6. The kit of claim 1 wherein each load cell is mounted on a plate, said plates being connected to form a quadrilateral grid with a load cell mounted on a plate at each corner of said grid, the distance between each plate being adjustable, each load cell being independently connected to said means for collecting and comparing. | Commwell Research and Development Ltd. | A61B005/00 A61B019/00 A47C007/62 A47C031/00
US07867177 | B2 | 2011-01-11 | Method for the identification, diagnosis and treatment of somatic dysfunctions | 2007-02-06 | A method for etiologic diagnosis of somatic dysfunctions of a patient and a therapeutic method for correction of said somatic dysfunctions, comprising the steps of identifying the presence of at least one body asymmetry or anomalous spatial arrangement while the patient is looking ahead in an erect position with relaxed shoulders; identifying at least one dynamic muscle which shows hyper-tonicity with respect to a corresponding contralateral muscle while the patient is in said position; identifying at least one peripheral receptor related to said muscular hyper-tonicity; and performing an exteroceptive, pressoceptive, and/or proprioceptive stimulation of at least one peripheral receptor identified as related to patient's somatic dysfunctions. | 1. A method for the etiologic diagnosis of somatic dysfunctions in a patient, comprising the steps of: (a) identifying the presence of at least one body asymmetry or anomalous spatial arrangement while the patient is looking ahead in an erect position with relaxed shoulders;(b) identifying at least one dynamic muscle which shows hyper-tonicity with respect to a corresponding contralateral muscle while the patient is in said position, which hypertonic muscle is responsible for said at least one asymmetry or anomalous spatial arrangement; and(c) identifying at least one peripheral receptor located in the patient's foot sole related to said muscular hyper-tonicity,wherein said step (c) is performed by applying material pieces at selected spots of the foot sole, one foot at a time and one piece at a time, while observing a possible consequent contraction of a sample muscle. 2. The method of claim 1, wherein said step (c) is performed with the patient in a lying position. 3. The method of claim 1, wherein said sample muscle is the deltoid. 4. The method of claim 1, wherein said spots are selected form a group consisting of: flexor hallucis brevis, adductor hallucis, flexor digiti quinti brevis, abductor hallucis, abductor digiti quinti, cuneous pronator, and cuneous supinator. 5. The method of claim 1, further comprising a step of exteroceptive, pressoceptive, and/or proprioceptive stimulation of at least one peripheral receptor located in the patient's foot sole based on the diagnosed muscular hyper-tonicity. 6. The method of claim 5, wherein said stimulation is carried out by a foot insole having elastic bulges at selected spots under the foot sole. 7. The method of claim 1, wherein the material pieces are made of cork. 8. The method of claim 1, wherein a thickness of the material pieces are not greater than 2 mm, and an area of about 4-5 mm2. 9. The method of claim 1, further comprising the steps of: selecting specific spots of the foot sole; andmildly touching the specifically selected spots of the foot sole,wherein the mild touching permits undistorted reaction of the sample muscle. 10. The method of claim 9, wherein a starting point of sensorial connections is identified as the spots of the foot sole that produces a contraction in the sample muscle. 11. The method of claim 1, wherein the patient is free of organic diseases. 12. The method of claim 1, wherein the method is capable of being conducted on-the-field. 13. The method of claim 1, wherein the at least one body asymmetry is one scapular cingolum being higher than a contralateral scapular cingolum. | - | A61B005/107 A61F005/01
US08180439 | B2 | 2012-05-15 | Ischemia detection using intra-cardiac signals | 2010-07-29 | An implanted cardiac rhythm management device is disclosed that is operative to detect myocardial ischemia. This is done by evaluating electrogram features to detect an electrocardiographic change; specifically, changes in electrogram segment during the early part of an ST segment. The early part of the ST segment is chosen to avoid the T-wave. | 1. A system for detecting ischemia comprising: at least one lead configured for placement within a patient, the at least one lead comprising a sensor operative to sense ventricular activity and to generate a corresponding signal; andan implantable medical device configured for connection to the at least one lead, the implantable medical device comprising circuitry operative to receive signals from the sensor corresponding to ventricular activity, to open a window a predetermined time after sensing the ventricular activity, wherein the window is shorter than a ST segment, the circuitry being further operative to determine an amplitude value within the window, and to process the amplitude value to detect ischemia;wherein the circuitry is operative to open a window by opening the window at a first time if the ventricular activity is intrinsic, and at a second time if the ventricular activity is paced. 2. The system of claim 1 wherein the first time is between 50 and 100 milliseconds after the onset of a Q wave. 3. The system of claim 1 wherein the first time is between 100 and 150 milliseconds after the onset of a Q wave. 4. The system of claim 1 wherein the circuitry is operative to compare the amplitude value to a threshold to detect ischemia. 5. The system of claim 4 wherein the threshold is a historical baseline value. 6. The system of claim 5 wherein the threshold is a long-term average value. 7. The system of claim 1 wherein the circuitry is operative to determine the amplitude value by measuring the amplitude value relative to an isoelectric baseline. 8. The system of claim 7 wherein the isoelectric baseline is measured during an interval between a T-wave and subsequent P-wave. 9. The system of claim 1 wherein the circuitry is operative to determine the amplitude value by taking a plurality of measurements and determining an average value. 10. The system of claim 9 wherein the plurality of measurements are taken over multiple cycles. 11. The system of claim 1 and wherein the circuitry is further operative to compute an ischemia burden value based on the processing and detecting of ischemia. 12. The system of claim 1 wherein the second time is greater than the first time. | Pacesetter, Inc. | A61B005/0452
US07862506 | B2 | 2011-01-04 | Diabetes management system | 2007-10-25 | A diabetes management system, comprising a glucose meter measuring blood glucose level data and a portable microprocessor-based unit in communication with the blood glucose meter such as to be capable of accessing the blood glucose level data. Input data corresponding to the blood glucose level data is used in a program of instructions running on the portable microprocessor-based unit or the glucose meter, or both. The program of instructions includes instructions to communicate a signal to inject insulin based upon the blood glucose level data. | 1. A diabetes management system, comprising: (a) a glucose meter for measuring a blood glucose level data; and(b) a portable microprocessor-based unit in communication with the glucose meter such as to be capable of (i) advising a user to measure said blood glucose level data and (ii) presenting to said user an interactive presentation;(c) wherein said blood glucose level data is used in a program of instructions running on the portable microprocessor-based unit or the glucose meter, or both, wherein said portable microprocessor-based unit comprises a multi-line display screen configured to present said blood glucose level data to said user; and(d) wherein said program of instructions includes instructions to (i) advise said user to inject insulin based upon a diabetes management plan, (ii) present to said user (a) said blood glucose level data and (b) information configured to educate said user regarding management of diabetes, (iii) receive from said user responses confirming said blood glucose level data has been measured, and (iv) positively reinforce compliance with said diabetes management plan by providing incentives in said interactive presentation in response to said blood glucose level data being within an optimal range. 2. The system of claim 1, said program of instructions including (i) said diabetes management plan, and (ii) instructions to send a signal from said glucose meter to said portable microprocessor-based unit to advance between steps of said diabetes management plan. 3. The system of claim 1, wherein said insulin is injected by said user. 4. The system of claim 3, wherein said user uses a syringe to inject said insulin. 5. The system of claim 1, said program of instructions including instructions to display on said multi-line display screen of the portable microprocessor-based unit an indicator to check a blood glucose level. 6. The system of claim 1, said program of instructions including instructions to display on said multi-line display screen of the portable microprocessor-based unit an indicator to select an insulin plan. 7. The system of claim 1, said program of instructions including instructions to display on said multi-line display screen of the portable microprocessor-based unit an indicator to get a menu of foods to eat. 8. The system of claim 1, said portable microprocessor-based unit being connectable with a remote communication unit. 9. The system of claim 1, said portable microprocessor-based unit being connectable with computer in a hospital, wherein said computer is remotely located from said portable microprocessor-based unit. 10. The system of claim 9, said portable microprocessor-based unit being connectable with said computer via a telephone line. 11. The system of claim 1, wherein said advising said user to said inject insulin based on said blood glucose level data is sent when said blood glucose level data does not remain in a predetermined range. 12. A method of diabetes management, compr... | Health Hero Network, Inc. | A61B005/00
US07894890 | B2 | 2011-02-22 | Devices and methods for monitoring physiological information relating to sleep with an implantable device | 2007-02-09 | Described here are implantable devices and methods for monitoring physiological information relating to sleep. The implantable devices are generally designed to include at least one sensor for sensing physiological information, a processor for processing the physiological information using low computational power to detect a sleep stage, and a battery. The detected sleep stage information may then be used to indicate sleep quality, identify or monitor a medical condition, or guide treatment thereof. | 1. An implantable device for monitoring physiological information relating to sleep comprising: at least one sensor configured to sense physiological information; anda processor coupled to the at least one sensor and configured to process the physiological information to detect at least one sleep state or stage, the processor configured to detect a plurality of half waves in the sensed physiological information based on a local maximum amplitude and a subsequent local minimum amplitude, or based on a local minimum amplitude and a subsequent local maximum amplitude,configured to determine a duration and an amplitude of each half wave of the plurality of detected half waves,configured to determine, for each half wave, whether the duration is within a range of a plurality of different duration ranges, and whether the amplitude of said half wave is within a range of a plurality of different amplitude ranges, each different duration range having a preset minimum duration threshold and a preset maximum duration threshold, each different amplitude range having a preset minimum amplitude threshold and a preset maximum amplitude threshold,configured, for each duration range of the plurality of different duration ranges, to count a number of occurrences of half waves within said each duration range for each amplitude range of the plurality of different amplitude ranges, andconfigured to determine a detection of an occurrence of the at least one sleep state or stage if the number of occurrences of half waves within specified amplitude ranges and specified duration ranges correspond to the at least one sleep state or stage. 2. The device of claim 1 further comprising a memory configured to store at least a portion of the physiological information or the processed physiological information. 3. The device of claim 1 further comprising a wireless system configured to transmit at least a portion of the physiological information or the processed physiological information to a remote location. 4. The device of claim 1, wherein the physiological information comprises electrographic information. 5. The device of claim 4, wherein the electrographic information is obtained from an electroencephalogram (EEG) signal. 6. The device of claim 4, wherein the electrographic information is obtained from an electrocorticogram (ECoG) signal. 7. The device of claim 1, wherein the processor is further configured to determine whether detected occurrences of the at least one sleep state or stage indicate a physiological condition and to generate a control signal in response to a determined physiological condition, the device further comprising at least one stimulation electrode coupled to the processor to provide a stimulus in response to the control signal. 8. The device of claim 1, wherein the at least one sensor is an electrode that is also configured to provide stimulation in response to a control signal, wherein the processor is further configured to determine whether d... | Neuropace, Inc. | A61B005/048
US07877277 | B1 | 2011-01-25 | System and method for predictive modeling in disease management | 2010-02-01 | A method and system for administering a disease management program to improve healthcare quality, reduce healthcare costs, and optimize delivery of healthcare services. A multi-condition risk assessment is conducted for all or a substantial portion of a population of program participants, and collected multi-condition risk assessment data are combined with claims data for predictive modeling of future healthcare risk and expense. Participants are risk-stratified into one or more classifications of future healthcare cost risk, and appropriate intervention or delivery of healthcare services is made based on the risk classification. | 1. A method of disease management embodied in a computer program product for execution on a computer system, comprising: a computer program product for execution on an instruction processing system, comprising a tangible storage medium readable by the instruction processing system and storing instructions for execution by the instruction processing system for performing the method;filtering claims data of program participants to define a population;conducting a multi-condition risk assessment of substantially all participants within the defined population;combining data gathered from said multi-condition risk assessment with claims data for participants within the defined population;classifying future healthcare risk of participants based on the combination of claims data and data gathered from said multi-condition risk assessment; andoutputting the classified future healthcare risk participants to the tangible storage medium. 2. The method of claim 1, further comprising directing a plan of care based on the classification of future healthcare risk. 3. The method of claim 1, wherein the step of conducting a multi-condition risk assessment comprises gathering data regarding one or more medical conditions of the program participant. 4. The method of claim 3, wherein the data regarding one or more medical conditions of the program participant are selected from blood pressure, hypertension, cholesterol levels, diabetes, elevated blood-glucose levels, swelling or inflammation, chest pain, fatigue, shortness of breath, depression, cancer, low back pain, cardiac disease, congestive heart failure, coronary artery disease, respiratory disease, asthma, chronic obstructive pulmonary disease, and combinations thereof. 5. The method of claim 1, wherein the step of conducting a multi-condition risk assessment comprises gathering data regarding one or more medical treatments of the program participant. 6. The method of claim 5, wherein the data regarding one or more medical treatments of the program participant are selected from prescription or non-prescription drugs taken, insulin therapy, oxygen therapy, compliance with prescribed treatment regimens, and combinations thereof. 7. The method of claim 1, wherein a multi-condition risk assessment is conducted for every member of the defined population of program participants. 8. The method of claim 1, wherein the multi-condition risk assessment is a preliminary multi-condition risk assessment conducted upon entry of the participant into a disease management program. 9. The method of claim 1, wherein the multi-condition risk assessment is conducted upon occurrence of a specified event. 10. The method of claim 1, wherein the multi-condition risk assessment is conducted via an Internet-based online question-and-answer session. 11. The method of claim 1, wherein the step of filtering claims data of program participants comprises identification of at least one potential indicator of increased risk of future ... | Matria Healthcare, Inc. | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07867176 | B2 | 2011-01-11 | Variable stiffness guidewire | 2005-12-27 | A guidewire navigable through body vessels of a human subject for delivery of a catheter or the like is provided. The guidewire has a tube which receives a corewire that protrudes beyond a distal portion of the tube. The protruding portion of the corewire is surrounded by a spring and shapeable into a curve or arc. The cross-sectional shape of the spring may be varied in order to promote bending flexibility and curvature or to favor curvature of a chosen type. The corewire is axially movable with respect to the tube, which compresses or stretches the spring to change the stiffness of the spring. The tube has a proximal portion comprised of a relatively rigid material, such as stainless steel, while a distal portion is comprised of a more flexible material, such as a nitinol. | 1. A variable stiffness guidewire for navigating through body vessels, comprising: a tube having a proximal portion substantially comprised of a first material and a distal portion substantially comprised of a second material, wherein said second material is different from and more flexible at human body temperature than said first material;an elongated corewire received by the tube and including a distal end, a proximal end, and a protruding portion extending beyond the distal portion of the tube;a unitary coiled spring surrounding at least said protruding portion of said corewire, said spring having a distal end associated with said distal end of said corewire and a proximal end associated with said distal portion of said tube, wherein said corewire is axially movable with respect to said tube to adjust the total length of the spring and thus the stiffness of said spring;the unitary coiled spring having a plurality of cross-sectional shapes, wherein a plurality of coil lengths of the spring have a first portion with a first cross-sectional shape that is substantially flat and a second portion with a second cross-sectional shape that is substantially circular with a given diameter, said substantially flat first cross-sections are diametrically opposed to said substantially circular second cross-sections, the first cross-sections being along a first longitudinal side of the coiled spring that is diametrically opposed to a second longitudinal side of the unitary coiled spring, and said substantially flat first cross-sections each have a length and a width shorter than the length, and said length extends substantially parallel to said corewire at its said protruding portion, said substantially parallel length being greater than the given diameter of said substantially circular second cross-sections, the second longitudinal side being longer than the first longitudinal side when respective flat first cross-section widths and circular second cross-sections of adjacent coil lengths engage each other, whereby the unitary coiled spring exhibits pre-shaped curvature bending; anda non-rotation system comprising a longitudinal pathway along one of said tube or corewire and a follower on the other of said tube or corewire, said follower being slidable with respect to said pathway and at least partially received within the pathway. 2. The guidewire of claim 1, wherein said second material has shape memory properties. 3. The guidewire of claim 1, wherein said second material is a nitinol material. 4. The guidewire of claim 1, wherein said first material is stainless steel and said second material is a nitinol material. 5. The guidewire of claim 4, wherein said proximal end of said corewire includes a portion having an increased diameter for contacting said proximal portion of said tube to prevent downstream movement of said corewire with respect to said tube. 6. The guidewire of claim 4, wherein said distal end of said spring is connected to and movable ... | Cordis Corporation | A61B005/00
US07867285 | B2 | 2011-01-11 | Sensing systems and methods for monitoring gait dynamics | 2006-02-02 | Systems and method for monitoring gait dynamics are disclosed. The performance of an orthotic or prosthetic device or other device associated with a limb may be measured based on the resistance of a bending sensor. Data from the sensors is gathered or processed, particularly for purposes of alignment, safety, failure, usage, selection, and artificial proprioception. Information relating to the device may be outputted visually or auditorily to an individual. | 1. A system for measuring performance of a prosthetic device, comprising: a prosthetic foot device having, an elongated heel member having a first end extending in a posterior direction from a second end, andan elongated foot member attached to a top portion of the second end of the elongated heel member, wherein the elongated foot member is configured to bend during locomotion by a user; anda bending force sensor comprising a thin elongated strip laminated on a top surface of the elongated foot member and extending in an anterior/posterior direction along both sides of a bending axis of the elongated foot member, wherein the bending force sensor is configured to bend with the elongated foot member during locomotion, and wherein the bending force sensor measures a degree in the bending of the elongated foot member during user locomotion by producing resistance output values that are correlated to the bending forces exerted on the elongated strip by the bending of the elongated foot member during locomotion. 2. The system of claim 1, wherein the bending force sensor is an elongate resistive strip. 3. The system of claim 2, wherein the resistive strip is configured to measure toe load and heel load on the prosthetic foot device. 4. The system of claim 2, wherein the resistive strip is configured to measure variations in bending over a designated step cycle. 5. The system of claim 1, wherein the bending force sensor is a variable printed resistor. 6. The system of claim 1, wherein the bending force sensor comprises resistive carbon technology. 7. The system of claim 1, further comprising a memory configured to store data related to the bending of the elongated foot member. 8. The system of claim 1, further comprising a user interface comprising headphones to receive data related to the bending of the elongated foot member. 9. The system of claim 1, wherein the bending force sensor comprises a length-to-width ratio of greater than about 5:1. 10. The system of claim 9, wherein the bending force sensor comprises a length-to-width ratio of greater than about 10:1. 11. The system of claim 1, further comprising a processor configured to receive the resistance output values from the bending force sensor and to identify a state of locomotion of the user based on a change in a current one of the resistance output values from a previous one of the resistance output values. 12. The system of claim 11, wherein the processor is further configured to dynamically adjust the prosthetic foot device during locomotion by the user based on the identified state of locomotion. 13. The system of claim 12, wherein the processor is configured to control movement of the prosthetic foot device. 14. The system of claim 1, wherein the resistance output value corresponds to a bend radius of the elongated foot member. 15. A system for measuring performance of a prosthetic device associated with a lower limb, comprising: a prosthetic foot device having, a heel member... | Össur hf | A61F002/48 A61F002/66 A61F002/68 A61B019/00 A61B005/00 G01L005/16 G01M019/00 G01M001/00
US08663128 | B2 | 2014-03-04 | Diagnostic method and apparatus | 2007-11-08 | The present invention provides a diagnostic method comprising continuously monitoring and transmitting urine output and urine flow rates of a catheterized patient to means which correlate the same with at least one of renal perfusion, renal function, fluid status, polyuria, oleguria, hypoperfusion, hemorrhage shock and GFR. | 1. A diagnostic method comprising: connecting a catheter of a catheterized patient to a low flow metering device, wherein said low flow metering device incorporates a drop generator and a droplet counter, continuously monitoring urine output and urine flow rates of said catheterized patient by measuring real time minute fluctations in renal flow and renal output by measuring drops of urine by said drop generator and droplet counter; transmitting said real time minute fluctations in renal flow and renal output to means which calculate a change of urine output per minute units to correlate with at least one of renal perfusion, renal function, fluid status, polyuria, oleguria, hypoperfusion, hemorrhage shock and GFR; and graphically displaying through a display said minute fluctations in renal flow and renal output, wherein calculating said change of urine output per minute units comprises calculating a slope. 2. A diagnostic method according to claim 1 for detection of a disease affiliated with abnormal body fluid status. 3. A diagnostic method according to claim 2, wherein said disease is affiliated with hypoperfusion. 4. A diagnostic method according to claim 2, wherein said disease is affiliated with hyperperfusion. 5. A diagnostic method according to claim 1, further comprising providing alarm means. 6. A diagnostic method according to claim 1, comprising continuously monitoring and transmitting urine flow rates per minute units of a catheterized patient to means which correlate the slope of urine output per minute units with at least one of renal perfusion, renal function, fluid status, polyuria, oleguria, hypoperfusion, hemorrhage shock and GFR. 7. A diagnostic method according to claim 6, comprising administering a diuretic to a catheterized patient and monitoring and displaying the slope of urine output per minute after administration thereof. 8. A diagnostic method according to claim 6, comprising administering fluid to a catheterized patient and monitoring and displaying the slope of urine output per minute unit after administration thereof and observing a parameter selected from the slope, a peak and a total time for the flow of urine to return to steady state flow in order to determine a state of the kidney. 9. The method of claim 8, wherein said administering said fluid comprises administering a predetermined amount of fluid. 10. A diagnostic method comprising: connecting a catheter of a catheterized patient to a low flow metering device, wherein said low flow metering device incorporates a drop generator and a droplet counter, continuously monitoring urine output and urine flow rates of said catheterized patient by measuring real time minute fluctations in renal flow and renal output by measuring drops of urine by said drop generator and droplet counter; transmitting said real time minute fluctations in renal flow and renal output to means which calculate a change of urine output per minute units to correlate with at least on... | Flowsense Ltd. | A61M001/00 A61B005/00
US08761873 | B2 | 2014-06-24 | Neural prosthesis system and method of control | 2009-01-30 | Multiple designs, systems, methods and processes for control using electrical signals recorded from clinically paralyzed muscles and nerves are presented. The discomplete neural prosthesis system and method for clinically paralyzed humans utilizes a controller. The controller is adapted to receive a volitional electrical signal generated by the human that is manifest below the lesion that causes the clinical paralysis. The controller uses at least the volitional electrical signal to generate a control signal that is output back to a plant to change the state of the plant, which in one aspect is one or more of the user's paralyzed muscles to achieve a functional result or to devices in the environment around the user that are adapted to receive commands from the controller. | 1. A device adapted to use signals from a nervous system of a human, comprising: a controller with an input means for generating an input of a volitional electrical potential generated by the human's body below a lesion in the nervous system and an output means for applying a control signal to a plant that is operably connected to said controller wherein said control signal is a function of said input. 2. The device of claim 1, wherein said input means comprises an electrode. 3. The device of claim 2, wherein said electrode is selected from the group consisting of intramuscular, epimysial, percutaneous, or surface. 4. The device of claim 1 wherein said controller further comprises a computer component to estimate a total energy of said volitional electrical potential. 5. The device of claim 4, wherein said control signal output is modulated proportionally as a function of said total energy. 6. The device of claim 1, wherein said volitional electrical potential is measured from a peripheral nerve. 7. The device of claim 1, whereby said input comprises a first volitional electrical potential and a second volitional electrical potential. 8. The device of claim 1, wherein said plant is a device external to the human. 9. The device of claim 1, wherein said control signal creates a functional movement of said plant and said functional movement of said plant is selected by the human via said volitional electrical potential. 10. The device of claim 1, wherein the volitional electrical potential is generated by a nerve of the human body. 11. The device of claim 1, wherein the volitional electrical potential is generated by a muscle of the human body. 12. A neural prostheses for a human, comprising: an electrode adapted to receive a volitional electrical potential generated by the human's body and manifested below a lesion in the human's nervous system;a controller operably connected to said electrode and adapted to measure said volitional electrical potential, wherein said controller generates a control signal that is a function of said volitional electrical potential; andan output operably connected to said controller to apply said control signal to a plant. 13. The device of claim 12, further comprising a command generation component that generates an estimated command that is a function of said volitional electrical potential. 14. The device of claim 13, further comprising an output control component adapted to accept said estimated command and generate said control signal as a function of said estimated command. 15. The device of claim 12, where said controller accepts a state measurement indicative of the human's state whereby said control signal is a function of said state measurement. 16. The device of claim 12, wherein a level of said control signal is modulated as a function of said volitional electrical potential. | Case Western Reserve University | A61N001/08 A61N001/36
US20110009770 | A1 | 2011-01-13 | Audiometric Testing and Calibration Devices and Methods | 2010-07-12 | An audiometric testing device is provided including a housing having one or more integral calibration couplers adapted to couple with a testing transducer. The testing device also includes a tone generator that generates tones of various frequencies and intensities during a hearing test. A calibration transducer is positioned proximate the coupler and converts an output of the testing transducer into a calibration signal. The calibration signal is measured by a signal measurement module within the housing, which generates a calibration measurement that can then be used to correct for undesired intensity level variations produced by the tone generator. Audiometric testing systems and calibration methods are also provided. In some cases a testing and/or calibration of an audiometric testing device is controlled by an external computing device coupled to the testing device. | 1. An audiometric testing device, comprising: a housing comprising a tone output and a coupler, the tone output adapted to couple with a lead of a testing transducer and the coupler adapted to receive the testing transducer;a tone generator positioned within the housing and coupled to the tone output, the tone generator adapted to generate a plurality of tones in a testing sequence and a plurality of tones in a calibration sequence;a calibration transducer positioned proximate to the coupler and adapted to convert an output of the testing transducer to a calibration signal; anda signal measurement module positioned within the housing, the signal measurement module coupled to the calibration transducer and adapted to measure the calibration signal resulting in a calibration measurement for adjusting operation of the tone generator. 2. The audiometric testing device of claim 1, wherein the coupler is an earphone coupler or a bone vibrator coupler. 3. The audiometric testing device of claim 1, wherein the tone output is adapted to couple to an external computing device for analyzing frequency characteristics of one or more tones. 4. The audiometric testing device of claim 1, further comprising a processor positioned within the housing and coupled to the tone generator and the signal measurement module, the processor adapted to receive calibration measurements from the signal measurement module and provide tone generation instructions to the tone generator. 5. The audiometric testing device of claim 4, wherein the housing further comprises a communication port coupled to the processor, the communication port adapted to couple the processor with an external computing device. 6. The audiometric testing device of claim 1, wherein the housing further comprises a communication port coupled to the tone generator and the signal measurement module, the communication port adapted to couple the tone generator and the signal measurement module with an external computing device. 7. An audiometric testing system, comprising the audiometric testing device of claim 6 and an external computing device coupled to the communication port, the external computing device having a processor and a non-transitory computer-readable storage medium comprising instructions for causing the processor to control the audiometric testing device based in part on inputs received through the external computing device. 8. An audiometric testing system, comprising the audiometric testing device of claim 1 and an earphone adapted to couple to the tone output and the coupler, wherein the earphone has a known input-output transfer function for calibrating the calibration transducer. 9. An audiometric testing system, comprising: one or more processors;a tone generator coupled with the one or more processors, the tone generator adapted to generate a plurality of tones in a testing sequence and a plurality of tones in a calibration sequence based on instructions from the one or more proc... | - | A61B005/00
US20110009762 | A1 | 2011-01-13 | PORTABLE PNEUMOTACHOGRAPH FOR MEASURING COMPONENTS OF AN EXPIRATION VOLUME AND METHOD THEREFOR | 2008-03-07 | A portable pneumotachograph for determining components of the expiration volume, comprising a processor (3), a PEEP valve (4) mounted on the pneumotachograph (1) on the exhaling side, a filter (5) mounted on the pneumotachograph (1) on the inhaling side for removing the portion of the component in the inspiration air to be determined, at least one sensor (7) mounted in or on the pneumotachograph tube (6), wherein in the sensor (7) disposed on the pneumotachograph tube (6), the pneumotachograph tube (6) having an opening (8) on the sensor side, and/or at least one means configured in the pneumotachograph tube (6) for sampling, and further an optical and/or acoustic control device (9) for the expiration flow. This enables a correlation of measurement data of one or more components of the exhaled flow detected with a lung function test, enabling the localization of the seat of the disease. | 1. A portable pneumotachograph for measuring components of an expiration volume consisting of a pneumotachograph (1) having an inspiration side and an expiration side, a processor (3), a PEEP valve (4) that is connected to the expiration side of the pneumotachograph (1), a filter (5) to remove the fraction of component to be determined in the inhaled air that is connected to the inspiration side of the pneumotachograph (1), a pneumotachograph tube (6), at least one sensor (7) mounted in or on the pneumotachograph tube (6), wherein for the case of the sensor (7) is mounted on the pneumotachograph tube (6) the pneumotachograph tube (6) features an opening on the sensor (8) side and/or at least one means to collect a sample that is located in the pneumotachograph tube (6) and furthermore a visual and/or acoustic control (9) of the expiration flow. 2. The portable pneumotachograph according to claim 1, wherein between the filter (5) that is connected to the inspiration side of the pneumotachograph (1) and the pneumotachograph tube (6) a valve (18) is configured, which closes the inlet opening pneumotachograph (1) during expiration. 3. The portable pneumotachograph according to claim 1, wherein the sensor (7) is selected from the group electrochemical sensor, chemiluminescence sensor, NO sensor, O2 sensor, H2O2 sensor, CO2 sensor, CO sensor, sensor for biomarkers or a combination of sensors from the aforementioned sensors. 4. The portable pneumotachograph according to claim 1, wherein, in the opening (8) between the sensor (7) and the pneumotachograph tube (6), a closed loop controllable and/or open loop controllable valve (15) is installed. 5. The portable pneumotachograph according to claim 4, wherein the valve (15) is “open” or “closed” according to the requirements of the inspiration or expiration volume or of a partial expiration volume that is passing through the pneumotachograph tube (6), whereas optionally the valve (15) is controlled by the processor (3). 6. The portable pneumotachograph according to claim 1, wherein the visual control (9) of the expiratory flow is selected from the group y-t graph, bar graph, and LED display with one or several light-emitting diodes. 7. The portable pneumotachograph according to claim 1, wherein through the PEEP valve (4) an expiratory resistance, preferably from 5 to 20 cm H2O, is adjustable and producible. 8. The portable pneumotachograph according to claim 1, that wherein the PEEP valve (4) exhibits on the inlet side of the pneumotachograph a spring-loaded check valve and on the outlet side (ambient) a pressure valve, such that the spring-loaded check valve will open when exceeding a first pressure value on the input side and the pressure valve on the outlet side closes when exceeding a second pressure value on the input side that is greater than the first pressure value. 9. The portable pneumotachograph according to claim 8, wherein the check valve consists of a valve plate (21) that pushes with ... | FILT Lungen-und Thoraxdiagnostik GmbH | A61B005/08
US20110020235 | A1 | 2011-01-27 | MR METHODS OF GRADING A TUMOR USING AN IMAGING MEDIUM THAT COMPRISES HYPERPOLARISED 13C-PYRUVATE | 2008-05-16 | The invention relates to MR methods of grading a tumor using an imaging medium which comprises hyperpolarized 13C-pyruvate. | 1. Method of determining the grade of a tumor, said method comprises a) MR detecting a tumor in a subject using an imaging medium that comprises hyperpolarized 13C-pyruvate wherein signals of 13C-pyruvate and its 13C-containing metabolites alanine, lactate and optionally bicarbonate are detected;b) using said signals and optionally the total carbon signal to generate a metabolic profile of the tumor;c) comparing the metabolic profile of the tumor generated in step b) to a known metabolic profile of a tumor of a certain grade; andd) determining the grade of the tumor based on the similarities and differences between said metabolic profile of the tumor and said known metabolic profile. 2. Method according to claim 1 wherein in step b) the signals of lactate and/or the total carbon signal are used. 3. Method according to claim 1 wherein the known metabolic profile has been obtained by MR detecting a tumor in a subject using an imaging medium that comprises hyperpolarized 13C-pyruvate wherein signals of 13C-pyruvate and its 13C-containing metabolites alanine, lactate and optionally bicarbonate are detected; and using said signals and optionally the total carbon signal to generate a metabolic profile of the tumor. 4. Method of determining the grade of a tumor, said method comprises a) MR detecting a tumor in a subject using an imaging medium that comprises hyperpolarized 13C-pyruvate wherein signals of 13C-pyruvate and its 13C-containing metabolites alanine, lactate and optionally bicarbonate are detected; andb) using said signals and optionally the total carbon signal to determine the grade of the tumor from a standard curve of such signals versus tumor grade. 5. Method according to claim 4 wherein in step b) the signals of lactate and/or the total carbon signal are used. 6. Method according to claim 1 wherein said method is a method of determining the grade of a prostate tumor. 7. Use of hyperpolarized 13C-pyruvate for the manufacture of an MR imaging medium used in a method of determining the grade of a tumor, said method comprising the steps of: a) MR detecting a tumor in a subject using said MR imaging that comprises hyperpolarized 13C-pyruvate wherein signals of 13C-pyruvate and its 13C-containing metabolites alanine, lactate and optionally bicarbonate are detected;b) using said signals and optionally the total carbon signal to generate a metabolic profile of the tumor;c) comparing the metabolic profile of the tumor generated in step b) to a known metabolic profile of a tumor of a certain grade; andd) determining the grade of the tumor based on the similarities and differences between said metabolic profile of the tumor and said known metabolic profile. 8. Use of hyperpolarized 13C-pyruvate for the manufacture of an MR imaging medium used in a method of determining the grade of a tumor, said method comprising the steps of: a) MR detecting a tumor in a subject using said imaging medium that comprises hyperpolarized 13C-pyruvate wherein sig... | - | A61B005/055
US20110009754 | A1 | 2011-01-13 | ARTERIAL BLOOD PRESSURE MONITORING DEVICES, SYSTEMS AND METHODS USING CARDIOGENIC IMPEDANCE SIGNAL | 2009-12-14 | Provided herein are implantable systems, and methods for use therewith, for monitoring a patient's arterial blood pressure. Electrode(s) implanting within and/or on the patient's heart are used to obtain a cardiogenic impedance (CI) signal indicative of cardiac contractile activity. Additionally, a signal (e.g., PPG or IPG signal) indicative of changes in arterial blood volume remote from the patient's heart is obtained using a sensor or electrodes that are implanted remote from the patient's heart. One or more metrics indicative of pulse arrival time (PAT) are determined, where each metric can be determined by determining a time from one of the detected features of the CI signal to one of the detected features of the signal indicative of changes in arterial blood volume. Based on at least one of the metric(s) indicative of PAT, arterial blood pressure is estimated, which can include determining values indicative of systolic blood pressure, diastolic blood pressure, pulse pressure and/or mean arterial blood pressure, and/or changes in such values. | 1. For use with an implantable system, a method for monitoring a patient's arterial blood pressure, the method comprising: (a) using one or more electrodes implanting within and/or on the patient's heart to obtain a cardiogenic impedance (CI) signal indicative of cardiac contractile activity;(b) using an implanted sensor or implanted electrodes, remote from the patient's heart, to obtain a signal indicative of changes in arterial blood volume remote from the patient's heart;(c) detecting one or more predetermined features of the CI signal;(d) detecting one or more predetermined features of the signal indicative of changes in arterial blood volume remote from the patient's heart;(e) determining one or more metrics indicative of pulse arrival time (PAT), each metric indicative of PAT determined by determining a time from one of the detected features of the CI signal to one of the detected features of the signal indicative of changes in arterial blood volume remote from the patient's heart; and(f) estimating the patient's arterial blood pressure based on at least one of the one or more metrics indicative of PAT. 2. The method of claim 1, wherein step (b) comprises one of the following: using an implanted photoplethysmography sensor, remote from the patient's heart, to obtain a photoplethysmography (PPG) signal indicative of changes in arterial blood volume remote from the patient's heart; andusing implanted electrodes, remote from the patient's heart, to obtain an impedance plethysmography signal (IPG) indicative of changes in arterial blood volume remote from the patient's heart. 3. The method of claim 1, wherein step (f) comprises determining a value indicative of systolic blood pressure (SBP) based on at least one of the one or more metrics indicative of PAT. 4. The method of claim 1, wherein step (f) comprises determining a value indicative of diastolic blood pressure (DBP) based on at least one of the one or more metrics indicative of PAT. 5. The method of claim 1, wherein step (f) comprises determining a value indicative of mean arterial blood pressure (MAP) based on at least one of the one or more metrics indicative of PAT. 6. The method of claim 1, wherein step (f) comprises: determining a value indicative of systolic blood pressure (SBP) based on at least one of the one or more metrics indicative of PAT; anddetermining a value indicative of diastolic blood pressure (DBP) based on at least one of the one or more metrics indicative of PAT. 7. The method of claim 6, wherein step (f) further comprises determining a value indicative of mean arterial blood pressure (MAP) based on the value indicative of SBP and the value indicative of DBP. 8. The method of claim 1, wherein the one or more predetermined features of the CI signal detected at step (c) is/are selected from the group consisting of: minimum amplitude of the CI signal;maximum upward slope of the CI signal;maximum amplitude of the CI signal; andmaximum downward slope of the CI sig... | - | A61B005/0215
US20110004074 | A1 | 2011-01-06 | PROTECTIVE GLOVE WITH WIRELESS PHYSIOLOGICAL SENSOR | 2009-07-06 | A multi-finger protective glove carries one or more physiological sensors in respective fingers. An elastomeric sleeve is positioned in a respective sensor carrying finger to maintain the sensor in proper position relative to a user's finger inserted into the glove, adjacent to at least one of the sensors. The sensor can be attached to a distal end of the sleeve, adjacent an end of the finger of the glove. | 1. An apparatus comprising: a multi-finger glove;at least one elastomeric sleeve position in one finger of the glove; andat least one physiological sensor carried in the finger of the glove, adjacent to an end thereof. 2. An apparatus as claim 1 where the sensor is selected from a class which includes at least a pulse oximeter, a thermal sensor, a blood pressure sensor and a pulse rate sensor. 3. An apparatus as in claim 2 which includes a transmitter carried by the glove. 4. An apparatus as in claim 3 where the transmitter is coupled to the sensor by a fiber optic link. 5. An apparatus as in claim 4 where the sensor is attached to an end of the sleeve. 6. An apparatus as in claim 5 which includes a displaced monitoring unit in wireless communication with the transmitter. 7. An apparatus as in claim 6 where the monitoring unit includes a computer readable storage medium where information as to sensor output is stored for processing. 8. An apparatus as in claim 7 where the monitoring unit includes a visual display device and the monitoring unit presents on the display device information as to one or more physiological parameters of an individual wearing the glove. 9. An apparatus as in claim 7 with a plurality of gloves in wireless communication with the monitoring unit, each of the gloves carries at least one physiological sensor in wireless communication with the unit via a respective glove carried transmitter. 10. An apparatus as in claim 4 which includes a plurality of physiological sensors carried in one of, a common finger of the glove, or, different fingers of the glove. 11. An apparatus as in claim 10 where at least some of the sensors are attached to a respective elastomeric sleeve. 12. An apparatus as in claim 11 where each of the sensors is coupled to the transmitter. 13. An apparatus as in claim 12 which includes a battery coupled to at least some of the sensors and the transmitter. 14. An apparatus as in claim 9 with a plurality of gloves in wireless communication with the monitoring unit, each of the gloves carries at least one physiological sensor in wireless communication with the unit via a respective glove carried transmitter. 15. An apparatus as in claim 14 which includes at least one glove with a plurality of physiological sensors carried in that glove in one of, a common finger of the glove, or, different fingers of the glove. 16. An apparatus as in claim 15 which includes a plurality of elastomeric sleeves carried in respective fingers of that glove. 17. A multi-finger protective glove comprising: a glove body from which extends a plurality of protective finger enclosures into which a user can insert a hand and associated fingers;at least one elastomeric sleeve position in one finger enclosure of the glove, the sleeve has a proximal end and a distal end which abuts an end of the finger enclosure;at least one physiological sensor selected from a class which includes at least a pulse oximeter, a thermal sensor, ... | Honeywell International Inc. | A61B005/02 A61B005/145 A61B005/01 A61B005/021 A61B005/024 A41D019/00
US07885698 | B2 | 2011-02-08 | Method and system for providing continuous calibration of implantable analyte sensors | 2006-02-28 | Method and system for providing continuous calibration of analyte sensors includes calibrating a first sensor, receiving data associated with detected analyte levels from the first sensor, and calibrating a second sensor based on a predetermined scaling factor and data associated with detected analyte levels from the first sensor, is disclosed. | 1. A method of providing continuous calibration of analyte sensors, comprising: calibrating a first sensor using one or more processors;receiving data associated with detected analyte levels from the first sensor; andcalibrating a second sensor with reference to one or more detected analyte levels from the first sensor using one or more processors, wherein calibrating the second sensor includes determining a scaling factor based on data from the first sensor and the second sensor, applying the scaling factor to the data from the second sensor, and determining a correlation level of data from the first sensor and from the second sensor; wherein determining the scaling factor includes comparing the data from the first sensor with the data from the second sensor, and determining the scaling factor based on a calculated scaling factor with the lowest level of average error between the data of the first sensor and the data of the second sensor. 2. The method of claim 1 further including the step of calibrating a third sensor based on a second scaling factor and data associated with detected analyte levels from the second sensor. 3. The method of claim 1 wherein the step of calibrating the second sensor starts after a predetermined time period has passed where the first sensor has been in fluid contact with an analyte of a patient. 4. The method of claim 3 wherein the predetermined time period includes at least approximately 90% of the life of the first sensor. 5. The method of claim 3 wherein the predetermined time period includes at least approximately 50% of the life of the first sensor. 6. The method of claim 1 further including the step of determining a sensitivity of the first sensor. 7. The method of claim 1 further including the step of receiving data associated with detected analyte levels from the second sensor. 8. The method of claim 1 wherein the step of calibrating the second sensor includes the steps of: determining an analyte level; andcomparing the determined analyte level with the data associated with the detected analyte level from the first sensor. 9. The method of claim 1 further including the step of comparing the correlation level with a predetermined correlation threshold defining an acceptable stability level of the second sensor. 10. The method of claim 1 wherein the first sensor and the second sensor are analyte sensors. 11. The method of claim 1 further including the step of removing the first sensor while retaining the second sensor in fluid contact with the analyte of a patient. 12. The method of claim 1 wherein the first sensor and the second sensor are subcutaneously positioned under a skin of a patient. 13. The method of claim 12 wherein at least a portion of the first sensor and at least a portion of the second sensor are in fluid contact with the patient's analyte. 14. A method, comprising: positioning a first sensor in fluid contact with an analyte of a patient;calibrating the first sensor using one or mo... | Abbott Diabetes Care Inc. | A61B005/00
US20110002513 | A1 | 2011-01-06 | Medical apparatus and method of using the same for generating a virtual road map guide of a patient's anatomy and annotate the images | 2009-07-01 | An apparatus for displaying and recording medical images and producing by sketching or drawing line segments creating an electronic virtual road map of a patient's anatomy for use by a physician during an intra-body medical procedure. The apparatus includes among its essential elements, an interventional device providing real time images of body structure of the patient; a medical image acquisition system computer adapted for acquisition of a plurality of medical imaging video signals wherein signals are acquired from a transmitting medical device in the medical device's native resolution having a central processing unit, a display, at least one input device, a storage device, an output device wherein the medical image is displayed on the display; and an image generating device, wherein physician uses the input device to sketch or draw line segments creating virtual road map of patient's anatomy over the medical image. The drawing features can also be used as a teaching tool, directing physicians as to what the plan for the procedure is. It can also be used to document the planned procedure.;A method of displaying and recording a medical image during an intra-body medical procedure, providing a medical image to a medical image acquisition system acquiring a plurality of medical imaging video signals wherein signals are acquired from a transmitting medical device in the medical device's native resolution having a central processing unit, a display, at least one input device, a storage device, an output device wherein the medical image is transmitting to a display, displaying the medical image on the display, viewing the medical image, the operator positioning the input device wherein a line segment appearing on the display, the operator, tracing segments using the input device, the segments appearing on the display over and along with the medical image creating curve segments coincident to the medical image. | 1. An apparatus for displaying and recording medical images and producing a virtual road map of a patient's anatomy for use by a physician during an intra-body medical procedure, comprising; an interventional device providing real time images of body structure of the patient;a medical image acquisition system computer adapted for acquisition of a plurality of medical imaging video signals wherein signals are acquired from a transmitting medical device in the medical device's native resolution having a central processing unit, a display, at least one input device, a storage device, an output device wherein the medical image is displayed on the display; andan image generating device for generating for a virtual road map, wherein the virtual road map comprises a two dimensional image illustrations representing anatomy structure as imputed by an operator. 2. An apparatus for displaying and recording medical images and producing a virtual road map of a patient's anatomy for use by a physician during an intra-body medical procedure as in claim 1 wherein the illustration represents operator's input annotations. 3. An apparatus for displaying and recording medical images and producing a virtual road map of a patient's anatomy for use by a physician during an intra-body medical procedure as in claim 1 where in the illustration represents vascular structure. 4. An apparatus for displaying and recording medical images and producing a virtual road map of a patient's anatomy for use by a physician during an intra-body medical procedure as in claim 1 wherein the medical images are pre-recorded images. 5. An apparatus for displaying and recording medical images and producing a virtual road map of a patient's anatomy for use by a physician during an intra-body medical procedure as in claim 1 wherein the virtual road map is sketched line segments. 6. An apparatus for displaying and recording medical images and producing a virtual road map of a patient's anatomy for use by a physician during an intra-body medical procedure as in claim 5 wherein the virtual road map is sketched line segments chosen from a viable pallet of colors and line weights. 7. A method of displaying and recording a medical image during an intra-body medical procedure, comprising; providing a medical image to a medical image acquisition system acquiring a plurality of medical imaging video signals wherein signals are acquired from a transmitting medical device in the medical device's native resolution having a central processing unit, a display, at least one input device, a storage device, an output device wherein the medical image is transmitting to a display;displaying the medical image on the display, viewing the medical image, the operator positioning the input device wherein a line segment appearing on the display, the operator, tracing segments using the input device, the segments appearing on the display over and along with the medical image creating curve segments coincident to t... | Foresight Imaging LLC | G06K009/00 A61B005/00
US20110028825 | A1 | 2011-02-03 | SYSTEMS AND METHODS FOR EFFICIENT IMAGING | 2010-06-03 | A system and methods for more efficient review, processing, analysis and diagnoses of medical imaging data is disclosed. The system and methods include automatically segmenting and labeling imaging data by anatomical feature or structure. Additional tools that can improve the efficiency of health care providers are also disclosed. | 1. A method for processing medical imaging data by a computer system, the method comprising: acquiring a first corpus data set comprising a three-dimensional data set;comparing reference database data to the three-dimensional data set;automatically segmenting the three-dimensional data set, wherein segmenting comprises identifying and labeling anatomical features within the three-dimensional data set;navigating through the three-dimensional data set;annotating the three-dimensional data set;generating an electronic file comprising inserting patient data in an electronic template and comprising automated detection diagnosing;altering the electronic file comprising adding observations to the electronic file after generating the electronic file; andtransmitting the electronic file from a first computer system to a second computer system. 2. The method of claim 1, wherein comparing further comprises scaling the acquired first corpus data with respect to the reference database data. 3. The method of claim 2, wherein scaling comprises deforming processing more than one distinct deformation vector to deform the acquired first corpus data. 4. The method of claim 3, further comprising comparing the first corpus data set to a past corpus data set for the same patient. 5. The method of claim 4, further comprising tagging and/or highlighting data in the first corpus data set, wherein highlighting comprises comparing with the past corpus data set. 6. The method of claim 1, further comprising altering the opacity of a feature in the three-dimensional data set. 7. The method of claim 1, further comprising navigating through the two-dimensional data set. 8. The method of claim 1, further comprising the first computer transmitting the electronic file to a third computer over a network, wherein the electronic file comprises data edited for a user of the third computer. 9. The method of claim 1, wherein annotating comprises linking observational text data to the data set. 10. The method of claim 1, wherein generating the electronic file comprises executing an algorithm to increase diagnostic efficiency. 11. A method for displaying medical corpus data, the method comprising: visually displaying at least one two-dimensional image concurrent with a three-dimensional volumetric image corresponding to the at least one two-dimensional image; andvisually displaying synchronous three-dimensional and two-dimensional navigation through the at least one two-dimensional image and the three-dimensional volumetric image. 12. The method of claim 11, further comprising: acquiring a first set of three-dimensional data points, constructing the at least one two-dimensional image from at least some of the three-dimensional data points; andconstructing the three-dimensional volumetric image from at least some of the three-dimensional data points. 13. The method of claim 11, further comprising measuring a geometric measurement of a feature in the first three-dimensional volu... | DATAPHYSICS RESEARCH, INC. | A61B005/05 G06K009/00
US07892185 | B2 | 2011-02-22 | Method and apparatus for body fluid sampling and analyte sensing | 2008-09-30 | A body fluid sampling system is provided for use on a tissue site. In one embodiment, the system comprises a cartridge; a penetrating member driver; a plurality of penetrating members arranged in a radial configuration on the cartridge wherein sharpened distal tips of the penetrating members point radially outward; wherein an active one of the penetrating members may be operatively coupled to the penetrating member driver, the penetrating member driver moving the active one along a path out of a housing having a penetrating member exit, into the tissue site, stopping in the tissue site, and withdrawing out of the tissue site; and a plurality of analyte detecting members, wherein at least one of the analyte detecting members is positioned to receive fluid from a wound created by the active one of the penetrating members, wherein the detecting members are not pierced by the active one of the penetrating members. | 1. A body fluid sampling system for use on a tissue site, the system comprising: a cartridge;a penetrating member driver;a plurality of penetrating members arranged in a radial configuration in an interior of the cartridge wherein sharpened distal tips of the penetrating members point radially outward;wherein an active one of said penetrating members is configured to be operatively coupled to said penetrating member driver, said penetrating member driver moving said active one along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site;a plurality of analyte detecting members positioned in the interior of the cartridge, wherein at least one of said analyte detecting members is positioned to receive the body fluid from a wound created by said active one of said penetrating members, wherein said analyte detecting members are not pierced by the active one of the penetrating members, and wherein said analyte detecting members are sized to receive no more than about 1.0 μL of the body fluid;a processor coupled to the penetrating member driver for controlling the penetrating member velocity into the tissue site; anda position sensor coupled to the processor and the penetrating member driver, the position sensor positioned to provide an indication of a position of the penetrating member during actuation, the position sensor coupled to a detection algorithm configured to provide feedback velocity information to the processor, wherein the processor modulates the penetrating member driver to control penetrating member velocity based on the feedback velocity information. | Pelikan Technologies, Inc. | A61B005/00 A61B017/14 A61B017/32 B65D081/00
US08024030 | B2 | 2011-09-20 | System and method for analyzing an electrocardiogram signal | 2009-08-12 | A system for analyzing an ECG signal is provided. The system comprises an interface that receives an ECG waveform associated with heart beat cycle of a patient. The system includes signal processor that determines a first isoelectric portion lying between a T-wave of a first heart beat cycle and a P-wave of a successive heart beat cycle, and a second isoelectric portion lying between a P-wave and a QRS complex of the first heart beat cycle. The signal processor determines a stability measure for each of said first and second portions and adaptively selects the first or the second portion as a baseline for the first heart beat cycle based on the stability measures. The signal processor determines a point of reference on an ST segment associated with the first heart beat cycle and evaluates a deviation of the point of reference on the ST segment from the selected baseline. | 1. A system for analyzing an electrocardiogram (ECG) signal, comprising: an interface that receives an ECG waveform associated with a series of heart beat cycles of a patient, said ECG waveform comprising a P-wave, a QRS complex, a T-wave and an ST segment associated with heart electrical activity of the patient for each heart beat cycle; anda signal processor that processes digitized data samples representing the ECG waveform to evaluate an ST segment deviation by: determining a first and a second isoelectric portion of the ECG waveform, the first isoelectric portion lying temporally between a T-wave of a first heart beat cycle and a P-wave of a successive heart beat cycle, the second isoelectric portion lying temporally between a P-wave of the first heart beat cycle and a QRS complex of the first heart beat cycle,determining a stability measure for each of said first and second isoelectric portions and adaptively selecting the first isoelectric portion or the second isoelectric portion as a baseline for the first heart beat cycle based on a comparison of the stability measures of said first and second isoelectric portions, anddetermining a point of reference on an ST segment on the ECG waveform associated with the first heart beat cycle and evaluating a deviation of the determined point of reference on the ST segment from the selected baseline. 2. The system according to claim 1, wherein the signal processor determines the stability measures of said first and second isoelectric portions respectively based on a number of representative data samples for the first isoelectric portion and a number of representative data samples for the second isolecetric portion. 3. The system according to claim 1, wherein the signal processor adaptively selects successive baselines for said series of heart beat cycles. 4. The system according to claim 1, wherein the signal processor further determines an R-peak, a T-peak and a P-peak for each heart beat cycle, wherein the signal processor: determines the T-peak of a heart beat cycle as a peak point in the ECG waveform between a temporal distance of 60 ms to 200 ms after the R-peak of that heart beat cycle, if a heart beat rate is greater than 120 beats per minute, and as a peak point in the ECG waveform between a temporal distance of 80 ms to 300 ms after the R-peak of that heart beat cycle, if the heart beat rate is lesser than or equal to 120 beats per minute, anddetermines the P-peak a heart beat cycle as a peak point in the ECG waveform between a temporal distance of 60 ms to 160 ms before the R-peak of that heart beat cycle, if the heart beat rate is greater than 120 beats per minute, and as a peak point in the ECG waveform between a temporal distance of 80 m to 200 ms before said R-peak of that heart beat cycle, if the heart beat rate lesser than or equal to 120 beats per minute. 5. The system according to claim 1, wherein the signal processor determines the first isoelectric portion by: determining expo... | Siemens Aktiengesellschaft | A61B005/04
US07865226 | B2 | 2011-01-04 | Specimen positioning system for imaging machines | 2006-02-03 | An animal holder is provided with a specialized coupling that is releasably mountable to a number of different imaging machines such as X-ray, CAT, MRI and PET machines. Composite images created from combining images from such different machines are particularly clear due to the predetermined alignment of the animal holder within the center of the field of view of each machine. | 1. A positioning system for an imaging machine, comprising: an animal holding assembly comprising a specimen alignment bed having a support surface configured to contact and hold an animal in a predetermined position on said support surface;a positioning receiver assembly connecting said animal holding assembly to an imaging machine;a releasable coupling having a first coupling portion coupled to said imaging machine and a second coupling portion coupled to said specimen alignment bed on said animal holding assembly, said releasable coupling releasably interconnecting said animal holding assembly and said positioning receiver assembly in a repeatable predetermined coupled position to facilitate registration and alignment of said specimen alignment bed and said animal in the imaging machine; andwherein said coupling comprises an interconnection panel having internal ports for communicating with wires and tubes. 2. The system of claim 1, further comprising a keying alignment feature provided within said coupling and comprising a male keying member guided into a female keying member and aligning said first and second coupling portions. 3. The system of claim 1, wherein said animal holding assembly comprises a hermetically sealed closed animal holding chamber. 4. The system of claim 1, further comprising a head holder selected from the group consisting of a bite bar, ear bars, and a head clamp. 5. The system of claim 1, further comprising an elongated closed tube internally receiving said specimen alignment bed and coupled to said animal holding assembly. 6. A method of imaging a specimen in a first imaging machine having an axial bore and a first receiver with a first receiver coupling portion for enabling production of accurate, repeatable images of the specimen taken over a period of time, said method comprising: mounting a specimen in a holder, said holder having a first holder coupling portion;providing an axially releasable coupling between said holder and said first receiver, said axially releasable coupling comprising said first holder coupling portion and said first receiver coupling portion aligned along an axis;mounting said holder on said first receiver in a predetermined axial position by axially engaging said first holder coupling portion with said first receiver coupling portion so that the specimen is positioned in a predetermined axial position on said first receiver;channeling a fluid through a closed flowpath through said first holder coupling portion and through said first receiver coupling portion and adjacent to the specimen;positioning said holder within a field of view of said first imaging machine, using said first receiver such that the axis of the releasable coupling extends axially within the axial bore;taking an image of said specimen with said first imaging machine;removing said holder from said field of view; andremoving said holder from said first receiver by axially releasing said coupling. 7. A positioning syst... | - | A61B005/05 A01K015/04
US07892186 | B2 | 2011-02-22 | Handle and articulator system and method | 2005-12-09 | The invention includes an assembly having a handle for articulating an articulator in an ergonomically improved manner. The handle can be adapted to articulate a distal portion of the articulator when squeezed. The handle can have at least one accessible attachment mechanism useful for reattachably attaching the articulator and the handle during a medical procedure. The invention also includes methods of making and using such a handle. | 1. An assembly including a handle for articulating an articulator, the handle comprising at least one accessible attachment mechanism useful for reattachably attaching the articulator and the handle during a medical procedure; and the handle further comprising: a first portion axially aligned with a second portion;a first actuating member pivotably coupled to the first portion and a second actuating member pivotably coupled to the second portion;wherein the first and second actuating members are pivotably coupled to each other and configured such that squeezing the first and second actuating members toward the first and second portions of the handle causes the articulating of the articulator; andwherein the articulator comprises a first wire having a distal portion and a proximal portion and a second wire having a distal portion and a proximal portion, the second wire being disposed in a lumen defined by the first wire and the distal portion of the second wire being functionally coupled to the distal portion of the first wire. 2. An assembly according to claim 1, wherein the handle provides a one-to-one torqueability ratio. 3. An assembly according to claim 1, wherein the handle is adapted to navigate a distal end of the articulator through a tortuous pathway. 4. An assembly according to claim 1, wherein the handle comprises a biasing member to bias a distal end of the articulator to a desired position. 5. An assembly according to claim 4, wherein the biasing member comprises a spring. 6. An assembly according to claim 1, wherein the handle comprises a biasing member to bias a distal end of the articulator to a straight position. 7. An assembly according to claim 1, wherein the handle comprises a biasing member to bias a distal end of the articulator to a curved position. 8. An assembly according to claim 1, wherein the handle is adapted to lock a distal end of the articulator in a desired position. 9. An assembly according to claim 1, wherein the articulator includes a first wire and a second wire, the articulator having a distal end being articulatable by moving the first wire relative to the second wire. 10. An assembly according to claim 9, the handle adapted to move the first wire relative to the second wire by applying a longitudinal force differential between the first wire and the second wire. 11. An assembly according to claim 10, wherein the first portion is functionally coupled to a proximal portion of the first wire and the second portion is functionally coupled to a proximal portion of the second wire, the first and second portions being movable relative to each other. 12. An assembly according to claim 10, wherein the first portion is reattachably attached to the proximal portion of the first wire and the second portion is reattachably attached to the proximal portion of the second wire. 13. An assembly including a handle for articulating an articulator, the handle comprising at least one accessible attachment mechanism ... | Heraeus Materials S.A. | A61M025/00 A61B005/00
US08301225 | B2 | 2012-10-30 | Magnetic resonance imaging apparatus, image processing apparatus, and image processing method | 2010-07-09 | A functional image generating unit that generates a functional image from a magnetic resonance image generated by collecting a magnetic resonance signal from a subject who executes a task with an intermission of a resting period. An image generation controlling unit controls the functional image generating unit so that a plurality of functional images along a time series are generated by separating a plurality of magnetic resonance images generated along the time series during a collecting period (in which the resting period and a task-execution period are repeated a plurality of times along the time series) into image groups, each corresponding to a predetermined temporal phase. A statistical process is performed on each of the image groups and a display controlling unit exercises control so that the functional images are displayed by a display unit. | 1. A magnetic resonance imaging apparatus comprising: an image generating unit configured to generate magnetic resonance images by acquiring magnetic resonance signals from an examined subject who repeatedly executes a task with an intermission resting period;a functional image generating unit configured to generate a functional-diagnosis-purpose magnetic resonance image from said magnetic resonance images;an image generation controlling unit configured to control the functional image generating unit to generate a plurality of functional images along a time series by separating a plurality of magnetic resonance images that have been generated along the time series during a collecting period into image groups, each of which corresponds to a predetermined temporal phase and by performing a statistical process on each of the separated image groups, the collecting period being configured so that the resting period and a task-execution period during which the task is executed are repeated therein a plurality of times along the time series; anda display controlling unit configured to cause the plurality of functional images along the time series that have been generated under control of the image generation controlling unit to be displayed by a display unit. 2. The magnetic resonance imaging apparatus according claim 1, wherein: the image generation controlling unit controls the functional image generating unit to generate the plurality of functional images along the time series by separating the plurality of magnetic resonance images that have been generated along the time series during the collecting period into image groups into which the magnetic resonance images are organized in such a manner that the magnetic resonance images corresponding to the resting periods and the task-execution periods are sequentially contained along the time series. 3. The magnetic resonance imaging apparatus according to claim 2, wherein: the image generation controlling unit controls the functional image generating unit to generate the plurality of functional images along the time series by separating the plurality of magnetic resonance images that have been generated along the time series during the collecting period into image groups that are obtained by using the resting period and the task-execution period that are positioned adjacent to each other along the time series as a basic unit. 4. The magnetic resonance imaging apparatus according to claim 2, wherein: the image generation controlling unit controls the functional image generating unit to generate the plurality of functional images along the time series by separating the plurality of magnetic resonance images that have been generated along the time series during the collecting period into image groups that are obtained by using a temporal phase unit employed for setting the resting periods and the task-execution periods as a basic unit. 5. The magnetic resonance imaging apparatus according to claim 2, wh... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | A61B005/055 G01V003/00
US07874995 | B2 | 2011-01-25 | Reducing or avoiding muscle cramps | 2006-11-01 | Muscle cramps can be reduced or avoided by monitoring the concentration of salt in perspiration to indicate the degree of salt depletion. Based on the degree of salt depletion, a sufficient amount of rehydration beverage is consumed to normalize the salt level. | 1. A product for reducing muscle cramps comprising: a packaging comprising an orally-consumable composition having at least one salt and a table providing an estimate of a quantity of sweat secreted during physical activity based on a level of sweat and a weight of the user;a tester associated with the packaging, said tester including a test portion configured to be exposed to body fluid comprising sweat for use in obtaining a result indicative of an estimated degree of salt depletion from a user;a test chart associated with the packaging configured to be compared with the tester after exposed to the body fluid to provide a test result of an estimated degree of salt loss; anddosage information regarding the orally-consumable composition configured to allow the user to determine a quantity of the orally-consumable composition to consume, based on the test result for salt replenishment and the estimate of a quantity of sweat secreted during physical activity, sufficient to substantially limit muscle cramping. 2. The product of claim 1, wherein the test chart comprises at least a portion of the dosage information. 3. The product of claim 1, wherein the at least one salt comprises an electrolyte selected from the group consisting of: sodium, potassium, magnesium, calcium, chloride, and combinations thereof. 4. The method of claim 1, wherein the consumable composition is selected from the group consisting of: a powder, a gel, a solid, a liquid, and combinations thereof. 5. The product of claim 1, further comprising a body fluid collection unit configured to be dermally mounted onto the skin of the test subject during physical exertion, wherein the tester is further configured to be compatible with the sample collection unit. 6. The product of claim 5, wherein the body fluid collection unit comprises: a first side substantially secured to a second side at a perimeter, such as to create an internal space between first and second sides; andan opening on the first side from an outside of the collection unit to the internal space, wherein the opening is configured to allow perspiration to flow within the collection unit through the opening during physical exertion. 7. The product of claim 6, wherein the body fluid collection unit further comprises an absorbent layer configured to be in substantial contact with the test subject during physical exertion, wherein the absorbent layer extends through the opening of the collection unit into the internal space, wherein during use, perspiration travels from the test subject through the absorbent layer and into the collection unit. 8. A product for reducing muscle cramps comprising: a packaging comprising an orally-consumable composition having at least one salt and a table providing an estimate of a quantity of sweat secreted during physical activity based on a level of sweat and a weight of the user;a tester associated with the packaging, said tester including a test portion configured to be exposed to bod... | Stokely-Van Camp, Inc. | A61B005/05 A61B005/00 A61K047/00
US08117777 | B2 | 2012-02-21 | Multi-layer wristband with removable labels incorporated into the wristband | 2010-04-23 | A bracelet that has two layers, one forming the outside of the bracelet and the other forming the inside of the bracelet. The layers have different properties. The bracelet has a head end and a tail. The layers are bonded to each other and there is at least a first portion of the band at the head end in which the first layer is bonded to the second layer by means of a contact adhesive and in which the second layer is a release liner with regard to the contact adhesive. The bracelet is formable into a loop by removing the second layer from the first layer in the first portion and attaching the exposed contact adhesive to the first layer in the tail. Also disclosed is a business form which contains a band that can be formed into the bracelet. | 1. A business form for a band that is formable into a loop and has a head end and a tail, the business form comprising: a first layer that when the business form's band is formed into the loop is the outside of the loop; anda second layer that when the business form's band is formed into the loop is the inside of the loop,the second layer having a first portion and a second portion which is separate from the first portion, the first portion being bonded to the head end of the band by a contact adhesive and being a release layer with respect to the contact adhesive and the second portion being bonded unreleasably to at least part of the first layer in the tail of the band andthe band being die cut in the business form such that only the first layer is cut for the head end and both layers are cut for the tail. 2. The business form set forth in claim 1 wherein: the first layer is made of a first material; andthe second layer is made of a second material. 3. The business form set forth in claim 2 wherein: the first material has a first property. 4. The business form set forth in claim 3 wherein: the second layer has a second property that is different from the first property. 5. The business form set forth in claim 1 wherein: the strengths of the layers when the band is made into the loop are such that the loop has a required strength. 6. The business form set forth in claim 3 wherein: the first property is that the first material is printable by a printing device. 7. The business form set forth in claim 4 wherein: the second property is that the second layer is chosen to reduce skin irritation. 8. The business form set forth in claim 1 wherein: the strengths of the layers when the band is made into the loop are such that the loop has a required strength. 9. The business form set forth in claim 8 wherein: in the head end, the first layer has by itself the required strength for the loop; andin the tail, the second layer has by itself the required strength for the loop. 10. The business form set forth in claim 8 wherein: the first layer has by itself the required strength for the loop. 11. The business form set forth in claim 1 wherein: in the tail, the first layer includes a second portion that is separable from the remainder of the first layer and the second layer includes a third portion under the second portion that is a release liner with regard to the second portion, whereby the second portion is separable from the remainder of the first layer. 12. The business form set forth in claim 11 wherein: the band has a first printed area which is not in the second portion and a second printed area which is in the second portion. 13. The business form set forth in claim 12 wherein: the first printed area contains identification information for an entity to which the loop made from the band is attached; andthe second printed area contains further information that is associated with the identification information. 14. The business form set fort... | Endur ID, Inc. | A44C005/00
US07885826 | B2 | 2011-02-08 | Medical service support system, medical service support server, and medical service support terminal for supporting medical service | 2008-08-25 | In the embodiment of the present invention, when identification information acquired from a drug or instrument to be used on a patient is not included in schedule information stored beforehand, a user is notified to confirm that the identification information is not included in the schedule information. After the notification is provided, in registering the acquired identification information as practice details, the name of a doctor who instructed to use the drug or instrument, which is different from the drug or instrument specified by schedule information, and the reason for the instruction are registered, while associated with the practice details. Accordingly, drugs or instruments necessary for patients can be used in an appropriate manner. | 1. A medical service support system comprising: a database operatively connected to a computer to store information of a drug or instrument scheduled to be used for a medical practice to be performed on a patient;an acquisition unit operatively connected to a computer to read identification information attached to a drug or instrument that will actually be used on the patient and operative to acquire the identification information which specifies the drug or instrument to be actually used;a determination unit operatively connected to a computer to determine whether the drug or instrument to be actually used that is specified by the identification information acquired by the acquisition unit matches the drug or instrument scheduled to be used for the medical practice stored in the database;a notification unit operatively connected to a computer to provide notification in case that the drug or instrument to be actually used is determined by the determination unit not to match with the drug or instrument scheduled to be used for the medical practice not included in the schedule information stored in the database; anda register operatively connected to a computer, when registering as practice details the identification information acquired by the acquisition unit after the notification unit made the notification, to register the name of a doctor who instructed the use of the drug or instrument to be actually used, which is different from the drug or instrument scheduled to be used for the medical practice, and the reason for the instruction whileassociating the name and the reason with the practice details. 2. The medical service support system according to claim 1, further comprising: an approval reception unit operatively connected to a computer, after the register makes the registration, to receive approval from a doctor who instructed the use of the drug or instrument to be actually used, which is different from the drug or instrument scheduled to be used for the medical practice, whereinthe register, in case the approval reception unit receives the approval, registers information indicating that the approval has been received, associating the information and the practice details with each other. 3. The medical service support system according to claim 1, wherein the database further stores substitute information that associates a drug or instrument and at least one alternative drug or instrument that can be used in place of the drug or instrument with each other,the determination unit, in case of having determined that the drug or instrument to be actually used does not match the drug or instrument scheduled to be used for the medical practice stored in the database, further determines whether the drug or instrument to be actually used is included in the alternative drugs or instruments that can be used in place of the drug or instrument scheduled to be used for the medical practice, andthe register, if the determination unit determines that t... | Olympus Medical Systems Corp. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US08636675 | B2 | 2014-01-28 | Method and device for the extraction of a body fluid | 2010-10-08 | The present invention provides a method and a device for piercing a body part and determining whether a sufficient volume of blood has been withdrawn. A lancing element is rapidly inserted into a body part in a forward phase and retracted quickly to a lesser puncturing depth. Subsequently, the lancing element is retracted slower than during the first retraction movement and the distance retracted during the second retraction is shorter than the first retraction movement. During the second retraction movement, body fluid is collected in a collection phase by a capillary structure of the lancing element. Contact between the lancing element and the body fluid is detected after the forward phase at the beginning and the end of a waiting period. | 1. A method of collecting body fluid with a lancing element having a tip and a capillary structure formed in the lancing element, comprising: inserting the tip of the lancing element into the skin of a body part to a puncture depth;retracting the tip from the puncture depth to a lesser depth than the puncture depth;retracting the tip from the lesser depth toward the skin surface, wherein the retracting of the tip from the puncture depth to the lesser depth is faster than the retracting of the tip from the lesser depth toward the skin surface;detecting contact between the lancing element and the body fluid after the tip has been retracted to the lesser depth;collecting body fluid with the lancing element; anddetermining from the detecting step whether a sufficient volume of body fluid has been collected by the lancing element. 2. The method of claim 1, further comprising: detecting contact between the lancing element and the body fluid when the tip is located at about the puncture depth; andthe determining step comprising comparing a first signal generated from the detecting of contact when the tip is located at about the puncture depth to a second signal generated from the detecting of contact after the tip has been retracted to the lesser depth. 3. The method of claim 2, wherein a sufficient volume of body fluid is determined to have been collected by the lancing element when the first signal and the second signal are substantially the same. 4. The method of claim 3, further comprising triggering an error state when the first signal and the second signal are substantially different. 5. The method of claim 1, further comprising: forming a lancing channel in the skin during the retraction of the tip from the puncture depth to the lesser depth;at least partially evacuating the lancing channel during the retraction of the tip from the puncture depth to the lesser depth;waiting an amount of time after the retracting of the tip to the lesser depth to allow the body fluid to enter the lancing channel; andthen performing the step of detecting contact between the lancing element and the body fluid. 6. The method of claim 1, further comprising applying a contact element against the skin, wherein the lancing element comprises electrically conductive material and the contact between the lancing element and the body fluid is detected by monitoring an electrical parameter between the contact element and the lancing element. 7. The method of claim 6, wherein the detecting step comprises measuring an electrical parameter. 8. The method of claim 7, wherein the electrical parameter comprises impedance. 9. The method of claim 7, further comprising detecting contact between the lancing element and the body fluid during the step of retracting the tip from the puncture depth to the lesser depth, wherein the electrical parameter reaches a reference value when the tip is located at about the puncture depth, then deviates from the reference value after the tip i... | Roche Diagnostics Operations, Inc. | A61B005/00 A61B017/14 A61B017/32 B65D081/00
US08554295 | B2 | 2013-10-08 | Set for determining blood type and cover body | 2009-02-26 | Disclosed is a set for determining whether blood obtained in puncture tools by puncturing a blood vessel is venous blood or arterial blood easily, and ensuring handling under a sterile state. A set for determining blood type has a cover body to be attached to a syringe set, and a blood type determination unit to be housed in the cover body. The blood type determination unit comprises a light projecting portion and a light receiving element, and a blue output LED and a red output LED showing a judgment result based on a signal obtained from the light receiving element. The cover body has a body and a lid that are combined with each other to cover the blood type determination unit, a partition for passing blood between the light projecting portion and the light receiving element, and a pair of measurement windows for transmitting light between the light projecting portion and the light receiving element through the partition. | 1. A set for determining blood type, comprising: a puncture tool;a cover body defining an exterior and an interior; anda blood type determination unit to be removably housed within an interior of the cover body, wherein the blood type determination unit includes a light projecting section and a light receiving section which are located so as to face each other;a judging section for determining whether blood between the light projecting section and the light receiving section is venous blood or arterial blood, based on a signal obtained from the light receiving section; andan output section for outputting the result of judgment made by the judging section;wherein the exterior of the cover body comprises a joint portion configured for attachment of the puncture tool thereto such that blood collected by the puncture tool flows into a space between the light projecting section and the light receiving section. 2. The set for determining blood type according to claim 1, wherein the cover body has a first cover member and a second cover member connected to the first cover member by a hinge, the first cover member including a measurement section, and the blood type determination unit being mounted in the first cover member. 3. The set for determining blood type according to claim 2, wherein the measurement section has an attachment portion to be attached to the puncture tool, anda measurement window provided on the attachment portion and through which light is transmitted between the light projecting section and the light receiving section. 4. The set for determining blood type according to claim 3, wherein at least a periphery of the measurement window of the measurement section is lightproof. 5. The set for determining blood type according to claim 3, wherein the puncture tool has a needle body and a syringe;the attachment portion has a blood inflow portion which is disposed between the needle body and the syringe and into which blood is introduced; andthe needle body and the syringe are connected to the blood inflow portion. 6. The set for determining blood type according to claim 3, wherein the puncture tool has a needle body and a syringe; andthe attachment portion is attached to an outer cylinder of the syringe. 7. The set for determining blood type according to claim 3, wherein the puncture tool has a needle body and a syringe; andthe attachment portion is attached to a blood inflow portion which is formed at a hub of the needle body and into which blood is introduced. 8. The set for determining blood type according to claim 1, wherein the output section has a first lamp for emitting light when the result of judgment by the judging section is venous blood, anda second lamp for emitting light when the result of judgment by the judging section is arterial blood. 9. The set for determining blood type according to claim 1, wherein the blood type determination unit has a power switch; andthe cover body has a switch operating section which, when... | Terumo Kabushiki Kaisha | A61B005/1455 A61B005/00
US20110040171 | A1 | 2011-02-17 | IMAGE GUIDED HIGH INTENSITY FOCUSED ULTRASOUND TREATMENT OF NERVES | 2010-09-21 | A method for using high intensity focused ultrasound (HIFU) to treat neurological structures to achieve a desired therapeutic effect. Depending on the dosage of HIFU applied, it can have a reversible or irreversible effect on neural structures. For example, a relatively high dose of HIFU can be used to permanently block nerve function, to provide a non-invasive alternative to severing a nerve to treat severe spasticity. Relatively lower doses of HIFU can be used to reversibly block nerve function, to alleviate pain, to achieve an anesthetic effect, or to achieve a cosmetic effect. Where sensory nerves are not necessary for voluntary function, but are involved in pain associated with tumors or bone cancer, HIFU can be used to non-invasively destroy such sensory nerves to alleviate pain without drugs. Preferably, ultrasound imaging synchronized to the HIFU therapy is used to provide real-time ultrasound image guided HIFU therapy of neural structures. | 1. A method for imaging a peripheral nerve, the method comprising the steps of: (a) positioning an ultrasound imaging device in the vicinity of a treatment site of a patient, wherein the treatment site comprises the peripheral nerve; and(b) imaging the treatment site comprising the peripheral nerve, thereby generating an ultrasound image in which the peripheral nerve is displayed. 2. The method of claim 1, wherein the step of imaging the treatment site comprising the peripheral nerve is performed in real-time. 3. The method of claim 1, further comprising the step of imaging the treatment site comprising the peripheral nerve using magnetic resonance imaging (MRI). 4. The method of claim 1, further comprising the step of relating the treatment site comprising the peripheral nerve to a three dimensional (3D) reference point. 5. The method of claim 1, wherein the imaging ultrasound comprises three dimensional (3D) ultrasound. 6. The method of claim 1, wherein the imaging ultrasound comprises two dimensional (2D) ultrasound. 7. The method of claim 1, wherein the ultrasound imaging device is disposed external to the patient. 8. The method of claim 1, wherein the ultrasound imaging device is disposed inside a body of the patient. 9. The method of claim 1, further comprising the step of providing therapy to the treatment site comprising the peripheral nerve. 10. The method of claim 9, wherein the step of providing therapy to the treatment site comprising the peripheral nerve comprises the step of administering a dose of high intensity focused ultrasound (HIFU) to the treatment site comprising the peripheral nerve. 11. The method of claim 10, wherein a HIFU transducer used to administer the dose of HIFU to the treatment site comprising the peripheral nerve and the treatment site in the ultrasound image are related to a common 3D spatial coordinate. 12. The method of claim 11, further comprising the step of utilizing a magnetic sensor to relate the HIFU transducer and the ultrasound image to the common 3D spatial coordinate. 13. The method of claim 10, further comprising the step of relating a focal point of a HIFU transducer used to administer the dose of HIFU to the treatment site comprising the peripheral nerve to a 3D spatial coordinate in the ultrasound image. 14. The method of claim 13, wherein the step of relating the focal point of the HIFU transducer to the 3D spatial coordinate comprises the step of utilizing a magnetic sensor to track a position of the HIFU transducer and an ultrasound imaging transducer used to generate the ultrasound image. | University of Washington | A61B005/055 A61N007/00
US20110009722 | A1 | 2011-01-13 | Historical Trend Icons For Physiological Parameters | 2008-12-24 | Embodiments relate to systems and methods for displaying graphical icons representing a detected medical condition or a sensor fault. Specifically, embodiments of relate to a monitoring system that includes a sensor configured to obtain a physiologic signal from a patient, and a monitor communicatively coupled to the sensor and configured to receive the signal. In an embodiment, the monitor includes a processor adapted to compute physiological data based on the signal and identify a pattern in the physiological data, wherein the pattern relates to a condition of the patient, the sensor, or the monitor. In an embodiment, the processor also selects a graphical icon indicative of the pattern and provides the selected icon to a display. | 1. A monitoring system, comprising: a sensor capable of obtaining a physiologic signal from a patient;a monitor communicatively coupled to the sensor and capable of receiving the signal, the monitor comprising a processor capable of: computing physiological data based on the signal;identifying a pattern in the physiological data, wherein the pattern relates to a condition of the patient, the sensor, or the monitor;selecting a graphical icon indicative of the pattern; anddisplaying the selected icon to a display. 2. The monitoring system of claim 1, wherein the pattern comprises a pattern indicative of airway instability, sleep apnea, substantially normal oxygen saturation level, slow oxygen desaturation, rapid oxygen desaturation, signal interference, weak signal, sensor off, sensor disconnect, bad sensor placement, arrhythmia, venous pulsation, sensor motion, and/or electromagnetic or radio frequency interference, and/or combinations thereof. 3. The monitoring system of claim 1, wherein the pattern comprises a trend of SpO2 data representative of airway instability and the graphical icon depicts characteristics of a generalized trend plot of SpO2 data representative of airway instability. 4. The monitoring system of claim 1, wherein the pattern comprises a trend of SpO2 data representative of sleep apnea and the graphical icon depicts characteristics of a generalized trend plot of SpO2 data indicative of sleep apnea. 5. The monitoring system of claim 1, wherein the pattern comprises a trend plot of SpO2 data representative of a substantially normal oxygen saturation level and the graphical icon depicts characteristics of a generalized trend plot of SpO2 data representative of a substantially normal oxygen saturation level. 6. The monitoring system of claim 1, wherein the pattern comprises a trend plot of SpO2 data representative of slow oxygen desaturation and the graphical icon depicts characteristics of a generalized trend plot of SpO2 data representative of a slow oxygen desaturation. 7. The monitoring system of claim 1, wherein the pattern comprises a trend of SpO2 data representative of rapid oxygen desaturation and the graphical icon depicts characteristics of a generalized trend of SpO2 data representative of rapid oxygen desaturation. 8. The monitoring system of claim 1, wherein the pattern comprises a trend of SpO2 data representative of signal interference and the graphical icon depicts general characteristics of signal interference. 9. The monitoring system of claim 1, wherein the pattern comprises a trend of SpO2 data representative of a weak signal and the graphical icon depicts general characteristics of a weak signal. 10. The monitoring system of claim 1, wherein the processor is capable of assigning a color to the graphical icon based on the pattern. 11. The monitoring system of claim 1, wherein the processor is capable of selecting a color of the graphical icon based on a severity associated with the pattern. 12. A me... | Nellcor Puritan Bennett LLC | A61B005/1455 G06F003/048
US20110046452 | A1 | 2011-02-24 | WIRELESS DEVICE AND SYSTEM FOR MONITORING PHYSIOLOGIC PARAMETERS | 2010-02-12 | A system for monitoring blood pressure and other physiologic parameters is provided. The system is designed such that it can be delivered to the patient with ease and minimal invasion. The system contains at least one self contained implantable sensing device consisting of a sensor, an electrical circuit for signal conditioning and magnetic telemetry, a biocompatible outer surface and seal, an anchoring method, and an external readout device. The implant is small in size so that it may be delivered to the desired location and implanted using a catheter, although direct surgical implantation is also possible. The circuit, sensor, and antenna for telemetry are packaged together and sealed hermetically to the biologic environment. The larger readout unit remains outside the body but proximal to the implant for minimizing communication distance. | 1. A miniature implant for measuring physiologic parameters for medical applications, comprising: Implantable self-contained sensor module for measuring physiologic parameters wherein the module: senses physiologic parameters with a sensor;receives operating power from a magnetic field with an inductor coil;conditions a sensor signal;transmits a conditioned sensor signal to an external signal detection system via magnetic telemetry;outer shell to house sensor module fit with anchoring mechanism for implantation. 2. The miniature implant of claim 1 wherein the implant is delivered and implanted with a catheter. 3. The miniature implant of claim 1 wherein the anchoring mechanism passes through a septum wall and opens on either side of a septal wall, clamping the implant to the wall. 4. The miniature implant of claim 1 wherein the anchoring mechanism is a helical screw. 5. The miniature implant of claim 1 wherein the anchoring mechanism is a tine that expands and catches on a trabeculated area of the heart. 6. The implant of claim 1 wherein said module is used for applications including but not limited to one or more of the following cardiovascular diseases,congestive heart failure,congenital heart diseases,diseases related to intracranial pressure,abdominal aortic aneurysm,intraocular pressure for eye related diseases including but not limited to glaucoma,urinary diseases,gastrointestinal diseases. 7. The implant of claim 1 wherein the sensor measures one or more of the following parameterspressure,temperature,flow,blood composition,blood gas content,chemical composition,acceleration,vibration. | INTEGRATED SENSING SYSTEMS, INC. | A61B005/00
US20110004070 | A1 | 2011-01-06 | DEVICE FOR COORDINATING PATIENT CARE IN HOSPITALS, NURSING HOMES, DOCTORS OFFICES OR THE LIKE AND METHOD FOR OPERATING SUCH A DEVICE | 2008-12-02 | The invention relates to a device for coordinating patient care in hospitals, nursing homes, doctors offices or the like, comprising a plurality of cross-departmental medical care stations (8, 9) that are combined into at least one care group (12, 13, 16, 17), each being stationary in different care regions (1 to 6), wherein at least one mobile care station support unit (10, 11, 14, 15) that is independent from the care groups is associated with each care group (12, 13, 16, 17). Furthermore, each care group (12, 13, 16, 17) is provided with a separate, particularly wireless communication device, with which communication exclusively within each care group (12, 13, 16, 17) and thus independent from communication within a possibly existing other care group (12, 13, 16, 17) can be carried out. A radio device is preferably provided as the communication device. By means of the particular wireless communication with the communication device associated with each care group (12, 13, 16, 17), the mobile care station support units (10, 11, 14, 15), the number thereof being smaller than the number of the stationary medical care stations (8, 9) of said care group (12, 13, 16, 17), can be directed as a function of predetermined care parameters to the stationary medical care stations (8, 9) that require support via the communication device. | 1. Device for coordinating patient care in hospitals, nursing homes, doctors offices or the like: a) comprising a plurality of cross-departmental medical care stations (8, 9) that are combined into at least one care group (12, 13, 16, 17), each being stationary in different care regions (1 to 6);b) wherein at least one mobile, care station support unit (10, 11, 14, 15) that is independent from the care groups is associated with each care group (12, 13, 16, 17); andc) whereby each care group (12, 13, 16, 17) is provided with a particularly wireless communication device, with which communication can be carried out within each care group (12, 13, 16, 17). 2. Device according to claim 1, characterized by the provision of a plurality of care regions (1 to 6) each associated with at least two functionally different stationary medical care stations (8, 9): a) whereby with regard to different care regions (1 to 6) multiple functionally similar stationary medical care stations (8, 9) are provided, each forming a cross-departmental care group (12, 13, 16, 17):b) whereby at least one mobile care station support unit (10, 11, 14, 15) that is independent from the care groups is associated with each care group (12, 13, 16, 17), andc) whereby the care groups (12, 13, 16, 17) are equipped with separate, particularly wireless communication devices that enable communication exclusively within each of the respective care groups (12, 13, 16, 17) and are therefore isolated from a communication within another care group (12, 13, 16, 17). 3. Device according to claim 2, characterized by the number of care station support units (10, 11, 14, 15) of each care group beings smaller than the number of stationary care stations (8, 9) within each care group (12, 13, 16, 17). 4. Device according to claim 3, characterized by the number of mobile care station support units (10, 11, 14, 15) within a care group (12, 13, 16, 17) being predefined in such a way that a single mobile care station support unit (10, 11, 14, 15) is associated with two or three or four functionally similar stationary medical care stations (8, 9) within the respective care group (12, 13, 16, 17). 5. Device according to claim 4, characterized by each cross-departmental care group (12, 13, 16, 17) having a predefined number, preferably between two and five, of stationary, functionally similar medical care stations (8, 9) so that if the number of total available stationary functionally similar medical care stations per care group (12, 13, 16, 17) exceeds this predefined number, the remaining stationary, functionally similar medical care stations (8, 9) of one or more additional care groups (12, 13, 16, 17) are associated as functionally similar care groups for the formation of multiple functionally similar care groups (12, 13, 16, 17) that are independent of each other, which in each case are equipped with different, particularly wireless communication devices so that communication exclusively within each o... | ANNETT WENDORF HANSJORG WAIBEL | A61B005/00
US20110004076 | A1 | 2011-01-06 | SYSTEM AND METHOD FOR COMMUNICATING WITH AN IMPLANT | 2009-01-30 | A system and method for communicating with a medical implant is disclosed. The system (10,210,310,410) includes on-board electronics, a signal generator (15,215), an amplifier (16,216), a coil (14,214), a receiver (22,222), and a processor (20,220). The on-board electronics (100, 110) include a power harvester, a sensor, a microprocessor, and a data transmitter. The signal generator (15,215) generates a first signal, the amplifier (16,216) amplifies the first signal, the coil (14,214) transmits the amplified signal, the power harvester receives the first signal and transmits a data packet (18,218) containing data, the receiver (22,222) receives the data packet (18,218), and the processor (20,220) either processes the data or sends the data to a data storage device. | 1. A system for communicating patient information, the system comprising: a medical implant, the medical implant having a first cavity and a second cavity, the first and second cavity connected by one or more apertures, the first cavity is adapted to receive on-board electronics, the on-board electronics comprising at least one sensor, a microprocessor, and a data transmitter, and the second cavity is adapted to receive an implant antenna;a signal generator adapted to generate a first signal;an amplifier electrically connected to the signal generator;at least one coil electrically connected to the amplifier;a receiver adapted to receive a data packet having data from the implant antenna; anda processor connected to the receiver;wherein the signal generator generates the first signal, the amplifier amplifies the first signal, the at least one coil transmits the amplified signal, the implant antenna receives the first signal and transmits a data packet containing data, the receiver receives the data packet, and the processor processes the data or sends the data to a data storage device. 2. The system of claim 1, wherein the processor is selected from the group consisting of a desktop computer, a laptop computer, a personal data assistant, a mobile handheld device, and a dedicated device. 3. The system of claim 1, wherein the receiver is an antenna with an adapter for connection to the processor. 4. The system of claim 1, wherein the on-board electronics comprise a plurality of sensor assemblies and a multiplexer. 5. (canceled) 6. The system of any of claim 1, wherein there are two coils, and the coils are housed within a paddle. 7. The system of claim 1, further comprising a control unit, and wherein the signal generator and the amplifier are housed within the control unit. 8-9. (canceled) 10. The system of claim 1, wherein the first cavity and the second cavity are orthogonal to one another. 11. The system of claim 1, wherein the first cavity and the second cavity are diametrically opposed. 12. The system of claim 1, wherein at least one of the first cavity and the second cavity further comprise a cover. 13. The system of claim 1, wherein the on-board electronics comprise an LC circuit, a bridge rectifier, a storage capacitor, a wake up circuit, a microprocessor, an enable measurement switch, an amplifier, a Wheatstone bridge assembly, and a modulation switch. 14. (canceled) 15. The system of claim 13, wherein the modulation switch modulates a load signal. 16. The system of claim 15, wherein the load signal is modulated at a frequency between 5 kHz and 6 kHz. 17. A method for communicating information from a medical device, the method comprising: determining that a value of stored voltage in a communication system of a medical device exceeds a threshold wake-up voltage value;enabling operation of a sensor device included in the medical device;receiving a reading from the sensor device, the reading indicating one or more parameters a... | SMITH & NEPHEW, INC. | A61B005/07 H04B005/00
US07869856 | B2 | 2011-01-11 | Encapsulated medical imaging device and method | 2001-01-10 | The present invention relates to an encapsulated medical imaging device (1) and method for imaging the gastroin-testinal tract in patients using optical scanning technologies. | 1. A medical imaging system comprising: an elongated capsule comprising: a housing of swallowable proportions and having a longitudinal axis with a first end and an oppositely positioned second end, comprising:two optical setups, a first optical setup positioned on the longitudinal axis at or near the first end and a second optical setup positioned on the longitudinal axis at or near the second end, for viewing sectors external to the housing and in opposite directions; at least one illuminator illuminating said sectors;at least one light sensitive array receiving light from one of the optical setups and generating image data;electronic circuitry for sampling image data from said light sensitive arrays and converting image data to digital data; anda transmitter receiving said digital data and transmitting said digital data,an external receiver receiving said digital data from said transmitter;an image processor capable of processing said digital data received by said external receiver into at least one image; anda display capable of displaying said at least one image. 2. The medical imaging system as claimed in claim 1, further comprising a coronal array of microlenses distributed coronally to at least one of the optical setups focusing light on a light sensitive array. 3. The medical imaging system as claimed in claim 2, wherein microlenses of said microlenses array have different foci distributed in a known distribution and wherein said image processor is capable of distinguishing focused image data from unfocused image data and disregarding at least a portion of said unfocused data. 4. The medical imaging system as claimed in claim 1, wherein said at least one illuminator comprises light emitting diodes. 5. The medical imaging system as claimed in claim 4, wherein said light emitting diodes emit light in different frequencies. 6. The medical imaging system as claimed in claim 5, wherein said light emitting diodes emit red, green, and blue light. 7. The medical imaging system as claimed in claim 1, wherein said longitudinal axis is from about 12 to 20 mm in length and said capsule housing has a diameter of from about 5 to 7 mm. 8. The medical imaging system as claimed in claim 1, wherein said capsule housing is made of biocompatible material. 9. The medical imaging system as claimed in claim 1, wherein said capsule housing comprises disposable material. 10. The medical imaging system as claimed in claim 9, wherein said capsule housing comprises dissolvable material sufficiently stable to pass through a patient's entire gastrointestinal tract without disintegrating. 11. The medical imaging system as claimed in claim 1, wherein said capsule is internally powered. 12. The medical imaging system as claimed in claim 1, wherein said capsule is externally powered. 13. The medical imaging system as claimed in claim 12, wherein said capsule is inductively powered. 14. The medical imaging system as claimed claim 1, wherein said capsule is fu... | Moshe Refael Avi Goldsobel Engel Research & Development (1993) Ltd. L.Z. Ashlang Finance Ltd. | A61B005/00
US20110034821 | A1 | 2011-02-10 | Increasing the information transfer rate of brain-computer interfaces | 2009-08-10 | Methods of increasing the rate of information transfer in brain-computer interface systems are disclosed. The present invention also discloses methods, devices and systems for the navigation of information representing neuronal or brain activity and the extraction of useful and/or actionable data from such information. | 1. An apparatus for extracting information from neurons or similar entities or simulation of same, said apparatus comprising one or more sensor elements responsive to signals from said neurons or similar entities and transforming said signals into one or more representative formats. 2. The apparatus recited in claim 1 wherein said one or more sensor elements is adapted to generate signals corresponding to the state of said neurons or similar entities or simulation of same. 3. The sensor elements recited in claim 2 wherein said signal corresponding to said state of said one or more neurons or similar entities or simulation of same is electromagnetic. 4. The sensor elements recited in claim 2 wherein said signal corresponding to said state of said one or more neurons or similar entities or simulation of same is acoustic or ultrasonic. 5. An apparatus for extracting one or more representations of salient features underlying the activities or states of neurons or similar entities or simulation of same. 6. An apparatus adapted to perform hierarchical decomposition of the feature space of features extracted from one or more representations of the activities or states of neurons or similar entities or simulation of same to provide a means of identifying and adaptively modifying/classifying simpler features (that are more likely to have characteristics common to most subjects) in parallel which are then re-combined to generate a signal or set of signals thus obviating or at least mitigating the need for extensive subject training. 7. A method of extracting information from neurons or similar entities or simulation of same, said method using input from one or more sensor elements or simulations thereof responsive to signals from said neurons or similar entities or simulation of same and transforming said signals into one or more representative formats. 8. The method recited in claim 7 wherein said one or more sensor elements or simulation thereof is adapted to generate signals corresponding to the state of said neurons or similar entities or simulation of same. 9. The sensor elements recited in claim 8 wherein said signal corresponding to said state of said one or more neurons or similar entities or simulation of same is electromagnetic. 10. The sensor elements recited in claim 8 wherein said signal corresponding to said state of said one or more neurons or similar entities or simulation of same is acoustic or ultrasonic. 11. A method of extracting one or more representations of salient features underlying the activities or states of neurons or similar entities or simulation of same. 12. A method of performing hierarchical decomposition of the feature space of features extracted from one or more representations of the activities or states of neurons or similar entities or simulation of same to provide a means of identifying and adaptively modifying/classifying simpler features (that are more likely to have characteristics common to most subject... | - | A61B005/0478 G06G007/60
US07892180 | B2 | 2011-02-22 | Head-stabilized medical apparatus, system and methodology | 2007-03-05 | Structure and methodology involving mountable and head-wearable frame structure which is positionally stabilized, during use, relative a human subject's head, and which carries a selection of positionally anchored data sensors, and stimuli deliverers, that are relevant to the diagnosis and treatment of vestibular disorders. Special configurations are provided for two types of stimulators, one for sound application and air-pressure modification, and the other for the introduction of fluids to the ear. Stabilization enables tight and accurate correlation of data which is quickly analyzable by a connected, properly algorithmed computer, which can also be used for feedback control in a designed “expert” system. The invention enables, among other things, practical and significant differentiation between physiological and pathological nystagmus. | 1. A method for preparing vestibular-response data for useful analysis regarding a desired diagnosis and prospective treatment of a subject's related vestibular disorder comprising: coupling an eye-enclosing, ambient light-excluding flame structure to and in a positionally stable non-relative-motion condition with a subject's head, wherein the frame structure includes a likewise positionally stable electronic video image-collecting sensor apparatus anchored thereto and positioned to observe one or both of the subject's eyes;applying to the subject a vestibular-response-inducing stimulus and producing a data record corresponding to the applied stimulus;collecting from the subject, concurrently with the applied stimulus, data indicating the subject's detectable vestibular-response relative to the applied stimulus;collecting from a geographical direction sensor data indicating a spatial position of the subject's head;producing a record corresponding to the vestibular-response data, wherein the vestibular-response data am thereafter selectively correlateable, whether individually or plurally, with respective corresponding contemporaneously-produced data of the stimulus data record; andelectronically analyzing, by a computer, the vestibular-response data and the geographical direction sensor data relative to one or more characteristics of the applied stimulus to evaluate the subject's related vestibular disorder. 2. The method of claim 1, wherein the collecting data indicating the subject's detectable vestibular-response comprises capturing, utilizing the electronic video image-collecting sensor apparatus, images of subject nystagmus behavior. 3. The method of claim 2, further comprising: utilizing the electronic video image-collecting device to observe nystagmus behavior while simultaneously delivering selected liquid to the ear employing a fluid-flow structure coupled with the ear. 4. The method of claim 3, further comprising: providing a controller for controlling the delivery of liquid to the ear by the fluid-flow structure, and affecting the controlling operation of the controller by utilizing data derived from observations of simultaneously observable nystagmus behavior made by the electronic video image-collecting device. 5. The method of claim 2, further comprising: identifying, by vector analysis of the collected nystagmus data relative to contemporaneously produced stimulus data, one or more originating semicircular canals of an observed nystagmus response. 6. The method of claim 1, wherein the stimulus is applied to the subject's head by a stimulus apparatus operatively coupled in a non-relative motion condition with the subject's head. 7. The method of claim 6, wherein the stimulus is applied using a stimulus apparatus comprises one or more selected from among (a) a sound deliverer, (b) an air-pressure modifier operatively coupled with the ear, (c) a fluid-flow structure operatively coupled with the ear, (d) a light-emitting structur... | Epley Research LLC | A61B005/00 A61B013/00 A61B005/103 A61B005/117 A61M031/00
US07892189 | B2 | 2011-02-22 | Movement analysis display apparatus and movement analyzing method | 2006-02-10 | A movement analysis display apparatus includes a analyzing unit for analyzing waveform data of a time series which is acquired from a movement sensor, and a display unit for displaying an analysis result that is analyzed by the analyzing unit. The analyzing unit includes a movement waveform generating section for generating a movement waveform from the waveform data, an energy balance value calculating means for calculating a movement energy ratio in opposed movement directions in a velocity waveform, an envelope curve generating section for generating an envelope curve at a given time width in a distance waveform, and a time integration curve analyzing section for integrating a movement waveform with respect to time. | 1. A movement function analysis display apparatus, comprising: a movement sensor for sensing and acquiring movement data of a live subject including waveform data of a time series;an analyzing unit coupled to the movement sensor, for analyzing the waveform data of the time series which is acquired from the movement sensor; anda display unit for displaying an analysis result that is analyzed by the analyzing unit,wherein the analyzing unit comprises:a movement waveform generating unit for generating a movement waveform corresponding to the waveform data from the waveform data, and for generating a velocity waveform from the movement waveform; andan energy balance value calculating unit for evaluating a movement energy ratio in positive and negative directions by using a positive velocity value and a negative velocity value from the velocity waveform,wherein the energy balance value calculating unit calculates a second-power integration value of the positive velocity value and a second-power integration value of the negative velocity value in the velocity waveform,wherein the energy balance value calculating unit calculates an energy balance value from a ratio of the second-power integration values after the second-power integration values are calculated, andwherein the analyzing unit evaluates the movement energy ratio in the positive and negative directions from the energy balance value, and determines a severity of a functional disorder of the live subject based on a magnitude of the movement energy ratio in the positive and negative directions so as to produce the analysis result for a visual display on the display unit. 2. The movement function analysis apparatus according to claim 1, wherein the live subject is a Parkinson's patient. 3. The movement function analysis apparatus according to claim 1, wherein the functional disorder of the live subject is a brain disorder which includes a Parkinson's disease. | Hitachi, Ltd. | A61B005/103 A61B005/117 A61B005/22 A61B001/24 A63B021/00 A63B023/16 G01L003/24 G01L005/00
US08257344 | B2 | 2012-09-04 | Deflectable catheter | 2010-10-07 | A deflectable catheter comprises a catheter body, a tip section, and a control handle for affecting deflection of the tip section. The tip section comprises a flexible tubing having proximal and distal ends and at least two lumens extending therethrough. The proximal end of the tip section is fixedly attached to the distal end of the catheter body. The flexible tubing comprises a plastic core, a metal braid over the core, a plastic skin over the braid, and at least two pieces of plastic tubing in longitudinal arrangement extruded over the plastic skin, wherein the at least two pieces of plastic tubing have different stiffnesses and are arranged with the stiffer plastic being proximal the less stiff plastic. The tip section is made by extruding a plastic core over at least one mandrel and braiding a metal over the plastic core. Thereafter, at least two pieces of plastic tubing having different degrees of stiffness are assembled along the length of the plastic core so that the piece of plastic tubing having the greatest stiffness is at the proximal end of the tip section and the piece of plastic having the lowest stiffness is near the distal end of the tip section. The pieces of plastic tubing are then melted to the plastic core. | 1. A deflectable catheter comprising: a catheter body having an outer wall, proximal and distal ends, and at least one lumen extending therethrough;a tip section comprising a flexible tubing having proximal and distal ends and at least two lumens extending therethrough, the proximal end of the tip section being fixedly attached to the distal end of the catheter body,a control handle mounted at the proximal end of the catheter body, the control handle comprising: a piston receiving assembly comprising a proximal barrel and a distal barrel, the distal barrel comprising a piston chamber at its distal end, anda piston longitudinally slidable within the piston chamber along a longitudinal axis of the control handle, wherein the distal barrel of the piston receiving assembly engages a proximal end of the piston to substantially prevent rotation of the piston relative to the piston receiving assembly; andmeans for deflecting the tip section by manipulation of the control handle. 2. A catheter according to claim 1, wherein the flexible tubing comprises at least three longitudinally arranged pieces of tubing, each piece of tubing having a different stiffness. 3. A catheter according to claim 1, wherein the flexible tubing comprises at least five longitudinally arranged pieces of tubing, each piece of tubing having a different stiffness. 4. A catheter according to claim 1, wherein the tip section has at least three lumens extending therethrough. 5. A catheter according to claim 1, further comprising a stiffening tube lining the outer wall, said stiffening tube having a distal end and a proximal end. 6. A catheter according to claim 5, wherein the distal end of the stiffening tube extends into the proximal end of the tip section. 7. A catheter according to claim 5, wherein the stiffening tube comprises polyimide. 8. A catheter according to claim 1, wherein the deflecting means comprises a puller wire having a proximal end and a distal end, the puller wire extending from the control handle, through the catheter body and into a lumen in the tip section. 9. A catheter according to claim 1, further comprising a tip electrode mounted on the distal end of the tip section. 10. A catheter according to claim 1, wherein the flexible tubing of the tip section comprises at least two pieces of tubing longitudinally arranged over a plastic housing, each piece of tubing having a different stiffness. 11. A deflectable catheter comprising: a catheter body having an outer wall, proximal and distal ends, and at least one lumen extending therethrough;a tip section comprising a flexible tubing having proximal and distal ends and at least two lumens extending therethrough, the proximal end of the tip section being fixedly attached to the distal end of the catheter body;a control handle mounted at the proximal end of the catheter body, the control handle comprising: a housing comprising a generally hollow interior,a piston receiving assembly at least partially within th... | Biosense Webster, Inc. | A61M025/00
US07885837 | B1 | 2011-02-08 | Pension Insurance Program Methods and Systems | 2010-05-20 | A pension insurance plan system and method protect the outstanding balance of a defined contribution plan loan secured by participant's plan assets, eliminating adverse tax consequences to the plan beneficiary in the event of the death or disability of the participant and ensuring that the full plan account balance passes intact to beneficiary in such event. | 1. A computer-implemented method for administering a premium, the method comprising: providing a debt mitigation strategy to a plan sponsor or a plan administrator of a retirement plan, wherein the plan sponsor or plan administrator amends a plan document to require a debt mitigation strategy as a condition of a plan loan;wherein a participant has a retirement account in the retirement plan,wherein the participant takes a loan from the assets of the retirement account and the plan loan is collateralized by the remaining assets in the retirement account, andwherein the debt mitigation strategy insures a participant's plan loan against non-repayment in the event of occurrence of a risk event subjecting the plan loan to potential default for non-repayment;calculating, by a computer, the consideration for the debt coverage based upon the outstanding balance of the plan loan and a state of domicile of the retirement plan issuing the plan loan;calculating, by the computer, a repayment schedule for the loan, wherein the repayment schedule reflects the consideration to be charged on a periodic basis to a payroll account of the participant or a participant's retirement account, based upon a choice by the plan sponsor or the plan administrator;billing, by a computer, periodically until the plan loan is repaid in full; andcollecting, by a computer, the consideration from the participant and maintaining on a computer a data structure regarding the collection of the consideration for insurance, self insurance, or debt cancellation agreements through the debt mitigation strategy,wherein, in the event of a risk event, the assets in the retirement account are preserved for the benefit of a beneficiary of the retirement account for rollover into another tax deferred retirement account or distribution to the beneficiary. 2. The computer-implemented method according to claim 1, wherein the debt mitigation strategy comprises an insurance policy. 3. The computer-implemented method according to claim 1, wherein the debt mitigation strategy comprises a self-insurance plan. 4. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises a debt cancellation agreement. 5. The computer-implemented method according to claim 1, wherein the debt mitigation strategy is provided in response to a requirement for insurance covering a balance of the plan loan, imposed as a pre-condition to making the plan loan. 6. The computer-implemented method according to claim 1, further comprising, upon occurrence of the risk event ensuring payment, of an amount equal to any unpaid plan loan balance, to the plan account. 7. The computer-implemented method according to claim 1, wherein the step of ensuring payment comprises the step of effecting payment. 8. The computer-implemented method according to claim 1, wherein the risk event comprises death of the participant. 9. The computer-implemented method according to claim 1, wherein the risk event com... | Pentech, LLC | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07881762 | B2 | 2011-02-01 | Clip-style medical sensor and technique for using the same | 2005-09-30 | A clip-style sensor is provided that includes a sliding clip, such as a flat spring that slides along the sensor to provide a closing force for the sensor. When the sliding clip is engaged, the sensor is secured to the patient. The sensor may be placed on a patient's finger, toe, ear, and so forth to obtain pulse oximetry or other spectrophotometric measurements. | 1. A sensor, comprising: a sensor body comprising a first portion and a second portion, wherein at least one of the first portion and the second portion comprises an alignment feature;at least one sensing element disposed on the sensor body; anda sliding clip adapted to slide along the sensor body to bias the first portion towards the second portion, wherein the sliding clip is adapted to apply a pressure to a patient's tissue sufficient to overcome the patient's venous blood pressure but not sufficient to overcome the patient's arterial blood pressure, and wherein the sliding clip is generally U-shaped and configured to be disposed along the alignment feature. 2. The sensor, as set forth in claim 1, wherein the sliding clip comprises a flat spring. 3. The sensor, as set forth in claim 1, wherein the sensor body comprises a restraining clip adapted to hold the sliding clip on the sensor body. 4. The sensor, as set forth in claim 3, wherein the sliding clip comprises a guard adapted to collide with the restraining clip to block the sliding clip from being detached from the sensor body. 5. The sensor, as set forth in claim 1, wherein the applied pressure is between about 10 mm Hg and about 80 mm Hg. 6. The sensor, as set forth in claim 1, wherein the applied pressure is between about 10 mm Hg and about 35 mm Hg. 7. A sensor, comprising: a sensor body comprising a first portion and a second portion, wherein at least one of the first portion and the second portion comprises an alignment feature;at least one sensing element disposed on the sensor body; anda sliding clip adapted to move between an engaged position and a disengaged position, wherein the first portion and the second portion are biased towards one another when the sliding mechanism is in the engaged position, wherein the sliding clip is adapted to apply a pressure to a patient's tissue sufficient to overcome the patient's venous blood pressure but not sufficient to overcome the patient's arterial blood pressure, and wherein the sliding clip is generally U-shaped and configured to be disposed along the alignment feature. 8. The sensor, as set forth in claim 7, wherein the sensor comprises at least one of a pulse oximetry sensor or a sensor for measuring a water fraction. 9. A pulse oximetry system, comprising: a pulse oximeter monitor; anda pulse oximetry sensor configured to communicate with the pulse oximetry monitor, the pulse oximetry sensor comprising: a sensor body comprising a first portion and a second portion, wherein at least one of the first portion and the second portion comprises an alignment feature;at least one sensing element disposed on the sensor body; anda sliding clip adapted to slide along the sensor body to bias the first portion towards the second portion, wherein the sliding clip is adapted to apply a pressure to a patient's tissue sufficient to overcome the patient's venous blood pressure but not sufficient to overcome the patient's arterial blood pressure,... | Nellcor Puritan Bennett LLC | A61B005/1455
US08111953 | B2 | 2012-02-07 | Optical power modulation vital sign detection method and measurement device | 2010-10-04 | A vital sign measurement device includes a sensor fixation device, a sensor frame, an optical sensing system, and an output unit. The sensor fixation device is adapted to be placed against an anatomical location of a subject. The optical sensing system includes an optical waveguide, an optical source device to supply optical energy to the optical waveguide, and an optical detector to detect an amount of optical energy exiting the optical waveguide. The optical sensing system is adapted to sense an arterial pulse from the compression or flexing of at least a portion of the optical waveguide resulting in reduction of the amount of light exiting the optical waveguide. The output unit is configured to receive a signal indicative of the amount of light exiting the optical waveguide and to generate a measure of the vital sign based at least in part on the received signal. | 1. A vital sign measurement device comprising: a sensor frame adapted to be held against an external skin surface of an anatomical location of a subject, within which is an artery;an optical sensing system held by the sensor frame and comprising an optical waveguide, an optical source device to supply optical energy to a first end of the optical waveguide, and an optical detector to detect an amount of optical energy exiting the second end of optical waveguide, the optical sensing system adapted to sense an arterial pulse from the compression or flexing of at least a portion of the optical waveguide resulting in reduction of the amount of optical energy exiting the second end of the optical waveguide; andan output unit configured to receive a signal indicative of the amount of light exiting the second end of the optical waveguide and to generate a measure of the vital sign based at least in part on the received signal. 2. The vital sign measurement device of claim 1, wherein the sensor frame is held against the external skin surface by a cuff comprising an inflatable bladder adapted to apply pressure to the anatomical location when inflated and thereby compress the artery within the anatomical location. 3. The vital sign measurement device of claim 1, further comprising a pressure sensor to detect a pressure applied to the anatomical location, wherein the output unit receives a pressure input indicative of the pressure applied to the anatomical location from the pressure sensor, and wherein the output unit generates the vital sign using the signal indicative of the optical signal received and the pressure input. 4. The vital sign measurement device of claim 1, wherein the anatomical location of the subject is an upper arm, and the sensor frame is configured on the sensor fixation device so that the optical sensing system is positioned to sense movement due to a pulse of a brachial artery resulting in the compression or flexing of at least a portion of the compressible optical waveguide. 5. The vital sign measurement device of claim 1, wherein the optical waveguide comprises an elastomer. 6. The vital sign measurement device of claim 5, wherein the elastomer is selected from the group consisting of polysiloxane, polyurethane, polybutadiene rubber, and combinations thereof. 7. The vital sign measurement device of claim 1, further comprising a waveguide support structure comprising a non-compliant surface, the waveguide support structure supporting at least a portion of the optical waveguide. 8. The vital sign measurement device of claim 1, further comprising a flexible and incompressible support surface that supports the optical waveguide over substantially all of its length. 9. The vital sign measurement device of claim 8, wherein the optical source device and the optical detector are mounted on the surface of the waveguide support surface. 10. The vital sign measurement device of claim 1, wherein the optical sensing system is configured ... | Tarilian Laser Technologies, Limited | G02B006/00 A61B005/02
US07895054 | B2 | 2011-02-22 | Pharmacy personal care account | 2005-08-04 | A method for providing insurance coverage to a subscriber comprises offering an insurance policy to said subscriber wherein said insurance policy comprises providing allowances to pay for one or more expenses covered by said insurance policy and providing monetary credits to said subscriber for any unused portion of a given allowance where said monetary credits can be used to cover the cost of a subsequent allowed expense. The insurance policy may be a pharmacy benefits plan. The covered expenses may be pharmacy expenses. The allowed expenses may be medical expenses which qualify under section 213 of the US Internal Revenue code for payment by an employer without said subscriber having to declare said payment as taxable income. | 1. A computerized method for an insurance company to provide insurance coverage to a subscriber, said method comprising; a) storing in a database at a server insurance policy data, said insurance policy data comprising; i. data for a plurality of allowances to pay for covered products or services, wherein each allowance for a covered product or service comprises to a maximum amount said insurance company pays for said covered product or service;ii. data of the accumulation of monetary credits in a personal care account assigned to said subscriber, said monetary credits being equal to the difference between the cost of said covered product or service and the allowance for said covered product or service;iii. a provision for said subscriber to cover the cost of at least a portion of an allowed expense using at least a portion of said accumulated monetary credits, said allowed expense specified in said insurance policy;b) accepting the payment of a premium for said insurance policy, said step of accepting said payment of said premium being necessary for said insurance policy to be in force;c) receiving at said server election information of said subscriber, said election information comprising personal identification data for said subscriber and said subscriber's election of said insurance policy;d) electronically transferring said election information to said insurance company;e) storing said election information of said subscriber in association with said personal care account within said electronic database;f) processing at a computer claim data for said covered product or service according to said insurance policy elected by said subscriber, said claim data comprising an actual cost for said covered product or service;g) accessing at said computer said database to locate an allowance for said covered product or service;h) calculating at said computer a monetary credit associated with said claim by calculating a difference between said allowance for said covered product or service and said actual cost for said covered expense; andi) if said difference is: i) positive, updating said database to increase by said monetary credit a balance in said personal care account for a subscriber; andii) negative or zero, leaving said balance in said personal care account for a subscriber unchanged. 2. The computerized method of claim 1 wherein said insurance policy comprises a pharmacy benefits plan wherein said covered products or services are prescription drugs and said allowed expenses are medical costs of said subscriber. 3. The computerized method of claim 2 wherein said allowances are determined by the anticipated impact of a given prescription drug on the anticipated subsequent health care costs that said subscriber would experience if said subscriber did not take said prescription drug when prescribed. 4. The computerized method of claim 2 wherein said medical costs are those that qualify for payment by an employer without said payment resulting in ... | Humana Inc. | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US08209001 | B2 | 2012-06-26 | Monitoring for mitral valve regurgitation | 2010-10-29 | Implantable systems, and methods for use therein, for monitoring for mitral valve regurgitation (MR) are provided. An electrogram (EGM) signal and a corresponding pressure signal are obtained, where the EGM signal is representative of electrical functioning of the patient's heart during a plurality of cardiac cycles, and the corresponding pressure signal is representative of pressure within the left atrium the patient's heart during the cardiac cycles. Windows of the pressure signal are defined, based on events detected in the EGM signal, and measurements from the windows are used to monitor for MR. | 1. In an implantable system, a method for monitoring a patient's heart for mitral valve regurgitation (MR), comprising: (a) obtaining an electrogram (EGM) signal and a corresponding pressure signal, the EGM signal representative of electrical functioning of the patient's heart during a plurality of cardiac cycles, the corresponding pressure signal representative of pressure within the left atrium the patient's heart during the cardiac cycles;(b) for each of a plurality of cardiac cycles represented in the EGM and pressure signals, (b.1) determining an average of a first window of the pressure signal,wherein a start of the window is defined relative to an event detected in the EGM signal, andwherein a length of the window is defined to include at least one of an a-wave and a c-wave, but not a v-wave, of the cardiac cycle represented in the pressure signal; and(b.2) determining an average of a second window of the pressure signal,wherein a start of the second window is defined relative to an event detected in the EGM signal or relative to the first window, andwherein a length of the second window is defined to include the v-wave of the cardiac cycle represented in the pressure signal;(c) comparing one or more average determined for a said first window to one or more average determined for a said second window; and(d) monitoring for MR based on results of the comparing. 2. The method of claim 1, wherein: step (c) includes determining a ratio of the average determined for one or more said first window to the average determined for one or more said second window, or vice versa; andstep (d) includes detecting MR by comparing the ratio to a threshold. 3. The method of claim 1, wherein: step (c) includes determining a difference between the average determined for one or more said first window to the average determined for one or more said second window; andstep (d) includes detecting MR when the difference exceeds a threshold. 4. The method of claim 1, wherein the length of the second window is substantially equal to a length of a cardiac cycle. 5. An implantable system for monitoring a patient's heart for mitral valve regurgitation (MR), comprising: one or more electrode to obtain an electrogram (EGM) signal representative of electrical functioning of the patient's heart during a plurality of cardiac cycles;one or more sensor to obtain a corresponding pressure signal representative of pressure within the left atrium the patient's heart during the cardiac cycles; andone or more processor to determine, for each of a plurality of cardiac cycles represented in the EGM and pressure signals, an average of a first window of the pressure signal, wherein a start of the window is defined relative to an event detected in the EGM signal, and wherein a length of the window is defined to include at least one of an a-wave and a c-wave, but not a v-wave, of the cardiac cycle represented in the pressure signal; andan average of a second window of the pressure signal... | Pacesetter, Inc. | A61B005/04
US07887489 | B2 | 2011-02-15 | System and method for detecting sleepiness | 2008-08-29 | The present invention relates to a drowsiness detection method. A heartbeat signal and a breathing signal are detected by exploiting together a scheme and an optical system scheme. The detected signals are applied to respective amplification units, noise signals are eliminated from the detected signals, and noise-free signals are amplified. The amplified signals are applied to a central processing unit, signal processing is processed on the signals, and processed signals are combined. The combined signal is counted, and a warning sound, voice message or vibration is output in a case where a value, obtained by subtracting a counted output value monitored one minute before a current time, from a counted output value monitored two minutes before the current time, falls within a detection range and where, with a passage of time, the value falling within the detection range is successively detected from two to ten times. | 1. A drowsiness detection method, comprising: a first step of detecting a heartbeat signal and a breathing signal without making direct contact with a target body by exploiting together a scheme using a principle of a Doppler radar and an optical system scheme using a light source/photodetector so that a heartbeat can be detected when a human body is moving;a second step of applying the signals, detected by the scheme using the principle of the Doppler radar and the optical system scheme using the light source/photodetector to respective amplification units, eliminating noise signals from the detected signals, respectively, and amplifying noise-free signals;a third step of applying signals output at the second step to a central processing unit, performing signal processing on the signals, and combining respective processed signals; anda fourth step of counting an output value output at the third step, and outputting a warning sound, voice message or vibration through an output unit in a case where a value, obtained by subtracting a counted output value which was monitored one minute before a current time, from a counted output value which was monitored two minutes before the current time, falls within a range from 2 to 10, or from −2 to −10 and where, with a passage of time, the value falling within the range from 2 to 10 or from −2 to −10 is successively detected from two to ten times within a same range. 2. The drowsiness detection method according to claim 1, wherein the optical system comprises a filter unit arranged upstream of a photodetection unit and configured to allow only light, having a specific wavelength and a predetermined light source bandwidth, to pass therethrough. 3. The drowsiness detection method according to claim 2, wherein the light source bandwidth is less than 100 nm. 4. The drowsiness detection method according to claim 2, wherein the light source is implemented using a Light Emitting Diode (LED) or a Laser Diode (LD). 5. The drowsiness detection method according to claim 2, wherein the light source uses visible, infrared or ultraviolet rays. 6. The drowsiness detection method according to claim 1, wherein the light source is implemented using a Light Emitting Diode (LED) or a Laser Diode (LD). 7. The drowsiness detection method according to claim 1, wherein the light source uses visible, infrared or ultraviolet rays. 8. A drowsiness detection system, comprising: a heart rate monitor including a radar system and an optical system, an amplification unit, a central processing unit and an output unit,wherein the radar system comprises: a signal generator for generating specific frequency signals;a coupler for receiving the signals generated by the signal generator and dividing a second signal, which is a reference frequency signal, and a first signal to be radiated to a heart of a target body, and outputting the first and second signals;a transmission antenna for receiving the first signal from the coupler and radia... | Telecommunication & Electronics Inc. | A61B005/02
US07885700 | B2 | 2011-02-08 | Electrodynamic sensors and applications thereof | 2002-12-09 | An electrodynamic sensor comprises a high input impedance electrometer adapted to measure small electrical potentials originating from an object under test and having a pair of input probes, characterized in that at least one of said pair of input probes has no direct electrical contact with said object, wherein the circuit arrangement of said electrometer comprises an amplifier which includes a combination of ancillary circuits cumulatively to increase the sensitivity of said electrometer to said small electrical potentials whilst not perturbing the electrical field associated therewith. | 1. An electrodynamic sensor comprising: an input probe for detecting small electrical potentials originating from an object under test and for generating detection signals, said input probe being configured for capacitive coupling with said object without direct electrical contact and comprising an electrode substrate having thereon a thin film of insulating material, anda high input impedance electrometer arranged to receive the detection signals and measure said electric potentials, the electrometer comprising:an amplifier,a dc input bias arrangement for providing a steady input bias current for said amplifier, said dc input bias arrangement including a biasing resistor,a dc to low frequency negative feedback loop comprising a low-pass filter circuit connected to receive amplified measurement signals and to feed a filtered stabilizing signal into the dc input bias arrangement for maintaining a stable dc operating point for the amplifier; anda combination of ancillary circuits configured cumulatively to increase the sensitivity of said electrometer to said small electrical potentials whilst not perturbing the electrical field associated therewith, the ancillary circuits including a guard circuit providing a shield surrounding said electrode substrate and input wiring for said amplifier, said guard circuit being driven by the output of the amplifier and being arranged to maintain the same potential on said one probe as on said shield, and a feedback arrangement between the output and the input of the amplifier, said feedback arrangement comprising a bootstrapping circuit for bootstrapping the biasing resistor of the dc input bias arrangement, and one of a neutralization circuit, a supply modulation circuit and an offset correction circuit for said amplifier. 2. An electrodynamic sensor according to claim 1 in which the feedback arrangement comprises at least bootstrapping and neutralization circuits. 3. An electrodynamic sensor according to claim 1 in which said one probe comprises a disc electrode, and in which said thin film of insulating material comprises an anodized oxide film. 4. An electrodynamic sensor according to claim 3 in which the thickness of said thin film is approximately 1 μm. 5. An electrodynamic sensor according to claim 3 in which the thickness of the oxide film is selected in dependence upon the surface area of the disc electrode and the material composing the disc electrode, in order to provide the capacitance required. 6. An electrodynamic sensor according to claim 1 in which said one probe comprises an electrode with a guarded coaxial structure. 7. An electrodynamic sensor according to claim 1 in which said one probe comprises a measurement electrode sandwiched between two guard planes in a tri-plate configuration. 8. An electrodynamic sensor according to claim 1 comprising a plurality of said probes configured as a linear or higher dimensional matrix array. 9. An electrodynamic sensor according to claim 1 in which... | The University of Sussex | A61B005/04
US07889070 | B2 | 2011-02-15 | Methods, systems, devices and computer program products for transmitting medical information from mobile personal medical devices | 2006-10-17 | A method of transferring medical information can include transmitting medical information from a mobile personal medical device responsive to determining that a remote system is authorized to receive the medical information. Related Systems, devices, and computer program products are also disclosed. | 1. A method of transferring medical information comprising: determining whether a syndrome associated with a body is detected;determining whether a measurement of a parameter associated with the body exceeds a threshold;transmitting medical information from a mobile personal medical device to a remote system responsive to determining that the remote system is authorized to receive the medical information wherein authorization for the remote system to receive the medical information is provided responsive to determining that the syndrome associated with the body detected by the mobile personal medical device indicates an emergency or that the measurement of the parameter exceeds the threshold, wherein determining that the remote system is authorized to receive the medical information further comprises:if the syndrome indicates no emergency or that the measurement does not exceed the threshold; thencomparing a key received at the mobile personal medical device as part of the authorization to a key previously transmitted to the mobile personal medical device; andauthorizing the transmission of the medical information to the remote system only if the key received at the mobile personal medical device matches the key previously transmitted to the mobile personal medical device. 2. A method according to claim 1 wherein transmitting comprises transmitting the measurement of the parameter associated with the body from the mobile personal medical device to a device that is outside the body for storage therein. 3. A method according to claim 2 wherein the outside device comprises a recognized emergency medical device, the method further comprising: transmitting the measurement of the parameter associated with the body from the mobile personal medical device to the recognized emergency medical device responsive to determining that an emergency medical provider is authorized to use the recognized emergency medical device. 4. A method according to claim 3 further comprising: transmitting the measurement of the parameter associated with the body from the recognized emergency medical device to a persistent long-term storage device. 5. A method according to claim 1 wherein transmitting further comprises: wherein authorization for the remote system to receive the medical information is provided responsive to receiving a request, at the mobile personal medical device, to transmit the medical information; andwherein transmitting comprises transmitting the medical information if the request includes an emergency medical provider authentication code. 6. A method according to claim 5 further comprising: storing the emergency medical provider authentication code associated with the request in the mobile personal medical device. 7. A method according to claim 1 wherein transmitting further comprises: wherein authorization for the remote system to receive the medical information is provided responsive to establishing a privacy relationship between the mobile person... | AT&T Intellectual Property I, L.P. | G08B001/08 A61B005/00
US07894877 | B2 | 2011-02-22 | System and method for adjusting image parameters based on device tracking | 2003-05-14 | A system for automatically adapting image acquisition parameters based on imaging and/or device tracking feedback is provided. An example system includes a subsystem for acquiring images (e.g., MR) of an object and tracking data for a device (e.g. catheter) inserted into the object and controllably moveable within the object. The system also includes an image processor for processing the images and a device tracking logic for computing device parameters (e.g., speed, direction of travel, rate of speed change, position, position relative to a landmark, device orientation). Based on the images and device parameter computations, a parameter control and adjustment logic can automatically update one or more image acquisition parameters that control the image acquisition subsystem. | 1. A system, comprising: an imaging system programmed to acquire image signal data from an object being imaged, the imaging being based on one or more image acquisition parameters, the object having a device inserted therein;a device tracking system programmed to acquire device tracking signal data associated with the device inserted in the object;a processor;a memory; andan interface to connect a group of logics, the group of logics comprising: logic programmed to track the device from the device tracking signal data and to determine one or more device properties from the device tracking signal data, where at least one device property is a velocity measurement for the device; andlogic programmed to automatically adjust one or more of the image acquisition parameters based, at least in part, on the one or more device properties, the one or more image acquisition parameters comprising: image slice position, slice orientation, slice thickness, slice rotation, scan pattern, field of view, image resolution, type of acquisition, tip angle, image contrast, temporal resolution, spatial resolution, bandwidth, trajectory, TR, and TE. 2. The system of claim 1, where the device is a catheter. 3. The system of claim 2, where the device tracking signal data is image signal data. 4. The system of claim 1, where the logic for tracking includes logic for comparing two or more sets of device tracking signal data and determining a location of the device in the object being imaged. 5. The system of claim 1, where the logic for tracking includes logic for comparing two or more images generated from the image signal data and determining a location of the device in an image. 6. The system of claim 1, where the one or more device properties include one or more of, a speed of the device, a rate of change of the speed of the device, and a direction of travel of the device. 7. The system of claim 1, where the one or more device properties are, the speed of the device, the rate of change of speed of the device, and the velocity measurement of the device. 8. The system of claim 1, where the one or more properties are the velocity of the device and the acceleration of the device. 9. The system of claim 1, where the image acquisition parameters are the field of view and the image resolution. 10. The system of claim 1, where the logic for adjusting one or more of the image acquisition parameters is configured to automatically adjust an image parameter by manipulating a pulse sequence used by one or more of, the imaging system, and the device tracking system. 11. The system of claim 1, where the logic programmed to automatically adjust performs adjustment automatically according to one or more of a sigmoidal function and a threshold function. 12. The system of claim 1, comprising a manual override that allows a user to manually override the logic programmed to automatically adjust. 13. The system of claim 1, where the logic programmed to automatically adjust one or ... | Case Western Reserve University | A61B005/05
US08924161 | B2 | 2014-12-30 | System for use in the treatment of diabetes | 2010-08-11 | A health-monitoring device assesses the health of a user based on levels of two analytes in a biological fluid. A first analyte that is utilized to assess a user's health is a fat metabolism analyte, such as ketones, free fatty acids and glycerol, which is indicative of fat metabolism. A second analyte that is utilized is a glucose metabolism analyte, such as glucose. The levels of the two analytes may be used to assess insulin sensitivity, to detect both recent hypoglycemia and the cause of high glucose levels, and/or to guide therapeutic intervention. The dual analyte model may calculate a discrepancy between an actual insulin activity level and a theoretical insulin activity level. The dual analyte model of the present invention may be used to identify individuals at risk for metabolic syndrome, insulin resistance and non-insulin dependent diabetes, and allows monitoring of the progression of those disease states, as well as progress made by therapeutic interventions. | 1. A system for treating diabetes, comprising: a health-monitoring device, comprising: a housing;a test port coupled to the housing and in communication with the housing, wherein the test port is configured to receive a test strip;a processor coupled to the housing and configured to receive one or more first signals indicative of the concentration of a first analyte and one or more second signals indicative of the concentration of a second analyte in a sample received at the test port, wherein the first analyte is glucose;a storage unit coupled to the housing and comprising one or more programs executable by the processor, the one or more programs comprising an algorithm for determining an insulin dosage for delivery with a medication delivery device, the algorithm involving the concentration of glucose indicated by the one or more first signals indicative of the concentration of a first analyte and the concentration of the second analyte indicated by the one or more second signals indicative of the concentration of a second analyte; anda display coupled to the housing to output data resulting from the one or more first and/or second signals and/or an insulin dosage determined by the algorithm. 2. The system of claim 1, wherein the test strip is an electrochemical test strip or a photometric test strip. 3. The system of claim 1, wherein the housing further comprises a data communication port operatively coupled to the processor. 4. The system of claim 3, wherein the data communication port is configured for one or more of electrical communication, acoustic communication, optical communication, or radio wave communication. 5. The system of claim 3, wherein the data communication port is configured for infrared communication. 6. The system of claim 3, wherein the data communication port is configured to communicate with a database at a remote site. 7. The system of claim 6, wherein the data communication port is configured to download data from the database at the remote site through a network. 8. The system of claim 6, wherein the data communication port is configured to upload data to the database at the remote site through a network. 9. The system of claim 8, wherein the data comprises the one or more signals indicative of the concentration of glucose and/or the one or more signals indicative of the concentration of the second analyte in the sample received at the test port. 10. The system of claim 6, wherein the data communication port is configured to download a program to the storage unit through a network. 11. The system of claim 1, wherein the glucose concentration is a blood glucose concentration. 12. The system of claim 11, wherein the display comprises an indicator activatable when the glucose concentration of the sample indicates prospective development of hypoglycemia. 13. The system of claim 11, wherein the display comprises an indicator activatable when the glucose concentration of the sample indicates prospective develop... | Abbott Diabetes Care Inc. | G01N033/48 G01N031/00 G06G007/48 G06G007/58 G06F019/00 A61B005/00 A61B005/145 G01N021/84
US08924160 | B2 | 2014-12-30 | Health-monitoring device | 2010-08-11 | A health-monitoring device assesses the health of a user based on levels of two analytes in a biological fluid. A first analyte that is utilized to assess a user's health is a fat metabolism analyte, such as ketones, free fatty acids and glycerol, which is indicative of fat metabolism. A second analyte that is utilized is a glucose metabolism analyte, such as glucose. The levels of the two analytes may be used to assess insulin sensitivity, to detect both recent hypoglycemia and the cause of high glucose levels, and/or to guide therapeutic intervention. The dual analyte model may calculate a discrepancy between an actual insulin activity level and a theoretical insulin activity level. The dual analyte model of the present invention may be used to identify individuals at risk for metabolic syndrome, insulin resistance and non-insulin dependent diabetes, and allows monitoring of the progression of those disease states, as well as progress made by therapeutic interventions. | 1. A health-monitoring device, comprising: a housing;a test port in communication with the housing, wherein the test port is configured to receive a test strip;a processor in communication with the housing and configured to determine a plurality of blood glucose concentrations over a plurality of days from blood samples deposited on test strips received at the test port;a storage unit in communication with the housing and comprising one ore more programs executable by the processor fordetermining the concentration of glucose in the blood samples deposited on the test strips, anddetecting a pattern in the subject's blood glucose levels determined from the blood samples deposited on the test strips over a same period of time over a plurality of days,wherein the period of time is a sub-period of time within a day; anda display coupled to the housing and configured to output information comprising the determined blood glucose concentration,an indication of the detected pattern in the subject's blood glucose levels over a same period of time over a plurality of days, andan indication of impending hypoglycemia or hyperglycemia in the subject based on the detected pattern in the subject's blood glucose levels over a same period of time over a plurality of days. 2. The health-monitoring device of claim 1, wherein the test strip is an electrochemical test strip or a photometric test strip. 3. The health-monitoring device of claim 1, wherein the housing further comprises a data communication port operatively coupled to the processor. 4. The health-monitoring device of claim 3, wherein the data communication port is configured for one or more of electrical communication, acoustic communication, optical communication, or radio wave communication. 5. The health-monitoring device of claim 3, wherein the data communication port is configured for infrared communication. 6. The health-monitoring device of claim 3, wherein the data communication port is configured to communicate with a database at a remote site. 7. The health-monitoring device of claim 6, wherein the data communication port is configured to download data from the database at the remote site through a network. 8. The health-monitoring device of claim 6, wherein the data communication port is configured to upload data to the database at the remote site through a network. 9. The health-monitoring device of claim 3, wherein the data communication port is configured to download a program to the storage unit through a network. 10. The health-monitoring device of claim 8, wherein the data comprises one or more of the determined blood glucose concentration,the indication of the detected pattern in the subject's blood glucose levels over a same period of time over a plurality of days, andthe indication of impending hypoglycemia or hyperglycemia in the subject based on the detected pattern in the subject's blood glucose levels over a same period of time over a plurality of days. 11. The health-mon... | Abbott Diabetes Care Inc. | G01N033/48 G01N031/00 G06G007/48 G06G007/58 G06F019/00 A61B005/00 A61B005/145 G01N021/84
US08388530 | B2 | 2013-03-05 | Personalized monitoring and healthcare information management using physiological basis functions | 2010-09-19 | Analysis of individual's serial changes, also referred to as the physiological, pathophysiological, medical or health dynamics, is the backbone of medical diagnosis, monitoring and patient healthcare management. However, such an analysis is complicated by enormous intra-individual and inter-individual variability. To address this problem, a novel serial-analysis method and system based on the concept of personalized basis functions (PBFs) is disclosed. Due to more accurate reference information provided by the PBFs, individual's changes associated with specific physiological activity or a sequence, transition or combination of activities (for example, a transition from sleep to wakefulness and transition from rest to exercise) can be monitored more accurately. Hence, subtle but clinically important changes can be detected earlier than using other methods. A library of individual's PBFs and their transition probabilities (which can be described by Hidden Markov Models) can completely describe individual's physiological dynamics. The system can be adapted for healthcare information management, diagnosis, medical decision support, treatment and side-effect control. It can also be adapted for guiding health, fitness and wellness training, subject identification and more efficient management of clinical trials. | 1. A method useful in healthcare information management comprising: collecting at least one primary element as a snapshot present at the time of recording of health data using at least one collection method selected from one-time, periodic, quasi-periodic and continuous monitoring, and electronically comparing said at least one primary element with at least one reference value selected from physiological basis functions representing physiological and pathophysiological states, and transitions between said physiological and pathophysiological states to detect changes in said at least one primary element and thereby identify any abnormal or unstable primary element (a first-level, low-resolution analysis); andanalyzing serial changes in said at least one primary element of health data using a dynamic serial analysis and processing unit employing at least one of the following methods selected from mathematical decomposition, mathematical modeling, computer modeling, signal processing, time-series analysis, statistical analysis, and methods of artificial intelligence for assessing changes in serial data, orthogonal decomposition, non-orthogonal decomposition (independent component analysis), multidimensional scaling based on non-metric distances and mapping techniques, non-orthogonal linear mappings, nonlinear mappings and other methods, that make use of projection, re-scaling (change of variables), methods from the theories of singularities, bifurcations, catastrophes, and dynamical systems, and other statistical estimators, linear and nonlinear correlation, analysis of variance, cluster analysis, factor analysis, canonical analysis, regression and discriminant function analyses, and probabilistic methods, Bayesian probability, Bayesian network, Euclidean, Markov model, hidden Markov model, and Mahalanobis distance, pattern recognition, fuzzy logic, neural networks, expert systems, and hybrid artificial intelligence systems to provide detailed characterization of serial changes in any abnormal or unstable primary element (a second-level, higher resolution serial analysis). 2. A method as set forth in claim 1 in which said physiological basis functions are adapted for representing a sequence of at least two physiological states with a transition between them represented by at least one of Markov and hidden Markov model with a transition probability matrix. 3. A method as set forth in claim 1 in which said physiological and pathophysiological states are selected from sleep, REM-sleep, non-REM sleep, wakefulness, physical activity, food intake, mental activity, exercise, physical exercise, mental stress and different body positions, including standing, sitting and supine position. 4. A method as set forth in claim 1 in which said physiological and pathophysiological states are adjusted for seasonal, circadian, geographical, environmental factors, age, sex, medical history, medications, genetic, biochemical, biophysical, physiological, occupational f... | - | A61N001/00
US08768520 | B2 | 2014-07-01 | Systems and methods for controlling a bedroom environment and for providing sleep data | 2008-11-14 | A system for controlling a bedroom environment includes an environmental data collector configured to collect environmental data relating to the bedroom environment; a sleep data collector configured to collect sleep data relating to a person's state of sleep; an analysis unit configured to analyze the collected environmental data and the collected sleep data and to determine an adjustment of the bedroom environment that promotes sleep of the person; and a controller configured to effect the adjustment of the bedroom environment. A method for controlling a bedroom environment includes collecting environmental data relating to the bedroom environment; collecting sleep data relating to a person's state of sleep; analyzing the collected environmental data and the collected sleep data; determining an adjustment to the bedroom environment that promotes sleep; and communicating the adjustment to a device that effects the bedroom environment. | 1. A system for controlling a bedroom environment, the system comprising: an environmental data collector configured to collect environmental data relating to the bedroom environment;a sleep data collector configured to collect sleep data relating to a person's state of sleep;an analysis unit configured to analyze the collected environmental data and the collected sleep data and to automatically determine an adjustment of the bedroom environment, including an adjustment to a firmness of a mattress on which the person sleeps, that promotes sleep of the person; anda controller configured to effect the automatically determined adjustment of the bedroom environment. 2. The system according to claim 1, wherein the controller is configured to control at least one of a cooling system and a heating system. 3. The system according to claim 1, wherein the controller is configured to control a humidity adjustment unit. 4. The system according to claim 1, wherein the controller is configured to control an audio system. 5. The system according to claim 1, wherein the controller is configured to control a bedding support member temperature adjustment unit. 6. The system according to claim 1, wherein the controller is configured to control a bedding support member humidity adjustment unit. 7. The system according to claim 1, wherein the controller is configured to control a lighting unit. 8. The system according to claim 1, wherein the sleep data collector is configured to collect sleep data from a sensor positioned in or on a mattress. 9. The system according to claim 1, wherein the sleep data collector is configured to collect sleep data from a sensor positioned in or on a head support member. 10. The system according to claim 1, wherein the sleep data comprises data relating to at least one of the person's weight, the person's weight distribution, the person's body position, the person's body movement, the person's breathing rate, the person's heart rate, the person's body temperature and a near body humidity of the person. 11. The system according to claim 1, wherein the environmental data of the bedroom comprises at least one of a temperature of the bedroom, a relative humidity of the bedroom, a lighting intensity of the bedroom, and a sound level of the bedroom. 12. A method for controlling a bedroom environment, the method comprising: collecting environmental data relating to the bedroom environment;collecting sleep data relating to a person's state of sleep;analyzing, using at least one processor, the collected environmental data and the collected sleep data;automatically determining, using the at least one processor, an adjustment to the bedroom environment, including an adjustment to a firmness of a mattress on which the person sleeps, that promotes sleep of the person; andcommunicating, using the at least one processor, the automatically determined adjustment to a device that effects the bedroom environment. 13. The method according to cl... | Kingsdown, Inc. | G06F019/00
US09131876 | B2 | 2015-09-15 | Portable sound source playing apparatus for testing hearing ability and method of testing hearing ability using the apparatus | 2010-08-10 | A method of testing a hearing ability of a user using a portable sound source playing apparatus includes; outputting a signal indicating at least one phoneme, obtaining a response of a user regarding a degree of audibility of the output signal, and testing the hearing ability of a user based on hearing characteristics of the user showing the degree of audibility of a frequency band comprising a frequency of the at least one phoneme based on the obtained response. | 1. A method of testing a hearing ability of a user using a portable sound source playing apparatus, the method comprising: extracting at least one phoneme from each of a plurality of groups comprising at least one phoneme each, wherein the plurality of groups are classified based on frequencies of phonemes;outputting a signal corresponding to the at least one extracted phoneme;obtaining a response of the user regarding a degree of audibility of the output signal;determining the degree of audibility of a predetermined frequency band, to which each of the at least one extracted phoneme corresponds, based on the response of the user; andmeasuring hearing characteristics of the user showing the degree of audibility for each frequency band based on a result of determination. 2. The method of claim 1, wherein the determining comprises: determining the degree of audibility of a plurality of predetermined frequency bands, based on the response of the user regarding the at least one phoneme included in the plurality of frequency bands corresponding to each of the plurality of groups comprising at least one phoneme each, andwherein the measuring comprises:measuring the hearing characteristics of the user showing the degree of audibility of each of the plurality of predetermined frequency bands. 3. The method of claim 1, wherein the outputting the signal comprises: outputting a signal representing a sentence to which at least one method of amplifying a sound source is applied,wherein the obtaining comprises:obtaining a preference regarding the signal from the user, andthe method further comprises determining a method of amplifying desired by the user, among the at least one method of amplifying a sound source. 4. The method of claim 3, wherein, in the at least one method of amplifying a sound source, a time when compression of the sound source starts is varied. 5. The method of claim 1, further comprising obtaining volume data from the user, wherein the hearing ability of the user is tested using a signal indicating the at least one phoneme at a volume set based on the obtained volume data. 6. The method of claim 1, further comprising displaying a result of testing the hearing of the user according to the hearing characteristics of the user. 7. The method of claim 1, further comprising: correcting a signal for playing a sound source by applying the hearing characteristics of the user to the signal for playing a sound source, and wherein the outputting includes outputting the corrected signal. 8. The method of claim 7, wherein, the correcting comprises: applying an amplification gain regarding a frequency of a sound source according to the hearing characteristics of the user and an amplification method desired by the user to correct the signal. 9. The method of claim 1, wherein obtaining the response of the user comprises at least one of recognizing a voice of the user to obtain the response, and obtaining an input signal, which corresponds to the re... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/00 A61B005/12 H04R025/00
US09031630 | B2 | 2015-05-12 | Analyte sensors and methods of use | 2010-11-01 | An analyte sensor system including a substrate, a first electrode disposed on a first surface of the substrate, a second electrode disposed on a second surface of the substrate, a third electrode provided in electrical contact with at least one of the first or second electrodes, where at least a portion of the first electrode and the second electrode are subcutaneously positioned in a patient, and where the third electrode is substantially entirely positioned external to the patient, and corresponding methods are provided. | 1. An analyte monitoring system, comprising: an analyte sensor comprising a substrate, at least one electrode disposed on a first side of the substrate and at least one electrode disposed on a second side of the substrate opposite the first side, wherein the electrodes are in fluid contact with an interstitial fluid of a user when at least a portion of the substrate is transcutaneously positioned within the user; andan on body unit operatively coupled to the analyte sensor and configured to receive one or more signals from the sensor corresponding to an analyte level of the user, wherein the on body unit is electrically connectable to the first side of the substrate by a first electrical connection and is electrically connectable to the second side of the substrate by a second electrical connection;wherein an external surface of the on body unit comprises at least one conductive point in electrical contact with an external electrode, wherein the at least one conductive point physically contacts the user's skin when the on body unit is operatively coupled to the transcutaneously positioned sensor. 2. The system of claim 1, wherein the external electrode is a counter electrode. 3. The system of claim 2, wherein at least one of the electrodes disposed on the substrate is a working electrode and another of the electrodes disposed on the substrate is a reference electrode. 4. The system of claim 1, wherein the external electrode is positionable external to the user. 5. The system of claim 1, wherein the external electrode is in electrical contact with a temperature probe. 6. The system of claim 1, wherein the external surface of the on body unit comprises a plurality of conductive points in electrical contact with the external electrode. 7. The system of claim 6, wherein the plurality of conductive points are disposed in a predetermined pattern. 8. The system of claim 1, further comprising an alarm unit configured to alert the user when the external electrode is no longer in electrical contact with the skin of the user. 9. The system of claim 8, wherein the alarm unit is integrated with the on body unit. 10. An analyte monitoring system, comprising: an analyte sensor comprising a substrate, at least one electrode disposed on a first side of the substrate and at least one electrode disposed on a second side of the substrate opposite the first side, wherein the electrodes are in fluid contact with an interstitial fluid of a user when at least a portion of the substrate is transcutaneously positioned within the user, the analyte sensor further comprising a conductive path through the substrate from the at least one electrode disposed on the first side to the second side; andan on body unit operatively coupled to the analyte sensor and configured to receive one or more signals from the sensor corresponding to an analyte level of the user, wherein the on body unit is electrically connectable to only the second side of the substrate;wherein an exte... | Abbott Diabetes Care Inc. | A61B005/05 A61B005/1473 A61B005/145 A61B005/00 A61B005/1486
US09492106 | B2 | 2016-11-15 | Respiratory sensor | 2010-02-24 | A sensor for measuring a respiration is disclosed herein. The sensor includes at least one housing (7) having a first cavity (12) with a first port (11) allowing a respiration gas flow, and a second cavity (22) with a second port (21) also allowing a respiration gas flow. The sensor also includes at least one breathing detector (51) for acquiring a signal indicative of the respiration gas flowing through the first cavity and the second cavity. The at least one housing is equipped with at least one additional port (31) for removing the respiration gas flow coming from the first cavity and the second cavity, which additional port being separate from the first port and the second port. | 1. A sensor for measuring a respiration comprising: at least one housing defining a first cavity with a first port configured for receiving a respiration gas flow exiting a patient's nose, and a second cavity with a second port, the second port configured for receiving a respiration gas flow exiting the patient's mouth;at least one additional port located between the first port and the second port for receiving the respiration gas flow exiting from said first cavity and said second cavity and opening to ambient air, said additional port being separate from said first port and said second port;a flow guide positioned in the middle of the additional port and between the first cavity and the second cavity for preventing respiration gas flow between the first cavity and the second cavity; andat least one breathing detector disposed proximate said additional port for acquiring a signal indicative of a thermal component of the respiration gas exiting, at least one of said first and second cavity. 2. The sensor according to claim 1, wherein a cross-sectional area of said at least one additional port is less than 10% of a combined cross-sectional area of said first port and said second port, more specifically less than 20% or even more specifically less than 50% of the combined cross-sectional area of said first port and said second port. 3. The sensor according to claim 1, wherein said first cavity and said second cavity are adjacent one another. 4. The sensor according to claim 1, wherein the flow guide is configured for guiding the respiration gas out from said first cavity and said second cavity through said at least one additional port. 5. The sensor according to claim 4, wherein said flow guide is configured to turn the respiration gas flow direction of said first cavity and said second cavity in a range between 30-90 degrees. 6. The sensor according to claim 1, wherein said breathing detector acquiring the signal indicative of the respiration gas flow is configured to be located at said at least one additional port. 7. The sensor according to claim 1, wherein said at least one additional port comprises two additional ports, one for said first cavity and another for said second cavity. 8. The sensor according to claim 7, wherein the breathing detector comprises two breathing detectors, one each being configured for sensing a thermal component of the respiration gas exiting the first and second cavities, respectively. 9. The sensor according to claim 1 further comprising a sticker for placing said housing on a skin of a subject below a nose but above an upper lip. 10. The sensor according to claim 1 wherein the breathing detector is positionable for acquiring a signal indicative of a thermal component of the respiration gas exiting only the first cavity. 11. The sensor according to claim 1 wherein the breathing detector is positionable for acquiring a signal indicative of a thermal component of the respiration gas exiting only the second c... | GENERAL ELECTRIC COMPANY | A61B005/08 A61B005/087 A61B005/00 A61B005/097
US09289133 | B2 | 2016-03-22 | Methods and apparatus for determining cardiac output | 2010-10-13 | The present invention provides methods and apparatus for determining a dynamical property of the systemic or pulmonary arterial tree using long time scale information, i.e., information obtained from measurements over time scales greater than a single cardiac cycle. In one aspect, the invention provides a method and apparatus for monitoring cardiac output (CO) from a single blood pressure signal measurement obtained at any site in the systemic or pulmonary arterial tree or from any related measurement including, for example, fingertip photoplethysmography.;According to the method the time constant of the arterial tree, defined to be the product of the total peripheral resistance (TPR) and the nearly constant arterial compliance, is determined by analyzing the long time scale variations (greater than a single cardiac cycle) in any of these blood pressure signals. Then, according to Ohm's law, a value proportional to CO may be determined from the ratio of the blood pressure signal to the estimated time constant. The proportional CO values derived from this method may be calibrated to absolute CO, if desired, with a single, absolute measure of CO (e.g., thermodilution). The present invention may be applied to invasive radial arterial blood pressure or pulmonary arterial blood pressure signals which are routinely measured in intensive care units and surgical suites or to noninvasively measured peripheral arterial blood pressure signals or related noninvasively measured signals in order to facilitate the clinical monitoring of CO as well as TPR. | 1. A method for determining cardiac output to within a constant scale factor, comprising: measuring a physiological signal of a patient using a sensor, the physiological signal indicative of cardiovascular system activity and measured over a plurality of cardiac cycles;determining, from the physiological signal, beat-to-beat variability of the physiological signal across multiple cardiac beats;determining cardiac output to within a constant scale factor using the beat-to-beat variability information extracted from the physiological signal; andpresenting the cardiac output on a display of a computing device; wherein determining beat-to-beat variability includes constructing a cardiac contractions timing signal from the physiological signal by forming an impulse train in which each impulse is located at the onset of upstroke of each pulse;determining an impulse response, which when convolved with the cardiac contractions timing signal, fits the physiological signal;determining a time constant by fitting an exponential function to a tail end of the impulse response; anddetermining cardiac output to within a constant scale factor using the time constant. 2. The method of claim 1 further comprises measuring the physiological signal invasively or non-invasively at any site in the systemic or pulmonary arterial tree. 3. The method of claim 1 wherein the physiological signal is arterial blood pressure. 4. The method of claim 1 wherein the physiological signal is pulmonary arterial blood pressure. 5. The method of claim 1 wherein the physiological signal is a signal related to arterial blood pressure. 6. The method of claim 1 wherein determining beat-to-beat variability further comprises estimating a function that represents the response of the physiological signal to a single cardiac contraction over a time period greater than a single cardiac cycle. 7. The method of claim 6 further comprises determining a characteristic time of the function and determining cardiac output to within a constant scale factor by dividing the magnitude of the physiological signal by the characteristic time. 8. The method of claim 1 further comprises determining vascular resistance to within a constant scale factor using the beat-to-beat variability information. 9. A method for determining a dynamical property of the systemic or pulmonary arterial tree, comprising: measuring a physiological signal of a patient using a sensor, the physiological signal indicative of cardiovascular system activity over a plurality of cardiac cycles;determining the timing of cardiac contractions in relation to each other, which occur in a sequence of cardiac contractions, where the determination is made from the physiological signal;determining, from the timing of cardiac contractions and the physiological signal, a mathematical relationship between the timing of cardiac contractions in the sequence and the physiological signal occurring over the plurality of cardiac cycles;determining a d... |  | A61B005/021 A61B005/02 A61B005/0402 A61B005/029 A61B005/0225 A61B005/0456 A61B005/00
US20110009764 | A1 | 2011-01-13 | DEVICES, SYSTEMS, AND METHODS FOR AIDING IN THE DETECTION OF A PHYSIOLOGICAL ABNORMALITY | 2010-07-13 | The present invention comprises a method for excluding the presence of a pulmonary embolism using a combination of a blood test for D-dimer and a carboximetry test measuring the ratio of exhaled CO2 to O2. If the measured D-Dimer value is at or above a threshold indicative of concern and the carboximetry value is equal to or greater than a pre-determined carboximetry ratio threshold, a clinically significant pulmonary embolism, or a pulmonary thrombus above a certain effective size, is unlikely to be present. If the measured D-Dimer value is at or above a threshold indicative of concern and the respiratory analysis yields a carboximetry ratio less than the carboximetry ratio threshold, test results are inconclusive and additional testing may be required to determine whether a clinically significant pulmonary embolism, or a pulmonary thrombus above a specified size is present. The present invention also comprises a method for diagnosing or excluding the presence of respiratory dysfunction in an individual by using carboximetry values corresponding to an end-tidal volume of an exhaled breath that is a pre-determined multiple of the individual's estimated pulmonary dead space. In one aspect, a linear regression function is used to represent the CO2/O2 ratio derived from the measured CO2 and O2 values in an exhaled breath, and the function is used to determine the CO2/O2 ratio corresponding to the pre-determined end-tidal volume. | 1. A method to aid in the diagnosis of respiratory dysfunction in an individual, comprising: a) estimating the pulmonary dead space in the individual;b) measuring the volume of air exhaled in each of a series of breaths by the individual;c) measuring a concentration of carbon dioxide and a concentration of oxygen in each of the exhaled breaths;d) calculating a final carbox ratio comprising an exhaled carbon dioxide concentration divided by an exhaled oxygen concentration at a pre-determined volume of each exhaled breath, the pre-determined volume being a multiple of the individual's pulmonary dead space; ande) comparing an average or median value of the final carbox ratios from the series of breaths to a known value; whereinthe presence of respiratory dysfunction is indicated if the average or median value is below the known value, and the absence of respiratory dysfunction is indicated if the average or median value is at or above the known value. 2. The method of claim 1, wherein the pre-determined volume of exhaled breath is within the range of about 2 times the pulmonary dead space to about 5 times the pulmonary dead space. 3. The method of claim 1, wherein the pre-determined volume of exhaled breath is within the range of about 2.5 times the pulmonary dead space to about 4 times the pulmonary dead space. 4. The method of claim 1, wherein the pre-determined volume of exhaled breath is about 3 times the pulmonary dead space. 5. The method of claim 1, wherein calculating the final carbox ratio for an exhaled breath comprises dividing the measured carbon dioxide concentration by the measured oxygen concentration corresponding to each of a series of exhaled breath volumes to generate a series of measurement-based carbox ratios, fitting an nth order equation to the series of measurement-based carbox ratios, and selecting the final carbox ratio for the pre-determined volume of exhaled breath from the fitted equation, whereinn is equal to or greater than 1. 6. The method of claim 5, wherein the nth order equation is a first order equation. 7. The method of claim 1, wherein calculating the final carbox ratio for an exhaled breath comprises fitting a first nth order equation to a series of measured carbon dioxide concentrations in each of a series of exhaled breaths, fitting a second nth order equation to a series of measured oxygen concentrations in each of the series of exhaled breaths, selecting a fitted value for carbon dioxide and a fitted value for oxygen corresponding to the pre-determined volume of exhaled breath, and dividing the fitted value for carbon dioxide by the fitted value for oxygen to yield a derived final carbox ratio for each exhaled breath, wherein n for the first equation is equal to or greater than 1, and n for the second equation is equal to or greater than 1. 8. The method of claim 7, wherein the first nth order equation and the second nth order equation are both first order equations. | DEKA Product Limited Partnership | A61B005/08
US07883471 | B2 | 2011-02-08 | Minimally invasive determination of collateral ventilation in lungs | 2005-12-07 | Minimally invasive methods, systems and devices are provided for qualitatively and quantitatively assessing collateral ventilation in the lungs. In particular, collateral ventilation of a target compartment within a lung of a patient is assessed by advancement of a catheter through the tracheobronchial tree to a feeding bronchus of the target compartment. The feeding bronchus is occluded by the catheter and a variety of measurements are taken with the use of the catheter in a manner which is of low risk to the patient. Examples of such measurements include but are not limited to flow rate, volume and pressure. These measurements are used to determine the presence of collateral ventilation and to quantify such collateral ventilation. | 1. A method for diagnosing collateral ventilation between lung compartments in a patient, said method comprising: isolating a target lung compartment from at least one adjacent lung compartment;allowing the patient to breathe air free from introduced markers;detecting, using a flow-measuring device, air flow or accumulation from the isolated lung compartment over time;calculating, using a computer, a value of collateral resistance; anddiagnosing collateral ventilation between the target lung compartment and at least one adjacent lung compartment based on the value of collateral resistance calculated by the computer. 2. A method as in claim 1, further comprising determining a treatment plan based on diagnosing the collateral ventilation. 3. A method as in claim 1, wherein isolating comprises introducing a catheter transtracheally to a main bronchus feeding into the target lung compartment and deploying an occlusion member on the catheter to isolate the target lung compartment. 4. A method as in claim 3, wherein detecting comprises measuring air flow through a lumen in the catheter while the patient exhales, and wherein diagnosing comprises concluding that air entered the isolated compartment via one or more collateral passages while the patient inhaled if the collateral resistance is less than a predetermined value. 5. A method as in claim 3, wherein detecting comprises accumulating air from the isolated compartment air from the isolated compartment through the catheter over a number of successive breathing cycles, wherein a continuous increase in accumulated air volume indicates collateral flow into the isolated compartment. 6. A method as in claim 3, wherein isolating comprises occluding air flow into the target lung compartment while permitting air flow from the target lung compartment through the catheter to the atmosphere. 7. A method as in claim 6, wherein the catheter comprises a one-way valve which blocks air flow into the target lung compartment and permits air flow from the compartment to the atmosphere. 8. A method as in claim 6, wherein detecting comprises generating at least one measurement of at least one characteristic of exhaled air from the target lung compartment. 9. A method as in claim 8, wherein the at least once characteristic includes volume. 10. A method as in claim 8, wherein the at least one characteristic includes volumetric flow rate. 11. A method as in claim 8, wherein the at least one characteristic includes pressure. 12. A method of diagnosing collateral ventilation between regions of a lung, comprising: positioning a catheter transtracheally into a bronchial passageway that provides direct fluid flow into a target compartment of the lung;deploying an occlusion member on the catheter to seal the catheter in the bronchial passageway; anddetecting, using a flow-measuring device, collateral ventilation between the target compartment of the lung and at least one adjacent lung compartment by measuring flow of a... | Pulmonx Corporation | A61B005/08
US07894870 | B1 | 2011-02-22 | Hermetic implantable sensor | 2004-08-23 | At least one conductor is formed at a preselected location on a substrate made of a first insulating material having a high temperature resistance. The conductor is made from a solidified electrically conductive thick film material. A coating made of a second insulating material is formed over the substrate to hermetically seal at least a portion of the conductor. An exposed distal region of the conductor provides a detection electrode. The conductor has a reduced porosity that inhibits migration of fluid or constituents thereof through the conductor. | 1. A sensor suitable for implanting in living tissue for monitoring the level of a biological agent in a tissue fluid, comprising: a substrate made of a first biocompatible insulating material having a high temperature resistance selected from the group consisting of ceramic and glass/ceramic combinations;at least one conductor disposed over a preselected external surface location on the substrate and made from an electrically conductive thick film material containing particles of a metal selected from the group consisting of platinum, palladium, rhodium, ruthenium, iridium and osmium;a coating made of a second biocompatible insulating material selected from the group consisting of glass, ceramic, glass/ceramic combinations, and polymers, and formed over at least a portion of the at least one conductor to hermetically seal at least a portion of the conductor leaving at least one exposed region thereof to provide a detection electrode; andthe at least one conductor consisting essentially of metal particles having been sintered by heating to a temperature within a range that extends from below a melting point of the metal particles to approximately three hundred degrees C. below the melting point of the metal particles, and having a reduced porosity that inhibits migration of the bodily fluid or any constituents thereof through the conductor; whereby instances of sensor failure and/or inaccuracy during implanted operation are minimized, wherein the at least one conductor has a helium leak value below about 1E-8 (atm cc/s). 2. The sensor of claim 1 wherein the thick film material contains platinum particles and wherein the thick film material is heated to a peak temperature between about 1500 degrees C. and about 1700 degrees C. 3. The sensor of claim 1 wherein the thick film material is heated to the peak temperature for a dwell time up to about sixty minutes. 4. An electrochemical sensor, comprising: a substrate made of a first insulating material having a high temperature resistance;at least one conductor disposed over a preselected external surface location on the substrate and made from an electrically conductive thick film material containing particles of a metal selected from the group consisting of platinum, palladium, rhodium, ruthenium, iridium and osmium;a coating made of a second insulating material formed over at least a portion of the at least one conductor to hermetically seal at least a portion of the at least one conductor leaving at least one exposed region to provide a detection electrode; andthe at least one conductor consisting essentially of metal particles having been sintered by heating to a temperature within a range that extends from below a melting point of the metal particles to approximately three hundred degrees C. below the melting point of the metal particles, and having a reduced porosity that inhibits migration of fluid or gaseous diffusion through the at least one conductor, wherein the at least one conductor ha... | Glysens, Incorporated | A61B005/00
US20110021941 | A1 | 2011-01-27 | SYSTEMS AND METHODS FOR RESPIRATION MONITORING | 2009-07-23 | According to embodiments, techniques for determining respiratory parameters are disclosed. More suitable probe locations for determining respiratory parameters, such as respiration rate and respiratory effort, may be identified. The most suitable probe location may be selected for probe placement. A scalogram may be generated from the detected signal at the more suitable location, resulting in an enhanced breathing band for determining respiratory parameters. Flexible probes that allow for a patient's natural movement due to respiration may also be used to enhance the breathing components of the detected signal. From the enhanced signal, more accurate and reliable respiratory parameters may be determined. | 1. A probe used in determining a respiratory parameter of a patient, the probe comprising: at least one energy emitter;at least one energy detector; andat least one flexible member connecting at least one of the at least one energy emitter and at least one of the at least one energy detector, wherein the at least one energy detector connected to the at least one energy emitter by the flexible member is free to move relative to the at least one energy emitter when positioned on the patient. 2. The probe of claim 1 wherein the at least one energy emitter comprises at least one light emitter. 3. The probe of claim 1 wherein the at least one energy detector comprises at least one photodetector. 4. The probe of claim 1 further comprising a wireless transmission device coupled to the at least one energy detector. 5. The probe of claim 1 wherein the flexible member comprises material selected from the group consisting of an elastoplastic, a rubber, a synthetic polymer, a coil, a spring, and a wire. 6. The probe of claim 1 further comprising a housing for the at least one energy detector, wherein at least one of the at least one energy emitter is rigidly coupled to the housing. 7. The probe of claim 6 wherein the at least one energy detector comprises a photodetector and the at least one energy emitter rigidly coupled to the housing comprises a red light emitter. 8. The probe of claim 6 wherein the at least one energy detector comprises a photodetector and the at least one energy emitter rigidly coupled to the housing comprises a infrared light emitter. 9. The probe of claim 1 wherein the flexible member comprises a hinge or pivot. 10. The probe of claim 9 wherein the hinge or pivot restrains motion of the at least one energy emitter relative to the at least one energy detector in one or more planes. 11. The probe of claim 1 wherein the at least one energy detector comprises at least two energy detectors in a flexible array covering a local area of the patient. 12. A method for determining at least one respiratory parameter of a patient, comprising: positioning at least one energy emitter on the patient;positioning, at a location exhibiting movement due to the respiration of the patient, at least one energy detector on the patient, wherein the at least one energy emitter and the at least one energy detector are connected by at least one flexible member;receiving a signal from the at least one energy detector; anddetermining at least one respiratory parameter from the received signal. 13. The method of claim 12 wherein the at least one respiratory parameter is selected from the group consisting of respiration rate and respiratory effort. 14. The method of claim 12 wherein positioning at least one energy emitter comprises positioning at least one light emitter. 15. The method of claim 12 wherein positioning at least one energy detector comprises positioning at least one photodetector. 16. The method of claim 12 wherein determining at least o... | Nellcor Puritan Bennett Ireland | A61B005/08
US20110009720 | A1 | 2011-01-13 | Continuous whole blood glucose monitor | 2007-11-02 | A portable continuous whole blood glucose monitor comprising, a mid-infrared quantum cascade laser and driver in optical communication with a transmission cell and a photo-conductive detector and pre-amplifier. The monitor further comprises a peristaltic pump connected to a single lumen catheter peripherally inserted into a patient's vein. The single lumen catheter, in combination with the peristaltic pump, is operable to automatically withdraw a fixed and metered amount of whole blood from a patient, then a tube delivers a fixed and metered amount of the saline/surfactant supply to the whole blood. Methods of enhancing measurement sensitivity are also provided. | 1. A system for measuring an analyte in a complex matrix, comprising: a mid-infrared quantum cascade laser emitting at least one wavelength; a photo-detector; a means for exposing a sample of the complex matrix to the laser and photo detector; and a means for enhancing measurement sensitivity. 2. The system in claim 1, wherein the analyte is glucose. 3. The system in claim 2, wherein the matrix is selected from the group consisting of whole blood, plasma, serum, and ultrafiltrate. 4. The system in claim 1, wherein the system further comprises a thermo-electric cooler; the laser and the photo-detector operate at room temperature being cooled by the thermoelectric cooler. 5. The system in claim 4, wherein: the quantum cascade laser contains an external trigger that emits many mid-infrared electromagnetic radiation shots; and the means for enhancing measurement comprises electronics; including a time gate that coincides with the trigger; a gated integrator designed to recover fast, repetitive, analog signals; a boxcar averager with an active baseline subtraction module; the photo detector converts the many radiation signals to analog signals and feeds the analog signals to the gated integrator that amplifies and integrates the analog signals present when the gate is open and ignores noise and interference during other times; the boxcar averager averages the output of the gated integrator, and the active baseline subtraction module cancels baseline drift. 6. The system in claim 5, wherein the means for exposing the sample to the photo detector, comprises a fluidic system comprising, a transmission cell, a peristaltic pump, and a catheter, and a saline supply for rinsing the transmission cell. 7. The system in claim 6, wherein the transmission cell has a pair of mid-infrared transmissive windows spaced apart in the path of the mid-infrared electromagnetic radiation shots and the photo detector; the peristaltic pump delivers the sample from the catheter to the transmission cell; the spacing between the windows defines the optimal path length that enhances measurement sensitivity. 8. The system in claim 7, wherein the transmissive windows are preferably Zinc Selenide, but not excluding Zinc Sulfide, Silicon, and Polypropylene. 9. The system in claim 8, wherein one end of the catheter is inserted in a peripheral vein of a patient, whereas the other catheter end is connected to the peristaltic pump. 10. The system in claim 9, wherein the analyte is glucose, and the catheter dimensions are also sized to minimize blood consumption and lag time between whole blood withdrawal and glucose measurements. 11. The system in claim 10, wherein a method for preventing clogging and non-homogeneity of the matrix sample, and for preventing air bubbles in the transmission cell, comprises the step of using known methods in the industry. 12. The system in claim 10, wherein the fluidic system contains a non-ionic surfactant and saline supply, a catheter, a tube, ... | - | A61B005/00 A61B005/15 A61M001/38 A61M001/34
US07885441 | B2 | 2011-02-08 | Systems and methods for implant virtual review | 2006-10-11 | Certain embodiments of the present invention provide systems and methods for virtual implant review. Certain embodiments provide a method for virtual implant review. The method includes registering a two-dimensional image and a three-dimensional image. The method includes detecting an implant or instrument in the two-dimensional image. The method also includes performing an initial implant registration of the implant or instrument with respect to the three dimensional image based on navigation information and refining the initial implant registration based on image data analysis to generate a refined implant registration. Additionally, the method includes displaying a representation of the implant or instrument with respect to the three-dimensional image based on the refined image registration and the refined implant registration. | 1. A method for virtual implant review, said method comprising: registering a two-dimensional image and a three-dimensional image;detecting an implant or instrument in said two-dimensional image;performing an initial implant registration of said implant or instrument detected in said two-dimensional image with respect to said three-dimensional image based on navigation information;refining said initial implant registration based on image data analysis of said two-dimensional image to generate a refined implant registration;displaying a representation of said implant or instrument with respect to said three-dimensional image based on said refined implant registration;computing a difference between said refined implant registration and a planned implant registration; and,adjusting a navigation system based on said difference between said refined implant registration and said planned implant registration. 2. The method of claim 1, wherein said step of registering said two-dimensional image and said three-dimensional image comprises at least one of contour matching, landmark matching, surface-volume matching and fiducial registration. 3. The method of claim 1, further comprising storing implant placement and distance measurement information for subsequent implant placement. 4. The method of claim 1, wherein said representation of said implant or instrument comprises a library representation of said implant or instrument. 5. The method of claim 1, wherein said implant or instrument is graphically rendered and overlaid on said three-dimensional image data set with trajectory information and distance information. 6. The method of claim 1, wherein said step of refining said initial image registration further comprises refining said initial image registration using image data from a plurality of two-dimensional images. 7. The method of claim 1, wherein said step of refining said initial implant registration further comprises correlating navigation information with imaging information from a plurality of two-dimensional images. 8. A user interface system for virtual implant review, said system comprising: a processor configured to determine a registration of a two-dimensional image of an object and a three-dimensional image of said object, said processor configured to determine an initial implant registration of an implant or instrument based on tracking data and to refine said initial implant registration based on said two-dimensional image of said implant or instrument within said three-dimensional image,wherein said processor further computes a difference between said refined implant registration and a planned implant registration;a navigation system obtaining tracking data relative to said image and said implant or instrument, said navigation system capable of being adjusted based on said difference between said refined implant registration and said planned implant registration; anda display configured to display a three-dimensional image includ... | General Electric Company | G06K009/00 A61B005/05
US08332016 | B2 | 2012-12-11 | Non-linear projections of 3-D medical imaging data | 2009-08-04 | The present invention improves projection displays of volume data. Using the Minimum Intensity Projection (MinIP), fluid filled regions or other regions of hyporeflective tissue are displayed. By limiting the projection to partial volumes within the volume, differences in the scattering intensity within specific regions are isolated. In this way, hyperreflectivity of weakly scattering tissue can be assessed. | 1. A method of generating an image of an anatomical region within an eye, said eye being examined by an optical coherence tomography (OCT) system, said OCT system including a light source for generating a beam of light, a sample arm, a reference arm, and a detector for measuring light combined from the sample and reference arms, said method comprising the steps of: a. scanning the beam of light over a region of the eye via the sample arm;b. combining the light from the sample and reference arms;c. measuring the combined light to generate three dimensional volume image data of the eye;d. identifying a first surface within the image data;e. identifying a subvolume of the image data, said subvolume having less than all the generated image data, said subvolume being identified based upon the identification of the first surface, said subvolume being either bounded by or including the first surface, said subvolume including the anatomical region;f. selecting a set of image points within the subvolume, said image points being selected by evaluating a plurality of ray projections extending through the subvolume and identifying one image point in each projection, wherein each of the identified image points has a common intensity attribute;g. generating an image based on information associated with the selected set of image points; andh. storing or displaying the image. 2. A method as recited in claim 1, wherein the common intensity attribute corresponds to the minimum intensity of the associated ray projection. 3. A method as recited in claim 1, wherein the common intensity attribute corresponds to the maximum intensity of the associated ray projection. 4. A method as recited in claim 1, wherein the common intensity attribute corresponds to the median intensity of the associated ray projection. 5. A method as recited in claim 1, wherein the generated image is based directly on the intensity of each of the selected set of image points. 6. A method as recited in claim 1, wherein the image generation is based on location information associated with each of the selected set of image points within the associated ray projection. 7. A method as recited in claim 6, wherein the generated image is a height map of the minimum intensity locations relative to a reference surface. 8. A method as recited in claim 7, wherein the height map is displayed as a surface within a 3-D volume. 9. A method as recited in claim 1, further comprising the step of identifying a second surface within the volume image of the sample and wherein the subvolume is identified based on the first and second surfaces. 10. A method as recited in claim 1, further comprising the step of identifying a second surface based on the locations of the selected set of image points. 11. A method as recited in claim 10, wherein the second surface is a smooth surface fit to the selected set of image points. 12. A method as recited in claim 1, wherein the generated image includes at least one regio... | Carl Zeiss Meditec, Inc. | A61B006/00
US07894871 | B2 | 2011-02-22 | Filtering method for surface modeling | 2006-12-29 | A system for mapping a tissue surface includes a probe for mapping a tissue surface, a localization system to measure a location data point indicative of the probe's location, a memory in which to store the location data point, a servo mechanism to move the probe along at least a portion of the tissue surface, a controller to move the probe to a plurality of locations and to record in the memory a plurality of location data points, and a contact-sensing processor to analyze the plurality of location data points and to identify a subset thereof on the tissue surface. A modeling processor generates a model of the tissue surface using the subset of location data points. The contact-sensing processor utilizes probe velocity, or a rate of change in the distance moved by the probe, to determine contact between the probe and the tissue surface. | 1. A method of mapping a tissue surface using a probe, comprising the steps of: placing a distal end of the probe in contact with the tissue surface;moving the probe along at least a portion of the tissue surface;periodically measuring position information for the distal end of the probe to create a measured location data point;determining a velocity of the distal end of the probe by determining a rate of change in the position information for the distal end of the probe; anddetecting a loss in contact between the probe and the tissue surface based upon the determined velocity of the distal end of the probe. 2. The method according to claim 1, further comprising the step of storing the measured location data point in a memory to create a plurality of measured location data points, the plurality of measured location data points defining a geometry of the tissue surface. 3. The method according to claim 2, further comprising suspending the step of storing the measured location data points for a geometry lockout interval upon detection of a loss in contact between the probe and the tissue surface, whereby the geometry lockout interval permits the probe to be repositioned on the surface of the tissue. 4. The method according to claim 3, wherein the geometry lockout interval starts upon detecting that the probe is repositioned on the surface of the tissue. 5. The method according to claim 3, wherein a user defines the geometry lockout interval to be a period of time between about 1 second and about 5 seconds. 6. The method according to claim 2, further comprising the step of generating a 3-D model of the tissue surface using a plurality of measured location data points that are measured while the probe is in contact with the tissue surface. 7. The method according to claim 2, further comprising: suspending the step of storing the measured location data points upon detection of a loss in contact between the probe and the tissue surface;detecting reestablishment of contact between the probe and the tissue surface based upon the determined velocity of the distal end of the probe; andresuming the step of storing the measured location data points upon detecting reestablishment of contact between the probe and the tissue surface. 8. The method according to claim 1, wherein the step of determining a velocity of the distal end of the probe comprises determining a velocity of the distal end of the probe for each measured location data point and associating the determined velocity with the measured location data point, and wherein the step of detecting a loss in contact between the probe and the tissue surface comprises identifying at least one measured location data point that is not on the tissue surface based on having an associated velocity in excess of a velocity threshold. 9. A method of generating a plurality of location data points defining a model of a tissue surface, comprising: periodically measuring position information for a probe moving wi... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/04
US07877127 | B2 | 2011-01-25 | Hat-based oximeter sensor | 2006-07-26 | A method for use and an improved oximeter sensor substrate that is conforming to the shape of the patient's forehead. In one embodiment, the present invention is an oximeter sensor, having a substrate with a shape similar to a shape of at least a portion of a patient's forehead and including a section adapted to substantially fit over a portion of a forehead of a patient; an emitter disposed on the substrate at a position located on the section; and a detector disposed on the substrate at a distance from the emitter. In one embodiment, the substrate includes a hat that holds the emitter and the detector in a spaced-part manner against the patient's forehead. | 1. A sensor comprising: a stocking cap;a substrate disposed on the stocking cap;at least one sensing component disposed on the substrate; andan electrical connector disposed on the substrate, wherein the electrical connector extends from the substrate through an open portion of the stocking cap configured to be proximate to a top of the stocking cap when the stocking cap is applied to the patient. 2. The sensor of claim 1, wherein the substrate is conformable to a forehead-contacting surface of the stocking cap. 3. The sensor of claim 1, wherein the substrate is adapted to be attached to the stocking cap. 4. The sensor of claim 1, wherein the substrate comprises an adhesive layer adapted to attach the substrate to the stocking cap. 5. The sensor of claim 1, wherein the stocking cap comprises a neonatal stocking cap. 6. The sensor of claim 1, wherein the at least one sensing component is adapted to be attached to a forehead-contacting surface of the stocking cap. 7. The sensor of claim 1, wherein the stocking cap comprises an opening adapted to accommodate the the at least one sensing component. 8. The sensor of claim 7, wherein the opening comprises a hole or a pocket. 9. The sensor of claim 1, wherein the electrical connector is secured by a portion of the stocking cap. 10. The sensor of claim 1, wherein at least a portion of the electrical connector comprises a flexible circuit. 11. The sensor of claim 1, wherein the at least a portion of the electrical connector comprises a cable. 12. The sensor of claim 1, wherein the sensing component comprises at least one of an emitter or a detector. 13. The sensor of claim 1, wherein the stocking cap is configured to position the substrate at a predetermined position on the patient's forehead. 14. The sensor of claim 13, wherein the predetermined position comprises a lower forehead region. 15. The sensor of claim 13, wherein the predetermined position comprises a position substantially centering the sensing component above a patient's iris. 16. A method of manufacturing a cap-based sensor comprising: providing a stocking cap;providing a substrate;providing at least one sensing component disposed on the substrate; andproviding an electrical connector disposed on the substrate, wherein the electrical connector extends from the substrate through an open portion of the stocking cap configured to be proximate to a top of the stocking cap when the stocking cap is applied to the patient. 17. The method of claim 16, wherein the substrate is conformable to a forehead-contacting surface of the stocking cap. 18. The method of claim 16, wherein the substrate is adapted to be attached to the stocking cap. 19. The method of claim 16, comprising providing an adhesive layer adapted to attach the substrate to the stocking cap. 20. The method of claim 16, wherein the stocking cap comprises a neonatal stocking cap. 21. The method of claim 16, wherein the at least one sensing component is adapted to be at... | Nellcor Puritan Bennett LLC | A61B005/1455
US07865227 | B2 | 2011-01-04 | Image reconstruction method for cardiac gated magnetic resonance imaging | 2006-09-21 | A cardiac gated acquisition of MR data during a breath-hold employs a hybrid PR pulse sequence to acquire projection views from which image frames may be reconstructed at a plurality of cardiac phases during each heartbeat. Composite images are reconstructed at each cardiac phase using interleaved projection views acquired during all the heartbeats. The composite images are used to reconstruct the highly undersampled image frames at the same cardiac phase using a highly constrained backprojection method. | 1. A method for producing an image of a subject positioned in a field of view (FOV) of a magnetic resonance imaging (MRI) system, the steps comprising: a) acquiring with the MRI system a set of projection views at each of a plurality of cardiac phases and during each of a plurality of heartbeats;b) producing a composite image for each cardiac phase with projection views acquired during the corresponding cardiac phase for a plurality of heartbeats, each composite image depicting the subject at the corresponding cardiac phase;c) reconstructing an image of the subject at a selected cardiac phase by: c)i) backprojecting a set of projection views acquired at the selected cardiac phase into the FOV and weighting the value backprojected into each image pixel by the value of the corresponding pixel in the composite image that depicts the subject during the selected cardiac phase; andc)ii) summing the backprojected values for each image pixel. 2. The method as recited in claim 1 in which the projection views acquired at each cardiac phase are interleaved. 3. The method as recited in claim 1 in which each image pixel backprojected value Sn is calculated in step c)i) asSn=(P×Cn)/∑n=1N⁢⁢Cnwhere: P=the projection view value being backprojected;Cn=corresponding pixel value in the composite image;Sn=the value of the nth pixel along the backprojecting path; andN=total number of pixels along the backprojection path. 4. The method as recited in claim 1 in which step b) includes editing the composite image to remove an object therein and to thereby substantially minimize the appearance of that object in the reconstructed image. 5. The method as recited in claim 1 in which step c) includes Fourier transforming each projection view prior to backprojecting the view. 6. The method as recited in claim 1 in which step c) is repeated to produce an image at each of a plurality of cardiac phases in each of a plurality of heartbeats. 7. The method as recited in claim 1 in which step a) is performed under the direction of a hybrid 2D PR pulse sequence that acquires a plurality of sets of projection views at each cardiac phase that depict a corresponding plurality of slices of the subject. 8. The method as recited in claim 7 in which step b) includes producing a composite image for each slice at each cardiac phase. 9. The method as recited in claim 8 in which step c) includes reconstructing an image of the subject at each slice during the selected cardiac phase using the composite images corresponding to the selected cardiac phase and slice. 10. The method as recited in claim 9 in which the number of slice images reconstructed at the selected cardiac phase in step c) is greater than the number of sets of projection views acquired at each cardiac phase using the hybrid 2D PR pulse sequence. 11. The method as recited in claim 10 which includes producing an additional set of projection views at the selected cardiac phase using projection views other than the set of proj... | Wisconsin Alumni Research Foundation | A61B005/05 G06K009/00
US07869855 | B2 | 2011-01-11 | Medical apparatus with releasable applicator | 2006-09-28 | A medical apparatus includes a substrate member and an applicator member. Substrate member includes a medical device and an adhesive layer for adhering to body tissue. Applicator member is releasably attached to the substrate member and is adapted for grasping by a clinician to facilitate application of the substrate member to body tissue. | 1. A medical apparatus, which comprises: a substrate member including opposed upper and lower surfaces and opposed longitudinal ends, and defining a reference plane extending between the longitudinal ends, the substrate member including a medical device and having an adhesive layer adjacent the lower surface for adhering to body tissue; andan applicator member releasably attached to the upper surface of the substrate member, the applicator member having a grasping segment depending at least radially outwardly relative to the reference plane to traverse the reference plane to define first and second opposing grasping surfaces, the first and second opposing grasping surfaces dimensioned and arranged to be engaged by a clinician during application of the substrate member to body tissue, the applicator member being releasable subsequent to disposition of the adhesive layer to the body tissue. 2. The medical apparatus according to claim 1 wherein the substrate member is an at least partially flexible. 3. The medical apparatus according to claim 1 including a protective member releasably secured to the adhesive layer, the protective member being removable prior to application of the substrate member to the body tissue. 4. The medical apparatus according to claim 1 where the medical device is selected from a group consisting of a sensor and a marker. 5. The medical apparatus according to claim 4, wherein the sensor includes one of an accelerometer, thermal energy sensor, oxygen sensor, electromagnetic sensor and/or mechanical sensor. 6. The medical apparatus according to claim 4, wherein the marker includes one of a group consisting of LEDs, light bulbs, lamps and reflective devices. 7. The medical apparatus according to claim 1 wherein the medical device includes at least one electrode. 8. The medical apparatus according to claim 1 wherein the medical device includes one or more electrodes on an electrode array. 9. The medical apparatus according to claim 1 wherein the applicator member is substantially flexible. 10. The medical apparatus according to claim 1 wherein the substrate member includes an intermediate segment disposed between the opposed longitudinal ends thereof and wherein the grasping segment of the applicator member is arranged to depend outwardly from a position adjacent the intermediate segment and spaced from the opposed longitudinal ends. 11. The medical apparatus according to claim 10 wherein the grasping segment is arranged in oblique relation to the reference plane. 12. The medical apparatus according to claim 10 wherein the grasping segment includes a handle, the handle extending in substantial orthogonal relation to the reference plane. 13. The medical apparatus according to claim 10 wherein the medical device is selected from a group consisting of a sensor and a marker. 14. The medical apparatus according to claim 10 wherein the medical device includes at least one electrode. 15. The medical apparatus according to... | Tyco Healthcare Group LP | A61B005/04
US08442613 | B2 | 2013-05-14 | Mapping probe assembly with skived tube body frame | 2010-07-21 | An embodiment of a mapping probe assembly includes a body frame with a lumen therein. The body frame includes a catheter shaft region, a loop section and a transition region between the catheter shaft region and a loop section. A plurality of mapping electrodes are attached around the loop section. Electrical conductors extend through the lumen of the body frame to the mapping electrodes. In some embodiments, the loop section is skived, where a portion of the body frame is removed toward the interior of the loop section. The loop section has a generally planar loop, and further has a loop center. In some embodiments, the catheter shaft has an alignment generally perpendicular to the loop section where the alignment of the catheter shaft is along a line that intersects the planar loop proximate to the loop center. | 1. A mapping probe assembly, comprising: a body frame with a lumen therein, the body frame including a catheter shaft region and a skived loop section where a portion of the body frame is removed toward the interior of the skived loop section; anda plurality of mapping electrodes attached around the skived loop section, wherein electrical conductors extend through the lumen of the body frame to the mapping electrodes. 2. The mapping probe assembly of claim 1, wherein the body frame is constructed from a nickel titanium alloy. 3. The mapping probe assembly of claim 1, wherein the skived loop section has a tapered distal portion. 4. The mapping probe assembly of claim 1, further comprising a control wire connected to the skived loop section, wherein pushing the control wire distally with respect to the skived loop section increases a diameter of the loop section and pulling the control wire proximally with respect to the skived loop section decreases the diameter of the loop section. 5. The mapping probe assembly of claim 1, further comprising a braided portion around a distal portion of the body frame, wherein the braided portion extends substantially around the skived loop section. 6. The mapping probe assembly of claim 5, further comprising an outer shaft portion around the braided portion, wherein the outer shaft portion has holes through which the electrical conductors extend. 7. The mapping probe assembly of claim 1, wherein the plurality of mapping electrodes includes a plurality of ring electrodes approximately equally spaced around the skived loop section. 8. The mapping probe assembly of claim 1, wherein the skived loop section has a generally planar loop and a loop center, and the catheter shaft is aligned with a line that is generally perpendicular to the planar loop and that intersects the planar loop proximate to the loop center. 9. The mapping probe assembly of claim 8, wherein the skived loop section is sized and configured to provide a generally uniform force in a radially outward direction to secure the mapping probe assembly in a pulmonary vein ostium. 10. A mapping probe assembly, comprising: a body frame with a lumen therein, the body frame including a catheter shaft region, a loop section, and a transition region between the catheter shaft region and the loop section;the loop section having a generally planar loop, and further having a loop center, wherein the loop section of the body frame is a skived portion in which a portion of the body frame toward the loop center has been removed;the catheter shaft having an alignment generally perpendicular to the loop section, wherein the alignment of the catheter shaft is along a line that intersects the planar loop proximate to the loop center; anda plurality of mapping electrodes attached around the loop section, wherein electrical conductors extend through the lumen of the body frame to the mapping electrodes. 11. The mapping probe assembly of claim 10, wherein the body f... | Boston Scientific Scimed, Inc | A61B005/042
US08465408 | B2 | 2013-06-18 | Systems and methods for modulating the electrical activity of a brain using neuro-EEG synchronization therapy | 2010-08-04 | Described are methods, devices, and systems for a novel, inexpensive, easy to use therapy for treatment of coma, post-traumatic stress disorder, Parkinson's disease, cognitive performance, and/or amblyopia. Described are methods and devices to treat coma, post-traumatic stress disorder, Parkinson's disease, cognitive performance, and/or amblyopia that involves no medication. Methods and devices described herein use alternating magnetic fields to gently “tune” the brain and affect symptoms of coma, post-traumatic stress disorder, Parkinson's disease, cognitive performance, and/or amblyopia. | 1. A method of treating Parkinson's disease, treating coma, treating post traumatic stress disorder (PTSD), treating amblyopia, and/or enhancing cognitive performance, in a subject, comprising: (a) adjusting a setting of a magnetic field such that the magnetic field is configured to do one or more of the following:move an intrinsic frequency of a specified EEG band of the subject toward a target intrinsic frequency of the specified EEG band using the magnetic field,move a Q-factor of the intrinsic frequency toward a target Q-factor of the intrinsic frequency using the magnetic field,move a coherence value of intrinsic frequencies among multiple sites in a brain of the subject within a specified EEG band toward a target coherence value using the magnetic field wherein if the coherence value is higher than the target coherence value, applying at least two asynchronous magnetic fields close to the head of the subject, and wherein if the coherence value is lower than the target coherence value, applying at least one synchronized magnetic field close to the head of the subject; andmove an EEG phase between two sites in the brain of a subject of a specified EEG frequency toward a target EEG phase of the specified EEG frequency using the magnetic field wherein the magnetic field comprises one or more magnetic field generators that are of the same frequency and are in-phase with each other, of the same frequency and out of phase with each other, or a combination thereof; and(b) applying said magnetic field close to the head of the subject. 2. A method of treating Parkinson's disease, treating coma, treating post traumatic stress disorder (PTSD), treating amblyopia, and/or enhancing cognitive performance in a subject, comprising moving at least one of: an intrinsic frequency of a brain of the subject within a specified EEG band toward a pre-selected intrinsic frequency within the same specified EEG band and a Q-factor of the intrinsic frequency toward a target Q-factor of the intrinsic frequency, by applying a magnetic field close to a head of the subject, wherein the magnetic field comprises at least one of(a) a single target frequency within the specified EEG band; (b) a plurality of frequencies within the specified EEG band; and(c) the intrinsic frequency of the brain of the subject within the specified EEG band. 3. The method of claim 1 or 2, further comprising the step of measuring EEG data of the subject after the applying step. 4. The method of claim 1 or 2, further comprising the steps of: adjusting frequency of said magnetic field based on the EEG data of the subject; and repeating the applying step with an adjusted frequency. 5. The method of claim 1 or 2, wherein the applying of the magnetic field applies the magnetic field to a diffuse area of the brain of the subject. 6. The method of claim 1 or 2, wherein the magnetic field is generated by movement of at least one permanent magnet. 7. The method of claim 6, wherein the strength of the ... | Neosync, Inc. | A61N002/06 A61N002/12 A61N002/02
US09149208 | B2 | 2015-10-06 | Device for assessing the spinal column | 2009-03-06 | A device for assessing the spinal column intends to assess the lateral curvatures and deviations in the spinal column, in a standing position. It can be a non-invasive instrument that allows reproducing the position of each vertebra in the spinal column, from the first cervical vertebra to the first sacral vertebra. The device can include or consist of two parts: a first part designated as body and a second part defined as support. The first part can include or can consist of a vertical piece (that fits in the support) and several horizontal pieces. Each horizontal piece will touch the spinal crest of each vertebra in the spinal column, from the first cervical vertebra to the first sacral vertebra. The support can include or can consist of two pieces: a vertical piece, wherein the instrument's body vertical piece is fitted and a base whereon the feet of the patient under assessment are placed. | 1. A device for providing a non-invasive assessment of lateral curvatures, deviations, and a position of each vertebra from a first cervical vertebra to a first sacral vertebra in a spinal column of a person in a standing position, the device comprising: a support including, a base on which a person under the assessment stands; anda vertical support having a top and a bottom, the bottom being attached to the base;a vertical body attached to the top of the vertical support; anda plurality of position indicators attached horizontally to the vertical body, each ore of the plurality of position indicators further comprising: a handle portion fixed on the vertical body; andan integral rod abutting against the handle portion, the integral rod configured to make contact with a spinal crest of a vertebra from the first cervical vertebra to the first sacral vertebra via a horizontal ratchet coupled. to two pinions disposed on the integral rod, the integral rod having a scale indicating a position of each vertebra;wherein the vertical body includes a cap enclosing a profile attached to a vertical ratchet embedded within the profile, the vertical body being attached via two spigots to a top support plate at a top of the vertical body and to a bottom support plate at a bottom of the vertical body where the vertical body attaches to the vertical support;wherein the attachment of the vertical body to the top and bottom support surf aces allows for the cap, the vertical ratchet, and the profile to be held in place and further provides for an adjustment in the position of the cap, the vertical ratchet, and the profile; andwherein the device is configured to assess the position of each vertebra, including assess a horizontal position of each vertebra and a vertical position of each vertebra from the first cervical vertebra to the first sacral vertebra,wherein the device is configured to assess the horizontal position by reading a position of each vertebra from the scale and assess the vertical position by a vertical movement of each one of the plurality of position indicators wherein the vertical movement of each position indicator being made possible through a pinion roll located on the vertical ratchet of the vertical body. 2. The device for providing a non-invasive assessment, according to claim 1, wherein each of the plurality of position indicators includes one piece, configured to also horizontally and transversally move, that allows undertaking horizontal and anterior-posterior movements, wherein data concerning the position and deviation of each vertebrae are gathered subsequent to the placement of each of the plurality of position indicators at the level of a corresponding spinal crest. 3. A method of use of a device for providing a non-invasive assessment of lateral curvatures, deviations, and a position of each vertebra from a first cervical vertebra to a first sacral vertebra in a spinal column of person, the device comprising: a support including, ... | UNIVERSIDADE NOVA DE LISBOA | A61B005/107 A61B005/00
US20110009708 | A1 | 2011-01-13 | System and apparatus for improved nutrition analysis | 2010-06-30 | A system having an incorporated software program attempts to tailor a person's diet with the person's desired nutritional goal and their current food consumption. | 1. A processing system comprising: a machine readable medium, said machine readable medium being encoded to control a processing device, whereby said processing device is adapted to plan a meal for a person, said meal meeting a nutritional need of said person, said processing device executing the steps of: receiving biographical information about said person;determining a daily energy expenditure of said person;determining recommended daily nutritional intake for said person; andplanning a meal for said person using said biographical information, said energy expenditure, and said recommended daily nutritional intake. 2. The processing system of claim 1, wherein said step of determining an energy expenditure is based on receiving a general description of the person's lifestyle. 3. The processing system of claim 1, wherein said step of determining an energy expenditure is based on receiving a somewhat descriptive indication of the person's lifestyle. 4. The processing system of claim 1, further comprising: receiving a food preference for the person,wherein said step of planning a meal uses said person's food preference. 5. The processing system of claim 1, further comprising: receiving exercise information for said person; andplanning a snack for the person using said exercise information, said biographical information, said energy expenditure, and said recommended daily nutritional intake. 6. The processing system of claim 1, further comprising: receiving exercise information for said person; andplanning a plurality of snacks for the person using said exercise information, said biographical information, said energy expenditure, and said recommended daily nutritional intake, where one of said plurality of snacks occurs at a first time period before said exercise begins, where a second of said plurality of snacks occurs at a second time period after said exercise ends. 7. The processing system of claim 6, wherein said first and second time periods are approximately an hour. 8. The processing system of claim 1, further comprising: receiving exercise information for said person; andplanning a plurality of snacks for the person using said exercise information, said biographical information, said energy expenditure, and said recommended daily nutritional intake, where one of said plurality of snacks occurs at a first time period before said exercise begins, where a second of said plurality of snacks occurs at a second time period after said exercise ends, where a third of said plurality of snacks occurs at a third time period before said exercise begins, where a fourth of said plurality of snacks occurs at a fourth time period after said exercise ends. 9. The processing system of claim 8, wherein said first and second time periods are approximately an hour, wherein said third and fourth time periods are approximately two hours. 10. The processing system of claim 1, wherein said meal planning incorporates the person's fitness goal. 11. The proc... | - | A61B005/00
US20110040574 | A1 | 2011-02-17 | Health Monitoring System with Biometric Identification | 2008-03-25 | A health monitoring method and system is disclosed. A sensor for use in a health monitoring system includes a first transducer for collecting biometric data for biometrically identifying an individual, a second transducer for measuring at least one physiological parameter and outputting data indicative thereof, and an output for outputting the data. In addition, a health monitoring system includes a server comprising a data processor and a memory, and one or more of the sensors, wherein the server is arranged to receive the data from the one or more sensors, to process the data associated with the respective individual from which the data was collected, and to output at least one result of the processing of the data. | 1. A portable sensor adapted to be secured to an individual for use in a health monitoring system, comprising: a first transducer for collecting biometric data for biometrically identifying an individual;a second transducer for measuring at least one physiological parameter and outputting data indicative thereof; andan output for outputting said data with or in association with collected biometric data, so that said data can be suitably associated with said individual indicated by said collected biometric data. 2. A portable sensor as claimed in claim 1, wherein said first transducer is a fingerprint sensor, a DNA analyser, a retinal pattern sensor, an iris pattern sensor, a face recognition sensor. 3. A portable sensor as claimed in claim 1, wherein said second transducer is a heartbeat rate monitor, a breathing rate monitor, a blood sugar level monitor, a body fat content monitor, a blood alcohol content monitor, or a nicotine level monitor. 4. A portable sensor as claimed in claim 1, wherein said sensor includes an encryption module for digitally signing and/or encrypting said data before said data is outputted. 5. A portable sensor as claimed in claim 1, wherein sensor is configured to store a registration biometric measurement for subsequent comparison with a biometric data collected by said first transducer, and arranged to collect or output said data only if said registration biometric measurement agrees with said collected biometric data. 6. A portable sensor as claimed in claim 1, wherein said sensor is adapted to be worn by the individual. 7. A health monitoring system, comprising: a server comprising a data processor and a memory; andone or more portable sensors as claimed in claim 1;wherein said server is arranged to receive said data from said one or more sensors, to process said data associated with said respective individual from which said data was collected, and to output at least one result of said processing of said data. 8. A system as claimed in claim 7, wherein said server includes a cryptographic unit for decrypting said data, for authenticating said data, or for both decrypting and authenticating said data. 9. A system as claimed in claim 7, wherein said server includes a processing module for processing said data and identifying individuals with higher or lower health risk. 10. A system as claimed in claim 7, wherein said server includes an incentives module with incentive data, for processing said data and determining any reward or entitlement for said individual from said data and from said incentive data. 11. A system as claimed in claim 7, wherein said server is adapted to analyze said data, and to generate and output health advice for the individual from whom the data was collected. 12. A system as claimed in claim 1, wherein a trusted service provider operates said health monitoring system, issues said sensors to registered individuals, and collect and validate data collected from said sensors, and releas... | - | A61B005/024 H04L009/32 A61B005/00 A61B005/08 A61B005/145 G06Q050/00
US20110009725 | A1 | 2011-01-13 | PROVIDING CONTEXTUALLY RELEVANT ADVERTISEMENTS AND E-COMMERCE FEATURES IN A PERSONAL MEDICAL DEVICE SYSTEM | 2009-07-09 | A medical device system includes a portable (patient-worn or patient-carried) medical device that can automatically initiate certain e-commerce and/or advertisement presentation features in response to the detection of one or more triggering events. A triggering even may, for example, be associated with the monitoring of the operation or status of the medical device, the operation or status of the medical device system, a medical condition of the patient, treatment of the patient, and/or the preferences of the patient or user. In one application, the medical device is an insulin infusion pump and the user is automatically prompted to purchase, order, or otherwise obtain replaceable or replenishable items associated with the operation of the insulin infusion pump, such as a continuous glucose sensor, an infusion set, or an insulin reservoir. | 1. A method of operating a portable medical device having one or more consumables associated therewith, the method comprising: automatically determining, with the portable medical device, when a consumable associated with the portable medical device requires replacement or replenishment; andin response to the automatically determining step, presenting an advertisement at the portable medical device, the advertisement identifying a supplier, seller, or vendor of the consumable. 2. The method of claim 1, wherein: the advertisement includes a displayable active link to a supplier, seller, or vendor of the consumable; andthe method further comprises initiating an electronic transaction for the consumable with the displayable active link. 3. The method of claim 1, further comprising the step of obtaining geographic position data that indicates a location of the portable medical device, wherein the step of presenting the advertisement is influenced by the geographic position data. 4. The method of claim 3, further comprising processing the geographic position data to find advertisers that maintain inventory of the consumable at facilities located within a predetermined distance from the location of the portable medical device, wherein the advertisement identifies one of the advertisers. 5. The method of claim 4, further comprising the step of providing directions to a facility of the advertiser, wherein the directions are influenced by the geographic position data. 6. The method of claim 3, wherein the obtaining step obtains the geographic position data from a global positioning system (GPS) receiver of the portable medical device. 7. The method of claim 1, wherein: the portable medical device comprises a medication delivery device that delivers a medication to a patient; andthe advertisement identifies a supplier, seller, or vendor of the medication. 8. The method of claim 1, wherein: the portable medical device comprises an insulin infusion pump; andthe advertisement identifies a supplier, seller, or vendor of an insulin reservoir for the insulin infusion pump. 9. The method of claim 1, wherein: the portable medical device comprises an infusion pump that cooperates with an infusion set for a patient; andthe advertisement identifies a supplier, seller, or vendor of a replacement infusion set. 10. The method of claim 1, wherein: the portable medical device comprises an insulin infusion pump that responds to glucose measurements obtained from a glucose sensor worn by a patient; andthe advertisement identifies a supplier, seller, or vendor of a replacement glucose sensor. 11. The method of claim 1, wherein: the portable medical device comprises a controller configured to control delivery of a medication to a patient via a medication delivery device that is worn or carried by the patient; andthe advertisement identifies a supplier, seller, or vendor of a quantity of the medication. 12. The method of claim 1, wherein: the portable medical device... | MEDTRONIC MINIMED, INC. | A61M005/168 A61B005/15 G01C021/00 G06Q030/00 G06Q010/00 G06Q050/00
US08011125 | B2 | 2011-09-06 | Business form and self-laminating wristband with overlapping lamination panels | 2010-07-30 | A business form including a wristband is provided with a print area forming a badge, the badge having a length less than and a width greater than a width of either of two straps used to attach the wristband to the wrist, the badge thus having a greater print area to improve the readability of the indicia displayed therein. A business form including a wristband has a bulged rectangularly shaped print area to allow information to be printed transversely across the print area to improve readability of the indicia displayed in the print area. Thus, the bulged rectangularly shaped print area can accommodate one or both of bar codes printed thereon in ladder as well as picket fence fashion. A security seal is disclosed which comprises a small, adhesive coated patch of laminate for sealing the joinder of the wristband ends and which has points of weakness such as X's die cut therein so that it separates or destructs upon removal to provide an indication of tampering with the wristband. Also disclosed is an embodiment of the wristband form that is imprinted with identifying indicia for collecting, tracking and administering a patient's blood collection and blood product compatibility. | 1. A business form including a sheet with a separable wristband die cut therein, said sheet comprising a lamination ply and a face stock ply, the face stock ply having a die cut defining a wristband print area, the lamination ply having a die cut defining a wristband laminating portion, said laminating portion comprising a pair of offset panels with a pair of straps extending to their sides, at least one of said panels having an asymmetrical transition between it and the other of said panels, wherein the strap associated with the panel having the asymmetrical transition being substantially adhesive free along its length except for a patch of adhesive at its end, and a die cut in said face stock ply to define a protective layer of face stock covering said adhesive patch so that the wristband may be separated from the sheet by separating it at the protective layer. 2. The business form of claim 1 wherein said asymmetrical transition comprises a curvilinear transition. 3. The business form of claim 1 wherein said asymmetrical transition comprises a non-orthogonal transition. 4. The business form of claim 1 wherein said asymmetrical transition comprises a shoulder extending between it and the strap extending therefrom. 5. The business form of claim 1 further comprising a plurality of self adhering labels die cut therein. 6. A business form including a sheet with a separable wristband, said sheet comprising a lamination ply and a face stock ply, the face stock ply having a die cut defining a wristband print area, the lamination ply having a die cut defining a wristband laminating portion comprising a pair of panels with a pair of straps extending to their sides, wherein at least one of said panels has an offset shoulder transition between at least it and the other of said panels to facilitate separation of said wristband from said sheet while maintaining the integrity of said laminating portion, wherein the strap associated with the panel having the offset shoulder transition being substantially adhesive free along its length except for a patch of adhesive at its end, and a die cut in said face stock ply to define a protective layer of face stock covering said adhesive patch so that the wristband may be separated from the sheet by separating it at the protective layer. 7. The business form of claim 6 wherein said offset shoulder transition extends outboard of the other panel's corresponding end. 8. The business form of claim 7 further comprising a fold line between said pair of panels, said fold line aligned with and extending substantially to an end of the offset shoulder transition. 9. The business form of claim 8 wherein each of said panels has an offset shoulder transition, said offset shoulder transitions extending between it and its associated strap. 10. The business form of claim 6 wherein the panel end adjacent the offset shoulder transition has a substantially straight edge leading into said offset shoulder transition, said straight ... | Laser Band, LLC | A44C005/00
US07883478 | B2 | 2011-02-08 | Patient support having real time pressure control | 2005-05-02 | A patient support having real time pressure control. The patient support includes a plurality of sensors located beneath a bladder including a plurality of upright cylindrical elements. The pressure within the bladder is controlled based on the pressure or force sensed by the plurality of sensors. | 1. A pressure adjustable mattress system, comprising: a plurality of air bladders; a plurality of force sensors, each of the plurality of force sensors subtending at least one of the plurality of air bladders to sense a force transmitted through the subtended air bladder; a network, a control unit operably coupled to the network, a plurality of outputs, to transmit a signal representative of a sensed force to the network, each of the plurality of outputs coupled to at least one of the plurality of force sensors, the control unit being configured to continuously monitor the outputs, determine whether a person is supported by the air bladders, determine a position of a person supported b the air bladders determine whether a person supported by the air bladders is moving, and adjust pressure in the air bladders according to the person's position and according to whether the person is moving. 2. The pressure adjustable mattress of claim 1, wherein the plurality of force sensors comprise a light responsive device disposed in a compressible medium. 3. The pressure adjustable mattress of claim 2, further comprising a converter operatively coupled to each of the plurality of outputs, the converter including a digitizer to digitize the signal representative of the sensed force, and a filter, to filter the signal representative of the sensed force. 4. A method for adjusting the pressure of the air bladders of claim 1, to a pressure value, the mattress system including a plurality of force sensors located underneath the air bladders, wherein each of the plurality of force sensors generates a force signal representative of a sensed force transmitted through the air bladders of the pressure adjustable mattress system to the sensor, comprising: detecting, by force sensors located underneath air bladders, sensed forces transmitted through the air bladders, generating force signals representative of the sensed forces transmitted through the air bladders, ordering the force signals in a predetermined order; calculating an indicator signal as a function of the ordered force signals; comparing the indicator signal to the predetermined value to generate a correction signal; and adjusting the pressure of the air bladders of the pressure adjustable mattress system based on the correction signal. 5. The method of claim 4, wherein the ordering step comprises ordering the force signals in a predetermined order by reading the force signals in a predetermined sequence. 6. The method of claim 5, wherein the calculating step comprises calculating an indicator signal as a function of the ordered force signals by detecting a bottoming-out trend. 7. The method of claim 6, wherein the calculating step comprises calculating the indicator signal as a function of a standard deviation of the generated force signals. 8. In a pressure adjustable support, including a bladder to support a patient and having a pressure therein, a sensor subtending the bladder to sense a force tran... | Hill-Rom Services, Inc. | A61B005/103 A61B005/117 A47C016/00 A47C027/08 A47C027/10
US07881770 | B2 | 2011-02-01 | Multiple cannula image guided tool for image guided procedures | 2004-03-16 | Apparatus and methods are disclosed for use within an image-guided surgical navigation system for facilitating the combined positioning and orientation of multiple surgical implements. A tool guide having multiple cannulas is tracked by a surgical navigation system in real time. Position data of the tool guide is registered and combined with pre-acquired images by the navigation computer. Concurrent graphical representations of the plurality of cannulas are superimposed over the images and displayed. The display allows the surgeon to place the tool guide into the patient's body and position and orient the plurality of cannulas which are then used to place each of the implements. | 1. An apparatus for the placement of surgical implements, comprising: a fixture that defines a plurality of attachment points that are spaced apart and fixed relative to each other;a plurality of cannulas coupled to the fixture with each of the plurality of cannulas coupled to a respective one of the plurality of attachment points;at least one trackable marker mounted on a frame that is coupled to the fixture to enable an instrument location system to detect a position of each of the plurality of cannulas throughout a surgical procedure without obstructing any of the cannulas during the procedure; anda plurality of surgical implement receivers provided on the fixture for receiving surgical implements, at least one of the plurality of receivers being substantially coaxially aligned with a respective one of the plurality of cannulas to enable a surgical implement to pass through a respective one of the plurality of receivers and a respective one of the plurality of cannulas. 2. The apparatus of claim 1, further comprising: a peripheral structure, coupled to the fixture to move the fixture and thereby maneuver the plurality of cannulas during a surgical procedure to place the surgical implements. 3. The apparatus of claim 2, wherein at least one of the surgical receivers is a flange configured to receive the surgical implements. 4. The apparatus of claim 2, wherein at least one of the plurality of cannulas is individually adjustable to vary its length. 5. The apparatus of claim 2, wherein at least one of the plurality of cannulas is individually adjustable to vary an angular position of at least one of the plurality of cannulas relative to the fixture and the peripheral structure. 6. The apparatus of claim 5, wherein the at least one angular position is an azimuth angular position measured in a horizontal reference plane associated with the fixture. 7. The apparatus of claim 5, wherein the at least one angular position is an elevation angular position measured in a vertical reference plane associated with the fixture. 8. The apparatus of claim 2, wherein the fixture is interchangeable from the peripheral structure, and wherein the fixture is configured to accommodate at least one of: a variable number of cannulas, cannulas having a plurality of relative placements, or cannulas having a plurality of fixed relative angles. 9. The apparatus of claim 2, wherein the peripheral structure is a pistol grip configured graspable handle. 10. The apparatus of claim 2, wherein the peripheral structure is machine operated. 11. The apparatus of claim 2, further comprising: a plurality of interchangeable frames, at least two of the frames having a different type of trackable marker. 12. The apparatus of claim 2, wherein the surgical implement is at least one of a surgical tool or an implant. 13. The apparatus of claim 2, wherein at least one of the plurality of cannulas can be interchanged with a cannula having at least one of a different inner diameter... | Medtronic Navigation, Inc. | A61B005/05
US07881778 | B2 | 2011-02-01 | Floating physiological data acquisition system with expandable ECG and EEG | 2006-09-28 | A floating physiological data acquisition system with expandable ECG and EEG. The data acquisition system for obtaining electrophysiological signals from a patient comprises a patient monitor and separate patient side acquisition units connectable to the patient monitor. Each of the patient side acquisition units comprises an analog-to-digital converter, a serial interface communicative towards a serial bus connecting the patient side data acquisition unit with the patient monitor, and a serial interface controller processor. Each of the patient side acquisition units further forms a part of a floating patient applied part of the data acquisition system. | 1. A data acquisition system for obtaining electrophysiological signals from a patient, the system comprising: a serial bus;a first connector element connected to the serial bus;a first signal acquisition unit connected to the first connector element and configured to receive a first analog signal, a second analog signal, a third analog signal, and a fourth analog signal from a first limb electrode, a second limb electrode, a third limb electrode, and a fourth limb electrode, respectively, and a fifth precordial analog signal from a fifth precordial electrode,to amplify said first analog signal, said second analog signal and said third analog signal to yield a first amplified analog signal, a second amplified analog signal and a third amplified analog signal, respectively,to compute an average voltage of said first amplified analog signal, said second amplified analog signal and said third amplified analog signal,to amplify said fifth precordial analog signal to yield a fifth amplified precordial analog signal,to obtain a fifth compared amplified precordial analog signal as the potential difference between said fifth amplified precordial analog signal and said average voltage,to convert said first amplified analog signal, said second amplified analog signal, said third amplified analog signal, and said fifth compared amplified precordial analog signal to digital form to yield a digital signal,to provide said digital signal via a connector element to said serial bus, andto provide said average voltage via said first connector element to said serial bus;a second connector element connected to the serial bus; anda second signal acquisition unit connected to the second connector element and configured to receive a first precordial analog signal, a second precordial analog signal, a third precordial analog signal, a fourth precordial analog signal and a sixth precordial analog signal from a first, a second, a third, a fourth and a sixth precordial electrode, respectively,to receive said average voltage from said serial bus,to amplify said first precordial analog signal, said second precordial analog signal, said third precordial analog signal, said fourth precordial analog signal, and said sixth precordial analog signal to yield a first amplified precordial analog signal, a second amplified precordial analog signal, a third amplified precordial analog signal, a fourth amplified precordial analog signal and a sixth amplified precordial analog signal, respectively,to obtain a first compared amplified precordial analog signal as the potential difference between said first amplified precordial analog signal and said average voltage,to obtain a second compared amplified precordial analog signal as the potential difference between said second amplified precordial analog signal and said average voltage,to obtain a third compared amplified precordial analog signal as the potential difference between said third amplified precordial analog signal and said avera... | The General Electric Company | A61B005/04
US07889895 | B2 | 2011-02-15 | Method and apparatus for identifying pathology in brain images | 2003-12-12 | A method and apparatus for identifying pathology in a brain image comprises the steps of firstly determining the location of the midsagittal plane (MSP) of the brain illustrated in the image under examination by identifying the symmetry of the two hemispheres based on the determination of up to 16 approximated fissure line segments (AFLSs). Those AFLSs with a larger angular deviation from the MSP than a predefined threshold are considered as outlier AFLSs while the rest are taken as inlier AFLSs. The ratio of the number of the outlier AFLSs to the number of inlier AFLSs is then calculated. A comparison of the ratio with a further predetermined threshold value is made and if the ratio exceeds the further predetermined threshold value, pathology is present in the brain image. | 1. A method for identifying pathology in a brain image comprising: (a) determining by a processing device the location of the midsagittal plane (MSP) by calculating up to 16 approximated fissure line segments (AFLSs) and removing outlier AFLSs, the outlier AFLSs having a larger angular deviation from the MSP than a predefined threshold;(b) calculating by the processing device the ratio of the number of outlier AFLSs to the number of inlier AFLSs, the inlier AFLSs having a smaller angular deviation from the MSP than the predefined threshold; and(c) comparing by the processing device the ratio with a further predetermined threshold value, the ratio exceeding the further predetermined threshold value when pathology is present in the brain image. 2. A method as claimed in claim 1, wherein the image comprises image data and the method further comprises: reformatting the image data to align the image with the midsagittal plane. 3. A method as claimed in claim 2, wherein the reformatting is performed using vector manipulation. 4. A method as claimed in claim 2, wherein the midsagittal plane has an equation x′+d′=0 where x′ is a co-ordinate in the O′X′ direction and d′ is a constant, the image having an original co-ordinate system OXYZ before reformatting and a reformatted co-ordinate system after reformatting comprising co-ordinates O′X′, O′Y′ and O′Z′, the reformatting comprising: (i) starting with an MSP equation: ax+by+cz+d=0,where d is less than 0 and (a,b,c) is the unit normal vector of the MSP;(ii) determining two points that are the intersections between the MSP and the volume of the brain shown in the image under examination, denoting the two points as A (xA, yA, 0) and B (xB, yB, 0), both of the two points being on the MSP;(iii) when a is not equal to 0, calculating A and B as follows: xA=−d/a,yA=0xB=−(d+b(ySize−1))/a,yB=ySize−1where ySize is the number of voxels in the Y direction of the original volume; (iv) when a is 0, h should not be 0, and calculating A and B as follows: xA=0,yA=−d/b xB=xSize−1,yB=−(d+a(xSize−1))/b where xSize is the number of voxels in the X direction of the original volume; and (v) changing the original coordinate system OXYZ to the reformatted co-ordinate system O′X′Y′Z′, the unit vectors of the co-ordinates O′X′, O′Y′ and O′Z′ being calculated in the following manner: O′X′=(a,b,c)=(nx1,ny1,nz1)O′Y′=((xA−xB)/¦A−B¦,(yA−yB)/¦A−B¦,0)=(nx2,ny2,nz2)O′Z′=O′X′xO′Y′=(nx3,ny3,nz3)where ¦A−B¦ the Euclidean distance between points A and B, the transformation between OXYZ and O′X′Y′Z′ being defined as follows: X′=nx1X+ny1Y+nz1Z Y′=nx2X+ny2Y+nz2Z Z′=nx3X+ny3Y+nz3Z O′=0. 5. A method according to claim 1 further comprising: processing one or more slices through the image to determine interhemispheric symmetry. 6. A method according to claim 5, wherein the processing comprises processing one or more slices having a plane substantially perpendicular to the midsagittal plane. 7. A method according to claim 6, further comprising de... | Agency for Science, Technology and Research | G06K009/00 A61B005/05
US07885699 | B2 | 2011-02-08 | Analyte monitoring device and methods of use | 2010-08-06 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. An in vivo glucose monitoring system comprising: two in vivo working electrodes; anda receiver unit external to a user in signal communication with the two in vivo working electrodes, the receiver unit comprising: one or more processors; anda computer-readable storage medium including instructions which, when executed by the one or more processors, causes the one or more processors to: receive signals from both of the working electrodes indicative of a glucose level of the user;process the signals received from both of the working electrodes; anddetermine the acceleration of the rate of change of the glucose level. 2. The system of claim 1, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the one or more processors to determine whether the acceleration of the rate of change of the glucose level exceeds or meets one or more threshold values. 3. The system of claim 2, wherein the receiver unit further comprises an alarm, and wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the one or more processors to activate the alarm when the acceleration of the rate of change of the glucose level exceeds or meets the one or more threshold values. 4. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 5. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when a single datapoint indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 6. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when a predetermined number of datapoints obtained over a period of time indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 7. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when an average of datapoints obtained over a period of time indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 8. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automaticall... | Abbott Diabetes Care Inc. | A61B005/00
US07881767 | B2 | 2011-02-01 | Method and system for registering a medical situation associated with a first coordinate system, in a second coordinate system using an MPS system | 2004-08-18 | System for registering a first image with a second image, the system including a first medical positioning system for detecting a first position and orientation of the body of a patient, a second medical positioning system for detecting a second position and orientation of the body, and a registering module coupled with a second imager and with the second medical positioning system, the first medical positioning system being associated with and coupled with a first imager, the first imager acquiring the first image from the body, the first imager producing the first image by associating the first image with the first position and orientation, the second medical positioning system being associated with and coupled with the second imager, the second imager acquiring the second image and associating the second image with the second position and orientation, the registering module registering the first image with the second image, according to the first position and orientation and the second position and orientation. | 1. System for re-positioning a portion of the body of a patient at the same therapeutic position and orientation as an initial therapeutic position and orientation relative to a therapeutic device for medically retreating a selected tissue of the body automatically, during multiple therapeutic sessions, the system comprising: a positioning user interface;a position and orientation detector for detecting a multiple axis electromagnetic field and producing a respective electrical analog signal associated with said multiple axis electromagnetic field, said position and orientation detector being located at a selected fixed location associated with said selected tissue;an electromagnetic medical positioning system, coupled with a storage unit, with said positioning user interface and with said position and orientation detector, said medical positioning system being predeterminedly calibrated with said therapeutic device, said medical positioning system including: a plurality of electromagnetic transmitters for transmitting said multiple axis electromagnetic field; anda position and orientation processor coupled with said transmitters for determining said position and orientation of said position and orientation detector relative to said therapeutic device, according to said respective electrical analog signal;said medical positioning system determining an initial position and orientation of said position and orientation detector, during a first therapeutic session, while said selected tissue is placed in said initial therapeutic position and orientation, said medical positioning system indicating, during the multiple therapeutic sessions, via said positioning user interface when said position and orientation detector is placed again in said initial position and orientation, thereby establishing that said selected tissue is placed again in said initial therapeutic position and orientation; at least one moving mechanism coupled with said therapeutic device; anda controller coupled with said at least one moving mechanism directing said at least one moving mechanism to move said therapeutic device such that said therapeutic device medically treats at least a portion of said selected tissue. 2. The system according to claim 1, further comprising said therapeutic device. 3. The system according to claim 1, further comprising said storage unit for recording said initial position and orientation. 4. The system according to claim 1, wherein said positioning user interface produces at least one indication to determine whether a currently detected position and orientation is substantially the same as said initial position and orientation detector. 5. The system according to claim 4, wherein said at least one indication includes a first indication respective of said currently detected position and orientation, and a second indication respective of said initial position and orientation. 6. The system according to claim 4, wherein said system further comprise... | MediGuide Ltd. | A61B005/05
US07865233 | B2 | 2011-01-04 | Subcutaneous cardiac signal discrimination employing non-electrophysiologic signal | 2004-02-23 | A cardiac monitoring and/or stimulation system includes a housing coupled to a plurality of electrodes configured for subcutaneous non-intrathoracic sensing. A signal processor receives a plurality of composite signals associated with a plurality of sources, separates a signal from the plurality of composite signals, and identifies the separated signal as a cardiac signal using information derived from a non-electrophysiologic sensor, such as an accelerometer or acoustic transducer. The signal processor may iteratively correlate separated signals from the plurality of composite signals with a non-electrophysiologic sensor signal until the cardiac signal is identified. | 1. A signal separation method, comprising: detecting a composite electrical signal at a subcutaneous non-intrathoracic location, the composite electrical signal associated with a plurality of sources;receiving information associated with a non-electrophysiological cardiac source;separating a signal from the composite electrical signal using source separation; andverifying that the separated signal is a cardiac signal using the separated signal and the non-electrophysiological cardiac source information. 2. The method of claim 1, wherein verifying that the separated signal is the cardiac signal comprises providing a detection window defined by a start time and a stop time determined using the non-electrophysiological cardiac source information. 3. The method of claim 2, further comprising detecting a QRS complex within the detection window. 4. The method of claim 1, wherein the non-electrophysiological cardiac source information comprises acoustic emission information. 5. The method of claim 1, wherein the non-electrophysiological cardiac source information comprises a temporal location of a peak heart-sound. 6. The method of claim 5, wherein verifying that the separated signal is the cardiac signal comprises providing a detection window defined by a start time preceding the temporal location of a peak heart-sound. 7. The method of claim 1, wherein the non-electrophysiological cardiac source information comprises blood-flow information. 8. The method of claim 1, wherein the non-electrophysiological cardiac source information comprises pulse pressure information. 9. The method of claim 1, wherein the non-electrophysiological cardiac source information comprises pulse oximetry information. 10. The method of claim 1, wherein the non-electrophysiological cardiac source information comprises transthoracic impedance information. 11. The method of claim 1, wherein verifying that the separated signal is the cardiac signal comprises providing a detection window within which the cardiac signal is correlated to a signal associated with the non-electrophysiological cardiac source. 12. The method of claim 1, further comprising determining a time separation between a peak of the separated signal and a peak of a signal associated with the non- electrophysiological cardiac source. 13. The method of claim 12, wherein the time separation is used to identify a cardiac signal. 14. The method of claim 1, wherein the signal is separated from the composite electrical signal using blind source separation. 15. The method of claim 14, wherein the blind source separation comprises an independent component analysis performed on the composite electrical signal. 16. The method of claim 1, further comprising detecting a cardiac condition using the separated signal. 17. The method of claim 1, further comprising detecting a cardiac condition using the separated signal by performing a correlation between the separated signal and a signal associated with the non-ele... | Cardiac Pacemakers, Inc. | A61B005/0452
US07873410 | B2 | 2011-01-18 | Implantable medical device with electromechanical delay measurement for lead position and ventricular | 2006-04-26 | The cardiac rhythm management system includes an implantable medical device (IMD) with leads carrying electrodes for sensing cardiac electrical activity, and a physiologic sensor for sensing cardiac mechanical activity. The IMD measures electromechanical delays between electrical activity sensed by the electrodes and mechanical activity sensed by the physiologic sensor. The measured electromechanical delays can be used to detect lead dislodgement and to assess dyssynchrony between two areas of the heart, such as the right ventricle and the left ventricle. | 1. An implantable medical device (IMD) comprising: a plurality of electrical sensors responsive to electrical activity of a heart; andproducing signals indicative thereof;a physiologic sensor responsive to mechanical activity of the heart and producing signals indicative thereof; andsignal processing circuitry responsive to the signals from the mechanical and electrical sensors and deriving a time delay measurement for each electrical sensor with respect to the physiologic sensor. 2. The IMD of claim 1, wherein the signal processing circuitry comprises: means for determining dislodgement of one of the sensors based upon a change in at least one of the time delay measurements. 3. The IMD of claim 1, wherein the signal processing circuitry comprises: means for determining ventricular dyssynchrony based upon values of the time delay measurements. 4. The IMD of claim 1, wherein the physiologic sensor comprises one of a pressure sensor, an accelerometer and an impedance sensor. 5. An implantable medical device (IMD) comprising: electrodes usable different heart chambers, responsive to electrical activity and producing signals indicative thereof;a physiologic sensor responsive to mechanical activity associated with one of the chambers and producing signals indicative thereof; andcircuitry responsive to the signals from the electrodes and the physiologic sensor, and producing an output based upon electromechanical delays between electrical activity sensed by the electrodes and the mechanical activity sensed by the physiologic sensor. 6. The IMD of claim 5, wherein the output is indicative of dislodgement of a sensor. 7. The IMD of claim 5, further comprising: means for providing right ventricular (RV) pacing and left ventricular (LV) pacing; andmeans for defining VV timing between RV pacing and LV pacing; andwherein the output is indicative of the VV timing for cardiac resynchronization therapy. 8. The IMD of claim 5, wherein output is indicative of a detected ventricular dyssynchrony. 9. The IMD of claim 5, wherein the physiologic sensor comprises one of a pressure sensor, an accelerometer, and an impedance sensor. 10. An implantable medical device (IMD) comprising: a right atrial (RA) lead including an RA electrode producing an RA Sense signal;a right ventricular (RV) lead including an RV electrode for producing an RV Sense signal;a left ventricular (LV) lead including an LV electrode producing an LV Sense signal;a sensor producing a signal representative of cardiac mechanical activity; andmeans for deriving a first electromechanical delay measurement based on the RA Sense signal and the mechanical signal, a second electromechanical delay measurement based on the RV Sense signal and the mechanical signal, and a third electromechanical delay measurement based on the LV Sense signal and the mechanical signal. 11. The IMD of claim 10 and further comprising: means for providing an output indicating a dislodgement of one of the electrodes based u... | Medtronic, Inc. | A61B005/02 A61N001/04
US08706205 | B2 | 2014-04-22 | Functional analysis of neurophysiological data | 2008-11-30 | A method for functional analysis of neurophysiological data by decomposing neurophysiological data and EEG signal to form a plurality of signal features. The signal features may then optionally be analyzed to determined one or more patterns. | 1. A method of analyzing neurophysiological data, comprising: using a data processor for:decomposing the data into a plurality of overlapping sets of waveforms;identifying data patterns in said waveforms, said data patterns being in at least three-dimensions according to location, latency and frequency of peaks within said waveforms; anddetermining brain activity patterns based on said data patterns. 2. The method according to claim 1, wherein said identifying said data patterns is further according to amplitude to define four-dimensional data patterns. 3. The method according to claim 2, further comprising clustering said data patterns according to said latency, said frequency and said amplitude. 4. The method according to claim 3, further comprising combining a collection of said clusters to determine a brain network activity (BNA). 5. The method according to claim 2, further comprising clustering said data patterns according to said latency, said amplitude and said frequency, to provide a plurality of clusters in at least three-dimensions. 6. The method according to claim 5, further comprising combining a collection of said clusters to determine a brain network activity (BNA). 7. The method according to claim 1, further comprising determining a causality relation among data pattern components, and utilizing said relation to determine an activity network among said data patterns. 8. The method according to claim 7, further comprising determining a brain network activity (BNA) correlation to said activity network. 9. The method according to claim 8, wherein said determination of said BNA correlation comprises determining synchronization, or lack thereof, between a plurality of areas of the brain. 10. The method according to claim 8, wherein said identifying said data patterns comprises identifying source localizations for said BNA. 11. The method according to claim 10, further comprising comparing said data patterns to a previously determined pattern and correcting said source localizations based on said comparison. 12. The method according to claim 1, further comprising comparing said data patterns to a previously determined pattern. 13. The method according to claim 1, further comprising searching through a database of previously determined data patterns and selecting from said database a pattern closest to at least one of said identified data patterns. 14. The method according to claim 1, wherein said neurophysiological data comprise data acquired from multiple subjects for a particular behavioral process, and wherein said brain activity patterns are associated with said behavioral process. 15. The method according to claim 1, wherein said neurophysiological data comprise data pertaining to a spontaneous brain activity. 16. The method according to claim 1, wherein said neurophysiological data comprise data acquired before performing a task and data acquired during or after performing said task. 17. The method according to clai... | Elminda Ltd. | G01N033/48
US08475409 | B2 | 2013-07-02 | Infusion pump | 2010-07-08 | Described is an infusion pump comprising electronic infusion regulating means with wired or wireless communication means and a power source, a medicament bag (7) and an infusion device, said infusion device being in fluid communication with said medicament bag (7) and comprising two valves (1, 2), a pressure cavity (4) provided between said two valves (1, 2) and a membrane (3) covering said cavity (4), wherein said infusion device comprises at least two active actuators, preferably either made of electroactive polymer so as to form a self actuating membrane (3) or made of shape memory alloy wire, one of said actuators being adapted to apply pressure to said membrane (3) for fluid displacement in said cavity (4), and the other of said actuators being adapted to operate one of said valves which is adapted to passively close in the flow direction of a fluid from said medicament bag (7) through the cavity (4), wherein the other of said valves is adapted to, in particular passively, close in a direction opposite to said flow direction. | 1. An infusion pump, comprising electronic infusion regulating means with wired or wireless communication means and a power source, a medicament bag and an infusion device, said infusion device being in fluid communication with said medicament bag and comprising first and second valves, a pressure cavity provided between said first and second valves and a membrane covering said cavity, wherein said infusion device comprises at least two active actuators, one of said actuators being adapted to apply pressure to said membrane for fluid displacement in said cavity, and the other of said actuators being adapted to operate said first valve which is adapted to passively close in the flow direction of a fluid from said medicament bag through the cavity, wherein the second valve is adapted to close in a direction opposite to said flow direction. 2. The infusion pump according to claim 1, wherein in case said second valve operates as an active valve, said second valve is adapted to be actuated by the same actuator as said first valve in a way that always said second valve is closed when said first valve is open. 3. The infusion pump according to claim 1, wherein said infusion regulating means is adapted to monitor an actuation travel of said one of said actuators being adapted to apply pressure to said membrane for fluid displacement in said cavity through electrical parameters of actuator changing with actuation stress/travel, as well as time of travel for safety and occlusion pressure alarm determination. 4. The infusion pump according to claim 1, wherein said infusion device with said medicament bag define a disposable part, and the electronic infusion regulating means and eventually the power source are located in a non-disposable infusion regulator which is attached to said disposable part and communicates with it via contacts for driving, infusion error reading and optionally for pressure measurement. 5. The infusion pump according to claim 4 , wherein said disposable part is calibrated during production for infusion rate errors and comprises communication means for transmitting information about an error and for correcting this error by the infusion regulator. 6. The infusion pump according to claim 4, wherein said disposable part is adapted to be prefilled with a drug. 7. The infusion pump according to claim 4, wherein said disposable part comprises at least a hole covered by an hydrophobic, watertight but air-permeable material to facilitate sterilization with ethylene oxide, and by a rubber gasket for filling the medicament bag through a drug introduction needle. 8. The infusion pump according to claim 4, wherein said disposable part further comprises bags, a glucose sensor microflow calibration means, an infusion needle for insulin infusion, a measurement needle for glucose measurement, a back side microdialysis circuit being an internal part of said measurement needle, and a glucose sensor provided on an external part of said measuremen... | Micrel Medical Devices S.A. | A61M005/20
US08798717 | B2 | 2014-08-05 | Patient support and/or transport means and magnetic resonance system | 2010-06-29 | A patient support and/or transport device is proposed. The patient support and/or transport device comprises a receptacle in particular for a push-in patient support plate. At least one metal detection device is provided on the receptacle adjacent to the patient support plate when the patient support plate is pushed in. | 1. A device for supporting and/or transporting a patient, comprising: a patient support plate;a receptacle for inserting the patient support plate;a metal detection device arranged on an upper surface of the receptacle and below the inserted patient support plate for detecting a metal object on the inserted patient support plate;a blocking device for locking the inserted patient support plate in place if the metal object is detected on the inserted patient support plate; anda control element for releasing the blocking device if the metal object is removed and no further metal object is detected on the patient support plate. 2. The device as claimed in claim 1, further comprising a plurality of metal detection devices which at least partly cover a surface of the inserted patient support plate. 3. The device as claimed in claim 2, wherein the metal detection devices are arranged as a matrix. 4. The device as claimed in claim 2, wherein the metal detection devices are able to be activated and/or read out individually and/or in groups. 5. The device as claimed in claim 1, wherein the metal detection device comprises a transceiver unit for wireless communication. 6. The device as claimed in claim 1, further comprising an analysis unit for processing a sensor data of the metal detection device. 7. The device as claimed in claim 1, further comprising an output device for outputting a warning when the metal object is detected. 8. The device as claimed in claim 7, wherein the output device comprises an optical output device comprising a screen and/or a warning light. 9. The device as claimed in claim 7, wherein the output device comprises an acoustic output device comprising a loudspeaker. 10. The device as claimed in claim 1, wherein the blocking device is an electrical or a mechanical device. 11. The device as claimed in claim 1, wherein the device comprises a transport wagon, a patient table of a magnetic resonance system, or a magnetoencephalography device. 12. A magnetic resonance system, comprising: a patient table according to claim 1; anda magnet unit comprising a patient chamber into which the inserted patient support plate can be moved from the patient table. 13. The magnetic resonance system as claimed in claim 12, further comprising a control device for communicating with a transceiver unit of the metal detection device and for activating as well as reading out sensor data of the metal detection device. 14. The magnetic resonance system as claimed in claim 12, further comprising an output device for outputting a warning when the metal object is detected. 15. The magnetic resonance system as claimed in claim 14, wherein the output device comprises an optical output device comprising a screen and/or a warning light. 16. The magnetic resonance system as claimed in claim 14, wherein the output device comprises an acoustic output device comprising a loudspeaker. | Siemens Aktiengesellschaft | A61B005/055 A61G013/10 A61G001/02
US10052048 | B2 | 2018-08-21 | Respiratory monitors and monitoring methods | 2009-05-04 | A respiratory monitor comprises: a first sensor (20, 70) configured to generate a respiration-related motion monitoring signal (72) indicative of respiration related motion; a second sensor (20, 22, 80, 82) configured to generate a sound monitoring signal (84) indicative of respiration-related sound; and a signals synthesizer (90) configured to synthesize a respiratory monitor signal (46) based on the respiration-related motion monitoring signal and the respiration-related sound monitoring signal. A sensor for use in respiratory monitoring comprises an accelerometer (30) and a magnetometer (32) together defining a unitary sensor (20) configured for attachment to a respiring subject (10) so as to move as a unit responsive to respiration related motion of the respiring subject. | 1. A respiratory monitor comprising: a sensor configured to generate a sensor signal indicative of both respiration-related motion and respiration-related sound, the sensor including: a single accelerometer configured for attachment to a respiring subject and measuring an acceleration A=G·cos(φ) where A is the acceleration measured by the accelerometer, G is the magnitude of the gravitation vector, and φ is an orientation angle of the accelerometer,a magnetometer coupled with the accelerometer such that the accelerometer and the magnetometer have a common orientation; anda signal processor configured to generate the sensor signal based on the acceleration measured by the accelerometer and a magnetometer signal output by the magnetometer;a first filter configured to generate a respiration-related motion monitoring signal indicative of respiration-related motion by extracting from the sensor signal a low frequency signal encompassing frequencies of respiration-related motion;a second filter configured to generate a sound monitoring signal indicative of respiration-related sound by extracting from the sensor signal a high frequency signal encompassing frequencies of respiration-related sound, the respiration-related sound monitoring signal being based on the high frequency signal; anda signals synthesizer configured to synthesize a respiratory monitor signal by combining the respiration-related motion monitoring signal and the respiration-related sound monitoring signal to form the respiratory monitor signal. 2. The respiratory monitor as set forth in claim 1, wherein the second filter further comprises: an envelope filter or extractor configured to generate an envelope signal corresponding to an envelope of the high frequency signal, the respiration-related sound monitoring signal being based on the envelope signal. 3. The respiratory monitor as set forth in claim 1, wherein the accelerometer and the magnetometer are monolithically integrated on a common substrate. 4. The respiratory monitor as set forth in claim 1 wherein: the signals synthesizer is configured to synthesize the respiratory monitor signal by computing a cross-correlation of the respiration-related motion monitoring signal and the respiration-related sound monitoring signal. 5. A sensor for use in respiratory monitoring, the sensor comprising: an accelerometer;a magnetometer defining together with the accelerometer a unitary sensor configured for attachment to a respiring subject so as to move as a unit responsive to respiration-related motion of the respiring subject; anda signal processor configured to generate a respiratory monitor signal based on a single signal comprising change as a function of time of the angle of a three degrees of freedom (3 DOF) rotation of the unitary sensor computed from an accelerometer signal output by the accelerometer and a magnetometer signal output by the magnetometer. 6. The sensor of claim 5, wherein the signal processor is configured comput... | KONINKLIJKE PHILIPS N.V. | A61B005/08 A61B005/113 A61B007/00 H04B017/00
US09521955 | B2 | 2016-12-20 | Subdural electro-optical sensor | 2008-05-05 | A subdural electro-optical sensor system may include a substrate to which is attached an array of electrodes, light emitters, and light detectors. The sensor system may be sufficiently thin, flexible, sterile and biocompatible to be positioned subdurally. | 1. A subdural electro-optical sensor system comprising: a subdural electro-optical sensor having a substrate to which is attached an array of electrodes and multiple integrated light emitter-detectors; wherein the subdural electro-optical sensor is sufficiently thin, flexible, sterile and biocompatible to be positioned subdurally, wherein the electrodes alternate with the multiple integrated light emitter-detectors in the array and wherein each integrated light emitter-detector comprises one light detector and two light emitters, each of the two light emitters are configured to emit light at different wavelengths. 2. The sensor system of claim 1, wherein the two light emitters are configured to selectively emit light at 570 nm and at 610 nm, respectively. 3. The sensor system of claim 1, wherein the one light detector in each of the integrated light emitter-detectors is a photodiode, and the system further comprises an amplifier coupled to the photodiode. 4. The sensor system of claim 1, wherein each wavelength is centered at a target wavelength and comprises a wavelength band having a bandwidth of less than 20 nm. 5. The sensor system of claim 1, wherein the electrodes are thin disc-shaped electrodes having a diameter of 0.1 mm to 10 mm. 6. The sensor system of claim 1, further comprising a processor connected to the electrodes and the integrated light emitter-detectors so as to receive signals from them indicative of light detected by the integrated light emitter-detectors and voltages measured by the electrodes. 7. The sensor system of claim 6, further comprising an electrical stimulator so coupled to the processor as to initiate electrical stimulation in response to the received signals. 8. The sensor system of claim 7, wherein the electrical stimulator is selected from a group comprising an electrical surface stimulator and an electrical deep brain stimulator. 9. The sensor system of claim 6, wherein the processor is further configured to control each of the two light emitters of the same integrated emitter-detector to sequentially emit the light. 10. The sensor system of claim 6, wherein the processor is further configured generate an alert based on the received signals. 11. The sensor system of claim 1, wherein a distance between two electrodes is in a range from 1 mm to 20 mm. 12. The sensor system of claim 11, wherein a distance between two electrodes is 10 mm. 13. The sensor system of claim 1, wherein the substrate has a thickness of about 300 microns. 14. The sensor system of claim 13, wherein the electrodes and the integrated light emitter-detectors have a thickness of 500 microns. 15. A method of monitoring a property of a subdural structure in a subject where a sensor system is subdurally and superficially implanted into the subject's pia mater, where the sensor system comprises a substrate to which is attached an array of electrodes and multiple integrated light emitter-detectors, wherein the sensor system is sufficie... | CORNELL RESEARCH FOUNDTION, INC. GEORGETOWN UNIVERSITY | A61B005/00
US09460266 | B2 | 2016-10-04 | System and a method for scoring functional abilities of a patient | 2009-02-19 | A method for scoring one or more functional abilities of a patient that comprises exercising a patient according to an exercising pattern, using at least one sensor for sensing one or more manipulations of the patient during the exercising, and scoring the one or more functional abilities according to the sensed manipulations. The scoring is performed according to a visual rehabilitation evaluation scale for scoring the one or more functional abilities. | 1. A computer-implemented method for computerized scoring at least one functional ability of a patient, comprising: using at least one computing unit, said computing unit including at least one processor in communication with at least one sensor located on a robotic rehabilitation platform, and with a non-transitory computer readable medium storing program code for implementation by said at least one processor, to perform the following:a) receiving an output of said at least one sensor configured for measuring an applied force applied by a patient exercising on said robotic rehabilitation platform configured to be manipulated by said patient by moving a mechanism module of said robotic rehabilitation platform according to a predefined exercising pattern defining a trajectory of a limb manipulation, wherein said trajectory extends between at least two points in a space;b) identifying at least one manipulation of at least one limb of said patient during said predefined exercising pattern by analyzing said received output of said at least one sensor;c) computing a functional evaluation score for said at least one functional ability according to said at least one identified manipulation, said functional evaluation score computed according to a function based on a plurality of sub-scores including at least one sub-score calculated from said measured applied force, each sub-score defining a functional component of said at least one manipulation of said at least one limb during said at least one trajectory performed by said patient compared to a full performance capability reference value, and said at least one sub-score includes at least one member selected from the group consisting of: a quality of movement score, an active range of motion score, and passive range of motion score; andd) converting said computed functional evaluation score to a respective score of a manual rehabilitation evaluation scale using at least one correlation value obtained from a database storing correlation values between functional ability scores and said manual rehabilitation evaluation scale, to generate an objective evaluation of said patient based on said manual rehabilitation evaluation scale;wherein said manual rehabilitation evaluation scale comprises a member selected from a group consisting of: an action research arm test (ARAT), a stroke impact scale (SIS), a motor assessment scale (MAS), a wolf motor function test (WMFT), a motor status score for the Upper Extremity (MSS), and a manual function test (MFT). 2. The method of claim 1, further comprising generating a single functional evaluation score according to said plurality of sub-scores. 3. The method of claim 1, further comprising diagnosing said patient with regard to said functional ability according to said manual rehabilitation evaluation scale. 4. The method of claim 1, further comprising using said manual rehabilitation evaluation scale for monitoring a therapy given to said patient with regard to sai... | Motorika Limited | A61B005/103 G06F019/00 G06Q050/22 G06Q050/24
US08343463 | B2 | 2013-01-01 | Optical imaging agent | 2010-08-02 | The near-infrared wavelengths (700 nm-900 nm) are a suitable optical window for light penetration and deep tissue imaging. The present invention provides a near-infrared fluorescent blood pool contrast agent. The agent is of use to detect and quantify pathological capillary leak in live animals, e.g., serially imaging of traumatized tissue (muscle) as well as tumors. | 1. A compound having the formula:in which X is a member selected from —CH3, —COOH, and —SO3Ha, b, c, and d are independently selected integers from 1 to 10; andCD is cyclodextrin. 2. A pharmaceutical formulation comprising the compound according to claim 1 and a pharmaceutically acceptable carrier. | Numira Biosciences Inc. | A61B005/00
US07865377 | B2 | 2011-01-04 | Connecting consumers with service providers | 2008-02-14 | A request is received from a consumer of services to consult with a service provider having a service provider profile that satisfies at least some attributes in a set of attributes that define a suitable service provider; an available service provider satisfying at least some of the attributes in the set of attributes is identified; and a communication channel is provided to establish a communication between the consumer of services and the identified service provider. | 1. A computer-implemented method, the method comprising: receiving from a consumer, by a computer configured to connect consumers with providers of medical services, a search for a provider of medical services based on search criteria;generating by the computer an ordered list of providers that satisfy the search criteria, the order of providers being based on a provider rating of each provider included in the list; andreceiving a selection of one of the providers from the consumer;establishing an engagement over the Internet between the consumer and the selected one of the providers to connect the consumer and the selected one of the providers in real time;subsequent to establishing the engagement between the consumer and the provider, requesting feedback from the consumer related to the engagement during a process for disconnecting the consumer from the provider;updating, by the computer, the provider rating based, at least in part, on the received feedback if the consumer provides feedback about the provider; andcharging an additional amount if the consumer fails to provide feedback about the provider. 2. The method of claim 1, further comprising displaying the provider rating in a provider profile for the provider. 3. The method of claim 1, wherein the provider rating reflects the quality of the provider's interaction with multiple, different consumers. 4. The method of claim 1, further comprising updating the provider rating at the end of each interaction between the provider and a consumer. 5. The method of claim 1, further comprising displaying a number of engagements of the provider over a period of time. 6. The method of claim 1, further comprising displaying a ratio of the number of return engagements to a total number of engagements for the provider. 7. The method of claim 1, further comprising performing an evaluation of the service provider based at least in part on the consumer generated provider rating. 8. The method of claim 7, further comprising prohibiting the service provider from entering into an engagement with a consumer if the provider rating is less than a threshold. 9. A computer-implemented method, the method comprising: receiving by a computer an enrollment request from a new service provider to participate as a probationary provider of services on a system that establishes real-time engagements over the Internet between consumers of medical services and providers of medical services;enrolling the probationary service provider as an eligible participate provider of services; andmonitoring a provider rating of the probationary service provider during a probationary period after the probationary service provider begins to participate in the system. 10. The computer-implemented method of claim 9, further comprising: receiving feedback from consumers related to engagements with the probationary service provider during the probationary period;generating a provider rating based, at least in part, on the received feed... | American Well Corporation | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US20110034792 | A1 | 2011-02-10 | Noninvasive Body Chemistry Monitor and Method | 2009-08-05 | There is disclosed a noninvasive body chemistry monitor and method. A sensor may be disposed in close gaseous communication with the subject's skin. The sensor may be configured to provide a sensor output signal indicative of a level of at least one gaseous material released through the subject's skin. An event detector may detect a body chemistry event based on the sensor output signal. | 1. A noninvasive body chemistry monitor, comprising: a sensor disposed in close gaseous communication with a subject's skin, the sensor configured to provide a sensor output signal indicative of a level of at least one gaseous material released through the subject's skinan event detector to detect a body chemistry event based on the sensor output signal. 2. The body chemistry monitor of claim 1, wherein the sensor is selected from the group consisting of a metal oxide gas sensor, a conductive polymer gas sensor, an optical gas sensor, a spectrographic gas sensor, an electrochemical gas sensor, and an electronic nose. 3. The body chemistry monitor of claim 2, wherein the sensor is a metal oxide gas sensor configured to detect acetone the body chemistry event is at least one of a high blood glucose level and a low blood glucose level. 4. The body chemistry monitor of claim 1, further comprising: an alarm coupled to the event detector, wherein the event detector is configured to activate the alarm when a body chemistry event is detected. 5. The body chemistry monitor of claim 1, further comprising: a wireless transmitter coupled to the event detector, wherein the event detector sends an event message via the wireless transmitter when a body chemistry event is detected. 6. The body chemistry monitor of claim 5, wherein the wireless transmitter sends the event message to a predetermined destination via a network. 7. The body chemistry monitor of claim 5, wherein the wireless transmitter sends the event message to a remote unit configured to perform, upon receipt of the message, at least one of activating an alarm and sending a relayed event message to a predetermined destination via a network. 8. The body chemistry monitor of claim 7, wherein at least one of the body chemistry monitor and the remote unit includes a GPS receiver to provide the location of the subject. 9. The body chemistry monitor of claim 7, wherein the remote unit is a portable device configured for wireless communications via a networkthe wireless transmitter sends the event message to the portable device via a Bluetooth connection. 10. The body chemistry monitor of claim 9, wherein the portable device contains an application which, upon receipt of the message, causes the portable device to send a relayed event message to a predetermined destination via the network. 11. The body chemistry monitor of claim 1, wherein the event detector detects a body chemistry event based on one or more of the level of the sensor output signal and the rate of change of the sensor output signal. 12. The body chemistry monitor of claim 1, wherein the sensor is a plurality of sensorsthe event detector detects a body chemistry event based on a corresponding plurality of sensor output signals from the plurality of sensors. 13. The body chemistry monitor of claim 12, wherein the event detector detects a body chemistry event based a difference or a ratio of two sensor output signals from two sen... | - | A61B005/145
US20110040208 | A1 | 2011-02-17 | INTEGRATED LANCET AND TEST STRIP AND METHODS OF MAKING AND USING SAME | 2009-08-11 | An integrated lancet and test strip for measuring an analyte level is provided. The test strip includes a protrusion at a sample receiving end of the test strip, and a notch at a body receiving end of the test strip. Further, a method of measuring an analyte level in a sample of bodily fluid using the integrated lancing and testing meter is also provided. Still further, a method of making a plurality of test strips for use in an integrated lancet and test strip is also provided. | 1. An integrated lancet and test strip for measuring an analyte level, comprising: a body comprising a test strip receiving end and a lancet end;a lancet needle coupled to the lancet end of the body; anda test strip coupled to the test strip receiving end of the body, wherein the test strip comprises: multiple electrodes and assay chemistry for determining an analyte level in a sample;a protrusion at a sample receiving end of the test strip; anda body receiving end of the test strip;wherein the test strip and lancet needle are configured such that the test strip is not in fluid communication with the lancet needle. 2. The integrated lancet and test strip of claim 1, wherein the test strip further comprises a notch at the body receiving end of the test strip, and wherein the shape of the notch and protrusion are substantially triangular. 3. The integrated lancet and test strip of claim 1, wherein the test strip further comprises a notch at the body receiving end of the test strip, and wherein the notch comprises a shape different than a shape of the protrusion at the sample receiving end of the test strip. 4. The integrated lancet and test strip of claim 1, wherein the test strip further comprises a notch at the body receiving end of the test strip, and wherein the notch comprises a shape substantially the same as a shape of the protrusion at the sample receiving end of the test strip. 5. The integrated lancet and test strip of claim 4, wherein the test strip further comprises a notch at the body receiving end of the test strip, and wherein the shape of the notch and protrusion are one selected from the group consisting of substantially triangular, substantially semi-circular, and substantially rectangular. 6. The integrated lancet and test strip of claim 1, wherein the body receiving end of the test strip is opposite the fluid receiving end of the test strip. 7. The integrated lancet and test strip of claim 1, wherein the test strip further comprises a notch at the body receiving end of the test strip, and wherein a location of the notch along the body receiving end corresponds to a location of the protrusion along the sample receiving end. 8. The integrated lancet and test strip of claim 7, wherein the test strip further comprises a notch at the body receiving end of the test strip, and wherein the location of the notch is in a side region of the body receiving end, and the location of the protrusion is in a side region of the sample receiving end. 9. The integrated lancet and test strip of claim 1, wherein the test strip further comprises a notch at the body receiving end of the test strip, and wherein a location of the notch along the body receiving end is offset from a location of the protrusion along the sample receiving end. 10. The integrated lancet and test strip of claim 9, wherein the test strip further comprises a notch at the body receiving end of the test strip, and wherein the location of the notch is in a center region of ... | Abbott Diabetes Care Inc. | A61B005/151
US20110046473 | A1 | 2011-02-24 | EEG TRIGGERED FMRI SIGNAL ACQUISITION | 2009-08-20 | Neuro-response data including Electroencephalography (EEG) and Functional Magnetic Resonance Imaging (fMRI) data is collected, filtered and/or analyzed to evaluate the effectiveness of stimulus materials such as marketing and entertainment materials. A data collection mechanism obtains fMRI signals indicating a hemodynamic response to marketing or entertainment stimuli. In certain embodiments, such signals include region-specific blood oxygen level dependent (BOLD) signals that correlate with region-specific neural activity. fMRI signal acquisition is triggered by one or more EEG signatures indicating neural activity in response to exposure to stimulus materials. | 1. A system, comprising: a data collection mechanism including a plurality of modalities operable to obtain response data from a subject exposed to stimulus material including marketing and entertainment stimulus material, the response data comprising electroencephalography (EEG) response data and functional magnetic resonance imaging (fMRI) response data;an fMRI data collection initiator operable to initiate fMRI data collection using an EEG signature included in the EEG response data; 2. The system of claim 1, wherein a filter connected to the data collection mechanism is operable to remove cross-modality interference from the EEG response data and the fMRI response data. 3. The system of claim 1, wherein a cross-modality response synthesizer is operable to analyze EEG response data and fMRI response data to evaluate effectiveness of the stimulus material, wherein EEG response data is combined with fMRI response data. 4. The system of claim 1 wherein the fMRI data collection initiator is operable to initiate fMRI data collection using an EEG spike. 5. The system of claim 1 wherein the fMRI data collection initiator is operable to initiate fMRI data collection using an EEG polyspike. 6. The system of claim 1 wherein the fMRI data collection initiator is operable to initiate fMRI data collection using a recognizable EEG pattern. 7. The system of claim 1, wherein the EEG signature comprises event related potential (ERP) data. 8. The system of claim 1, wherein removing cross-modality interference comprises removing EEG generated artifacts from fMRI response data and fMRI generated artifacts from EEG response data. 9. The system of claim 1, wherein EEG response data is aligned with fMRI response data, wherein aligning EEG response data with fMRI response data comprises time and phase shifting. 10. The system of claim 1, wherein fMRI measures are aligned and combined with electroencephalography (EEG) to enhance estimates of effectiveness. 11. The system of claim 1, wherein EEG response data is combined with fMRI response data to determine attention, emotional engagement, and memory retention. 12. A method, comprising: obtaining response data using a plurality of modalities, the response data obtained from a subject exposed to stimulus material including marketing and entertainment stimulus material, the response data comprising electroencephalography (EEG) response data and functional magnetic resonance imaging (fMRI) response data, wherein obtaining response data using a plurality of modalities comprises triggering fMRI response data collection using an EEG signature. 13. The method of claim 12, further comprising removing cross-modality interference from the EEG response data and the fMRI response data. 14. The method of claim 12, further comprising analyzing EEG response data and fMRI response data to evaluate effectiveness of the stimulus material, wherein EEG response data is combined with fMRI response data. 15. The method of clai... | NeuroFocus, Inc. | A61B005/055 A61B005/0476
US20110015508 | A1 | 2011-01-20 | Analyte Monitoring and Management System and Methods Therefor | 2010-09-23 | Method and apparatus for providing multiple data receiver units in a data monitoring and management system such as analyte monitoring system where a first data receiver includes all of the functionalities for the data monitoring and management system receiver unit, and a second data receiver unit is configured with a limited functions to provide application specific convenience to the user or patient is disclosed. | 1. (canceled) 2. A system for continuous measurement of an analyte in a host, the system comprising: a continuous analyte sensor configured to measure a concentration of an analyte in a host; anda sensor electronics module physically connected to the continuous analyte sensor during operation of the continuous analyte sensor, wherein the sensor electronics module is configured to process a data stream associated with an analyte concentration measured by the continuous analyte sensor and to generate displayable sensor information based on at least some sensor data of the data stream, wherein the sensor electronics module comprises: a storage device configured to store at least some of the displayable sensor information; anda telemetry module configured to wirelessly transmit a first portion of the displayable sensor information to a first display device, wherein the first portion of displayable sensor information is formatted for display on the first display device, and to wirelessly transmit a second portion of the displayable sensor information to a second display device, wherein the second portion of displayable sensor information is formatted for display on the second display device. 3. The system of claim 2, wherein the first and second portions of the displayable sensor information each comprise a common subset of the displayable sensor information and the first portion also includes an additional subset of the displayable sensor information that is not included in the second portion. 4. The system of claim 2, wherein the telemetry module is further configured to wirelessly transmit a third portion of the displayable sensor information to a third display device, wherein the third portion of displayable sensor information is formatted for display on the third display device. 5. The system of claim 2, wherein the sensor electronics module is configured to communicate simultaneously with the first and second display devices. 6. The system of claim 2, wherein the first portion of the displayable sensor information is transmitted via a data network and the second portion of the displayable sensor information is transmitted via at least a cellular telephone network. 7. The system of claim 6, wherein the second portion of the displayable sensor information comprises at least one message for display. 8. The system of claim 2, wherein the first portion of displayable sensor information is selected based at least on a type of the first display device. 9. The system of claim 4, wherein display devices types include one or more of hand-held display devices, mobile devices, mobile phones, single point analyte monitor, drug delivery device, medical device, smart phones, laptop computer, and desktop computers. 10. The system of claim 2, wherein the sensor electronics module is further configured to determine if one or more sets of alert conditions are satisfied by the sensor data. 11. The system of claim 10, wherein each set of alert conditions i... | Abbott Diabetes Care Inc. | A61B005/05
US20110015509 | A1 | 2011-01-20 | Analyte Monitoring and Management System and Methods Therefor | 2010-09-23 | Method and apparatus for providing multiple data receiver units in a data monitoring and management system such as analyte monitoring system where a first data receiver includes all of the functionalities for the data monitoring and management system receiver unit, and a second data receiver unit is configured with a limited functions to provide application specific convenience to the user or patient is disclosed. | 1. (canceled) 2. A system for continuous measurement of a blood glucose level of a host, the system comprising: a continuous analyte sensor configured to determine a blood glucose level of a host;a storage device for storing a plurality of blood glucose levels of the host at each of the plurality of sample times;a sensor electronics module physically connected to the continuous analyte sensor during operation of the continuous analyte sensor, wherein the sensor electronics module is configured to determine whether at least some of the blood glucose levels match one or more requirements associated with a hypoglycemia or near hypoglycemia condition; andin response to determining that the one or more requirements associated with the hypoglycemia or near hypoglycemic condition are matched by the at least some of the blood glucose levels of the host, generate a first data package for transmission to a first device associated with the host, wherein the first data package includes displayable data indicating that the one or more requirements associated with the hypoglycemia condition or near hypoglycemic condition are matched; and substantially concurrently generate a second data package for transmission to a second device associated with a caretaker of the host, wherein the second data package includes displayable data indicating that the one or more requirements associated with the hypoglycemia condition are matched, wherein the first data package comprises data content customized for display on the first device and the second data package comprises data content customized for display on the second device. 3. The system of claim 2, wherein the caretaker of the host comprises one of a parent, a relative, a guardian, a doctor, and a nurse. 4. The system of claim 2, wherein the data indicating that the one or more requirements associated with the hypoglycemia condition are matched comprises an arrow, a numeric value and/or a graphical illustration. 5. The system of claim 2, further comprising: a telemetry module configured to wirelessly transmit the first data package to the first device and to wirelessly transmit the second data package to the second device. 6. The system of claim 5, wherein the telemetry module is configured to transmit data packages using one or more of: radio frequency (RF), infrared (IR), Bluetooth, frequency hopping, IEEE 802.11, paging network, and inductive coupling communication protocols. 7. The system of claim 6, wherein the first data package is transmitted via a first communication protocol and the second data package is transmitted via a second communication protocol that is different than the first communication protocol. 8. The system of claim 7, wherein the first communication protocol comprises Bluetooth and the second communication protocol comprises wireless local area network. 9. The system of claim 2, wherein the requirements include at least a threshold blood glucose level. 10. The system of claim 2, where... | Abbott Diabetes Care Inc. | A61B005/05
US20110021908 | A1 | 2011-01-27 | METHODS AND SYSTEMS FOR SPATIALLY IDENTIFYING ABNORMAL CELLS | 2010-05-27 | The present invention provides compositions and methods for imaging tumor resections. | 1. An method for spatially determining tissue heterogeneity in a subject undergoing surgery or a minimally invasive intervention, the method comprising: (a) administering a composition comprising a molecular imaging probe to said subject; and(b) obtaining an in situ image of said tissue wherein the image allows for the detection of a diseased cell, if present in said tissue. 2. A method of tumor resection comprising: (a) administering a composition comprising a molecular imaging probe to a tissue of a subject undergoing surgery or a minimally invasive intervention;(b) obtaining an in situ image of said tissue wherein the image allows for the detection of a diseased cell, if present in the tissue(c) removing said diseased cell detected in step (b)(d) repeating steps (b) and step (c) until no diseased cell is detected in said tissue. 3. A method of confirming the surgical margin of an excised tumor or growth comprising; (a) topically administering a composition comprising a molecular imaging probe to the surface of the excised tumor or growth;(b) obtaining an in situ image of said tumor or growth wherein the image allows for the detection of a diseased cell, if present in the tissue. 4. The method of claim 1, wherein said composition is administered systemically to said subject or topically to said tissue. 5. The method of claim 3, wherein said topical administration is by spraying or painting. 6. The method of claim 1, wherein said composition is administered on a film or sponge. 8. The method of claim 1, wherein said diseases cell is a tumor cell. 9. The method of claims 1, wherein said molecular imaging probe comprising a targeting moiety and an imaging moiety. 10. The method of claim 9, wherein said imaging moiety is a chromophore, a fluorochrome or a chemoluminescent moiety. 11. The method of claim 10, wherein said fluorochrome is a near infrared fluorochrome. 12. The method of claim 11, wherein said near infrared fluorochrome is Cy5.5, Cy5, Cy 7, Alexa, Fluoro 680, Alexa Fluoro 750, IRD41, IRD700, NIR-1, LaJolla Blue, indocyanine green (ICG), indotricarbocyanine (ITC), or a chelated lanthanide compound. 13. The method of claim 10, wherein said chemoluminescent moiety is a bioluminescent moiety. 12. The method of claim 10, wherein said bioluminescent moiety is a combination of luciference and luciferin, Guassia luciferase and colenterazine, or luminol and peroxide. 13. The method of claim 9, wherein said targeting moiety binds specifically to CD20, CD33, carcinoembryonic antigen (CEA), alpha fetoprotein (AFP), CA125, CA19-9, prostate specific antigen (PSA), human chorionic gonadotropin (HCG), acid phosphatase, neuron specific enolase, galacatosyl transferase II, immunoglobulin, CD326, her2NEU, EGFR, PSMA, TTF1, Muc, immature glycoslytaion, or an EMT marker. 14. The method according to claim 1, wherein the composition further comprises a UV dye or a fluorescent compound. 15. The method according claim 1, wherein the composition f... | LUMICELL DIAGNOSTICS, INC. | A61B005/00 A61K049/00
US20110009753 | A1 | 2011-01-13 | Respiration Rate Trending for Detecting Early Onset of Worsening Heart Failure | 2010-05-26 | Patient respiration is sensed from which respiration measurements are made, including a median respiration rate (MedRR) and a maximum respiration rate (MaxRR). Determinations are made as to whether an abnormality exists in MedRR and in MaxRR. An output indicative of the patient's tachypnea status is generated in response to determining the abnormality in MedRR and MaxRR. | 1. A method, comprising: sensing respiration and measuring a median respiration rate (MedRR) and a maximum respiration rate (MaxRR) using the sensed respiration;determining whether an abnormality exists in MedRR;determining whether an abnormality exists in MaxRR; andgenerating an output indicative of the patient's tachypnea status in response to determining the abnormality in MedRR and MaxRR. 2. The method of claim 1, wherein determining the abnormality in MedRR and MaxRR comprises: computing a baseline for MedRR and MaxRR, respectively;computing a near-term measure for MedRR and MaxRR, respectively; anddetermining the abnormality in MedRR and MaxRR based on a comparison of a difference between the baseline and near-term measure relative to a predetermined threshold for MedRR and MaxRR, respectively. 3. The method of claim 2, wherein: the baseline comprises one of a mean of a first predetermined period, a median of the first predetermined period, a weighted average of the first predetermined period, a predetermined value based on a population, and a predetermined value based on a patient's history; andthe near-term measure comprises one of a mean of a second predetermined period and a median of the second predetermined period. 4. The method of claim 3, wherein the first predetermined period is longer than the second predetermined period or overlaps with at least a portion of the second pre-determined period. 5. The method of claim 2, wherein determining the abnormality in MedRR and MaxRR comprises one of: determining whether the difference exceeds the predetermined threshold respectively for MedRR and MaxRR for a predetermined duration of time; anddetermining whether the difference exceeds the predetermined threshold respectively for MedRR and MaxRR for a predefined percentage of the predetermined duration of time. 6. The method of claim 2, wherein determining the abnormality in MedRR and MaxRR comprises determining whether the respective abnormality in MedRR and MaxRR occur within a predetermined time window. 7. The method of claim 2, wherein: determining the abnormality in MedRR and MaxRR comprises at least one of: determining whether the difference exceeds the predetermined threshold respectively for MedRR and MaxRR for a predetermined duration of time;determining whether the difference exceeds the predetermined threshold respectively for MedRR and MaxRR for a predefined percentage of the predetermined duration of time; anddetermining whether the respective abnormality in MedRR and MaxRR occur within a predetermined time window; andthe method comprising adjusting at least one of the predetermined threshold, predetermined duration, predefined percentage, and predetermined time window based on desired performance requirements. 8. The method of claim 2, comprising adjusting the respective predetermined threshold based on the respective baseline. 9. The method of claim 1, comprising determining whether the patient is engaged in patient act... | - | A61B005/0205 A61B005/08
US08971996 | B2 | 2015-03-03 | Medical imaging method and associated device | 2010-06-30 | A medical imaging method and associated device for generating an image data record of a recording region of a patient, which region is influenced by a cyclical cardiac motion, in which an EKG signal is used to derive a series of recording pulses matched to the cardiac motion, by which pulses the imaging is actuated in a pulsed fashion. In at least one embodiment, a time window of a future recording pulse is calculated taking into account at least one dispersion parameter characterizing the variation in the cycle duration and a location parameter characterizing the expected value of the cycle duration, wherein the dispersion parameter is included into the calculation of the time window using a weighting determined on the basis of the location parameter. Taking into account the dispersion parameter depending on the location parameter allows a reduction in the applied X-ray dose in the situations with irregular heart rates when a low mean heart rate is present at the same time, without significantly reducing the image quality of the generated image data record. | 1. An imaging method comprising: receiving an image data record of a recording region of a patient, the region being influenced by a cyclical cardiac motion which cycles at a cycle duration;registering an EKG signal of the cardiac motion of the patient;determining a dispersion parameter, the dispersion parameter indicating a variation in the cycle duration,determining a weight given to the dispersion parameter based on an expected value of the cycle duration,deriving a series of recording pulses from the EKG signal, matched to the cardiac motion by calculating a time window of a future recording pulse such that a size of the time window is based on at least (i) the dispersion parameter, (ii) the determined weight given thereto, and (iii) a cardiac frequency of the cyclical cardiac motion, andactuating a CT scanner in a pulsed fashion according to the requirements of the derived recording pulses to generate an image. 2. The imaging method as claimed in claim 1, wherein at least one of the dispersion parameter and the variation in the cycle duration is determined statistically by analyzing a number of preceding cardiac cycles. 3. The imaging method as claimed in claim 2, wherein the variation in the cycle duration is a mean value or a median value of the cycle duration of the analyzed cardiac cycles. 4. The imaging method as claimed in claim 3, wherein at least one of a minimum duration and a maximum duration of the analyzed cardiac cycles is regarded as the at least one dispersion parameter. 5. The imaging method as claimed in claim 4, wherein a start time of the time window is determined taking into account the minimum duration and an end time of the time window is determined taking into account the maximum duration of the analyzed cardiac cycles. 6. The imaging method as claimed in claim 3, wherein a standard deviation of the cycle duration of the analyzed cardiac cycles is regarded as the at least one dispersion parameter. 7. The imaging method as claimed in claim 3, wherein a standard deviation of the cycle duration compared to a trend of the cycle duration of the analyzed cardiac cycles is regarded as the at least one dispersion parameter. 8. The imaging method as claimed in claim 2, wherein at least one of a minimum duration and a maximum duration of the analyzed cardiac cycles is regarded as the at least one dispersion parameter. 9. The imaging method as claimed in claim 8, wherein a start time of the time window is determined taking into account the minimum duration and an end time of the time window is determined taking into account the maximum duration of the analyzed cardiac cycles. 10. The imaging method as claimed in claim 2, wherein a standard deviation of the cycle duration of the analyzed cardiac cycles is regarded as the at least one dispersion parameter. 11. The imaging method as claimed in claim 2, wherein a standard deviation of the cycle duration compared to a trend of the cycle duration of the analyzed cardiac cycle... | Siemens Aktiengesellschaft | A61B005/05 A61B006/00 A61B005/0456 A61B005/00
US08728065 | B2 | 2014-05-20 | Apparatus and methods for contactless electrophysiology studies | 2009-07-02 | An electrophysiology catheter includes an elongate catheter body having an elastically-deformable distal region predisposed to assume a spiral shape and a first plurality of electrodes disposed thereon. Each of the first plurality of electrodes includes an electrically active region limited to the inner surface of the spiral shape for use in non-contact electrophysiology studies. A second plurality of electrodes may also be disposed on the distal region interspersed (e.g., alternating) with the first plurality of electrodes, with each of the second plurality of electrodes having an electrically active region extending into the outer surface of the spiral shape for use in contact electrophysiology studies. The distal region may be deformed into a straight configuration for insertion into and navigation through the patient's vasculature, for example via use of a tubular introducer. As the distal region deploys beyond the distal end of the introducer, it resumes the spiral shape. | 1. An electrophysiology catheter, comprising: an elongate catheter body comprising an elastically-deformable distal region including a portion predisposed to assume a spiral shape comprising an inner surface on an inside of the spiral shape configured not to contact tissue and an outer surface on an outside of the spiral shape capable of contacting tissue when in a relaxed condition, the distal region terminating in a distal tip; and a plurality of electrodes disposed on the distal region, each of the plurality of electrodes comprising an electrically active region, wherein there are no electrically active regions on the outer surface of the spiral shaped portion of the distal region. 2. The electrophysiology catheter according to claim 1, further comprising a shape memory material extending through the distal region of the catheter body to predispose the spiral shaped portion of the distal region into the spiral shape. 3. The electrophysiology catheter according to claim 2, wherein the shape memory material comprises a shape memory metal wire. 4. The electrophysiology catheter according to claim 3, further comprising a polymeric tube, and wherein at least a portion of the shape memory metal wire is encased in the polymeric tube. 5. The electrophysiology catheter according to claim 2, wherein the shape memory material comprises a shape memory polymer. 6. The electrophysiology catheter according to claim 1, further comprising a spring element extending through the distal region of the catheter body to predispose the spiral shaped portion of the distal region into the spiral shape. 7. The electrophysiology catheter according to claim 1, wherein the distal region of the catheter body comprises a polymeric material thermoset into the spiral shape. 8. The electrophysiology catheter according to claim 1, wherein at least some of the plurality of electrodes comprise ring electrodes, wherein portions of the ring electrodes on the outer surface of the spiral shape are covered by an electrically insulating material. 9. The electrophysiology catheter according to claim 1, wherein at least some of the plurality of electrodes extend about a fraction of the inner surface of the spiral shape. 10. The electrophysiology catheter according to claim 1, wherein the spiral shape has a central axis, the elongate catheter has a longitudinal axis, and the central axis of the spiral shape extends continuously from the longitudinal axis of the elongate catheter. 11. A kit for conducting electrophysiology studies, comprising: at least one of a guidewire and a tubular introducer comprising an elongate introducer body comprising a distal end, a proximal end, and a lumen extending from the proximal end to the distal end; and an electrophysiology catheter comprising an elongate catheter body adapted to be one of deployed through the lumen of the introducer or introduced over the guidewire and comprising an elastically-deformable distal region predisposed to assume a ... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B018/18 A61B005/0452 A61B005/0402
US09931078 | B2 | 2018-04-03 | Near-infrared spectroscopy sensor with light sheet | 2010-07-08 | A sensor includes a sensor pad configured to be disposed on a portion of a patient's body. A light sheet is disposed on the sensor pad and has a first substrate and a second substrate spaced from one another. The light sheet further includes a light source configured to emit near-infrared light and a light detector configured to detect near-infrared light. The light source and the light detector are disposed between the substrates. The sensor pad is configured to allow light generated by the light source to travel through the portion of the patient's body to the light detector. The light received by the light detector is indicative of oxygen saturation of the portion of the patient's body through which the light travelled. | 1. A sensor comprising: a sensor pad configured to be disposed on a portion of a patient's body; anda light sheet disposed on the sensor pad and having a first substrate and a second substrate at least partially spaced from one another, wherein the light sheet includes a light source and a light detector and wherein the light source and the light detector are disposed between the first and second substrates such that the light source is configured to emit near-infrared light through the first substrate into the patient's body and the light detector is configured to detect near-infrared light reflected from the patient's body through the first substrate,wherein the first substrate includes a first layer having a conductive pattern and a second layer having a conductive pattern and wherein the first layer and the second layer are separated by a non-conductive spacer material,wherein the light source and the light detector are both electrically connected between the second substrate and the second layer of the first substrate. 2. The sensor as set forth in claim 1, further comprising a non-conductive adhesive disposed between the substrates. 3. The sensor as set forth in claim 2, wherein the non-conductive adhesive is configured to prevent the light detector from receiving light generated by the light source through the non-conductive adhesive. 4. The sensor as set forth in claim 1, further comprising a light blocker disposed between the substrates and configured to prevent the light detector from receiving light travelling between the substrates. 5. The sensor as set forth in claim 1, wherein at least the first substrate is transparent to near-infrared light. 6. The sensor as set forth in claim 1, wherein the second substrate includes an electrically conductive pattern. 7. The sensor as set forth in claim 6, wherein the light source and the light detector are in electrical communication with one another via the electrically conductive pattern. 8. The sensor as set forth in claim 1, further comprising a light-blocking pattern disposed on at least one of the substrates. 9. The sensor as set forth in claim 8, wherein the sensor pad is non-conductive. 10. The sensor as set forth in claim 9, wherein the sensor pad is configured to prevent ambient light from interfering with the light detector. 11. A light sheet comprising: a first substrate having a first layer having a conductive pattern and a second layer having a conductive pattern, wherein the first layer and the second layer are separated by a non-conductive spacer material;a second substrate at least partially spaced from the first substrate;a light source disposed between the first substrate and the second substrate and configured to emit near-infrared light through the first substrate; anda light detector disposed between the first substrate and the second substrate and configured to detect near-infrared light through the first substrate;wherein the light source and light detector are ... | COVIDIEN LP | A61B005/00 A61B005/1455
US10463838 | B2 | 2019-11-05 | Vascular access methods and devices | 2010-08-19 | Methods and devices for detecting positioning of a probe in a tissue of a patient. A method can include providing a detection device; advancing a device coupled probe through the tissue of the patient and toward the patient's target tissue; detecting a change in pressure about the distal portion of the coupled probe during advancing, where the detected pressure change indicates probe positioning in a vein or artery of the patient; outputting the detected pressure change or indication of probe positioning to a visual display. | 1. A method of distinguishing between positioning of a probe in a vein or an artery of a patient, comprising: providing a device comprising a housing having a proximal housing portion and a distal housing portion, the distal housing portion coupled to the probe, the device further comprising,a pressure sensing system at least partially carried by the housing and comprising a processing unit coupled with a pressure sensor so as to receive a signal from the pressure sensor and determine from the signal a pressure value of an environment about a distal portion of the probe, andan output unit carried by the housing intermediate the proximal housing portion and the distal housing portion, the output unit comprising a visual display, the output unit coupled to the pressure sensing system so as to receive a pressure value signal and output on the visual display a reporting signal indicating the determined pressure value and/or positioning of the distal portion of the probe in the vein or the artery of the patient;advancing the device distally such that the distal portion of the probe advances through tissue of the patient and into the vein of the patient;detecting a change in pressure about the distal portion of the probe during advancing so as to indicate positioning of the distal portion of the probe;distinguishing between positioning of the distal portion in the vein of the patient in response to the reporting signal on the visual display;introducing a guidewire through the housing proximally to distally generally align with the probe into the vein following distinguishing of probe positioning, wherein the guidewire extends out a distal end of the probe; anddrawing blood toward a syringe coupled to the proximal housing portion of the housing for blood visualization. 2. The method of claim 1, wherein the probe is rigidly coupled to the device so as to allow positioning of probe by position manipulation of the device housing by a user; wherein advancing includes advancing a probe tip of the probe through the tissue of the patient, wherein detecting a change in pressure includes transmitting pressure from the probe tip through a bore in the probe to the pressure sensing system; and wherein advancing includes positioning the output unit adjacent a patient's body during advancing so that the output unit is in a line of sight by the user advancing the device. 3. The method of claim 1, wherein the probe is directly coupled to the device. 4. The method of claim 1, wherein the vein is an internal jugular vein, a subclavian vein, or a femoral vein. 5. The method of claim 1, further comprising detecting an increase in pressure upon probe entry into the vein and discontinuing advancement. 6. The method of claim 1, further comprising detecting an increase of pressure indicating that the probe has entered an artery, and altering device advancement based on the detected pressure. 7. The method of claim 1, further comprising withdrawing the probe and the housi... | Medline Industries, Inc | A61M025/06 A61B005/0215 A61B005/00 A61M005/42 A61B090/00 A61B008/08 A61M005/168 A61M005/46 A61B017/34 A61B008/00
US20110015517 | A1 | 2011-01-20 | MRI BIOPSY DEVICE LOCALIZATION FIXTURE | 2010-09-24 | A localization mechanism, or fixture, is used in conjunction with a breast coil for breast compression and for guiding a core biopsy instrument during prone biopsy procedures in both open and closed Magnetic Resonance Imaging (MRI) machines. The localization fixture includes a three-dimensional Cartesian positionable guide for supporting and orienting an MRI-compatible biopsy instrument, and, in particular, a sleeve to a biopsy site of suspicious tissues or lesions. A z-stop enhances accurate insertion, prevents over-insertion or inadvertent retraction of the sleeve. The sleeve receives a probe of the MRI-compatible biopsy instrument and may contain various features to enhance its imagability, to enhance vacuum and pressure assist therethrough, etc. | 1. A compression assembly, comprising: (a) a frame; and(b) a plurality of adjustable compression members positioned within the frame, where the adjustable compression members and the frame are sized and configured to compress and hold a patient's breast against a fixed structure, wherein at least one adjustable compression member of the plurality of adjustable compression members is movable relative to the frame to open an access point to a patient's breast. 2. The compression assembly of claim 1, wherein the adjustable compression members comprise bars. 3. The compression assembly of claim 2, wherein the bars are substantially cylindraceous. 4. The compression assembly of claim 3, wherein at least one of the bars has a first end, a second end, and an annular flange at each of the first and second ends. 5. The compression assembly of claim 1, wherein the adjustable compression members are removable from the frame. 6. The compression assembly of claim 5, wherein the frame defines at least one L-shaped slot, wherein at least one of the adjustable compression members is selectively engagable with the frame by traveling along the L-shaped slot. 7. The compression assembly of claim 5, wherein each adjustable compression member comprises a resilient portion configured to deform as the adjustable compression member is selectively coupled with or decoupled from the frame. 8. The compression assembly of claim 7, wherein each resilient portion is resiliently biased to assume a bowed configuration. 9. The compression assembly of claim 8, wherein the frame includes slots configured to receive the resilient portions of the adjustable compression members. 10. The compression assembly of claim 1, wherein the frame extends along a horizontal dimension and along a vertical dimension 11. The compression assembly of claim 10, wherein the adjustable compression members are oriented to extend along the vertical dimension. 12. The compression assembly of claim 10, wherein the adjustable compression members are movable relative to the frame along the horizontal dimension. 13. The compression assembly of claim 10, wherein the frame is substantially rectangular. 14. The compression assembly of claim 1, wherein the adjustable compression members are substantially parallel to each other. 15. The compression assembly of claim 1, further comprising at least one fixed compression member, wherein the at least one fixed compression member is fixed to the frame. 16. The compression assembly of claim 1, wherein the adjustable compression members comprise vanes. 17. The compression assembly of claim 1, wherein the frame includes a ramped surface, wherein each adjustable compression member comprises a ramped surface configured to complement the ramped surface of the frame. 18. The compression assembly of claim 1, further comprising a tray configured to engage the frame. 19. A compression assembly, comprising: (a) a first compression fixture, wherein the first comp... | DEVICOR MEDICAL PRODUCTS, INC. | A61B005/055
US20110046469 | A1 | 2011-02-24 | Glucose Measuring Device for Use In Personal Area Network | 2010-10-27 | A glucose measuring system includes a glucose meter that incorporates wireless communication capabilities. The meter can advantageously be a low cost meter by eliminating expensive components, such as the display. The user nevertheless is able to retrieve and view his or her glucose values by referring to displays within the glucose measuring local area network. Feedback via these displays can advantageously be used by the diabetic to create a higher level of confidence and safety. | 1-20. (canceled) 21. An analyte monitoring system, comprising: a client unit having a test strip port to mechanically receive an analyte test strip, an analog interface to communicate with the test strip, a processor coupled to the analog interface, and a wireless data transmitter coupled to the processor to send analyte data to one or more devices on a personal area network;a server unit in data communication with the client unit; anda data processing terminal in direct electronic contact with the server unit to receive analyte data from the server unit. 22. The analyte monitoring system of claim 21, wherein the analyte test strip is a glucose test strip. 23. The analyte monitoring system of claim 21, wherein the client unit is a blood glucose meter. 24. The analyte monitoring system of claim 21, wherein the server unit is in data communication with the client unit via a wireless data transmission protocol. 25. The analyte monitoring system of claim 24, wherein the wireless data transmission protocol is selected from a group consisting of: a 802.11 protocol, a Bluetooth protocol, an RF protocol, and an IrDA protocol. 26. The analyte monitoring system of claim 21, wherein the server unit includes a display. 27. The analyte monitoring system of claim 21, wherein the client unit includes a user input interface having one or more push buttons. 28. The analyte monitoring system of claim 21, wherein the client unit does not include a display. 29. The analyte monitoring system of claim 21, wherein the client unit does not include a user interface. 30. The analyte monitoring system of claim 21, wherein the client unit includes a password protection protocol. 31. The analyte monitoring system of claim 21, wherein a user must enter an identity code in order to verify that the user is an authorized user. 32. The analyte monitoring system of claim 21, wherein the data processing terminal further includes a wireless data receiver. 33. The analyte monitoring system of claim 32, wherein the data processing terminal receives data from the server via a wireless data receiver. 34. The analyte monitoring system of claim 21, wherein the server unit includes a display and an input device. 35. The analyte monitoring system of claim 34, wherein the display is a liquid crystal display (LCD). 36. The analyte monitoring system of claim 34, wherein the input device is a button, keyboard, mouse, or touch-screen. 37. The analyte monitoring system of claim 34, wherein a user can control the client unit via the input device on the server unit. 38. The analyte monitoring system of claim 21, wherein the server unit receives instructions from the data processing terminal and transmits the instructions to the client unit. 39. The analyte monitoring system of claim 21, wherein the one or more devices are selected from a group consisting of: a desktop, a laptop, a handheld computer, a printer, a telephone, a navigation computer in a car, a car computer system, a ... | Abbott Diabetes Care Inc. | A61B005/145
US20110046468 | A1 | 2011-02-24 | Method and Device for Utilizing Analyte Levels to Assist in the Treatment of Diabetes | 2010-08-23 | A health-monitoring device assesses the health of a user based on levels of two analytes in a biological fluid. A first analyte that is utilized to assess a user's health is a fat metabolism analyte, such as ketones, free fatty acids and glycerol, which is indicative of fat metabolism. A second analyte that is utilized is a glucose metabolism analyte, such as glucose. The levels of the two analytes may be used to assess insulin sensitivity, to detect both recent hypoglycemia and the cause of high glucose levels, and/or to guide therapeutic intervention. The dual analyte model may calculate a discrepancy between an actual insulin activity level and a theoretical insulin activity level. The dual analyte model of the present invention may be used to identify individuals at risk for metabolic syndrome, insulin resistance and non-insulin dependent diabetes, and allows monitoring of the progression of those disease states, as well as progress made by therapeutic interventions. | 1.-20. (canceled) 21. A health-monitoring device comprising: a testing device configured to measure a level of one or more analytes in a sample from a user and generate a signal indicative of the level of the one or more analytes in the sample;a processor comprising a calculator comprising programming to calculate a concentration of the one or more analytes in the sample based on the signal and a correlator comprising an algorithm configured to determine a correlation between the level of the one or more analytes in the sample and a health parameter indicative of the user's health; anda user interface comprising an input configured to receive information from the user and a display configured to display at least one of the concentration of the one or more analytes in the sample and the health parameter. 22. The device of claim 21, wherein the health parameter comprises an insulin sensitivity factor, a medication dosage, an assessment of metabolic syndrome, a likelihood of the user developing hypoglycemia, a likelihood of the user developing hyperglycemia and a likelihood that the user recently developed hypoglycemia. 23. The device of claim 22, wherein the health parameter comprises the insulin sensitivity factor. 24. The device of claim 22, wherein the medication dosage comprises an insulin dosage. 25. The device of claim 21, wherein the algorithm includes the information from the user to determine the correlation between the level of the one or more analytes in the sample and the health parameter. 26. The device of claim 25, wherein the information from the user comprises one or more of a type of medication dosage, an amount of medication dosage, a type of food consumption, an amount of food consumption, a weight of the user, and a body composition of the user. 27. The device of claim 21, further comprising a data communication port configured to connect the health-monitoring device to a remote site. 28. The device of claim 27, wherein the data communication port is configured to connect the health-monitoring device to the remote site over a network. 29. The device of claim 27, wherein the data communication port is configured to provide a wireless connection between the health-monitoring device and the remote site. 30. The device of claim 29, wherein the wireless connection comprises an acoustic connection, an optic connection or a radio wave connection. 31. The device of claim 30, wherein the optic connection comprises an infrared connection. 32. The device of claim 21, wherein the testing device is configured to measure the level of a glucose analyte in the sample. 33. The device of claim 32, wherein the glucose analyte comprises glucose, pyruvate, glucose-6-phosphate or lactate. 34. The device of claim 21, wherein the testing device is configured to measure the level of a fat analyte in the sample. 35. The device of claim 34, wherein the fat analyte comprises a ketone, a free fatty acid, glycerol or an analyte indicative of l... | Abbott Diabetes Care Inc. | A61B005/145
US07894900 | B2 | 2011-02-22 | Device and method for monitoring cardiac pacing rate | 2007-08-30 | A device for monitoring cardiac pacing rate having a measuring unit for receiving an electrical signal representing the patient's cardiac demand, and a computing unit for determining the myocardial energy balance by calculating energy consumed by the myocardium for both an external dynamic work for pumping blood into a vascular system, and an internal static work of the myocardium. Volume and time based measurements are used, and in one embodiment, volumes are estimated and volume ratios are calculated from volume estimates. In another embodiment, volumes are estimated from bioimpedance measurements. A further aspect is a rate adaptive pacemaker, wherein the maximum pacing rate is determined from the myocardial energy balance such that the energy supplied to the myocardium approximately equals the energy consumed by the myocardium for both an external dynamic work for pumping blood into a vascular system and an internal static work of the myocardium. | 1. A device for monitoring a cardiac pacing rate, the device comprising: a measuring unit adapted for interaction with a subject for obtaining an electrical signal representing cardiac demand of said subject; anda computing unit supplied with said signal for determining a relation between an energy supplied to a myocardium and an energy consumed by the myocardium, and calculating a maximum pacing rate from said relation so that said maximum pacing rate corresponds to a myocardial energy balance where said energy supplied to the myocardium equals to said energy consumed by the myocardium for both an external dynamic work for pumping blood into a vascular system, and an internal static work of the myocardium, wherein a ratio of a diastolic pressure and a systolic pressure is calculated and said ratio is used to calculate said energy consumed by the myocardium. 2. The device as in claim 1, wherein said energy consumed by the myocardium is calculated from a stroke volume SV and an end-systolic volume Ves. 3. The device as in claim 1, wherein said myocardial energy balance is calculated from a stroke volume SV, an end-systolic volume Ves, and a duration of a diastole tdiast. 4. The device as in claim 1, wherein said internal static work of the myocardium is calculated from its unloaded volume V0 and end-systolic volume Ves. 5. The device as in claim 1, wherein said electrical signal is obtained through bioimpedance measurement. 6. The device as in claim 2, wherein said volumes are estimated based on bioimpedance measurement. 7. The device as in claim 1, wherein said ratio of said diastolic pressure and said systolic pressure is calculated from a ratio of a systolic volume and a diastolic volume in a blood vessel. 8. The device as in claim 7, wherein said blood vessel is an aorta or an artery. 9. The device as in claim 8, wherein said blood vessel is subclavian artery. 10. The device as in claim 7, comprising at least a pair of electrodes on said blood vessel for measuring of a systolic bioimpedance and a diastolic bioimpedance and calculating said ratio based on said systolic bioimpedance and said diastolic bioimpedance. 11. The device as in claim 2, comprising ultrasound sensors for volume measurements. 12. The device as in claim 3, wherein said duration of a diastole tdiast is measured from an electrocardiogram. 13. The device as in claim 3, wherein said duration of a diastole tdiast is measured from an impedance cardiogram. 14. The device as in claim 10, said pair of electrodes arranged to measure and to estimate a velocity of a pulse wave propagation. 15. The device as in claim 10, said pair of electrodes arranged to measure and to estimate an elasticity of arteries. 16. The device as in claim 10, said pair of electrodes arranged to characterize a blood flow into peripheral arteries and to estimate their hydraulic resistance to the blood flow. 17. The device as in claim 5, comprising at least one intracardiac implanted electrode fo... | Smartimplant OÜ | A61B005/08 A61N001/08 A61N001/365
US07878207 | B2 | 2011-02-01 | Locating an implanted object based on external antenna loading | 2005-07-20 | In general, the invention is directed to techniques for locating an implanted object using an external antenna. The implanted object may be, for example, an internal antenna that facilitates recharging of and/or communication with an implantable medical device. An external device coupled to the external antenna drives the antenna with a plurality of waveforms. Asymmetry in the loading profile of the external antenna when it is driven by the plurality of waveforms allows the external device or another device to determine the location of the implanted object relative to the external antenna. The external device or other device may provide information to a user based on the determined location of the implanted object relative to the external antenna, such as information to help a user position the external antenna with respect to an internal antenna in embodiments in which the implanted object is an internal antenna. | 1. A method comprising: generating a first waveform in an external device;driving an external antenna with the first waveform, wherein the external antenna has an asymmetrical loading profile;measuring a first value of an electrical parameter associated with the external antenna while the external antenna is driven with the first waveform;generating a second waveform in the external device, the second waveform having at least one different waveform parameter than the first waveform;driving the external antenna with the second waveform;measuring a second value of the electrical parameter while the external antenna is driven with the second waveform; anddetermining a location of an implanted object relative to the external antenna based on a difference between the measured first and second values of the electrical parameter, wherein the difference between the first and second values of the electrical parameter is caused at least in part by the asymmetrical loading profile of the external antenna. 2. The method of claim 1, wherein the first and second waveforms have at least one of different amplitudes, different frequencies, different pulse widths, or different shapes. 3. The method of claim 1, wherein the implanted object comprises at least one of an implantable medical device, an antenna associated with an implantable medical device, or a fill port associated with an implantable medical device. 4. The method of claim 1, further comprising providing an indication of the location of the implanted object relative to the external antenna to a user. 5. The method of claim 4, wherein providing an indication to the user comprises providing at least one of a direction or a distance. 6. The method of claim 1, wherein the external antenna is located at a first location when driven with the first and second waveforms, the method further comprising providing instructions to a user to move the external antenna to a second location that is closer to the implanted object than the first location based on the determined location of the implanted object. 7. The method of claim 1, wherein determining the location of the implanted object relative to the external antenna based on the first and second electrical parameter values comprises applying first and second equations to the first and second electrical parameter values, andwherein the first equation defines a first position variable as a function of the electrical parameter, and the second equation defines a second position variable as a function of the electrical parameter. 8. The method of claim 7, wherein the first position variable comprises a radial distance, and the second position variable comprises a depth. 9. The method of claim 7, further comprising: generating a third waveform in the external device, the third waveform having at least one different waveform parameter than the first and second waveforms;driving the external antenna with the third waveform; andmeasuring a third value of the elec... | Medtronic, Inc. | A61B019/00 A61B005/00
US07892178 | B1 | 2011-02-22 | Monitoring device for an interactive game | 2010-07-17 | A system and method (1000) for an interactive game is disclosed herein. The system (1000) preferably includes monitoring device (20) monitoring the vital signs of a user, an interface (1115), a game console (1010) and an accessory (1020). The monitoring device (20) is preferably an article (25) having an optical sensor (30) and a circuitry assembly (35), and a pair of straps (26a and 26b). The monitoring device (20) preferably provides for the display of the following information about the user: pulse rate; blood oxygenation levels; calories expended by the user of a pre-set time period; target zones of activity; time; distance traveled; and/or dynamic blood pressure. The article (25) is preferably a band worn on a user's wrist, arm or ankle. | 1. A system for a playing an interactive video game, the system comprising: a game console configured to operate an interactive video game, the interactive video game having a representation of a player of the interactive video game;a video monitor for displaying the interactive video game operated on the game console;an interface device in communication with the game console to allow the player of the interactive video game to act upon an object in the interactive video game by movements of the interface device; anda monitoring device for monitoring at least a real-time heart rate of the player, the monitoring device in wireless communication with the interface device or the game console to allow the real-time heart rate of the player to be utilized in the interactive video game, the monitoring device comprising an article having an interior surface and an exterior surface, a processor disposed within the article, an optical sensor positioned on the interior surface of the article, the optical sensor measuring blood flow through an artery of an arm of the player, a motion sensor disposed within the article, a wireless transceiver for transmitting information from the monitoring device to the interface device or the game console, the motion sensor and the optical sensor electrically connected to the processor, the processor configured to determine the real-time heart rate of the player based on the blood flow through the artery of the player and the processor configured to detect motion of the player based on a signal from the motion sensor, the optical sensor comprising a LED and a photodetector, and a power source for providing power to the processor, the optical sensor and the motion sensor;wherein the processor determines a real-time heart rate of the user by generating a first heart rate value from a first filtered pulse data value which is generated from subtracting an average value of a second set of sample data values from an average value of a first set of sample data values, and wherein the processor uses a cascade adaptive resonant filter to generate subsequent heart rate values;wherein an activity of the representation of the player on the video monitor is partially controlled by the real-time heart rate of the player monitored by the optical sensor and the motion of the player detected by the motion sensor. | Impact Sports Technologies, Inc. | A61B005/02 A61B005/103 A63F013/06
US07881777 | B2 | 2011-02-01 | Method and apparatus for performing intra-operative angiography | 2007-11-28 | Method for assessing the patency of a patient's blood vessel, advantageously during or after treatment of that vessel by an invasive procedure, comprising administering a fluorescent dye to the patient; obtaining at least one angiographic image of the vessel portion; and evaluating the at least one angiographic image to assess the patency of the vessel portion. Other related methods are contemplated, including methods for assessing perfusion in selected body tissue, methods for evaluating the potential of vessels for use in creation of AV fistulas, methods for determining the diameter of a vessel, and methods for locating a vessel located below the surface of a tissue. | 1. A device for monitoring movement of a fluorescent dye contained in blood moving through a cardiovascular bypass graft and vasculature connected to the graft of a subject during a surgical procedure, the device comprising: an illumination source emitting radiation of which a portion is absorbable by the moving fluorescent dye;a camera having an image capturing device configured to capture images of fluorescent radiation emitted from the fluorescent dye moving through the bypass graft and connected vasculature while the subject's heart is beating, at an acquisition rate of at least 15 angiographic images per second, the captured images including at least an optimal image of fluorescent dye initially moving through the bypass graft and connected vasculature, as well as images leading and trailing the optimal image;wherein the fluorescent dye has a peak absorption in the range of 800 to 850 nm; anda display device that transforms the captured images into viewable images. 2. The device according to claim 1, wherein the camera captures images at an increased acquisition rate of at least 30 images per second and the dye is indocyanine green. 3. The device according to claim 1, wherein the illumination source and the camera are positioned with respect to one another such that an angle between an optical axis of the camera and an axis of a radiant energy beam produced by the illumination source is less than 85°. 4. The device according to claim 1, wherein the illumination source is a laser, and the device further comprises optics positioned to diverge a radiant energy beam produced by the laser. 5. The device according to claim 4, wherein the laser and the camera are positioned with respect to one another such that an angle between an optical axis of the camera and an axis of the radiant energy beam from the laser is between 20° and 70°. 6. The device according to claim 4, wherein the optics are adjustable, permitting variation in a field of illumination. 7. The device according to claim 4, further comprising a bandpass filter positioned relative to the radiant energy beam to limit the radiant energy beam to one of a substantially uniform wavelength. 8. The device according to claim 4, wherein the laser is pulsed and synchronized with the acquisition rate of the camera. 9. The device according to claim 1, further comprising a bandpass filter positioned to prevent the camera from capturing radiation other than that emitted by the fluorescent dye. 10. The device according to claim 1, wherein the camera comprises a lens system for magnifying a field of view. 11. The device according to claim 10, wherein the lens system is capable of being engaged by the laser to correspondingly adjust a field of illumination provided by the laser as a function of the field of view of the camera. 12. The device according to claim 4, further comprising a distance sensor that provides distance information from at least one of the camera and the laser to the cardio... | National Research Council of Canada | A61B005/00
US07996187 | B2 | 2011-08-09 | Method and system for health monitoring | 2008-01-02 | A personal health monitor, including: (a) a physiological data input device operative to gather physiological data; and (b) a multi-purpose personal data accessory, whereas the multi-purpose personal data accessory is adapted to execute health monitoring software such as to enable the multi-purpose personal data accessory to receive the physiological data, process the physiological data to provide processed physiological data and control a long range transmission of the processed physiological data to a remote entity; wherein the physiological data input device controls a transmission of the physiological data to the multi-purpose personal data accessory. | 1. A personal health monitor comprising: a physiological data input device operative to gather physiological data; and a multi-purpose personal data accessory, whereas the multi-purpose personal data accessory is adapted to execute health monitoring software such as to enable the multi-purpose personal data accessory to receive the physiological data, process the physiological data to provide processed physiological data and control a long range transmission of the processed physiological data to a remote entity; wherein the physiological data input device controls a transmission of the physiological data to the multi-purpose personal data accessory; wherein the physiological data input device comprises a receiver that is maintained in a low power mode during idle periods that are determined by the physiological data input device. 2. The personal health monitor of claim 1 wherein the physiological data input device determines whether physiological data can be sent from the physiological data input device to the multi-purpose personal data accessory. 3. The personal health monitor of claim 2 wherein the physiological data input device generates an alert in response to a determination that a transmission of physiological data to the multi-purpose personal data accessory is expected to fail. 4. A personal health monitor comprising: a physiological data input device operative to gather physiological data and a multi-purpose personal data accessory, whereas the multi-purpose personal data accessory is adapted to execute health monitoring software such as to enable the multi-purpose personal data accessory to receive the physiological data, process the physiological data to provide processed physiological data and control a long range transmission of the processed physiological data to a remote entity; wherein the multi-purpose personal data accessory is adapted to receive physiological data input device related information and to provide, in response to the physiological data input device related information, a physiological data input device instruction to the physiological data input device; wherein the physiological data input device is adapted to send configuration information to the multi-purpose personal data accessory in response to the physiological data input device instruction. 5. A personal health monitor comprising: a physiological data input device operative to gather physiological data; and a multi-purpose personal data accessory, whereas the multi-purpose personal data accessory is adapted to execute health monitoring software such as to enable the multi-purpose personal data accessory to receive the physiological data, process the physiological data to provide processed physiological data and control a long range transmission of the processed physiological data to a remote entity; wherein the multi-purpose personal data accessory is adapted to receive physiological data input device related information and to provide, in response t... | Card Guard Scientific Survival Ltd. | G06F019/00
US20110028805 | A1 | 2011-02-03 | INFORMATION PROCESSING APPARATUS, METHOD, AND PROGRAM | 2010-07-16 | An information processing apparatus includes: reproduction means for reproducing 3D content; and detection means for detecting a biological signal associated with a viewer who is viewing the 3D content, wherein the reproduction means attenuates the 3D effect of the 3D content when the viewer's biological signal produced when the 3D content is being reproduced exceeds a threshold value. | 1. An information processing apparatus comprising: reproduction means for reproducing 3D content; anddetection means for detecting a biological signal associated with a viewer who is viewing the 3D content,wherein the reproduction means attenuates the 3D effect of the 3D content when the viewer's biological signal produced when the 3D content is being reproduced exceeds a threshold value. 2. The information processing apparatus according to claim 1, further comprising storage means for storing a standard value of the viewer's biological signal and viewer information representing the characteristics of the viewer, the standard value and the viewer information related to each other,wherein the reproduction means changes the threshold value based on a viewing pattern related to a pattern according to which at least one of video and audio signals representing the 3D content changes, the viewing pattern stored in advance on a 3D content storage medium on which the 3D content is stored. 3. The information processing apparatus according to claim 2, wherein the viewer information contains at least the age of the viewer, andwhen a plurality of viewers are viewing the 3D content and the viewer whose biological signal shows a value greater than the threshold value is of a predetermined age or under, the reproducing means places a priority on the viewer and attenuates the 3D effect accordingly. 4. The information processing apparatus according to claim 3, wherein the reproduction means further changes the threshold value over the period during which the 3D content is reproduced. 5. The information processing apparatus according to claim 4, wherein the reproduction means transmits information obtained by relating to one another the viewer information, the viewer's biological signal, and the time along which the 3D content is reproduced via a network. 6. An information processing method used with an information processing apparatus including reproduction means and detection means, the method comprising the steps of: reproducing 3D content by using the reproduction means;detecting a biological signal associated with a viewer who is viewing the 3D content by using the detection means; andattenuating the 3D effect of the 3D content by using the reproducing means when the viewer's biological signal produced when the 3D content is being reproduced exceeds a threshold value. 7. A program that instructs a computer to function as: reproduction means for reproducing 3D content; anddetection means for detecting a biological signal associated with a viewer who is viewing the 3D content,wherein the reproduction means attenuates the 3D effect of the 3D content when the viewer's biological signal produced when the 3D content is being reproduced exceeds a threshold value. 8. An information processing apparatus comprising: a reproduction unit configured to reproduce 3D content; anda detection unit configured to detect a biological signal associated with a viewer who is... | Sony Corporation | A61B005/00 H04N013/00 A61B005/024 A61B005/11 A61B005/026
US20110028807 | A1 | 2011-02-03 | Contact lens for collecting tears and detecting at least one analyte | 2010-07-08 | Utilization of a contact device placed on the eye in order to detect physical and chemical parameters of the body as well as the non-invasive delivery of compounds according to these physical and chemical parameters, with signals being transmitted continuously as electromagnetic waves, radio waves, infrared and the like. One of the parameters to be detected includes non-invasive blood analysis utilizing chemical changes and chemical products that are found in the conjunctiva and in the tear film. A transensor mounted in the contact device laying on the cornea or the surface of the eye is capable of evaluating and measuring physical and chemical parameters in the eye including non-invasive blood analysis. The system utilizes eye lid motion and/or closure of the eye lid to activate a microminiature radio frequency sensitive transensor mounted in the contact device. The signal can be communicated by wires or radio telemetered to an externally placed receiver. The signal can then be processed, analyzed and stored. Several parameters can be detected including a complete non-invasive analysis of blood components, measurement of systemic and ocular blood flow, measurement of heart rate and respiratory rate, tracking operations, detection of ovulation, detection of radiation and drug effects, diagnosis of ocular and systemic disorders and the like. | 1. A contact device for detecting at least one analyte in eye fluid, said contact device comprising a housing for obtaining eye fluid from the eye,an analyte detecting element of the housing for detecting the at least one analyte in the eye fluid from the eye,said housing including at least one microfluidic channel,said housing being less than 5.0 mm in thickness. 2. The contact device of claim 1, wherein said analyte detecting element is one of an enzyme, antibody, antigen, chromogen, a fluorescent element and an analyte detecting sensor. 3. The contact device as claimed in claim 1, wherein the housing further includes surface charges. 4. The contact device as claimed in claim 3, wherein the surface charges include a negatively charged perfluorinated ionomer. 5. The contact device as claimed in claim 3, wherein the housing includes at least one of a surface charge for repulsing analytes and a surface charge for increasing attraction to analytes of interest. 6. The contact device as claimed in claim 3, wherein the surface charges include at least one of a positive surface charge and a negative surface charge. 7. The contact device as claimed in claim 1, wherein the housing includes at least one coating for preferential selection of said at least one analyte. 8. The contact device as claimed in claim 1, wherein said housing includes a site for said at least one analyte. 9. The contact device as claimed in claim 8, wherein said site is coated with antibodies for generating a detectable signal after interaction with said at least one analyte. 10. The contact device as claimed in claim 1, wherein said analyte detecting element detects an analyte selected from the group consisting of antibodies, hormone, drugs, enzymes, nucleic acids, glucose, cholesterol, proteins, toxic compounds, amino acids, inflammatory mediators, hemoglobin, cancer markers, genetic markers, heart markers, electrolytes, gases, antigens, alcohol, urea, creatinine and infectious agents. 11. The contact device as claimed in claim 1, wherein said housing generates a detectable signal according to one of presence and amount of said at least one analyte in said eye fluid. 12. The contact device as claimed in claim 1, wherein the housing detects at least one hormone in the eye fluid. 13. The contact device as claimed in claim 12, wherein said housing detects a health status of a mammal. 14. The contact device as claimed in claim 1, wherein the housing detects ovulation. 15. The contact device as claimed in claim 1, wherein the analyte detecting element generates a detectable signal. 16. The contact device as claimed in claim 15, wherein the detectable signal includes optically detectable signals. 17. The contact device as claimed in claim 16, wherein the optically detectable signal includes at least one of fluorescent light, change in color and change of an amount of radiation of said analyte detecting element. 18. The contact device as claimed in claim 1, wherein the h... | - | A61B005/00
US08965075 | B2 | 2015-02-24 | System and method for predicting future fractures | 2010-05-13 | A method of predicting bone or articular disease in a subject includes determining one or more micro-structural parameters, one or more macroanatomical parameters or biomechanical parameters of a joint in the subject and combining at least two of the parameters to predict the risk of bone or articular disease. | 1. A method of evaluating a bone condition in a subject using a computer system, comprising: obtaining an image of a bone of the subject at an initial time,placing one or more reference markers on the image in their home states using an automatic algorithm,selecting one or more regions of interest from the image based on the one or more reference markers placed on the image,obtaining one or more parameters from the one or more regions of interest, wherein the one or more parameters are selected from micro-structural parameters, macroanatomical parameters and biomechanical parameters,combining at least two parameters by univariate, bivariate or multivariate statistical analysis, anddetermining the bone condition using the one or more parameters, which includes the combined parameter. 2. The method of claim 1, wherein the one or more reference markers are geometric markers. 3. The method of claim 1, wherein the location of the one or more reference markers is based on anatomical landmarks. 4. The method of claim 1, wherein the one or more reference markers are placed on the image using user a-priori input. 5. The method of claim 1, wherein the at least two parameters include one or more microstructural parameters and one or more macroanatomical parameters. 6. The method of claim 1, wherein the at least two parameters include one or more microstructural parameters and one or more biomechanical parameters. 7. The method of claim 1, wherein the bone is in a region selected from the group consisting of leg, knee, hip, ankle, spine, radius and arm. 8. The method of claim 1, wherein the bone condition is osteoporosis, bone strength or fracture risk. 9. The method of claim 1, wherein the image is selected from the group consisting of an X-ray image, a CT image, an ultrasound image and an MRI. 10. The method of claim 1, wherein the step of determining includes comparing the one or more parameters to data derived from a reference database of known bone condition parameters. 11. The method of claim 1, wherein the reference database includes bone condition parameters derived from images from bones of one or more known conditions. 12. The method of claim 11, wherein the one or more known conditions include osteoporosis, a fracture or bone failure load. 13. The method of claim 1, further comprising: selecting a candidate agent based on the bone condition of the subject,administering the candidate agent to the subject,obtaining a second image of the bone of the subject at a later time,selecting one or more regions of interest from the second image using the one or more reference markers,obtaining a second set of one or more parameters from the one or more regions of interest from the second image, anddetermining the bone condition at the later time using the one or more parameters. 14. A method of evaluating a bone condition in a subject using a computer system, comprising: obtaining an image of a bone of the subject,placing one or more reference ma... | ImaTx, Inc. | G06K009/00 G06T007/00
US20110046497 | A1 | 2011-02-24 | SENSOR PLATFORM FOR RESPIRATORY GAS ANALYSIS | 2010-08-19 | An appliance is disclosed for measuring at least one gas analyte in exhaled air. In at least one embodiment, the appliance includes an inlet opening for the introduction of exhaled air (mouthpiece), at least one measuring chamber for receiving a sensor unit, and a conduit which provides a fluid connection from the opening to the measuring chamber, and wherein, depending on the gas analyte that is to be measured, a sensor unit with a suitable gas sensor can be introduced into the receiver. | 1. An appliance for measuring at least one gas analyte in exhaled air, said appliance comprising: an inlet opening for the introduction of exhaled air;at least one measuring chamber for receiving at least one sensor unit; anda conduit to provide a fluid connection from the inlet opening to the measuring chamber, wherein, depending on the gas analyte that is to be measured, a sensor unit with a gas sensor is introduceable into the receiver. 2. The appliance as claimed in claim 1, wherein the at least one measuring chamber includes a plurality of measuring chambers. 3. The appliance as claimed in claim 1, further comprising a valve for selectively connecting the at least one measuring chamber to the inlet opening. 4. The appliance as claimed in claim 1, further comprising at least one gas-conditioning device. 5. The appliance as claimed in claim 1, further comprising a partitioning device for partitioning the incoming exhaled air and, optionally, for delivering a portion to the at least one measuring chamber. 6. The appliance as claimed in claim 1, further comprising a device for delivering a defined volume of the exhaled gas to a defined measuring chamber. 7. The appliance as claimed in claim 1, further comprising a particle filter. 8. The appliance as claimed in claim 1, further comprising: a one-way valve, to prevent air exhaled into the appliance from being sucked out and re-inhaled by a user. 9. The appliance as claimed in claim 4, wherein the at least one gas analyte is nitrogen monoxide, and wherein the device for gas conditioning is a device for oxidation of nitrogen monoxide to nitrogen dioxide. 10. The appliance as claimed in claim 9, wherein the gas sensor is chosen from the group comprising NO2-sensitive FET sensor, IR sensor, metal oxide sensor. 11. The appliance as claimed in claim 1, further comprising: a calibration gas device for supplying the measuring chamber with at least one calibration gas. 12. The appliance as claimed in claim 1, further comprising: a temperature control device for controlling the temperature of at least one measuring chamber. 13. The appliance as claimed in claim 1, wherein at least one measuring chamber is thermally insulated. 14. The appliance as claimed in claim 2, further comprising a valve for selectively connecting a subgroup of the plurality of measuring chambers to the inlet opening. 15. The appliance as claimed in claim 2, further comprising a partitioning device for partitioning the incoming exhaled air and, optionally, for delivering a portion to a measuring chamber of the plurality of measuring chambers. 16. The appliance as claimed in claim 1, wherein the inlet opening is a mouthpiece. | - | A61B005/097
US08478395 | B2 | 2013-07-02 | Method of measuring electrical resistance value of corneal trans-epithelium | 2009-03-03 | The invention provides a method for evaluating a corneal disorder quantitatively and is applicable to living eyes. In particular, the invention provides a method for measuring a corneal transepithelial electric resistance, which method comprises: (1) a step of placing a first electrode on the cornea and a second electrode on the conjunctiva; and (2) a step of flowing an electric current between the first electrode and the second electrode to measure the electric resistance. The invention also provides a device for measuring a corneal transepithelial electric resistance value. | 1. A method for measuring a corneal transepithelial electric resistance value, comprising (1) placing electrodes on a cornea and a conjunctiva, wherein the electrodes are connected to a voltage-electric resistance value measuring device, and wherein the device flows an electric current through corneal epithelium, and(2) measuring the corneal transepithelial electric resistance value based on the electric current flowing through the corneal epithelium. 2. A method for measuring a corneal transepithelial electric resistance value comprising: (1) a step of preparing a voltage-electric resistance value measuring device, and a first electrode and a second electrode which are connected to the device;(2) a step of placing the first electrode on the cornea and the second electrode on the conjunctiva; and(3) a step of measuring the corneal transepithelial electric resistance value by flowing an electric current between the first electrode and the second electrode and through corneal epithelium to measure an electric resistance. 3. The method according to claim 1, wherein a periphery of the electrode placed on the cornea is insulated. 4. The method according to claim 1 which is carried out on a living eye. 5. A device for measuring a corneal transepithelial electric resistance on a living eye, which comprises: (1) a first electrode suitable to be placed on a cornea,(2) a second electrode suitable to be placed on a conjunctiva, ˜d(3) an insulator that insulates the first electrode, and(4) a voltage-electric resistance value measuring device connected to the first electrode and the second electrode, wherein the device is capable of flowing an electric current through corneal epithelium,wherein the first electrode and the insulator are integrated, andwherein the insulator is made of silicone rubber, and the insulator has a contact surface to the cornea having the same curvature as that of the cornea. 6. The device according to claim 5, wherein the first electrode and the insulator are integrated via a conductor. 7. The device according to claim 5, wherein the second electrode is further integrated. 8. The device according to claim 5, wherein the second electrode is formed into a sheet shape. 9. The device according to claim 6, wherein the conductor is a gel of hyaluronic acid or atelocollagen. 10. The method according to claim 3 which is carried out on a living eye. 11. The method according to claim 2, wherein a periphery of the electrode placed on the cornea is insulated. 12. The method according to claim 2 which is carried out on a living eye. | Nagasaki University Tohoku University | A61B005/053
US07878979 | B2 | 2011-02-01 | Methods and systems for determining dynamic hyperinflation | 2006-05-19 | This invention provides methods and systems for non-invasively determining the presence (and amount) or absence of dynamic hyperinflation in a subject. The invention is based on a novel combination of respiratory parameters that can be measured in a way that is non-invasive and unobtrusive to the subject. Dynamic hyperinflation is often a significant factor in the quality of life of patients suffering from a variety of obstructive pulmonary diseases, and this invention permits simple, routine tracking and management of dynamic hyperinflation in affected patients. | 1. A computer-implemented method for assessing dynamic hyperinflation (DH) in a monitored subject, comprising: receiving respiratory data monitored with a physiological monitoring system during a current period of monitoring the subject, said monitoring system having at least one sensor, said respiratory, data including respiratory volumes during rib cage(RC) expansion and contraction;determining from said received data at one or more times during said current monitoring period a first parameter (M % RC) with processor means, said first parameter comprising median % rib cage, said first parameter (M % RC) reflecting relative contributions of expansions and contractions of the rib cage (RC) to tidal volumes (VT) of the subject;determining from said received data at one or more times during said current monitoring period a second parameter (MqdEELV) with said processor means, said second parameter (MqdEELV) comprising median cumulative changes in end-expiratory lung volume; andassessing DH in the subject during said current monitoring period in dependence on said determined MqdEELV and M % RC parameters. 2. The computer-implemented method of claim 1, wherein said current monitoring period comprises a plurality of sub-periods, and wherein said M % RC parameter and said MqdEELV parameter are determined for each of said plurality of sub-periods only from received respiratory data from that sub-period. 3. The computer-implemented method of claim 1, wherein DH is present during said current monitoring period if M % RC parameter values increase and if MqdEELV parameter values decrease during said current monitoring period. 4. The computer-implemented method of claim 1, wherein DH is present during said current monitoring period in comparison with a previous monitoring period if M % RC parameter values during said current monitoring period are greater than during said, previous monitoring period and if said MqdEELV parameter values during said current monitoring period are less than during said previous monitoring period. 5. The computer-implemented method of claim 4, wherein said previous monitoring period comprises a baseline period during which DH is known to be absent. 6. The computer-implemented method of claim 1, wherein said determination of said M % RC parameter comprises dividing said respiratory volumes during rib cage (RC) expansion and contraction by tidal volume (VT). 7. The computer-implemented method of claim 6, wherein said received data comprises rib cage (RC) sizes at one or more selected positions on the subject, and wherein said respiratory volumes during rib cage (RC) expansion and contraction are determined in dependence on one or more rib cage (RC) sizes. 8. The computer-implemented method of claim 1, wherein determination of said MqdEELV parameter comprises determining a breath-by-breath change in expiratory lung volume (EELV) for a first period of time,determining a cumulative change in said EELV in dependence on plurality ... | adidas AG | A61B005/08
US08814808 | B2 | 2014-08-26 | Body fluid sampling device | 2010-09-10 | Body fluid sampling device comprising a skin-piercing element having a collection zone for receiving body fluid. In one embodiment, the collection zone includes a plurality of holes, e.g., 3 or more holes, or 50 to 100 holes. The hole size may be small, e.g., having a diameter of 0.01 to 0.5 mm. The holes may or may not extend through the lancet. In some embodiments, the holes have a depth of 50 to 500 μm. In other embodiments, the collection zone is porous, or is a roughened area. The collection zone can take up a very small volume of body fluid, e.g., 3 to 10 nL, in a very short time period, e.g., less than 0.5 seconds. In other embodiments, the device further comprises a fluid receiving means spaced apart from the collection zone so that body fluid in the collection zone will not contact the fluid receiving means initially. | 1. A body fluid system, comprising: a skin-piercing element having a sharpened end for piercing tissue and a collection zone for receiving body fluid from the tissue, wherein the thickness of the skin-piercing element is below 500 μm and the collection zone comprises 3 or more blind holes open to the environment for receiving the body fluid; anda test zone for receiving body fluid from the collection zone of said skin-piercing element. 2. The system according to claim 1, wherein the sample volume of the collection zone is less than 0.5 μL (microliters). 3. The system according to claim 1, wherein the thickness of the skin-piercing element is below 250 μm (micrometers). 4. Device according to claim 1, wherein the collection zone comprises 50 to 100 blind holes. 5. A testing system comprising: a lancet having a sharpened end adapted to penetrate the skin surface and obtain a sample of body fluid, the lancet including a plurality of blind holes open to the environment and forming a collection zone, wherein the diameter of each hole is 0.01 mm (millimeters) to 0.5 mm (millimeters); anda test zone for receiving body fluid from the collection zone of said lancet. 6. The system of claim 5 wherein the depth of each hole is 50 to 500 μm (micrometers). 7. The apparatus of claim 5 wherein the plurality of holes includes 50 to 100 blind holes. 8. The system of claim 5 wherein the sample volume of the collection zone is less than 0.5 μL (microliters). 9. The apparatus of claim 5 wherein the sample volume of the collection zone is 3 to 10 nL (nanoliters). 10. A testing system comprising: a lancet having a sharpened end adapted to penetrate a skin surface to obtain a sample of body fluid, the lancet including a collection zone adapted to hold a sample of bodily fluid as the lancet is withdrawn from the skin, wherein the collection zone includes a plurality of blind holes open to the environment and having diameters of 0.01 mm to 0.5 mm; anda test zone for receiving body fluid from the collection zone of said lancet. 11. The apparatus of claim 10 wherein the collection zone includes 50 to 100 blind holes. 12. The apparatus of claim 10 wherein the sample volume of the collection zone is 3 to 10 nL (nanoliters). 13. The system of claim 10 wherein the collection zone is configured and adapted to fill with a sample of body fluid in less than 0.1 seconds. 14. The system of claim 10 wherein the collection zone is configured and adapted to fill with a sample of body fluid in 1 to 10 ms (milliseconds). 15. The sampling system of claim 1 and further comprising: a meter with a detection unit for receiving signals from said test zone to determine the presence and/or concentration of an analyte in the body fluid based on the received signals. 16. The system of claim 5 and further comprising: a meter with a detection unit for receiving signals from said test zone to determine the presence and/or concentration of an analyte in the body fluid based on the receiv... | Roche Diagnostics Operations, Inc. | A61B005/151 A61B005/15 A61B005/145
US08845535 | B2 | 2014-09-30 | Device for detecting a medical condition or disease | 2009-02-23 | The invention relates to a capsule or chip or sensor comprising a marker/detector and signalling device/method associated with the development of a medical condition/disease and to its use. This In Situ Lab On a Chip Signalling device (ISLOGS device) is used for detecting a medical condition/disease in situ or in vivo (inside the body), for example by swallowing it for testing in the digestive tract or by putting it into the vagina, mouth or nose. All reactions takes place in the body and when the test result is positive the device will notify the subject by colorizing the tested body fluid or biomaterial (for example faeces, urine or saliva) or the device itself. The detection is made outside of the body when the biomaterial or ISLOCS device has left the body. | 1. A device for detecting a marker associated with a medical condition or disease in a subject, wherein the device is capable of entering the subject and comprises (a) a detector which is able to specifically react in situ to the marker present in a body fluid or biomaterial of the subject (b) a chamber wherein a visualisable substance is received, and (c) a signalling unit arranged and constructed to emit the visualisable substance from the chamber upon detection of the marker into the body fluid, the visualisable substance being visible outside the subject, wherein the device additionally comprises an extraction unit and/or a purification unit for extracting and/or purifying the marker from the body fluid or biomaterial of the subject. 2. The device according to claim 1, wherein the device further comprises a swallowable housing for entering a subject. 3. The device according to claim 1, wherein the device further comprises a housing for leaving a subject. 4. The device according to claim 1, wherein the detector comprises an automated miniature analysis system. 5. The device according to claim 1, wherein the signalling unit comprises a release mechanism for releasing the substance upon detection. 6. The device according to claim 1, wherein the visualisable substance is a dye. 7. The device according to claim 6, wherein the dye has a predetermined colour corresponding with the detection of a specific marker. 8. The device according to claim 1, wherein the visualisation of the detection is seen in a colouring of the body fluid or biomaterial which has been in contact with the device or in a colouring of the device itself. 9. The device according to claim 1, wherein the detector is able to specifically react in the digestive tract and/or vagina of a subject. 10. The device according to claim 9, wherein the device is an ingestible device and/or is resistant against degradation that occur within the digestive tract. 11. The device according to claim 1, wherein the medical condition or disease is a medical condition or disease of the digestive tract or vagina and/or detectable in the digestive tract or vagina. 12. The device according to claim 11, wherein the medical condition or disease of the vaginal tract is a HPV cancer. 13. The device according to claim 1, wherein the medical condition or disease of the digestive tract is a cancer, a colon cancer or a stomach cancer. 14. The device according to claim 1, wherein the subject is a human being. 15. The device according to claim 1, wherein the chamber has a volume of 1-100 microliter. 16. The device according to claim 1, wherein the device comprises two or more detectors for different markers. 17. The device according to claim 16, wherein upon detection of a combination of markers the signalling unit emits the visualisable substance. 18. The device according to claim 1, comprising multiple chambers, wherein a different visualisable substance is received in each of the chambers. 19. ... | Herbert Michael Pinedo Roderik Adriaan Kraaijenhagen Universiteit Twente | A61B005/1455 A61B005/145 A61B005/07 A61B010/00 A61B005/00
US09687656 | B2 | 2017-06-27 | Arterial blood pressure monitoring devices, systems and methods for use while pacing | 2009-12-14 | Provided herein are implantable systems, and methods for use therewith, for monitoring a patient's arterial blood pressure while a patient's heart is being paced. A signal (e.g., PPG or IPG signal) indicative of changes in arterial blood volume remote from the patient's heart is obtained using a sensor or electrodes that are implanted remote from the patient's heart. One or more metrics indicative of pulse arrival time (PAT) are determined, where each metric can be determined by determining a time from a paced cardiac event to one or more predetermined features of the signal indicative of changes in arterial blood volume. Based on at the metric(s) indicative of PAT, arterial blood pressure is estimated, which can include determining values indicative of systolic blood pressure, diastolic blood pressure, pulse pressure and/or mean arterial blood pressure, and/or changes in such values. | 1. For use with an implantable system, a method for monitoring a patient's arterial blood pressure, the method comprising: (a) producing a ventricular pacing signal, which includes a plurality of ventricular pacing pulses, and pacing a patient's heart using the ventricular pacing signal and at least one electrode implanted within or on the patient's heart;(b) using an implanted sensor or implanted electrodes, remote from the patient's heart, to obtain a signal indicative of changes in arterial blood volume remote from the patient's heart while the patient's heart is being paced;(c) detecting one or more predetermined features of the signal indicative of changes in arterial blood volume remote from the patient's heart obtained at step (b);(d) determining one or more metrics indicative of pulse arrival time (PAT), by determining a time from a ventricular pacing pulse of the ventricular pacing signal to one or more of the predetermined features of the signal indicative of changes in arterial blood volume remote from the patient's heart detected at step (c); and(e) estimating the patient's arterial blood pressure based on at least one of the one or more metrics indicative of PAT. 2. The method of claim 1, further comprising: determining a value indicative of electromechanical delay (EMD) between delivery of pacing at step (a) and a mechanical cardiac contraction resulting from the pacing; andwherein step (d) comprises determining at least one of the one or more metrics indicative of PAT also based on the value indicative of EMD. 3. The method of claim 2, wherein step (d) comprises determining at least one of the metrics indicative of PAT by determining a time from a ventricular pacing pulse of the ventricular pacing signal to a predetermined feature of the signal indicative of changes in arterial blood volume remote from the patient's heart, detected at step (c), minus the determined value indicative of EMD. 4. The method of claim 2, further comprising adjusting the pacing at step (a) to reduce variance of the value indicative of electromechanical delay (EMD). 5. The method of claim 1, wherein the signal indicative of changes in arterial blood volume remote from the patient's heart obtained at step (b) comprises a photoplethysmography (PPG) signal or an impedance plethysmography (IPG) signal, and the one or more predetermined features of the PPG or IPG signal detected at step (d) is/are selected from the group consisting of: minimum amplitude of the PPG or IPG signal;maximum upward slope of the PPG or IPG signal;maximum amplitude of the PPG or IPG signal;dicrotic notch of the PPG signal or IPG;maximum downward slope of the PPG or IPG signal prior to the dicrotic notch; andmaximum downward slope of the PPG or IPG signal following the dicrotic notch. 6. The method of claim 5, wherein step (d) includes: filtering the PPG or IPG signal to reduce effects of respiratory noise, motion artifacts and baseline drift;grouping a plurality of cycles of the P... | Pacesetter, Inc. | A61B005/02 A61N001/365 A61B005/0215 A61B005/0295 A61B005/1459 A61N001/08 G03F007/00 B82Y010/00 B82Y040/00 A61B005/021 A61B005/053 A61B005/00
US20110004083 | A1 | 2011-01-06 | Method and Device for Measuring Blood Flow Velocity | 2008-04-24 | A method for measuring blood flow velocity comprises the steps of: placing a first and second blood vessel signal detectors on a body of a person to be measured in such a manner that the first and second blood vessel signal detectors are located a predetermined distance from each other; using the first and second blood vessel signal analyzers to record the blood signal at a predetermined time interval; setting an interval of time from the moment a specific blood vessel signal appears in a record of the first blood vessel signal analyzer to the moment the specific blood vessel signal appears in a record of the second blood vessel signal analyzer to be a predetermined time period; and dividing a value of the predetermined distance by a value of the predetermined time period can get a blood flow velocity. | 1. A method for measuring blood flow velocity, comprising the steps of: placing a first and second blood vessel signal detectors on a body of a person to be measured, wherein the first blood vessel signal detector is electrically connected to a first blood vessel signal analyzer, the second blood vessel signal detector is electrically connected to a second blood vessel signal analyzer, and the first and second blood vessel signal detectors are located a predetermined distance from each other;using the first and second blood vessel signal analyzers to record the blood signal of the person to be measured at a predetermined time interval;setting an interval of time from the moment a specific blood vessel signal appears in a record of the first blood vessel signal analyzer to the moment the specific blood vessel signal appears in a record of the second blood vessel signal analyzer to be a predetermined time period; anddividing a value of the predetermined distance by a value of the predetermined time period can obtain a blood flow velocity between the first and second blood vessel signal detectors. 2. The method for measuring blood flow velocity as claimed in claim 1, wherein the blood flow velocity obtained is an average blood flow velocity between the first and second blood vessel signal detectors. 3. The method for measuring blood flow velocity as claimed in claim 1, wherein the first and second blood vessel signal detectors are connected by a connecting member. 4. The method for measuring blood flow velocity as claimed in claim 2, wherein the connecting member is a length-adjustable link rod or telescopic rod. 5. The method for measuring blood flow velocity as claimed in claim 1, wherein the first and second blood vessel signal detectors are of invasive or non-invasive type. 6. The method for measuring blood flow velocity as claimed in claim 1, wherein the first and second blood vessel signal detectors are of clamp or loop type. 7. The method for measuring blood flow velocity as claimed in claim 1, wherein the first and second blood vessel signal analyzers are used to analyze signals produced by blood flow and selected from the group consisting of oxy-meter, oxyhemoglobin analyzer, deoxyhemoglobin analyzer, carboxyhemoglobin analyzer, methemoglobin analyzer, and carbon dioxide meter. 8. The method for measuring blood flow velocity as claimed in claim 1, wherein the first and second blood vessel signal analyzers are used to analyze wave signals of blood. 9. The method for measuring blood flow velocity as claimed in claim 1, wherein a light beam of a predetermined wavelength is projected into blood vessel of the person to be measured to enable a specific type of blood cells or blood ingredient to produce an optical rotation signal which is used as a blood vessel signal value of the person to be measured and received by the blood vessel signal analyzers. 10. A device for measuring blood flow velocity comprising: a controller;a first blood v... | - | A61B005/0265 A61B005/1455 A61B006/00
US20110034787 | A1 | 2011-02-10 | Device for Interstitial Fluid Extraction, Production Process Thereof and Analyzing Process of Interstitial Fluid Using the Device | 2010-08-03 | A device for interstitial fluid extraction, having a base material formed from a synthetic resin film, a pressure sensitive adhesive layer, a hydrogel layer formed from at least one hydrophilic polymer selected from the group consisting of polyvinyl alcohol and polyvinyl pyrrolidone, and a release layer, wherein the hydrogel layer has an area of a size that the pressure sensitive adhesive layer is exposed from around the hydrogel layer, does substantially not contain a sodium ion and causes no water separation. | 1. A device for interstitial fluid extraction, comprising a base material formed from a synthetic resin film;a pressure sensitive adhesive layer arranged on one surface of the base material;a hydrogel layer arranged on the surface of the pressure sensitive adhesive layer and formed from at least one hydrophilic polymer selected from the group consisting of polyvinyl alcohol and polyvinyl pyrrolidone; anda release layer covering exposed surfaces of both pressure sensitive adhesive layer and hydrogel layer, whereinthe hydrogel layer has an area of a size that the pressure sensitive adhesive layer is exposed from around the hydrogel layer, does substantially not contain a sodium ion and causes no water separation when measured according to the measuring method described in the present description. 2. The device for interstitial fluid extraction according to claim 1, wherein the hydrogel layer further contains an osmotic pressure control agent composed of at least one compound selected from the group consisting of potassium chloride, potassium phosphate, magnesium chloride, magnesium phosphate, calcium chloride, calcium phosphate, amino acids, urea, acetic acid, ammonia, tricine, thiamine, riboflavin, nicotinic acid amide, pyridoxine, cyanocobalamine and ascorbic acid in a proportion that an osmolarity falls within a range of from 0.05 to 0.94 osmoles. 3. The device for interstitial fluid extraction according to claim 1, wherein the hydrogel layer is formed by irradiating a coating film of an aqueous solution of the hydrophilic polymer with radiation to crosslink the hydrophilic polymer. 4. The device for interstitial fluid extraction according to claim 1, wherein the hydrogel layer is obtained by irradiating a coating film of an aqueous solution of the hydrophilic polymer, said aqueous solution having a hydrophilic polymer concentration within a range of from 7 to 30% by weight, containing an osmotic pressure control agent composed of at least one compound selected from the group consisting of potassium chloride, potassium phosphate, magnesium chloride, magnesium phosphate, calcium chloride, calcium phosphate, amino acids, urea, acetic acid, ammonia, tricine, thiamine, riboflavin, nicotinic acid amide, pyridoxine, cyanocobalamine and ascorbic acid in a proportion that an osmolarity falls within a range of from 0.05 to 0.94 osmoles, substantially containing no sodium ion, and satisfying, assuming that the concentration of the hydrophilic polymer is b % by weight, and the osmolarity of the osmotic pressure control agent is a osmole(s), the relationship represented by the following expression (A): a≦0.1b−0.6 (A)with radiation selected from the group consisting of α-ray, electron beam and γ-ray to crosslink the hydrophilic polymer. 5. The device for interstitial fluid extraction according to claim 4, wherein the hydrogel layer is obtained by irradiating the coating film of the aqueous solution of the hydrophilic polymer with electron beam, whose ac... | SYSMEX CORPORATION NICHIBAN CO., LTD. | A61B005/00 B32B037/02 B29C065/14
US07881516 | B2 | 2011-02-01 | Method and system of image fusion for radiation therapy | 2007-12-14 | An image fusion method which is used for medical applications. The method includes: (a) acquiring a first image with a planned radiation region; (b) acquiring a second image with actual radiation region; (c) determining if user defined landmarks have been placed on the first and second images, if user defined landmarks are present go to step (d), if not go to step (e); (d) pre-transforming the first image or second image or both images; and (e). performing a first delineation step on the actual radiation region. The method further includes: (f) determining if the delineation is correct, if yes go to step (g), if not go to step (h); (g) fusing the first and second image and exit process; and (h) selecting multiple guide points around the actual radiation region in the second image, wherein the guide points are position adaptable points; which are placed near but not on the boundaries of the region; and (i) performing a second delineation step on the actual radiation region and go to step (f). | 1. An image fusion method for medical applications, comprising; a. acquiring a first image with a planned radiation region;b. acquiring a second image with an actual radiation region;c. determining if user defined landmarks have been placed on said first and second images, if user defined landmarks are present go to step (d), if not go to step (e);d. pre-transforming said first image or second image or both images;e. performing a first delineation step on said actual radiation region using active contour model with adaptive weighting on said second region;f. determining if said delineation is correct, if yes go to step (g), if not go to step (h);g. fusing said first and second image and exit process, wherein fusing comprises the step of forming a mask that encloses said second region using said region contour vector with distinct values inside and outside of said mask, the mask tagging pixels in the radiation and non-radiation regions differently; andh. selecting multiple guide points around said actual radiation region in said second image, wherein said guide points are position adaptable points, which are placed near but not on the boundaries of the region; andi. performing a second delineation step on said actual radiation region and go to step (f). 2. A method as in claim 1 wherein the landmarks in said first image correspond to the landmarks in said second image. 3. A method as in claim 1 wherein said first delineation step comprises: e.1. applying a level set contour finding algorithm with an adaptive weight to said second image to delineate said second region; ande.2. forming a region contour vector of said second region. 4. A method as in claim 1 wherein: h.1. said multiple guide points are selected near the boundaries of said second region as position adaptable seed-target pairs; and wherein said second delineation step comprises:i.1. applying an intelligent scissors algorithm with adaptive banding method to said second image using said position adaptable seed-target pairs, wherein said adaptable seed and target positions are adjusted based on the cost of traveling between said adaptable seed and target so that the adjusted seed and target are on or closer to the boundaries of said second region; andi.3. forming a region contour vector of said second region. 5. A method as in claim 1 wherein fusion comprises the steps of: defining a mask feature point;defining a feature point for said first region in said first image; andsearching maximum similarity of said first and second images in multi-resolution, multi-dimensional spaces using information extracted from the inside or outside or both of said mask region. 6. A method as in claim 5 wherein said mask feature point is derived from mask geometry or selected landmarks of said second region. 7. A method as in claim 5 comprising the steps of: decomposing said mask, said first and second images into a plurality of masks and images with different resolutions levels;transforming said deco... | Carestream Health, Inc. | G06K009/36 G06K009/00 A61B008/00 A61B005/05
US20110028856 | A1 | 2011-02-03 | System for Heart Performance Characterization and Abnormality Detection | 2010-03-30 | A system improves analysis, diagnosis and characterization of cardiac function signals (including surface ECG signals and intra-cardiac electrograms) based on cardiac electrophysiological activity momentum computation, characterization and mapping. The system calculates an electrophysiological signal momentum of different portions of cardiac signals including a timing, location and severity of cardiac pathology and improves reliability of diagnosis, detection, mapping to an identified medical condition, and characterization. The system improves identification of cardiac disorders, differentiation of cardiac arrhythmias, characterization of pathological severity, prediction of life-threatening events and supports evaluation of drug administration effects. | 1. A system for heart performance characterization and abnormality detection, comprising: an interface for receiving an electrical signal indicating electrical activity of a patient heart over at least one heart beat cycle;a signal processor for calculating a signal characteristic value comprising a summation of values of rate of change of amplitude of said electrical signal over at least a portion of a heart beat cycle;a comparator for comparing the calculated signal characteristic value with a threshold value to provide a comparison indicator; anda patient monitor for in response to said comparison indicator indicating the calculated signal characteristic value exceeds the threshold value, generating an alert message associated with the threshold. 2. A system according to claim 1, wherein said patient monitor, in response to said comparison indicator indicating the calculated signal characteristic value lies in a predetermined value range, generates an alert message associated with the value range. 3. A system according to claim 2, wherein said patient monitor substantially continuously monitors said comparison indicator for at least a 24 hour period. 4. A system according to claim 1, wherein said threshold value is derived from recorded electrical signal data for said patient. 5. A system according to claim 1, wherein said threshold value is derived from recorded electrical signal data for a population of patients. 6. A system according to claim 5, wherein said population of patients has similar demographic characteristics including at least two of (a) age, (b) weight, (c) gender and (d) height, to those of said patient. 7. A system according to claim 1, wherein said signal processor dynamically adjusts said threshold value in response to a determined sensitivity of arrhythmia detection. 8. A system according to claim 1, wherein said signal processor calculates said signal characteristic value for a predetermined portion of a heart beat cycle in response to a synchronization signal. 9. A system according to claim 8, wherein said predetermined portion of said heart beat cycle includes an ST segment. 10. A system according to claim 1, wherein said signal processor calculates said signal characteristic value as an averaged value over a plurality of heart beat cycles. 11. A system according to claim 1, wherein said signal processor calculates said signal characteristic value in response to a heart rate synchronization signal. 12. A system according to claim 1, including a repository of mapping information, associating ranges of the signal characteristic value or values derived from the signal characteristic value, with corresponding medical conditions andsaid comparator compares the calculated signal characteristic value with said ranges to provide a comparison indicator identifying a medical condition andsaid patient monitor generates an alert message identifying said medical condition. 13. A system according to claim 12, wherein said... | SIEMENS MEDICAL SOLUTIONS USA, INC. | A61B005/0402
US20110009757 | A1 | 2011-01-13 | BLOOD PRESSURE INFORMATION MEASUREMENT DEVICE | 2010-09-17 | A sphygmomanometer separately includes a cuff accommodating an air bladder and a main body provided with an expanding/contracting mechanism arranged to expand and contract the air bladder, and the cuff includes a tubular cuff main body portion into which an upper arm is insertable and a gripping portion provided on an outer peripheral surface of the cuff main body portion. The cuff main body portion has a tightening belt wrapped around the outer side of the air bladder and a tightening length adjustment mechanism arranged to variably adjust a tightening length of the tightening belt over the upper arm. The tightening length adjustment mechanism includes a winding roller arranged to wind and feed the tightening belt, a geared motor arranged to drive and rotate the winding roller, and an electromagnetic brake arranged to apply a braking force to the winding roller. With such a configuration, in the sphygmomanometer in which the cuff and the main body are separated from each other, the cuff can be easily attached to a measurement site, and reliable winding of the cuff around the measurement site can be repeated in every measurement. | 1. A blood pressure information measurement device, comprising: a cuff arranged to be attached to a measurement site at a time of measurement; anda main body, separate from the cuff, and to be placed on a supporting surface for use at the time of measurement; whereinthe cuff includes: a tubular cuff main body portion including a hollow opening portion into which the measurement site is insertable in an axial direction; anda gripping portion provided on an outer peripheral surface of the cuff main body portion; whereinthe cuff main body portion includes: a fluid bag arranged to compress the measurement site;a tightening belt wrapped around an outer side of the fluid bag; anda tightening length adjustment mechanism arranged to variably adjust a tightening length of the tightening belt over the measurement site;the main body includes: an expanding/contracting mechanism arranged to expand and contract the fluid bag; andthe tightening length adjustment mechanism includes: a winding roller arranged to wind and feed the tightening belt;an electric motor arranged to drive and rotate the winding roller in a forward direction and a backward direction; anda brake arranged to apply a braking force to the winding roller at a time of stopping the electric motor. 2. The blood pressure information measurement device according to claim 1, wherein the gripping portion includes a base portion attached to the cuff main body portion, andthe tightening length adjustment mechanism is accommodated in the base portion. 3. The blood pressure information measurement device according to claim 1, wherein a winding operation unit arranged to receive a command to start a winding operation of the tightening belt by the winding roller is provided in the gripping portion. 4. The blood pressure information measurement device according to claim 3, wherein the winding operation unit is defined by a push button. 5. The blood pressure information measurement device according to claim 1, further comprising: a tightening force detection mechanism arranged to detect a tightening force of the tightening belt over the measurement site, whereinthe tightening length adjustment mechanism is arranged to stop a winding operation of the tightening belt by the winding roller when the tightening force detection mechanism detects a predetermined magnitude of the tightening force. 6. The blood pressure information measurement device according to claim 5, wherein the tightening force detection mechanism is arranged to detect an internal pressure of the fluid bag at a time of the winding operation of the tightening belt by the winding roller in a state in which a predetermined amount of fluid is injected into the fluid bag by the expanding/contracting mechanism, so as to detect the tightening force of the tightening belt over the measurement site. 7. The blood pressure information measurement device according to claim 5, wherein the tightening force detection mechanism is arranged to detect a ro... | OMRON HEALTHCARE CO., LTD. | A61B005/022
US20110034812 | A1 | 2011-02-10 | PULMONARY ARTERY PRESSURE BASED SYSTOLIC TIMING INTERVALS AS A MEASURE OF RIGHT VENTRICULAR SYSTOLIC PERFORMANCE | 2010-07-22 | Systems and methods include identifying a first portion and a second portion of a pulmonary artery pressure (PAP) signal during a cardiac cycle. A first timing interval between the first portion and the second portion is obtained and data related to the first timing interval is trended to provide a chronic physiological prognostic indicator. In an embodiment, a second timing interval is obtained from a third portion and a fourth portion of the PAP signal. Then, a function of the first and second timing intervals is trended to provide the chronic physiological prognostic indicator. In one instance, a ratio of the first interval to the second interval is calculated to provide an estimated right ventricle ejection fraction (RVEF) and the RVEF is trended. | 1. A system, comprising: a machine-readable media; andone or more processors communicatively coupled to the machine-readable media, the machine-readable media including instructions, which when executed by the one or more processors, cause the one or more processors to: identify a first portion of a pulmonary artery pressure (PAP) signal during a cardiac cycle;identify a second portion of the PAP signal during the cardiac cycle;obtain a first timing interval between the first portion and the second portion; andtrend data related to the first timing interval to provide a chronic physiological prognostic indicator. 2. The system of claim 1, wherein the instructions to identify the first portion of the PAP signal during the cardiac cycle further comprises instructions to: identify a feature of an intrinsic electrical cardiac signal during the cardiac cycle; andcorrelate the feature of the intrinsic electrical cardiac signal with a position on the PAP signal to identify the first portion of the PAP signal. 3. The system of claim 2, wherein the instructions to identify the feature of the intrinsic electrical cardiac signal during the cardiac cycle further comprises instructions to identify a feature indicative of the start of depolarization of ventricles. 4. The system of claim 1, wherein the instructions to identify the first portion of the PAP signal during the cardiac cycle further comprises instructions to identify a start of a right ventricle pre-ejection period, and wherein the instructions to identify the second portion of the PAP signal during the cardiac cycle further comprises the instructions to identify an end of the right ventricle pre-ejection period. 5. The system of claim 1, wherein the instructions to identify the first portion of the PAP signal during the cardiac cycle further comprises the instructions to identify a start of a right ventricle ejection time, and wherein the instructions to identify the second portion of the PAP signal during the cardiac cycle further comprises the instructions to identify an end of the right ventricle ejection time. 6. The system of claim 1, wherein the machine-readable media includes instructions, which when executed by the one or more processors, cause the one or more processors to: identify a third portion of the PAP signal during the cardiac cycle;identify a fourth portion of the PAP signal during the cardiac cycle; andobtain a second timing interval between the third portion and the fourth portion, and wherein the instructions to trend data related to the timing interval further comprises the instructions to trend a function of the first and second timing intervals to provide the chronic physiological prognostic indicator. 7. The system of claim 6, wherein the instructions to trend the function of the first and second timing intervals further comprises instructions to: calculate a ratio of the first interval to the second interval to provide an estimated right ventricle ejection fraction (... | - | A61B005/021 A61N001/365
US20110028859 | A1 | 2011-02-03 | Methods, Systems and Devices for Monitoring a Target in a Neural System and Facilitating or Controlling a Cell Therapy | 2009-09-01 | Methods, systems and devices for monitoring a target in a neural system involve delivering inputs to and recording output responses from the neural system and obtaining I/O characterizations of the neural system both before and during or after a change is introduced to the system, for example, a change associated with a stem cell therapy. Baseline and therapeutic I/O characterizations are compared, and interpretations of the results of the comparison may be used to adjust some aspect of the therapy to facilitate it or control it. An implantable neurostimulator with the capability of delivering electrical stimulation to an electrode and of measuring and recording responses sensed from electrodes, where the electrodes are strategically positioned at or near the target, can be used to deliver an electrical stimulation input and record the output responses for the I/O characterizations. An implantable neurostimulator also may be used to affect a cell therapy by delivering electrical stimulation to the neural system to encourage stem cell differentiation. | 1. A method for monitoring a target in a neural system in connection with a cell therapy for a human patient comprising: identifying a target;selecting a technique for acquiring input/output characterizations of the neural system;selecting an input to use for the characterizations;positioning a least one sensor in or near the neural system to sense an output response for the I/O characterizations;applying the input, sensing an output and performing a baseline characterization of the neural system with a signal processor based on the applied input, the sensed output and the selected characterization technique;introducing a change to the neural system associated with the cell therapy;applying the input, sensing the output and performing a therapeutic characterization of the neural system with the signal processor based on the input and output and using the selected characterization technique;displaying the results of the baseline characterization and the therapeutic characterization to permit an observer to compare the characterizations. 2. The method of claim 1 wherein the target includes a population of a particular neuronal cell type within the neural system. 3. The method claim 2 wherein the neuronal cell type is one of the group including excitatory neurons and inhibitory neurons. 4. The method of claim 2 wherein the neuronal cell type is one of the group including granule cells, pyramidal cells, and interneurons. 5. The method claim 1 wherein the selected technique for acquiring the characterizations is based one of linear or nonlinear systems theory. 6. The method of claim 5 wherein the selected technique for acquiring the characterizations is further based on one or more systems models. 7. The method of claim 6 wherein the systems models are selected from the group including a feedback model, an addition model, and a cascade model. 8. The method of claim 1 wherein positioning at least one sensor includes positioning a first sensor at a first location in a hippocampal pathway in the patient's brain and positioning a second sensor at a second location in the patient's brain. 9. The method of claim 8 wherein each of the first and second sensors is an electrode configured to sense electrical activity from the patient's brain. 10. The method of claim 1 wherein each of the first and second sensors is an electrode configured to sense electrical activity from the patient's brain and further comprising positioning a stimulating electrode in or near the neural system to deliver the input for the I/O characterizations, generating an electrical stimulus corresponding to the selected input, and delivering the generated electrical stimulus through the stimulating electrode. 11. The method of claim 10 wherein the selected technique for acquiring the I/O characterizations is based one of linear or nonlinear systems theory and is further based on one or more systems models. 12. The method of claim 11 wherein the one or more systems models are fro... | NeuroPace, Inc. | A61B005/05
US20110030256 | A1 | 2011-02-10 | Adjustable and reusable identification bracelet with concealed compartment and secure locking mechanism | 2010-04-02 | Adjustable and Reusable Identification Bracelet with Concealed Compartment and Secure Locking Mechanism for identifying individuals. The device comprises one individual strap of polypropylene webbing which can fasten securely around an individual's wrist. The opposite ends of the strap lock together to form a continuous identification bracelet around an individual's wrist. An adjustment means is incorporated onto the strap to lengthen or shorten the bracelet to ensure a proper fit. The bracelet further comprises a concealed fabric compartment beneath the bracelet to secure identification information on a plastic card. Accordingly, the device may be used as a means of identification and identification storage. | 1. Identification bracelet made of polypropylene webbing or flexible material with an adjustable sliding mechanism made of plastic enabling the size of the bracelet to be reduced or enlarged so it is adjustable and better able to fit various wrist sizes. Adjustable sliding mechanism also serves to secure the identification in place in concealed storage compartment. 2. Identification bracelet comprising of a sliding lock mechanism made of plastic which functions as a closure device and enables the identification bracelet or anklet to be attached or removed manually from the body part of the wearer. 3. Identification bracelet with concealed compartment made of polypropylene webbing sewn to fabric for identification storage made of fabric beneath identification bracelet with a slide in and out, removable water proof identification component made of plastic. Adjustable sliding mechanism also serves to secure the identification in place in concealed storage compartment. 4. The fabric concealed compartment of claim 3 is not transparent. 5. Identification bracelet with concealed compartment of claim 3 wherein the information tag is not viewable from the bracelet unless it is physically removed. 6. Identification bracelet with concealed compartment of claim 3 wherein it is sealed shut with the same adjustable sliding mechanism that also enables the bracelet to be enlarged or reduced in size. 7. Identification bracelet with concealed compartment of claim 3 wherein the identification information can be accessed by first responders or caring individuals by sliding back manually the movable lock located on the female clasp and then sliding back the adjustable sliding element and manually pulling out the information card to reveal the contact or medical information which is discreetly located. | - | A44C005/00 A61B005/117
US20110015496 | A1 | 2011-01-20 | PORTABLE MEDICAL DEVICE | 2009-07-14 | A housing configured to at least partially encapsulate a mobile communication device, a sensor disposed on the housing, the sensor being configured to collect medical information relating to a patient, and an interface coupled to the sensor and configured to communicate the collected medical information to the mobile communication device. | 1. A medical device comprising: a housing configured to at least partially encapsulate a mobile communication device;a sensor disposed on the housing, the sensor being configured to collect medical information relating to a patient; andan interface coupled to the sensor and configured to communicate the collected medical information to the mobile communication device. 2. The medical device of claim 1 wherein the sensor includes a plurality of sensors configured to detect at least one of electrical activity in the patient, physiological parameters, and environmental parameters. 3. The medical device of claim 2 wherein the plurality of sensors are configured to collect information used to generate an electrocardiogram. 4. The medical device of claim 1 wherein the housing is configured as a skin that at least partially encapsulates the mobile communication device without substantially interfering with the operation of the mobile communication device. 5. The medical device of claim 1 wherein the housing includes a feedback device configured to provide to the patient feedback relating to the operation of the medical device. 6. The medical device of claim 5 wherein the feedback device is configured to provide feedback relating to at least one of a duration to collect the medical information and a quality of medical information collected. 7. The medical device of claim 1 wherein the interface is at least one of a wired interface configured to connect with an interface on the mobile communication device and a wireless interface. 8. The medical device of claim 1 further comprising a processor configured to control the operation of at least one of the sensor and the interface. 9. The medical device of claim 8 wherein the processor is configured to perform digital signal processing on information provided by the sensor. 10. The medical device of claim 1 wherein the interface comprises at least one of a radio frequency identification tag, a tone generator, and a vibration generator. 11. A tangible computer readable medium including instructions that, when executed, cause a mobile communication device to: receive medical information relating to a patient from an external device;perform digital signal processing on the medical information;display on a display of the mobile communication device a visual indication that is derived from the medical information; andstore the medical information in a memory. 12. The tangible computer readable medium of claim 11 wherein the instructions are further configured to cause the mobile communication device to display an electrocardiogram relating to the patient on the display. 13. The tangible computer readable medium of claim 11 wherein the medical information is received via a local interface of the mobile communication device, wherein the local interface is at least one of a wired and a wireless interface. 14. The tangible computer readable medium of claim 11 wherein the instructions are further configured ... | - | A61B005/00 G08B023/00
US20110046460 | A1 | 2011-02-24 | METHOD AND MEASUREMENT DEVICE FOR RECORDING MEASUREMENT SIGNALS FROM VITAL TISSUE | 2009-02-25 | The invention is directed to a measuring method and apparatus for taking spectrometric readings from vital tissue. The object of the invention is to provide solutions by which, in the course of a spectrometric measurement, it is possible to generate measurement values that, compared to above-described previous recording methods, provide more comprehensive information. According to the invention, this object is attained by a method that generates spectrometric signals and in which light is directed into a sample of vital tissue to be examined, remission light exiting the tissue sample to be examined is fed to a spectrometer, and the spectrometer generates output signals that represent the intensity of the remission light with respect to the wavelength, the measurement being performed in such a manner that it extends over a period of time during which the presence of hemoglobin in the examined tissue sample is actively changed by changing pressure on the tissue, wherein the spectra determined in succession for different tissue pressures are used in order to determine changes in the spectra induced by the hemoglobin concentration, and, from these different spectra, the concentrations of selected substances in the tissue sample are calculated. In this manner, it is advantageously possible to generate, in relatively quick succession, a plurality of spectra relating to the tissue sample under examination, wherein these spectra have differences (“distortions”) caused by actively induced changes of pressure on the tissue, and the differences are sufficient for determining the concentration of substances to be detected in the vascular parts of the tissue system. | 1. A method of generating spectrometric readings, the method comprising the steps of: a) directing light into a sample of vital tissue to be examined,b) feeding remission light exiting the tissue sample to be examined to a spectrometer,c) generating with the spectrometer output spectra that represent the intensity of the remission light with respect to the wavelength,d) repeating steps a to c over a period of time during which pressure on the tissue is changed to actively alter the presence of hemoglobin in the examined tissue sample, ande) comparing the spectra determined in succession for different tissue pressures to determine changes in the spectra induced by the hemoglobin concentration, andf) calculating from these different spectra the concentrations of selected substances in the tissue sample. 2. The method according to claim 1, wherein pressure on the tissue is changed by locally applying pressure with a gaseous medium. 3. The method according to claim 2, wherein the pressure is applied by using an elastic bellows. 4. The method according to claim 3, wherein the bellows is connected to a measuring head that has a light source usable for step a and a light sensor usable for step b. 5. The method according to claim 1 wherein pressure on the tissue is increased. 6. The method according to claim 1, wherein for each tissue pressure level, a spectrum of the remission light is recorded. 7. The method according to claim 1 wherein for recording the successively recorded spectra, a data field is created for data for each resolved wavelength value and the intensity or optical density and pressure on the tissue. 8. A mobile spectrometer comprising: a storage device,an evaluation device,means for directing light into a sample of vital tissue to be examined,means for feeding remission light exiting the tissue sample to be examined to a spectrometer,means for generating output signals that represent the intensity of the remission light with respect to its wavelength,means for changing the pressure in the tissue sample to be examined, andmeans connected to the storage device and to the evaluation device for recording a plurality of spectra detected at different tissue pressures. 9. An apparatus for generating spectrometric readings, the apparatus comprising: means for directing light into a sample of vital tissue to be examined, whereby remission light is caused to issue from the tissue sample,a spectrometer;means for feeding the remission light exiting the tissue sample to the spectrometer,means for changing pressure on the sample over a predetermined period of time;means in the spectrometer for generating at different pressure levels within the period of time spectra that represent the intensity of the remission light with respect to the wavelength at the respective pressure level; andevaluating means for comparing these spectra to each other and determining changes in the spectra and for suppressing the changes in the spectra, whereby only re... | - | A61B005/1455
US20110027189 | A1 | 2011-02-03 | BIOCOMPATIBLE PRODUCTS FOR MAGNETIC PARTICLE IMAGING | 2009-03-26 | The present invention relates to materials and methods for non-invasive imaging of biocompatible products. Biocompatible products are provided that can be visualized in vivo and in vitro using magnetic particle imaging methods. It was found that these materials can be employed for monitoring development, remodeling or degradation of biocompatible products. In another aspect, the inventive products enable substance delivery tracking and can be used to visualize targeted delivery of active agents. | 1. A biocompatible product comprising at least one magnetic particle imaging (MPI) tracer agent. 2. The biocompatible product of claim 1, wherein the biocompatible product is an artificial tissue construct. 3. The biocompatible product of claim 1, wherein the biocompatible product is a microcarrier or a microcontainer. 4. The biocompatible product of claim 1, wherein the biocompatible product is an implant. 5. The biocompatible product of claim 1, wherein the at least one MPI tracer agent is incorporated into the biocompatible product. 6. The biocompatible product of claim 1, wherein the at least one MPI tracer agent is adhered to the surface of the biocompatible product. 7. The biocompatible product of claim 1, wherein the at least one MPI tracer agent is surrounded by a pharmaceutically acceptable shell. 8. The biocompatible product of claim 7, wherein the at least one MPI tracer agent surrounded by a pharmaceutically acceptable shell further comprises at least one ligand which is immobilized on the surface of the shell. 9. The biocompatible product of claim 1, wherein the at least one MPI tracer agent has a particle size between 5 and 5000 nm, preferably between 20 and 50 nm. 10. A method for visual monitoring a biocompatible product comprising at least one magnetic particle imaging (MPI) tracer agent, the method, comprising: providing the biocompatible product comprising the at least one MPI tracer agent;introducing the biocompatible product comprising the at least one MPI tracer agent to a target site; anddetecting signals arising from the at least one MPI tracer agent by magnetic particle imaging. 11. The method according to claim 10, wherein the step of detecting signals comprises the following steps: localizing the signals;quantifying the signals; andevaluating the magnetic particle imaging data obtained for the biocompatible product. 12. A method comprising use of an MPI tracer agent for visual monitoring a biocompatible product comprising at least one magnetic particle imaging (MPI) tracer agent using magnetic particle imaging. | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/055 A61K049/18 C12Q001/02
US07883476 | B2 | 2011-02-08 | Selectively detachable outer cannula hub | 2006-02-28 | A disposable tissue removal device comprises a “tube within a tube” cutting element mounted to a handpiece. The inner cannula of the cuffing element defines an inner lumen and terminates in an inwardly beveled, razor-sharp cutting edge. The inner cannula is driven by both a rotary motor and a reciprocating motor. At the end of its stroke, the inner cannula makes contact with the cutting board to completely sever the tissue. An aspiration vacuum is applied to the inner lumen to aspirate excised tissue through the inner cannula and into a collection trap that is removably mounted to the handpiece. The rotary and reciprocating motors are hydraulically powered through a foot pedal operated hydraulic circuit. The entire biopsy device is configured to be disposable. In one embodiment, the cutting element includes a cannula hub that can be connected to a fluid source, such as a valve-controlled saline bag. | 1. A medical device for attaching to a handpiece defined by a distal end and having an inner cutting element including an inner cannula that has an open distal end, said medical device comprising: an outer cannula having a proximal end and a distal end, said outer cannula being sized to accept the inner cutting element including the inner cannula; anda cannula hub having a bore therethrough, said cannula hub being fixedly attached to said proximal end of said outer cannula and supporting said outer cannula, wherein said cannula hub further comprises a cooperating fitting member that is configured for selective positive engagement with a mating fitting disposed on the distal end of the handpiece so as to selectively and removably secure said cannula hub to said handpiece, the cooperating fitting member including a flange having a plurality of wings separated by a plurality of recessed regions, at least one of the wings having a larger circumferential length than at least one of the remaining wings;wherein an instrument is selectively insertable through said bore and into said outer cannula. 2. The device of claim 1, wherein a marker, a therapeutic agent, a fluid, or a drug is provided through said bore when the handpiece is removed from said cannula hub. 3. The device of claim 1, wherein said cannula hub is configured to facilitate delivery of a marker. 4. The medical device of claim 1 wherein the distal end of the outer cannula is closed. 5. The medical device of claim 1 further comprising an irrigation fitting on the cannula hub, the irrigation fitting comprising an irrigation lumen in communication with the bore in the cannula hub. 6. The medical device of claim 1 wherein the proximal end of the outer cannula extends partially through the bore in the cannula hub. 7. The medical device of claim 1 wherein said flange of said cooperating fitting member is received within a recess of said mating fitting of said handpiece. 8. The medical device of claim 7, wherein said cooperating fitting member of said cannula hub and said mating fitting of said handpiece are configured as Luer fittings that are selectively engageable by a twisting motion. 9. The medical device of claim 1, wherein said flange of said cooperating fitting member includes at least one of a protuberance and a retention dimple engageable with a correspondingly configured feature arranged on said mating fitting of said handpiece. 10. A medical device comprising: a handpiece having an inner cutting instrument including an inner cannula, wherein said handpiece is further defined by a distal end, said distal end having a fitting disposed thereon;an outer cannula having a proximal end and a distal end, said proximal and distal ends sized for selective insertion of said inner cutting instrument including the inner cannula; anda cannula hub having a bore therethrough, said cannula hub being fixedly attached to said proximal end of said outer cannula, wherein said cannula hub further c... | Suros Surgical Systems, Inc. | A61B005/117 A61B005/103
US08241229 | B2 | 2012-08-14 | Drill device and method for forming microconduits | 2010-10-29 | The present invention relates to methods and devices for formation of microconduits in tissue, particularly using an impedance sensing drill to form microconduits. One embodiment of the invention is an impedance sensing drill comprising a drilling assembly, a control module, mechanically connected to the drilling assembly for controlling the depth of drilling by the drilling assembly; and a sensor, electrically connected to the drilling assembly and control module for detecting a change in an electrical impedance of a material being drilled. Another embodiment is a method of forming a microconduit in a material, which comprises the steps of drilling into the material, monitoring an electrical impedance of the material, and stopping the drilling into the material when a change in the impedance is detected, thereby forming microconduit. | 1. A method of removing an analyte, which comprises the steps of: painlessly drilling into a target area with a drilling assembly, monitoring an electrical impedance of the target area, stopping the drilling into the target area when a change in the electrical impedance is detected, thereby forming a microconduit, and removing the analyte via the formed microconduit; wherein the drilling assembly includes a motor, an impedance sensor, and a cutter, and wherein the drilling into the tissue is stopped and automatically reversed by a control module when a change in the electrical impedance of the target area is detected by the sensor. 2. The method of claim 1, wherein the electrical impedance being measured is electrical resistance. 3. The method of claim 1, wherein the impedance sensing drill is computer controlled. 4. The method of claim 1, wherein the drilling assembly further comprises a housing adapted to be placed in contact with a tissue being drilled; a motor and a cutter rotated by the motor which is mobile in translation in the housing and which is adapted to be moved into and out of the tissue being drilled for the formation of microconduits; a control module mechanically connected to the drilling assembly for controlling the translation of the cutter in the housing and into and out of the tissue being drilled; and a sensor, electrically connected to the drilling assembly and to the control module for detecting a change in an electrical impedance between the cutter and the surface of the tissue being drilled when the cutter is translated in the tissue being drilled; wherein the sensor, upon detection of a change in the electrical impedance corresponding to a transition between portions of the material being drilled, sends a signal to the control module to automatically reverse the direction of the cutter, thereby withdrawing the cutter from the tissue being drilled, and thereby controlling the depth of the drilling. 5. A method of removing the stratum corneum from the epidermis to permit extraction of analyte from the body, which comprises the steps of: drilling into a target area with a drilling assembly, monitoring an electrical impedance of the target area, stopping the drilling into the target area when a change in the electrical impedance is detected, thereby forming a microconduit, and extracting analyte from the body through the microconduit; wherein the drilling assembly includes a motor, an impedance sensor, and a cutter, and wherein the drilling into the tissue is stopped and automatically reversed by a control module when a change in the electrical impedance of the target area is detected by the sensor. 6. The method of claim 5, wherein the analyte is a blood sample. 7. The method of claim 5, wherein the analyte is a tissue sample. 8. The method of claim 5, wherein the analyte an interstitial fluid sample. 9. The method of claim 5, wherein the electrical impedance being measured is electrical resistance. 10. The method... | Path Scientific, LLC Galderma Pharma S.A. | A61B005/00
US07883469 | B2 | 2011-02-08 | Device for determining cardiac function parameters | 2006-09-07 | A device for detecting the state of a heart on the basis of intracardial impedance measurement. The device has an impedance measuring unit, which has electrical terminals, configured to be electrically connected or to be connected to electrodes for delivering and detecting a current or voltage, and is implemented to ascertain an impedance on the basis of the dimension of the delivered current or voltage and the voltage drop caused by the current or the current caused by the voltage, as well as an analysis unit, which is connected to the impedance measuring unit and is implemented to derive a cardiac function parameter from a time curve of the impedance ascertained using the impedance measuring unit. The analysis unit analyzes the impedance curve assigned to a diastole and derives a cardiac function parameter characterizing the behavior of a heart during the diastole (filling phase of the ventricle). | 1. A device for detecting the state of a heart on the basis of intracardial impedance measurement comprising: at least two electrodes;an impedance measuring unit which is electrically connected to the at least two electrodes;wherein said impedance measuring unit is configured to deliver a delivered current or apply an applied voltage and detect a voltage or a current;wherein said impedance measuring unit is configured to ascertain an impedance based on a quantity of said delivered current or said applied voltage and a voltage drop caused by the current, or the current caused by the voltage and derive an impedance curve;an analysis unit, connected to said impedance measuring unit and configured to identify a diastole in said impedance curve as a period between maximum and minimum impedance, and to derive a maximum negative gradient of the impedance curve during said diastole; and,wherein said analysis unit is configured to determine a period of time between a beginning of an impedance reduction in said impedance ascertained by said impedance measuring unit and a following occurrence of the maximum negative gradient of said impedance curve derived by said impedance measuring unit as a cardiac function parameter through analysis of said impedance curve. 2. The device according to claim 1 wherein said analysis unit is configured to determine either a period of time between an occurrence of the maximum negative gradient and an apex of a parabolic approximation function of said impedance curve, beginning with said occurrence of said maximum negative gradient, or a time constant of a next potential approximation function of said impedance curve, beginning with said occurrence of said maximum negative gradient, as a value of the cardiac function parameter, namely a filling deceleration time, through analysis of said impedance curve. 3. The device according to claim 1 wherein said analysis unit is configured to determine a ratio of the maximum negative gradient of said impedance curve between the beginning of the impedance reduction and a following atrial contraction to said maximum negative gradient of the impedance curve between this atrial contraction and a following further increase of said impedance curve as a value of the cardiac function parameter through analysis of said impedance curve. 4. The device according to claim 1 wherein said analysis unit is configured to determine a duration of a particular diastole as the cardiac function parameter value through analysis of said impedance curve, in that said analysis unit determines a period of time between the beginning of the impedance reduction after exceeding the maximum impedance and an occurrence of the minimum impedance of said impedance curve, through analysis of said impedance curve. 5. The device according to claim 1 wherein said analysis unit is configured to analyze said impedance curve assigned to a systole and to calculate the cardiac function parameter value that characterizes behavi... | Biotronik CRM Patent AG | A61B005/0452
US08460212 | B2 | 2013-06-11 | Body fluid sampling unit | 2010-10-21 | A body fluid sampling unit includes a puncture device including a needle having a sharp needle point and a drive mechanism for operating the needle to puncture a living body surface with the needle point, a tip including an introducing section for introducing a body fluid flowing out from a puncture portion of the living body surface punctured with the needle point, a test paper for detecting a predetermined component in the body fluid, and a tip body mounted to a component measuring device for measuring the quantity and/or property of the predetermined component detected by the test paper; and a case including a mechanism containing section for containing the puncture mechanism so that the needle point of the puncture mechanism can be moved by the drive means, and a sampling implement containing section for containing the body fluid sampling implement so that the tip can be taken out. | 1. A body fluid sampling unit comprising: a puncture mechanism comprising a needle possessing a sharp needle point at a distal end of the needle, and drive means for operating the needle to puncture a living body surface with the needle point;a body fluid sampling implement comprising an introducing section for introducing a body fluid flowing out from a puncture portion of the living body surface punctured with the needle point, a detecting section for detecting a predetermined component in the body fluid introduced through the introducing section, and a mounted section configured to be mounted to a component measuring device for measuring at least one of a quantity and a property of the predetermined component detected by the detecting section;a container comprising a mechanism containing section containing the puncture mechanism so that the needle point of the puncture mechanism is movable by the drive means, and a sampling implement containing section containing the body fluid sampling implement in a manner allowing the body fluid sampling implement to be removed from the sampling implement containing section; andthe puncture mechanism is put into a state of being prepared to perform an operation of puncturing the living body surface by removal of the body fluid sampling implement from the sampling implement containing section. 2. The body fluid sampling unit according to claim 1, wherein the drive mean comprises a support section positioned inside the mechanism containing section and an engaging piece forming a part of the support section and engaging an inner periphery of an opening in the container, the needle possessing a proximal end fitted in the support section so that the needle and the support section move together. 3. The body fluid sampling unit according to claim 2, wherein the drive means comprises a spring which contacts the support section and applies a biasing force to the support section. 4. The body fluid sampling unit according to claim 2, further comprising a plate piece comprising a proximal end portion possessing an inner periphery contacted by an outer periphery of the body fluid sampling implement, the plate piece also comprising a distal end portion at which is located the opening. 5. The body fluid sampling unit according to claim 4, wherein plate piece is supported by a wall positioned intermediate the proximal end portion and the distal end portion so that a force applied in one direction to the proximal end portion of the plate piece causes movement of the distal end portion of the plate piece in an opposite direction, the contact between the outer periphery of the body fluid sampling implement and the inner periphery of the proximal end portion of the plate piece preventing the distal end portion of the plate piece from moving in a manner causing the engaging piece to become disengaged from the opening when a force is applied to the proximal end portion of the plate piece. 6. The body fluid sampling unit acc... | Terumo Kabushiki Kaisha | A61B005/00 A61F002/02 B65D081/00
US08386006 | B2 | 2013-02-26 | Medical electrode, electrode bundle and electrode bundle array | 2008-12-03 | A medical microelectrode includes portions capable of movement relative to each other when implanted in or inserted into soft tissue, so as to increase or decrease their distance along the electrode. The electrode is at least partially embedded in a substantially rigid biocompatible matrix that is soluble or biodegradable a body fluid. Also disclosed are uses of the microelectrode; microelectrode bundles and arrays of microelectrode bundles and their uses; methods for inserting or implanting microelectrodes, microelectrode bundles and arrays of microelectrode bundles in soft tissue. | 1. A medical electrode comprising: a base section,a tip section,an electrode body disposed between the base section and the tip section,the electrode body including one of: (a) an electrically conducting core and an insulating layer on the core, or(b) an electrically non-conducting core, an electrically conducting layer on the core and an insulating layer on the electrically conducting layer;the electrode body further including portions configured for movement relative to each other when implanted in or inserted into soft tissue, so as to increase or decrease their distance along the electrode, anda substantially rigid biocompatible matrix that is soluble or biodegradable in a body fluid, at least the tip section and the electrode body embedded in the matrix, the matrix forming a matrix body of a shape configured to follow a straight insertion track line when inserting the electrode deep into soft tissue. 2. The electrode of claim 1, wherein the base section, tip section, electrode body and substantially rigid biocompatible matrix form the matrix body. 3. The electrode of claim 2, wherein the electrode body comprises one or several passages through the insulating layer perpendicular to the core permitting electrical contact with the electrically conducting core or the electrically conducting layer. 4. The electrode of claim 2, wherein the electrode body is embedded in the matrix in a configuration in which its length is greater by 2 percent or more than the distance between its ends. 5. The electrode of claim 1, further comprising an anchor disposed at the tip section. 6. The electrode of claim 5, wherein the anchor is embedded in the matrix. 7. The electrode of claim 1, wherein the matrix body comprises two sections differing in their dissolution or degradation rate. 8. The electrode of claim 1, wherein a diameter of the electrode body is from about 10−7 to about 10−4 m. 9. The electrode of claim 1, further comprising a dissolution retardation coating on the matrix body. 10. A medical electrode bundle comprising: two or more bundle electrodes, each bundle electrode including:base section,a tip section,an electrode body disposed between the base section and the tip section,the electrode body including one of: (a) an electrically conducting core and an insulating layer on the core, or(b) an electrically non-conducting core, an electrically conducting layer on the core and an insulating layer on the electrically conducting layer; andportions of the electrode body configured for movement relative to each other when implanted in or inserted into soft tissue, so as to increase or decrease their distance along the electrode, anda substantially rigid biocompatible matrix that is soluble or biodegradable in a body fluid, at least the tip section and the electrode body embedded in the matrix, the matrix forming a bundle matrix body of a shape configured to follow a straight insertion track line when inserting the electrode deep into soft tissue. ... | Neuronano AB | A61B005/04 A61N001/05
US08565846 | B2 | 2013-10-22 | Blood oxygen saturation measuring apparatus | 2010-02-22 | A blood oxygen saturation measuring apparatus includes: a light emitter adapted to emit, to living tissue of a subject, at least two light beams which have different wavelengths from each other; a light receiver adapted to receive the light beams which are transmitted through or reflected from the living tissue, the light receiver which converts the received light beams to electric signals corresponding to receiving intensities of the light beams, respectively; an extractor which extracts frequency components corresponding to changes of the receiving intensities according to a variation of a blood volume of a vein in the living tissue due to respiration of the subject, from the electric signals, respectively; an attenuation ratio calculator which calculates a first attenuation ratio based on amplitudes of the extracted frequency components; and an oxygen saturation calculator which calculates a first blood oxygen saturation in the vein based on the first attenuation ratio. | 1. A blood oxygen saturation measuring apparatus comprising: a light emitter adapted to emit, to living tissue of a subject, at least two light beams which have different wavelengths from each other;a light receiver adapted to receive the light beams which are transmitted through or reflected from the living tissue, the light receiver which converts the received light beams to electric signals corresponding to receiving intensities of the light beams, respectively;an extractor which extracts frequency components corresponding to changes of the receiving intensities according to a variation of a blood volume of a vein in the living tissue due to respiration of the subject, from the electric signals, respectively;an attenuation ratio calculator which calculates a first attenuation ratio based on amplitudes of the extracted frequency components;an oxygen saturation calculator which calculates a first blood oxygen saturation in the vein based on the first attenuation ratio; anda determiner which determines that the subject is in an obstructive apnea condition, when the first blood oxygen saturation in the vein shows a decreasing trend for a predetermined time period or longer, whereinthe extractor extracts pulsation components corresponding to changes of the receiving intensities according to pulsation of an artery in the living tissue, from the electric signals, respectively,the attenuation ratio calculator calculates a second attenuation ratio based on amplitudes of the pulsation components,the oxygen saturation calculator calculates a second blood oxygen saturation in the artery based on the second attenuation ratio, andthe determiner determines that the subject is in a central apnea condition, when the amplitudes of the frequency components are smaller than a predetermined value for a predetermined time period or longer and the second blood oxygen saturation in the artery does not show a decreasing trend. 2. The blood oxygen saturation measuring apparatus according to claim 1, further comprising: a body motion detector which detects a change of body motion of the subject,wherein the oxygen saturation calculator corrects the calculated first blood oxygen saturation based on the change of the body motion. | Nihon Kohden Corporation | A61B005/00
US08594759 | B2 | 2013-11-26 | Systems and methods for resolving the continuous wavelet transform of a signal | 2009-07-30 | According to an embodiment, techniques for estimating scalogram energy values in a wedge region of a scalogram are disclosed. A pulse oximetry system including a sensor or probe may be used to receive a photoplethysmograph (PPG) signal from a patient or subject. A scalogram, corresponding to the obtained PPG signal, may be determined. In an approach, energy values in the wedge region of the scalogram may be estimated by performing convolution-based or convolution-like operations on the obtained PPG signal, or a transformed version thereof, and the scalogram may be updated according to the estimated values. In an approach, a deskewing technique may be used to align data prior to adding the data to the scalogram. In an approach, one or more signal parameters may be determined based on the resolved and estimated values of the scalogram. | 1. A method for processing a scalogram, the method comprising: receiving a signal from a signal generator;using processing equipment for: sampling the signal to generate a set of samples,generating the scalogram based at least in part on a continuous wavelet transform of the set of samples,updating the set of samples with at least one new sample of the signal to generate an updated set of samples,convolving the updated set of samples with wavelet coefficients to produce skewed scalogram samples, wherein the skewed scalogram samples are skewed in time relative to the scalogram,deskewing the skewed scalogram samples to produce scalogram update samples, wherein deskewing comprises aligning the skewed scalogram samples in time relative to the scalogram,updating the scalogram by adding the scalogram update samples to the scalogram, anddetermining information from the signal based at least in part on the updated scalogram; andoutputting the information to an output device. 2. The method of claim 1, wherein the signal comprises a photoplethysmograph (PPG) signal. 3. The method of claim 2, wherein the information comprises a patient respiration rate. 4. The method of claim 2, wherein the information comprises a patient oxygen saturation level. 5. The method of claim 2, wherein the information comprises a patient respiration effort level. 6. The method of claim 2, wherein the sampling is performed at a substantially constant time-rate. 7. The method of claim 1, wherein the updating is performed at a refresh rate that depends at least in part on an expected time-rate of change of at least one patient parameter. 8. The method of claim 1, wherein the output device comprises a display. 9. The method of claim 1, wherein the convolving is performed at least in part by at least one processor, and wherein the at least one processor performs one or more matrix multiplications. 10. The method of claim 1, wherein the wavelet coefficients correspond to a truncated Morlet wavelet. 11. The method of claim 1, wherein the wavelet coefficients are stored in a wavelet table and wherein the wavelet coefficients for at least two scales are misaligned in time. 12. The method of claim 1, wherein the skewed scalogram samples comprise samples for at least two scales and wherein the deskewing comprises mapping the skewed scalogram samples for each scale to a different time position to produce the scalogram update samples. 13. A system for processing a scalogram, the system comprising: a sensor for receiving a signal from a signal generator;one or more processors coupled to the sensor, wherein the one or more processors are configured to: sample the signal to generate a set of samples,generate the scalogram based at least in part on a continuous wavelet transform of the set of samples,update the set of samples with at least one new sample of the signal to generate an updated set of samples,convolve the updated set of samples with wavelet coefficients to produce skewed ... | Nellcor Puritan Bennett Ireland | A61B005/00 G01N033/48 G01N033/50 G06F017/14
US08417310 | B2 | 2013-04-09 | Digital switching in multi-site sensor | 2009-08-10 | A system includes a flexible sensor configured to be placed into a first configuration and a second configuration, wherein the sensor is configured to measure a physiological characteristic. The sensor may include a first memory device configured to store a first set of calibration data and a second memory device configured to store a second set of calibration data. The system may further include a monitor coupled to the sensor, wherein the first memory device is accessible by the monitor in the first configuration and the second memory device is accessible by the monitor in the second configuration. | 1. A system, comprising: a monitor;a sensor configured to be coupled to the monitor, the sensor comprising: a sensor body configured to flex between a first configuration in which the sensor body is disposed in a substantially flat position and adapted to be applied to a first tissue site and a second configuration in which the sensor body is disposed in a flexed position and adapted to be applied to a second tissue site, wherein the sensor is configured to measure a physiological characteristic;a first memory device storing a first set of calibration data; anda second memory device storing a second set of calibration data, wherein the first memory device is accessible by the monitor in the first configuration and the second memory device is accessible by the monitor in the second configuration. 2. The system of claim 1, wherein the monitor comprises a pulse oximetry monitor. 3. The system of claim 1, wherein the sensor comprises a sensor connector. 4. The system of claim 3, wherein the sensor connector is configured to house the first memory and the second memory. 5. The system of claim 3, wherein the sensor body is configured to house the first memory and the second memory. 6. The system of claim 1, wherein the sensor comprises a first switch configured to allow activation of the first memory device in the first configuration and activation of the second memory device in the second configuration. 7. The system of claim 6, wherein the sensor comprises a first emitter configured to transmit light in the first configuration and a second emitter configured to transmit light in the second configuration. 8. The system of claim 7, wherein the sensor comprises a second switch configured to allow activation of the first emitter in the first configuration and activation of the second emitter in the second configuration. 9. The system of claim 8, wherein the sensor comprises a third switch configured to allow sensing of the first emitter in the first configuration and sensing of the second emitter in the second configuration. 10. A medical sensor, comprising: a sensor body comprising: a first sensor configuration device configured to store a first set of sensor parameters;a second sensor configuration device configured to store a second set of sensor parameters; anda first conductive path coupled to the first sensor configuration device and a second conductive path coupled to the second sensor configuration device, wherein the first conductive path is configured to be electrically connected to a third conductive path in a first configuration of the sensor and the second conductive path is configured to be electrically connected to the third conductive path in a second configuration of the sensor. 11. The sensor of claim 10, comprising a first switch configured to connect the first conductive path coupled to the first sensor configuration device in the first configuration and the second conductive path coupled to the second sensor configuration d... | Covidien LP | A61B005/1455
US08998828 | B2 | 2015-04-07 | Visualization testing and/or training | 2009-07-09 | The visual and/or visualization skills of a subject may be tested and/or trained using systems and methods in accordance with the present invention. Visual stimuli may be presented to a test subject. Measurements relating to the performance of the subject, such as stability information, eye movement data, physiological information, or other information, may be made both with and without visual stimuli being provided to a subject. Similar measurements may be made while the subject visualizes the visual stimuli, and the collected data may be compared to evaluate the subject's ability to visualize visual stimuli. Secondary stimuli may also be provided to subject as part of this process, or mental activities may be required by subject as well. | 1. A system for testing or training the stability of an individual, the system comprising: a display device that provides visual stimuli to the individual for a first period of time;at least one stability measurement device that measures the balance of the individual; anda control unit coupled to the at least one stability measurement device configured to: (1) collect a first set of measurement data while the display device provides the visual stimuli to the individual for the first period of time,(2) prevent the individual from being exposed to the visual stimuli and collect a second set of measurement data while the individual is instructed to form a mental image of the visual stimuli, the second set of measurement data collected during a second period of time not coextensive with the first period of time, and(3) compare the first set of measurement data to the second set of measurement data to evaluate the individual. 2. The system for testing or training the stability of an individual of claim 1, wherein the at least one stability measurement device comprises at least one pressure sensitive surface stood upon by the individual during the first period of time and the second period of time. 3. The system for testing or training the stability of an individual of claim 1, wherein the at least one stability measurement device comprises a plurality of inertial sensors that are affixed to the individual during the first period of time and the second period of time. 4. The system for testing or training the stability of an individual of claim 1, wherein the second period of time occurs after the first period of time. 5. The system for testing or training the stability of an individual of claim 4, wherein the at least one stability measurement device and control unit is further configured to: measure the stability of the individual during a third period of time not coextensive with the first period of time and not coextensive with the second period of time, wherein the third period of time occurs before the first period of time, and wherein the display device does not provide the visual stimuli to the individual during the third period of time; andcompare the first, second and third set of measurement data to evaluate the individvual. 6. The system for testing or training the stability of an individual of claim 1, wherein the display device comprises a display screen. 7. The system for testing or training the stability of an individual of claim 1, wherein the display device comprises display goggles. 8. The system for testing or training the stability of an individual of claim 1, wherein the visual stimuli displayed by the display device simulates an activity. 9. A computerized method stored on a non-transient storage medium carried out by a visualization training system having at least one processor for testing or training the visualization ability of an individual, the method comprising: measuring performance of the individual while the indi... | NIKE, Inc. | A61B003/032 A61B005/103 A61B005/117 A61B005/11 A63B023/025 A63B024/00 A63B071/06 A61B003/113 A61B005/0402 A61B005/0476 A61B005/0496 A63B026/00 A63B069/00
US08538505 | B2 | 2013-09-17 | Method for 3-D data collection with a biplane C-arm system with biplane acquisition multiplexing | 2010-08-13 | A method for collecting three-dimensional data of an object from a series of projection images recorded by a biplane C-arm system is provided. A cardiac activity is recorded. The cardiac frequency and a start cardiac phase are determined for calculating parameters of the C-arm planes. The C-arm planes are set with the parameters and data is acquired in the start cardiac phase. The C-arm planes are uniformly rotated at a same speed in a forward motion over an angular area and record data at different angular areas at different cardiac phases. Data is acquired in the start cardiac phase after termination of the forward motion. The C-arm planes are uniformly rotated at a same speed in a backward motion over an angular area and records data at different angular areas at different cardiac phases. The captured data are reconstructed after termination of the backward motion upon completed acquisition. | 1. A method for collecting a three-dimensional data of an object to be examined from a series of projection images of the object recorded by a biplane C-arm system having two separate C-arm planes, comprising: recording a cardiac activity of the object by an ECG;determining a cardiac frequency and a start cardiac phase;calculating parameters of the two C-arm planes from the cardiac frequency and the start cardiac phase;setting the C-arm planes with the parameters;acquiring data in the start cardiac phase after setting of the parameters;uniformly rotating the two C-arm planes at a speed in a forward motion over an angular area;acquiring data at different angles in the angular area in different cardiac phases during rotating in the forward motion;acquiring data in the start cardiac phase after terminating the forward motion;uniformly rotating the two C-arm planes at the speed in a backward motion over the angular area;acquiring data at the different angles in the angular area in the different cardiac phases during rotating in the backward motion; andreconstructing all of the acquired data for collecting the three-dimensional data of the object after terminating the backward motion upon completed acquisition,wherein the two separate C-arm planes are spaced by a fixed intermediate angle between each other. 2. The method as claimed in claim 1, wherein the two C-arm planes are aligned in the start cardiac phase. 3. The method as claimed in claim 1, wherein the angular area is from 0° to 200° plus minimal of the intermediate angle. 4. The method as claimed in claim 1, wherein the angular area is from 0° to 200° minus minimal of the intermediate angle. 5. The method as claimed in claim 1, wherein the speed for rotating the two C-arm planes is calculated. 6. The method as claimed in claim 1, wherein the two C-arm planes are rotated over the angular area in a cyclical and a sequential arrangement. 7. The method as claimed in claim 6, wherein the two C-arm planes are rotated over the angular area in the cyclical and the sequential arrangement of: rotating one of the two C-arm planes in the forward motion,rotating the one of the two C-arm planes in the backward motion,rotating another one of the two C-arm planes in the forward motion, androtating the another one of the two C-arm planes in the backward motion. 8. The method as claimed in claim 1, wherein the fixed intermediate angle equals two fans. | Siemens Aktiengesellschaft | A61B005/055 A61B006/03
US07862514 | B2 | 2011-01-04 | Blood pressure measurement | 2006-02-06 | The diastolic blood pressure and a systolic blood pressure of a vertebrate are measured using an inflatable pressure cuff disposed about a limb of the vertebrate and connected to an inflation device through a single conduit. A pressure sensor is pneumatically connected to the cuff through the same common conduit. The sensed pressures at the diastolic and systolic points are corrected for a pressure differential associated with the flow of an inflation fluid through the common conduit. The sensed pressures at the diastolic point and the systolic point may be measured during inflation of the cuff or during controlled deflation of the cuff. | 1. A method for determining blood pressure in an artery of a vertebrate at a diastolic point and at a systolic point using an inflatable pressure cuff disposed about a limb of the vertebrate and connected to an inflation device and a pressure sensor through a single conduit having a distal end in fluid communication with said cuff and a proximal end in fluid communication with said inflation device and in pneumatic communication with said pressure sensor, comprising the steps of: inflating said cuff by passing a flow of an inflation fluid through said single conduit;sensing fluid pressure within the proximal end of said conduit at the diastolic point and at the systolic point;determining a pressure differential between said cuff and said pressure sensor at the diastolic point and a pressure differential between said cuff and said pressure sensor at the systolic point; andcorrecting, using a microprocessor, the sensed pressure at the diastolic point using the pressure differential at the diastolic point and correcting the sensed pressure at the systolic point for the pressure differential at the systolic point. 2. A method as recited in claim 1 wherein the step of determining a pressure differential between said cuff and said pressure sensor at the diastolic point and a pressure differential between said cuff and said pressure sensor at the systolic point comprises the steps of: stopping the inflation of said cuff;measuring the pressure differential between said cuff and said pressure sensor after a time delay following the stopping of the cuff inflation process; andcalculating the respective pressure differentials associated with the sensed pressures at the diastolic and systolic points as a function of the measured pressure differential. 3. A method as recited in claim 2 wherein the step of measuring the pressure differential between said cuff and said pressure sensor after a time delay following the stopping of the cuff inflation process comprises the steps of: terminating the inflation of said cuff at a cuff pressure in excess of the systolic point;sensing the fluid pressure within the proximal end of said conduit immediately prior to termination of the cuff inflation process;providing a time delay after termination of the cuff inflation process sufficient for fluid pressure within said cuff and within the proximal end of said conduit to equalize;sensing the fluid pressure within the proximal end of said conduit at the end of said time delay; andsubtracting the fluid pressure sensed at the end of the time delay from the fluid pressure sensed immediately prior to termination of the cuff inflation process. 4. A method as recited in claim 3 wherein the time delay is approximately one second. 5. A method as recited in claim 3 further comprising the step of venting said cuff to atmosphere after said time delay. 6. A method as recited in claim 1 further comprising the step of inflating said cuff to a cuff pressure in excess of the systolic poin... | Welch Allyn, Inc. | A61B005/02
US09795326 | B2 | 2017-10-24 | Continuous analyte measurement systems and systems and methods for implanting them | 2010-07-22 | Low profile continuous analyte measurement systems and systems and methods for implantation within the skin of a patient are provided. | 1. A continuous analyte measurement system, comprising: a base unit configured for mounting on a skin surface;an analyte sensor comprising a first side edge and a second side edge defining a width of the analyte sensor, and a proximal portion and a distal portion defining a length of the analyte sensor, the analyte sensor further comprising a first and a second functional sides on opposing surfaces of the analyte sensor, a first conductive layer disposed on the first functional side at the distal portion of the analyte sensor and disposed from the first side edge to the second side edge, and a sensing layer disposed on the first conductive layer in a continuous band orthogonal to the first side edge and the second side edge to define an analyte sensor active area, the active area having a rectilinear polygon shape defined by the intersection of the first conductive layer and the sensing layer disposed on the first conductive layer, the proximal portion configured for positioning within the base unit and the distal portion configured for insertion through the skin surface;wherein the base unit comprises a two-piece conductive member including a first piece and a second piece; andwherein the base unit comprises a side wall including a slit through which the width of the proximal portion of the analyte sensor is positioned so that the first and second pieces of the conductive member are electrically coupled with the corresponding first and the second functional sides of the proximal portion of the analyte sensor, the base unit further including one or more features on at least one orthogonally adjacent wall to the side wall for exerting a compressive force on each of the first piece and the second piece of the two-piece conductive member, the one or more features aligned with the two-piece conductive member and comprising one or more exterior flexures of the at least one orthogonally adjacent wall to the side wall. 2. The system of claim 1, wherein at least the proximal portion of the analyte sensor has a planar configuration and wherein the conductive member is mechanically and electrically coupled to the first and the second functional sides of the analyte sensor. 3. The system of claim 1, further comprising an alignment pin extending through the proximal portion of the analyte sensor. 4. The system of claim 1, wherein the base unit comprises a non-conductive compressible material. 5. The system of claim 4, wherein the non-conductive compressible material comprises silicone. 6. The system of claim 1, wherein the conductive member comprises a conductive compressible material. 7. The system of claim 6, wherein the conductive compressible material comprises carbon-doped silicone. 8. The system of claim 1, further comprising a transmitter coupled to the base unit, wherein the base unit further comprises a first pair of receptacles for receiving a corresponding first pair of conductive pins of the transmitter, wherein the conductive pins contact... | Abbott Diabetes Care Inc. | A61B005/145 A61B005/1486 A61B005/00
US20110028828 | A1 | 2011-02-03 | T1 RHO MAGNETIC RESONANCE IMAGING FOR STAGING OF HEPATIC FIBROSIS | 2010-07-30 | Methods for diagnosis of fibrotic diseases, staging of fibrotic diseases and monitoring treatment of fibrosis. The presence of fibrotic tissue is detected. First, a T1ρ relaxation time of tissue is determined using magnetic resonance imaging. The determined T1ρ relaxation time is then compared to a baseline T1ρ relaxation time indicative of healthy tissue, and the presence of fibrotic tissue is then determined based on results of said comparison step. To determine a stage of fibrosis, a T1ρ relaxation time of tissue is determined and compared to one or more calibrated T1ρ relaxation times indicative of one or more stages of fibrosis, and the stage of fibrosis is determined based on results of said comparison step. The proposed technology offers a non-invasive MRI technique based on T1ρ contrast that is sensitive enough to detect small changes of ECM hepatic protein concentration and architecture in all stages of hepatic fibrosis. | 1. A method for detecting a presence of fibrotic tissue comprising the steps of: determining a T1ρ relaxation time of tissue using magnetic resonance imaging,comparing the determined T1ρ relaxation time to a baseline T1ρ relaxation time indicative of healthy tissue, anddetermining the presence of fibrotic tissue based on results of said comparison step. 2. A method as claimed in claim 1 wherein the presence of fibrotic tissue is determined from measurement of a T1ρ relaxation time longer than the baseline T1ρ relaxation time. 3. A method as claimed in claim 2, wherein the tissue is liver tissue. 4. A method as claimed in claim 3, wherein the presence of fibrotic tissue indicates liver fibrosis. 5. A method as claimed in claim 3, wherein the presence of fibrotic tissue indicates liver cirrhosis. 6. A method as claimed in claim 1, further comprising the step of determining a stage of fibrosis by comparing the determined T1ρ relaxation time to one or more calibrated T1ρ relaxation times indicative of one or more stages of fibrosis. 7. A method as claimed in claim 6, wherein the step of determining a stage of fibrosis further comprises the step of determining the concentration and architecture of matrix proteins of said fibrotic tissue and comparing the determined concentration and organization to one or more calibrated concentrations and architectures indicative of one or more stages of fibrosis. 8. A method for determining a stage of fibrosis comprising the steps of: determining a T1ρ relaxation time of tissue using magnetic resonance imaging,comparing the determined T1ρ relaxation time to one or more calibrated T1ρ relaxation times indicative of one or more stages of fibrosis, anddetermining the stage of fibrosis based on results of said comparison step. 9. A method as claimed in claim 8 wherein the stage of fibrosis is determined from measurement of a T1ρ relaxation time that falls within a predetermined range of T1ρ relaxation times indicative of a particular stage of fibrosis. 10. A method as claimed in claim 9, wherein the tissue is liver tissue. 11. A method as claimed in claim 9, wherein the fibrosis is liver fibrosis. 12. A method as claimed in claim 9, wherein the fibrosis is liver cirrhosis. 13. A method as claimed in claim 9, wherein the step of determining a stage of fibrosis further comprises the step of determining the concentrations and architecture of said fibrotic tissue and comparing the determined concentrations and architecture to one or more calibrated concentrations and architecture indicative of one or more stages of fibrosis. 14. A method of monitoring progress of a therapeutic treatment of fibrosis comprising the steps of: measuring a baseline value of a T1ρ relaxation time of fibrotic tissue prior to or upon initiation of said treatment using magnetic resonance imaging,measuring one or more additional T1ρ relaxation times after initiation of said treatment,comparing at least one of the one or more additional T... | - | A61B005/055
US20110009760 | A1 | 2011-01-13 | Hospital Readmission Alert for Heart Failure Patients | 2010-05-26 | Physiologic data is collected for a heart failure patient during hospitalization and after discharge of the heart failure patient from a hospital, and analyzed for a change in the collected physiologic data. Deterioration of the patient's heart failure status is detected based at least in part on detecting an acute deleterious change in the collected physiologic data indicative of an abnormality in one or more physiologic parameters represented by the physiologic data. A readmission alert is generated in response to the acute deleterious change in the collected physiologic data. The readmission alert indicates that further deterioration in the patient's heart failure status may require readmission of the patient to the hospital. | 1. A method, comprising: collecting respiration data for a heart failure patient at least during hospitalization and a time after discharge of the heart failure patient from a hospital;measuring, using the collected respiration data, a median respiration rate (MedRR) and a maximum respiration rate (MaxRR);analyzing the MedRR and MaxRR for a change in the MedRR and MaxRR;detecting further deterioration of the patient's heart failure status based at least in part on detecting an acute deleterious change in the MedRR and MaxRR indicative of an abnormality in MedRR and MaxRR; andgenerating a readmission alert in response to detecting the acute deleterious change in the MedRR and MaxRR, the readmission alert indicating that the further deterioration in the patient's heart failure status may require readmission of the patient to the hospital. 2. The method of claim 1, wherein: detecting further deterioration of the patient's heart failure status comprises: determining whether the acute deleterious change in the MedRR exceeds a first threshold; anddetermining whether the acute deleterious change in the MaxRR exceeds a second threshold; andgenerating the readmission alert comprises generating the readmission alert in response to the acute deleterious change in the MedRR and MaxRR exceeding the first and second thresholds, respectively. 3. The method of claim 2, wherein: the first threshold is developed using the MedRR collected before discharge of the heart failure patient from the hospital;determining whether the acute deleterious change in the MedRR exceeds the first threshold is based on a comparison of at least an MedRR collected after discharge of the heart failure patient from the hospital relative to the MedRR collected before discharge;the second threshold is developed using the MaxRR collected before discharge of the heart failure patient from the hospital; anddetermining whether the acute deleterious change in the MaxRR exceeds the second threshold is based on a comparison of at least a MaxRR collected after discharge of the heart failure patient from the hospital relative to the MaxRR collected before discharge. 4. The method of claim 2, wherein each of the first and second thresholds is developed using respiration data collected during a specified period of time prior to patient discharge from the hospital, and the acute deleterious change in the MedRR and MaxRR is detected within about one day to about two months after patient discharge from the hospital. 5. The method of claim 2, wherein each of the first and second thresholds is developed using respiration data collected during a specified period of time prior to patient discharge from the hospital, and the acute deleterious change in the MedRR and MaxRR is detected within about one day to about one month after patient discharge from the hospital. 6. The method of claim 1, comprising: developing a baseline and a threshold for each of the MedRR and MaxRR using respiration data collecte... | - | A61B005/08
US20110021933 | A1 | 2011-01-27 | Methods and Systems for Use in Determination of a Patient's Heart Condition | 2009-03-29 | A method and system are presented for use in monitoring a subject's heart conditions such as ejection fraction and cardiac synchrony. For determination of the ejection fraction, ECG measured data is processed. This ECG measured data comprises a first measured data portion indicative of a first time profile of an electrical signal measured on a planar pair of electrodes applied on a patient during a measurement session and a second measured data portion indicative a second time profile of an electrical signal measured on a diagonal pair of electrodes applied to said patient during said measurement session. The processing comprises comparing the first and second data portions to identify a time shift between a predetermined event in the first and second time profiles. For the determination of heart synchrony, first and second measured data portions corresponding to respectively an ECG measurement and an acoustic measurement both obtained from a patient during the same measurement session of a duration including at least two cardiovascular cycles, are processed. Such processing comprises identifying predetermined first and second events in each of said at least two cycles in the first and second measured data portions, respectively, determining a relation between the first and second events for each of the at least two cycles, determining a degree of fluctuation of said relation in the at least two cycles. | 1. A method for use in monitoring a subject's ejection fraction, the method comprising: processing ECG measured data which comprises a first measured data portion indicative of a first time profile of an electrical signal measured on a planar pair of electrodes applied on a patient during a measurement session and a second measured data portion indicative a second time profile of an electrical signal measured on a diagonal pair of electrodes applied to said patient during said measurement session, said processing comprising comparing the first and second data portions and upon identifying a time shift between a predetermined event in the first and second time profiles, generating output data indicative of the ejection fraction for said subject. 2. A method according to claim 1, wherein said event is a signal peak. 3. A method according to claim 1, wherein a time delay of appearance of said event in the second time profile as compared to that of the first time profile is indicative of reduction in the ejection fraction. 4. A method according to claim 1, wherein said processing comprises processing third measured data indicative of acoustic signals recorded by an array of acoustic receivers from a planar region of the chest or back skin surface of the subject overlying a heart chamber whose ejection fraction is to be determined, and generating output data indicative of the ejection fraction of said heart chamber. 5. A monitoring system for monitoring a subject's ejection fraction, the system comprising: a data processing and analyzing utility configured for processing ECG measured data which comprises a first measured data portion indicative of a first time profile of an electrical signal measured on a planar pair of electrodes applied to a patient during a measurement session and a second measured data portion indicative a second time profile of an electrical signal measured on a diagonal pair of electrodes applied to said patient during said measurement session, said processing comprising comparing the first and second measured data portions and upon identifying a time shift between a certain event in the first and second time profiles, generating output data indicative of the ejection fraction for said patient. 6. A system according to claim 5, comprising a measurement unit comprising an electrodes arrangement configured and operable for carrying out the ECG measurement session. 7. A system according to claim 5, wherein said data processing and analyzing utility is configured for processing third measured data indicative of acoustic signals recorded by an array of acoustic receivers from a planar region of the chest or back skin surface of the subject overlying a heart chamber whose ejection fraction is to be determined, and generating output data indicative of the ejection fraction of said heart chamber. 8. A method for use in monitoring a patient's heart condition, the method comprising: processing first and second measured data portion... | DeepBreeze Ltd. | A61B005/0452 A61B007/04
US08380275 | B2 | 2013-02-19 | Systems and methods for making and using medical ablation systems and having mapping catheters with improved anchoring ability | 2010-08-11 | A mapping catheter includes an elongated body for inserting into patient vasculature. A distal end of the elongated body includes a distal portion that includes a plurality of electrodes, a proximal portion disposed proximal to the distal portion, and a reduced-dimension portion disposed between the proximal and distal portions. The distal end is formed, at least in part, from a memory shape material that bends into a preformed shape upon release from a confined space. The preformed shape includes a first loop formed, at least in part, by the distal portion. The first loop is transverse to a longitudinal axis of the proximal portion. The reduced-dimension portion is configured and arranged to bend such that the reduced-dimension section advances distally through the first loop when the first loop is held in a fixed position and a force is applied distally along the longitudinal axis of the proximal portion. | 1. A mapping catheter comprising: an elongated body configured and arranged for insertion into patient vasculature, wherein adistal end of the elongated body comprises: a distal portion comprising a plurality of electrodes,a proximal portion disposed proximal to the distal portion, anda reduced-dimension portion disposed between the proximal portion and the distal portion, the reduced-dimension portion having a cross-sectional dimension that is less than corresponding cross-sectional dimensions of both a proximally-positioned adjacent section of the proximal portion and a distally-positioned adjacent section of the distal portion;wherein the distal end is formed, at least in part, from a memory shape material that bends into a preformed shape upon release from a confined space, the preformed shape comprising a first loop formed, at least in part, by the distal portion, wherein the first loop is transverse to a longitudinal axis of the proximal portion,wherein the reduced-dimension portion is configured and arranged to bend such that the reduced-dimension portion advances distally through the first loop when the first loop is held in a fixed position and a force is applied along the longitudinal axis of the proximal portion in a distal direction. 2. The mapping catheter of claim 1, wherein the distal portion of the mapping catheter tapers in a distal direction. 3. The mapping catheter of claim 1, wherein the reduced-dimension portion of the mapping catheter tapers in a proximal direction. 4. The mapping catheter of claim 1, wherein the reduced-dimension portion is configured and arranged to bend preferentially in a dimension such that the proximally-positioned adjacent section of the proximal portion advances distally through the first loop when the first loop is held in a fixed position and a force is applied along the axis of the proximal portion in a distal direction. 5. The mapping catheter of claim 1, wherein the preformed shape comprises a second loop formed, at least in part, by the reduced-dimension portion. 6. The mapping catheter of claim 5, wherein the first loop and the second loop are parallel. 7. The mapping catheter of claim 5, wherein the first loop has a transverse diameter that is larger than a transverse diameter of the second loop. 8. The mapping catheter of claim 1, wherein the reduced-dimension portion is disposed proximal to the plurality of electrodes. 9. The mapping catheter of claim 1, wherein the reduced-dimension portion has a transverse cross-sectional shape that is round. 10. The mapping catheter of claim 1, wherein the reduced-dimension portion has a transverse cross-sectional shape that is rectangular and comprises a first dimension and a second dimension that is perpendicular to the first dimension and that is no larger than the first dimension. 11. The mapping catheter of claim 10, wherein the first dimension of the reduced-dimension portion is no more than ten times larger than the second dimension. 12... | Boston Scientific Scimed, Inc. | A61B005/042 A61B018/02 A61B018/14
US08983158 | B2 | 2015-03-17 | System for performing coronary digital subtraction angiography (DSA) | 2010-03-31 | An image data subtraction system receives an electrical signal representing a heart cycle electrical waveform during multiple heart cycles and acquires data representing a first image set comprising multiple temporally sequential individual mask images of vessels of a portion of patient anatomy during the multiple heart cycles in the absence of a contrast agent. The system acquires data representing a second image set comprising a multiple temporally sequential individual contrast enhanced images of vessels of the portion of patient anatomy during the multiple heart cycles in the presence of a contrast agent. An image data processor automatically uses the electrical signal to identify temporally corresponding pairs of images comprising a mask image and a contrast enhanced image acquired substantially at a same point within a heart cycle. The image data processor, for the corresponding pairs, automatically subtracts data representing a mask image of a corresponding pair from a contrast enhanced image of the corresponding pair, to provide subtracted images. | 1. An image data subtraction system to enhance visualization of vessels subject to movement, comprising: an interface for receiving an electrical signal representing a heart cycle electrical waveform occurring over a plurality of heart cycles;an imaging system for,acquiring data representing a first image set comprising a plurality of temporally sequential individual mask images of vessels of a portion of patient anatomy during said plurality of heart cycles in the absence of a contrast agent, andacquiring data representing a second image set comprising a plurality of temporally sequential individual contrast enhanced images of vessels of said portion of patient anatomy during said plurality of heart cycles in the presence of a contrast agent, wherein each individual mask image of the first image set is acquired at one of multiple segments within the heart cycle and each contrast enhanced image of the second image set is acquired at one of the same segments within the heart cycle; andan image data processor for automatically,using said electrical signal to identify corresponding pairs of images, wherein each of the corresponding pairs comprises a mask image and a contrast enhanced image acquired at a same segment within the heart cycle, andfor each of said corresponding pairs, subtracting data representing a mask image of a corresponding pair from a contrast enhanced image of said corresponding pair, to provide a plurality of subtracted images providing enhanced visualization of vessels. 2. A system according to claim 1, wherein said vessels are Coronary vessels. 3. A system according to claim 1, wherein said imaging system acquires said individual mask images in response to a synchronization signal triggering image acquisition at predetermined segments during the heart cycle and synchronized with said electrical signal. 4. A system according to claim 1, wherein said portion of patient anatomy comprises a heart and said vessels comprise coronary vessels. 5. A system according to claim 1, including a display processor for processing data representing said subtracted images for display. 6. A system according to claim 1, wherein said electrical signal is an ECG signal. 7. An image data subtraction system to enhance visualization of vessels subject to movement, comprising: an interface for receiving an electrical signal representing a heart cycle electrical waveform during a plurality of heart cycles;an imaging system for,acquiring data representing a first image set in the absence of a contrast agent and comprising a plurality of temporally sequential individual mask images of vessels of a portion of patient anatomy, wherein each individual mask image of the first image set is acquired at one of multiple segments within a heart cycle based on said electrical signal, andacquiring data representing a second image set in the presence of a contrast agent and comprising a plurality of temporally sequential individual contrast enhanced images of ves... | Siemens Medical Solutions USA, Inc. | G06K009/00 A61B006/00 A61B005/02 G06T005/50 A61B005/00
US09216262 | B2 | 2015-12-22 | Device for detecting the observance by a patient of an oxygen-therapy treatment | 2009-04-10 | The invention relates to a device for implementing a system for detecting the observance by a patient of an oxygen-therapy treatment involving an oxygen supply, and for recording the data of said treatment and transmitting said data. The invention also relates to the use of said device, and to a method for tracking the observance of an oxygen-therapy treatment. | 1. A method for tracking the observance of an oxygen-therapy treatment by measuring the duration of treatment of the patient, wherein the method comprises the steps of: a) providing a device for detecting the observance of an oxygen-therapy treatment with feeding of oxygen to a patient via a cannula delivering oxygen to the airways of the patient, the device allowing the recording of the data of the treatment and the transmission of the data, the device comprising two sensors of absolute pressure, one for measuring the pressure in the cannula, the other for measuring the atmospheric pressure, or a differential pressure sensor; a micro-processor implementing an algorithm capable of transforming the pressure measurements into respiratory frequencies of the patient and of deducing therefrom the duration of the daily oxygen-therapy treatment; a memory chip making it possible to record the data; a radiofrequency antenna making it possible to ensure the wireless transmission of the data; and a battery allowing the system to operate in complete autonomy;b) detecting an oxygen flow rate by measuring a pressure in the cannula supplying the patient's airways with oxygen,c) detecting the patient's respiration by measuring a pressure variation in the cannula, andd) when an oxygen-therapy treatment is detected, transforming the pressure values into respiratory frequency data with the aid of an algorithm and transmitting the respiratory frequency data with the radiofrequency antenna, wherein the method further comprises a step of determining a time period t1, wherein the data regarding the pressure in the cannula are measured during a first plurality of successive daily oxygen-therapy treatment periods (t1),the data measured during said first plurality of daily oxygen-therapy treatment periods (t1) are utilized by said algorithm during a second plurality of daily oxygen-therapy treatment period (t2) during which no other additional pressure datum is recorded, the second plurality of daily oxygen-therapy treatment periods (t2) requiring consumption of energy for the calculation serving to transform the pressure values into respiratory frequency, and wherein t2 is a multiple of t1 and wherein each of the first plurality of successive daily oxygen-therapy treatment periods (t1) is followed in sequence by one of the second daily oxygen-therapy treatment periods (t2). 2. The method of claim 1, wherein the transformation of the pressure values into respiratory frequency is done with the aid of an FFT algorithm or of a TDS algorithm. 3. The method of claim 1, wherein the algorithm is designed to filter the interferences caused by the oxygen sources. 4. The method of claim 1, wherein the measurement period (t1) lies between 20 and 120 seconds. 5. The method of claim 2, wherein the algorithm is designed to filter the interferences caused by the oxygen sources. 6. The method of claim 5, wherein the measurement period (t1) lies between 20 and 120 seconds. | L'Air Liquide, Société Anonyme pour l'Etude et l'Exploitation des Procédés Georges Claude | A61B005/02 A61M016/00 A61B005/08 A61M016/06 A61M016/10 A61B005/00
US07877130 | B2 | 2011-01-25 | System and method for using delayed enhancement magnetic resonance imaging and artificial intelligence to identify non-viable myocardial tissue | 2003-10-02 | A system and method for imaging and identifying non-viable myocardial tissue in a patient's myocardium is disclosed. Images of a section of the myocardium are obtained. An endocardial border and epicardial border of the section of the myocardium is segmented. The section of the myocardium is divided into sectors. One or more selected features of the sectors of the myocardial wall are measured and applied to a decision surface. A determination is made as to whether each sector contains viable or non-viable myocardial tissue. An image that shows each sector of the myocardial wall and an indication of its viability is displayed. | 1. A computer implemented method for imaging and identifying non-viable myocardial tissue in a patient's heart, the method performed by the computer comprising the steps of: obtaining images of a section of the myocardium;segmenting an endocardial border and epicardial border of the section of the myocardium;dividing the segment of the section of the myocardium into sectors;measuring one or more a plurality of selected features of the sectors of the myocardium;applying the measured plurality of one or more features to a decision function to create a decision surface, wherein the decision function that creates said decision surface is a function ƒ(x) that maps a multidimensional vector x formed of said selected feature measurements to a scalar real value indicative of whether a tissue sector is viable or non-viable;determining from said decision surface whether each sector contains viable or non-viable myocardial tissue; anddisplaying an image that shows each sector of the myocardium and an indication of its viability,wherein said steps of obtaining images, segmenting an endocardial border and epicardial border, dividing the section into sectors, measuring a plurality of selected features, determining whether each sector contains viable or non-viable myocardial tissue, and displaying an image are perfoimed by a computer processor. 2. The method of claim 1 wherein the images obtained are using Delayed Enhanced Magnetic Resonance (DEMR) imaging. 3. The method of claim 1 wherein the section of the myocardium is a left ventricle. 4. The method of claim 1 wherein said features are selected from a group comprising image intensity, thickness of the myocardium, differential thickness of the myocardium, homogeneity of the myocardium, transmurality of scar (non-viable) tissue, a position of the sector of tissue, a rate of contrast agent uptake, and material strain information from tagged-MR. 5. The method of claim 1 wherein said decision surface separates measurements identifying viable tissue sectors from non-viable tissue sectors. 6. The method of claim 1 wherein sectors of viable tissue are displayed as a first color and sectors of non-viable tissue are displayed as a second color. 7. A program storage device readable by a computer, tangibly embodying a program of instructions executable by the computer to perform the method steps for imaging and identifying non-viable myocardial tissue in a patient's heart, the method comprising the steps of: obtaining images of a section of the myocardium;segmenting the section of the patient's myocardium into an endocardial border and an epicardial border;dividing the a myocardial wall of the section of the myocardium into sectors;measuring one or more selected features of the sectors of the myocardial wall;applying the measured one or more features to a decision function to create a decision surface, wherein the decision function that f creates said decision surface is a function ƒ(x) that maps a multidimension... | Siemens Medical Solutions USA, Inc. | A61B005/05
US08750954 | B2 | 2014-06-10 | Medical monitoring patch device and methods | 2009-03-31 | Embodiments described herein may include systems and methods for monitoring physiological parameters of a patient. Specifically, embodiments disclose the use of a generally self-contained pulse oximeter that is small and lightweight, such that it may be comfortably affixed to a patient to provide physiological data pertaining to the patient. Embodiments also provide methods of using and manufacturing a pulse oximetry patch. | 1. A pulse oximetry patch adapted for placement adjacent to tissue to be tested, comprising: an upper layer comprising a display window;a bottom layer configured to be disposed on a tissue measurement site of a patient; anda middle layer disposed between the upper layer and the bottom layer comprising: one or more light emitters and one or more light detectors, the emitters and detectors configured to acquire physiological data from the patient;a drive engine configured to activate the one or more light emitters;an oximetry engine configured to receive a signal corresponding to the physiological data acquired by the detectors and generate an output corresponding to the physiological data; anda display adapted to receive the output corresponding to the physiological data; andwherein the upper layer and the bottom layer are configured to be separated from each other to enable the removal of the middle layer, and wherein the upper and bottom layers are disposable and the middle layer is reusable. 2. The pulse oximetry patch of claim 1, wherein the bottom layer comprises a layer of adhesive for attaching the pulse oximetry patch to the patient. 3. The pulse oximetry patch of claim 1, wherein the display comprises a display configured to display a numerical value. 4. The pulse oximetry patch of claim 1, wherein the display comprises one or more light emitting diodes. 5. The pulse oximetry patch of claim 4, wherein one of the one or more light emitting diodes blinks correlatively with the heartbeat of the patient. 6. The pulse oximetry patch of claim 4, wherein one of the one or more light emitting diodes is activated or deactivated in response to a comparison of a measured blood-oxygen saturation value and a programmed blood-oxygen saturation value. 7. The pulse oximetry patch of claim 6, wherein the one or more light emitting diodes are color coded. 8. The pulse oximetry patch of claim 1, comprising a wireless transmitter configured to transmit patient parameters to an external monitor, display, and/or other device. 9. The pulse oximetry patch of claim 1, wherein the drive engine, the oximetry engine, the display, or any combination thereof, are from a used pulse oximetry patch, and wherein the upper layer, the bottom layer, or a combination thereof, are new. 10. The pulse oximetry patch of claim 1, wherein the bottom layer comprises a window configured to align with the one or more light emitters and the one or more light detectors. 11. A pulse oximeter comprising: a single processor chip comprising: one or more first converters configured to activate one or more signal emitters for transmission of an emitted signal into tissue of a patient;one or more second converters configured to receive a signal generally correlative to the emitted signal from the tissue of the patient from one or more signal detectors;a core processor configured to calculate physiological parameters of the patient from the signal from the one or more signal detectors... | Covidien LP | A61B005/1455
US09113802 | B2 | 2015-08-25 | Method of stimulating and recording from ulnar dorsal cutaneous sensory nerve | 2009-07-12 | A device to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a patient may include a first surface electrode to apply a stimulating signal to the ulnar dorsal cutaneous nerve, a second surface electrode to record the response from the stimulating signal after applying the stimulating signal to the ulnar dorsal cutaneous nerve, and a third surface electrode to reference the stimulating signal. The second surface electrode may be placed over the fifth metacarpal joint, and the device may include a ground electrode positioned over the dorsal surface of the hand. The device may include a inactive stimulating electrode positioned near to the first surface electrode. The inactive stimulating electrode may be positioned within 3 cm of the first surface electrode, and the first electrode may be placed substantially above the tendon the flexor carpi ulnaris muscle. The first electrode may be placed substantially at the level of styloid, and the second surface electrode may be placed in an area supplied by the ulnar dorsal cutaneous nerve. The first surface electrode may be placed substantially proximal to the styloid. | 1. A method to stimulate and measure a response from an ulnar dorsal cutaneous nerve of a hand, comprising the steps of: positioning a first surface electrode in a medial area of a forearm of the hand between a tendon of a flexor carpi ulnaris muscle and a styloid process of an ulna bone during flexion of a wrist of the hand at an acute angle with respect to the forearm to increase contact area with the first surface electrode;positioning a second surface electrode on a dorsal surface of the hand over a fifth metacarpal joint;using the first surface electrode to apply a stimulating signal to the ulnar dorsal cutaneous nerve;using the second surface electrode to record the response from the stimulating signal after applying the stimulating signal to the ulnar dorsal cutaneous nerve; andusing a third surface electrode to reference the stimulating signal. 2. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 1, wherein the method includes the step of positioning a ground electrode over the dorsal surface of the hand. 3. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 1, wherein the method includes the step of positioning an inactive stimulating electrode near to the first surface electrode. 4. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 3, wherein the method includes the step of positioning the inactive stimulating electrode within 3 cm of the first surface electrode. 5. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 1, wherein the flexion of the wrist at the acute angle with respect to the forearm renders the ulnar dorsal cutaneous nerve more superficial. 6. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 5, wherein the first surface electrode is placed substantially at the level of styloid. 7. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 5, wherein the first surface electrode is placed substantially proximal to styloid. 8. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 1, wherein the second surface electrode is placed in an area supplied by the ulnar dorsal cutaneous nerve. 9. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 1, wherein the flexion of the wrist at the acute angle with respect to the forearm is in a range of 1 to 60 degrees. 10. A method to stimulate and measure a response from the ulnar dorsal cutaneous nerve of a hand as in claim 1, wherein the first surface electrode and the second surface electrode are reversed in position. |  | A61B005/05 A61B005/04 A61N001/00 A61B005/11 A61B005/00
US07881780 | B2 | 2011-02-01 | Biological interface system with thresholded configuration | 2005-12-30 | A system and method for a biological interface system that processes multicellular signals of a patient and controls one or more devices is disclosed. The system includes a sensor that detects the multicellular signals and a processing unit for producing the control signal based on the multicellular signals. The system further includes an automated configuration routine that is used to set or modify the value of one or more system configuration parameters. | 1. A biological interface system comprising: a sensor comprising a plurality of electrodes for detecting multicellular signals emanating from one or more living cells of a patient;a processing unit configured to receive the multicellular signals from the sensor, process the multicellular signals to produce a processed signal, and transmit the processed signal to a controlled device;the controlled device for receiving the processed signal; andan integrated patient training routine performed to generate one or more system configuration parameters used by the processing unit to produce the processed signal, the patient training routine comprising: a system configuration plan comprising a first configuration of steps to be performed;means for collecting a first set of patient data, the first set of patient data collected prior to patient control of the controlled device;means for analyzing the first set of patient data and providing an output of the analysis; andmeans for modifying the system configuration plan; wherein the system configuration plan is modified when the output of the analysis of the first set of patient data falls below a success threshold value, andwherein the success threshold value is set to require at least one modification of the configuration plan prior to the patient controlling the controlled device. 2. The system of claim 1, wherein the first set of patient data does not include data regarding control of the controlled device. 3. The system of claim 1, wherein the first configuration of steps includes a cell sorting algorithm to select a set of cells for processing. 4. The system of claim 3, wherein the cell sorting algorithm is a neuron spike sorting algorithm. 5. The system of claim 3, wherein the success threshold value is a minimum number of cells selected for processing. 6. The system of claim 1, wherein the first configuration of steps includes a measurement and analysis of one or more cellular signal modulation rates. 7. The system of claim 1, wherein the first configuration of steps includes a measurement and analysis of neural signals. 8. The system of claim 7, wherein the output of the analysis includes a quantification of neural spikes and the threshold value is a minimum number of spikes. 9. The system of claim 7, wherein the output of the analysis includes a cellular signal modulation rate quantification and the threshold value is a minimum level for the quantification. 10. The system of claim 7, wherein the output of the analysis includes a multicellular signal modulation rate quantification and the threshold value is a minimum level for the quantification. 11. The system of claim 1, wherein the first configuration of steps includes a measurement and analysis of one or more patient physiologic parameters. 12. The system of claim 11, wherein the patient physiologic parameter is selected from the group consisting of: EKG; respiration; blood glucose; temperature; blood pressure; EEG; perspiration; and c... | BrainGate Co., LLC | A61B005/00
US07874994 | B2 | 2011-01-25 | Method and apparatus for penetrating tissue | 2006-10-16 | An improved penetrating tissue system is provided that reduces pain when penetrating tissue by providing controlled depth and velocity of penetration, stimulation and pressure to the targeted tissue without breaching a sterile environment. The system has user interfaces, human interfaces and low volume sample chambers. | 1. A body fluid sampling system for use on a tissue site, the system comprising: a controllable electromagnetic based penetrating member driver;a plurality of penetrating members housed in or on a cartridge; anda transport mechanism configured to carry the plurality of penetrating members on top of a belt, and move the plurality of penetrating members in its entirety to a position for one of the penetrating members to be aligned with the controllable electromagnetic based penetrating member driver to drive the penetrating member along a path into the tissue site, wherein each of a penetrating member is operatively engaged with the penetrating member driver when moved into position by the transport mechanism. 2. The system of claim 1, wherein the transport mechanism has a configuration selected from at least one of a disk, a drum, a bandolier and a magazine. 3. A body fluid sampling system for use on a tissue site, the system comprising: a controllable electromagnetic drive force generator;a penetrating member operatively coupled to said force generator, housed in or on a cartridge; said force generator moving said member along a path out of a housing having the penetrating member exit; into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site; anda tissue stabilizer device suitable for stretching a surface of a tissue site, said tissue stabilizer at least partially surrounding the penetrating member exit;a transport mechanism configured to carry the penetrating member on top of a belt, and move the penetrating member to a position for the penetrating member to be aligned with the controllable electromagnetic drive force generator to drive the penetrating member along a path into the tissue site, wherein each of the penetrating member is operatively engaged with the controllable electromagnetic drive force generator when moved into position by the transport mechanism. 4. The system of claim 3, wherein the transport mechanism has a configuration selected from at least one of a disk, a drum, a bandolier and a magazine. 5. A body fluid sampling system for use on a tissue site, the system comprising: a controllable electromagnetic drive force generator;a penetrating member operatively coupled to said force generator and housed in or on a cartridge, said force generator moving said member along a path out of a housing having the penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site; anda pressure applicator coupled to the housing;a transport mechanism configured to carry the penetrating member on top of a belt, and move the penetrating member to a position for the penetrating member to be aligned with the controllable electromagnetic drive force generator to drive the penetrating member along a path into the tissue site, wherein the penetrating member is operatively engaged with the controllable electromagnetic drive force generator when moved into po... | Pelikan Technologies, Inc. | A61B005/00
US08255027 | B2 | 2012-08-28 | Multiple wavelength sensor substrate | 2010-07-19 | A physiological sensor has emitters configured to transmit optical radiation having multiple wavelengths in response to corresponding drive currents. A thermal mass is disposed proximate the emitters so as to stabilize a bulk temperature for the emitters. A temperature sensor is thermally coupled to the thermal mass. The temperature sensor provides a temperature sensor output responsive to the bulk temperature so that the wavelengths are determinable as a function of the drive currents and the bulk temperature. | 1. A physiological sensor comprising: a plurality of emitters configured to transmit optical radiation having a plurality of wavelengths in response to a corresponding plurality of drive currents;a thermal mass disposed proximate the emitters so as to stabilize a bulk temperature for the emitters; anda temperature sensor thermally coupled to the thermal mass, the temperature sensor providing a temperature sensor output responsive to the bulk temperature so that the wavelengths are determinable as a function of the drive currents and the bulk temperature; andand an optical filter, wherein the optical filter attenuates at least a portion of a wavelength spectrum so as to compensate for an imbalance in an attenuation of the plurality of wavelengths by the tissue site. 2. The physiological sensor according to claim 1 further comprising a substrate having a first side and a second side, wherein the emitters are mounted to the first side, andwherein the temperature sensor is mounted to the second side. 3. The physiological sensor according to claim 2 wherein the temperature sensor is a thermistor and the emitters are LEDs. 4. The physiological sensor according to claim 3: wherein the thermal mass is a plurality of layers of the substrate, andwherein each of the layers is substantially copper clad. 5. A physiological sensor capable of emitting light into tissue and producing an output signal usable to determine one or more physiological parameters of a patient, the physiological sensor comprising: a thermal mass disposed within a substrate;a plurality of light emitting sources disposed on the substrate proximate the thermal mass and thermally coupled to the thermal mass, the sources having a corresponding plurality of operating wavelengths;a temperature sensor disposed on the substrate proximate the thermal mass and thermally coupled to the thermal mass and capable of determining a bulk temperature for the thermal mass, the operating wavelengths dependent on the bulk temperature;an optical filter configured to attenuate at least one of the plurality of light emitting sources; anda detector capable of detecting light emitted by the light emitting sources after tissue attenuation, wherein the detector is capable of outputting a signal usable to determine one or more physiological parameters of a patient based upon the operating wavelengths. 6. The physiological sensor according to claim 5 wherein the temperature sensor comprises a thermistor. 7. The physiological sensor according to claim 6 wherein the light emitting sources are disposed on a first side of the substrate and the temperature sensor is disposed on a second side of the substrate. | Cercacor Laboratories, Inc. | A61B005/1455
US07878985 | B2 | 2011-02-01 | Guidewire having linear change in stiffness | 2003-09-11 | The invention is directed to a guidewire having a distal section with multiple distally tapered core segments with at least two contiguous distally tapering core segments in which the most distal tapered core segment preferably has a greater degree of taper than the proximally contiguous tapered core segment. The invention is also directed to an elongated intracorporeal device, preferably a guidewire or section thereof, that has a core member with a plurality of contiguous tapered segments having taper angles that are configured to produce a linear change in stiffness over a longitudinal section of the device. The device may also have a core section with a continuously changing taper angle to produce a curvilinear profile that preferably is configured to produce a linear change in stiffness of the core over a longitudinal section of the device. | 1. An intracorporeal device comprising an elongated member having a longitudinal length and means for causing a substantially linear change in bending stiffness over the entire longitudinal length of the elongated member, wherein the elongated member has a plurality of tapered segments configured to produce the substantially linear change in bending stiffness over the length of the member. 2. The device of claim 1 wherein the length of the elongated member has a continuously changing taper angle producing a curvilinear profile that is configured to produce the substantially linear change in bending stiffness over said length. 3. The device of claim 1 wherein each tapered segment has a substantially constant taper angle. 4. The device of claim 1 wherein the elongate core member comprises at least 3 to about 100 tapered segments. 5. The device of claim 1 wherein the elongated core member comprises a material with changing hardness in a longitudinal direction configured such that the change in hardness produces a substantially linear change in bending stiffness along the length of the core member. 6. The device of claim 1 wherein the elongated member tapers distally to a more flexible distal portion. 7. A guidewire comprising an elongate core member with at least one longitudinal section having a tapering diameter and a substantially linear change in bending stiffness over a longitudinal length thereof and defined substantially by the formula:DL=[64⁢CLE⁢⁢π+D04]14Where DL is the diameter of the elongate core member at length L from a position of starting diameter D0, L is a length greater than zero, E is the modulus of elasticity of the core member material, and C is a constant that depends on the boundary conditions of the longitudinal section. 8. A guidewire as defined in claim 7, wherein said core member is formed of one of the group constituting stainless steel, NiTi alloys and combinations thereof. 9. A guidewire as defined in claim 7, wherein said core member has a proximal core section, said proximal core section being coated with a lubricious coating. 10. A guidewire as defined in claim 7, wherein said core member has a distal core section, said distal core section being coated with a lubricious coating. 11. A guidewire as defined in claim 7, wherein said core member has a distal core section, and a flexible body disposed about and secured to the distal core section. 12. A guidewire as defined in claim 7, wherein the guidewire comprises at least 3 to about 100 tapered segments. 13. A guidewire as defined in claim 7, wherein the guidewire comprises at least 5 to about 20 tapered segments. 14. A guidewire comprising an elongate core member with at least one longitudinal section having a substantially linear change in bending stiffness over a longitudinal length thereof and a moment of inertia defined substantially by the formula:IL=C⁢⁢LE+I0Where IL is the moment of inertia of the longitudinal section at length L from a position of ... | Advanced Cardiovascular Systems, Inc. | A61B005/00
US07883472 | B2 | 2011-02-08 | Method and system of measuring IAP using a naso-enteric tube | 2009-07-08 | A method and apparatus for measuring the intra-abdominal pressure (IAP) within a patient having a naso-enteric tube. The system senses the IAP of the patient by means of a sensing conduit that is filled with a sensing fluid. The sensing fluid in the sensing conduit may be saline, water or feeding solution. After the sensing conduit is filled with the sensing fluid, the sensing conduit is elevated and a portion of the sensing fluid enters the patient. The remaining portion of the sensing fluid is related to the IAP and the IAP may be determined from the height of the sensing fluid within the sensing conduit. | 1. A method of determining the intra-abdominal pressure (IAP) of a patient having a naso-enteric tube to supply a liquid feed solution to a patient, the method comprising the steps of: providing a supply conduit to supply the liquid feed solution from a supply of liquid feed solution to the patient via the naso-enteric tube;positioning a sensing conduit in fluid communication with the naso-enteric tube;filling the sensing conduit with a supply of a sensing fluid separate from the supply of liquid feed solution;elevating at least a portion of sensing conduit above a mid-axillary line of the patient;determining the height of the sensing fluid within the sensing conduit above the mid-axillary line of the patient; anddetermining the IAP based upon the height of the sensing fluid. 2. The method of claim 1 wherein the sensing conduit extends between a first end and a second end, the first end being in communication with the naso-enteric tube and the second end being vented to atmosphere. 3. The method of claim 1 wherein the sensing fluid is injected into the sensing conduit intermittently during the supply of the liquid feed solution to the patient. 4. The method of claim 3 wherein the sensing fluid is injected into the sensing conduit from a syringe. 5. The method of claim 1 further comprising the steps of: positioning an adapter between the naso-enteric feed tube and the supply conduit that receives the liquid feed solution;connecting the sensing conduit to a first port contained on the adapter; andinjecting the sensing fluid through a second access port contained on the adapter such that the sensing fluid fill the sensing conduit. 6. The method of claim 5 wherein the sensing fluid is injected into second access port from a syringe. 7. The method of claim 1 further comprising the step of temporarily interrupting the supply of the liquid feed solution to the patient and filling the sensing conduit with the supply of sensing fluid during the temporary interruption. 8. The method of claim 6 wherein the second access port is opened by the syringe. 9. A method of determining the intra-abdominal pressure (IAP) of a patient having a naso-enteric feeding tube to supply a liquid feed solution to the patient, the method comprising the steps of: pumping the liquid feed solution into the patient through the naso-enteric feeding tube;positioning a sensing conduit in fluid communication with the supply of liquid feed solution pumped into the patient;interrupting the pumping of the liquid feed solution into the patient;elevating at a least a portion of the sensing conduit above a mid-axillary line of the patient;determining the height of the feed solution within the sensing conduit above the mid-axillary line of the patient during the interruption of the pumping of the liquid feed solution; anddetermining the IAP based upon the height of the feed solution. 10. The method of claim 9 wherein the sensing conduit extends between a first end and a second end, t... | Holtech Medical | A61B005/00
US07883464 | B2 | 2011-02-08 | Integrated transmitter unit and sensor introducer mechanism and methods of use | 2005-09-30 | Method and apparatus for providing an integrated transmitter unit and sensor insertion mechanism is provided. | 1. A system, comprising: an integrated housing having electronics including a transmitter unit and a sensor; the integrated housing configured for mounting onto a skin of a patient; anda sensor introducer assembly mounted onto the integrated housing, the introducer assembly configured to position at least a portion of the sensor under the skin of the patient; wherein the transmitter unit is coupled to and in electrical contact with the sensor, and configured to transmit one or more signals corresponding a respective one or more signals received from the sensor;wherein a portion of the electronics is configured to be selectively disabled after a predetermined time period has elapsed since coupling to the sensor;wherein the integrated housing further comprises at least one sensor insertion location for extending the portion of the sensor therethrough; and furtherwherein the integrated housing further comprises a plurality of sensor insertion locations. 2. The system of claim 1 further including a receiver unit configured to receive the one or more signals transmitted by the transmitter unit. 3. The system of claim 2 wherein the receiver unit includes one of an infusion pump, a monitoring device, a personal digital assistant, a pager, or a mobile telephone. 4. A system, comprising: a housing configured for mounting onto a skin of a patient and comprising at least one sensor insertion location therethrough;a sensor having a portion configured to be positioned through the at least one sensor insertion location and then within the skin of a patient; anda transmitter integrally provided within the housing, wherein the transmitter is electrically coupled to the sensor by an electrical connection within the housing extending between the transmitter and the at least one sensor insertion location when the sensor is operatively positioned within the skin and wherein the electrical connection is configured to be selectively disabled after a predetermined time period has elapsed since coupling to the sensor;wherein the housing comprises a plurality of sensor insertion locations and a plurality electrical connections, and further wherein there is a one for one correlation between the locations and the number of electrical connections. 5. The system of claim 4 further comprising a sensor introducer configured to transcutaneously position at least a portion of the sensor within the skin of the patient. 6. The system of claim 5 wherein the introducer is configured to be removed from the system after insertion of the sensor within the skin. 7. The system of claim 5 wherein the introducer is configured to be substantially housed within the system after insertion of the sensor within the skin. 8. The system of claim 4 wherein the housing is flexible to substantially conform to the shape of a body portion to which it is mounted. 9. A method, comprising: securing a housing substantially around a body part of a patient, the housing comprising an integrated elect... | Abbott Diabetes Care Inc. | A61B005/00
US08535223 | B2 | 2013-09-17 | Wireless patient monitoring using streaming of medical data with body-coupled communication | 2009-02-19 | The invention relates to a system and method for wirelessly monitoring a patient (4), the system comprising a medical sensor (1) for collecting medical data and a coupling interface (2) for coupling with a monitor (3). The medical sensor (1) is attached to the patient's (4) body, and the medical data collected by the medical sensor (1) is transmitted to the coupling interface (2) via body-coupled communication (5). Preferably, the collected medical data is streamed to the coupling interface (2) via body-coupled communication (5). This way, a secure and reliable system and method for wireless patient monitoring are achieved which only require little power. | 1. A system for wirelessly monitoring a patient, the system comprising: at least one medical sensor which collects medical data from the patient and wirelessly streams the medical data over a patient body using body-coupled communication with a streaming protocol including at least one of real time transport protocol (RTP) and user datagram protocol (UDP), the medical sensor including at least one of: an electrocardiogram (ECG);a photoplethysmograph (PPG);a pulse oximetry (SpO2); andan electroencephalography (EEG);a coupling interface which receives the streamed medical data which is streamed over the patient body via the body-coupled communication; anda monitor which receives the streamed medical data from the coupling interface and displays the received streamed medical data. 2. The system according to claim 1, wherein the medical data is collected by multiple medical sensors and the coupling interface acts as a gateway for collecting the medical data streamed from the multiple medical sensors and retransmitting the medical data to the monitor. 3. The system according to claim 1, wherein the monitor is coupled with the coupling interface via an ohmic contact and an existing conductive facility of one of a bed, a wheelchair and a walker. 4. The system according to claim 1, wherein the monitor is coupled with the coupling interface via a wireless connection. 5. The system according to claim 1, wherein the monitor includes a mobile device. 6. The system according to claim 1, wherein the coupling interface includes a conductive plate or sheet in body-coupled communication with the patient and the at least one medical sensor; andthe monitor receives the streamed medical data via the conductive plate or sheet. 7. The system according to claim 6, further including: a RF fallback system including an alternative communication path which transmits the streamed medical data from the at least one medical sensor to the monitor when the body-coupled communication between the medical sensor and the conductive plate or sheet of the coupling interface breaks down. 8. A method for wirelessly monitoring a patient, the method comprising: streaming collected medical data which includes at least one of: electrocardiogram (ECG) data, photoplethysmograph (PPG) data, pulse oximetry (SpO2) data, and electroencephalography (EEG) data over a patient's body to a coupling interface via body-coupled communication using a streaming protocol including at least one of real time transport protocol (RTP) and user datagram protocol (UDP); anddisplaying the streaming medical data on a monitor which receives the streaming medical data from the coupling interface. 9. The method according to claim 8, wherein the data is further streamed to the monitor in real-time. 10. The method according to claim 8, further including: establishing a bidirectional body-coupled communication between the medical sensor and the coupling interface. 11. The method according to claim 10, further ... | Koninklijke Philips N.V. | A61B005/00
US20110015547 | A1 | 2011-01-20 | METHOD AND A DUAL-ARRAY TRANSDUCER PROBE FOR REAL TIME MECHANICAL IMAGING OF PROSTATE | 2010-09-20 | The present invention relates to a transrectal probe and method for real time mechanical imaging of a prostate. The probe is equipped with dual-array pressure sensors—one on the probe head and another on the shaft of the probe spaced away from the head with an angular and linear offset forming an S-shaped transition between the shaft and the head of the probe. The addition of the shaft pressure sensor array together with orientation tracking sensors allows precise calculation of the current head position throughout the examination of the prostate. Display means are used to guide the user in the proper manipulation of the probe in order to reduce the forces on surrounding tissues and organs and to minimize patient's discomfort. | 1. A transrectal probe for mechanical imaging of a prostate comprising: a probe head sized to fit inside a rectum and equipped with a head pressure sensor array, said probe head having a distal end and a proximal end, said probe head defining a longitudinal probe head axis, anda probe shaft defining a longitudinal probe shaft axis, said probe shaft having an S-shaped distal end attached to said proximal end of the probe head, said S-shaped distal end defining a linear offset between said longitudinal probe head axis and said longitudinal probe shaft axis,whereby said S-shaped distal end of said probe shaft configured to minimize disturbance of surrounding tissues and organs when said prostate is compressed by said probe head during examination procedure. 2. The transrectal probe as in claim 1, wherein said probe head including a straight segment containing said head pressure sensor array, said straight segment being about 25 mm (1 inch) to about 50 mm (2 inches) long. 3. The transrectal prove as in claim 1, wherein said linear offset ranging between about 4 mm ( 5/32 of an inch) and 15 mm ( 19/32 of an inch). 4. The transrectal probe as in claim 3, wherein said linear offset is about 10 mm (⅜ of an inch). 5. The transrectal probe as in claim 1, wherein said S-shaped distal end of the probe shaft shaped to define an angular offset between said longitudinal probe head axis and said longitudinal probe shaft axis. 6. The transrectal probe as in claim 5, wherein said angular offset ranging between about 4 degrees and 12 degrees. 7. The transrectal probe as in claim 1, wherein said probe shaft further including a shaft pressure sensor array, said head pressure array spaced away from said shaft pressure sensor array by about 40 mm (1 9/16 inches) to 80 mm (3⅛ inches) as measured between the centers of the respective pressure sensor arrays. 8. A dual-array transrectal probe for mechanical imaging of a prostate, said probe comprising: a probe head sized to fit inside a rectum and equipped with a two-dimensional head pressure sensor array, anda probe shaft attached to said probe head, said probe shaft including a shaft pressure sensor array. 9. The dual-array transrectal probe as in claim 8, wherein said shaft pressure sensor array comprising at least one linear pressure sensor array. 10. The dual-array transrectal probe as in claim 9, wherein said shaft pressure sensor array comprising at least two linear pressure sensor arrays. 11. The dual-array transrectal probe as in claim 10, wherein said shaft pressure array comprising three linear pressure sensor arrays. 12. The dual-array transrectal probe as in claim 11, wherein said three linear pressure sensor arrays are parallel to each other and to said probe shaft and spaced equally about the periphery thereof. 13. The dual-array transrectal probe as in claim 8, wherein said shaft pressure sensor array is a two-dimensional pressure sensor array. 14. A dual-array transrectal probe system for mech... | ARTANN LABORATORIES, INC. | A61B005/103
US20110015499 | A1 | 2011-01-20 | PATIENT BED WITH MONITORING AND THERAPY DEVICE | 2009-01-05 | The sick bed (1) can be rotated about an axis (9) in a direction indicated by an arrow (P), and consequently propped up. A patient strapped to the bed (4) will likewise be propped up. The verticalisation of the patient has effect on the physiological parameters of the same. According to the invention, the angle of inclination of the bed (4) relative to the horizontal is set according to the measured physiological parameters of the patient fixed on the bed (4). | 1. A patient bed comprising a frame, a bed connected to the frame and having a support surface for a patient, with at least one subsection of the bed arranged such that it can be moved relative to the frame, and at least one drive for setting the relative orientation of the at least one subsection of the bed relative to the frame, wherein the patient bed has at least one of a controller and a regulator for activating the at least one drive as a function of at least one measured physiological parameter of the patient. 2. The patient bed in accordance with claim 1, wherein the at least one drive is configured for setting the inclination of the support surface of the bed or a subsection of the support surface of the bed relative to the frame. 3. The patient bed in accordance with claim 2, wherein the inclination can be adjusted to an angle approx. −30° and +90° with respect to the horizontal position of the support surface. 4. The patient bed in accordance with claim 3, the inclination of the support surface can be adjusted such that the head area of the bed is inclined downwards. 5. The patient bed in accordance with claim 1, wherein the at least one drive is configured for changing the orientation of a part of the support surface of the bed for the purpose of mobilising an area of the patient's body lying on the support surface. 6. The patient bed in accordance with claim 1, wherein the patient bed comprises a measuring device for measuring the at least one physiological parameter of the patient, particularly a circulation parameter, such as heart rate, blood pressure, pulse, ECG, blood sugar, body temperature, intracranial pressure, etc. 7. The patient bed in accordance with claim 1, wherein the controller and regulator, respectively, has at least one of a control unit, a simple regulator, and an automatic regulator. 8. The patient bed in accordance with claim 1, wherein the controller/regulator is arranged at the patient bed. 9. The patient bed in accordance with claim 1, wherein the patient bed is be equipped such that the support surface can be cooled, warmed, put into a harder or softer setting, and/or can convey sensory stimuli. 10. A monitoring and therapy device, comprising: at least a patient bed comprising a frame, a bed connected to the frame and having a support surface for a patient, with at least one subsection of the bed arranged such that it can be moved relative to the frame, and at least one drive for setting the relative orientation of the at least one subsection of the bed relative to the frame; andat least one of a controller and a regulator for activating the at least one drive as a function of at least one measured physiological parameter of the patient. 11. The monitoring and therapy device in accordance with claim 10, wherein the bed is preferably connected to the frame such that the support surface can be inclined relative to the frame up to a predetermined angle. 12. The monitoring and therapy device in accord... | - | A61G007/015 A61G007/002 A61G007/005 A61B005/02 A61B005/01 A61G015/00
US07883545 | B2 | 2011-02-08 | Method and device for determining the change in an object | 2007-09-19 | A method is provided for determining a dimensional change in a first object, wherein in an initial state said first object is connected to a second object via a joint, and wherein a first reference system is rigidly connected to the second object. The method includes using a non-invasive coupling device to couple a second reference system to the first object; ascertaining and/or storing a first location corresponding to a location of the second reference system relative to the first reference system, and a second location corresponding to a location between at least one reference point on the first object relative to the first reference system; separating the first object from the second object at the joint; altering the first object; re-assembling the altered first object and the second object to form the joint, wherein the first location is used to align the altered first object to the second object; after alignment, ascertaining or storing a third location corresponding to a location of the at least one reference point relative to the first reference system; and determining the dimensional change in the first object due to the alteration based on a change in location between the second location and the third location of the at least one reference point. | 1. A method for determining a dimensional change in a first object, wherein in an initial state said first object is connected to a second object via a joint, and wherein a first reference system is rigidly connected to the second object, comprising: using a non-invasive coupling device to couple a second reference system to the first object;ascertaining and/or storing a first location corresponding to a location of the second reference system relative to the first reference system, and a second location corresponding to a location between at least one reference point on the first object relative to the first reference system, wherein ascertaining the second location includes using a pointer to ascertain the second location;separating the first object from the second object at the joint;altering the first object;re-assembling the altered first object and the second object to form the joint, wherein the first location is used to align the altered first object to the second object;after alignment, ascertaining or storing a third location corresponding to a location of the at least one reference point relative to the first reference system; anddetermining the dimensional change in the first object due to the alteration based on a change in location between the second location and the third location of the at least one reference point. 2. The method according to claim 1, wherein coupling the second reference system to the first object includes indirectly coupling the second reference system to the first object. 3. The method according to claim 1, wherein using the non-invasive coupling device includes using a belt or elastic strip as the coupling device. 4. The method according to claim 1, wherein coupling the second reference system to the first object includes coupling the second reference system on a mantle that surrounds the first object. 5. The method according to claim 1, wherein ascertaining or storing the first, second and third locations includes using a navigation system to ascertain said locations. 6. The method according to claim 1, wherein the joint is a ball joint. 7. The method according to claim 1, wherein altering the first object includes at least one of cutting or rasping the first object, or introducing or attaching an implant, joint implant or femoral head to the first object. 8. The method according to claim 1, wherein determining a dimensional change in the first object includes determining a change in length of the first object. 9. The method according to claim 1, wherein ascertaining or storing the third location includes ascertaining or storing a location of a reference point different from the at least one reference point, said location of the different reference point being relative to the first reference system, said different reference point having a known positional relationship to the at least one reference point, and deducing a location of the at least one reference point relative to the first reference system... | BrainLAB AG | A61F002/30 A61B005/103
US07873399 | B2 | 2011-01-18 | Monitoring of physiological analytes | 2007-01-30 | Methods and devices are provided for measuring the concentration of target chemical analytes present in a biological system. Device configuration and/or measurement techniques are employed in order to reduce the effect of interfering species on sensor sensitivity. One important application of the invention involves a method and device for monitoring blood glucose values. | 1. A method for measuring an amount or concentration of an analyte present in one or more samples in a biological system, the method comprising: pre-treating a sample site to enhance permeability of the sample site;providing one or more samples comprising said analyte from the sample site, wherein said samples are provided using iontophoresis;sensing each of said one or more samples with a sensor device;obtaining respective measurement signals from the analyte in each of said samples based on said sensing, said measurement signals specifically related to the analyte amount or concentration in the biological system;operating, with a microprocessor, said sensor device to provide said measurement signals;applying one or more processes, with said microprocessor, for selectively favoring analyte-specific signal components over other signal components due to interfering species, said processes selected from the group consisting of (a) a differential signal process which subtracts non-analyte signal components from the measurement signal, (b) a selective detection process which is performed during the sensing, (c) a purge performed before and after said obtaining step, wherein said purge substantially removes residual signal components from the sensor device, and (d) any combinations thereof; andrepeating, with said microprocessor, the operating and applying steps at least once over time. 2. The method of claim 1, wherein said pre-treating is performed with a laser. 3. The method of claim 1, wherein said pre-treating is performed with a plurality of micro-needles. 4. The monitoring system of claim 1, wherein the analyte comprises glucose. | Animas Corporation | A61B005/05 A61B005/00
US20110034785 | A1 | 2011-02-10 | Biopolymer resistant devices using coatings with surface energy gradient | 2010-06-30 | A medical device or analytical device comprising a fluid-impervious surface comprising a base surface, at least one distinct region of the base surface covered by a mixed monolayer film, the mixed monolayer film comprising a species having a functional group M1 and a species having a functional group M2 where M1 and M2 have different surface energies, the mixed monolayer forming a surface energy gradient wherein at least one of the species used to form the monolayer on the surface comprises a biopolymer-resistant domain. | 1) A medical device comprising one or more tubes, at least one tube comprising a fluid-impervious surface comprising a base surface, at least one distinct region of length L and width W of the base surface covered by a mixed monolayer film, the mixed monolayer film comprising a species having a functional group M1 and a species having a functional group M2 where M1 and M2 have different surface energies, the mixed monolayer forming a surface energy gradient from a proximal location to a distal location within the region wherein 1 equals the distance from the proximal location to the distal location, W is 20 nanometers or greater, and the ratio of L/W is greater than 2 and wherein any portions of the surface that border the at least one distinct region along the dimension L have substantially equal surface energies. 2) The device of claim 1, wherein the mixed monolayer is formed from species X1-J1-M1 and X2-J2-M2 wherein X1, X2, M1, and M2 represent separate functional groups where M1 and M2 have different surface energies and J1 and J2 represents spacer moieties, the species X1-J1-M1 and X2-J2-M2 forming a self-assembled monolayer onto the base surface from solution. 3) The device of claim 1 wherein the molar concentration of the species comprising the functional group M2 continuously increases relative to the concentration of the species comprising functional group M1 from the proximal location to the distal location of the region having the mixed monolayer. 4) The device of claim 2 wherein at least one of the species comprises a biopolymer-resistant domain. 5) The device of claim 1 wherein any portions of the surface that border the at least one distinct region along the dimension L are covered by substantially uniform monolayers. 6) The device of claim 5 wherein the substantially uniform monolayers comprise biopolymer-resistant domains. 7) The device of claim 1 where the ratio of L/W is greater than 10. 8) An analytical device comprising a fluid-impervious surface comprising a base surface, at least one distinct region of length L and width W of the base surface covered by a mixed monolayer film, the mixed monolayer film comprising a species having a functional group M1 and a species having a functional group M2 where M1 and M2 have different surface energies, the mixed monolayer forming a surface energy gradient from a proximal location to a distal location within the region wherein L equals the distance from the proximal location to the distal location, W is 20 nanometers or greater, and the ratio of L/W is greater than 2 and wherein any portions of the surface that border the at least one distinct region along the dimension L have substantially equal surface energies and wherein at least one of the regions comprising the surface energy gradient is defined in a series of channels and passages. 9) The device of claim 8, wherein the mixed monolayer is formed from species X1-J1-M1 and X2-J2-M2 wherein X1, X2, M1, and M2 represent separa... | - | A61B005/145
US20110046475 | A1 | 2011-02-24 | TECHNIQUES FOR CORRECTING TEMPERATURE MEASUREMENT IN MAGNETIC RESONANCE THERMOMETRY | 2009-08-24 | Techniques for correcting temperature measurement in MR thermometry are disclosed. In particular, phase shifts that arise from factors other than temperature changes are detected, facilitating correction of temperature measurements. | 1. A method of correcting proton resonance frequency (PRF) based magnetic resonance (MR) temperature measurement, the method comprising the steps of: detecting at least one first MR response of one or more micro-coils located in or near a region of interest, the detection being performed approximately when a first PRF image of the region of interest is acquired;detecting at least one second MR response of the one or more micro-coils approximately when a second PRF image of the region of interest is acquired;determining a temperature-invariant difference between the at least one second MR response and the at least one first MR response, the temperature-invariant difference being caused by factors unrelated to a temperature change in or near the region of interest; andcorrecting, based on the temperature-invariant difference, a temperature measurement of the region of interest made from the second PRF image and the first PRF image. 2. The method of claim 1, wherein (i) the one or more micro-coils consist of a single micro-coil, and (ii) the temperature-invariant difference determined from the at least one second MR response and the at least one first MR response of the single micro-coil provides a spatially uniform correction to the first PRF image. 3. The method of claim 1, wherein (i) the one or more micro-coils consist of four micro-coils arranged in a non-coplanar fashion, and (ii) the temperature-invariant difference determined from the at least one second MR response and the at least one first MR response of the four micro-coils provides a zero-th order and a first order corrections to the first PRF image. 4. The method of claim 1, wherein (i) the one or more micro-coils consist of five or more micro-coils arranged in a non-coplanar fashion, and (ii) the temperature-invariant difference determined from the at least one second MR response and the at least one first MR response of the five or more micro-coils provides a zero-th order, a first order, and at least a partial second order corrections to the first PRF image. 5. The method of claim 1, wherein the one or more micro-coils are filled with a substance that provides temperature-insensitive MR responses. 6. The method of claim 5, wherein the one or more micro-coils are filled with oil or other non-aqueous fluid. 7. The method of claim 1, wherein the one or more micro-coils are filled with water or water-based substance, and wherein temperature(s) at the locations of the micro-coil(s) are constant. 8. The method of claim 1, wherein the one or more micro-coils are filled with water or water-based substance and temperature(s) at the locations of the micro-coil(s) are known, and further comprising adjusting for a temperature-dependent component of the second MR response in determining the temperature-invariant difference between the at least one second MR response and the at least one first MR response. 9. The method of claim 1, further comprising: independently determining temperature... | - | A61B005/055
US20110040160 | A1 | 2011-02-17 | ANALYSIS DEVICE HAVING A DISCARDING MECHANISM | 2007-01-21 | An analyzing device which fits an analyzing instrument 2 thereinto for use and has a disposal mechanism 5 for disposing of the analyzing instrument 2 after completion of analysis, wherein the disposal mechanism 5 inputs a load in a direction different from a disposal direction D1 of the analyzing instrument 2 to dispose of the analyzing instrument 2. In the analyzing device 1 having a connector portion 4 having a terminal to be brought into contact with a terminal of the analyzing instrument 2, the disposal mechanism 4 is arranged below the connector portion 4. | 1. An analyzing device which fits an analyzing instrument thereinto for use and has a disposal mechanism for disposing of the analyzing instrument after completion of analysis, wherein the disposal mechanism inputs a load in a direction different from a disposal direction of the analyzing instrument to dispose of the analyzing instrument. 2. An analyzing device comprising: a connector portion which fits thereinto an analyzing instrument which can analyze a specimen by an electrochemical method and has a terminal brought into contact with a terminal of the analyzing instrument; anda disposal mechanism for disposing of the analyzing instrument after completion of analysis, wherein the disposal mechanism is arranged below the connector portion. 3. The analyzing device having the disposal mechanism according to claim 2, wherein the disposal mechanism inputs a load in a direction different from a disposal direction of the analyzing instrument to dispose of the analyzing instrument. 4. The analyzing device having the disposal mechanism according to claim 1 or 3, wherein the disposal mechanism has a moving body which has a contacting portion brought into close contact with the analyzing instrument and can be reciprocated between a standby position and a disposal position which can dispose of the analyzing instrument,wherein when a load is inputted by the contacting portion and the moving body is moved from the standby position to the disposal position, the analyzing instrument is moved together with the moving body. 5. The analyzing device having the disposal mechanism according to claim 4, wherein the analyzing instrument contacted with the contacting portion is moved together with the moving body due to frictional resistance between the analyzing instrument and the contacting portion. 6. The analyzing device having the disposal mechanism according to claim 5, wherein the contacting portion is formed of an elastic member. 7. The analyzing device having the disposal mechanism according to claim 4, wherein when the moving body is in the standby position, the contacting portion is not brought into contact with the analyzing instrument, and when the moving body is moved from the standby position to the disposal position, the contacting portion is brought into contact with the analyzing instrument. 8. The analyzing device having the disposal mechanism according to claim 7, wherein the contacting portion is shifted upward while the moving body is moved from the standby position to the disposal position. 9. The analyzing device having the disposal mechanism according to claim 8, further comprising an engaging portion which engages the moving body and can be moved relatively to the moving body, wherein one of the moving body and the engaging portion has a guide groove and the other has a convex portion which engages the guide groove,wherein when the moving body is reciprocated relatively to the housing, the convex portion is moved in the guide groove t... | Arkray, Inc. | A61B005/145
US08657830 | B2 | 2014-02-25 | Pelvic plane locator and patient positioner | 2010-09-09 | The invention includes both a method for acquiring the position of a patient's anterior pelvic plane, and a patient positioner or holder (apparatus) for securing a patient and acquiring the position of the patient's anterior pelvic plane. The method and apparatus facilitate definition of the anterior pelvic plane (APP) in the context of computer assisted hip surgery. | 1. A pelvic locating frame, adapted to adjustably mount in opposition to an adjustable backrest, and suitable for use in connection with computer assisted surgery for finding the orientation of a patient's pelvic plane, comprising: a support assembly, comprising: at least two anterior superior iliac spine (ASIS) locating features, adapted to press against and superficially engage the patient's ASIS;at least one pubic locating feature, adapted to engage the patient's pubic symphysis, said ASIS and pubic locating features defining an anterior pelvic plane (APP);a group of index features, said group constrained to maintain a predetermined rotational relationship to said anterior pelvic plane (APP) defined by said ASIS and pubic locating features; anda plurality of elongated members to which said at least two ASIS locating features, said at least one pubic locating feature, and said group of index features are mounted;a horizontal adjustment rod distinct from said support assembly to which said support assembly is mounted, said horizontal adjustment rod having an associated first axis along its length; anda vertical adjustment rod distinct from said support assembly to which said horizontal adjustment rod is mounted, said vertical adjustment rod being substantially perpendicular to said horizontal adjustment rod and having an associated second axis along its length, said elongated members arranged such that each can move independently of said vertical adjustment rod,said horizontal and vertical adjustment rods arranged such that said support assembly is capable of being rotated about said first and second axes,said pelvic locating frame arranged such that the position of said second axis does not move when said support assembly is rotated about said first axis, and does not move when said support assembly is rotated about said second axis. 2. The pelvic locating frame of claim 1, wherein said support assembly is adapted to allow engagement with two ASIS substantially vertically disposed, to accept a patient in a lateral decubitus position. 3. The pelvic locating frame of claim 1, wherein said group of index features comprises a non-collinear set of at least three features, said set defining a coordinate system in predetermined rotational relationship with said APP. 4. The pelvic locating frame of claim 3, wherein at least one of said index features comprises an indentation adapted to receive a complementary probe tip on a trackable probe. 5. The pelvic locating frame of claim 3, wherein said index features are constrained to move in a predetermined, coordinated relationship with said APP, so that the orientation of the APP can be calculated from known positions of the index features. 6. The pelvic locating frame of claim 5, wherein said index features are carried on a rigid plate; and wherein said rigid plate is attachable to said pelvic locating frame by an apparatus which constrains the plate to rotate in concert with said pelvic locating fram... | Kinamed, Inc. | A61B005/103 A61B005/117 A61B017/58 A61B017/60 A61B019/00
US20110009726 | A1 | 2011-01-13 | TAPE CASSETTE FOR A MEDICAL HANDHELD DEVICE AND BLOOD SUGAR MEASURING SYSTEM | 2010-07-10 | A tape cassette for a medical handheld device. The inventive tape cassette includes a carrier tape with a plurality of functional elements which may take the form of test fields for analyzing a sample of a human or animal body fluid and/or a plurality of lancets. During use, the carrier tape slides past friction points. To reduce friction between the carrier tape and at least one of the friction points, the at least one friction point includes a surface made of a fluoropolymer. The tape cassette may have a housing that includes a sealed compartment with a tape exit opening wherein the friction point is disposed proximate the tape exit opening. A cover element may also be positioned to cover the tape exit opening. The cover element may also include a fluoropolymer material engageable with the carrier tape. A method of manufacturing such tape cassettes is also disclosed. | 1. A tape cassette for a medical handheld device, the tape cassette comprising: a carrier tape having disposed thereon a plurality of functional elements, wherein the functional elements include at least one of either a plurality of test fields for analyzing a sample of a human or animal body fluid or a plurality of lancets, whereby, during use, the carrier tape slidingly engages friction points disposed on the tape cassette, and whereinat least one of the friction points includes a surface made of a fluoropolymer to thereby reduce friction between the carrier tape and the at least one friction point. 2. The tape cassette according to claim 1 wherein the fluoropolymer contains a material selected from the group consisting of PTFE and PVDF. 3. The tape cassette according to claim 1 wherein the fluoropolymer is a coating. 4. The tape cassette according to claim 3 wherein the fluoropolymer is applied as varnish. 5. The tape cassette according to claim 3 wherein the fluoropolymer is applied as a foil. 6. The tape cassette according to claim 1 wherein the at least one friction point having a surface made of a fluoropolymer is formed in a two-component injection molding process. 7. The tape cassette according to claim 1 wherein the tape cassette includes cassette housing defining a tape exit opening and wherein the at least one friction point having a surface made of a fluoropolymer is positioned proximate the tape exit opening. 8. The tape cassette according to claim 7 wherein the at least one friction point having a surface made of a fluoropolymer is formed by a sealing lip positioned proximate the tape exit opening. 9. The tape cassette according to claim 7 wherein the at least one friction point having a surface made of a fluoropolymer is formed by a cover element that covers the tape exit opening. 10. A blood sugar measuring system comprising: a handheld device containing a carrier tape;the carrier tape having disposed thereon a plurality of functional elements wherein the plurality of functional elements include at least one of either a plurality of test fields for analyzing a sample of a body fluid or a plurality of lancets, anda transport mechanism that is operable to move the plurality of functional elements successively into a usage position, wherein the carrier tape slidingly engages friction points when moved by the transport mechanism; andwherein at least one of the friction points includes a surface made of a fluoropolymer to thereby reduce friction between the carrier tape and the at least one friction point. 11. The tape cassette according to claim 10 wherein the fluoropolymer contains a material selected from the group consisting of PTFE and PVDF. 12. The tape cassette according to claim 10 wherein the at least one friction point having a surface made of a fluoropolymer is formed in a two-component injection molding process. 13. The tape cassette according to claim 12 wherein the at least one friction point having a surface... | - | A61B005/00 B23P011/00 B29C045/16
US08548547 | B2 | 2013-10-01 | Aviation physiological health monitoring system and method | 2010-06-29 | An aircraft cockpit with health monitoring system and method enables an occupant (e.g. a pilot) and/or a ground station to monitor the physical condition of the occupant during a flight. A sensor such as a pulse oximetry sensor is in predetermined location to an occupant seat (e.g. an aircraft cockpit seat), within the field of reach of a person seated in the seat. The sensor is supported in the predetermined location in a stable manner that provides sufficient reaction to insertion of a finger to enable insertion of a finger in the sensor while the person is in a seated position, and the sensor is in communication with a control box that processes data from the sensor, and presents the data, e.g. at a the multi functional display in an aircraft cockpit. During a flight, a pilot seated in the aircraft cockpit seat will be prompted to insert a finger into the pulse oximetry sensor to produce output at the multi functional display. | 1. A sensing and reporting system for an aircraft, comprising a display, and a sensor configured to provide data related to at least one human vital condition, within the reach of an occupant of the cockpit, the sensor in communication with a processor for data from the sensor, the processor providing the display with output related to the human vital condition based on the data from the sensor, and means for prompting the cockpit occupant to engage the sensor, wherein the means for prompting the cockpit occupant to engage the sensor is configured to prompt the cockpit occupant based on the sensed altitude of the aircraft and time at that sensed altitude. 2. The system of claim 1, wherein the said display comprises a multifunctional display and the sensor and the said multifunctional display communicate via a control box system. 3. The system of claim 2, wherein the said control box system communicates with the said sensor wirelessly. 4. The system of claim 1, wherein the said display comprises a multifunctional display that is part of an aircraft instrument panel, and wherein the sensor is fixed to one of the instrument aircraft panel and seat. 5. The system of claim 4, wherein the said sensor comprises a pulse oximeter that is supported in a fixed, stabilized position in an aircraft panel, within reach of an occupant of the cockpit. 6. The system of claim 1, wherein the sensor data is transmitted and/or recorded on the aircraft or at a ground station. 7. The system of claim 6, wherein the said data is recorded in at least one of aircraft black box or the said control box system. 8. The system of claim 1, wherein the said control box system communicates with a portable device, and data from the control device is recorded in the portable device. 9. A method comprising a. providing a system for an aircraft, comprising a display, and a sensor within the reach of a cockpit occupant, the sensor configured be engaged by a cockpit occupant seated in the cockpit to provide data related to at least one human vital condition, the sensor in communication with a processor for data from the sensor, and the processor providing the display with output related to the human vital condition based on the data from the sensorb. prompting a cockpit occupant seated in the cockpit to engage the sensor to produce data as to the human vital condition, that is communicated to the processor, and produces output at the display related to the human vital condition, further including prompting the cockpit occupant by sending a reminder to the cockpit occupant that is designed to remind the cockpit occupant to engage the sensor to produce the data related to the at least one human condition. 10. The method of claim 9, including providing the reminder to the cockpit occupant based on the sensed altitude of the aircraft and time at that sensed altitude. 11. The method of claim 10, further including providing periodic reminders to the cockpit occupant when the sensed a... | - | A61B006/00
US08226558 | B2 | 2012-07-24 | Analyte monitoring device and methods of use | 2010-09-27 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. A glucose monitoring system comprising: a subcutaneous in vivo sensor for monitoring a glucose level of an interstitial fluid of a user;a sensor control unit coupled to the sensor to output one or more signals associated with the monitored glucose level of the interstitial fluid of the user; anda receiver/display unit in communication with the sensor control unit and comprising: a receiver to receive the outputted one or more signals associated with the monitored glucose level of the interstitial fluid of the user from the sensor control unit;an input device to receive an input from the user, wherein the input comprises, a medication sensitivity, drug delivery information, a food intake, and an exercise level;a processor to evaluate the one or more signals received by the receiver to produce a first set of results comprising: (i) the glucose level in interstitial fluid of the user,(ii) a rate of change of the glucose level, and(iii) an acceleration in the rate of change of the glucose level,wherein the processor further evaluates the input from the user and the first set of results to determine a dose of a medication; anda display in communication with the processor. 2. The system of claim 1, wherein the medication comprises insulin. 3. The system of claim 1, wherein the display displays a warning indication to the user if one or more of the glucose level, the rate of change of the glucose level and the acceleration in the rate of change of the glucose level meets or exceeds a threshold level. 4. The system of claim 3, wherein the warning indication comprises one of more of: an indication of hypoglycemia; an indication of impending hypoglycemia; an indication of hyperglycemia; an indication of impending hyperglycemia; an indication of a malfunction; and an indication of an error condition. 5. The system of claim 1, wherein the processor tags the one or more signals with an event marker. 6. The system of claim 5, wherein the event marker comprises one or more of a hypoglycemic episode, a hyperglycemic episode, the food intake, the drug delivery information, and the exercise level. 7. The system of claim 1, wherein the system comprises an integrated unit including the sensor, the sensor control unit and the receiver/display unit. 8. The system of claim 1, wherein the sensor control unit and the receiver/display unit are separate units. 9. The system of claim 8, wherein the sensor control unit comprises a transmitter to transmit the one or more signals from the sensor control unit to the receiver/display unit. 10. The system of claim 1, wherein the receiver/display unit further comprises a data storage unit to store at least one of the one or more signals and the result. 11. The system of claim 1, wherein the receiver/display unit further comprises a port to connect the receiver/display unit to a computer. 12. The system of claim 11, wherein the receiver/display unit transmits at least one of the one or more signals and the result from... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/05
US07894878 | B2 | 2011-02-22 | Anatomically-referenced fiducial marker for registration of data | 2005-12-30 | An anatomical data registration device may include a dental device. The dental device may configurable to be subject-specific. A dental device may be positionable in a substantially fixed spatial relationship relative to a portion of the subject. An anatomical data registration device may include one or more radiopaque markers. At least one of the radiopaque markers may be positioned in the dental device. The radiopaque marker may be configured to facilitate alignment of two or more images of the portion of the subject. The data registration device may be positionable within a subject's mouth. | 1. An anatomical data registration device, comprising: a dental device configurable to be subject specific such that upon configuration, the dental device is positionable in a substantially fixed spatial relationship relative to a portion of the subject; andone or more radiopaque markers positioned in the dental device, wherein at least one of the radiopaque markers is configured to facilitate alignment of two or more images of the portion of the subject;wherein at least a portion of the dental device is positionable within a subject's mouth, and wherein one or more of the radiopaque markers are positioned in the portion of the dental device that is positioned in the subject's mouth during use. 2. The device of claim 1, wherein the device is configured to allow the subject freedom of movement when positioned in the subject's mouth. 3. The device of claim 1, further comprising a grasping device coupled to the dental device. 4. The device of claim 1, wherein the dental device comprises a bite plate. 5. The device of claim 1, wherein the dental device is configured to be reproducibly positioned in the subject's mouth. 6. A method for aligning two or more images of a subject, comprising: obtaining a dental device wherein the dental device comprises impressions of at least a portion of one or more teeth of a subject;positioning at least a portion of the dental device in the subject's mouth using the impressions formed in the dental device to position the dental device in a substantially fixed spatial relationship relative to at least a portion of the subject, wherein the dental device comprises one or more radiopaque markers, and wherein one or more of the radiopaque markers are positioned in the portion of the dental device that is positioned in the subject's mouth;obtaining two or more images of at least the portion of the subject and at least a portion of at least one of the radiopaque markers; andusing at least a portion of an image of at least one of the radiopaque markers as a reference to align at least a part of two or more of the images. 7. The method of claim 6, using at least a portion of at least one of the radiopaque markers as a reference to align at least a part of two or more of the images after the images have been obtained. 8. The method of claim 6, allowing the subject at least a minimum of freedom of movement of at least the portion of the subject relative to an imaging system and relative to a position of the portion of the subject during a plurality of images. 9. The method of claim 6, wherein at least the portion of the subject is in a first position relative to an imaging system when a first image is obtained, wherein at least the portion of the subject is in a second position relative to an imaging system when a second image is obtained, and wherein the first position is different from the second position. 10. The method of claim 6, further comprising using at least a portion of at least one of the radiopaque markers ... | Board of Regents, The University of Texas System | A61B005/05
US07885705 | B2 | 2011-02-08 | System and method for facilitating hip surgery | 2006-02-10 | A system and method for use during hip surgery includes a position frame that is removably secured to the patient's femur. The system further includes a tracking unit and navigation engine that establish a pelvic coordinate system and, with the patient's leg placed in a predetermined orientation, capture the location of the femoral position frame within the pelvic coordinate system. The system further includes an image generator configured to display one or more images for directing the surgeon to return the patient's leg to the predetermined orientation following a trial reduction based upon the previously captured position. A leg change determination engine compares the current, post-reconstruction position of the femoral position tracking frame with the previously captured, pre-reconstruction position to derive leg length, offset, and anterior-posterior change values. The change values are then provided to the surgeon for evaluating the suitability of the trial reduction. | 1. A system for performing hip surgery on a patient having a leg, a pelvis and a femur, the system comprising: a tracking unit configured to detect a femoral position tracking frame removably secured to the patient's femur;a navigation engine coupled to the tracking unit, the navigation unit configured to establish a pelvic coordinate system at the patient's pelvis and to derive a first, pre-reconstruction location of the femoral position tracking frame relative to the pelvic coordinate system and a pre-reconstruction orientation of the femoral position tracking frame relative to the pelvic coordinate system while the leg is in a predetermined orientation;an image generator coupled to the tracking unit and the navigation engine; anda display screen coupled to the image generator, wherein the navigation engine is configured to utilize the previously derived pre-reconstruction orientation of the femoral position tracking frame to compute a plurality of orientation change values following a trial reduction as the patient's leg is moved, and to provide the computed orientation change values to the image generator,the image generator is configured to display on the display screen one or more images including the computed orientation change values to assist in returning the patient's leg to the predetermined orientation following a trial reduction,the navigation engine is further configured, with the patient's leg returned to the predetermined orientation, to derive a second, post-reconstruction location of the femoral position tracking frame, and to compute one or more of a leg length change value, an offset change value and an anterior-posterior (AP) change value, based on the first, pre-reconstruction location,the image generator displays the one or more computed leg length change value, the offset change value, and the AP change value, andthe system is free from establishing a coordinate system on the patient's femur. 2. The system of claim 1 wherein the plurality of computed orientation change values displayed by the image generator include an indication of flexion, an indication of internal rotation and an indication of abduction of a current orientation of the patient's leg relative to the predetermined orientation. 3. The system of claim 2 wherein the indication of flexion, the indication of internal rotation and the indication of abduction are displayed in at least near real-time. 4. The system of claim 1 wherein the image generator displays each of the computed leg length change value, the offset change value, and the AP change value. 5. A system for performing hip surgery on a patient having a leg, a pelvis and a femur, the system comprising: a femoral position tracking frame configured for attachment to the patient's femur;a pelvic position tracking frame configured for attachment to the patient's pelvis;means for calculating a pre-reconstruction orientation and position of the femoral position tracking frame relative to the pelvic posi... | - | A61B005/05
US07890159 | B2 | 2011-02-15 | Cardiac activation sequence monitoring and tracking | 2004-09-30 | Cardiac monitoring and/or stimulation methods and systems provide monitoring, diagnosis, and defibrillation and/or pacing therapies. A signal processor receives a plurality of composite signals associated with a plurality of sources, performs a source separation, and produces one or more cardiac signal vectors associated with all or a portion of one or more cardiac activation sequences based on the source separation. A method of signal separation involves detecting a change in a characteristic of the cardiac signal vector relative to a baseline. One or more vectors and/or activation sequences may be selected, and information associated with the vectors and/or activation sequences may be stored and tracked. | 1. A cardiac method, comprising: sensing a plurality of composite cardiac signals over time at a plurality of patient-internal locations;performing one or more source separations using one or more of the sensed plurality of composite signals;producing one or more cardiac signal vectors indicative of a cardiac depolarization sequence of an area of the heart based on the one or more source separations;establishing a baseline of a characteristic of the one or more cardiac signal vectors;performing at least one subsequent source separation using at least one of the sensed plurality of composite signals;monitoring the characteristic of the one or more cardiac signal vectors based on the at least one subsequent source separation;detecting a change in the characteristic of the one or more cardiac signal vectors relative to the baseline of the characteristic of the one or more cardiac signal vectors; andidentifying a cardiac condition based on the detected change, wherein performing the one or more source separations, producing the one or more cardiac signal vectors, establishing the baseline, performing the at least one subsequent source separation, monitoring, detecting, and identifying are performed at least in part by circuitry of a medical system. 2. The method of claim 1, wherein the baseline is established based on the characteristic of the one or more cardiac signal vectors over multiple cardiac cycles. 3. The method of claim 1, wherein the change in the characteristic comprises one or more of a magnitude change of the one or more cardiac signal vectors, a variance change of the one or more cardiac signal vectors, a power spectral density change of the angle of the one or more cardiac signal vectors, a power spectral density change of the magnitude of the one or more cardiac signal vectors, a trajectory change of the one or more cardiac signal vectors, an activation velocity change, and a temporal profile change of the one or more cardiac signal vectors. 4. The method of claim 1, wherein the change in the characteristic comprises one or more of a rate of change of angle of the one or more cardiac signal vectors, a rate of change of magnitude of the one or more cardiac signal vectors, a rate of change of variance of the one or more cardiac signal vectors, a rate of change of velocity of the one or more cardiac signal vectors and a rate of change of temporal profile of the one or more cardiac signal vectors. 5. The method of claim 1, wherein the change in the characteristic comprises one or more of a trend of an angle of the one or more cardiac signal vectors, a trend of a magnitude of the one or more cardiac signal vectors, a trend of a variance of the one or more cardiac signal vectors, a trend of a velocity of the one or more cardiac signal vectors and a trend of a temporal profile of the one or more cardiac signal vectors. 6. The method of claim 1, wherein the change is detected using a window configured to enhance a morphological feature ... | Cardiac Pacemakers, Inc. | A61B005/0428 A61B005/0452
US08241213 | B2 | 2012-08-14 | Microprocessor system for the analysis of physiologic datasets | 2010-07-19 | A system and method for organization and analysis of complex and dynamically interactive time series is disclosed. One example comprises a processor based system for relational analysis of physiologic signals for providing early recognition of catastrophic and pathologic events such as pathophysiologic divergence. The processor is programmed to identify pathophysiologic divergence of at least one of first and second physiologic parameters in relationship to the other and to output an indication of the divergence. An object-based method of iterative relational processing waveform fragments in the time domain is described wherein each more complex waveform object inherits the characteristics of the waveform objects from which it is derived. The first physiologic parameter can be the amplitude and frequency of the variation in chest wall impedance or nasal pressure and the second parameter can be a measure or indication of the arterial oxygen saturation. | 1. A microprocessor system for evaluation of a patient, the system comprising: monitor having a first and a second sensor, wherein the monitor is positioned for positioning adjacent said patient, the first sensor obtaining first data associated a first physiologic parameter and the second sensor obtaining second data associated with a second physiologic parameter, wherein the second physiologic parameter comprises a measure of arterial oxygenation;a processor in communication with the monitor, the processor programmed to: receive the first data associated with the first physiologic parameter from the monitor;receive the second data associated with the second physiologic parameter from the monitor;produce a first timed waveform based on the first physiologic parameter of a patient;produce a second timed waveform based on the second physiologic parameter which is generally subordinate to said first physiologic parameter, so that said second physiologic parameter normally changes in response to changes in said first physiologic parameters;identify pathophysiologic divergence of at least one of said first and second physiologic parameters in relationship to the other of said physiologic parameters, the pathophysiologic divergence comprising a change that is both unexpected and pathological;automatically determine a possible dysfunction associated with the change; andoutput an indication of the dysfunction. 2. The system, as in claim 1, wherein the at least one processor is further programmed to: a) calculate an index of said pathophysiologic divergence and, b) provide an indication based on said index. 3. The system of claim 1, wherein said first physiologic parameter comprises at least one of an indication and measure of a magnitude of timed ventilation of the patient. 4. The system of claim 3, wherein said first physiologic parameters includes at least one of an amplitude and a frequency of variation in chest wall impedance. 5. The system of claim 3, wherein said first physiologic parameter includes at least one of an amplitude and a frequency of variation in nasal pressure. 6. The system of claim 3, wherein the first physiologic parameter is a measure of timed volume of at least one gaseous component of ventilation. 7. The system of claim 1, wherein said first physiologic parameter is at least one of a measure and an indication of carbon dioxide. 8. A method of monitoring a patient comprising: a processor monitoring a patient with a first sensor to produce a first timed waveform of a first physiologic parameter and a second sensor to produce a second timed waveform of a second physiologic parameter, wherein the second physiologic parameter comprises a measure of arterial oxygenation;the processor transforming the sensed physiological inputs into visual data, wherein the visual data representing a first and second waveform, and wherein the second physiologic parameter being physiologically subordinate to said first physiologic parameters;the... |  | A61B005/00 A61B005/08
US07894888 | B2 | 2011-02-22 | Device and method for measuring three-lead ECG in a wristwatch | 2008-09-24 | A wristwatch worn by a user for measuring a three-lead ECG includes three electrodes placed separately on the front, either side, and back or strap thereof. The wristwatch further includes an electrode panel having the electrode on the front or either side of the watch, sensing elements, pressure, infrared or impedance detectors, and circuits. The electrode panel is capable of sensing the contact or press of fingers to trigger the ECG measuring. While the electrode in the back-side of the watch contacts the hand wearing the watch, the electrode and electrode panel on the front or either side of the watch are pressed by fingers from the other hand, and the electrode in the strap contacts the abdomen or left leg simultaneously. Thus, a three-lead ECG can be measured. ECG data can be transmitted to a personal or hospital computer by wireless networks or flash memory. | 1. A wristwatch for measuring a three-lead ECG, comprising: a case including a front side and a back side opposite to the front side;a strap attached to both ends of the case and including an inside surface and an outside surface, with the inside surface being intermediate the back side and the outside surface;a circuit board disposed in the case;a first electrode positioned on the outside surface of the strap;a second electrode positioned on the front side of the case;a third electrode positioned on the back side of the case, with the first, second and third electrodes being electrically separate for detecting voltage differences generated by cardiac systems, with the circuit board detecting voltage differences generated by the first and second electrodes, by the first and third electrodes, and by the second and third electrodes to obtain a set of ECG data comprising ECG lead I, ECG lead II, and ECG lead III;a display disposed on the case;a control panel comprising one of the electrodes and a plurality of sensing elements; andat least one control button for inputting information, changing function, and modifying user's mode of the wristwatch. 2. The wristwatch of claim 1, wherein the electrode and the sensing elements of the control panel are adapted to be touched separately. 3. The wristwatch of claim 1, wherein the control panel further includes an integrated unit covering both the electrode and the sensing elements of the control panel to allow the electrode and the sensing elements of the control panel to be activated simultaneously by one touch; wherein the integrated unit is an interface activating both the electrode and the sensing elements of the control panel by one touch. 4. The wristwatch of claim 1, further comprising an alerting device disposed in the case alerting a user to practice the ECG measurement. 5. The wristwatch of claim 4, wherein the alerting device emits light, makes sound, or vibrates. 6. The wristwatch of claim 4, wherein the alerting device is an LED or OLED emitting light. 7. The wristwatch of claim 4, wherein the alerting device is an OLED or LCD flashing light. 8. The wristwatch of claim 4, wherein the alerting device is a buzzer or speaker. 9. The wristwatch of claim 4, wherein the alerting device is a miniature vibrating motor for vibrating. 10. The wristwatch of claim 4, further comprising a sensing element measuring temperature changes as a reference value to activate the alerting device. 11. The wristwatch of claim 4, further comprising a sensing element measuring a user's activity levels as a reference value to activate the alerting device. 12. The wristwatch of claim 1, further comprising a transmission device for data transmission after the set of ECG data is obtained; with the transmission device comprising at least one of a flash memory and a wireless network for transmitting data to a remote personal or hospital computer. 13. The wristwatch of claim 1, further comprising a wireless module re... | Chang Gung University | A61B005/04
US20110015543 | A1 | 2011-01-20 | FLUID SAMPLE COLLECTION SYSTEM | 2010-07-19 | A sample collection container adapted to be connected to, and disconnected from, a sample collection adapter and a sample collection container cap. The sample collection adapter has a downward sloping interior surface that provides a free-flowing path for samples provided by a subject to the sample collection container. The sample collection adapter may have a three-dimensional rim surface that is curved in two different non-parallel planes. | 1. A sample collection system comprising: a sample collection container;a sample collection adapter for receiving a sample, wherein said sample collection adapter comprises a downward sloping interior surface and is adapted to be connected to, and disconnected from, said sample collection container; anda sample collection container cap, wherein said sample collection container cap is adapted to be connected to and disconnected from said sample collection container. 2. The sample collection system of claim 1 wherein said sample collection adapter is a conical funnel. 3. The sample collection system of claim 1 wherein said sample collection adapter is an ergonomic funnel. 4. The sample collection system of claim 1 wherein said sample collection adapter comprises a rim having a curvature in a first plane and a concave curvature in a second plane. 5. The sample collection system of claim 1 wherein said sample collection adapter, comprises a female thread. 6. The sample collection system of claim 1 wherein said sample container comprises a male thread. 7. The sample collection system of claim 1 wherein said sample collection container cap comprises a female thread. 8. The sample collection system of claim 1 wherein said sample collection adapter comprises an aperture that has a diameter equivalent to an inner diameter of an opening in said sample collection container. 9. A sample collection system comprising: a sample collection container;a sample collection adapter for receiving a sample, wherein said sample collection adapter comprises a conical interior surface and is adapted to be connected to, and disconnected from, said sample collection container; anda sample collection container cap, wherein said sample collection container cap is adapted to be connected to and disconnected from said sample collection container. 10. The sample collection system of claim 9 wherein said sample collection adapter comprises a female thread. 11. The sample collection system of claim 9 wherein said sample container comprises a male thread. 12. The sample collection system of claim 9 wherein said sample collection container cap comprises a female thread. 13. The sample collection system of claim 9 wherein said sample collection adapter comprises an aperture that has a diameter equivalent to an inner diameter of an opening in said sample collection container. 14. A sample collection system comprising: a sample collection container;a sample collection adapter for receiving a sample, wherein said sample collection adapter comprises an ergonomic funnel and is adapted to be connected to, and disconnected from, said sample collection container; anda sample collection container cap, wherein said sample collection container cap is adapted to be connected to and disconnected from said sample collection container. 15. The sample collection system of claim 14 wherein said sample collection adapter comprises a female thread. 16. The sample collection system of cla... | - | A61B005/00
US20110046515 | A1 | 2011-02-24 | METHODS AND APPARATUS FOR EXPRESSING BODY FLUID FROM AN INCISION | 2010-11-03 | A sample of a body fluid such as blood or interstitial fluid is obtained from a body by lancing a portion of a user's skin, preferably in an area other than a finger tip, to form an incision. After the needle has been removed from the incision, a force is applied to depress the skin in a manner forming a ring of depressed body tissue in surrounding relationship to the incision, causing the incision to bulge and the sides of the incision to open, whereby body fluid is forced out through the opening of the incision. A stimulator member is mounted to an end of a lancet-carrying housing for applying the force. The stimulator member can be movable relative to the housing, and can be either heated or vibrated to promote movement of the body fluid. | 1. A method of obtaining a sample of body fluid from a body, comprising the acts of: A) applying a skin-lancing device against the skin of a user to form an incision in the skin;B) removing the skin-lancing device from the incision; and thereafterC) applying a force to depress the skin in a manner forming a ring of depressed body tissue in surrounding relationship to the incision wherein a stimulating member and a plurality of levers act to bulge and stretch the incision open, whereby body fluid is expressed from the opening of the incision. 2. The method according to claim 1, wherein forces between the stimulating member and the skin cause the stimulating member to retain the skin in a stretched position. 3. The method according to claim 1 wherein act C includes applying the force in a direction inclined generally toward the bulged incision. 4. The method according to claim 1 wherein act C comprises applying the force progressively closer to the incision. 5. The method according to claim 1 wherein act C includes applying heat in a region of the incision. 6. The method according to claim 1 wherein act C includes applying ultrasonic frequency to a region of the incision. 7. The method according to claim 1 wherein act A comprises lancing a region of the body of the user other than a finger tip. 8. The method according to claim 1 wherein act A comprises applying a lancet against the skin. 9. A method, comprising the acts of: forming an incision in skin with an incision forming mechanism of a sampling device; andsqueezing the skin around the incision with a stimulation mechanism of the sampling device after said forming the incision to express body fluid from the incision. 10. The method of claim 9, wherein: the stimulation mechanism includes a plurality of levers pivotally mounted on the sampling device, each of the levers has an end for squeezing the skin; andsaid squeezing includes pivoting the levers so that the ends of the levers move in a radially inward direction. 11. The method of claim 10, wherein: the stimulation mechanism includes an inner member slidably disposed in the sampling device, the inner member has a cam face that engages the levers; andsaid squeezing includes pressing the sampling device against the skin so that cam face of the inner member pivots the levers. 12. A method of obtaining a sample of capillary whole blood from a target tissue, comprising the acts of: providing a penetrating system that includes a tissue stabilizing member;applying skin stimulation to a skin surface site with the tissue stabilizing member;introducing a penetrating member through the skin surface site to form an incision; andcollecting blood from the incision in the penetrating system. 13. The method of claim 12, wherein the skin stimulation is a vibratory motion applied to the skin surface. 14. The method of claim 13, wherein the skin stimulation is an ultrasonic vibratory motion applied to the skin surface. 15. The method of claim 12, ... | - | A61B005/151
US07890160 | B2 | 2011-02-15 | Heart monitoring system | 2007-08-30 | Heart monitoring system includes implantable medical device and service center. Implantable medical device includes stimulation pulse generator, ventricular sensing stage, activity sensor, impedance determination unit with a constant current or voltage source to generate sub-threshold measuring current pulses having constant current strength or constant voltage, measuring unit for measuring a voltage corresponding to a current fed through a body, impedance value determination unit connected with measuring unit adapted to determine impedance value for each measuring current pulse, and perform intrathoracic impedance measurement, a control unit adapted to collect data representing values of changes, and initiate data transmission, implant transceiver unit communicating with the service center with a data evaluation module including data trending of stored data with a user interface and said data evaluation module adapted to allow a physician to set for each data trend trigger criteria for decompensation detection and generate a decompensation indicator signal. | 1. A heart monitoring system comprising: an implantable medical device;a service center;a central data base;a user interface;said implantable medical device comprising a stimulation pulse generator configured to generate electric stimulation pulses and configured to connect to at least a ventricular stimulation electrode configured to deliver said electric stimulation pulses to at least a ventricle of a heart;a ventricular sensing stage configured to connect to an electrode configured to pick up electric potentials inside said at least said ventricle of said heart, said sensing stage configured to sense an excitation or a contraction of ventricular myocardium;an activity sensor;an impedance determination unit comprising a constant current source or a constant voltage source having current feed terminals that are configured to connect to electrodes for intracorporeal placement wherein said constant current source or said constant voltage source is configured to generate sub-threshold measuring current pulses having a constant current strength or a constant voltage, respectively, that cause a current to be fed through a body via intracorporeally placed electrodes;a measuring unit configured to measure a voltage corresponding to said current fed through said body by said constant current source or a current strength of said current fed through said body by said constant voltage source, respectively;an impedance value determination unit that is connected to said constant current source or said constant voltage source wherein said measuring unit is configured to determine an impedance value for each of said sub-threshold measuring current pulses;said impedance determination unit configured to perform an intrathoracic impedance measurement via a first and a second electrode combination wherein said second electrode combination is different from said first electrode combination;a battery;a control unit that is connected to said battery, said stimulation pulse generator, said ventricular sensing stage, said activity sensor and said impedance determination unit and that is configured to collect data representing values of: DC thoracic impedance from said first electrode combination,DC thoracic impedance from said second electrode combination,cardiac output,heart rate at a resting state of a patient,heart rate during exercise of said patient,activity level of said patient,ventricular extra systoles or VES per time unit,breathing rate,tidal volume;an implant transceiver unit configured to communicate with said service center;said service center comprising another data communication interface configured to allow at least an indirect data communication with said implantable medical device wherein said another data communication interface is connected to said central data base and to said user interface;wherein said control unit is configured to initiate transmission of said collected data to said service center;said service center further comprising a data ev... | Biotronik CRM Patent AG | A61B005/04
US08489371 | B2 | 2013-07-16 | Method and device for determining transfer functions of the HRTF type | 2009-02-17 | The invention relates to a method for determining transfer functions of the HRTF type for an individual, that includes: measuring, for a first number of directions, the transfer functions of the HRTF type specific to the individual; matching the directivity functions associated with the measured functions of the HRTF type, with reference directivity functions associated with reference transfer functions of the HRTF type, the reference functions of the HRTF type being determined for a second number of directions higher that the first number of directions and reconstructing the measured directivity functions from the reference directivity functions. | 1. A computer-implemented method for determining head related transfer functions (HRTFs) for an individual comprising: measuring, for a first number of directions, HRTFs specific to said individual;extracting measured directivity functions from said measured HRTFs for said individual;evaluating, using a computer, spatial similarities between the measured directivity functions and representative directivity functions of groups of reference directivity functions;associating a measured directivity function with a group of reference directivity functions, wherein the measured directivity function is associated with the group from which originates the representative function with which the evaluation of the spatial similarity is maximal;determining, for the measured directivity function, a reconstruction directivity function by interpolating from a reference directivity function of the associated group;determining reconstruction directivity functions for additional directions by interpolating from reference directivity functions of the associated group; andobtaining HRTFs for said individual for a second number of directions greater than the first number of directions from the reconstructed directivity functions. 2. The method according to claim 1, wherein the method further comprises a preliminary phase comprising: determining reference HRTFs for a plurality of individuals, according to a plurality of frequencies and said second number of directions;evaluating spatial similarities between directivity functions associated with said reference HRTFs;classifying said directivity functions into groups according to their spatial similarities;selecting a representative directivity function for each group; andmodifying the directivity functions in order to minimize a spatial shift with respect to their respective representative directivity functions and to form the reference directivity functions. 3. The method according to claim 2, wherein said evaluation of similarities between the directivity functions is based on a similarity criterion representative of independent similarities with respect to rotational shifts of said directivity functions. 4. The method according to claim 2, wherein said associating comprises: evaluating spatial similarities between the measured directivity functions and the representative directivity functions of the groups of reference directivity functions;associating the measured directivity functions with the groups of reference directivity functions according to said evaluation of similarities; andmodifying the measured directivity functions in order to minimize a spatial shift with respect to the representative directivity functions of the associated groups. 5. The method according to claim 4, further comprising modifying the measured directivity functions, after said determining reconstruction directivity functions for additional directions, in order to at least partially compensate for the minimization of the spatial shift... | France Telecom | G06G007/48 G06F007/60 G06F017/10 H04R005/00
US09049988 | B2 | 2015-06-09 | Automated vessel repair system, devices and methods | 2010-10-26 | Systems, devices and automated methods for minimally invasive surgery. A device is fabricated of bio-compatible semiconductor elements, and can be assembled, delivered, navigated and implanted by automated methods, using Nuclear Magnetic Resonance (NMR) technology. | 1. A method of implanting a device in a vessel or duct of a patient, said method comprising the steps of: scanning the patient to provide a map of a target surgical location of a defect in the vessel or duct;programming coordinates of the target surgical location relative to the map in a computer controller;inserting the device through a small opening in the patient, so that no portion of the device extends out of the patient and the device is not connected to anything that extends out of the patient;wherein the device comprises a plurality of semiconductor cells held in assembly primarily by magnetic attraction forces;moving and guiding the device through an anatomy of the patient to the target surgical location, wherein guiding is performed by registering a device mover with the map to facilitate navigation of the device through the anatomy and to the target surgical location; andfixing the device to the target surgical location. 2. The method of claim 1, wherein said moving and guiding are performed by a nuclear magnetic resonance machine, controlled by the computer controller and registered to the map to provide magnetic fields to move the device through the anatomy according to coordinates referenced to the map. 3. The method of claim 1, wherein said scanning is performed by a CAT scanner. 4. The method of claim 1, wherein said fixing comprises fusing at least a portion of the device to a tissue of the vessel or duct wall. 5. The method of claim 1, further comprising assembling the device from a plurality of cells, wherein said plurality of cells are guided, by a nuclear resonance machine, into position for assembly. 6. The method of claim 5, wherein the plurality of cells are assembled by polarizing portions thereof to attract to one another. 7. The method of claim 5, wherein the device is assembled to have a size determined to cover the defect, as determined by the coordinates of a location of a defect dimension. 8. A minimally invasive method of implanting a device in a patient, said minimally invasive method comprising the steps of: scanning the patient to provide a map of a target surgical location in the patient;programming coordinates of the target surgical location relative to the map in a computer controller;inserting the device through a small opening in the patient, so that no portion of the device extends out of the patient;wherein the device comprises a plurality of semiconductor cells held in assembly primarily by magnetic attraction forces;moving and guiding the device through an anatomy of the patient to the target surgical location, wherein guiding is performed by registering a device mover with the map to facilitate navigation of the device through the anatomy and to the target surgical location; andfixing the device to the target surgical location. 9. The method of claim 8, wherein said moving and guiding are performed by a nuclear magnetic resonance machine, controlled by the computer controller and registered to ... | - | A61B005/05 A61B005/00 G01R033/28 G01R033/48
US09351659 | B2 | 2016-05-31 | Biomedical electrode configuration for suppressing movement artifact | 2009-08-24 | A biomedical sensor for detecting EMG signals may include a non-conducting fixed framework supporting a symmetrical arrangement of four electrode surfaces, configured as two signal detection contacts, each with a respective associated signal reference contact. The mechanical and electrical configuration act together to suppress movement artifact. | 1. A biomedical sensor for sensing phenomena of a muscle, comprising: a substrate; a plurality of conductive areas arranged to be spaced along a length of the substrate and configured as signal detection contacts each with a proximate, distinct respective associated signal reference contact, forming a plurality of signal detection/reference contact pairs spaced along the length, each one in the plurality of signal detection/reference contact pairs being aligned in a direction transversely to the length and in parallel with respect to each other so that each signal detection/reference contact pair produces movement induced artifacts in unison in response to relative movement between the plurality of signal detection/reference contact pairs and the muscle; and a differential sensing circuit that receives input from each of the signal detection/reference contact pairs and provides an output that is an arithmetic difference between the signals detected at detection contacts measured with respect to the reference contacts, the reference contacts being directly electrically coupled together and to a reference input of the differential sensing circuit, and the detection contacts being provided at separate inputs to the differential sensing circuit, wherein each signal detection/reference contact pair is configured to react in response to the relative movement, wherein at least one common electrical component caused by the relative movement is cancelled out by the differential sensing circuit, and wherein the at least one common electrical component caused by the relative movement includes the movement induced artifacts that are suppressed by the differential sensing circuit. 2. The biomedical sensor as in claim 1, wherein the substrate is rigid. 3. The biomedical sensor as in claim 1, further comprising: a framework coupled to the substrate, wherein the substrate is a body-directed electrically insulating bottom surface of the framework. 4. The biomedical sensor as in claim 1, wherein the plurality of signal detection/reference contact pairs are spaced apart so that the signal detection contact is aligned to be on the same side of each pair. 5. The biomedical sensor as in claim 1, wherein each of the conductive areas is formed in the shape of a conductive rod aligned in parallel on a plane of the substrate, and projecting from the substrate. 6. The biomedical sensor as in claim 1, wherein each of the conductive areas is formed in the shape of a disc attached to and coplanar with the substrate. 7. The biomedical sensor as in claim 1, wherein each of the conductive areas is formed in a shape of a dome whose cross-sectional profile exhibits a concave curved shape attached to and projecting from the substrate. 8. The biomedical sensor as in claim 1, wherein each of the conductive contact areas are formed of substantially identical material composition, size and shape. 9. The biomedical sensor as claim 1, wherein each of the conductive areas is forme... | Altec, Inc. | A61B005/04 A61B005/0488 A61B005/0492 A61B005/00
US20110028865 | A1 | 2011-02-03 | Inertial Sensor Kinematic Coupling | 2009-08-03 | A method is disclosed for measuring the motion of an object, composed of multiple segments connected by joints, via the estimation of the 3D orientation of the object segments relative to one another without dependence on a magnetic field as a reference for heading. The method includes first applying a plurality of inertial sensor units to the segments of the object, e.g., a user thigh, shank, foot, etc. Next an approximation of the distance between each inertial sensor unit and at least one adjacent joint is provided and the joint is subjected to an acceleration, e.g., as the user takes a step or two. The relative orientations of the segments are calculated and the orientations are used to form an estimation of the 3D orientation of the object segments relative to one another without using the local magnetic field as a reference for heading. | 1. A method of measuring the motion of an object composed of multiple segments connected by joints via the estimation of the 3D orientation of the object segments relative to one another, without dependence on the Earth magnetic field as a reference for heading, the method comprising: applying a plurality of inertial sensor units to respective ones of the multiple segments;subjecting the joint to an acceleration;calculating the relative orientation of each segment with respect to each other based on data from the sensor units; andusing the orientations of the segments to form an estimation of the 3D orientation of the object segments relative to one another without using the local magnetic field as a reference for heading. 2. The method of measuring the motion of an object according to claim 1, wherein calculating the relative orientation of each segment further comprises comparing the measured accelerations from a first inertial sensor and a second inertial sensor at the location of the joint. 3. The method of measuring the motion of an object according to claim 1, further comprising calculating the distance between each sensor and each adjacent joint based on data from the sensors. 4. The method of measuring the motion of an object according to claim 1, wherein using the orientations of the segments without using the local magnetic field as a reference for heading comprises calculating position and orientation of the object. 5. The method of measuring the motion of an object according to claim 1, wherein the object is a human body. 6. The method of measuring the motion of an object according to claim 1, further comprising providing the estimation of 3D orientation to one of a motion capture system, Virtual Reality system and an Augmented Reality system. 7. The method of measuring the motion of an object according to claim 1, wherein the object is a robotic device. 8. A computer-readable medium having thereon computer-executable instructions for measuring the motion of an object composed of multiple segments connected by joints via the estimation of the 3D orientation of the object segments relative to one another, without dependence on the Earth magnetic field as a reference for heading, the object having a plurality of inertial sensor units affixed thereto upon respective ones of the multiple segments, the computer-executable instructions comprising: instructions for receiving data from one or more of the inertial sensor units indicating that that one or more joints have been subjected to acceleration;instructions for calculating the relative orientation of each segment with respect to each other based on the received data from the sensor units; andinstructions for using the orientations of the segments to form an estimation of the 3D orientation of the object segments relative to one another without using the local magnetic field as a reference for heading. 9. The computer-readable medium according to claim 8, wherein the instructions... | Xsens Technologies, B.V. | A61B005/11
US20110009736 | A1 | 2011-01-13 | Localization Using Non-Metallic Implantable Fiducial Markers | 2009-07-09 | A system may include a biocompatible housing and an electrical circuit disposed within the housing, the electrical circuit to emit a radiofrequency signal, and wherein the electrical circuit does not comprise a metal. Also included may be a signal generator to generate a wireless signal to trigger the electrical circuit, a receiver to receive a wireless response signal generated by the triggered electrical circuit, a processor to determine a location of the biocompatible housing based on the wireless response signal, and a ion beam source to deliver a ion beam to a patient volume including the biocompatible housing. | 1. A system comprising: a biocompatible housing; andan electrical circuit disposed within the housing, the electrical circuit to emit a radiofrequency signal,wherein the electrical circuit does not comprise a metal. 2. A system according to claim 1, wherein the electrical circuit comprises: a carbon nanocoil; anda capacitor electrically coupled to the carbon nanocoil. 3. A system according to claim 2, wherein the capacitor comprises carbon nanotubes. 4. A system according to claim 1, further comprising: a signal generator to generate a wireless signal to trigger the electrical circuit;a receiver to receive a response signal generated by the triggered electrical circuit; anda processor to determine a location of the biocompatible housing based on the response signal. 5. A system according to claim 4, the processor to determine that the location conforms to a treatment plan, the system further comprising: a ion beam source to deliver a ion beam associated with the treatment plan. 6. A system according to claim 5, wherein the electrical circuit comprises: a carbon nanocoil; anda capacitor electrically coupled to the carbon nanocoil. 7. A system according to claim 1, further comprising: a signal generator to generate a wireless signal to trigger the electrical circuit;a receiver to receive a wireless response signal generated by the triggered electrical circuit; anda processor to determine a difference between the wireless response signal and a previous wireless response signal received from the electrical circuit. 8. A system according to claim 7, the processor to determine that the difference exceeds a threshold, the system further comprising: an imaging system to generate an image of a patient volume in which the biocompatible housing is located in response to the determination; anda positioning system to change a position of the patient volume with respect to a ion beam source based on the image. 9. A system according to claim 7, the processor to determine that the difference exceeds a threshold, the system further comprising: an imaging system to generate an image of a patient volume in which the biocompatible housing is located in response to the determination; anda positioning system to change one or more characteristics of a ion beam to be delivered by a ion beam source based on the image. 10. A system according to claim 7, wherein the electrical circuit comprises: a carbon nanocoil; anda capacitor electrically coupled to the carbon nanocoil. 11. A method comprising: generating a wireless signal to trigger an electrical circuit within a biocompatible housing disposed in a target volume; andreceiving a wireless response signal from the electrical circuit,wherein the electrical circuit does not comprise a metal. 12. A method according to claim 11, further comprising: determining a location of the biocompatible housing based on the wireless response signal. 13. A method according to claim 12, further comprising: determining that the ... | - | A61B005/05 G08B001/08 A61N005/00
US07884604 | B2 | 2011-02-08 | Systems and methods for reconstruction of sensitivity encoded MRI data | 2009-03-30 | Methods and systems in a parallel magnetic resonance imaging (MRI) system utilize sensitivity-encoded MRI data acquired from multiple receiver coils together with spatially dependent receiver coil sensitivities to generate MRI images. The acquired MRI data forms a reduced MRI data set that is undersampled in at least a phase-encoding direction in a frequency domain. The acquired MRI data and auto-calibration signal data are used to determine reconstruction coefficients for each receiver coil using a weighted or a robust least squares method. The reconstruction coefficients vary spatially with respect to at least the spatial coordinate that is orthogonal to the undersampled, phase-encoding direction(s) (e.g., a frequency encoding direction). Values for unacquired MRI data are determined by linearly combining the reconstruction coefficients with the acquired MRI data within neighborhoods in the frequency domain that depend on imaging geometry, coil sensitivity characteristics, and the undersampling factor of the acquired MRI data. An MRI image is determined from the reconstructed unacquired data and the acquired MRI data. | 1. A method for reconstructing an actual image from sensitivity encoded magnetic resonance imaging (MRI) data capable of being represented in a frequency domain, the frequency domain comprising: at least a phase-encoding direction and a slice-encoding direction, the phase-encoding and slice-encoding directions corresponding to a first spatial dimension and a second spatial dimension, respectively, the method comprising: based on a reduced MRI data set, acquired from each of two or more receiver coils, the reduced MRI data set configured to be represented in a first portion of the frequency domain, wherein the MRI data set is undersampled in at least the phase-encoding direction;determining reconstruction coefficients for each of the two or more receiver coils, the reconstruction coefficients being spatially varying in at least the second spatial dimension which corresponds to the slice-encoding direction;determining a reconstructed MRI data set, configured to be represented in a second portion of the frequency domain, for each of the two or more receiver coils by combining the reconstruction coefficients and the reduced MRI data set from at least one of the two or more receiver coils, the second portion comprising representative locations in the frequency domain wherein the reduced MRI data set has not been acquired;determining a complete MRI data set for each of the two or more receiver coils by combining the reduced MRI data set and the reconstructed MRI data set; andoutputting at least a portion of the complete MRI data set into an output device configured for display or a non-transitory computer readable medium. 2. The method of claim 1, wherein the reduced MRI data set comprises an auto-calibration signal data set. 3. The method of claim 1, wherein the act of determining reconstruction coefficients further comprises: dividing at least a portion of the reduced MRI data set into overlapping blocks along at least the second spatial dimension corresponding to the slice-encoding direction;determining block reconstruction coefficients for each of the overlapping blocks; andinterpolating among the block reconstruction coefficients to determine values of the reconstruction coefficients at a location in the second spatial dimension. 4. The method of claim 1, wherein the act of determining reconstruction coefficients further comprises using a weighted least squares technique having one or more weights in order to determine the reconstruction coefficients, the weights being inversely related to energies of the detected signals of the reduced MRI data set or related to distances from a center of the frequency domain. 5. The method of claim 1, wherein the act of determining reconstruction coefficients further comprises using a robust least squares technique in order to determine the reconstruction coefficients, wherein the robust least squares technique excludes a portion of the reduced MRI data set from the act of determining reconstruction coeffici... | University of Utah Research Foundation | G01V003/00 A61B005/05
US07881573 | B2 | 2011-02-01 | Optical imaging probe connector | 2009-10-02 | This document discusses, among other things, a connector for an optical imaging probe that includes one or more optical fibers communicating light along the catheter. The device may use multiple sections for simpler manufacturing and ease of assembly during a medical procedure. Light energy to and from a distal minimally-invasive portion of the probe is coupled by the connector to external diagnostic or analytical instrumentation through an external instrumentation lead. Certain examples provide a self-aligning two-section optical catheter with beveled ends, which is formed by separating an optical cable assembly. Techniques for improving light coupling include using a lens between instrumentation lead and probe portions. Techniques for improving the mechanical alignment of a multi-optical fiber catheter include using a stop or a guide. | 1. An apparatus comprising: an optical coupler, the coupler comprising a first central body and a plurality of first optical fibers disposed about and extending longitudinally along the first central body, the first central body and the plurality of first optical fibers providing a first commonly beveled end;wherein the first optical fibers extend away from the first beveled end, the coupler configured to accept an elongated member configured for imaging within an object, the elongated member including a second central body and a plurality of second optical fibers disposed about and extending longitudinally along the second central body, the elongated member having a second commonly beveled end;wherein the coupler and the elongated member are configured to permit an end-user to insert a beveled end of the elongated member into the coupler, with the elongated member extending away from the second beveled end and the coupler, such that the second beveled end of the elongated member butts against and mates in end-user-provided contacting self-alignment to the first beveled end to couple light between the plurality of the first optical fibers and the contacting self-aligned plurality of the second optical fibers of the elongated member; andwherein the pluralities of the first and second optical fibers are cut, at their respective beveled ends, from a single optical fiber assembly comprising a central body and a plurality of optical fibers disposed about the central body, to permit the end-user-provided contacting self-alignment, wherein the plurality of first optical fibers are disposed about the central body in an encapsulation placement, wherein the encapsulation placement permits relative spatial variations in the placement of the first optical fibers along a length of the optical fiber assembly according to equipment or process variation of the encapsulation. 2. The apparatus of claim 1, further comprising the elongated member configured for imaging within an object, the elongated member including a second central body and a plurality of second optical fibers disposed about and extending longitudinally along the second central body, the elongated member having a second commonly beveled end. 3. The apparatus of claim 2, wherein the elongated member comprises a second commonly beveled end forming a common second plane extending obliquely through both the second central body and the plurality of second optical fibers; and wherein the first central body and the plurality of first optical fibers provide a first commonly beveled end forming a common first plane extending obliquely through both the first central body and the first optical fibers. 4. The apparatus of claim 2, in which the elongated member includes proximal and distal ends and comprises at least one acousto-optical transducer at or near the distal end. 5. The apparatus of claim 1, in which the encapsulation comprises a plastic matrix. 6. The apparatus of claim 1, in which the coupler... | Vascular Imaging Corporation | G02B006/26 C23C016/00 A61B005/05
US07881761 | B2 | 2011-02-01 | Method and circuit for storing and providing historical physiological data | 2006-06-01 | Embodiments of the present invention relate to a sensor for facilitating detection of a physiological characteristic of a patient. Specifically, one embodiment includes a detector configured to obtain signals from the patient, the signals being indicative of the physiological characteristic, an interface configured to send the signals to a monitor, and a sensor memory coupled with the sensor and separate from the monitor, the sensor memory storing sensor expiration data. | 1. A pulse oximeter system, comprising: a sensor and a monitor, the sensor comprising: a detector configured to obtain data indicative of a physiological characteristic from a patient; anda sensor memory coupled with the sensor and separate from the monitor, the sensor memory storing sensor expiration data; andthe monitor comprising: a first component configured to read the data and derive patient physiological data from the data; anda second component configured to read the sensor expiration data and configured to disable the sensor based on the sensor expiration data, wherein the second component is configured to pass a high current through the sensor to disable the sensor based on the sensor expiration data. 2. The system of claim 1, wherein the second component is configured to provide a status of the sensor to a screen of the monitor for display. 3. The system of claim 2, wherein the second component is configured to provide an indication that the sensor is expired as the status. 4. The system of claim 1, wherein the sensor expiration data comprises a warranty date. | Mallinckrodt Inc. | A61B005/1455
US07877128 | B2 | 2011-01-25 | Simulation of invasive procedures | 2005-08-02 | A method for pre-planning and performing a cardiac procedure on a heart includes acquiring an image or map of the heart; displaying the image or map of the heart; marking at least one feature on the image or map; calculating dimensions of the at least one feature; identifying one or more points on or within the heart for treatment; determining paths to the one or more points on or within the heart for treatment; simulating insertion of a sheath into the heart; simulating insertion of a medical device through the sheath and within the heart; verifying that the one or more points on or within the heart can be accessed for treatment; and performing a medical procedure on or within the heart. | 1. A method for pre-planning a cardiac procedure on a heart comprising the steps of: acquiring an image or map of the heart;displaying the image or map of the heart;marking at least one feature on the image or map;calculating dimensions of the at least one feature;identifying one or more points on or within the heart for treatment;determining paths based on the at least one marked feature and the dimensions of the at least one marked feature to the one or more points on or within the heart for treatment;simulating insertion of a sheath into the heart through one of the paths;simulating insertion of a medical device through the sheath and within the heart;verifying that the one or more points on or within the heart can be accessed by the medical device for treatment;measuring a penetration depth of the sheath in the heart using the step of simulating insertion of the sheath;verifying that the penetration depth does not exceed a specified measurement;confirming that the insertion of the sheath avoids injury to the heart at the penetration depth of the sheath using the step of simulating insertion of the sheath; andsimulating insertion of selected ablation catheters through the sheath. 2. The method of claim 1, further comprising: selecting ablation catheters having different radii of curvature for the step of simulating insertion of the selected ablation catheters through the sheath; anddetermining the paths to the heart of the ablation catheters in relation to the penetration depth of the sheath. 3. A method for developing a plan for a cardiac procedure comprising the steps of acquiring an image or map of the heart;displaying the image or map of the heart;marking at least one feature on the image or map;calculating dimensions of the at least one feature;identifying one or more points on or within the heart for treatment;determining paths based on the at least one marked feature and the dimensions of the at least one marked feature to the one or more points on or within the heart for treatment;simulating insertion of a sheath into the heart through one of the paths;simulating insertion of a medical device through the sheath and within the heart; andverifying that the one or more points on or within the heart can be accessed by the medical device for treatment;measuring a penetration depth of the sheath in the heart using the step of simulating insertion of the sheath;verifying that the penetration depth does not exceed a specified measurement;confirming that the insertion of the sheath avoids injury to the heart at the penetration depth of the sheath using the step of simulating insertion of the sheath; andsimulating insertion of selected ablation catheters through the sheath. 4. A method for pre-planning and performing a cardiac procedure on a heart comprising the steps of: acquiring an image or map of the heart;displaying the image or map of the heart;marking at least one feature on the image or map;calculating dimensions of the at least one f... | Biosense Webster, Inc. | A61B005/05
US07883479 | B1 | 2011-02-08 | Stability augmentation system | 2005-08-05 | A stability augmentation system for stabilizing human limbs from against limb tremors. The stability augmentation system includes: and arm brace; a plurality of piezoelectric sensors fixed on said arm brace for sensing angular motion about a first axis of a limb induced by alternating perturbations; and a plurality of piezoelectric actuators coupled to the limb and piezoelectric sensor, said piezoelectric actuator being responsive to the piezoelectric sensors for selectively applying a torque to the limb representative of the sensed angular motion such that the angular motion about the first axis induced by the perturbation is cancelled, said plurality of piezoelectric actuators concurrently applying a reactive torque, equal and opposite to the applied torque, to the limb. | 1. A stability augmentation system for stabilizing human limbs against limb tremors, comprising: a brace mountable to a limb; a plurality of sensors mounted on said brace for sensing angular motion about a first axis of a limb, induced by alternating perturbations; a series of actuators coupled to said limb and one or more of said sensors, said actuators being responsive to said sensors for selectively applying a reactive torque to said limb, representative of the sensed angular motion such that the angular motion about said first axis, induced by the perturbation, is cancelled or dampened, said reactive torque being applied without pushing against an external mass member. 2. The stability augmentation system of claim 1 having an inertial sensor which includes a feedback control system for measuring the angular rotation about said limb. 3. The stability augmentation system of claim 2 in which said brace is a hook and loop strap which holds and aligns said actuators along the path of the expected tremor. 4. The stability augmentation system of claim 3 in which said actuators include a series of piezoelectric stacks as piezoelectric actuators, each separated by a brace element such that said piezoelectric actuators have the dual ability to apply expansion or contraction force based on electrical stimulation and to create a voltage differential based on mechanical displacement. 5. A stability augmentation system for stabilizing human limbs against limb tremors, comprising: a brace mountable to a limb; a series of piezoelectric sensors/actuators mounted on said brace, producing sensor signals from sensing angular motion about a first axis of a limb, induced by alternating perturbations, said series of piezoelectric sensors/actuators coupled to said brace and said limb, said piezoelectric sensors/actuators being responsive to said sensor signals for selectively applying a reactive torque to said limb, representative of the sensed angular motion, such that the angular motion about the first axis, induced by the perturbation is cancelled or dampened, said reactive torque being applied without pushing against an external mass member. 6. The stability augmentation system of claim 5 in which the sensors/actuators include a feedback control system for measuring the angular rotation about said limb. 7. The system of claim 1 wherein said brace is a resilient pad which is adapted for snug fit to said limb, said actuators being rigidly connected to each other in series to cumulatively expand and contract, which series is mounted to said brace pad so that said actuators and pad can expand and contract on said limb to cancel or dampen tremors thereof. | The United States of America as represented by the Secretary of the Air Force | A61B005/103 A61B005/117
US08106781 | B2 | 2012-01-31 | Device for monitoring the condition of a human being | 2010-08-13 | A device for continuously monitoring one or more indicators relating to the health or comfort of a person 12 includes one or more sensing means for measuring a value of one or more desired indicators and a communicating means 13 communicating an alarm signal when the measured value is above or below a predetermined value or range of values range, wherein the device 10 is adapted to be attached to an item of clothing 14 of the person 12. | 1. A health monitoring device for monitoring one or more parameters associated with a body of a person, the health monitoring device comprising: a housing; anda clip having a first end integral with and cantilevered to the housing and a second end remote from the first end, the clip being substantially parallel to the housing to define a slot with the housing for receiving a portion of a garment therein, whereby the health monitoring device is capable of being secured to the garment, the clip comprising a pressure sensing arrangement to convert abdominal pressure exerted on the clip due to breathing into a signal, where the housing comprises a circuit to activate an alarm based on the signal. 2. The health monitoring device of claim 1, wherein the pressure sensing arrangement comprises a piezoelectric element that generates the signal, wherein the signal is a fluctuating electronic potential corresponding to the abdominal pressure exerted on the clip, the signal is measured by the circuit, wherein the circuit activates the alarm when the signal is abnormal. 3. The health monitoring device of claim 2, wherein the pressure sensing arrangement comprises an outer flexible membrane, the outer flexible membrane is capable of transmitting the abdominal pressure exerted on the clip to piezoelectric element. 4. The health monitoring device of claim 3, wherein the outer flexible membrane is made of a polyester or polycarbon. 5. The health monitoring device of claim 1, further comprises a stimulating device coupled to the circuit and located in the housing, wherein the circuit activates the stimulating device when the signal is abnormal. 6. The health monitoring device of claim 5, the stimulating device comprises an offset weight and a motor, the motor being capable of energizing the offset weight to cause vibration when the circuit activates the stimulating device. 7. The health monitoring device of claim 1, further comprising a first temperature sensor incorporated in the clip for sensing and measuring body temperature, wherein the temperature sensor is coupled to the circuit, and wherein the circuit activates the alarm when the body temperature is abnormal. 8. The health monitoring device of claim 7, further comprising a second temperature sensor attached to the housing, the second temperature sensor measures a second temperature parameter, wherein the body temperature and the second temperature parameter are compared continuously to calculate a rate of heat loss, wherein the circuit activates the alarm when the rate of heat loss is abnormal. 9. The health monitoring device of claim 1, further comprising a first humidity sensor incorporated in the clip for sensing and measuring body humidity, wherein the humidity sensor is coupled to the circuit, and wherein the circuit activates the alarm when the body humidity is abnormal. 10. The health monitoring device of claim 9, further comprising a second humidity sensor attached to the housing member, t... |  | G08B023/00
US09131885 | B2 | 2015-09-15 | Analyte sensors and methods of manufacturing same | 2010-07-01 | Analyte sensors and methods of manufacturing same are provided, including analyte sensors comprising multi-axis flexibility. For example, a multi-electrode sensor system 800 comprising two working electrodes and at least one reference/counter electrode is provided. The sensor system 800 comprises first and second elongated bodies E1, E2, each formed of a conductive core or of a core with a conductive layer deposited thereon, insulating layer 810 that separates the conductive layer 820 from the elongated body, a membrane layer deposited on top of the elongated bodies E1, E2, and working electrodes 802′, 802″ formed by removing portions of the conductive layer 820 and the insulating layer 810, thereby exposing electroactive surface of the elongated bodies E1, E2. | 1. A continuous analyte sensor configured for in vivo use, the continuous analyte sensor comprising: an elongated conductive body comprising a working electrode, wherein the elongated conductive body has a Young's modulus of from about 170 GPa to about 220 GPa and a yield strength from about 800 MPa to about 3,000 MPa; anda membrane covering at least a portion of the working electrode, wherein the membrane comprises an enzyme layer. 2. The continuous analyte sensor of claim 1, wherein the elongated conductive body comprises an elongated core, an insulating layer covering at least a portion of the elongated core, a conductive layer comprising a reference electrode or a counter electrode and covering at least a portion of the insulating layer, and a membrane covering at least a portion the working electrode. 3. The continuous analyte sensor of claim 2, wherein the elongated core comprises an elongated body and a layer of conductive material covering at least a portion of the elongated body. 4. The continuous analyte sensor of claim 3, wherein the elongated body comprises at least one material selected from the group consisting of stainless steel, titanium, tantalum, and a polymer. 5. The continuous analyte sensor of claim 3, wherein the layer of conductive material covering at least a portion of the elongated body comprises a conductive material selected from the group consisting of platinum, platinum-iridium, gold, palladium, iridium, alloys thereof, graphite, carbon, and a conductive polymer. 6. The continuous analyte sensor of claim 2, wherein the conductive layer comprising the reference electrode or the counter electrode comprises a silver-containing material. 7. The continuous analyte sensor of claim 6, wherein the silver-containing material has a particle size associated with a maximum particle dimension that is less than about 100 microns. 8. The continuous analyte sensor of claim 7, wherein the silver-containing material has a particle shape that is substantially spherical. 9. The continuous analyte sensor of claim 2, wherein the insulating layer comprises at least one polymer selected from the group consisting of polyurethane and polyimide. 10. The continuous analyte sensor of claim 2, wherein a ratio of a thickness of the conductive layer to a thickness of insulating layer is from about 1:5 to about 1:1. 11. The continuous analyte sensor of claim 2 wherein the conductive material is a silver-containing material. 12. The continuous analyte sensor of claim 1, wherein the membrane comprises a polymer having a Shore hardness of from about 70 A to about 55 D. 13. The continuous analyte sensor of claim 1, having an ultimate tensile strength of from about 80 kPsi to about 500 kPsi. 14. The continuous analyte sensor of claim 1, having an ultimate tensile strength of from about 150 kPsi to about 280 kPsi. 15. The continuous analyte sensor of claim 1, having a fatigue life of at least 1,000 cycles of flexing of from about 28° to abou... | DexCom, Inc. | A61B005/05 A61B005/1473 A61B005/145 A61B005/1486 B05C003/10
US08401603 | B2 | 2013-03-19 | Imaging device, method thereof, and program | 2010-09-28 | To realize an imaging device, imaging method and program capable of improving image quality. This invention controls light sources so as to irradiate a body with irradiation light of a luminance level higher than that in the air arriving at the body, adjusts the imaging sensitivity of a solid imaging element for performing photoelectric conversion on pattern scattered light obtained through the inside of the body, and masks the luminance level deterioration region of a pattern image obtained from the solid imaging element to treat the masked image as an image to be composed with a reference image. | 1. An imaging device comprising: an ambient light measurement receiving unit configured to receive a measurement of a luminance level of light in the air arriving at a body;a light source control unit configured to control light sources to irradiate the body with irradiation light of a luminance level greater than the measurement of the luminance level of air surrounding the body;an imaging element configured to perform photoelectric conversion on a pattern of light scattered through an inside of the body;a sensitivity adjustment unit configured to adjust imaging sensitivity of the imaging element; andan image generating unit configured to generate an image that excludes luminance level deterioration regions by replacing the luminance level deterioration regions with a reference image generated at a different sensitivity. 2. The imaging device according to claim 1, wherein: the light source control unit is configured to control the light sources so as to sequentially irradiate the body with irradiation light every unit time; andthe image generating unit is configured to remove the luminance level deterioration regions from the image every unit time and to compose the image with the reference image. 3. The imaging device according to claim 1, wherein: the light source control unit is configured to control the light sources so as to sequentially irradiate the body every unit time at luminance levels different from each other; andthe image generating unit is configured to remove the luminance level deterioration region of the image every unit time and compose the image with the reference image. 4. The imaging device according to claim 1, wherein: the sensitivity adjustment unit is configured to adjust the imaging sensitivity by sequentially changing an amount of charge signal accumulated in the imaging element as a result of the photoelectric conversion for each unit time; andthe image generating unit is configured to remove the luminance level deterioration region of the image based on the amount of charge obtained from the imaging element and to compose the image with the reference image. | Sony Corporation | A61B005/00 A61B006/00
US09101303 | B2 | 2015-08-11 | Analyte sensor ports | 2010-08-18 | The present disclosure provides a contamination resistant sensor port which includes one or more sealing members positioned so as to limit and/or prevent internal contamination of the sensor port with fluids and/or particles present in the environment outside the sensor port. | 1. A replaceable sensor port for use with an analyte meter, comprising: a housing comprising a top portion and a bottom portion, wherein the top portion and bottom portion are joined to form a body and a protruding member of the housing;an entry slot formed by the top and bottom portions to receive an analyte sensor and comprising a first edge having an internal beveled face to prevent ingress of external contaminants through the entry slot when an analyte sensor is received within the entry slot, and a second edge without an internal beveled face, wherein the second edge is in opposite and facing configuration with the first edge;an electrical interface disposed within an internal opening in the body of the housing, wherein the electrical interface is electrically coupled to an analyte sensor when the sensor is received within the entry slot;a first sealing member extending around the protruding member and positioned wholly between the entry slot and the body of the housing, wherein the first sealing member provides a seal between the sensor port and an analyte meter into which the sensor port is integrated when the sensor port is coupled to the analyte meter; anda plurality of electrical contacts extending from an exterior surface of the housing to electrically couple with the analyte meter. 2. The replaceable sensor port of claim 1, further comprising a second sealing member disposed between the plurality of electrical contacts and the exterior surface of the housing extending around an edge of the body of the housing. 3. The replaceable sensor port of claim 1, wherein the first sealing member is an oval shape, or an elliptical shape. 4. The replaceable sensor port of claim 1, wherein the analyte sensor is a glucose test strip or a ketone test strip. 5. The replaceable sensor port of claim 1, wherein the entry slot further comprises a third edge having an internal beveled face and a fourth edge having an internal beveled face. 6. The replaceable sensor port of claim 1, wherein the distal end of the protruding member comprises a single flush surface. 7. The replaceable sensor port of claim 6, wherein the single flush surface has the same shape as the first sealing member. 8. The replaceable sensor port of claim 6, wherein the entry slot further comprises a third edge and a fourth edge in an opposite and facing configuration with the third edge, and wherein the single flush surface comprises the first edge, second edge, third edge and fourth edge. 9. The replaceable sensor port of claim 1, wherein the distal end of the protruding member defines a periphery having an elliptical shape. 10. The replaceable sensor port of claim 1, wherein the entry slot further comprises a third edge that is curved and a fourth edge that is curved and on an opposite side of the entry slot than the third edge. 11. The replaceable sensor port of claim 1, wherein the entry slot comprises a flush surface about the circumference of the entry slot. 12. A replac... | Abbott Diabetes Care Inc. | A61B005/05 A61B005/145 A61B005/1455 A61B005/1468 H01R013/52
US08535060 | B2 | 2013-09-17 | System and method for detecting a specific cognitive-emotional state in a subject | 2007-08-27 | A system and method is disclosed which detects a specific cognitive-emotional mental state (e.g., passion) in a subject. The system and method utilizes different Internal Personalized Explorer (IPE) stimuli subliminally exposed to the subject on a visual display. Following exposure, the subject performs a target-specific cognitive task and a related behavior (e.g., deciding whether a briefly displayed string of letters is a word or not, and pressing a button to indicate the decision). The decision reaction lime for the behavior is recorded in a computer database. The reaction times are interpreted quantitatively to detect the relative intensity of the specific cognitive-emotional mental state in the subject. A reaction time for a specific individual IPE that is shorter than the mean reaction time for an IPE that does not evoke the emotional state indicates that the specific individual IPE did evoke the emotional state in the subject. | 1. An un-blinded behavioral method for detecting a specific cognitive-emotional mental state in a subject, the method comprising the steps of: consciously providing at least two different conscious and subjective stimuli, selected from one or more of a group of conscious, subjective stimuli designed to match a particular cognitive-emotional state in a particular subject, at least one of the group of conscious subjective stimuli being a test stimulus for testing whether the subject associates the test stimulus with a subsequent target-specific cognitive task that incorporates non-conscious components designed to match the particular cognitive-emotional state sought to be detected, said group including also a neutral conscious stimulus for which the subject is neutrally acquainted and an indifferent conscious, subjective stimulus which represents no particular significance to the subject, such stimuli referred to hereinafter as Internal Personalized Explorer (IPE) stimuli;exposing the subject to a visual display of an individual IPE stimulus on a visual display device for a subliminal time duration, the visual display device being an output device of a computer system;masking the visual display with a mask for a brief non-subliminal time duration;following the masking step with the subject performing a meaningful target-specific cognitive task and a related behavior, wherein such target includes a sensory stimulus that incorporates non-conscious components designed to match the particular cognitive emotional state sought to be detected;presenting on the visual display a randomly varied response period having at least a response period duration sufficient to complete the behavior, and recording a reaction time, whether voluntary or involuntary, and the accuracy of the behavior via an input device of the computer system in a database on the computer system;repeating the exposing step a sufficient number of times with different individual IPEs to generate a database sufficient to determine a mean reaction time of the subject relative to each of the individual IPEs; andinterpreting the reaction time and accuracy of data of the data set, to detect the relative intensity of the specific cognitive-emotional mental state in the subject relative to the individual IPEs. 2. The un-blinded behavioral method of claim 1, wherein the providing step further comprises the IPEs being selected from the classes of IPE consisting of: a favored IPE, a neutral IPE and an indifferent IPE. 3. The un-blinded behavioral method of claim 1, wherein the exposing step further comprises the exposing of the individual IPE stimulus on the visual display for a time duration that is below the threshold of conscious perception of the stimulus and above the time duration for unconscious perception of the stimulus by the subject. 4. The un-blinded behavioral method of claim 1, wherein the exposing step further comprises the exposing of the individual IPE stimulus on the visual displa... | Brain & Science LLC | G09B023/28
US08535282 | B2 | 2013-09-17 | Wound healing sensor techniques | 2009-07-14 | The present disclosure is directed at methods and apparatus to evaluate and monitor healing progress of wound and/or to generate data to modify or identify a treatment protocol. Fluid exudate removed from the wound by a negative pressure therapy device may now be analyzed to identify the progress of wound healing, such as whether healing is progressing in a positive manner or experiencing one or more impediments. The fluid exudates may be specifically analyzed for one or more analytes indicative of one or more of the biochemical reactions that may occur during wound recovery. In addition, one may separately utilize optical sensors integrated into a dressing enclosure, optionally in those dressings employed in a negative pressure therapy device. Such optical sensors may then illuminate the wound and collect information regarding the wound via a light scattering type response. | 1. An apparatus to evaluate and monitor healing progress of a wound and to generate data to modify or identify a treatment protocol comprising: a dressing enclosure for administering reduced pressure treatment to a wound, wherein the enclosure is configured to cover the wound and to maintain reduced pressure at the wound location;a device capable of supplying reduced pressure;a microfluidic device including a closed-channel continuous flow system connected between said dressing enclosure for said wound and said device capable of supplying reduced pressure to withdraw fluid exudates containing one or more analytes from said wound wherein said microfluidic device includes first and second compartments wherein in said first compartment an analyte that is one of said one or more analytes binds with one or more labeled antibodies or aptamers and in a second compartment the bound analyte is detected; anda sensor in said second compartment of said microfluidic device capable of analyzing said fluid exudates for one or more analytes indicative of said progress of wound healing and said analyzing of said fluid exudates is done continuously without disruption of said reduced pressure at said wound location and providing real time analysis of wound healing and providing real-time treatment protocol adjustments. 2. The apparatus of claim 1 wherein said closed-channel continuous flow system includes flow passages having a width or dept of 1 μm to 500 μm and wherein said exudate fluid is configured to flow in said flow passages at a rate of 1 μL/min to 100 μL/min. 3. The apparatus of claim 1 wherein said sensor analyzes said fluid exudate and provided an indication of fluid pH. 4. The apparatus of claim 1 wherein said sensor analyzes said fluid exudate and provides an indication of fluid oxygen level. 5. The apparatus of claim 1 wherein said sensor analyzes said fluid exudates for the presence of wherein said sensor analyzes said fluid exudates for the presence of one or more of the following: biomarkers of active bacterial infection, pro-inflammatory cytokines, proteases, and tissue growth factors. 6. The apparatus of claim 1 wherein a pH sensitive chromophore is introduced into said microfluidic device. 7. The apparatus of claim 1 wherein an oxygen sensitive chromophore is introduced into said fluid trap. 8. The apparatus of claim 1 wherein said sensor analyzes said fluid exudates for the presence of one or more of the following: biomarkers of active bacterial infection, pro-inflammatory cytokines, proteases, and tissue growth factors. 9. The apparatus of claim 1 wherein said sensor analyzes said fluid exudate for matrix metallo proteinase. 10. The apparatus of claim 1, further including an optical sensor in said dressing enclosure, wherein said optical sensor is configured to provide varying wavelengths of light to said wound location. 11. The apparatus of claim 1 wherein said optical sensor is configured to provide light at wavelengths of 350 nm ... | Southwest Research Institute | A61M001/00
US08383044 | B2 | 2013-02-26 | Blood sampling device | 2010-07-07 | A blood sampling device useful for collecting a blood sample from a separate vascular access device is described herein. The blood sampling device includes a body shaped and sized for partial insertion into a separate vascular access device. The body includes a reservoir defined within the body, which has an internal volume sufficient to contain enough blood for use in a diagnostic blood test. The body also includes a gas permeable vent disposed on the body, in which the gas permeable vent is in gaseous communication with the reservoir. When connected to a separate vascular access device the blood sampling device collects a blood sample as blood flows into the reservoir from the separate vascular access device and as gases pass out the reservoir via the gas permeable vent. | 1. A blood sampling device, comprising: a body having a proximal end, a distal end, and a distal opening, the distal end being configured for coupling with a separate vascular access device;a reservoir defined within the body, the distal opening comprising an opening of the reservoir;a gas permeable vent comprising a membrane, the gas permeable vent being disposed on the proximal end of the body, a distal surface of the gas permeable vent being in direct communication with the reservoir, and a proximal surface of the gas permeable vent being in direct communication with an exterior environment; anda diagnostic component in fluid communication with the reservoir. 2. The blood sampling device of claim 1, wherein the diagnostic component includes an on-board diagnostic cartridge in fluid communication with a chamber of the reservoir. 3. The blood sampling device of claim 1, wherein the diagnostic component includes a diagnostic reagent disposed within a chamber of the reservoir. 4. The blood sampling device of claim 1, wherein the reservoir includes multiple chambers, and wherein at least two of the multiple chambers includes a separate diagnostic component. 5. The blood sampling device of claim 1, wherein the diagnostic component includes a preservative. 6. The blood sampling device of claim 1, the body further comprising a flexible body which is compressible, wherein when the flexible body is compressed, an internal volume of the reservoir is reduced. 7. The blood sampling device of claim 1, wherein the body further comprises a time indicator. 8. The blood sampling device of claim 1, wherein the reservoir is in fluid communication with a separate vascular access device via the distal opening when the distal end of the body is coupled to the separate vascular access device. 9. The blood sampling device of 1, wherein the distal end of the body further comprises a male luer adapter. 10. The blood sampling device of claim 1, wherein the distal opening includes a flow restrictor. 11. The blood sampling device of claim 1, wherein the distal end of the body comprises a luer adapter. | Becton, Dickinson and Company | G01N015/06 G01N033/00 G01N033/48 A61B005/00 B65D081/00
US20110004089 | A1 | 2011-01-06 | EAR-WORN EEG MONITORING DEVICE | 2009-02-05 | An ear-worn EEG monitoring device is disclosed. The EEG monitoring device includes plural electrodes, an EEG signal acquisition circuitry with RF module, at least a housing for accommodating the EEG signal acquisition circuitry with RF module, and at least an ear-worn structure, wherein the ear-worn structure can be implemented to mount on one or two ear(s) of the user, so that during the EEG monitoring process, the ear-worn structure can be used to hold and support the housing and the EEG signal acquisition circuitry therein at position(s) above and including the neck of the user, and the RF module is used for achieving a wireless communication with an external apparatus, so as to transmit acquired signals thereto. | 1. An ear-worn EEG monitoring device, comprising: plural electrodes;an EEG signal acquisition circuitry with RF module;a housing, for accommodating the EEG signal acquisition circuitry; andan ear-worn structure, mounted on one single ear of the user;wherein during the EEG monitoring process,the ear-worn structure is used to hold and support the housing and the EEG signal acquisition circuitry therein at a position around said ear, andthe RF module is used for achieving a wireless communication with an external apparatus, so as to transmit acquired signals thereto. 2. The monitoring device as claimed in claim 1, wherein the housing is positioned behind the ear. 3. The monitoring device as claimed in claim 1, wherein the number of the ear-worn structure is implemented to be two. 4. The monitoring device as claimed in claim 3, wherein the number of the housing is implemented as two, and each housing is respectively attached with one single ear-worn structure. 5. The monitoring device as claimed in claim 4, wherein the EEG signal acquisition circuitry is separated into two electrically connected modules for locating in two housings. 6. The monitoring device as claimed in claim 3, wherein the housing is formed to connect with both single ear-worn structures. 7. The monitoring device as claimed in claim 1, wherein the ear-worn structure is integrated with at least one of the electrodes. 8. The monitoring device as claimed in claim 7, wherein the electrode integrated with the ear-worn structure is disposed on the ear. 9. The monitoring device as claimed in claim 7, wherein the electrode integrated with the ear-worn structure is disposed on the mastoid. 10. The monitoring device as claimed in claim 7, wherein the electrode integrated with the ear-worn structure is a reference electrode. 11. The monitoring device as claimed in claim 1, further comprising an electrode-positioning element, for fixing the electrodes. 12. The monitoring device as claimed in claim 1, wherein the external apparatus is used for monitor, analysis and/or storage. 13. The monitoring device as claimed in claim 1, wherein based on the RF module, the external apparatus configures the EEG monitoring device. 14. The monitoring device as claimed in claim 1, wherein the ear-worn EEG monitoring device further comprises a memory for storing EEG signals. 15. The monitoring device as claimed in claim 14, wherein the memory is a removable memory. 16. The monitoring device as claimed in claim 1, wherein the ear-worn EEG monitoring device further comprises a wired communication interface for communicating with the external apparatus. 17. An ear-worn EEG monitoring device, comprising: plural electrodes;an EEG signal acquisition circuitry with a RF module;a housing, for accommodating the EEG signal acquisition circuitry; andan ear-worn structure, mounted on both the user's ears, whereinduring the EEG monitoring process,the ear-worn structure is used to hold and support the housing ... | - | A61B005/0478
US20110015536 | A1 | 2011-01-20 | EEG-based method for determining a subject's compatibility with a work environment | 2009-07-17 | This invention provides an EEG-based method of determining whether a test subject possesses an attitude compatible with the work environment at a predetermined organization. | 1. A method of determining whether a test subject possesses an attitude compatible with employment at a predetermined organization comprising: a) recording a first electroencephalograph (EEG) from the test subject over a first period of time and performing a power component analysis on brain wave activity from the EEG of the test subject so as to thereby determine if the subject is alert enough to proceed to step b) of the method, and if so, proceeding to step b);b) recording a second EEG from the test subject over a second period of time and exposing the test subject during the second period of time to a series of visual images, which images have been predetermined to be emotionally neutral in a reference population of subjects, and quantitating the levels of alpha and beta wave activity in the test subject during the second period of time so as to determine baseline alpha and beta wave activity levels;c) recording a third EEG over a third period of time from the test subject and exposing the test subject during the third period of time to a series of pairs of sensory stimuli, wherein the pairs each consist of a visual stimulus and a corresponding auditory stimulus, and quantitating the levels of alpha and beta wave activity in the test subject during exposure of the subject to the series of pairs of sensory stimuli;d) subtracting the baseline levels of alpha and beta wave activity quantitated during step b) from the levels of alpha and beta wave activity, respectively, quantitated during step d), so as to determine corrected alpha wave activity and corrected beta wave activity in the subject;e) determining the ratio of corrected alpha wave activity to corrected beta wave activity,f) determining if the ratio of corrected alpha wave activity to corrected beta wave activity is higher in the RCH of the test subject or in the LCH of the test subject,wherein a higher alpha/beta ratio in the RCH than the LCH indicates that the test subject has an attitude compatible with employment at the predetermined organization, and wherein a higher alpha/beta ratio in the LCH than the RCH indicates that the test subject does not possesses an attitude compatible with employment at the predetermined organization. 2. The method of claim 1 wherein (1) a higher alpha/beta ratio in the RCH than LCH and an alpha/beta ratio in excess of two standard deviations of baseline indicates that the subject is content and/or calm and possesses an attitude compatible with the predetermined organization, and (2) a higher alpha/beta ratio in the RCH than LCH and an alpha/beta ratio below two standard deviations of baseline indicates that the subject is excited and/or happy and possesses an attitude compatible with the predetermined organization. 3. The method of claim 1, wherein the first, second and third EEGs are recorded using a bilateral electrode with one pole of the bilateral electrode positioned over the left cerebral hemisphere of the test subjects' brain and with another ... | - | A61B005/0484
US08150493 | B2 | 2012-04-03 | Patient infusion and imaging system | 2010-07-12 | A patient infusion and imaging system includes a magnetic resonance imaging (MRI) scan room shielded from electromagnetic interference and image acquisition equipment positioned within the MRI scan room. The system further includes an IV liquid infusion apparatus configured to be positioned at any location within the MRI scan room without introducing image-degrading RF interference and without being affected by magnetic fields produced by the magnetic source. The IV liquid infusion apparatus includes a pump disposed to receive a liquid conduit and to transfer IV liquid through the liquid conduit from a liquid IV container source, a non-magnetic ultrasonic motor, and a non-magnetic controller configured to selectively actuate the motor to drive the pump, and thereby transfer liquid from the IV container source through the liquid conduit at a controllable volumetric rate. | 1. A patient infusion and imaging system, the system comprising: a magnetic resonance imaging (MRI) scan room shielded from electromagnetic interference;image acquisition equipment positioned within the MRI scan room, the image acquisition equipment comprising a magnetic source; andan IV liquid infusion apparatus configured to be positioned at any location within the MRI scan room without introducing image degrading RF interference and without being affected by magnetic fields produced by the magnetic source, the IV liquid infusion apparatus comprising: a pump disposed to receive a liquid conduit and to transfer IV liquid through the liquid conduit from a liquid IV container source;a non-magnetic ultrasonic motor configured to provide mechanical movement to the pump in response to applied multiphasic electrical signals to provide positive displacement of the IV liquid from the IV container source; anda non-magnetic controller configured to selectively actuate the motor to drive the pump, and thereby transfer liquid from the IV container source through the liquid conduit at a controllable volumetric rate. 2. The system of claim 1, wherein the non-magnetic controller is disposed within an electrically conductive shielding structure to inhibit radiation of radio frequency interference during operation of the controller. 3. The system of claim 1, wherein the liquid infusion apparatus is housed in one RF shielded housing. 4. The system of claim 1, wherein the IV liquid infusion apparatus further comprises a display comprising non-magnetic and RF-shielding material. 5. The system of claim 1, wherein the pump comprises a peristaltic pump. 6. The system of claim 1, wherein the liquid infusion apparatus is not anchored to the ground. 7. The system of claim 1, wherein the liquid infusion apparatus does not comprise any ferrous materials. 8. The system of claim 1, wherein the pump, the motor, and the controller are housed in an RF-shielded housing. 9. A patient infusion and imaging system, the system comprising: a magnetic resonance imaging (MRI) scan room shielded from electromagnetic interference;image acquisition equipment positioned within the MRI scan room, the image acquisition equipment comprising a magnetic source; andan IV liquid infusion apparatus configured to be positioned at any location within the MRI scan room without introducing image degrading RF interference and without being affected by magnetic fields produced by the magnetic source, the IV liquid infusion apparatus comprising: a peristaltic pump disposed to receive a liquid conduit and to transfer IV liquid through the liquid conduit from a liquid IV container source;a non-magnetic ultrasonic motor configured to provide mechanical movement to the peristaltic pump in response to applied multiphasic electrical signals to provide positive displacement of the IV liquid from the IV container source; anda non-magnetic controller configured to selectively actuate the motor to drive the... | Iradimed Corporation | A61B005/055
US20110015537 | A1 | 2011-01-20 | METHOD, APPARATUS AND COMPUTER PROGRAM FOR MONITORING SPECIFIC CEREBRAL ACTIVITY | 2009-07-15 | A method, apparatus and computer program product are disclosed for monitoring specific cerebral activity of a subject. Occurrences of a predetermined event are detected in each of at least two brain wave signals acquired from a subject, wherein the predetermined event is indicative of a specific cerebral activity. Based on the detecting, an event information signal is produced for each of the at least two brain wave signals, thereby to obtain at least two event information signals, each event information signal being indicative of occurrences of the predetermined event in respective brain wave signal. Based on the at least two event information signals, an overall event information signal is generated, which is indicative of an occurrence of the predetermined event when at least one of the at least two event information signals is indicative of an occurrence of the predetermined event within a given time window, thereby to obtain a signal representing a general view of the specific cerebral activity of the subject within a measurement area covered by the at least two brain wave signals. | 1. A method for monitoring specific cerebral activity of a subject, the method comprising: detecting occurrences of a predetermined event in each of at least two brain wave signals acquired from a subject, wherein the predetermined event is indicative of a specific cerebral activity;producing, based on the detecting, an event information signal for each of the at least two brain wave signals, thereby to obtain at least two event information signals, each event information signal being indicative of occurrences of the predetermined event in respective brain wave signal; andgenerating, based on the at least two event information signals, an overall event information signal that indicates an occurrence of the predetermined event when at least one of the at least two event information signals is indicative of an occurrence of the predetermined event within a given time window, thereby to obtain a signal representing a general view of the specific cerebral activity of the subject within a measurement area covered by the at least two brain wave signals. 2. The method according to claim 1, further comprising determining an overall activity index based on the overall event information signal, wherein the overall activity index is indicative of a degree of occurrences of the predetermined event, said occurrences being indicated by the overall event information signal. 3. The method according to claim 1, wherein the producing includes producing the at least two event information signals, in which each event information signal assumes a logic true value when an occurrence of the predetermined event is detected and a logic false value otherwise. 4. The method according to claim 3, wherein the generating includes applying a logical OR operation to the at least two event information signals, thereby to obtain the overall event information signal. 5. The method according to claim 1, wherein the producing includes producing the at least two event information signals, in which each event information signal assumes credibility values indicative of credibility of the occurrences of the predetermined event. 6. The method according to claim 5, wherein the generating includes applying a MAX operation to the at least two event information signals, thereby to obtain the overall event information signal. 7. The method according to claim 5, further comprising examining whether the credibility values fulfill a predetermined criterion; andconverting the credibility values that fulfill the predetermined criterion to a predetermined value. 8. The method according to claim 1, wherein the detecting further includes detecting at least one further predetermined event in each of the at least two brain wave signals;the producing further includes producing at least two further event information signals for each of the at least one further predetermined event; andthe generating further includes generating a dedicated overall event information signal for each of the at least on... | GENERAL ELECTRIC COMPANY | A61B005/0476
US07876945 | B2 | 2011-01-25 | Method for processing slice images | 2004-11-12 | A method for processing slice images wherein information indicative of at least one quantitative measure is formed from image sets. The image sets are acquired from at least two imaging directions. At least two of the image sets relating to the imaging directions are used for forming the information. A method for reducing misalignments, a module, a system and a computer program product for implementing the method. | 1. A method for processing slice images, the method comprising: using image sets from at least two imaging directions, wherein an image set is a set of slice images in one imaging direction;automatically correcting misalignment of a first image slice with respect to a second image slice in a target area, wherein the first image slice and the second image slice belong to the same image set in one imaging direction, and wherein said correcting misalignment is carried out by using data from at least one other image set in at least one other imaging direction;defining at least one object of interest from corrected image sets; andforming information indicative of at least one quantitative measure of the at least one object of interest. 2. The method according to claim 1, further comprising: segmenting the target area from said image sets. 3. The method according to claim 1, further comprising: utilizing information about at least two imaging directions for computing a quantitative measure. 4. The method according to claim 1, wherein the slice images comprise cardiac image sets acquired from a heart. 5. The method according to claim 4, wherein the cardiac image sets are acquired from short-axis and long-axis of the heart. 6. The method according to claim 1, wherein the misalignments are corrected between different slice images separately on a same phase. 7. The method according to claim 1, wherein said information is used to compute at least one quantitative measure or said information itself is at least one quantitative measure related to at least one of structure, function, gray value or color. 8. The method according to claim 7, wherein the information related to structure comprises at least one of indicators for pose, size or shape. 9. The method according to claim 1, wherein at least one image set contains at least two slice images. 10. The method according to claim 1, wherein said slice images are produced by magnetic resonance imaging, ultrasound imaging or confocal microscope. 11. A method, comprising: selecting a slice image from a first image set in one imaging direction, wherein an image set is a set of slice images in one imaging direction; registering the selected slice image of the first image set to slice images of a second image set in another imaging direction; selecting another slice image from said first image set; registering said other selected slice image from said first image set to slice images of the second image set; and iterating said selecting and registering for slice images of said first image set in order to automatically correct misalignment of slice images of the first image set with respect to other slice images of the first image set. 12. The method according to claim 11, further comprising: optimizing a similarity measure between a slice of said first image set and said second image set by moving said slice of said first image set to optimize said similarity measure. 13. A module for processing slice imag... | Valtion Teknillinen Tutkimuskeskus | G06K009/00 A61B005/00 A61B005/05 A61B006/00
US07890162 | B2 | 2011-02-15 | Method and apparatus for sensing improvement using pressure data | 2007-04-30 | Method and apparatus for sensing improvement using pressure data. The method and apparatus may be used in an implantable medical device to confirm that an EGM event signifies a true mechanical cardiac activity and not just electrical oversensing. The mechanical activity may be used to create a mechanical marker channel in the implantable medical device. | 1. A method of confirming that an EGM deflection represents an R-wave comprising: sensing an EGM via an implantable medical device (IMD);sensing an intracardiac pressure via the IMD;determining a dP/dt of the sensed intracardiac pressure;detecting an electrical event using the EGM;opening a time window after detecting the electrical event;determining a maximum positive dP/dt of the intracardiac pressure sensed during the open time window;detecting the presence of a mechanical event responsive to the maximum positive dP/dt being greater than a first threshold; anddetermining that the electrical event includes an R-wave responsive to the presence of the mechanical event;determining a minimum negative dP/dt of the intracardiac pressure sensed during the open time window; anddetecting the presence of an oversensed T-wave or noise using the minimum negative dP/dt, the oversensed T-wave or noise being deemed present when the mechanical event is deemed absent and an absolute value of the minimum negative dP/dt is greater than a second threshold. 2. A method according to claim 1, further comprising withholding defibrillation when the EGM indicates a tachyarrhythmia and the oversensed T-wave or noise is deemed present. 3. A method according to claim 1, further comprising detecting the presence of an R-wave with no corresponding mechanical event, the R-wave without a corresponding mechanical event being deemed present when the mechanical event is deemed absent and an absolute value of the minimum dP/dt is not greater than a second threshold. 4. A method according to claim 1, further comprising recording the presence of the mechanical event by the device on a marker channel. 5. A method according to claim 1, wherein the first threshold is approximately 100 mmHg/s. 6. A method according to claim 1, wherein the time window is between approximately 400 and approximately 500 milliseconds or equal to the duration between two sequentially detected R-waves, which ever is smallest. | Medtronic, Inc. | A61B005/0452
US07871377 | B2 | 2011-01-18 | Method of evaluating human subconscious response to smell | 2006-01-20 | The present invention provides a method of evaluating the human subconscious response to smell and a method of evaluating a subject's subconscious response to an aroma stimulus. | 1. A method of comparing measurements of corona discharge of a subject prior to, during and after exposure to an external stimulus in order to evaluate the subject's subconscious response to sensory stimulation measured using the external stimulus, comprising the steps of: (a) evaluating the subject's corona discharge by measuring the corona discharge using a specialized camera to capture at least one image of the corona discharge from at least one fingertip of the subject prior to the exposure to the external stimulus, using digital software to project the at least one image from the at least one fingertip as at least one digitized image, and interpreting the at least one image of the corona discharge as at least one quantitative data point;(b) exposing the subject to the external stimulus for a predetermined period of time, during which the subject's corona discharge is measured at predetermined intervals using the specialized camera and the digital software;(c) measuring the subject's corona discharge using the specialized camera and the digital software after the exposure to the external stimulus; and(d) comparing the measurements obtained before, during and after exposure to the external stimulus to evaluate the subject's overall subconscious response to the external stimulus. 2. The method of claim 1 wherein step (a) further comprises the steps of: determining a baseline measurement of the subjects energy fields based on the corona discharge prior to exposure to the external stimulus; determining a balance standard for energy fields using statistical analysis measured from thousands of subjects wherein the balance standard depicts a healthy individual with relatively balanced chakras; and determining the baseline relative balance of the subject's energy fields. 3. The method of claim 2 further including repeating steps (a) through (d) until the desired balance against the defined standard of energy fields is achieved. 4. The method of claim 2 further including choosing different external stimuli to find the desired balance for the subject. 5. The method of claim 1, wherein the external stimulus is selected from the group consisting of aroma stimulus, visual stimulus, audio stimulus and touch stimulus. 6. The method of claim 5, wherein the external stimulus is an aroma stimulus. 7. The method of claim 1 wherein the step (c) further comprises reading the subject's corona discharge after exposure to the aroma stimulus at the predetermined intervals for a predetermined period of time. 8. The method of claim 1 wherein the specialized camera is a GDV electrophotography camera. 9. The method of claim 1 wherein the corona images are taken of each fingertip of the subject. 10. The method of claim 1 wherein the quantitative data points representing each reading of the corona discharge of the subject's fingertip are interpreted in terms of parameters selected from the group consisting of area of the image of the corona discharge, intensity of... | Aveda Corporation | A61B005/08 A61B005/05 A61B013/00 A61B005/00
US07894879 | B2 | 2011-02-22 | Device enabling an individual to self-monitor his breathing in order to assist in the control of a radiotherapy or imaging unit | 2006-04-04 | A device enabling an individual to self-monitor his breathing in order to assist in the control of a radiotherapy or imaging unit includes elements for acquiring a current metabolic volume signal, a signal-processing unit and a communication interface. The device also includes elements capable of detecting regular breathing during rest and then of forcing the communication interface to communicate a signal representing a predetermined setpoint value. | 1. Self-monitoring device enabling an individual to self-monitor his breathing in order to assist in the control of a radiotherapy or imaging unit comprising: means for acquiring at least one current metabolic volume signal, which is representative of the variation in the volume of air contained in the lungs of an individual during his breathing cycles;a signal-processing unit connected to said means for acquiring, said processing unit comprising means for accessing at least one rest value, which is representative of a metabolic volume for an individual of a breathing cycle during rest, and at least one setpoint value, which is representative of a metabolic volume for an individual of a predetermined target inhalation and/or exhalation; andat least one interface for communication with an individual, said communication interface being connected to said signal-processing unit and comprising communication means capable of communicating a signal to the individual,wherein the signal-processing unit comprises means for detecting regular breathing during rest by comparing the current metabolic volume signal with said rest value(s) and is capable of controlling said communication means in order that said means transmit a signal that is representative of said setpoint value(s) if said signal-processing unit detects regular breathing during rest,wherein the signal-processing unit further comprises means for detecting non-regular breathing and means for detecting inhalation and/or exhalation reaching the setpoint value(s), by comparing the current metabolic volume signal with said rest value(s) and with said setpoint values(s) respectively,and is capable of controlling said communication means in order that said means do not transmit a signal that is representative of said setpoint value(s) if said processing unit detects non-regular breathing and if no inhalation and/or exhalation reaching the setpoint value(s) is/are detected. 2. Self-monitoring device according to claim 1, wherein said regular breathing during rest is defined by at least three successive breathing cycles during rest, and said non-regular breathing is defined by less than three breathing cycles during rest. 3. Self-monitoring device according to claim 1, wherein the signal-processing unit is capable of controlling said communication means in order that said means continuously transmit a signal correlated with said current metabolic volume signal. 4. Self-monitoring device according to claim 1, wherein the means for acquiring said current metabolic volume signal comprise a spirometer. 5. Self-monitoring device according to claim 1, wherein said rest and setpoint values are a ventilation level when an individual is at rest at the end of his breathing cycle during rest and a volume of air contained in the lungs of an individual corresponding to said predetermined target inhalation and/or said predetermined target exhalation, respectively. 6. Self-monitoring device according to claim 1, ... | - | A61B005/05
US08622904 | B2 | 2014-01-07 | Use of heat sensitive color changing formula to detect and prevent the onset of decubitus ulcers | 2010-07-28 | A method for diagnosing the onset of necrosis, decubiti, and/or decubitus ulcers in the skin of a patient. Included is a topical product accessed in liquid form and applied to the skin. The product displays thermochromic temperature change when the product comes into contact with human skin tissue to identify body parts at risk for tissue necrosis by the color differences. | 1. A method of assessing the skin diagnosis of a patient, comprising: providing a water washable topical solution containing thermochromic material;applying said water washable topical solution to the skin of the patient;recording the color of the water washable topical solution as it appears on the skin of the patient;analyzing changes in the color of the water washable topical solution as affected by the skin of the patient after a period of predetermined time; andcorrelating locations of changes in color of the water washable topical solution to an increase temperature of the skin at the color change location and to the risk for the onset of decubiti. 2. The method of claim 1, wherein the thermochromic material is a plurality of liquid crystals. 3. The method of claim 1, wherein the thermochromic material is a leuco dye. 4. The method of claim 1, further including providing a color key indicating the correlation between the water washable topical solution color and ranges of temperature of the skin proximate to the water washable topical solution. 5. The method of claim 4, wherein the water washable topical solution is in a liquid form. 6. The method of claim 5, wherein the water washable topical solution can be applied as a spray. 7. The method of claim 6, wherein the period of predetermined time is less than 10 seconds. 8. The method of claim 7, wherein the period of predetermined time is less than 5 seconds. 9. A kit for the early assessment of decubitus ulcers, the kit containing; a water washable liquid solution including a thermochromic material and soap, the thermochromic material reacting to temperature variants to display a color change with increased temperature, the temperature variants apparent when the water washable liquid solution is in contact with human skin tissue; anda color key indicating the correlation between the color change and ranges of temperature of that skin area proximate to the water washable liquid solution exhibiting the color change; ranges in the color key identifying a risk for tissue necrosis. 10. The kit of claim 9, wherein the thermochromic material is a plurality of liquid crystals. 11. The kit of claim 9, wherein the thermochromic material is a leuco dye. 12. A method of assessing the risk for the onset of decubitus ulcers in the skin of a patient, comprising: providing a water washable topical solution having a first color and containing thermochromic material, the water washable topical solution absorbable through the skin of the patient;applying the water washable topical solution to the skin of the patient;recording the color of the water washable topical solution as adjusted by variances in the temperature of the skin;analyzing changes in the color of the water washable topical solution as affected by the skin of the patient after a period of predetermined time;providing a color key indicating the correlation between the water washable topical solution color and ranges of temperature of ... | - | A61B005/00
US08559589 | B2 | 2013-10-15 | Waveguide and computed tomography system with a waveguide | 2009-03-12 | The invention relates to a waveguide (24) for focusing electromagnetic energy on an area of interest, said waveguide (24) comprising outer walls forming a channel with a width (w) and a height (h) for the propagation of the electromagnetic waves. In order to be able to focus the EM waves on an area of interest, i.e. the chest of a patient (20), so that influences which other moving objects in the environment have on the waves, are shielded away, the channel is provided with a first opening and a second opening wherein said walls are non-metal hollow walls filled with an electrically conductive liquid. The waveguide (24) can be used for example with a computed tomography system (12) CT with a rotating gantry (16) and a patient table (18). The waveguide (24) will not produce artefacts in the CT image by scattering the X-ray, like a metal wall would do it. Nevertheless the conductive liquid makes the construction work as a waveguide for EM waves. | 1. An apparatus comprising a radar sensor emitting electromagnetic waves said radar sensor including an antenna, a control unit and at least one waveguide for the focusing of the electromagnetic waves on an area of interest of a subject to be examined, wherein the waveguide includes outer walls forming a channel with a width and a height for propagation of the electromagnetic waves, the channel having a first opening and a second opening, wherein the walls are non-metal hollow walls filled with an electrically conductive liquid or gel, and the at least one waveguide having the first opening oriented towards the radar sensor. 2. The waveguide according to claim 1, wherein the width- and the height of the channel and the channel openings are tailored to the frequency of the electromagnetic waves that are to be transported through the channel. 3. The waveguide according to claim 2, wherein the width is basically equal to two times the height. 4. The apparatus according to claim 1, wherein the radar sensor is a Doppler radar sensor. 5. The apparatus according to claim 1, wherein a distance piece is located between the radar antenna and the subject to be examined, whereby said distance piece is formed by said at least one waveguide. 6. The apparatus according to claim 1, wherein a support device is provided upon which the subject to be examined is placed temporarily. 7. The apparatus according to claim 6, wherein the subject to be examined is a patient and the support device is a patient table to receive the patient. 8. The apparatus according to claim 7, wherein the at least one waveguide is located within the patient table. 9. The apparatus according to claim 8, wherein the radar sensor is installed in the patient table. 10. The apparatus according to claim 7, wherein the control unit is installed in the patient table. 11. The apparatus according to claim 1, wherein multiple waveguides form an array of waveguides. 12. The apparatus according to claim 11, wherein multiple radar sensors are used feeding the waveguide array. 13. A computed tomography system with a rotating gantry, said computed tomography system comprising: a system control, and an apparatus, comprising: a radar sensor emitting electromagnetic waves, the radar sensor including an antenna, a control unit and at least one waveguide for focusing electromagnetic energy on an area of interest, the waveguide comprising outer walls forming a channel with a width and a height for propagation of the electromagnetic waves, the channel having a first opening and a second opening, and the at least one waveguide having the first opening oriented towards the radar sensor, wherein the walls are non-metal hollow walls filled with an electrically conductive liquid or gel, wherein the apparatus detects motion in a patient due to respiration or due to blood pumping of the heart and provides a gating signal to a system controlling the gating signal triggering an acquisition of CT data. 14. A ... | Koninklijke Philips N.V. | A61B006/00
US08527035 | B2 | 2013-09-03 | System, optode and cap for near-infrared diffuse-optical function neuroimaging | 2009-04-23 | An optode for a functional infrared diffuse optical neuroimaging system uses a GRIN lens and prism for coupling to an optical fiber at a right-angle to the lens. The lens is inserted into a grommet for attachment to an elastomeric cap worn by a subject. In an embodiment, the cap also has an array of electroencephalographic electrodes, with optodes arranged such that an optical path exists between a transmitter and a receiver optode beneath each electroencephalographic electrode. In an embodiment, quadfurcated optical cabling is used to permit illumination of each of multiple optodes with wavelengths provided from each of multiple lasers, and permitting each laser to illuminated multiple optodes. | 1. An optode for interfacing light to biological tissue comprising: at least one flexible light guide;a prism optically coupled to the flexible light guide;a graded incidence (GRIN) refractive lens having a circumferential surface, a convex surface and a flat surface, the flat surface optically coupled to the prism; anda body for holding the optical fiber, prism, and GRIN lens, the body configured to maintain the fiber, prism, and GRIN lens in fixed relationship having a right angle between an axis of the fiber and an axis of the GRIN lens;wherein the convex surface of the graded incidence refractive lens is capable of interfacing light to biological tissue. 2. The optode of claim 1 further comprising an elastomeric grommet having a hole, the circumferential surface of the GRIN lens secured in the hole of the elastomeric grommet, the grommet and body retaining the optode in a hole of an optode-positioning device. 3. The optode of claim 1 wherein the body has a tubular extension, the extension having an inner surface in contact with the circumferential surface of the GRIN lens and an outer surface in contact with an elastomeric grommet, the grommet and body retaining the optode in a hole of an optode-positioning device. 4. The optode of claim 1 wherein the GRIN lens and flexible light guide are optimized for transparency in the infrared. 5. The optode of claim 1 wherein the flexible light guide comprises at least one optical fiber. 6. The optode of claim 1 wherein the flexible light guide comprises at least two optical fibers, and wherein the flexible light guide is at least bifurcated into a first and a second subguide. 7. An optode for interfacing light to biological tissue comprising: at least one flexible light guide;a prism optically coupled to the flexible light guide;a GRIN lens having at least a convex surface and a second surface, the second surface optically coupled to the prism; anda body for holding the flexible light guide, prism, and GRIN lens, the body configured to maintain the guide, prism, and GRIN lens in fixed alignment having right angle between an axis of the fiber and an axis of the GRIN lens;wherein the convex surface of the lens is capable of interfacing light to biological tissue; andwherein the flexible light guide comprises a plurality of optical fibers and divides into a plurality of subguides, and wherein at least a first optical fiber of the plurality of fibers extends from the prism to an end of a first subguide of the subguides, and wherein at least a second optical fiber of the plurality of fibers extends from the prism to an end of a second subguide of the subguides. 8. The optode of claim 7 wherein the body has a tubular extension, the extension having an inner surface in contact with a circumferential surface of the lens and an outer surface for contacting an elastomeric grommet. 9. The optode of claim 7 wherein the GRIN lens and flexible light guide are optimized for transparency in the infrared. | The Trustees of Dartmouth College | A61B006/00
US08594773 | B2 | 2013-11-26 | Denoising and artifact rejection for cardiac signal in a sensis system | 2010-09-28 | A system denoises and rejects artifacts from cardiac signals, by accepting a cardiac signal from a patient, processing the cardiac signal from the patient using a frequency band width controllable choke to separate the cardiac signal into predefined frequencies, filtering each of the predefined frequencies to remove dynamic common noise, joining each of the predefined frequencies into a cardiac signal without the dynamic common noise, and providing feedback control of the filtering of each of the predefined frequencies. | 1. A method for denoising and rejecting artifacts from cardiac signals, comprising: accepting a cardiac signal from a patient;separating the cardiac signal into signals in predefined frequency bands in the time domain using a frequency band width controllable choke;filtering each of the signals in the predefined frequency bands to remove dynamic noise;joining filtered signals of each of the predefined frequency bands into a cardiac signal without the dynamic noise; andproviding a feedback control signal to control the filtering of each of the predefined frequency bands. 2. The method according to claim 1, wherein the filtering of each of the predefined frequency bands to remove dynamic noise is accomplished by identification of interference signals from medical instruments and elimination of noise related to the medical instruments. 3. The method according to claim 1, wherein the dynamic noise removed is a non-linear signal. 4. The method according to claim 1, wherein the signals from the patient are controlled through a computer programmable selection that decreases common mode noise. 5. The method according to claim 1, wherein the filtering of each of the predefined frequencies is through empirical mode decomposition processing. 6. A method for denoising and rejecting artifacts from cardiac signals, comprising: accepting a cardiac signal from a patient;separating the cardiac signal from the patient into signals in predefined frequency bands in the time domain using a frequency band width controllable choke;filtering each of the signals in the predefined frequency bands to remove dynamic noise; andjoining filtered signals of each of the predefined frequency bands into a cardiac signal without the dynamic noise. 7. The method according to claim 6, wherein the filtering of each of the predefined frequencies is through empirical mode decomposition processing. 8. The method according to claim 6, wherein the filtering of each of the predefined frequency bands to remove dynamic noise is accomplished by identification of interference signals from medical instruments and elimination of noise related to the medical instruments. 9. The method according to claim 6, wherein the dynamic noise removed is a non-linear signal. 10. A program storage device readable by machine, tangibly embodying a program of instructions executable by the machine to perform method steps for denoising and rejecting artifacts from cardiac signals, comprising: accepting a cardiac signal from a patient;separating the cardiac signal from the patient into signals in predefined frequency bands in the time domain using a frequency band width controllable choke;filtering each of the signals in the predefined frequency bands to remove dynamic noise;joining filtered signals of each of the predefined frequency bands into a cardiac signal without the dynamic noise; andproviding a feedback control signal to control the filtering of each of the predefined frequency bands. 11. The dev... | Siemens Medical Solutions USA, Inc. | A61B005/0428
US20110040196 | A1 | 2011-02-17 | BLOOD PRESSURE MONITOR | 2009-11-17 | A blood pressure monitor for measuring a blood pressure of a human body includes a main body, a connecting mechanism, and a sensing unit. The main body surrounds the human body and has a first end portion and a second end portion. The connecting mechanism connects the first end portion with the second end portion at a first position or a second position. When the first end portion and the second end portion are connected at the first position, the human body surrounded thereby has a first circumference, and when the first end portion and the second end portion are connected at the second position, the human body surrounded thereby has a second circumference. The sensing unit, disposed on the body, detects the first position or the second position to determine the first circumference or the second circumference respectively. | 1. A blood pressure monitor, measuring a blood pressure of a human body, comprising: a main body, for surrounding the human body, having a first end portion and a second end portion;a connecting mechanism, for connecting the first end portion with the second end portion at a first position or a second position at a first position or at a second position, when the first end portion and the second end portion are connected at the first position, the human body surrounded thereby has a first circumference, and when the first end portion and the second end portion are connected at the second position, the human body surrounded thereby has a second circumference; anda sensing unit, disposed on the body and detecting the first position or the second position to determine the first circumference or the second circumference. 2. The blood pressure monitor as claimed in claim 1, wherein the main body comprises: a first body having a first connecting end portion; anda second body having a second connecting end portion pivotally connected to the first connecting end portion. 3. The blood pressure monitor as claimed in claim 2, wherein the first end portion is opposite to the first connecting end portion, and the second end portion is opposite to the second connecting end portion. 4. The blood pressure monitor as claimed in claim 1, wherein the connecting mechanism comprises: a first connecting portion disposed on the first end portion and having a first engaging portion and a second engaging portion; anda second connecting portion disposed on the second end portion and having a corresponding connecting portion, the corresponding connecting portion selectively engages with the first engaging portion or the second engaging portion;wherein when the corresponding engaging portion engages with the first engaging portion, the first end portion and the second end portion are connected at the first position, and when the corresponding engaging portion engages with the second engaging portion, the first end portion and the second end portion are connecting at the second position. 5. The blood pressure monitor as claimed in claim 4, wherein the first engaging portion and the second engaging portion are notches, and the corresponding engaging portion is movably disposed on the second end portion. 6. The blood pressure monitor as claimed in claim 4, wherein the second connecting portion further comprises an opening, and after the first connecting portion penetrates the opening, the corresponding engaging portion engages with the first engaging portion or the second engaging portion. 7. The blood pressure monitor as claimed in claim 5, wherein the corresponding engaging portion is disposed at the edge of the opening. 8. The blood pressure monitor as claimed in claim 4, wherein the first connecting portion further comprises: a first sensing point, corresponding to the sensing unit when the first end portion and the second end portion being connected at the first po... | QUANTA COMPUTER INC. | A61B005/022
US07883480 | B2 | 2011-02-08 | Motion monitoring apparatus | 2007-11-13 | A motion monitoring apparatus and method for monitoring a patient under medical care. A sensor arrangement is provided in the form of a pad which the patient lies on. The sensor arrangement provides a signal which can be monitored to observe motion of the patient and provide an alarm should the motion meet certain predetermined conditions. The invention is particularly applicable for monitoring patients under sedation, recovering from anaesthesia, or in intensive care. The device is particularly useful for veterinary patients. | 1. A method of monitoring a patient under medical care, the method comprising the steps of: providing a sensor arrangement which is arranged to detect motion of a head, limbs and trunk of the patient, and produce a motion signal in response to the motion of a patient, andanalysing the motion signal to determine whether or not the patient is displaying signs of painfulness,wherein if the analysis of the motion signal determines that the motion includes rhythmic, repetitive motion of the head, the limbs or the trunk, a determination is made that the patient is displaying signs of painfulness. 2. The method in accordance with claim 1, wherein the step of detecting motion of the patient includes detecting respiratory motion of the patient and producing a motion signal in response to the respiratory motion. 3. The method in accordance with claim 2, further comprising the step of monitoring the respiratory motion and the bodily motion of the head, the limbs or the trunk of the patient by the same sensor arrangement. 4. The method in accordance with claim 2, further comprising the step of monitoring the respiratory motion and the bodily motion of the head, the limbs or the trunk of the patient by separate sensor arrangements. 5. The method in accordance with claim 2, further comprising the step of generating a single motion signal, as the motion signal, from a combination of the bodily motion of the head, the limbs or the trunk of the patient and the respiratory motion. 6. The method in accordance with claim 2, further comprising the step of generating the motion signal to include a first motion signal, generated from the bodily motion of the head, the limbs or the trunk of the patient, and a second motion signal, generated from the respiratory motion. 7. The method in accordance with claim 1, further comprising the step of using a non-human animal patient as the patient. 8. The method in accordance with claim 1, further comprising the step of using a human patient as the patient. 9. The method in accordance with claim 1, comprising the further step of controlling operation of a peripheral device in response to the analysis. 10. A device for monitoring a patient under medical care, the device comprising a sensor arrangement which is arranged to detect motion of the head, limbs and trunk of the patient and produce a motion signal in response to the motion, and a control means which is arranged to process the motion signal and analyse the motion signal, to determine whether or not the patient is displaying signs of painfulness, the control means is arranged to determine that the patient is displaying signs of painfulness if the motion signal is determined to include rhythmic repetitive motion. 11. A device for monitoring a patient under medical care, the device comprising a sensor arrangement which is arranged to detect motion of the head, limbs and trunk of the patient and produce a motion signal in response to the motion, and a control means w... | - | A61B005/08 A61B005/103 A61B005/117
US08221325 | B2 | 2012-07-17 | Detection of hypovolemia using implantable medical device | 2010-09-14 | An implantable medical device receives a physiological signal indicative of circulatory blood volume and detects hypovolemia from that physiological signal. In one embodiment, an implantable pulmonary artery pressure (PAP) senses a PAP signal, and the implantable medical device detects hypovolemia from the PAP signal. | 1. A system for monitoring a body having a pulmonary artery and a cardiac wall, the system comprising: an hypovolemia detection module configured to receive a pulmonary artery pressure (PAP) signal via wireless telemetry, the hypovolemia detection module including a signal processor configured to process the PAP signal and a hypovolemia detector configured to produce a diastolic PAP (DPAP) using the processed PAP signal and produce at least one of a hypovolemia detection signal and a hypovolemia prediction signal using at least the DPAP, the hypovolemia detection signal indicative of a detection of hypovolemia, the hypovolemia prediction signal indicative of an elevated risk of hypovolemia. 2. The system of claim 1, wherein the hypovolemia detector is configured to produce a systolic PAP (SPAP) using the processed PAP signal and produce the at least one of the hypovolemia detection signal and the hypovolemia prediction signal using the DPAP and the SPAP. 3. The system of claim 2, wherein the hypovolemia detector is configured to produce a mean PAP (MPAP) using the DPAP and SPAP and produce the at least one of the hypovolemia detection signal and the hypovolemia prediction signal using the MPAP. 4. The system of claim 1, comprising an implantable pulmonary artery pressure (PAP) sensor configured to be placed in the body to sense the PAP signal. 5. The system of claim 4, comprising an anchoring structure coupled to the implantable PAP sensor and configured to anchor the implantable PAP sensor within the pulmonary artery. 6. The system of claim 4, comprising an anchoring structure coupled to the implantable PAP sensor and configured to anchor the implantable PAP sensor to the cardiac wall. 7. The system of claim 1, comprising an implantable medical device including the hypovolemia detection module, and the hypovolemia detection module is configured to be communicatively coupled to the implantable PAP sensor via an ultrasonic telemetry link. 8. The system of claim 7, wherein the implantable medical device further comprises an edema detection module adapted to detect an edema and produce a detection signal indicative of a detection of the edema. 9. The system of claim 8, wherein the hypovolemia detector is configured to produce at least the hypovolemia detection signal, and the implantable medical device further comprises: a therapy delivery device adapted to deliver one or more therapies; anda therapy controller adapted to control the delivery of the one or more therapies using at least the hypovolemia detection signal and the edema detection signal. 10. A method for monitoring a patient, the method comprising: receiving a pulmonary artery pressure (PAP) signal from an implantable PAP sensor via wireless telemetry;producing a diastolic PAP (DPAP) using the PAP signal using a processor; andproducing at least one of a hypovolemia detection signal and a hypovolemia prediction signal using at least the DPAP using the processor, the hypovolemia d... | Cardiac Pacemakers, Inc. | A61B005/484
US07873401 | B2 | 2011-01-18 | System and method for a magnetic catheter tip | 2006-01-13 | A system whereby a magnetic tip attached to a surgical tool is detected, displayed and influenced positionally so as to allow diagnostic and therapeutic procedures to be performed rapidly, accurately, simply, and intuitively is described. The tools that can be so equipped include catheters, guidewires, and secondary tools such as lasers and balloons, in addition biopsy needles, endoscopy probes, and similar devices. The tip position and orientation information and the dynamic body part position information are also utilized to provide a display that allows three dimensional viewing of the magnetic tip position and orientation relative to the body part. | 1. A catheter tip assembly, comprising; a flexible portion proximate to a distal end of a catheter;a rigid portion disposed at said distal end of said catheter and adjacent to said a distal end of said flexible portion;a permanent magnet responsive to an applied magnetic field to move said catheter tip assembly, said permanent magnet disposed in said rigid portion; anda plurality of piezoelectric rings disposed in said relatively rigid portion, said piezoelectric rings responsive to an applied ultrasonic signal, said piezoelectric rings producing data to measure a rotation of a north pole of said permanent magnet, a first piezoelectric ring of said plurality of piezoelectric rings disposed proximate to a distal end of said permanent magnet and a second piezoelectric ring of said plurality of piezoelectric rings disposed proximate to a proximal end of said permanent magnet. 2. The assembly of claim 1 further comprising a computer that determines an angle of said rotation. 3. The assembly of claim 2 wherein said computer further determines said rotation relative to a stereotactic framing. 4. The assembly of claim 1 wherein said distal end comprises one or more magnetic field sensors and one or more temperature sensors. 5. The assembly of claim 4 wherein said temperature sensors measure a temperature of said one or more magnetic field sensors. 6. The assembly of claim 1 further comprising a control system configured to calculate respective torque and associated current for a magnetic source to configure a magnetic field to move said distal end to a desired location. 7. The assembly of claim 1 further comprising a display of a visual representation of said distal end in substantially real time as said distal end moves through a body. 8. The assembly of claim 1 further comprising controlling one or more electromagnets to produce an external magnetic field. 9. The assembly of claim 2, wherein said computer determines a current position of said distal end in comparison to a desired location. 10. The assembly of claim 9, wherein said computer computes a position error of said distal end. 11. The assembly of claim 10 further comprising a tactile feedback system that is configured to produce a tactile feed back if said position error exceeds a predetermined amount along at least one axis. | Magnetecs, Inc. | A61B005/05
US20110015539 | A1 | 2011-01-20 | METHODS FOR PHYSIOLOGICAL MONITORING, TRAINING, EXERCISE AND REGULATION | 2010-01-08 | A computer assisted method comprising: measuring activity of one or more internal voxels of a brain; employing computer executable logic that takes the measured brain activity and determines one or more members of the group consisting of: a) what next stimulus to communicate to the subject, b) what next behavior to instruct the subject to perform, c) when a subject is to be exposed to a next stimulus, d) when the subject is to perform a next behavior, e) one or more activity metrics computed from the measured activity, f) a spatial pattern computed from the measured activity, g) a location of a region of interest computed from the measured activity, h) performance targets that a subject is to achieve computed from the measured activity, i) a performance measure of a subject's success computed from the measured activity, j) a subject's position relative to an activity measurement instrument; and communicating information based on the determinations to the subject in substantially real time relative to when the activity is measured. | 1. A computer assisted method comprising: measuring activity of one or more internal voxels of a brain;employing computer executable logic that takes the measured brain activity and determines one or more members of the group consisting of: a) what next stimulus to communicate to the subject, b) what next behavior to instruct the subject to perform, c) when a subject is to be exposed to a next stimulus, d) when the subject is to perform a next behavior, e) one or more activity metrics computed from the measured activity, f) a spatial pattern computed from the measured activity, g) a location of a region of interest computed from the measured activity, h) performance targets that a subject is to achieve computed from the measured activity, i) a performance measure of a subject's success computed from the measured activity, j) a subject's position relative to an activity measurement instrument; andcommunicating information based on the determinations to the subject in substantially real time relative to when the activity is measured. 2. A method according to claim 1 wherein measuring brain activity is performed by fMRI. 3. A method according to claim 1 wherein the determinations are made in less than 10 seconds relative to when the activity is measured. 4. A method according to claim 1 wherein the determinations are made in less than 1 second relative to when the activity is measured. 5. A method according to claim 1 wherein the determinations are made in less than 0.5 second relative to when the activity is measured. 6. A method according to claim 1 wherein the information is determined while the instrument used for measurement remains positioned about the subject. 7. A method according to claim 1 wherein the activity measurements are made using an apparatus capable of taking measurements from one or more internal voxels without substantial contamination of the measurements by activity from regions intervening between the internal voxels being measured and where the measurement apparatus collects the data. 8. A method according to claim 1 wherein measurements are made from at least 100 separate internal voxels, and these measurements are made at a rate of at least once every five seconds. 9. A method according to claim 1 wherein measurements are made from a set of separate internal voxels corresponding to a scan volume including the entire brain. 10. A method according to claim 1 wherein the size of the internal voxels have a total three dimensional volume of 5×5×5 cm or less. 11. A method according to claim 1 wherein the size of the internal voxels have a total three dimensional volume of 1×1×1 cm or less. 12. A method according to claim 1 wherein the method further comprises selecting one or more of the internal voxels to correspond to a region of interest for the subject and using the selected internal voxels of the region of interest to make the one or more determinations. 13. A method according to claim 12 wherein the region of in... | - | A61B005/04
US20110015530 | A1 | 2011-01-20 | LIVING BODY TISSUE THREE-DIMENSIONAL MODEL AND PRODUCTION METHOD THEREFOR | 2010-09-27 | An internal tissue including a lesion region in the human body is modeled as a three-dimensional model. By reconstructing thickness or flexibility of a lumen wall portion including the lesion region and making it possible to confirm a motion of the lumen wall or a flow of fluid in the inside of the lumen wall, a state of the lesion region in the lumen can be confirmed clearly by visual inspection or the like. As a result, the diagnosis in the lumen can be made easier. | 1. A living body tissue three-dimensional model comprising a three-dimensional model of a lumen wall portion of a lumen portion of an actual living body so the three-dimensional model is a three-dimensional model of the lumen portion of the actual living body, the three-dimensional model being configured to reconstruct a thickness of the lumen wall portion of the lumen portion including a reconstructed lesion region of the actual living body. 2. The living body tissue three-dimensional model according to claim 1, wherein the three-dimensional model includes a penetrating portion for receiving an operation instrument into the reconstructed lesion region, the penetrating portion being formed on a side face of the lumen portion. 3. The living body tissue three-dimensional model according to claim 1, wherein the three-dimensional model is configured from at least first and second parts connected together, the first and second parts connected together each including a connecting portion, the connecting portion of the first part possessing a lumen and the connecting portion of the second part possessing a lumen, the connecting portion of the first part at least partially overlapping the connecting portion of the second part with no offset between an inside surface of the lumen of the first connecting part and an inside surface of the lumen of the second. 4. The living body tissue three-dimensional model according to claim 1, wherein the lumen portion is a blood vessel. 5. The living body tissue three-dimensional model according to claim 1, wherein the lesion region is a blood vessel. 6. A method for producing a living body tissue three-dimensional model comprising: obtaining three-dimensional data of a lesion region of a living body tissue;producing, based on the three-dimensional data of the lesion region of the living body tissue, lesion region tomographic data in which a thickness of a lumen wall portion of a lumen portion including the lesion region is reconstructed; andlamination shaping the living body tissue three-dimensional model using the lesion region tomographic data. 7. The method according to claim 6, further comprising coupling tomographic data three-dimensional data to the lesion region to form a penetrating portion of an operation instrument into a lesion region which has been reconstructed on a side face or an end face of the lumen portion. 8. The method according to claim 6, further comprising coupling three-dimensional data of an overlapping coupling portion with the lesion region tomographic data so that no offset is provided on a lumen inside face region in a state in which a coupling portion of the lumen for coupling the adjacent parts to each other is coupled. 9. A living body tissue three-dimensional model comprising a three-dimensional model of an actual living body lumen produced using tomographic image data of the actual living body lumen so the three-dimensional living body lumen model is a three-dimensional model o... | TERUMO CORPORATION | A61B005/02 G06G007/60 G06F019/00
US07877129 | B2 | 2011-01-25 | Magnetic resonance imaging apparatus and RF reception coil apparatus | 2002-09-13 | An RF receiving means, which comprise respective at least three RF receiving coils (203, 201 (or 202), 204), (205, 202 (or 201), 206), (207, 201 (or 202), 208) disposed in at least two orthogonal directions (x, y, z), are set to a patient 209, any one direction (x) of the two directions is set as a phase encoding direction, magnetic resonance signals are measured by executing a pulse sequence while thinning phase encoding steps, and the aliasing of an image is eliminated by an arithmetic operation executed using the magnetic resonance signals measured by at least the three RF receiving coils (205, 202 (or 201), 206) disposed in the thus set phase encoding direction and using the sensitivity distributions of the respective RF receiving coils. | 1. A magnetic resonance imaging apparatus comprising: magnetic field generation means for applying gradient magnetic fields and high frequency magnetic fields to a patient placed in static magnetic fields in a predetermined pulse sequence while thinning phase encoding steps, a receiving coil group comprising a plurality of RF receiving coils for receiving the nuclear magnetic resonance signals generated from the patient, and image reconstruction means for reconstructing an image by processing the nuclear magnetic resonance signals, whereinthe receiving coil group comprises at least one main RF receiving coil having a high sensitivity region in a diagnosing portion of the patient and at least two sets of auxiliary RF receiving coil groups each set including at least two auxiliary RF receiving coils disposed in confrontation with each other in either of at least two orthogonal directions, and the high sensitivity region of the respective auxiliary RF receiving coils is formed narrower than that of the main RF receiving coil; andthe image reconstruction means eliminates the aliasing of an image by an arithmetic operation executed based on the magnetic resonance signals received by the main RF receiving coil and at least one set of the respective auxiliary RF receiving coils disposed in confrontation with each other in at least a phase encoding direction and based on the sensitivity distributions of these RF receiving coils. 2. A magnetic resonance imaging apparatus according to claim 1, characterized in that three sets of the auxiliary RF receiving coil groups are disposed in three orthogonal directions. 3. A magnetic resonance imaging apparatus according to claim 1, characterized in that at least one of the RF receiving coils is a solenoid coil. 4. A magnetic resonance imagine apparatus according to claim 1, characterized in that at least one of the RF receiving coils is a saddle coil. 5. A magnetic resonance imaging apparatus according to claim 1, characterized in that at least one of the auxiliary RF receiving coils is a solenoid coil. 6. A magnetic resonance imaging apparatus according to claim 1, characterized in that at least one of the auxiliary RF receiving coils is a 8-shaped coil. 7. A magnetic resonance imaging apparatus comprising: magnetic field generation means for applying gradient magnetic fields and high frequency magnetic fields to a patient placed in static magnetic fields in a predetermined pulse sequence while thinning phase encoding steps, a receiving coil group comprising a plurality of RE receiving coils for receiving the nuclear magnetic resonance signals generated from the patient, and image reconstruction means for reconstructing an image by processing the nuclear magnetic resonance signals, whereinthe receiving coil group comprises at least three RF receiving coils disposed in each of at least two orthogonal directions, and any one of the two directions is set as a phase encoding direction; andthe image reconstructi... | Hitachi Medical Corporation | A61B005/00
US07874984 | B2 | 2011-01-25 | Methods for delivering a drug to a patient while preventing the exposure of a foetus or other contraindicated individual to the drug | 2005-04-12 | Novel methods for delivering a drug to a patient while preventing the exposure of a foetus or other contraindicated individual to the drug. Embodiments are provided in which the involved prescribers, pharmacies and patients are registered in one or more computer databases. Embodiments are also provided in which registered patients receive counseling information concerning the risks attendant to foetal exposure to the drug. Male patients and female patients who are not pregnant may, in certain circumstances, receive the drug. | 1. A method of treating a male patient, suffering from erythema nodosum leprosum, with thalidomide, said method comprising (a) permitting thalidomide to be distributed by a pharmacy only after the pharmacy has received a prescription and (b) administering said thalidomide to the patient, wherein the generation of the prescription comprises the following steps: (1) via a computer readable storage medium, registering a prescriber and the pharmacy with a distributor of thalidomide;(2) determining whether the patient is able to understand and carry out instructions;(3) upon determination that the patient is able to carry out the instructions, providing verbal and written warnings of the hazard of taking thalidomide and exposing fetus to the drug;(4) further providing verbal and written warnings of the risk of possible contraception failure and of the need to use barrier contraception when having sexual intercourse with women of child bearing potential;(5) obtaining acknowledgement of said warnings from the patient;(6) via a computer readable storage medium, registering the patient with the distributor; and(7) upon obtaining the acknowledgement and registrations, generating via a computer readable storage medium the prescription to be received by the pharmacy before thalidomide is distributed. 2. The method of claim 1, wherein the acknowledgement requires the patient's acknowledgement of one or more of the following: (a) the understanding that thalidomide must not be taken if unprotected sex cannot be avoided;(b) the understanding of potential birth defects;(c) that the patient has been advised of the need for barrier contraception by the prescriber;(d) the obligation to inform the prescriber if the patient's sexual partner is suspected of becoming or being pregnant;(e) that thalidomide is solely for the use of the patient himself and must not be shared with any other person;(f) that the patient has read the information brochure or viewed the information film on thalidomide;(g) that the semen or blood must not be donated during the thalidomide treatment;(h) that all of the patient's inquiries regarding thalidomide treatment have been answered by the prescribing physician; or(i) the patient's understanding that participation in a survey and patient registry is required during the thalidomide treatment. 3. The method of claim 1 further comprising obtaining a written authorization by the prescriber. 4. The method of claim 1, wherein the acknowledgement is a written informed consent and wherein the pharmacy fills the prescription only upon patient's presentation of the written informed consent to the pharmacy. | Celgene Corporation | A61B005/00
US07894889 | B2 | 2011-02-22 | ECG signal power vector detection of ischemia or infarction | 2009-06-12 | A method comprising sensing at least one cardiac signal representative of cardiac activity of a subject using an implantable medical device (IMD), calculating, from the cardiac signal, a first dominant vector corresponding to a direction and magnitude of maximum signal power of an ST-T first segment of a cardiac cycle and a second dominant vector corresponding to a direction and magnitude of maximum signal power of a P-QRS second segment of a cardiac cycle, measuring a change in the first dominant vector, measuring a change in the second dominant vector, and subtracting the change in the second dominant vector from the measured change in the first dominant vector to form a difference. | 1. A method comprising: sensing at least one cardiac signal representative of cardiac activity of a subject using an implantable medical device (IMD);calculating, from the cardiac signal, a first dominant vector corresponding to a direction and magnitude of maximum signal power of an ST-T first segment of a cardiac cycle and a second dominant vector corresponding to a direction and magnitude of maximum signal power of a P-QRS second segment of a cardiac cycle, wherein the P-QRS segment includes the P-wave and extends to and includes the S-wave;measuring a change in the first dominant vector; measuring a change in the second dominant vector;subtracting the change in the second dominant vector from the measured change in the first dominant vector to form a difference; anddeeming an ischemic event to have occurred at least in part from whether the difference exceeds a specified threshold value. 2. The method of claim 1, further including activating an alarm indicative of the deemed ischemic event. 3. The method of claim 1, wherein deeming an ischemic event to have occurred further includes deeming an ischemic event if the difference exceeds the specified threshold value in X of Y consecutive cardiac cycles, wherein X and Y are integers and X is less than or equal to Y. 4. The method of claim 1, comprising establishing a baseline vector for at least one of the first and second dominant vectors and measuring a change from the baseline vector. 5. The method of claim 4, wherein establishing a baseline vector includes calculating a central tendency for at least one of the first and second dominant vectors. 6. The method of claim 4, comprising reducing noise related to posture by subtracting the measured change in the second dominant vector from the measured change in the first dominant vector to form a difference. 7. The method of claim 4, wherein measuring the change from the baseline vector includes: subtracting a first baseline vector from a later-measured first dominant vector to obtain a first change representative of a change due to posture during the first segment;subtracting a second baseline vector from a later-measured second dominant vector to obtain a second change representative of a change due to posture and ischemia during the second segment; andsubtracting the first change from the second change to obtain a third change representative of a change due to ischemia. 8. The method of claim 4, comprising measuring a trend in at least one of the first and second dominant vectors. 9. The method of claim 8, including storing an ischemia indication using the measured trend. 10. The method of claim 1, wherein measuring a change in the first dominant vector and measuring a change in the second dominant vector comprises measuring at least one of a change in angle, a change in magnitude, a variance of the angle or the magnitude, a power spectral density, a rate of change of the angle, a rate of change of the magnitude, and a rate of change of... | Cardiac Pacemakers, Inc. | A61B005/0452
US08140151 | B2 | 2012-03-20 | Implantable seizure monitor | 2010-04-29 | An implantable seizure monitor can include at least one sensing electrode and an electronics module configured to detect, record and/or log neurological events. For example, the electronics module can be configured to detect brainwaves indicative of seizures, such as, for example, epileptic seizures, and to create a log indicating when such seizures occur. The implantable seizure monitor can include a cushioning member made of a soft material and configured to be implantable between the epidermis and cranium of a patient. | 1. A method for detecting neurological events comprising: sensing brainwaves of a patient;generating a signal indicative of the sensed brainwaves;amplifying the signal in an amplifier, the amplifier having a gain adjustable to predetermine an amplitude of the signal that will place the amplifier in saturation; anddetermining, by a processor, if a neurological event has occurred based on occurrences of saturation of the amplifier. 2. The method according to claim 1, wherein determining, by the processor, if the neurological event has occurred includes determining if the neurological event has occurred based on the frequency at which the amplifier saturates. 3. The method according to claim 1, wherein determining, by the processor, if the neurological event has occurred includes determining if the neurological event has occurred based on the number of times the amplifier saturates in a predetermined number of time windows. 4. The method according to claim 1, wherein determining, by the processor, if the neurological event has occurred includes: determining, for each time window in a plurality of contiguous time windows, whether the amplifier is saturated; anddetermining whether a ratio of a number of time windows in which the amplifier is saturated and the number of contiguous time windows exceeds a predetermined ratio. 5. A method for calibrating an implantable device including an adjustable gain amplifier, the method comprising: setting an initial gain of the amplifier;sensing brainwaves of a patient using the implantable device;generating a signal indicative of the sensed brainwaves using the implantable device;amplifying the signal in an amplifier, the amplifier having a gain adjustable to predetermine an amplitude of the signal that will place the amplifier in saturation;determining, by a processor, occurrences of saturation of the amplifier during normal brain activity;determining, by the processor, whether the occurrences of saturation of the amplifier during normal brain activity exceed a threshold;adjusting the gain of the amplifier if the occurrences of saturation of the amplifier during normal brain activity exceed the threshold;repeating determining, by the processor, occurrences of saturation of the amplifier during normal brain activity;determining, by the processor, whether the occurrences of saturation of the amplifier during normal brain activity exceed the threshold, and adjusting the gain, until the occurrences of saturation of the amplifier during normal brain activity do not exceed the threshold; andif the occurrences of saturation of the amplifier during normal brain activity do not exceed the threshold, determining, by the processor, if a neurological event has occurred based on occurrences of saturation of the amplifier. 6. The method of claim 5, wherein determining, by the processor, whether the occurrences of saturation of the amplifier during normal brain activity exceed a threshold includes determining whether a num... | NeuroPace, Inc. | A61B005/04
US08734359 | B2 | 2014-05-27 | Method and system for determining an individual's state of attention | 2010-08-03 | The present invention relates to a method and a system designed to determine an individual's state of attention from an individual's respiratory signal. The method of the invention comprises a first learning step, wherein the characteristics of an individual's normal state are determined by selecting a respiratory signal fragment that is considered to be normal according to a pre-established criterion, and a second analysis phase, wherein the individual's state of attention is determined from parameters extracted from the respiratory signal on the basis of some pre-defined rules and the individual's normal state previously characterised. The method of the invention is implemented in two embodiments. A first embodiment is based on the identification of pre-defined patterns in the respiratory signal and, in the second embodiment, an index indicative of the variability of the respiratory signal is defined. | 1. A method for determining an individual's state of attention, which comprises the following operational steps: a) providing a respiratory signal capture apparatus,b) obtaining an individual's respiratory signal with the respiratory signal capture apparatus,c) segmenting said respiratory signal into a plurality of segments,d) extracting at least one parameter from each segment of the respiratory signal,e) characterising a normal state of attention for the individual by selecting a respiratory signal fragment that is considered to be normal according to a pre-established criterion and, in the event that no respiratory signal fragment is found that meets said pre-established criterion, characterising the normal state of attention for the individual on the basis of certain data stored in a memory device, andf) determining the individual's state of attention from the at least one parameter extracted from each segment on the basis of pre-defined rules and the normal state of attention for the individual, wherein step d) comprises obtaining at least one fuzzy parameter from the at least one parameter extracted from each segment, and step f) comprises: evaluating a degree of similarity of each segment with respect to each one of a plurality of pre-determined phases, by comparing the at least one fuzzy parameter of each segment to characteristic parameters for said pre-determined phases, and determining the individual's state of attention from the degree of similarity of the segments of the signal with respect to said pre-determined phases on the basis of certain pre-defined rules. 2. The method for determining an individual's state of attention, according to claim 1, characterised in that said certain data stored in the memory device are data pertaining to the subject him/herself and/or data for a set of subjects that statistically represent a population. 3. The method for determining an individual's state of attention, according to claim 1, characterised in that step e) comprises: determining the variability of the respiratory signal in the signal fragment considered, or evaluating the degree of similarity of the respiratory signal fragment considered with respect to a model respiratory signal pattern, or evaluating the homogeneity of certain given variables in the respiratory cycles within the respiratory signal fragment in question. 4. The method for determining an individual's state of attention, according to claim 1, characterised in that the at least one parameter is selected from the group constituted by amplitude, frequency, minimum value, maximum value, amplitude symmetry and frequency symmetry. 5. The method for determining an individual's state of attention, according to claim 1, characterised in that step c), wherein the respiratory signal is segmented, comprises using a watershed relevant peak algorithm in order to reduce the amount of information to be processed. 6. The method for determining an individual's state of attention, accor... | Fico Mirros, S.A. | A61B005/08
US08942781 | B2 | 2015-01-27 | Medical system comprising a percutaneous probe | 2009-04-09 | The percutaneous probe, made in MRI-compatible materials, comprises: a body percutaneously inserted into the tissue of a patient's body organ (8) having a region (10) to be analyzed, treated and monitored during a single medical procedure; at least one information collection sensing device (30,33,34); treatment application transducers (30) 360° disposed to emit focused or defocused therapeutic ultra-sound waves. The computerized system comprises a parametrizable command device (50) adapted to simulate then command a generation of the therapeutic ultra-sound waves, and to monitor the treatment by thermal MRI images. | 1. A percutaneous probe of MRI-compatible materials, the percutaneous probe comprising: a body made of MRI-compatible materials and having an insertion end shaped to be percutaneously inserted into tissue of a patient's body organ having a region to be analyzed, treated and monitored during a single medical procedure, the body of the percutaneous probe including a cavity, the cavity being configured to insert fluid into the patient through an opening in the insertion end of the body, to collect a part of said tissue from the patient's body through the opening, and to decrease in volume said region of the organ after treatment by allowing substances to be aspirated through the opening into the cavity,a rigid mandrel shaped to be inserted into the cavity,at least one information collection sensing device, adapted to collect information about the region of the organ, anda plurality of treatment application transducers disposed all around a circumference and a length of a part of the body of the probe which is to be inserted in the region to be treated, the plurality of treatment application transducers being operable as a phased-array, adapted to emit both focused and not-focused therapeutic ultra-sound waves to the region of the organ. 2. The percutaneous probe according to claim 1, wherein the information collection sensing device comprises at least one ultra-sound transducer adapted to perform ultra-sound imaging of the body organ, and wherein each transducer is adapted to be operated in an imaging mode, wherein it operates as the information collection sensing device, and in a therapeutic mode, wherein it operates as the treatment application transducer. 3. The percutaneous probe according to claim 1, comprising an aspiration device adapted to collect the part of said tissue from the patient's body through the cavity of the probe body. 4. The percutaneous probe according to claim 1, wherein the information collection sensing device comprises at least one of the following: at least one electro-encephalography electrode adapted to collect brain physiological signals,at least one biosensor adapted to collect tissue physiological signals,at least one thermal electrode adapted to measure the temperature of the region,a biopsy needle adapted to collect tissue. 5. The percutaneous probe according to claim 1, comprising fluid evacuation circuitry, adapted to evacuate fluid from the organ through the probe body. 6. The percutaneous probe according to claim 1 comprising fluid circuitry adapted to allow a cooling fluid to flow in the probe body. 7. The percutaneous probe according to claim 1 wherein treatment application transducers are adapted to generate a temperature rise of at least 30° C., within less than 1 s, in a treatment region of spherical shape of at least 2 mm in diameter neighbouring the treatment application device. 8. The percutaneous probe according to claim 1 wherein the treatment application transducers are adapted to generate a t... | Universite Pierre et Marie Curie (Paris 6) Assistance Publique Hopitaux de Paris | A61B005/05 A61B008/08 A61B005/00 A61B008/12 A61N007/02 A61B005/01 A61B005/0402 A61B005/0476 A61B005/055 A61B017/00 A61B019/00 A61N007/00
US08610583 | B2 | 2013-12-17 | Abnormality detecting device for heart assist device, method for detecting abnormality of heart assist device, and abnormality detecting program | 2009-07-13 | In order to provide an abnormality detecting device for a heart assist device, a method for detecting an abnormality of a heart assist device and an abnormal state of a heart assist device detecting program which can early detect an abnormal state, the abnormality detecting device for a heart assist device according to the present invention includes a user information acquiring means for obtaining user information that indicates an operation state of the heart assist device implanted in a body of a user or a biological state of the user measured by the heart assist device and associating said user information with time to store in a history information storing means, and an abnormal state judging means for referring said history information storing means and judging whether or not an abnormal state is present based on a history of said user information. | 1. An abnormality detecting device for a heart assist device, comprising: a user information acquiring means for obtaining user information that indicates an operation state of the heart assist device implanted in a body of a user or a biological state of the user measured by the heart assist device and associating said user information with time to store in a history information storing means; andan abnormal state judging means for referring to said history information storing means and judging whether or not an abnormal state is present based on a history of said user information,wherein said heart assist device comprises a liquid pump that is controlled by electric power such that a rotation speed is a desirable rotation speed,wherein said user information acquiring means acquires consumed power of said liquid pump as said user information, andwherein said abnormal state judging means judges whether or not the abnormal state is present in a flow path, based on a history of said consumed power of the liquid pump. 2. The abnormality detecting device for a heart assist device according to claim 1, further comprising: a wireless communication means for sending an alarm signal indicating the abnormal state to a contact receiver device located at a distant place, when the abnormal state judging means judges that the abnormal state is present. 3. The abnormality detecting device for a heart assist device according to claim 2, further comprising: a location information acquiring means for acquiring location information that indicates a current location,wherein said wireless communication means sends a signal including said location information as said alarm signal. 4. The abnormality detecting device for a heart assist device according to claim 2, wherein said user information acquiring means, said history information storing means, and said abnormal judging means are provided in a controller device that is connected to said heart assist device with a wire line, wherein said wireless communication means comprises:a first wireless communication means provided in said controller device; anda second wireless communication means provided in a mobile terminal, andwherein said first wireless communication means sends said alarm signal to said contact receiver device via said second wireless communication means. 5. The abnormality detecting device for a heart assist device according to claim 4, further comprising: an emergency treatment informing means for generating emergency treatment information that indicates contents of emergency treatment to inform around of the user, when the abnormal judging means judges that the abnormal state is present. 6. The abnormality detecting device for a heart assist device according to claim 5, wherein said emergency treatment informing means informs said emergency treatment information with an audio output device that outputs sound, and wherein said audio output device is provided in either one of said mobile termina... | Mitsubishi Heavy Industries, Ltd. | G08B001/08 G08B023/00 A61N001/00 A61M016/00
US20110021928 | A1 | 2011-01-27 | METHODS AND SYSTEM OF DETERMINING CARDIO-RESPIRATORY PARAMETERS | 2010-07-14 | Embodiments of the present invention provide noninvasive methods and systems of determining and monitoring an individual's respiration pattern, respiration rate, other cardio-respiratory parameters or variations thereof. | 1. A method of determining an individual's respiration pattern, respiration rate, other cardio-respiratory parameters or variations thereof, the method comprising sensinga) mechanical movements of the individual's chest; and/orb) acoustic waves generated by the individual's heart beats;processing said mechanical movements and/or acoustic waves to obtain one or more respiration-dependent parameters for subsequent computational analysis of the individual's cardio-respiratory parameters, whereinone respiration-dependent parameter is a variation in S1-S2 intervals between two consecutive heart beats and another respiration-dependent parameter is a variation in heart sound amplitude. 2. The method of claim 1, wherein the computational analysis comprises estimating S1-S2 interval variations between a beat and its preceding beat; whereinbeats include both first (S1) and second (S2) heart sounds;quantifying the similarity between the preceding beat and versions of the beat;identifying maximum similarity throughout the versions and assigning a corresponding S1-S2 interval variation to the beat based on identified version. 3. The method of claim 2, wherein the heart beat is detected by its first heart sound S1, and wherein assessing a degree of similarity only includes the second sound S2. 4. The method of claim 2, wherein the heart beat is detected by its second heart sound S2, and wherein assessing a degree of similarity only includes the first sound S1. 5. The method of claim 1, wherein an additional respiration-dependent parameter is a variation in S1-S1 intervals. 6. The method of claim 1, wherein an additional respiration-dependent parameter is chest wall motion. 7. The method of claim 1, wherein the computational analysis is carried out with a combined plurality of respiration-dependent parameters. 8. The method of claim 1, wherein the computational analysis is carried out with one respiration-dependent parameter. 9. A method of detecting respiratory disorders in an individual, the method comprising sensinga) mechanical movements of the individual's chest; and/orb) acoustic waves generated by the individual's heart beats;processing said mechanical movements and/or acoustic waves to obtain one or more respiration-dependent parameters for subsequent computational analysis of the individual's respiration pattern, respiration rate, other cardio-respiratory parameters or variations thereof, wherein one respiration-dependent parameter is a variation in S1-S2 intervals between two consecutive heart beats and another respiration-dependent parameter is a variation in heart sound amplitude. 10. The method of claim 9, wherein the computational analysis comprises estimating S1-S2 interval variations between a beat and its preceding beat; whereinbeats include both first (S1) and second (S2) heart sounds;quantifying the similarity between the preceding beat and versions of the beat;identifying maximum similarity throughout the versions and assigning a co... | The Boards of Trustees of the Leland Stanford Junior University | A61B005/0205
US08764187 | B2 | 2014-07-01 | Determination of the effective lens position of an intraocular lens using aphakic refractive power | 2010-07-13 | An ophthalmic method for determining a relationship between aphakic ocular power and estimated effective lens position (ELP) of an intraocular lens (IOL) to be implanted in a patient's eye. The method can be used to determine an estimate of the ELP of an IOL given the aphakic ocular power of the patient's eye, for example, without measurement of the corneal curvature or axial length of the patient's eye. The estimate of ELP can then be used to determine a suitable value of optical power for the IOL to be implanted in the patient's eye. | 1. A method for determining the optical power of an intraocular lens to be inserted into the eye of a patient, the method comprising: receiving as an input from an input device an aphakic refractive power value of the patient's eye;determining, with a processor, an estimate of the post-surgical effective lens position (ELP) of the intraocular lens for the patient's eye, the estimate being based on the aphakic refractive power value of the patient's eye and on a relationship between aphakic refractive power and post-surgical intraocular lens ELP; anddetermining an amount of optical power for the intraocular lens to be inserted into the patient's eye based on the estimate of the post-surgical ELP of the intraocular lens in the patient's eye. 2. The method of claim 1, wherein the relationship between aphakic refractive power and post-surgical intraocular lens ELP is determined using regression analysis. 3. The method of claim 1, wherein the estimate of the post-surgical ELP of the intraocular lens is not based on measured corneal power of the patient's eye. 4. The method of claim 1, wherein the estimate of the post-surgical ELP of the intraocular lens is not based on measured axial length of the patient's eye. 5. The method of claim 1, wherein the aphakic refractive power value of the patient's eye comprises an intraoperative measurement of the aphakic refractive power of the patient's eye. 6. The method of claim 5, wherein the intraoperative measurement of the aphakic refractive power of the patient's eye comprises the spherical equivalent power of the patient's eye. 7. The method of claim 1, further comprising outputting the estimate of the post-surgical ELP of the intraocular lens. 8. The method of claim 1, further comprising outputting the amount of optical power for the intraocular lens. 9. The method of claim 1, wherein the estimate of the post-surgical ELP of the intraocular lens is based on the intraocular lens being inserted into the capsular bag of the patient's eye. 10. The method of claim 1, wherein determining the estimate of the post-surgical ELP of the intraocular lens comprises calculating the estimate using a mathematical function that comprises the relationship between aphakic refractive power and intraocular lens ELP. 11. The method of claim 1, further comprising performing a measurement using the input device to determine the aphakic refractive power value of the patient's eye. 12. The method of claim 1, wherein the input device comprises a wavefront sensor. 13. A non-transitory machine-readable medium that, when read by a machine, causes the machine to perform a method comprising: receiving as an input an indication of the aphakic refractive power value of the patient's eye;determining an estimate of the post-surgical effective lens position (ELP) of the intraocular lens based on the aphakic refractive power value of the patient's eye and on a relationship between aphakic refractive power and intraocular lens ELP; an... | Wavetec Vision Systems, Inc. | G02C007/02 A61F002/16
US08457726 | B2 | 2013-06-04 | Heart failure detecting medical device | 2008-04-29 | An implantable medical device has an event detector that detects a predetermined cardiac event during a heart cycle of a subject. A reference time is assigned to this detected cardiac event. An onset detector detects the onset of ventricular filling of the heart during the heart cycle. The relative time of the detected filling onset is determined based on the assigned time reference. An increased risk of heart failure of the subject is automatically determined based on the determined relative time for the filling onset. Generally, a reduction in the relative time, as determined at different points in time, indicates an increased heart failure risk or the presence of a heart failure condition. | 1. An implantable medical device comprising: an event detector that detects occurrence of a predefined heart event in a detection time window from start of a heart cycle of a heart up to opening of aortic valves in said heart cycle, said detected predefined event is assigned a time reference;an onset detector that detects an onset of ventricular filling of said heart in said heart cycle;a timing processor that determines a relative time of said detected onset of ventricular filling based on said assigned time reference;a heart failure processor that generates detection data indicative of an increased risk of heart failure of a subject based on said determined relative time;a signal applier that applies a current or voltage signal over at least a portion of said heart via two electrodes;a signal measurer that measures a resulting voltage or current signal over between said two electrodes;an impedance processor that determines an impedance signal based on said applied current or voltage signal and said measured resulting voltage or current signal, wherein said onset detector detecting said onset of ventricular filling based on said impedance signal.a time assigner that assigns a zero time reference to said detected predefined event:said timing processor comprising a sample counter that counts a number of samples in said impedance signal from said occurrence of said predefined event up to said onset of ventricular filling; andsaid timing processor comprising a time converter that determines said relative time of said detected onset of ventricular filling by dividing said counted number of samples with a sampling frequency of said impedance processor. 2. The device according to claim 1, further comprising a time assigner that assigns said time reference to said detected predefined event. 3. The device according to claim 1, wherein said event detector comprises a QRS detector that detects occurrence of a QRS complex during said detection time window based on an electric signal sensed from a heart of said subject. 4. The device according to claim 1, wherein said signal applier applies said current or voltage signal during at least one respiratory cycle of said subject and wherein said impedance processor determines a respective impedance signal for each complete heart cycle occurring during said at least one respiratory cycle and determines an average impedance signal based on said respective impedance signals. 5. The device according to claim 1, wherein said signal applier applies said current or voltage signal during multiple respiratory cycles of said subject, said implantable medical device further comprising a heart cycle detector that detects, based on said impedance signal, multiple heart cycles occurring at a predefined respiratory phase, and wherein said impedance processor determines a respective impedance signal for each heart cycle of said multiple heart cycles and determines an average impedance signal based on said respective impedanc... | St. Jude Medical, AB | A61B005/04
US20110040170 | A1 | 2011-02-17 | ENDOSCOPIC IMAGING PHOTODYNAMIC THERAPY SYSTEM AND METHODS OF USE | 2009-01-28 | The invention provides an endoscopic imaging photodynamic therapy system (EIPS) for focused tissue ablation by illumination of a photosensitizer drug in a target tissue, said system comprising an endoscopic assembly, a real-time imaging component for locating the target tissue and monitoring the ablation intervention, a therapeutic light system and, optionally, a drug delivery module, wherein said imaging component comprises a flexible transducer with an operative channel for insertion of a flexible light guide of the therapeutic light system and, optionally, a flexible drug delivery catheter of the drug delivery module. This EIPS may be used in various medical applications where tissue ablation is required and photodynamic therapy may be applied, in particular, in the treatment of extrauterine pregnancy (EUP). | 1. An endoscopic imaging photodynamic therapy system for focused tissue ablation of a feto-placental unit in an extrauterine pregnancy (EUP) by illumination of a photosensitizer drug in said feto-placental unit, said system comprising an endoscopic assembly designed for entry into a vaginal tract of a female patient, a real-time imaging component for locating the feto-placental unit and monitoring the ablation intervention, a therapeutic light system and a drug delivery module, wherein said imaging component comprises a flexible transducer with an operative channel for insertion of a flexible light guide of the therapeutic light system and a flexible drug delivery catheter of the drug delivery module. 2. The endoscopic imaging photodynamic therapy system of claim 1, wherein: (a) said endoscopic assembly comprises a control handle, an operation handle and an application adaptor;(b) said real-time imaging component comprises means for guidance for location of said feto-placental unit and monitoring of the ablation intervention in said feto-placental unit, and a flexible transducer with an operative channel;(c) said therapeutic light system consists of a light source, a flexible light guide and an operating switch for the light system; and(d) said drug delivery module comprises a flexible drug delivery catheter adapted for injecting a photosensitizer drug to the feto-placental unit, a drug delivery means and a photosensitizer drug in an injectible form. 3. The endoscopic imaging photodynamic therapy system of claim 2, wherein the flexible light guide of the therapeutic light system, and the flexible drug delivery catheter of the drug delivery module are inserted into the operative channel of the flexible transducer of the real-time imaging component. 4. The endoscopic imaging photodynamic therapy system of claim 3, wherein said insertion is via a flexible service catheter. 5. The endoscopic imaging photodynamic therapy system of claim 2, wherein: (ii) said control handle of the endoscopic assembly is manual or computer-controlled and comprises a proximal grip and at least one service opening; or(iii) said control handle of the endoscopic assembly has a second service opening adapted for washing the tissue, suction from the tissue, or insertion of needle biopsies to sample cells from an abnormal area for laboratory testing. 6. (canceled) 7. The endoscopic imaging photodynamic therapy system of claim 2, wherein said operation handle of the endoscopic assembly comprises means for aiming and bending the flexible transducer with the flexible drug delivery catheter and the flexible light guide towards the feto-placental unit. 8. The endoscopic imaging photodynamic therapy system of claim 1, wherein: said real-time imaging component is ultrasound (US), magnetic resonance imaging (MRI), computerized tomography (CT), positron emission tomography (PET), light-based video, or any combination thereof; or(ii) said real-time imaging component also measures... | YEDA RESEARCH AND DEVELOPMENT CO., LTD. | A61N005/00 A61B001/00 A61B008/12 A61B005/055
US20110040198 | A1 | 2011-02-17 | Invasive Cardiology Digital Signal Amplifier and Acquisition Device | 2010-10-27 | The digital signal amplifier and acquisition system includes at least one catheter input module (CIM) configured to receive, route and digitize incoming analog cardiac signals from a number of catheters, as well as outgoing stimulator pulses. A plurality of CIMs are mechanically stacked and electrically daisy-chained and coupled with a mounting. The mounting platter is clamped to the bedrail and provides a single digital output cable to the base. The system also includes an acquisition device and the aforementioned base, which is configured to collect, filter and distribute the acquired data. Lastly an analog output module receives filtered digital signals from the base and is configured to reconstruct analog representations of such filtered digital signals for outside devices. | 1.-9. (canceled) 10. A method of cardiology signal acquisition and amplification for distribution to at least a display, the method comprising the steps of: collecting a set of analog cardiac data from a patient catheter with a cardiac input module;converting the set of analog cardiac data directly to a set of digital data within the cardiac input module without amplification;removably coupling a mount platter to a patient bed, such that the mount platter is coupled with he catheter input module and holds the catheter input module;communicating the set of digital data from the cardiac input module to the mount platter;communicating the set of digital data from the mount platter to a base unit, wherein the base unit is coupled with the patient bed and is configured to filter and distribute the set if digital data for illustration on the display,wherein the mount platter is further configured to support a plurality of catheter input modules in a stacked configuration, further wherein the mount platter and each of the plurality of catheter input modules are electrically coupled to one another. 11. The method as claimed in claim 10, further comprising receiving a filtered set of digital data in an analog output module from the base unit, wherein the analog output module is configured to convert the filtered set of digital data to a reconstructed analog signal. 12. The method as claimed in claim 10, further comprising coupling an acquisition device with the base unit, the acquisition device configured to control the collection of a plurality of physiological parameters from a patient. 13. The method as claimed in claim 10, wherein the mount platter includes an emergency stimulator connector configured to provide a pulse to the patient catheter in the event of total system failure. 14. The method as claimed in claim 10, wherein the mount platter includes an auxiliary reference input configured to collect a non-catheter patient input. 15. The method as claimed in claim 10, wherein the catheter input module includes a plurality of inputs. 16. The method as claimed in claim 15, wherein the mount platter includes a test signal generator configured to test the functionality of any of the plurality of inputs of the catheter input module. 17. (canceled) 18. The method as claimed in claim 10, wherein the mount platter is fastened to a patient bed. | THE GENERAL ELECTRIC COMPANY | A61B005/0428
US20110021929 | A1 | 2011-01-27 | SYSTEMS AND METHODS FOR CONTINUOUS NON-INVASIVE BLOOD PRESSURE MONITORING | 2009-07-27 | Systems and methods are disclosed herein for continuous non-invasive blood pressure (CNIBP) monitoring. Multiple reference blood pressure values may be obtained using a calibration device. These multiple reference blood pressure values may be used as calibration points for determining a relationship between the blood pressure of a patient and photoplethysmograph (PPG) signals. | 1. A system for monitoring blood pressure of a patient, the system comprising: a signal generator for generating at least two PPG signals from at least two respective sensors attached to the patient;a processor coupled to the signal generator, wherein the processor is capable of: determining at least two reference blood pressure values based at least in part on a calibration device coupled to the patient and to the processor, andupdating a relationship between blood pressure of the patient and the at least two PPG signals based at least in part on the at least two reference blood pressure values,calculating a blood pressure value based at least in part on the updated relationship; andan output device coupled to the processor. 2. The system of claim 1, wherein the processor is further capable of: identifying at least two points in the at least two PPG signals, wherein the at least two points occur after the reference blood pressure value is obtained;determining a time difference between the at least two points; andcalculating the blood pressure value based at least in part on the time difference and the updated relationship. 3. The system of claim 2, wherein the relationship is P=a+b·ln(T)or a mathematical equivalent thereof, wherein P is the blood pressure value, T is the time difference, and a and b are constants determined based at least in part on the at least two reference blood pressure values. 4. The system of claim 1, wherein the processor is further capable of identifying a reference blood pressure value as an outlier. 5. The system of claim 4, wherein the processor is further capable of determining a new reference blood pressure value to verify the outlier. 6. The system of claim 1, wherein the processor is further capable of: associating weighting factors with the at least two reference blood pressure values; andupdating the relationship based at least in part on the at least two reference blood pressure values and the weighting factors. 7. The system of claim 1, wherein the processor is further capable of: identifying a blood pressure event;determining at least two further reference blood pressure values after the blood pressure event occurs;resetting the relationship between blood pressure of the patient and the at least two PPG signals; andupdating the relationship based at least in part on the at least two further reference blood pressure values. 8. The system of claim 7, wherein the blood pressure event is a change in vascular compliance. 9. The system of claim 7, wherein the blood pressure event is a blood pressure change that exceeds a threshold stored in the processor. 10. A method for monitoring blood pressure of a patient, the method comprising: determining using a processor at least two reference blood pressure values based at least in part on a calibration device coupled to the patient and to the processor;obtaining at least two PPG signals from at least two respective sensors attached to the patient;updating a rela... | Nellcor Puritan Bennett Ireland | A61B005/02
US09014787 | B2 | 2015-04-21 | Bioabsorbable target for diagnostic or therapeutic procedure | 2010-05-28 | An implantable device has a body that is substantially rigid and has an imageable shape. The body is further bioabsorbable and may contain permanent metallic elements to aid in its imaging. When the device is implanted in a resected cavity in soft tissue, it can cause the cavity to conform substantially to a known imageable shape. The implantable device is further imageable due to its attenuation properties being different from those of soft tissue such that the boundaries of the tissue corresponding to the predetermined shape can be determined. | 1. A device for targeting external beam radiation to a region of tissue surrounding a lumpectomy cavity comprising: an implant body, the implant body being formed of a bioabsorbable polymer, and comprising a periphery; anda plurality of radiopaque markers attached to the implant body about the periphery to define orthogonal x, y, and z axes of the implant body, the orthogonal x, y, and z axes having an axial length between the plurality of radiopaque markers,wherein the implant body includes spaced apart elements that define an open architecture that permits partial encapsulation of the cavity wall around the elements upon initial placement of the implant body within the cavity while the radiopaque markers remain in position for targeting the region of tissue surrounding the lumpectomy cavity,and wherein the x, y, and z axes comprise a range of axial lengths so that the implant body can be sized to fit an irregular surgical cavity. 2. The device of claim 1, wherein the radiopaque markers comprise metallic clips. 3. The device of claim 1, wherein the bioabsorbable polymer has an absorption time of at least 2 weeks. 4. The device of claim 3, wherein the bioabsorbable polymer has an absorption time of at least 6 months. 5. The device of claim 1, wherein the bioabsorbable polymer is thermoplastic. 6. The device of claim 1, further comprising one or more protrusions for bluntly penetrating into surrounding tissue. 7. The device of claim 1, wherein the radiopaque markers allow visualization by radiography. 8. The device of claim 7, wherein the radiography includes mega-voltage x-rays. 9. The device of claim 7, wherein the radiography includes kilo-voltage x-rays. 10. The device of claim 7, wherein the radiography includes fluoroscopy. 11. A device for targeting external beam radiation to regions of target tissue surrounding a tumor resection cavity comprising: an implant body, the implant body having an open architecture and being formed of a bioabsorbable material, and comprising a periphery; anda plurality of radiopaque markers attached to the implant body about the periphery to define orthogonal x, y, and z axes of the implant body, the orthogonal x, y, and z axes having an axial length between the plurality of radiopaque markers, wherein the open architecture implant body supports a wall of the cavity yet also permits partial encapsulation of the cavity wall around the body upon initial placement of the body within the cavity, andwherein the x, y, and z axes comprise a range of axial lengths so that the implant body can be sized to fit an irregular surgical cavity. 12. The device of claim 11, wherein the radiopaque markers comprise surgical clips. 13. The device of claim 11, wherein the bioabsorbable material is a thermoplastic polymer. 14. The device of claim 11, further comprising one or more protrusions for bluntly penetrating into surrounding tissue. 15. The device of claim 11, wherein the bioabsorbable material has an absorption ... | Focal Therapeutics, Inc. | A61B005/05 A61B019/00 A61N005/10 A61B017/00
US08337476 | B2 | 2012-12-25 | Real time urine monitoring system | 2009-08-20 | A disposable real-time closed catheter urine output monitoring system. The monitoring system of the invention can measure urine outputs at both high and low urine flow in real time, using a pressure chamber in the form of a hollow body having an input port for coupling to a tube connected to a catheter for collection of urine from a patient, an exit port coupled to a collection container and a vent port, such that air may flow through the vent, but urine is blocked by the vent. A sensor measures flow from the exit port of the pressure chamber into the collection container, and a pressure sensor may be included to measure pressure in the pressure chamber. The invention also provides a method of using the apparatus in a real-time monitoring system. | 1. A urine collection monitoring system comprising: a) a pressure chamber, comprising: i) a hollow body having an input port for coupling to a tube connected to a catheter for collection of urine from a patient, an exit port and at least one vent port;ii) a vent located at the vent port, such that air may flow through the vent, but urine is blocked by the vent;b) a flow sensor for measuring a quantity of urine flowing from the exit port of the pressure chamber at both a continuous flow and a discrete flow wherein the flow sensor is configured to measure: (i) the quantity flowing at the continuous flow, based upon a pressure in the hollow body, and (ii) the quantity flowing at the discrete flow, by counting discrete units of fluid flowing from the exit port;c) a urine collection container coupled to the exit port of the pressure chamber; andd) a controller having an input coupled to the flow sensor. 2. The system of claim 1, further comprising an adaptor located at the inlet port of the hollow body, having an inlet for coupling to the tube, and an outlet opening into the hollow body through the inlet port. 3. The system of claim 2, in which the outlet of the adaptor is in the form of a frustum of a cone having a diameter which increases from the inlet to the outlet. 4. The system of claim 2, further comprising a catheter coupled to the inlet of the adaptor by a tube. 5. The system of claim 1, further comprising a nozzle connected to the exit port of the pressure chamber. 6. The system of claim 1, in which the vent comprises a check valve having a floating check element which floats into a blocking position in the presence of liquid, in which the floating check element blocks the vent, preventing urine flow through the vent. 7. The system of claim 6, in which the check valve comprises a v-shaped seat, the floating check element is a ball, and the blocking position is a position in which the ball floats into the v-shaped seat to seal the check valve. 8. The system of claim 1, in which the vent further comprises a bidirectional membrane on an outlet of the vent that permits air to flow through the vent, but limits urine leakage from the outlet of the vent. 9. The system of claim 1, in which the urine collection container is a collapsible urine collection bag. 10. The system of claim 9, in which the urine collection bag further comprises a plurality of depressions on a surface, such that when the bag is empty, a gap will be maintained and urine will be able to fully pass through the collecting bag. 11. The system of claim 9, in which the urine collection bag further comprises a bidirectional air membrane for venting air into and out of the bag. 12. The system of claim 1, in which the urine collection container further comprises a drain valve for emptying the container. 13. The system of claim 1, in which the pressure chamber is enclosed in a pocket in the urine collection container. 14. The system of claim 1, in which the controller compr... | Greenwald Technologies, LLC | A61M001/00 A61B005/00 B65D081/00 G01F007/00 G01F001/20 G01F013/00
US07881773 | B2 | 2011-02-01 | Multi-mode internal imaging | 2006-07-12 | The invention described herein provides systems and methods for multi-modal imaging with light and a second form of imaging. Light imaging involves the capture of low intensity light from a light-emitting object. A camera obtains a two-dimensional spatial distribution of the light emitted from the surface of the subject. Software operated by a computer in communication with the camera may then convert two-dimensional spatial distribution data from one or more images into a three-dimensional spatial representation. The second imaging mode may include any imaging technique that compliments light imaging. Examples include magnetic resonance imaging (MRI) and computer topography (CT). An object handling system moves the object to be imaged between the light imaging system and the second imaging system, and is configured to interface with each system. | 1. A system for providing multiple types of imaging data for an internal portion of an object, the system comprising: a light imaging system for providing light data related to light emitted from a light source located inside the object, the light imaging system including: an imaging chamber comprising an interior cavitya first light seal disposed on an exterior wall thereof, anda camera configured to capture the light data from the object while the object is in the interior cavity;a second imaging system configured to provide imaging data for the internal portion of the object and having a receiving area that allows the internal portion to be imaged while the object is in the receiving area; andan object handling system adapted to transfer the object between the interior cavity of the light imaging system and the receiving area of the second imaging system, wherein the object handling system is configured to move the object within the interior cavity of the imaging chamber when the imaging chamber is sealed in a light tight manner, wherein the object handling system includes: a portable stage that is adapted to support the object,a manipulator configured to move the stage between the interior cavity of the light imaging system and the receiving area of the second imaging system, anda second light seal disposed on said manipulator and configured to interface with the first light seal on the exterior wall of the light imaging system to produce a substantially light tight seal between the interior cavity of the light imaging system and the space outside the light imaging system while the stage is within the interior cavity. 2. The system of claim 1 wherein the object handling system is configured to transfer the object between the interior cavity of the light imaging system and the receiving area of the second imaging system without substantially moving the object on the portable stage. 3. The system of claim 1 further comprising a gas delivery system configured to deliver an anesthesia gas to at least one gas delivery outlet that provides the anesthesia gas to the object while the object rests on the portable stage. 4. The system of claim 1 wherein the second imaging system is configured to produce spatial information for the internal portion. 5. The system of claim 4 wherein the light imaging system is configured to produce functional information for the internal portion. 6. The system of claim 5 wherein the second imaging system is configured to produce three dimensional spatial information for the internal portion and the light imaging system is configured to produce three dimensional functional information for the internal portion. 7. The system of claim 6 wherein an exterior wall of the imaging chamber comprises a moveable interface having: a hole that receives the portable stage of the object handling system therethrough;a first circular bearing disposed between the exterior wall and a wall of the moveable interface and permits rotatio... | Xenogen Corporation | A61B005/00
US07881774 | B2 | 2011-02-01 | Apparatus for obtaining ultrasonic image and method of obtaining ultrasonic image | 2006-07-18 | A transmitting/receiving circuit drives an ultrasonic probe in accordance with a trigger signal based on an electrocardiographic waveform under the control of a control device so as to scan a plurality of regions, and obtains scan data for each region. During the scanning, the circuit causes the probe to scan the respective regions so that scan data in which time phases of the electrocardiographic waveform substantially coincide are obtained in the vicinities of the boundary between at least two adjacent regions among the regions. That is, in adjacent regions, the circuit causes the probe to scan the respective regions in the main or sub-scanning directions set to be reverse to each other. Further, an image processor generates ultrasonic image data of the range in which the regions are joined, on the basis of the scan data obtained by starting scanning at the same time phase. | 1. An apparatus for obtaining an ultrasonic image comprising: an ultrasonic probe that scans an ultrasonic beam in a main scanning direction and a sub-scanning direction;a scan device that receives a trigger signal based on an electrocardiographic waveform, drives the ultrasonic probe in accordance with the trigger signal so as to scan a plurality of regions, and obtains scan data for each region; andan image generating device configured to combine the scan data obtained for each region so as to generate ultrasonic image data of a range in which the plurality of regions are joined,wherein the scan device is configured to cause the ultrasonic probe to scan the plurality of regions so that scan data in which time phases of the electrocardiographic waveform substantially coincide are obtained in the vicinities of the boundary between at least two adjacent regions among the plurality of regions, and to divide the regions by boundaries drawn extending outward from a center point of the regions in at least two directions passing through the center point. 2. The apparatus for obtaining an ultrasonic image according to claim 1, wherein, whenever receiving the trigger signal, the scan device causes the ultrasonic probe to scan a different region so as to obtain scan data in which the time phases of the electrocardiographic waveform substantially coincide in the vicinities of the boundary, thereby obtaining scan data for each region, andthe image generating device combines scan data obtained by scanning the plurality of regions and the scan data obtained by starting scanning at the same time phase so as to generate ultrasonic image data of the range in which the plurality of regions are joined. 3. The apparatus for obtaining an ultrasonic image according to claim 1, wherein, in adjacent regions, the scan device causes the ultrasonic probe to scan the plurality of regions in the main or sub-scanning directions set to be reverse to each other so as to obtain scan data in which the time phases of the electrocardiographic waveform substantially coincide. 4. The apparatus for obtaining an ultrasonic image according to claim 1, wherein the image generating device combines the scan data obtained by scanning the plurality of regions and the scan data obtained by starting scanning at the same time phase so as to generate ultrasonic image data of the range in which the plurality of regions are joined. 5. The apparatus for obtaining an ultrasonic image according to claim 1, wherein, in adjacent regions among the plurality of regions, the scan device causes the ultrasonic probe to scan the plurality of regions by changing the main or sub-scanning directions so that the main or sub-scanning directions are axisymmetric with each other with a predetermined dividing line set to a symmetric axis to obtain scan data in which the time phases of the electrocardiographic waveform substantially coincide. 6. The apparatus for obtaining an ultrasonic image according to claim... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | A61B005/05
US07881769 | B2 | 2011-02-01 | Method and system for mounting an MPS sensor on a catheter | 2003-04-07 | Catheter for performing a medical operation on an organic lumen, the catheter including an elongated member, a medical operational element located at a distal end of the elongated member, an electromagnetic field detector located at the distal end, and a wiring for coupling the electromagnetic field detector with a medical positioning system, wherein the medical positioning system determines the position and orientation of the distal end. | 1. Catheter for performing a medical operation on an organic lumen, the catheter comprising: an elongated member comprising substantially flexible material;a medical operational element located at a distal end of said elongated member;an electromagnetic field detector located at said distal end, said electromagnetic field detector operative to detect an electromagnetic field generated by a medical positioning system that is proximal of said elongated member distal end; anda wiring for coupling said electromagnetic field detector with said medical positioning system, said wiring operative to transmit a signal respective of said detected electromagnetic field to said medical positioning system, said wiring being separate and isolated from said medical operational element and from said electromagnetic field detector, said wiring embedded within substantially the entirety of said elongated member between said medical positioning system and said electromagnetic field detector such that said wiring is positioned within the wall of said elongated member and is surrounded, in said embedded portion, by said substantially flexible material of said elongated member,wherein said medical positioning system determines the position and orientation of said distal end, relative to a reference coordinate system, according to said transmitted signal, andwherein said elongated member has a mechanical property of one of rigidity, elasticity, and plasticity, and wherein said mechanical property of the part of said elongated member between said medical positioning system and said electromagnetic field detector where said wiring is embedded is increased or decreased due to said mechanical property of said wiring, irrespective of any modified mechanical property resulting from said medical operational element or from said electromagnetic field detector, such that the increase or decrease in said mechanical property of the part of said elongated member between said medical positioning system and said electromagnetic field detector where said wiring is embedded enhances the maneuverability of said catheter within said organic lumen throughout the stages of: insertion of said catheter into the organic lumen, performance of the medical operation with said medical operational element, and removal of said catheter from the organic lumen. 2. Position and orientation determination system comprising: a guiding catheter;a guiding catheter electromagnetic field detector located at a guiding catheter distal end of said guiding catheter, said guiding catheter electromagnetic field detector being coupled with a medical positioning system; andat least one medical catheter located within said guiding catheter, said at least one medical catheter comprising: an elongated member comprising substantially flexible material;a medical operational element located at least one medical catheter distal end of a respective one of said at least one medical catheter;a medical catheter electromagneti... | MediGuide Ltd. | A61B005/05
US07894873 | B2 | 2011-02-22 | Method for determining the rectum distance in the colon | 2006-04-07 | A method is disclosed for automatically determining the rectum distance in a patient's colon, the path in the colon between a selected site, preferably a lesion, and the rectum being determined by processing 3D image data records. At least two 3 image data records are recorded with the patient being differently positioned. Further, the path between the rectum and selected site is determined from the common perusal of the at least two 3D image data records. | 1. A method for automatically determining a rectum distance in a colon of a patient, comprising: recording at least two 3D image data records of the colon, each recorded with the patient being differently positioned; anddetermining a path in the colon between a selected site and the rectum by processing the 3D image data records, the path between the rectum and selected site including at least one blockage in the colon and being determined from a common perusal of the at least two 3D image data records, the processing excluding registration between the at least two 3D image data records; wherein for each of the at least two 3D image data records, an automatic path calculation is carried out in relation to a plurality of salient points in the colon,subpath lengths between the plurality of salient points are determined, anda total path length between the selected site and the rectum is determined from all determined subpath lengths. 2. The method as claimed in claim 1, wherein the total path is calculated from a sum of at least one of mean values, meridian values, and arithmetic means of the determined subpath lengths. 3. The method as claimed in claim 2, wherein the automatic path calculation begins both at the rectum and at the selected site in the colon. 4. The method as claimed in claim 2, wherein the automatic path calculation begins exclusively at the selected site in the colon. 5. The method as claimed in claim 2, wherein a Dijkstra method is used for the automatic path calculation. 6. The method as claimed in claim 1, wherein the automatic path calculation begins both at the rectum and at the selected site in the colon. 7. The method as claimed in claim 1, wherein the automatic path calculation begins exclusively at the selected site in the colon. 8. The method as claimed in claim 1, wherein a Dijkstra method is used for the automatic path calculation. 9. A method for automatically determining a rectum distance in a colon of a patient, comprising: recording at least two 3D image data records of the colon, each recorded with the patient being differently positioned; anddetermining a path in the colon between a selected site and the rectum by processing the 3D image data records, the path between the rectum and selected site including at least one blockage in the colon and being determined from a common perusal of the at least two 3D image data records, the processing excluding registration between the at least two 3D image data records; wherein for at least one of the at least two 3D image data records, an automatic path calculation is carried out in the colon in relation to at least one salient point whose mean statistical distance from the rectum is at least one of, constant and independent of the patient, anda function of other patient parameters, anda path length to the rectum is determined from a sum of the automatically calculated path and a distance, which does not depend on the patient. 10. The method as claimed in claim 9, ... | Siemens Aktiengesellschaft | A61B005/05
US07871456 | B2 | 2011-01-18 | Membranes with controlled permeability to polar and apolar molecules in solution and methods of making same | 2006-08-10 | A membrane for use in an implantable glucose sensor including at least one crosslinked substantially hydrophobic polymer and at least one crosslinked substantially hydrophilic polymer; wherein the first and second polymers are different polymers and substantially form an interpenetrating polymer network, semi-interpenetrating polymer network, polymer blend, or copolymer. The membranes are generally characterized by providing a permeability ratio of oxygen to glucose of about 1 to about 1000 in units of (mg/dl glucose) per (mmHg oxygen). Three methods of making membranes from hydrophobic and hydrophilic monomers formed into polymer networks are provided, wherein according to at least two of the methods, the monomers may be substantially immiscible with one another. | 1. A membrane having controlled permeability to polar and non-polar solutes, for use in an implantable sensor, comprising different first and second polymers, the first polymer comprising at least one crosslinked substantially hydrophobic polymer and the second polymer comprising at least one crosslinked substantially hydrophilic polymer; wherein the hydrophobic and hydrophilic polymers are present in a mass ratio from between about 50:50 to 5:95 of the total mass of the hydrophobic and hydrophilic polymers, and form an interpenetrating polymer network, semi-interpenetrating polymer network or polymer blend; andwherein further the membrane is characterized by a permeability ratio of oxygen to glucose of about 1 to about 1000 in units of (mg/dl glucose) per (mmHg oxygen). 2. The membrane of claim 1, wherein the hydrophobic and hydrophilic polymers are present in a ratio of 12.5:87.5. 3. The membrane of claim 1, wherein the permeability ratio of oxygen to glucose is about 1 to about 100. 4. The membrane of claim 3, wherein the permeability ratio of oxygen to glucose is about 10 to about 60. 5. The membrane of claim 1, wherein the membrane has a water uptake of about 1 wt. % to about 20 wt. %. 6. The membrane of claim 1, wherein the membrane has a water uptake of about 5 wt. % to about 20 wt. %. 7. The membrane of claim 1, wherein the membrane has a water uptake of about 10 wt. % to about 20 wt. %. 8. The membrane of claim 1, wherein the first polymer comprises a silicone-containing polymer. 9. The membrane of claim 1, wherein the second polymer comprises an acrylate or methacrylate polymer. 10. The membrane of claim 1, wherein the first polymer and the second polymer are not mutually soluble. 11. The membrane of claim 1, wherein the first polymer is prepared from a first monomer, and the second polymer is prepared from a second monomer, and the first and second monomers are not miscible. 12. The membrane of claim 1, wherein the first polymer comprises a silicone-containing polymer, the second polymer comprises an acrylate or methacrylate polymer, and the permeability ratio of oxygen to glucose is about 10 to about 60 in units of (mg/dl glucose) per (mmHg oxygen). 13. The membrane of claim 1, wherein the first polymer comprises a silicone-containing polymer, the second polymer comprises an acrylate or methacrylate polymer, the membrane has a water uptake of about 10 wt. % to about 20 wt. %; and the first polymer and the second polymer are not mutually soluble. 14. The membrane of claim 1, wherein the first polymer is present in units of % wt. of the overall wt. of polymer present in the membrane in a range from 2% wt. to 50% wt. 15. The membrane of claim 14, wherein the first polymer is present in a range from 5% wt to 30% wt. 16. The membrane of claim 15, wherein the first polymer is present in a range from 7% wt to 20% wt. 17. The membrane of claim 1, wherein the second polymer is present in units of % wt. of the overall wt. of pol... | The Regents of the University of California | B01D053/22 B01D071/06 A61B005/00
US07879211 | B2 | 2011-02-01 | Analyzing instrument, lancet-integrated attachment for concentration measuring device provided with analyzing instrument, and body fluid sampling tool | 2002-07-11 | The present invention relates to an analyzing instrument (1A) which includes a capillary (31), a fluid feed port (20) for introducing a sample liquid to the capillary (31), and a fluid feed promoter (6) for promoting the introduction of the sample liquid into the feed port (20). The capillary (31) of the analyzer (1A) may be formed on a substrate (2) for example. The fluid feed promoter (6) may include at lease one of a water-absorbing layer having a higher water-absorbing capacity than the substrate (2) and an adhesive layer having a greater adhesion to the skin than the substrate (2). | 1. An analyzing instrument comprising: a flat cover plate having an air vent and a needle insertion port,a flat substrate having a fluid feed port for introducing a sample liquid, the substrate also having a reacting portion located downstream from the fluid feed port,a capillary formed between the substrate and the cover plate to extend from the fluid feed port and the needle insertion port toward the air vent for transporting the sample liquid to the reacting portion, anda fluid absorber fitted entirely within the fluid feed port of the substrate upstream from the capillary, the fluid absorber having a needle insertion opening,wherein the needle insertion opening of the fluid absorber is aligned with the needle insertion port of the cover plate for allowing insertion of a lancing needle through the needle insertion opening of the fluid absorber, the capillary and the needle insertion port of the cover plate. 2. The analyzing instrument according to claim 1, wherein the fluid absorber includes a water-absorbing layer having higher water-absorbing capacity than the substrate. 3. The analyzing instrument according to claim 1, wherein the fluid absorber includes an adhesive layer having greater adhesion to the skin than the adhesion that the substrate has to the skin. 4. The analyzing instrument according to claim 1, wherein the fluid absorber has higher elasticity than the substrate. 5. The analyzing instrument according to claim 1, further comprising a spacer interposed between the substrate and the cover plate, wherein the spacer is formed with a through-hole aligned with the liquid feed port of the substrate and the needle insertion opening of the cover plate. 6. The analyzing instrument according to claim 1, wherein the fluid absorber has a thickness equal to a thickness of the substrate. 7. The analyzing instrument according to claim 6, wherein the fluid absorber comprises a ring. 8. The analyzing instrument according to claim 1, wherein the fluid absorber is disposed around the fluid feed port. 9. A lancet-integrated attachment for mounting to a concentration measuring device, the attachment comprising a lancet and an analyzing instrument, wherein the analyzing instrument comprises a flat cover plate having an air vent and a needle insertion port, a flat substrate having a fluid feed port for introducing a sample liquid, the substrate also having a reacting portion located downstream from the fluid feed port, a capillary formed between the substrate and the cover plate to extend from the fluid feed port and the needle insertion port toward the air vent for transporting the sample liquid to the reacting portion, and a fluid absorber fitted entirely within the fluid feed port of the substrate upstream from the capillary, the fluid absorber having a needle insertion opening,wherein the needle insertion opening of the fluid absorber is aligned with the needle insertion port of the cover plate for allowing insertion of the lancet through ... | Arkray, Inc. | B01J019/12 C25D017/00 C12M001/00 C12N009/00 C12N011/00 C25B009/00 C25B011/00 C25B013/00 A61B005/00 A61B017/14 A61B017/32 B65D081/00 G01N001/00 G01N027/26 G01N031/00 G01N033/487
US20110028860 | A1 | 2011-02-03 | NEUROMONITORING SYSTEM WITH WIRELESS INSTRUMENTATION | 2010-07-29 | A method and apparatus is provided that allow the surgeon and the IONM specialist to continuously get feedback on the relative proximity of instrumentation to a nerve and the health of this nerve. By using a coupling device that allows every instrument to be electrically stimulated, the surgeon, through the instruments, continuously sends a small electrical input into the patient's body. Depending on the relative proximity between this electrical stimulation and a nerve, a neuromonitoring modality is recorded through electrodes placed on the patient. This modality is read and analyzed by the IONM specialist. The characteristics of this response allow the IONM specialist to determine the health status and or location of the nerves. | 1. A neurophysiological stimulator comprising: a source of electrical stimulation current;an electrically conductive instrument; anda medical glove made of an electrically insulated material where at least a portion of the outside surface of the glove has at least one electrically conductive open path and at least one electrical connection for connecting the said electrically conductive path to the said source of current; whereinthe electrically conductive path transmits electrical stimulation current to the said electrically conductive instrument when contact is made between the glove and the instrument. 2. An intra-operative neuromonitoring system for ensuring safety of a patient during a surgical procedure comprising: a source of electrical stimulation current;an electrically conductive surgical instrument;a surgical glove made of an electrically insulated material where at least a portion of the outside surface of the surgical glove has at least one electrically conductive open path and at least one electrical connection for connecting the said electrically conductive path to the said source of current; whereinthe electrically conductive path transmits electrical stimulation current to the said electrically conductive surgical instrument when contact is made between the glove and the instrument. 3. The system of claim 2, wherein the neuromonitoring system has sensing electrodes for recording electromyography activity in the patient's body. 4. The system of claim 2, wherein the neuromonitoring system has mechanical sensors for recording movements of the patient's body. 5. The system of claim 2, wherein the neuromonitoring system uses algorithms with pre-set threshold currents that can automatically determine the proximity of the surgical instrument to a nervous system. 6. The system of claim 2, wherein the neuromonitoring system has a wireless remote screen that can be placed anywhere in the operating room. 7. The system of claim 2, wherein the neuromonitoring system has a wireless connection with its main connection box where the sensing electrodes and the sources of electrical stimulation are connected to. 8. The system of claim 2, wherein the said source of current is transmitted to the glove electrical connection through an electrical pathway mounted on a surgical gown. 9. The system of claim 2, wherein the surgical instrument has an electrically conductive handle, an insulated electrically conductive structural elongated portion and an electrically conductive open end effector that transmit the said electrical stimulation current. 10. The system of claim 2, wherein the surgical instrument has an electrically non-conductive handle on which a single use thin and flexible electrically conductive sock is placed, an insulated electrically conductive structural elongated portion and an electrically conductive open end effector that transmit the said electrical stimulation current. 11. The system of claim 2, wherein the surgical glove ... | - | A61B005/05
US07878989 | B2 | 2011-02-01 | Knee prostheses | 2006-12-27 | A total knee prosthesis comprises a femoral component and a patellar component. The femoral component has a patellar flange into which is built a twist or roller coaster configuration. The patellar component has an off-set dome with respective lateral and medial facets. The configuration of the patellar flange resists forces from a medial to a lateral direction, whilst the lateral and medial facets of the patellar component are respectively shaped to allow area contact and line contact respectively with the femoral component. Also disclosed is a device for intra-operative checking of patellar tracking to allow correct insertion of the femoral and patellar components. | 1. A device for intra-operative checking of patellar tracking with a total knee replacement, the device comprising a first part having a polymeric patellar component with articular surfaces matching a femoral component of the knee replacement, and an extension member for registering with a second part of the device, the second part having a portion sized to a cut surface of a patellar and with means for engagement with the cut surface, and an extension member for registering with said first part of the device, relative sliding movement being possible, in use, between said first and second parts and wherein the extension members of the first and second parts are flexible and the extension member of the first part is in a form of a flexible wire which extends through a flexible tube of said second part. 2. A device as claimed in claim 1, wherein a scale on one of the first or second parts of the device indicates a relative degree of sliding between the first part and the second part. 3. A device as claimed in claim 2, wherein the scale is configured to indicate an amount of patellar mal-tracking when a knee is flexed through a whole range of movement while the patellar component of the first part is applied to the femoral component of the knee replacement residing within the knee. 4. A device as claimed in claim 2, wherein, a free end of the flexible wire registers with said scale on the second part to indicate the relative degree of sliding between the first part and the second part. 5. A device as claimed in claim 4, wherein said scale is formed on a reference portion at an end of the flexible tube remote from said portion of the second part sized to the cut surface of the patella. 6. A device as claimed in claim 5, wherein the reference portion defines a flat surface carrying said scale and having an open-topped channel therein, in which said free end of the flexible wire is slidably received. 7. A device as claimed in claim 4, wherein there is a flexible strap part between said portion of the second part sized to the cut surface of the patella and said flexible tube. 8. A device as claimed in claim 7, wherein the tube extends from a sleeve at one end of the strap. 9. A device as claimed in claim 1, wherein the means of the second part for engagement with the cut surface are a plurality of spikes extending from an opposite surface of the second part to that which effects relative sliding movement with the first part. 10. A device as claimed in claim 1, wherein the patellar component of the first part has a recess in which is received part of the portion of the second part sized to the cut surface of the patellar. 11. A device as claimed in claim 10, in which said recess has undercut sides. 12. A device as claimed in any claim 1, wherein the second part is of plastics material. 13. A device claimed in claim 1, wherein the second part is a metallic base plate part of the device. 14. A device as claimed in claim 4, wherein the device i... | - | A61B005/06 A61B005/103 A61B005/117 A61B017/58 A61B017/60 A61F005/00 A61F002/00 A61F002/28 A61F002/30 A61F002/38
US08290567 | B2 | 2012-10-16 | Cardiac diagnostics using wall motion and perfusion cardiac MRI imaging and systems for cardiac diagnostics | 2010-09-20 | Cardiac information of a patient is displayed by obtaining a plurality of MRI cine loops of the heart of the patient at a plurality of heart rates, the plurality of cine loops including both wall motion cine loops and at least one perfusion cine loops and simultaneously displaying both the wall motion cine loops and the at least one perfusion cine loop. | 1. A workstation for providing cardiac diagnostic data to a clinician during a patient stress test, comprising: a circuit in communication with an MRI System configured to generate: (i) a plurality of temporally synchronized wall motion cine loops of MRI cardiac images of a patient at different heart beat rates during the cardiac stress test including, (a) a first baseline wall motion cine loop, (b) a second wall motion cine loop obtained after administration of a contrast agent, and (c) a third recovery wall motion cine loop obtained after the first and second cine loops; (ii) at least one MRI perfusion cine loop of MRI perfusion cardiac images of the patient, wherein a frame rate of the at least one MRI perfusion cine loop is substantially the same as a frame rate of the wall motion cine loops, and wherein the wall motion cine loops and the at least one perfusion cine loop are co-registered so that images of respective cine loops correspond to a substantially common cardiac location of the patient; and (iii) at least one delayed enhancement MRI image of cardiac tissue of the patient; anda display in communication with the circuit for displaying the cine loops and the at least one delayed enhancement MRI image of the cardiac tissue of the patient,wherein the circuit is configured to provide the temporally synchronized wall motion cine loops, the at least one perfusion cine loop and the at least one delayed enhancement image to the display so that the wall motion cine loops, the at least one perfusion cine loop and the at least one delayed enhancement image are concurrently shown in different windows on the display during a cardiac stress test while a patient is undergoing the cardiac stress test while in the MRI System for image data acquisition. 2. The workstation of claim 1, wherein the display has a user interface that is configured to accept a user input to rapidly adjust at least one of contrast, brightness, or gamma of the baseline cine loop which is automatically propagated to cine loops of different windows on the display to thereby provide a clinician with patient information for safer administration of the cardiac stress test. 3. The workstation of claim 1, wherein the at least one cine loop of MRI perfusion images is a plurality of perfusion cine loops, one cine perfusion loop for each of a plurality of different locations in the heart, and wherein the perfusion cine loops are registered to the wall motion cine loops so that the MRI perfusion cine loops correspond to substantially the same cardiac locations as the wall motion cine loops. 4. The workstation of claim 3, wherein the plurality of perfusion cine loops correspond to three short axis wall motion cine loops. 5. A workstation in communication with an MRI System, comprising: a circuit configured to generate (a) a plurality of MRI cine loops of a heart of a patient at a plurality of heart rates, while the patient is undergoing a cardiac stress test in an MR Scanner, the plur... | Wake Forest University Health Sciences | A61B005/05
US07885703 | B2 | 2011-02-08 | System and method for scanning a patient | 2004-09-27 | Systems and methods for using a patient display is described. The described system includes a medical imaging system. The medical imaging system includes a gantry, a patient display, and a computer. The gantry at least partially circumscribes a patient viewing area. The patient display is viewable from the patient viewing area. The patient display is communicatively coupled to the computer. The computer is programmed to receive scan protocol instructions and control the patient display according to the predetermined scan protocol. It is further programmed to prompt the patient, using the patient display, to perform a bodily action that facilitates the scan and to provide the patient with at least one of entertainment and distraction using the patient display. | 1. A method of scanning a patient using a medical imaging device, said method comprising: receiving patient demographic data comprising an approximate mental age of the patient, wherein the approximate mental age of the patient is different than the intelligence quotient of the patient;positioning a display within a field of view of the patient;controlling the display according to a predetermined scan protocol and the received patient demographic data;selecting a display script based on the received patient demographic data and the scan protocol;prompting the patient, using the display script, to perform a bodily action that facilitates a scan, the patient prompted by receiving instructions indicative of the bodily action;providing the patient with at least one of entertainment and distraction using the display script; andgenerating an image of the patient from data acquired during the scan. 2. A method in accordance with claim 1 wherein positioning a display within a field of view of the patient comprises at least one of coupling the display to a medical imaging device gantry and positioning a movable display stand proximate the gantry. 3. A method in accordance with claim 2 wherein coupling the display to a medical imaging device gantry comprises at least one of magnetically coupling the display to the medical imaging device gantry and mechanically fastening the display to the medical imaging device gantry. 4. A method in accordance with claim 1 wherein providing the patient with at least one of entertainment and distraction using the display comprises providing the patient with at least one of entertainment and distraction during periods of patient agitation. 5. A method in accordance with claim 1 wherein receiving patient demographic data comprises receiving at least one of a native language of the patient, a chronological age of the patient, an entertainment preference of the patient, and a gender of the patient. 6. A method in accordance with claim 1 wherein prompting the patient to perform a bodily action that facilitates the scan using the display comprises prompting the patient to perform a bodily action using at least one of text, sound, and graphics through the display. 7. A method in accordance with claim 6 wherein prompting the patient to perform a bodily action using graphics comprises displaying the prompted bodily action using international symbols. 8. A method in accordance with claim 6 wherein prompting the patient to perform a bodily action using graphics comprises displaying the prompted bodily action using moving images. 9. A method in accordance with claim 8 wherein displaying the prompted bodily action using moving images comprises displaying the prompted bodily action using pre-recorded video. 10. A method in accordance with claim 8 wherein displaying the prompted bodily action using moving images comprises displaying the prompted bodily action using animation. 11. A method in accordance with claim 6 wherein prom... | General Electric Company | A61B005/05
US20110038801 | A1 | 2011-02-17 | METHODS AND COMPOSITIONS FOR IDENTIFYING LUNG CANCER OR A HUMORAL IMMUNE RESPONSE AGAINST LUNG CANCER | 2008-12-19 | Methods and compositions are provided for identifying lung cancer or a humoral immune response against lung cancer. Also disclosed are methods for determining whether a subject is responding or is likely to respond to lung cancer immunotherapy. | 1. A method for identifying whether a subject is afflicted with lung cancer, comprising detecting an immune response against an antigen identified in Table 2, 3 or 4, wherein detection of the immune response indicates that the subject is afflicted with lung cancer. 2. The method of claim 1, wherein the lung cancer is non-small cell lung cancer. 3. The method of claim 1, wherein the subject is a mammal. 4. The method of claim 1, wherein the subject is a human. 5. The method of claim 1, wherein the immune response is a humoral immune response. 6. The method of claim 1, wherein the immune response is a cellular immune response. 7. The method of claim 1, wherein an immune response is detected against 1, 2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14, or 15 of the antigens in Table 2, 3 or 4. 8. A method for determining whether a subject is likely to respond to lung cancer therapy with a composition comprising cancer cells that have been rendered proliferation-incompetent and have been genetically engineered to express GM-CSF, comprising detecting an immune response against an antigen listed in Table 2, 3 or 4, wherein detecting the immune response indicates that the subject is likely to respond to said lung cancer therapy. 9. The method of claim 8, wherein the lung cancer therapy is for the treatment of non-small cell lung cancer. 10. The method of claim 8, wherein the subject is a mammal. 11. The method of claim 8, wherein the subject is a human. 12. The method of claim 8, wherein the cancer cells are autologous. 13. The method of claim 8, wherein the cancer cells are allogeneic. 14. The method of claim 8, wherein an immune response is detected against one, two, three, four, five, six, seven, eight, nine, ten, eleven, twelve, thirteen, fourteen, fifteen, or more of the antigens listed in Table 2, 3 or 4. 15. The method of claim 8, wherein responsiveness to the cancer therapy is measured by decreased serum concentrations of tumor specific markers, increased overall survival time, increased progression-free survival, decreased tumor size, decreased metastasis marker response, increased impact on minimal residual disease, increased induction of antibody response to the cancer cells that have been rendered proliferation-incompetent, increased induction of delayed-type-hypersensitivity (DTH) response to injections of autologous tumor, increased induction of T cell response to autologous tumor or candidate tumor-associated antigens, increased impact on circulating T cell and dendritic cell numbers, phenotype, and function, cytokine response, reduced metastasis as measured by bone scan/MRI, increased time to progression, decreased serum concentrations of ICTP, decreased concentrations of serum C-reactive protein or decreased numbers of circulating tumor cells. 16. The method of claim 8, wherein responsiveness to the cancer therapy is measured by decreased serum concentrations of tumor specific markers. 17. The method of claim 8, wherein respons... | BioSante Pharmaceuticals, Inc. | A61B005/055 A61K049/00 C40B030/04 G01N033/53 C40B060/12 A61P035/00 A61P037/00 G06F019/00
US20110033607 | A1 | 2011-02-10 | AEROSOLIZED FINGERPRINT POWDER COMPOSITIONS | 2010-08-03 | An aerosolized fingerprint compositions for spraying onto a surface for the purpose of indentifying latent fingerprints that comprises about 1 to about 75% fingerprint dusting powder and about 25 to about 99% non-CFC propellant. | 1. An aerosolized fingerprint powder composition, comprising (by weight %): about 1 to about 75% fingerprint dusting powder; andabout 25 to about 99% non-CFC dry propellant. 2. The composition of claim 1, wherein at least about 75% of the fingerprint powder has a particle size of about 30 microns or less. 3. The composition of claim 1, wherein at least about 95% of the fingerprint powder has a particle size of about 50 microns or less. 4. The composition of claim 1, in combination with a pressurized container, wherein the fingerprint powder and the propellant are stored in the can under pressure. 5. The composition of claim 1, wherein the fingerprint powder is a metallic fingerprint powder. 6. The composition of claim 1, consisting essentially of the propellant and the fingerprint powder. 7. The composition of claim 1, comprising about 1 to about 50 weight % fingerprint powder. 8. The composition of claim 1, comprising about 50 to about 96 weight % propellant. 9. The composition of claim 1, wherein the non-CFC propellant is a hydrocarbon propellant. 10. The composition of claim 9, wherein the hydrocarbon propellant is selected from the group consisting of isobutene, butane, and mixtures thereof 11. The composition of claim 1, with a ratio of about 1 to 10, powder to propellant. 12. The composition of claim 1, with a ratio of from about 1 to 5, powder to propellant to about 1 to 25, powder to propellant. 13. A method of obtaining fingerprints, comprising: (1) providing an aerosolized spray fingerprint powder composition, comprising a fingerprint powder and a dry propellent;(2) providing an isolation device that comprises a small opening defined by at least one wall and a large opening defined by at least one wall;(3) identifying a surface that may contain a latent fingerprint;(4) placing the large opening of the isolation device over the surface;(5) spraying the aerosolized fingerprint powder into the small opening of the isolation device to actualize latent print; and(6) analyzing actualized print. 14. The method of claim 13, wherein the isolation device is conical. 15. The method of claim 13, wherein after the spraying step, the small opening of the isolation device is covered to allow the particles to swirl around the print and attach to the print. 16. A kit, comprising: at least one aerosolized container that comprises a fingerprint formulation and a propellant;a finger print brush;an isolation device that is defined by at least one wall and has a fingerprint surface opening and a spray opening. 17. The kit of claim 16, wherein the isolation device is foldable into a flat form. | Enamelite, LLC | A61B005/117 C09D004/00 B65D071/00
US20110015500 | A1 | 2011-01-20 | METHOD FOR DETERMINING CUSTOMIZED ESSENTIAL OIL BLEND | 2010-07-16 | A scientific method and system for determining a customized essential oil blend through accessing a physiolocial data from a user are disclosed in the present invention. The physiological data is heart rate vairiability (HRV), which is associated with cardiac diseases and the autonomic nervous system. | 1. A method for determining a customized essential oil blend, comprising (a) measuring a heart rate variability (HRV) of a user; and(b) generating a formulation of a customized essential oil blend according to an HRV-essential oil database and the HRV of the user. 2. The method of claim 1, which further comprises a step of inputting a personal profile. 3. The method of claim 2, wherein the personal profile comprises at least one of subjective feeling about essential oils, personal information, and physiological condition. 4. The method of claim 3, wherein the physiological condition comprises at least one of medical history, an allergic essential oil, and sleep time. 5. The method of claim 3, wherein the personal information comprises at least one of age, gender, and race. 6. The method of claim 1, wherein the HRV is measured by electrocardiography, photoplethysmograph, or by using a microphone. 7. The method of claim 1, wherein the heart rate variability comprises at least one of high frequency (HF), low frequency (LF), very low frequency (VLF), low frequency/high frequency (LF/HF), total power (TP), stress index (SI), standard deviation of the NN interval (SDNN), and the square root of the mean of the sum of the squares of differences between adjacent NN intervals (RMSSD). 8. The method of claim 1, wherein the customized essential oil blend comprises at least one of a pure essential oil and a commercial essential oil blend collection. 9. A system for determining a customized essential oil blend, comprising (a) a heart rate variability (HRV) monitor, which accesses HRV parameters of a user;(b) an HRV-essential oil database; and(c) a data processing device for determining the customized essential oil blend. 10. The system of claim 9, wherein the heart rate variability monitor is a electrocardiography monitor, a photoplethysmograph monitor, or an acoustic blood flow detector. 11. The system of claim 9, wherein the data processing device is used for analyzing signals accessed by the HRV monitor. 12. The system of claim 9, wherein physiological signals of the user are the HRV parameters. 13. The system of claim 9, wherein the HRV parameters comprises at least one of high frequency (HF), low frequency (LF), very low frequency (VLF), low frequency/high frequency (LF/HF), total power (TP), stress index (SI), standard deviation of the NN interval (SDNN), and the square root of the mean of the sum of the squares of differences between adjacent NN intervals (RMSSD). 14. The system of claim 9, wherein the data processing device is a computer, a notebook computer, a mobile phone, a console, or a personal digital assistant. 15. The system of claim 9, wherein the data processing device further comprises a data processing algorithm which calculates the customized essential oil blend based on the HRV parameters, a personal profile and the HRV-essential oil database. 16. The system of claim 15, wherein the personal profile comprises age, gender, r... | SUSTINEO BIOTECHNOLOGY CO., LTD. | A61B005/00 A61B005/02 A61B005/0402 A61B005/0295 A61B005/026
US07866212 | B2 | 2011-01-11 | High-sensitivity sensors for sensing various physiological phenomena, particularly useful in anti-snoring apparatus and methods | 2005-06-09 | A mechanical vibration sensor adapted to be brought into contact with an object for sensing mechanical vibrations in the object, includes a body of a soft elastomeric material having high transmissivity and low attenuation properties with respect to a preselected type of energy waves; and a pair of transducers mounted, by mounting members having high attenuation properties with respect to the energy waves, in spaced relationship to each other to define a transmission channel between the transducers. Such sensor is particularly useful in a method and apparatus for controlling snoring by a person, by utilizing a stimulus device effective, when sensing snoring, to immediately produce a response in the person tending to interrupt the person's snoring. | 1. Apparatus for controlling snoring or other predetermined body condition by a person, comprising: a sensor for sensing said predetermined body condition of the person while sleeping; and for outputting a sensor signal corresponding to the sensed condition; a stimulus device effective, when actuated, to immediately produce a response in the person tending to interrupt said sensed body condition; and a processor for processing said outputted sensor signal to determine whether it indicates the presence of said sensed body condition, and if so, for actuating said stimulus device to produce said response tending to interrupt said sensed body condition; characterized in that said sensor is a mechanical vibrations sensor and is included between, and in contact with, a pair of plates of larger size than said sensor for sensing mechanical vibrations in an external part of the person's body and for transmitting said mechanical vibrations to said sensor for generating and outputting said sensor signal. 2. The apparatus according to claim 1, wherein said processor: analyzes said sensor signal for frequency components within a predetermined frequency band;accumulates the energy of the sensor signal within said predetermined frequency band for a preselected time interval;compares said accumulated energy for each predetermined time interval with the average of the accumulated energy in a plurality of the preceding preselected time intervals;and controls the actuation of said stimulus device as a result of such comparison. 3. The apparatus according to claim 2, wherein said processor controls the actuation of said stimulus device in accordance with a determination of whether or not the accumulated energy in the preselected time interval exceeds by a predetermined percentage the average of the accumulated energy in said plurality of preceding preselected time intervals. 4. The apparatus according to claim 3, wherein said processor actuates said stimulus device when the accumulated energy in each of at least two successive preselected time intervals exceeds by said predetermined percentage the average of the accumulated energy in the plurality of preceding preselected time intervals. 5. The apparatus according to claim 3, wherein said apparatus includes a manual control for varying said predetermined percentage, and thereby for varying the sensitivity of the apparatus. 6. The apparatus according to claim 1, wherein said preselected time interval is substantially less than the inhale and exhale half-cycle times of the person's respiration cycle time such that the actuation of the stimulus device is closely time-coordinated with the sensed body condition during the person's respiration cycle. 7. The apparatus according to claim 1, wherein said sensor senses said mechanical vibrations in the part of the person's body by causing said mechanical vibrations sensor to change the transit time of an acoustic wave transmitted from a transmitter to a receiver in an a... | Nexense Ltd. | G01N029/00 A61B005/08
US07892179 | B2 | 2011-02-22 | Apparatus, method and system for determining a physiological condition within a mammal | 2010-05-12 | A system, method and apparatus for determining a physiological condition within a mammal based on pH and/or temperature measurements. The invention is for determining vaginal health and/or fertility status in a female mammal based on pH and/or temperature measurements of vaginal and/or cervical fluids or tissue. The invention is suited for taking either in vivo or external measurements. In one aspect, the invention is an apparatus comprising an elongated probe portion and a handle portion. In one embodiment, the circuitry of the apparatus is strategically split between the probe portion and the handle portion to facilitate a plug-and-play type of device. In another embodiment, an elongated probe is provided that has a hermetically sealed internal cavity that houses the circuitry while leaving the temperature sensor and the pH sensor exposed for direct contact with the desired biological fluid. In a further embodiment, an ISFET is used as the pH sensor. | 1. An apparatus for measuring a physiological condition within a mammal comprising: a probe for insertion into a body lumen of the mammal, the probe comprising: an elongated tubular housing having a first internal cavity;a first circuit board located within the first internal cavity;a pH sensor for measuring pH within the body lumen located at a distal portion of the elongated tubular housing, the pH sensor generating a pH signal indicative of the measured pH and operably coupled to the first circuit board;a memory device storing parametric data unique to the pH sensor, the memory device located within the first internal cavity and operably coupled to the first circuit board;a temperature sensor for measuring temperature within the body lumen and generating a temperature signal indicative of the measured temperature, the temperature sensor located at the distal portion of the elongated tubular housing and operably coupled to the first circuit board; anda first interface connector located at a proximal portion of the elongated tubular housing and operably coupled to the first circuit board;a handle for manipulating tile probe, the handle comprising: a second housing having a second internal cavity;a second circuit board located within the second internal cavity;a microprocessor for processing the pH signal and generating an output signal based on the processing of the pH signal, the microprocessor located within the second internal cavity and operably coupled to the second circuit board;a display device on the second housing for displaying the output signal generated by the microprocessor and operably coupled to the second circuit board; anda second interface connector located at the second housing and operably coupled to the second circuit board;the probe connected to the handle in a manner that allows the probe and the handle to be repetitively engaged and disengaged from each other;wherein when the probe is connected to the handle, the first and second interface connectors are in electrical connection so that the microprocessor retrieves the parametric data from the memory device of the probe and receives the pH signal;wherein the microprocessor processes the temperature signal and generates the output signal based on the processing of the pH signal and the temperature signal; andwherein the output signal represents the measured physiological condition and comprises vaginal health, cervical health, and/or fertility status. 2. The apparatus of claim 1 further comprising: wherein the first internal cavity is separated into a first internal chamber and a second internal chamber by a transverse wall;wherein the first circuit board is located within the first internal chamber and wherein the second chamber is filled with an electrolyte solution buffered at a known pH and hermetically sealed; andwherein the probe further comprises a diaphragm located at the distal portion of the elongated tubular housing having a first portion in contact with the el... | PH Diagnostics, Inc. | A61B010/00 A61B005/103 A61B005/117 A61B005/05 A61B005/00 B65D081/00
US20110015518 | A1 | 2011-01-20 | METHOD AND INSTRUMENT FOR SURGICAL NAVIGATION | 2010-07-07 | An operating microscope including an optical unit for forming an image of an object plane in oculars of the microscope, an optoelectronic image receiver coupled to the microscope and optics to form images, of objects placed in a region between a front objective of the microscope and an object plane of the microscope, on the optoelectronic image receiver. The microscope has a magnification factor of the optics to form images, and a system to detect optical markings forming a markings pattern placed on a surgical instrument or an object placed in the region between the front objective of the microscope and the object plane of the microscope. The system calculates a geometrical position and orientation of the markings pattern in relation to the optoelectronic image receiver, relative to the microscope. | 1. An operating microscope, comprising: detecting means for detecting a topography of a situs relative to the microscope; andtransmitting means for transmitting the topography of the situs to a navigational system, the navigational system referencing and displaying the topography within location-correlated preoperative diagnostic datasets of the situs. 2. The microscope according to claim 2, further comprising: a photonic mixer device (PMD), as the detecting means, to detect the topography of the situs relative to the microscope; anda modulated illumination device associated with the PMD to provide modulated illumination light. 3. The microscope according to claim 2, wherein the PMD is arranged such that a sensor axis of the PMD is coaxial or substantially coaxial to the optics. 4. The microscope according to claim 2, further comprising: optical filters positioned to protect the PMD from intensive illumination by unmodulated light. 5. The microscope according to claim 2, wherein the PMD detects optical markings forming a markings pattern placed on a surgical instrument or an object placed in a region between a front objective of the microscope and the object plane of the microscope. 6. The microscope according to claim 2, wherein a front objective and the PMD are positioned at a common distance from the situs. 7. The microscope according to claim 1, wherein the detecting means includes a stereo camera pair with a computing system to perform stereoscopic analysis. 8. The microscope according to claim 7, further comprising: means for projecting a light pattern onto the situs, the light pattern captured by the stereo camera pair for improved correspondence analysis of the stereoscopic images. 9. The microscope according to claim 1, wherein the detecting means is a single camera with a computing system to perform stereoscopic analysis by operating the single camera as an inverse camera, the operating microscope further comprising: means for projecting a light pattern onto the situs. 10. The microscope according to claim 1, wherein the topography is transmitted to the navigational system, the navigational system correlating the topography as a starting point for correction of a brain shift. 11. The microscope according to claim 1, wherein a focus of the operating microscope is adjusted based on the topography. 12. The microscope according to claim 1, wherein a position of a surgical instrument or a pointer is extracted from the topography relative the detecting means, respective the microscope, andthe position of the surgical instrument or the pointer is transmitted to the navigational system for referencing or displaying the position of the surgical instrument or the pointer in the topography within the location-correlated preoperative diagnostic datasets of the situs. 13. An operating microscope, comprising: optical means for forming an image of an object plane in oculars of the microscope;an optoelectronic image receiver coupled to the ... | - | A61B005/00
US08894933 | B2 | 2014-11-25 | Portable handheld medical diagnostic devices | 2009-08-06 | A portable handheld medical diagnostic device includes a front housing and a rear housing opposite the front housing. The front housing and the rear housing form a protective enclosure. A main circuit board is located in the protective enclosure. The main circuit board includes a controller facilitating a physiologic measurement. A display device is connected to the main circuit board that displays information related to the physiologic measurement. A frame is located in the protective enclosure that carries the display device and locates the display device adjacent the front housing such that the display device can be viewed from outside the protective enclosure. The frame includes a strip port formed integrally therewith that is accessible from outside the protective enclosure. | 1. A portable handheld medical diagnostic device, comprising: a front housing;a rear housing opposite the front housing, the front housing and the rear housing forming a protective enclosure;a main circuit board having printed circuits connecting electronic components located in the protective enclosure, the main circuit board including a controller facilitating a physiologic measurement;an electronic paper display device laminated on a display board connected to the main circuit board that displays information related to the physiologic measurement;touch sensors located on the display board; anda single-piece frame located in the protective enclosure that carries the display device and locates the display device adjacent the front housing such that the display device can be viewed from outside the protective enclosure, the frame being located between the front housing and the main circuit board, the frame including a strip port formed integrally therewith that is accessible from outside the protective enclosure. 2. The portable handheld medical diagnostic device of claim 1, wherein the frame and strip port are molded together. 3. The portable handheld medical diagnostic device of claim 1, wherein the main circuit board includes an opening that at least partially receives a battery located in the protective enclosure such that a portion of the battery is located at opposite sides of the main circuit board. 4. The portable handheld medical diagnostic device of claim 3 further comprising a battery contact comprising a spring arm arranged and configured to move the battery from an insertion position to a lock position. 5. The portable handheld medical diagnostic device of claim 4, wherein movement of the battery from the insertion position to the lock position moves the battery beneath overhang structure. 6. The portable handheld medical diagnostic device of claim 1, wherein the frame is sized and configured to align the strip port with a strip reader carried by the main circuit board. 7. The portable handheld medical diagnostic device of claim 1 further comprising an audio component carried by the frame. 8. The portable handheld medical diagnostic device of claim 1 further comprising a wireless communication board carried by the frame, the wireless communication board configured to connect to an external device. 9. The portable handheld medical diagnostic device of claim 1 wherein the touch sensors comprise, a button formed by a deflectable component mounted within an opening through the front housing, the deflectable component configured to move relative to the front housing; anda button actuation assembly located within the protective enclosure, the button actuation assembly including a membrane that spans the opening. 10. The portable handheld medical device of claim 9, wherein the deflectable component is formed of a material that is harder than material forming the membrane. 11. The portable handheld medical diagnostic device of clai... | Roche Diagnostics Operations, Inc. | G01N033/00 A61B005/00 G06F003/0488 A61B005/145 G01N033/487
US08891851 | B2 | 2014-11-18 | Home healthcare management system and hardware | 2010-07-13 | A home healthcare management system wherein a patient conducts self diagnoses and self testing, and manages their own medical records at home. A digital microscope is utilized as part of the system that is smaller, of lower cost, faster, of a higher dynamic range, and has a higher resolution than conventional microscopes. The microscope consists of an illumination source, a spatial sub-sampling device and a detector device. The digital microscope provides a vectored method of collecting images from a digital microscope that is independent of the optical resolution, and a slide based coordinate system, and a method of displaying images and communicating such images over the Internet in a file format that does not require a header or prior knowledge of magnification, coordinate system, or tiling structure. The system further includes an interface for physiological monitoring devices and a connection to the Internet for more comprehensive services. | 1. A method for self-diagnostic testing of a patient's whole blood sample comprising; placing a whole blood sample on a slide, wherein said slide includes colorimetric indicators and an imaging area for constituents;placing said slide into a digital microscope having a digital camera wherein said digital microscope performs morphological and colorimetric functionality by said attached digitalcamera, wherein said digital camera includes x, y, and z software controlled translation means;scanning, by said digital camera, the whole blood sample deposed on said slide andcapturing, by said digital camera, multiple images from different locations wherein said locations incorporates nuclei;analyzing said multiple images, by a processor, for high optical density nuclei and constituents using contrast detection algorithms;determining, by said processor, diagnostic information based on said analysis of said high optical density nuclei and said constituents; anddisplaying said diagnostic information. | Glenn F. Spaulding | G06K009/00 A61B005/0205 G02B021/00 G06Q050/22 G06F019/00 A61B005/00
US08647285 | B2 | 2014-02-11 | Fluid sample collection system | 2010-10-15 | A sample collection container adapted to be connected to, and disconnected from, a sample collection adapter and a sample collection container cap. The sample collection adapter has a downward sloping interior surface that provides a free-flowing path for samples provided by a subject to the sample collection container. The sample collection adapter may have a three-dimensional rim that is curved in two different non-parallel planes. | 1. An oral sample collection system comprising: An ergonomic oral fluid sample funnel for receiving sample, wherein said ergonomic funnel is characterized by:an effective funnel opening width of between 25 and 80 millimeters, a minor axis of between 20 and 60 millimeters, a major axis of between 40 and 120 millimeters, wherein said minor axis is less than said major axis and bounded by a sampling edge adapted to be contacted by a user's lip and a privacy edge located opposite of said sampling edge, configured to provide privacy when the user expectorates into said funnel, a sampling rim height of between 10 and 80 millimeters associated with said sampling edge, an overall height of between 35 and 100 millimeters associated with said privacy edge, wherein said sampling rim height is less than said overall height and is the minimum height of the funnel rim, a funnel front slope angle of between 15 degrees and 70 degrees, and a funnel opening angle and a privacy angle wherein said funnel opening angle is not equal to said privacy angle;A threaded sample collection container adapted to be connected to, and disconnected from, threads located on said ergonomic sample funnel; andA threaded sample collection container cap adapted to be connected to and disconnected from said sample collection container. 2. The sample collection system of claim 1 further comprising a closure liner for providing a seal at an opening of said sample container. 3. The sample collection system of claim 2 wherein said closure liner and said sample collection container are sorption neutral. 4. The sample collection system of claim 3 further comprising a compression disk for backing up said closure liner. 5. The sample collection system of claim 4 wherein said sample collection container cap comprises a hole for enabling sample extraction. 6. The sample collection system of claim 1 further comprising a label, said label comprising a minimum fill marking and a maximum fill marking. 7. The sample collection system of claim 1 further comprising a sheath for substantially surrounding and protecting said sample collection container. 8. The sample collection system of claim 7 further comprising a window in said sheath for permitting viewing the interior of said sample collection container. 9. The sample collection system of claim 1 wherein said ergonomic sample funnel comprises a thermoplastic polymer. 10. The sample collection system of claim 1 wherein said sample collection container comprises a glass. 11. An oral sample collection system comprising: An ergonomic oral fluid sample funnel for receiving a said oral fluid sample, wherein said oral fluid ergonomic funnel is characterized by: an effective funnel opening width of between 30 and 50 millimeters, a minor axis of between 25 and 35 millimeters, a major axis of between 50 and 90 millimeters, a funnel front slope angle of between 35 degrees and 55 degrees, an upper planar rim with a single curved cutout adapted for cont... | Biophor Diagnostics, Inc. | A61B005/00 B65D081/00
US09084819 | B2 | 2015-07-21 | Magnetic microstructures for magnetic resonance imaging | 2010-04-02 | The present invention relates to a magnetic resonance structure with a cavity or a reserved space that provides contrast and the additional ability to frequency-shift the spectral signature of the NMR-susceptible nuclei such as water protons by a discrete and controllable characteristic frequency shift that is unique to each MRS design. The invention also relates to nearly uniform solid magnetic resonance T2* contrast agents that have a significantly higher magnetic moment compared to similarly-sized existing MRI contrast agents. | 1. A magnetic resonance contrast agent comprising: a plurality of solid disk-shaped magnetic resonance structures, the magnetic resonance structures being of uniform size and magnetic moment and each comprising a pair of disks comprising magnetic material, wherein the disks are separated via an internal non-magnetic spacer, wherein each magnetic resonance structure has a volume of 5×10−22 m3 to 5×10−15 m3, and wherein the magnetic resonance structures are dispersed in a fluid medium. 2. The magnetic resonance contrast agent of claim 1, wherein each magnetic resonance structure of the plurality of magnetic resonance structures has a magnetic moment from about 10−14 A·m2 to about 10−11 A·m2. 3. The magnetic resonance contrast agent of claim 1, wherein each magnetic resonance structure of the plurality of magnetic resonance structures varies in size by less than about 5% of the average size of the plurality of magnetic resonance structures. 4. The magnetic resonance contrast agent of claim 3, wherein each magnetic resonance structure of the plurality of magnetic resonance structures varies in magnetic moment by less than about 5% of the average magnetic moment of the plurality of magnetic resonance structures. 5. The magnetic resonance contrast agent of claim 1, wherein the magnetic material comprises a ferromagnetic material, paramagnetic material, superparamagnetic material, magnetic alloy material, magnetic compound, or any combination thereof. 6. The magnetic resonance contrast agent of claim 5, wherein the magnetic resonance structures further comprise a coating selected from the group consisting of an oxidation barrier, a corrosion barrier, a mechanical strengthening layer, a spatial buffer coating to inhibit the clumping of magnetic resonance structures, a non-toxic coating, a biologically inert coating, a coating to facilitate bioconjugation, a cell-specific antibody and ligand coating, and combinations thereof. 7. The magnetic resonance contrast agent of claim 1, wherein each disk has a diameter ranging from about 0.1 μm to about 10 μm and a thickness ranging from about 0.1 μm to about 10 μm. | The United States of America, as Represented by the Secretary, Department of Health and Human Services, Office of Technology Transfer, National Institutes of Health THE UNITED STATES OF AMERICA, REPRESENTED BY THE SECRETARY OF COMMERCE | A61K049/18 A61B005/055 B82Y005/00 B82Y030/00 G01R033/56 A61B005/05 G01R033/563
US08942817 | B2 | 2015-01-27 | Broadcasting a signal indicative of a disease, disorder, or symptom determined in response to contactlessly acquired information | 2009-07-28 | In an embodiment, a system includes a sensor device operable to detect a property of a mammal without a direct physical contact with the mammal. The system also includes a patient assessment device operable to evaluate the detected property of the mammal for an indicator of a disease. The system further includes a patient information device including a computer-readable medium, and configured to maintain and to provide access to information corresponding to the indication of the disease. The system further includes a transmitter device operable to broadcast a signal indicative of the disease. | 1. A system comprising: a sensor device operable to detect a property of a mammal without a direct physical contact with the mammal;a patient assessment device operable to evaluate the detected property of the mammal for an indicator of a disease, disorder, or symptom potentially treatable by a neuromodulation treatment regimen;a patient information device including a computer-readable medium and configured to maintain and to provide access to information corresponding to the indicator of the disease, disorder, or symptom potentially treatable by the neuromodulation treatment regimen;a transmitter device operable to broadcast a signal indicative of the disease, disorder, or symptom potentially treatable by the neuromodulation treatment regimen; anda neuromodulation device programmed to administer the neuromodulation treatment regimen to the mammal in response to the signal from the transmitter device. 2. The system of claim 1, wherein the sensor device operable to detect a property of a mammal without a direct contact with the mammal includes: a sensor device operable to detect a property of a mammal absent a direct conductive, electrical, or physical contact with the mammal. 3. The system of claim 1, wherein the sensor device operable to detect a property of a mammal without a direct contact with the mammal includes: a sensor device operable to at least one of electronically, sonically, acoustically, thermally, radiatively, or chemically detect a property of a mammal without a direct contact with the mammal. 4. The system of claim 1, wherein the patient assessment device includes: a patient assessment device physically distinct and spatially apart from the sensor device. 5. The system of claim 1, wherein the disease, disorder, or symptom includes at least one of a neurological disease, nervous system disorder, heart rhythm, or heart disease. 6. The system of claim 1, wherein the patient assessment device operable to evaluate the detected property of the mammal for the indicator of the disease, disorder, or symptom potentially treatable by the neuromodulation treatment regimen includes: a patient assessment device operable to receive a signal indicative of the detected property of the mammal, and to evaluate the detected property of the mammal for the indicator of the disease, disorder, or symptom potentially treatable by the neuromodulation treatment regimen. 7. The system of claim 1, wherein the patient assessment device operable to evaluate the detected property of the mammal for the indicator of a disease, disorder, or symptom potentially treatable by the neuromodulation treatment regimen includes: a patient assessment device operable to evaluate the detected property of the mammal for a physiological characteristic of the mammal, and to identify the disease, disorder, or symptom potentially treatable by the neuromodulation treatment regimen in response to the evaluated physiological characteristic of the mammal. 8. The system of claim... | The Invention Science Fund I, LLC | A61N001/37 A61B005/00 A61N001/372 A61N001/36 G06F019/00 A61N001/08
US20110046504 | A1 | 2011-02-24 | DISTRIBUTED NEURO-RESPONSE DATA COLLECTION AND ANALYSIS | 2010-07-29 | Distributed mechanisms are provided for collecting neuro-response data from subjects exposed to the stimulus material in multiple settings. Stimulus material may include marketing and entertainment materials. Neuro-response data collection mechanisms such as Electroencephalography (EEG) and Electrooculography (EOG) are used to collect data from subjects in laboratory and corporate settings. Neuro-response data is transmitted over a network to a data analyzer. The neuro-response data is processed at the data analyzer and effectiveness data for the stimulus material is received. | 1. A method, comprising: receiving neuro-response data from a plurality of neuro-response data collection mechanisms associated with a plurality of users exposed to marketing and entertainment materials, the neuro-response data collection mechanisms comprising a plurality of electroencephalography (EEG) electrodes, wherein the neuro-response data is collected in a plurality of disparate settings and received over a network;analyzing the neuro-response data to determine the effectiveness of the marketing and entertainment materials that the plurality of users are exposed to in the plurality of disparate settings;sending effectiveness data to the plurality of disparate settings. 2. The method of claim 1, wherein analyzing the neuro-response data comprises identifying evidence of the occurrence or non-occurrence of specific time domain difference event-related potential (DERP) components. 3. The method of claim 2, wherein neuro-response data is synchronized with stimulus material. 4. The method of claim 1, wherein neuro-response data is received from distributed neuro-response data collection mechanisms associated with the plurality of users in a plurality of geographic locations. 5. The method of claim 1, wherein the plurality of disparate settings comprise a plurality of corporate settings. 6. The method of claim 1, wherein the plurality of disparate settings comprise a plurality of laboratory settings. 7. The method of claim 1, wherein the neuro-response data collection mechanism further comprises electrooculography (EOG) sensors. 8. The method of claim 1, wherein the plurality of EEG electrodes are EEG dry electrodes. 9. A system, comprising: an interface configured to receive neuro-response data from a plurality of neuro-response data collection mechanisms associated with a plurality of users exposed to marketing and entertainment materials, the neuro-response data collection mechanism comprising a plurality of electroencephalography (EEG) electrodes, wherein the neuro-response data is collected in a plurality of disparate settings and received over a network;a processor configured to analyze the neuro-response data to determine the effectiveness of the marketing and entertainment materials that the plurality of users are exposed to in the plurality of disparate settings;wherein effectiveness data is sent to the plurality of disparate settings. 10. The system of claim 9, wherein analyzing the neuro-response data comprises identifying evidence of the occurrence or non-occurrence of specific time domain difference event-related potential (DERP) components. 11. The system of claim 10, wherein neuro-response data is synchronized with stimulus material. 12. The system of claim 9, wherein neuro-response data is received from distributed neuro-response data collection mechanisms associated with the plurality of users in a plurality of geographic locations. 13. The system of claim 9, wherein the plurality of disparate settings comprise a pl... | NEUROFOCUS, INC. | A61B005/0476 A61B005/0496
US20110046502 | A1 | 2011-02-24 | DISTRIBUTED NEURO-RESPONSE DATA COLLECTION AND ANALYSIS | 2009-08-20 | Distributed mechanisms are provided for identifying stimulus material and collecting neuro-response data from subjects exposed to the stimulus material in a variety of environments. Neuro-response data collection mechanisms such as Electroencephalography (EEG), Electrooculography (EOG), eye tracking, and/or facial emotion encoding are used to collect data from subjects in environments such as home, work, and recreational settings. Cameras, audio recorders, video recorders, processors, etc., identify stimulus material the subject is exposed to. The data collected may be minimally or extensively processed prior to transmission to a neuro-response data analyzer. | 1. A method, comprising: receiving neuro-response data from a neuro-response data collection mechanism associated with a user, the neuro-response data collection mechanism including electroencephalography (EEG) electrodes;identifying video data in the frame of view of the user, the video data including marketing and entertainment materials observed by the user in non-laboratory settings;identifying marketing and entertainment materials included in the video data;analyzing the neuro-response data synchronized with the video data to determine the effectiveness of the marketing and entertainment materials that the user is exposed to in non-laboratory settings. 2. The method of claim 1, wherein identifying video data comprises receiving video data from a camera included in the neuro-response data collection mechanism. 3. The method of claim 1, wherein neuro-response data is received from distributed neuro-response data collection mechanisms associated with a plurality of users in a plurality of geographic locations. 4. The method of claim 1, wherein the neuro-response data collection mechanism is shaped in the form of a baseball cap. 5. The method of claim 1, wherein the neuro-response data collection mechanism comprises EEG electrodes and a video camera. 6. The method of claim 1, wherein the neuro-response data collection mechanism comprises EOG sensors and an audio recorder. 7. The method of claim 1, wherein the neuro-response data collection mechanism comprises EEG electrodes, a video camera, persistent storage, and a network interface integrated in a cap configured for wear in non-laboratory settings. 8. The method of claim 7, wherein the neuro-response data collection mechanism comprises an externally oriented video camera recording data in the frame of view of the user and an internally oriented video camera recording facial emotion expressions and pupillary dilation. 9. The method of claim 1, wherein video data recording is triggered by detection of salient neuro-responses. 10. The method of claim 1, wherein determining the effectiveness of marketing and entertainment materials comprises measuring attention, emotion, and memory retention. 11. A neuro-response data collection mechanism, comprising: a plurality of electroencephalography (EEG) dry electrodes arranged on an interior surface of a neuro-response data collection mechanism;a camera having a lens arranged on an exterior surface of the neuro-response collection mechanism, the camera configured to capture data in the frame of view of a user wearing the neuro-response data collection mechanism;a network interface configured to send data collected by the plurality of EEG dry electrodes and the camera to a data analyzer. 12. The neuro-response data collection mechanisms of claim 11, further comprising electrooculography (EOG) sensors. 13. The neuro-response data collection mechanisms of claim 11, further comprising a second camera having a lens arranged to detect facial expressi... | NeuroFocus, Inc. | A61B005/0476 G06Q010/00 H04N007/18 A61B005/103
US20110032104 | A1 | 2011-02-10 | MEDICINE MANAGEMENT APPARATUS FOR MANAGING PATIENT'S MEDICATION PERIOD THROUGH MESSAGE TRANSMISSION | 2009-11-20 | A medicine management apparatus comprises a medicine storage casing for storing medicines, a display unit for displaying medication times for the medicines and temperature and humidity within the medicine storage casing, an audio unit for informing the medication times on a periodic basis, a timer for setting the medication times on a periodic basis and outputting the medication times through the audio unit or the display unit, a message transmission module for transmitting a message to an administrator over a wireless communication network when there is no open signal detected when doors provided in the medicine storage casing are opened using the open signal as a medication signal, and a control unit for performing an overall control operation. | 1. A medicine management apparatus for managing a patient's medication period through message transmission, the apparatus comprising: a medicine storage casing for storing medicines;a display unit for displaying medication times for the medicines and temperature and humidity within the medicine storage casing;an audio unit for informing the medication times on a periodic basis;a timer for setting the medication times on a periodic basis and outputting the medication times through the audio unit or the display unit;a message transmission module for transmitting a message to an administrator over a wireless communication network when there is no open signal detected when doors provided in the medicine storage casing are opened using the open signal as a medication signal; anda control unit for performing an overall control operation of the medicine management apparatus. 2. The medicine management apparatus as claimed in claim 1, further comprising a wired communication module for transmitting an emergency signal to a management server over a wired communication network when there is no open signal using the open signal as the medication signal. 3. The medicine management apparatus as claimed in claim 1, wherein the control unit controls the audio unit or the display unit such that the audio unit or the display unit outputs a medication notification signal again based on the period set in the timer in response to a periodic signal when there is no open signal. 4. The medicine management apparatus as claimed in claim 3, wherein the message transmission module and the wired communication module transmit the message and an emergency signal to the administrator and a management server, respectively, when there is no open signal although the medication notification signal is outputted again by a predetermined number of times in response to the periodic signal. 5. The medicine management apparatus as claimed in claim 1, wherein: the medicine storage casing is partitioned into a hot storage room, a cold storage room, and a room-temperature storage room with a barrier rib interposed therebetween, and temperature and humidity within each of the storage rooms are controlled by the control unit. 6. The medicine management apparatus as claimed in claim 5, wherein the timer comprises a first timer, a second timer, and a third timer configured to set the respective medication times of the medicines stored in the hot storage room, the cold storage room, and the room-temperature storage room. 7. The medicine management apparatus as claimed in claim 1, further comprising a blood sugar measuring instrument for enabling a diabetic patient himself to measure his blood sugar. 8. The medicine management apparatus as claimed in claim 7, wherein a blood sugar value measured by the blood sugar measuring instrument is transmitted to the administrator and the management server through the message transmission module and the wired communication module, respectively. 9.... | - | G08B023/00 A61B005/15 G04B047/00
US20110034788 | A1 | 2011-02-10 | METHODS AND APPARATUS FOR USING MULTIPLE SENSORS TO MEASURE DIFFERENTIAL BLOOD TRANSPORT TIME IN A PATIENT | 2009-08-04 | The difference in time that it takes for blood to flow to two locations in a patient's body can be measured by placing, on the patient's body, sensors for non-invasively detecting a variable characteristic of the content of the blood at each of those locations. The time of occurrence of a change in that characteristic as detected by one of the sensors is compared to the time of occurrence of that same change as detected by the other sensor. The difference between these two times is a measure of the difference in time that it takes for blood to flow to the locations of the two sensors. | 1. A method of measuring differential blood transport time in a patient comprising: placing a first sensor adjacent a first location on the patient's body;placing a second sensor adjacent a second location on the patient's body;operating the first sensor to detect a variable characteristic of the patient's blood at the first location;operating the second sensor to detect the variable characteristic of the patient's blood at the second location; anddetermining elapsed time between when a change in the variable characteristic is detected by the first and second sensors at the first and second locations, respectively. 2. The method defined in claim 1 wherein each of the sensors comprises an oximetry sensor. 3. The method defined in claim 1 wherein the variable characteristic comprises SxO2. 4. The method defined in claim 1 wherein the first sensor is placed adjacent the patient's forehead. 5. The method defined in claim 4 wherein the second sensor is placed adjacent an extremity of the patient other than the forehead. 6. The method defined in claim 1 wherein each of the first and second sensors comprises a non-invasive blood characteristic sensor. 7. The method defined in claim 1 wherein the characteristic is a characteristic of the content of the patient's blood. 8. The method defined in claim 1 wherein each of the first and second sensors produces a respective one of first and second electrical output signals indicative of variation of said characteristic as detected by the respective one of said sensors. 9. The method defined in claim 8 wherein the determining comprises: electrically processing said first and second output signals to facilitate determination of said elapsed time. 10. Apparatus for measuring differential blood transport time in a patient comprising: a first sensor adapted for placement adjacent a first location on a patient's body and operable to produce a first electrical output signal indicative of a variable characteristic of the patient's blood at the first location;a second sensor adapted for placement adjacent a second location on the patient's body and operable to produce a second electrical output signal indicative of the variable characteristic of the patient's blood at the second location; andcircuitry for outputting indications of the first and second signals in a way that enables determination of elapsed time between when a change in the variable characteristic is detected by each of the first and second sensors. 11. The apparatus defined in claim 10 wherein each of the sensors comprises an oximetry sensor. 12. The apparatus defined in claim 10 wherein the variable characteristic comprises SxO2. 13. The apparatus defined in claim 10 wherein the first sensor is adapted for placement adjacent to the patient's forehead. 14. The apparatus defined in claim 13 wherein the second sensor is adapted for placement adjacent an extremity of the patient other than the forehead. 15. The apparatus defined in claim 10 wh... | Nellcor Puritan Bennett LLC | A61B005/1455 A61B005/02
US07883463 | B2 | 2011-02-08 | Combination sensor for physiological parameters | 2004-09-28 | A compact combination sensor is provided for physiological parameters, which can be worn behind the ear. The combination sensor has a heat flux sensor (1) for detecting two temperatures (Th1, Th2) for determining the body core temperature (Tc), an acoustic sensor (2) for measuring the heart rate (HR) and for detecting speech signals, as well as a measured signal evaluating unit (4, 5, 7). A communications interface (8) is provided for transmitting the processed data from the measured signal evaluating unit (4, 5, 7) to a communications unit (9) are located in the sensor carrier (11). | 1. A combination sensor for physiological parameters, the combination sensor comprising: a communications unit;a sensor carrier;a heat flux sensor for detecting a first temperature and a second temperature, said first temperature being the temperature at or below the skin of a person and said second temperature being the temperature of the environment;an acoustic sensor for measuring the heart rate and for detecting speech signals;a first measured signal evaluating unit and a second measured signal evaluating unit, said first measured signal evaluating unit determining a body core temperature using said first detected temperature and said second detected temperature; anda communications interface for transmitting the processed data from said first measured signal evaluating unit to said communications unit, said heat flux sensor, said acoustic sensor, said first measured signal evaluating unit and said second measured signal evaluating unit and said communications interface being each located at closely spaced locations in said sensor carrier to form a compact arrangement of compact sensors, wherein said sensor carrier is embedded in a pliable plastic material and is adapted to the shape of the human ear. 2. A combination sensor in accordance with claim 1, wherein an additional measured signal evaluating unit is provided for calculating a characteristic number for the physiological stress from the body core temperature and the heart rate and for transmission to the communications unit. 3. A combination sensor in accordance with claim 1, wherein said communications unit sends signals to a receiving unit, which is equipped with a downstream display unit for displaying the heart rate, the body core temperature and/or a characteristic number for the physiological stress. 4. A combination sensor in accordance with claim 1, wherein said sensor carrier has a loudspeaker or alarm which is connected with the communications interface and/or with the measured signal evaluating unit and can be activated when said processed data deviates from a preset limit value. 5. A combination sensor in accordance with claim 1, wherein said data transmission between the communications interface and the communications unit is wireless transmission. 6. A combination sensor in accordance with claim 1, wherein said heat flux sensor is shaped for being in contact with skin of the person. 7. A combination sensor for physiological parameters, the combination sensor comprising: a communications unit;a sensor carrier;a temperature sensor for detecting two temperatures at or near the user and for the determination of a body core temperature;an acoustic sensor for measuring the heart rate of the user and for detecting speech signals of the user;a measured signal evaluating means for receiving signals from said temperature sensor and said acoustic sensor and providing one or more processed data signals;a loudspeaker connected with the communications interface and/or with the me... | Dräger Safety AG & Co. KGaA | A61B005/00
US07865235 | B2 | 2011-01-04 | Method and system for detecting and classifying the mental state of a subject | 2005-09-12 | A method of detecting and classifying mental states, comprising the steps of: detecting bio-signals from one or more than one bio-signal detector; transforming the bio-signals into one or more than one different representations; detecting values of one or more than one property of the transformed bio-signal representations; and applying one or more than one mental state detection algorithm to the detected properties in order to classify whether the bio-signals indicate the presence of a predetermined response by a subject. | 1. A method of detecting and classifying mental states, comprising the steps of: detecting bio-signals from a bio-signal detector, the bio-signals including at least electroencephalograph (EEG) signals from a subject;using a processing device, transforming the electroencephalograph signals into a plurality of different representations, including separating the electroencephalograph signals into different epochs and generating representations of the electroencephalograph signal epochs in a plurality of different frequency domain representations, a plurality of different temporal domain representations and a plurality of different spatial domain representations;detecting values of a plurality of different properties of each of the plurality of different frequency domain representations, plurality of different temporal domain representations and plurality of different spatial domain representations; andapplying a mental state detection algorithm to the detected values of the properties in order to classify whether the electroencephalograph signals indicate the presence of a predetermined mental state of the subject. 2. The method according to claim 1, wherein the different frequency domain representations are obtained by dividing each electroencephalograph signal epoch into distinguishable frequency bands. 3. A method of detecting and classifying mental states, comprising the steps of: detecting bio-signals from a bio-signal detector, the bio-signals including at least electroencephalograph (EEG) signals from a subject;using a processing device, transforming the electroencephalograph signals into a plurality of different representations, including separating the bio-signals into different epochs, generating representations of the electroencephalograph signal epochs into a plurality of different domains, the different domains including different temporal domain representations, wherein the different temporal domain representations are obtained by dividing each electroencephalograph signal epoch into a plurality of time segments;detecting values of properties of the transformed representations; andapplying one or more than one mental state detection algorithm to the detected values of the properties in order to classify whether the electroencephalograph signals indicate the presence of a predetermined mental state of the subject. 4. The method according to claim 3, wherein the plurality of time segments in each epoch are temporally overlapping. 5. A method of detecting and classifying mental states, comprising the steps of: detecting bio-signals from a bio-signal detector, the bio-signals including at least electroencephalograph (EEG) signals from a subject;using a processing device, transforming the electroencephalograph signals into a plurality of different representations, including separating the electroencephalograph signals into different epochs, generating representations of the electroencephalograph signal epochs into a plurality of differ... | - | A61B005/04
US08335559 | B2 | 2012-12-18 | ECG rhythm advisory method | 2010-10-20 | A method of automatically determining which type of treatment is most appropriate for a cardiac arrest victim, the method comprising transforming one or more time domain electrocardiogram (ECG) signals into a frequency domain representation comprising a plurality of discrete frequency bands, combining the discrete frequency bands into a plurality of analysis bands, wherein there are fewer analysis bands than discrete frequency bands, determining the content of the analysis bands, and determining the type of treatment based on the content of the analysis bands. | 1. A method of automatically determining which type of treatment is most appropriate for a patient, the method comprising: transforming a time domain electrocardiogram (ECG) measurements signal measured from the patient into frequency domain data;processing the frequency domain data to determine the variance of the energy distributed across a selected frequency range; anddetermining the type of treatment based on the determined variance. 2. The method of claim 1 further comprising providing a prompt to a user of the result of the determination of which type of treatment is most appropriate. 3. The method of claim 1 wherein determining the type of treatment based on the determined variance comprises comparing the determined variance to a threshold. 4. The method of claim 1 wherein determining the type of treatment based on the determined variance comprises combining the determined variance with one or more of the following in a multidimensional space: a number of peaks in the selected frequency range;relative values of multiple peaks in the selected frequency range; andrelative bandwidths of multiple peaks in the selected frequency range. 5. The method of claim 1 wherein the selected frequency range excludes a low-frequency portion in which a chest compression artifact dominates. | ZOLL Medical Corporation | A61B005/04
US07867170 | B2 | 2011-01-11 | Method and apparatus for non-invasively measuring hemodynamic parameters using parametrics | 2004-05-03 | An improved method and apparatus for non-invasively assessing one or more hemodynamic parameters associated with the circulatory system of a living organism. In one aspect, the invention comprises a method of measuring a hemodynamic parameter (e.g., arterial blood pressure) by applanating or compressing portions of tissue proximate to the blood vessel of concern until a desired condition is achieved, and then measuring the hemodynamic parameter. Such applanation effectively mitigates transfer and other losses created by the tissue proximate to the blood vessel, thereby facilitating accurate and robust tonometric measurement. An algorithm adapted to maintain optimal levels of applanation is also described. Methods and apparatus for scaling such hemodynamic parameter measurements based on subject physiology, and providing treatment to the subject based on the measured parameters, are also disclosed. | 1. Apparatus adapted to measure the blood pressure of a living subject, comprising: a pressure sensor adapted to generate a signal relating to the pressure applied to a blood vessel of said subject, said applied pressure being applied at least in part by tissue interposed between said sensor and a blood vessel of said subject;a bias element disposed proximate to said pressure sensor and adapted to mitigate the effects of said tissue on said applied pressure; anda digital processor configured to scale said signal using a scaling factor, said scaling factor comprising said subject's weight divided by said subject's height squared. 2. The apparatus of claim 1, wherein said bias element substantially envelops at least a portion of said pressure sensor. 3. The apparatus of claim 2, further comprising applanation apparatus adapted to controllably press at least a portion of said pressure sensor and said bias element against said tissue. 4. The apparatus of claim 1, further comprising a user interface operative to receive said data relating to said subject's weight and said subject's height from a user. 5. The apparatus of claim 1, wherein said scaling factor comprises a wrist circumference. 6. The apparatus of claim 1, wherein said living subject comprises a human being, and said bias element uses at least said pressure sensor to at least partly applanate the radial artery of said human being, said at least partial applanation being performed while said signal is being generated. 7. A method of measuring the blood pressure of a living subject, comprising: utilizing a pressure sensor to applanate a blood vessel of said living subject;utilizing said pressure sensor to measure a plurality of pressure values of said applanated blood vessel;determining a first physiological parameter of said living subject;determining a BMI of said living subject;deriving a scaling factor based at least in part on said first physiological parameter and said BMI;scaling said plurality of pressure values using said scaling factor, said act of scaling comprising using at least one algorithm run on a processor associated with said pressure sensor; andutilizing said scaled pressure values to generate a blood pressure value for presentation to a user. 8. The method of claim 7, wherein said first physiological parameter comprises pulse pressure (PP). 9. The method of claim 7, wherein said first physiological parameter comprises wrist circumference (WC). 10. The method of claim 7, wherein said act of determining said BMI of said living subject comprises utilizing one or more inputs from a user. 11. The method of claim 7, wherein said living subject comprises a human being, and said act of applanating is performed substantially contemporaneously with said act of measuring said plurality of pressure values. 12. Apparatus for measuring the blood pressure of a living subject, comprising: a first means for generating a signal relating to the pressure in a blood vessel of said... | Tensys Medical, Inc. | A61B005/02
US07873398 | B2 | 2011-01-18 | Saccharide-measuring fluorescent monomer, saccharide-measuring fluorescent sensor substance, and implantable, saccharide-measuring sensor | 2008-08-12 | A saccharide-measuring fluorescence sensor substance including a copolymer of at least the following two compounds (I) and (II) is provided: (I) a fluorescent monomer compound represented by the following formula (1): formula (1); wherein: Q, Q′ and D3 may be the same or different, may be combined together into a fused ring, and are each a substituent selected from the group consisting of a hydrogen atom, a halogen atom, a hydroxyl group, and substituted or unsubstituted alkyl, acyl, oxyalkyl, carboxyl, carboxylate ester, carboxamido, cyano, nitro, amino and aminoalkyl groups; andD1, D2 and D4 each represent a substituent, wherein at least one of D1, D2 and D4 is a substituent group comprising a vinyl group at an end thereof, and wherein the substituent group comprising a vinyl group at an end thereof enables the fluorescent monomer compound to be soluble in water, and(II) at least one hydrophilic, polymerizable monomer having a vinyl group. | 1. A saccharide-measuring fluorescence sensor substance comprising a copolymer of at least the following two compounds (I) and (II): (I) a fluorescent monomer compound represented by the following formula (1):wherein: Q, Q′ and D3 may be the same or different, may be combined together into a fused ring, and are each a substituent selected from the group consisting of a hydrogen atom, a halogen atom, a hydroxyl group, and substituted or unsubstituted alkyl, acyl, oxyalkyl, carboxyl, carboxylate ester, carboxamido, cyano, nitro, amino and aminoalkyl groups;D1, D2 and D4 each represent a substituent, wherein at least one of D1, D2 and D4 is a substituent group comprising a vinyl group at an end thereof, and wherein the substituent group comprising a vinyl group at an end thereof enables the fluorescent monomer compound to be soluble in water, and (II) at least one hydrophilic, polymerizable monomer comprising a vinyl group, wherein said polymerizable monomer (II) is a compound represented by the following formula (6):wherein A is a hydrogen atom or a methyl group, and U and U′ may be the same or different and are each a substituted or unsubstituted alkyl group. 2. A saccharide-measuring fluorescence sensor substance comprising a copolymer of at least the following two compounds (I) and (II): (I) a fluorescent monomer compound represented by the following formula (1):wherein: Q and Q′ may be the same or different, may be combined together into a fused ring, and are each a substituent selected from the group consisting of a hydrogen atom, a halogen atom, a hydroxyl group, and substituted or unsubstituted alkyl, acyl, oxyalkyl, carboxyl, carboxylate ester, carboxamido, cyano, nitro, amino and aminoalkyl groups;and in D1, D2, D3 and D4,(i) D1 and D2 may be the same or different and are each a substituted or unsubstituted alkyl group, D3 is a hydrogen atom, and D4 is a substituent group represented by the following formula (2), or(ii) D1 and D2 may be the same or different and are each a substituent group represented by the following formula (3), and D3 and D4 may be the same or different, may be combined together into a fused ring, and are each a substituent selected from the group consisting of a hydrogen atom, a halogen atom, a hydroxyl group, and substituted or unsubstituted alkyl, acyl, oxyalkyl, carboxyl, carboxylate ester, carboxamido, cyano, nitro, amino and aminoalkyl groups:wherein: X, in a case of the definition (i), is a substituent group selected from the group consisting of —COO—, —OCO—, —CH2NR—, —CH2S—, —CH2O—, —NR—, —NRCO—, —CONR—, —SO2NR—, —NRSO2—, —O—, —S—, —SS—, —NRCOO—, —OCONR— and —CO—, and in a case of the definition (ii), represents a C1 to C30 alkylene group comprising at least one substituent group selected from the group consisting of —COO—, —OCO—, —CH2NR—, —NR—, —NRCO—, —CONR—, —SO2NR—, —NRSO2—, —O—, —S—, —SS—, —NRCOO—, —OCONR— and —CO—, in which R represents a hydrogen atom or a substituted or unsubstituted alkyl group,Y is a... | Terumo Kabushiki Kaisha | A61B005/00
US07865230 | B1 | 2011-01-04 | Method and system for detecting sentinel lymph nodes | 2003-07-11 | According to one embodiment of the present invention, a method for detecting lymph nodes in a human includes introducing a fluorescent contrast agent into a lymph node system of a body, directing near-infrared time-varying excitation light into the tissue of the body, causing the near-infrared time-varying excitation light to contact a lymph node of the lymphatic system, whereby a redshifted and time-varying emission light is generated, detecting the time-varying emission light at a surface of the body, filtering the time-varying emission light to reject excitation light re-emitted from the lymph node, and imaging the lymph node of the lymphatic system. | 1. Method for detecting lymph nodes in a human, comprising: introducing a fluorescent contrast agent into a lymphatic system of a body;directing time-varying excitation light into the tissue of the body;causing the time-varying excitation light to contact a lymph node of the lymphatic system, whereby a time-varying emission light is generated, wherein the generated time-varying excitation light comprises a spatial distribution selected from the group consisting of a planar wave, a series of lines of illumination, concentric circles of illumination, and a ronchi rule pattern;detecting the time-varying emission light at a surface of the body;using frequency-domain photon migration, filtering light that is not time-varying;filtering the time-varying emission light to reject excitation light re-emitted from the lymph node; andimaging the lymph node of the lymphatic system. 2. The method of claim 1, wherein introducing the fluorescent contrast agent comprises introducing indocyanine green (“ICG”). 3. The method of claim 1, wherein directing time-varying excitation light into the tissue of the body comprises directing time-varying excitation light into the tissue of the body with a light source selected from the group consisting of a pulse, a series of pulses, pseudo random modulation, sinusoidally modulated light, and a square wave. 4. The method of claim 1, wherein the spatial distribution is the ronchi rule pattern. 5. The method of claim 1, further comprising modulating the excitation light to obtain a wavelength between approximately 700 nm and 900 nm. 6. The method of claim 1, wherein causing the excitation light to contact a lymph node of the lymphatic system comprises causing the excitation light to contact a sentinel lymph node of the lymphatic system. 7. A system for detecting lymph nodes in a human, comprising: a laser diode operable to direct time-varying near-infrared excitation light into the tissue of a body;one or more optical components adapted to provide the time-varying near-infrared excitation light a spatial distribution selected from the group consisting of a planar wave, a series of lines of illumination, concentric circles of illumination, and a ronchi rule pattern;an image intensifier operable to detect, at a surface of the body, a redshifted and time-varying emission light generated by the near-infrared time-varying excitation light contacting a lymph node of the lymphatic system;one or more optical filters operable to reject excitation light re-emitted from the lymph node; andan imaging apparatus operable to image the lymph node of the lymphatic system. 8. The system of claim 7, further comprising a fluorescent contrast agent adapted to be injected into a lymphatic system of a body, the fluorescent contrast agent selected from the group consisting of a non-specific fluorescent contrast agent and a specific fluorescent contrast agent. 9. The system of claim 7, wherein the spatial distribution is the ronchi rule pattern... | Texas A&M University System | A61B005/00
US07879057 | B2 | 2011-02-01 | Method of setting lancing member to lancing device, lancing device and cam mechanism | 2003-07-01 | A lancing member (2) is integrally provided with a cap (29) for covering a needle (21). Upon or after being attached to a holder (5) of a lancing device (A), the lancing member (2) is rotated relative to a cap (29) by a rotator (C). As a result, the boundary (28) between the cap (29) and the lancing member (2) is broken easily to expose the needle (21). | 1. A cam mechanism for rotating a lancet relative to a cap, the lancet including a body and a lancing needle projecting from the body, the cap being initially integral with the body of the lancet and subsequently removable from the body by rotating the cap relative to the lancet, the cam mechanism comprising: a cylindrical housing separate from the cap;a movable member contained in the housing for moving reciprocally in first and second directions parallel to a longitudinal axis of the housing, the movable member holding the lancet and configured to allow the lancet to rotate relative to the cap that covers the lancing element of the lancet;a first groove inclined relative to the axis of the housing and a second groove connected to the first groove and extending linearly in parallel to the axis of the housing, the first and the second grooves being provided at one of the housing and the movable member; anda protrusion provided at the other one of the housing and the movable member, the protrusion extending into the first groove and the second groove;wherein the protrusion moves in the first groove when the movable member moves in the first direction, and moves in the second groove when the movable member moves in the second direction. 2. A lancing device comprising: a cylindrical housing;a lancet holder that holds a lancet including a body and a lancing needle projecting from the body, the lancing needle being covered by a cap that is separate from the housing and initially integral with the body of the lancet, the lancet holder being reciprocally movable along a longitudinal axis of the housing;a moving mechanism for advancing the lancet holder along the longitudinal axis of the housing; anda cam mechanism for converting a retracting movement of the lancet holder along the longitudinal axis of the housing into rotation of the lancet holder and the lancet relative to the cap for removing the cap from the body of the lancet. 3. The lancing device according to claim 2, wherein the lancet holder holds the lancet in a manner such that the lancet holder and the lancet are not rotatable relative to each other, wherein the cam mechanism causes the lancet holder to rotate together with the lancet. 4. The lancing device according to claim 2, wherein the cam mechanism includes a first groove which is provided on the housing and is inclined relative to the longitudinal axis of the housing, and a protrusion which is provided on the lancet holder and is fitted in the first groove. 5. The lancing device according to claim 4, wherein the cam mechanism further includes a second groove connected to the first groove and extending in parallel to the longitudinal axis of the housing, and wherein the protrusion passes through the second groove when the lancet holder advances. 6. The lancing device according to claim 2, further comprising a holding portion for removably holding an analyzer used for analyzing a sample taken by a piercing process. 7. The lancing d... | ARKRAY, Inc. | A61B017/14 A61B017/32 A61B017/34 A61B005/00 B65D081/00
US20110004091 | A1 | 2011-01-06 | Methods for Using Resonant Acoustic and/or Resonant Acousto-EM Energy to Detect And/Or Effect Structures | 2010-09-17 | The present invention makes use of resonant acoustic and/or acousto-EM energy applied to inorganic or biologic structures for the detection and/or identification, and for augmentation and/or disruption of function within the biologic structure. In particular, the invention provides a method of generating resonant acoustic and/or acousto-EM energy in biologic structures such as virus, bacteria, fungi, worms and tumors for the detection and disruption of these structures. Moreover, the invention provides a method of augmenting functions of biologic structures such as bone through the generation of resonant acoustic and/or acousto-EM energy in the structure. Systems are also provided for the generation and detection of resonant acoustic and/or resonant acousto-EM energy. | 1. A method for detecting a biologic structure comprising utilizing acousto-EM spectroscopy. 2. A method for detecting a biologic structure comprising: a) inducing acoustic resonance in the structure; andb) detecting at least one electromagnetic energy pattern of the structure. 3. A method for detecting a biologic structure comprising: a) inducing acoustic resonance in the structure with resonant acousto-EM energy; andb) detecting at least one acoustic signature of the structure. 4. The method according to claim 2, further comprising comparing said detected at least one electromagnetic energy pattern of the structure with at least one reference signature. 5. The method according to claim 3, further comprising comparing said detected at least one acoustic signature of the structure with at least one reference signature. 6. The method according to claim 2, wherein said at least one electromagnetic energy pattern comprises at least one acousto-EM signature which is produced by at least one source selected from the group of sources consisting of acoustic energy and electromagnetic (EM) properties and/or fields. 7. The method according to claim 2, wherein acoustic resonance is induced by utilizing at least one energy selected from the group consisting of acoustic energy including at least one resonant acoustic frequency of the structure, electromagnetic energy which is substantially equivalent to at least one resonant acoustic frequency of the structure, electromagnetic energy which is substantially equivalent to at least a portion of at least one acousto-EM signature of the structure and electromagnetic energy which is substantially equivalent to at least one substantially complete acousto-EM signature of the structure. 8. A system for identifying a structure by determining at least one resonant acoustic signature of the structure comprising: a) electromagnetic means for inducing acoustic resonance in the structure;b) means for detecting at least one acoustic signature of the structure; andc) means for comparing said at least one detected acoustic signature of the structure with at least one reference acoustic signature of the structure. 9. The system according to claim 8, further comprising a means for detecting at least one acousto-EM signature of the structure. 10. The system according to claim 8, wherein said structure comprises at least one member selected from the group consisting of inorganic and biologic structures. 11. The method of claim 2, wherein said biologic structure comprises at least one member selected from the group consisting of virus, bacteria, fungi, tissue masses, worms, arthropods, chitins, plants, animals, microorganisms, multicellular organisms, protozoa, liver, muscle, feet, brain, kidney, spleen, blood, lung, lens of eye, aqueous humor, vitreous humor, animal cell, plant cell, proteins, molecules, cell wall, capsule, spore, pili, plasma membrane, organ, portions of structures, components of structures flagellum, ... | GR Intellectual Reserve, LLC | A61B005/00 G01N029/22
US07889899 | B2 | 2011-02-15 | Directional statistics visualization tool | 2007-01-22 | A computer-implemented method for visualization of diffusion tensor images includes providing a diffusion tensor image input and providing a volume of interest within the diffusion tensor image input. The method includes determining a plurality of direction-based classifications of the volume of interest, wherein classes are defined by a set of parameters, An optimal solution is then selected within the classifications by using a criterion defined as a ratio of inertia indicators. To represent the chosen classification, a cone graph is determined for each of the directional classes to be displayed or stored, each cone pair being the geometrical interpretation of the class parameters. The method further includes determining a spherical scatterplot of the volume of interest augmented with a cone graph for visualization of at least one of the directional classes, and displaying and/or storing the scatterplot. | 1. A computer-implemented method for visualization of diffusion tensor images comprising: providing a diffusion tensor image input produced by a magnetic resonance imaging device, the diffusion tensor image input corresponding to a brain;providing a volume of interest within the diffusion tensor image input;determining a plurality of directional classes of the volume of interest, wherein the directional classes are defined by a set of parameters;determining an optimal number of the directional classes using a criterion defined as a ratio of inertia indicators;determining a spherical scatterplot of the volume of interest augmented with a cone graph for visualization of at least one of the directional classes; anddisplaying the scatterplot. 2. The computer-implemented method of claim 1, wherein the volume of interest is a set of voxels selected by a user. 3. The computer-implemented method of claim 2, further comprising determining a principal diffusion direction for each voxel. 4. The computer-implemented method of claim 3, further comprising plotting each voxel and its principal diffusion direction on the spherical scatterplot indicating a direction. 5. The computer-implemented method of claim 3, further comprising plotting each voxel and its principal diffusion direction on the spherical scatterplot indicating a weight according to a fractional anisotropy of a respective voxel. 6. The computer-implemented method of claim 3, wherein each voxel and its principal diffusion direction are plotted on the spherical scatterplot, the method further comprising: indicating a direction by color; andindicating a fractional anisotropy by opacity. 7. The computer-implemented method of claim 1, further wherein each cone graph is a geometrical interpretation of the set of parameters defining a respective directional class. 8. A non-transitory computer readable medium embodying an application program executable by a processor to perform a method for visualization of diffusion tensor images, the method comprising: providing a diffusion tensor image input;providing a volume of interest within the diffusion tensor image input;determining a plurality of directional classes of the volume of interest, wherein the directional classes are defined by a set of parameters;determining an optimal number of the directional classes using a criterion defined as a ratio of inertia indicators; anddetermining a spherical scatterplot of the volume of interest augmented with a cone graph for visualization of at least one of the directional classes. 9. The method of claim 8, wherein the volume of interest is a set of voxels selected by a user. 10. The method of claim 9, further comprising determining a principal diffusion direction for each voxel. 11. The method of claim 10, further comprising plotting each voxel and its principal diffusion direction on the spherical scatterplot indicating a direction. 12. The method of claim 10, further comprising plotting each voxel and i... | Siemens Medical Solutions USA, Inc. | G06K009/00 G06K009/46 G06K009/62 G01V003/00 A61B005/05
US07892233 | B2 | 2011-02-22 | Magnetically navigable telescoping catheter and method of navigating telescoping catheter | 2006-11-28 | A magnetically navigable catheter includes a sheath having a proximal end and a distal end, and an extension member having a proximal end and a distal end, slidably mounted in the sheath so that the distal end portion of the extension member telescopes from the distal end of the sheath. The distal end portion of the extension member being relatively more flexible than the distal end of the sheath. There may be one or more electrodes on the distal end of the extension member. There is also at least one magnet on the distal end portion of the extension member to allow the distal end of extension member to be oriented by the application of an externally applied magnetic field. According to the method of this invention, the distal end of the electrode catheter is introduced into the part of the body where the electrode will be used to contact the specific body structures, and the electrode is moved into contact with the body structure by applying an external magnetic field and selectively telescoping the extension member relative to the sheath to bring the electrode on the distal end of the extension member into contact with the specific body structure. | 1. A magnetically navigable electrode catheter comprising: a sheath having a proximal end portion and a distal end portion;a flexible tube having a proximal end and a distal end slidably mounted in and telescoping from the sheath, and a lumen extending from the proximal end outside of a subject's body to the distal end, for delivery of fluids therethrough;a first electrode at the distal end of the flexible tube having at least one passage therethrough for delivery of fluids to the distal end of the flexible tube;a second electrode spaced proximally from the first electrode; andat least one magnet at the distal end portion of the flexible tube that is configured to orient the distal end portion of the flexible tube by the application of an externally applied magnetic field, the magnet having a passage therethrough for permitting delivery of fluids to the distal end of the flexible tube, wherein the at least one magnet is disposed on the distal end portion of the flexible tube such that a flex point is formed in between the at least one magnet and the distal end portion of the sheath so that the flexible tube bends at the flex point between the at least one magnet and the distal end of the sheath, andwherein the distal end of the flexible tube is adapted to be slidably mounted in the sheath so that the distal end portion telescopes from the distal end of the sheath. 2. The electrode catheter according to claim 1 wherein there are a plurality of magnets on the distal end portion of the flexible tube. 3. The electrode catheter according to claim 1 wherein the second electrode on the distal end portion of the flexible tube has at least one passage therethrough for delivery of fluids to the distal tip of the catheter. 4. The electrode catheter according to claim 1 wherein the second electrode on the distal end portion of the flexible tube is electrically insulated from the first electrode. 5. The electrode catheter according to claim 1 wherein the distal end portion of the flexible tube comprises a hollow flexible tube and a plurality of magnets disposed in the hollow tube. 6. The electrode catheter according to claim 5 wherein the plurality of magnets are relatively closely spaced within the flexible tube, and are spaced proximally from the at least one electrode on the distal end, to form a second flex point in the flexible tube between the magnets and the at least one electrodes. 7. The electrode catheter according to claim 5 wherein the plurality of magnets are relatively closely spaced within the hollow flexible tube, and wherein the tube can telescope out of a the sheath beyond the most proximal of the magnets, to form the flex point in the flexible tube between the magnets and the sheath. 8. The catheter of claim 1, wherein the distal end portion of the flexible tube containing the at least one magnet is relatively rigid such that the application of a magnetic field causes the flexible tube to bend at a point proximal to the at least one ... | Stereotaxis, Inc. | A61B018/14 A61B005/04
US07890154 | B2 | 2011-02-15 | Selection of ensemble averaging weights for a pulse oximeter based on signal quality metrics | 2008-12-03 | A method and a system for ensemble averaging signals in a pulse oximeter, including receiving first and second electromagnetic radiation signals from a blood perfused tissue portion corresponding to two different wavelengths of light, obtaining an assessment of the signal quality of the electromagnetic signals, selecting weights for an ensemble averager using the assessment of signal quality, and ensemble averaging the electromagnetic signals using the ensemble averager. | 1. A method of ensemble averaging signals in a pulse oximetry system, comprising: receiving, via a sensor, first and second electromagnetic radiation signals from a blood perfused tissue portion corresponding to two different wavelengths of light; andusing an oximeter, obtaining an assessment of the signal quality of the electromagnetic signals;calculating weights for an ensemble averager from a continuous weighting function based at least in part upon the assessment of signal quality; andensemble averaging the electromagnetic signals using the ensemble averager. 2. The method of claim 1, wherein the obtaining an assessment of the signal quality comprises obtaining a measure of a degree of arrhythmia of the signals. 3. The method of claim 2, wherein the obtaining an assessment of the signal quality further comprises obtaining a measure of a degree of similarity or correlation between the first and second signals. 4. The method of claim 1, wherein the obtaining an assessment of the signal quality comprises obtaining a measure of a degree of motion artifact present in the signals. 5. The method of claim 4, wherein the obtaining a measure of the degree of motion artifact comprises obtaining a ratio of a current pulse amplitude to a long-term average pulse amplitude of the signals. 6. The method of claim 1, wherein the obtaining an assessment of the signal quality comprises obtaining a ratio of a current pulse amplitude to a previous pulse amplitude of the signals. 7. The method of claim 1, wherein the obtaining an assessment of the signal quality comprises obtaining a measure of a degree of the overall signal quality metric for a single pulse, which comprises a combination of other metrics. 8. The method of claim 1, wherein the obtaining an assessment of the signal quality comprises obtaining a ratio of a current pulse period to an average pulse period of the signals. 9. The method of claim 1, wherein the calculating weights comprises forming a combination of one or more parameters comprising a measure of a degree of arrhythmia of the signals, a measure of the degree of similarity or correlation between the first and second electromagnetic radiation signals, a measure of a degree of motion artifact by obtaining a ratio of a current pulse amplitude to a long-term average pulse amplitude of the signals, a ratio of a current pulse amplitude to a previous pulse amplitude of the signals, and/or a ratio of a current pulse period to an average pulse period of the signals. 10. A pulse oximetry system, comprising: a sensor configured to receive first and second electromagnetic signals from a blood perfused tissue portion corresponding to two different wavelengths of light; anda pulse oximeter configured to: obtain an assessment of a signal quality of the first and second electromagnetic signals;calculate a first weight associated with the first and second electromagnetic signals for an oxygen saturation ensemble averager and/or calculate a second we... | Nellcor Puritan Bennett LLC | A61B005/1455
US08346354 | B2 | 2013-01-01 | Determining a neuromodulation treatment regimen in response to contactlessly acquired information | 2009-07-28 | Described embodiments include a system, an apparatus, and a method. A described system includes a sensor device configured to sense a property of a mammal without physically contacting the mammal. The system also includes a signal generator configured to generate a signal indicative of the sensed property of the mammal. The system further includes a treatment decision device configured to determine in response to the signal indicative of the sensed property of the mammal a neuromodulation treatment regimen for administration to a nervous system component of the mammal. The system also includes a computer-readable medium configured to maintain and to provide access to information corresponding to the determined neuromodulation treatment regimen. | 1. A system comprising: a sensor device configured to sense a property of a mammal without physically touching the mammal;a signal generator configured to generate a signal indicative of the sensed property of the mammal;a patient confirmation device configured to confirm that the signal indicative of the sensed property of the mammal was originated in response to the mammal;a feature interpretation device configured to generate data indicative of a physiological characteristic of the mammal in response to the signal indicative of the property of the mammal;a treatment decision device configured to determine in response to the data indicative of a physiological characteristic of the mammal a neuromodulation treatment regimen for administration to a nervous system component of the mammal;wherein the patient confirmation device is further configured to authorize administration of the neuromodulation treatment regimen to the mammal in response to confirmation that the signal indicative of the sensed property of the mammal was originated in response to the mammal; anda patient records device configured to store information corresponding to the determined neuromodulation treatment regimen in a computer-readable medium and to provide electronic access to the stored information. 2. The system of claim 1, wherein the patient confirmation device comprises a RFID tag associated with the mammal. 3. The system of claim 1, wherein the signal generator configured to generate a signal indicative of the sensed property of the mammal includes: a signal generator configured to generate and wirelessly transmit a signal indicative of the sensed property of the mammal. 4. The system of claim 1, wherein the feature interpretation device configured to generate data indicative of a physiological characteristic of the mammal in response to the signal indicative of the property of the mammal includes: a feature interpretation device configured to generate data indicative of a brain waveform, cardiac waveform, neural activity, pupil diameter, breathing, temperature, movement, acoustic, or body position characteristic of the mammal in response to the signal indicative of the property of the mammal. 5. The system of claim 1, wherein the feature interpretation device configured to generate data indicative of a physiological characteristic of the mammal in response to the signal indicative of the property of the mammal includes: a feature interpretation device configured to generate data indicative of a static or a varying physiological characteristic of the mammal in response to the signal indicative of the property of the mammal. 6. The system of claim 1, wherein the feature interpretation device configured to generate data indicative of a physiological characteristic of the mammal in response to the signal indicative of the property of the mammal includes: a feature interpretation device not physically coupled with the sensor device and configured to generate data indi... | The Invention Science Fund I, LLC | A61N001/00
US07862519 | B1 | 2011-01-04 | Easy-to-use multi-use body fluid specimen collection and analyte sensing assembly | 2003-06-16 | A multiple use analyte sensing assembly that includes a long, thin conductor. A plurality of sensing sites are spaced along the conductor and each sensing site includes a membrane system adapted to create a current when placed into contact with body fluid containing the analyte. The assembly also includes a housing, having an aperture, an uptake spool and a payout spool, the conductor being wrapped about the payout spool before the sensing assembly is used. A sensor positioning actuator turns the uptake spool to move each sensing site, in sequence, to the housing aperture and then to the uptake spool. The assembly also includes a skin broaching assembly, having a multiplicity of lancets and a lancet positioning actuator adapted to move each lancet to the aperture. In addition, a lancet use actuator is adapted to move each lancet at least partially through the aperture and then back again. | 1. A multiple use analyte sensing assembly, comprising: (a) a long, thin electrical conductor;(b) a plurality of sensing sites spaced along said electrical conductor, each said sensing site including a membrane system adapted to create a current in conjunction with said electrical conductor when each said sensing site is placed into contact with body fluid containing an analyte;(c) a housing, having a housing aperture;(d) an uptake spool and a payout spool, located within said housing, said electrical conductor being wrapped about said payout spool and threaded to said uptake spool prior to said sensing assembly being used;(e) a sensor positioning actuator for turning said uptake spool so that each said sensing site is moved, in sequence, to said housing aperture and to said uptake spool; and(f) a skin broaching assembly, having: (i) a multiple lancet holding mechanism;(ii) a multiplicity of lancets held by said multiple lancet holding mechanism;(iii) a lancet positioning actuator adapted to move, in sequence, each lancet to a position coincident to said housing aperture; and(iv) a lancet use actuator adapted to move a lancet that is positioned coincident to said housing aperture, at least partially through said housing aperture and then back again. 2. The multiple use analyte sensing assembly of claim 1, wherein said membrane system is a single membrane system that is continuous along said electrical conductor. 3. The multiple use analyte sensing assembly of claim 1, wherein said electrical conductor is covered with insulation, and wherein said insulation is absent in a plurality of regions, said plurality of regions being the locations of said sensing sites where said membrane system has been applied. 4. The multiple use analyte sensing assembly of claim 1, wherein said membrane system is covered by an absorbent layer. 5. The multiple use analyte sensing assembly of claim 4, wherein tight bands are interspersed about said absorbent layer between sensing sites to prevent blood entering a first sensing site of said plurality of sensing sites from a second sensing site of said plurality of sensing sites. 6. A method of taking a sequence of blood analyte measurements, comprising: (a) providing an analyte sensing assembly, including: (i) a lengthwise electrical conductor; and(ii) a first measurement site and a second measurement site located along said lengthwise electrical conductor, each said measurement site including a membrane assembly adapted to create an electrical current flowing out of said lengthwise electrical conductor in the presence of an analyte when a voltage is placed on said lengthwise electrical conductor relative to ground;(b) forming a first analyte measurement by placing a voltage on said lengthwise electrical conductor, exposing said first measurement site to body fluid, and measuring electrical current in said lengthwise electrical conductor; and(c) subsequent to forming said first analyte measurement, forming a second a... | Isense Corporation | A61B005/00
US08638100 | B2 | 2014-01-28 | Surface coil for a magnetic resonance tomography apparatus with interchangeable operating elements and a simultaneously mechanical and electrical locking element | 2010-08-03 | An assembly for mounting and unmounting a surface coil of a magnetic resonance tomography apparatus includes an upper part, a lower part, a locking element to fix the upper part and the lower part to one another, and contact elements at the upper part and at the lower part. The contact elements at the upper part can be brought into an electrically conductive connection with contact elements at the lower part and the upper part and the lower part can be fixed to one another by a movement of the locking element. | 1. A surface coil configured for a magnetic resonance tomography apparatus, comprising: a coil housing;a magnetic resonance surface coil contained inside coil housing;said coil housing comprising an upper part and a lower part that are separable from each other and that respectively have mating surfaces that engage said upper part and said lower part with each other;said coil housing upper part comprising upper part electrical contact elements and said coil housing lower part comprising lower part electrical contact elements that face each other when said coil housing upper part is placed on said coil housing lower part;a manually operable locking element located at one of said coil housing upper part or said coil housing lower part and being configured, when manually operated, to proceed through a movement, that produces a relative movement between said coil housing upper and coil housing lower parts in order to simultaneously bring said coil housing upper part electrical contact elements into electrically conductive connection with said coil housing lower part electrical contact elements and mechanically fix said coil housing upper part and said coil housing lower part in engagement with each other. 2. A surface coil as claimed in claim 1 wherein said movement comprises a rotation movement, and wherein said rotation movement brings said coil housing upper part electrical contact elements into said electrically conductive connection with said coil housing lower part electrical contact elements. 3. A surface coil as claimed in claim 1 wherein said locking element is configured to simultaneously bring said coil housing upper part electrical contact elements into said electrically conductive connection with said coil housing lower part electrical contact elements, and configured to mechanically fix said coil housing upper part and said coil housing lower part in engagement with each other, by only a single movement of said locking element. 4. A surface coil as claimed in claim 1 wherein said movement is a locking movement, and wherein said locking element is configured to execute an unlocking movement in a direction opposite to said locking movement that separates said coil housing upper part electrical contact elements from said coil housing lower part electrical contact elements, and that also mechanically disengages said coil housing upper part and said coil housing lower part from each other. 5. A surface coil as claimed in claim 1 wherein said coil housing upper part is displaced onto said coil housing lower part by said movement of said locking element, and wherein said coil housing upper part comprises a plurality of retention elements that positively fit said coil housing upper part onto said coil housing lower part as said coil housing upper part and said coil housing lower part are brought into engagement with each other by said movement of said locking element. 6. A surface coil as claimed in claim 1 wherein said coil housing upper ... | Siemens Aktiengesellschaft | G01R033/36
US08798934 | B2 | 2014-08-05 | Real time management of data relating to physiological control of glucose levels | 2010-07-23 | Continuous glucose monitoring (CGM) data and insulin delivery data are used to generate more reliable projected alarms related to a projected glucose levels. A memory stores endogenous data related to measurements of glucose level in a patient, and also stores exogenous data, such as insulin on board, both of which are used by a processor to create projected alarms. Profiles of CGM data are created for use in tuning patient-specific insulin data, such at basal rate, carb ratio, and insulin sensitivity. A processor searches for patterns in the data profiles and if found, recommended changes to patient-specific insulin data are provided to permit more accurate control over a patient's glucose levels. | 1. An integrated glucose management system for tuning patient-specific insulin data, comprising: a memory configured to record and store glucose measurement data representing measurements of physiological glucose levels, and to store exogenous data in a form of attributes tagged to the stored glucose measurement data;a user interface comprising a visual display and an input device configured to receive and communicate user input data and instructions; anda processor comprising computer-executable instructions to: record multiple series of glucose measurement data into the memory during defined time periods;tag each of the recorded series of glucose measurement data with exogenous attributes including a profile name, wherein the name of the profile is selected to identify the data recording as belonging to a particular category of patient conditions;access the memory to retrieve a plurality of profiles having a same profile name;compare the retrieved plurality of profiles to detect a persistent pattern of undesirable measured glucose levels existing in the plurality of profiles;provide an alarm at the user interface if the persistent pattern is detected in the retrieved plurality of profiles; andprovide a recommended change to patient-specific insulin data as a result of the detected persistent pattern, and display the recommended change on the user interface;wherein the computer-executable instructions include an excess insulin manager configured to provide a plurality of alternative recommended changes to the patient-specific insulin data; and furtherwherein the excess insulin manager is further configured to prioritize a recommended increase in carbohydrate intake lower than recommended changes to insulin delivery. 2. The integrated glucose management system of claim 1, wherein the patient-specific insulin data comprises a basal rate, a carb ratio, and an insulin sensitivity. 3. The integrated glucose management system of claim 1, wherein the plurality of profiles includes a skip-meal profile, a meal test profile, and a correction bolus test profile. 4. The integrated glucose management system of claim 1, wherein the processor comprises a further computer-executable instruction to require that a minimum number of the retrieved plurality of profiles must be outside a safe range before providing the recommended change to the patient-specific insulin data. 5. The integrated glucose management system of claim 1, wherein the processor comprises a further computer-executable instruction to require that the plurality of profiles retrieved for comparison must have been recorded within a selected time period. 6. The integrated glucose management system of claim 1, wherein the processor comprises a further computer-executable instruction to require that a recommendation for change of a basal rate, a carb ratio, and an insulin sensitivity cannot exceed a predetermined amount. 7. The integrated glucose management system of claim 1, wherein the proces... | Abbott Diabetes Care Inc. | G01N033/48 G01N031/00 G06G007/48 G06G007/58 A61B005/145 G06F019/00 A61M005/172 A61M005/142
US08357106 | B2 | 2013-01-22 | Analysis device | 2008-04-29 | This invention is regard to an analysis device that analyze a specific component in bodily fluid extracted from the skin using an analysis tool. The analysis device comprise a laser beam oscillation section that emits a laser beam for extracting the bodily fluid from the skin, and a detection mechanism that detects whether or not the analysis tool exists at a target position. The analysis device is adapted to emit the laser beam from the laser beam oscillation section when the analysis tool is detected by the detection mechanism. | 1. An analysis device adapted to analyze a specific component in bodily fluid extracted from skin and operative in conjunction with an analysis tool, the analysis device comprising: an extraction mechanism configured to extract the bodily fluid from the skin,a detection mechanism configured to detect whether the analysis tool is located at a target position, andan analyzing element configured to analyze the specific component in the bodily fluid extracted from the skin,wherein the extraction mechanism is configured to activate when the analysis tool is detected by the detection mechanism as located at the target position. 2. The analysis device according to claim 1, wherein the detection mechanism comprises a switch configured to turn ON when the analysis tool is located at the target position, and the extraction mechanism is configured to activate when the switch is turned ON. 3. The analysis device according to claim 2, further comprising a moving portion capable of moving together with the analysis tool, wherein the switch is configured to be turned ON by movement of the moving portion to where the analysis tool is located at the target position. 4. The analysis device according to claim 3, wherein the moving portion is configured to move by applying an external force thereto. 5. The analysis device according to claim 3, wherein the detection mechanism further comprises an elastic member that is configured to contact the switch when the moving portion is moved to where the analysis tool is located at the target position. 6. The analysis device according to claim 3, wherein: the extraction mechanism comprises a pricking element for pricking the skin; andthe moving portion comprises a pass-through portion that is sized to allow the pricking element to pass therethrough. 7. The analysis device according to claim 1, wherein the detection mechanism is configured to optically detect a specific portion of the analysis tool. 8. The analysis device according to claim 1, wherein the extraction mechanism is capable of emitting a laser beam to irradiate the skin, and is configured to emit the laser beam when the analysis tool is detected by the detection mechanism as being located at the target position. 9. An analysis apparatus, comprising: an analysis tool including a plurality of electrodes;an analysis device configured to analyze a specific component in bodily fluid extracted from skin and operative in conjunction with the analysis tool, the analysis device including an extraction mechanism configured to extract the bodily fluid from the skin and a detection mechanism configured to detect whether the analysis tool is located at a target position; andan analyzing element configured to analyze the specific component in the bodily fluid extracted from the skin,wherein the extraction mechanism is configured to activate when the analysis tool is detected by the detection mechanism as located at the target position. 10. The analysis apparatus accor... | Arkray, Inc. | A61B005/00
US08496598 | B2 | 2013-07-30 | Apparatus for early diagnosis of carpal tunnel syndrome and method and device for acquiring information on the skin sensitivity of a finger of a patient | 2009-02-02 | The apparatus (2) for early diagnosis of carpal tunnel syndrome, of the type comprising: receiving means which define a receiving zone (8) for receiving a patient's finger,means for measuring the cutaneous sensibility of a finger (10) disposed in said receiving means, comprising a device for determining a sensation threshold of the finger pulp and a device for measuring the innervation density of the finger pulp, each device comprising protruding portions (22) adapted to be brought into contact with the finger (10), is characterized in that the protruding portions (22) of the device for measuring the innervation density are identical with those of the device for determining the sensation threshold. | 1. An apparatus (2) for the early diagnosis of carpal tunnel syndrome, of the type comprising: receiving means which define a receiving zone (8) for receiving a patient's finger,means for measuring the cutaneous sensibility of a finger (10) disposed in said receiving means, comprising a device for determining a sensation threshold (16) of the finger pulp and a device for measuring the innervation density (18) of the finger pulp, each device comprising protruding portions (22) adapted to be brought into contact with the finger (10),characterized in that the protruding portions (22) of the device for measuring the innervation density (18) are identical with those of the device for determining the sensation threshold (16), and the device for determining the sensation threshold (16) comprises a first set (20) of at least two filaments (22) of different diameters, means for selecting a filament of said first set (20), and means for displacing the selected filament (22) towards the receiving zone (8), parallel to the direction of said selected filament (22). 2. The apparatus (2) according to claim 1, characterized in that the device for measuring the innervation density (18) of the finger (10) comprises a pair of identical sets (20, 28) of filaments (22), each set (20, 28) comprising at least two filaments (22) of different diameters, means for selecting a pair of filaments (22) of identical diameters and positioning them parallel with one another, each filament (22) of said pair belonging to one of said sets (20, 28) of filaments (22), means for adjusting the distance between the filaments (22) of the selected pair, and means for displacing the selected pair of filaments (22) towards the receiving zone (8) parallel to the direction of said filaments (22) of the selected pair. 3. The apparatus (2) according to claim 2, characterized in that it comprises a shaft (36) and at least one support means (24, 30), which is mounted freely in rotation about the shaft (36) and is displaceable in translation in a direction perpendicular to the shaft (36), each set (20, 28) of filaments (22) being disposed on a support means (24, 30), said filaments (22) being oriented in radial direction relative to the shaft (36). 4. The apparatus (2) according to claim 3, characterized in that the selection and positioning means, respectively, comprise a second cam (68) which is arranged to cause each support means (24, 30) to rotate about the shaft (36). 5. The apparatus (2) according to claim 1, characterized in that the set (20) of the filaments (22) of the device for determining the sensation threshold (16) is one of the two sets (20, 28) of filaments (22) of the device for measuring the innervation density (18). 6. The apparatus (2) according to claim 1, characterized in that it comprises a shaft (36) and at least one support means (24, 30), which is mounted freely in rotation about the shaft (36) and is displaceable in translation in a direction perpendicular to the s... | Technologia | A61B019/00
US08100532 | B2 | 2012-01-24 | Eye and body movement tracking for testing and/or training | 2009-07-09 | The eye movement, body movement, and/or physiological performance of a subject may tracked while the subject performs a task, such as participating in an activity or sport. The collected data may then be used to identify correlations between the subject's eyesight and the subject's body movement exists and/or physiology. Such a correlation may be analyzed (e.g., over time) to determine any delays or gaps in the subject's ability to track an object, such as a ball, while participating in a sport or other activity. Further, a subject's performance may be compared to data collected from other individuals. The eye movement, body movement, and/or physiological performance data may be used to test and/or train the visual and cognitive abilities of an individual. | 1. A method for testing or training the eye and body movement of a subject while the subject performs a task, the method comprising: tracking the subject's eye movement while the subject performs the task to produce eye movement data associated therewith;tracking the subject's body movement while the subject performs the task to produce body movement data associated therewith; anddetermining a correlation between the subject's eye movement data and the subject's body movement data. 2. The method of claim 1, wherein the subject's eye movement is tracked by an eye tracking system. 3. The method of claim 1, wherein the subject's body movement is tracked by a body tracking system. 4. The method of claim 1, wherein tracking the subject's eye movement includes tracking the lens of at least one of the subject's eyes. 5. The method of claim 3, further comprising: collecting physiological data while the subject performs the task to produce physiological data associated therewith; anddetermining a correlation between the subject's physiological data and the subject's eye movement data and the subject's body movement data. 6. The method of claim 3, wherein the body tracking system is at least one of an optical system and a non-optical system. 7. The method of claim 1, wherein determining a correlation between the subject's eye movement and the subject's body movement includes plotting the eye movement and body movement versus time. 8. The method of claim 1, further comprising: storing the eye movement data and the body movement data on a storage device. 9. The method of claim 1, further comprising: sending the eye movement data and the body movement data to an output device. 10. A system for testing or training the eye and body movement of a subject, the system comprising: an eye tracking system that tracks the eye movement of a subject, wherein the eye tracking system collects eye movement data associated with the subject's eye movement;a body tracking system that tracks the body movement of a subject, wherein the body tracking system collects body movement data associated with the subject's body movement; anda computing device that correlates the subject's eye movement with the subject's body movement. 11. The system of claim 10, further comprising: a physiological data collection system, wherein the physiological data collection system collects physiological data associated with the subject's physiological condition; and whereinthe computing device further analyzes the physiological data. 12. The system of claim 10, wherein the eye tracking system tracks the subject's eye movement by tracking the lens of at least one of the subject's eyes. 13. The system of claim 10, wherein the body tracking system is an optical system. 14. The system of claim 10, wherein the body tracking system is a non-optical system. 15. The system of claim 10, further comprising an output device. 16. The system of claim 15, wherein the computing device plots the eye ... | Nike, Inc. | A61B003/14
US07881766 | B2 | 2011-02-01 | Apparatus and method for attaching biosignal measurement sensor to subject | 2006-05-16 | An apparatus for attaching a biosignal measurement sensor to a subject including a housing having an open space at one side, an installation portion installed in the open space, on which the biosignal measurement sensor is installed, a plurality of detection electrodes installed at an open side of the housing to contact the subject, and arranged along a circumferential edge of the installation portion to form pairs of detection electrodes facing each other with respect to a center of arrangement, an actuator installed in the housing to rotate and linearly move the installation portion, and a circuit portion searching for the detection electrodes that provide optimal detection data based on a difference in signals between the pairs of detection electrodes, and driving the actuator to allow a pair of sensor electrodes provided at the biosignal measurement sensor to be arranged in the same direction as the searched detection electrodes. | 1. An apparatus for attaching a biosignal measurement sensor to a subject, the apparatus comprising: a housing having an open space at one side;an installation portion rotatably installed in the open space, on which the biosignal measurement sensor is installed, the biological measurement sensor having a pair of sensor electrodes;a plurality of detection electrodes on the one side of the housing so as to be arranged along a circumferential edge of the installation portion, and forming pairs of detection electrodes that face each other with respect to a center of arrangement along the circumferential edge;an actuator installed in the housing, the actuator operative to rotate and linearly move the installation portion; anda circuit portion that searches for detection electrodes that provide optimal detection data based on a difference in signals between each pair of detection electrodes, and drives the actuator to allow the pair of sensor electrodes provided at the biosignal measurement sensor to be arranged in a same direction in which the searched detection electrodes are arranged,wherein the detection electrodes comprise at least twelve detection electrodes arranged along the circumferential edge of the installation portion at a constant interval. 2. The apparatus as claimed in claim 1, wherein the optimal detection data is detection data having a maximum correlation coefficient with respect to reference data. 3. The apparatus as claimed in claim 2, wherein the circuit portion comprises a microcontroller operative to search the detection electrodes having the detection data having the maximum correlation coefficient by calculating correlation coefficients between the detection data and the reference data, and the microcontroller is operative to provide a control signal to the actuator to arrange the sensor electrodes in the same direction in which the searched detection electrodes are arranged. 4. The apparatus as claimed in claim 3, wherein the circuit portion comprises: a detection signal selection portion operative to sequentially select two detection signals output from each pair of the detection electrodes; anda detection signal processing portion operative to amplify a difference between the detection signals received from the detection signal selection portion and the detection signal processing portion is operative to digitalize the amplified signal difference to generate detection data and provide the generated detection data to the microcontroller. 5. The apparatus as claimed in claim 4, wherein the detection signal processing portion comprises: a differential amplification portion operative to amplify the signal difference between the detection signals received from the detection signal selection portion; andan A/D conversion portion operative to correct an analog signal received from the differential amplification portion into a digital signal. 6. The apparatus as claimed in claim 2, wherein the reference data is data obtained b... | Samsung Electronics Co., Ltd. | A61B005/04
US08777923 | B2 | 2014-07-15 | Patient monitoring and drug delivery system and method of use | 2010-10-26 | Disclosed is a patient monitoring and drug delivery system and associated methods for use during diagnostic, surgical or other medical procedures. The functionality of the invention enables many time consuming and laborious activities to be minimized or moved to a part in the procedure where time is not as critical. The invention is capable of increasing practice efficiency in patient care facilities through system architecture and design into two separate units. A patient unit receives input signals from patient monitoring connections and outputs the signals to a procedure unit. The procedure unit is operational during the medical procedure and controls the delivery of drugs to the patient. | 1. A method for monitoring a patient and delivering at least one drug during a medical procedure in a medical care facility comprising the steps of: a) locating the patient in a first location of the medical care facility and connecting to the patient at least one sensor for monitoring at least one physiological parameter of the patient;b) obtaining a first housing having a first microprocessor-based patient unit having at least one first connection point for receiving input signals from the at least one sensor and at least one second connection point for outputting patient physiological parameters;c) locating the patient in a second location of the medical care facility, connecting the at least one second connection point to a second housing having a second microprocessor-based procedure unit, and delivering the drug to the patient and performing a medical procedure on the patient, wherein at least a portion of the medical procedure is controlled by the second microprocessor-based procedure unit and;d) delivering oxygen to the patient by way of the second connection point. 2. The method of claim 1 further comprising the steps of inputting to the first microprocessor-based patient unit physical attributes of the patient. 3. The method of claim 1 further comprising the step of querying the patient to determine a level of consciousness of the patient. 4. The method of claim 3 further comprising the step of the patient activating a response device. 5. The method of claim 1 further comprising the step of delivering the drug via an IV tube and priming the IV tube before delivering the drug to the patient. 6. The method of claim 5, wherein priming the IV tube is automatically performed by the procedure unit. 7. A method for monitoring a patient and delivering at least one drug during a medical procedure in a medical care facility comprising the steps of: a) locating the patient in a first location of the medical care facility and connecting to the patient at least one sensor for monitoring at least one physiological parameter of the patient;b) obtaining a first housing having a first microprocessor-based patient unit having at least one first connection point for receiving input signals from the at least one sensor and at least one second connection point for outputting patient physiological parameters;c) locating the patient in a second location of the medical care facility, connecting the at least one second connection point to a second housing having a second microprocessor-based procedure unit, and performing a medical procedure on the patient, wherein at least a portion of the medical procedure is controlled by the second microprocessor-based procedure unit and;d) inputting exhaled respiratory gases from the patient to the procedure unit by way of the second connection point. 8. The method of claim 7 further comprising the step of delivering oxygen to the patient by way of the second connection point. | Ethicon Endo-Surgery, Inc. | A61M031/00
US08652058 | B2 | 2014-02-18 | Blood vessel pressing cuff, blood pressure measuring apparatus including the blood vessel pressing cuff, and blood pressure measuring method using the blood pressure measuring apparatus | 2010-04-01 | A blood vessel pressing cuff includes a strap surrounding a body part, a first actuator disposed on the strap and including a first shape memory alloy which changes to a first shape memorized in advance, at a temperature equal to or higher than a first temperature, and a second actuator disposed on the strap and including a second shape memory alloy which changes to a second shape memorized in advance, at a temperature equal to or higher than a second temperature that is different from the first temperature. If the first shape memory alloy changes to the first shape, pressure applied to the body part surrounded by the strap is increased. Even when the first shape memory alloy changes to the first shape, if the second shape memory alloy changes to the second shape, the pressure applied to the body part surrounded by the strap is reduced. | 1. A blood vessel pressing cuff comprising: a strap adapted to surround a body part; a first actuator disposed on the strap, and comprising a first shape memory alloy which changes to a first shape memorized in advance, at a temperature equal to or higher than a first temperature; and a second actuator disposed on the strap and separated from the first actuator along a length of the strap, and comprising a second shape memory alloy which changes to a second shape memorized in advance, at a temperature equal to or higher than a second temperature which is different from the first temperature, wherein the first actuator is adapted to apply an increased pressure at a target part of the body part when the first shape memory alloy changes to the first shape, and wherein the first actuator is adapted to apply a decreased pressure to the target part of the body part due to the change of the second shape memory alloy to the second shape. 2. The blood vessel pressing cuff of claim 1, wherein the first shape is a shape of the first shape memory alloy, in which a length of the first actuator is reduced when the first shape memory alloy changes to the first shape, and wherein the second shape is a shape of the second shape memory alloy, in which a length of the second actuator is increased when the second shape memory alloy changes to the second shape. 3. The blood vessel pressing cuff of claim 1, wherein the second shape memory alloy returns to a third shape memorized in advance, at a temperature equal to or lower than a third temperature which is between the first and second temperatures, and has no shape memorized in advance at a temperature between the second and third temperature. 4. The blood vessel pressing cuff of claim 3, wherein the third shape is a shape of the second shape memory alloy, in which a length of the second actuator is reduced when the second shape memory alloy returns to the third shape. 5. The blood vessel pressing cuff of claim 1, wherein the first temperature is higher than room temperature, andwherein the second temperature is higher than the first temperature. 6. The blood vessel pressing cuff of claim 1, wherein the first shape is a shape of the first shape memory alloy, in which the first actuator protrudes toward the body part when the first shape memory alloy changes to the first shape, andwherein the second shape is a shape of the second shape memory alloy, in which a length of the second actuator is increased when the second shape memory alloy changes to the second shape. 7. The blood vessel pressing cuff of claim 1, wherein the first shape memory alloy or the second shape memory alloy is self-heated due to a supplied current. 8. The blood vessel pressing cuff of claim 1, wherein the first actuator comprises wires of the first shape memory alloy, the wires of the first shape memory alloy being longitudinally extended in a length direction of the first actuator. 9. The blood vessel pressing cuff of claim 1, wherein t... | Samsung Electronics Co., Ltd. | A61B005/02
US07862620 | B2 | 2011-01-04 | Sensing systems and methods for monitoring gait dynamics | 2006-05-08 | Systems and method for monitoring gait dynamics are disclosed. The performance of an orthotic or prosthetic device or other device associated with a limb may be measured based on the resistance of a bending sensor. Data from the sensors is gathered or processed, particularly for purposes of alignment, safety, failure, usage, selection, and artificial proprioception. Information relating to the device may be outputted visually or auditorily to an individual. | 1. A method of gathering information regarding use of a prosthetic foot, comprising: providing a prosthetic foot having at least one flexible portion and at least one elongate deformable sensor associated therewith, wherein the elongate deformable sensor is configured to bend with said flexible portion and is positioned on a top portion of said flexible portion and extending along both sides of a bending axis of said flexible portion;measuring with said elongate deformable sensor a performance characteristic of said prosthetic foot while in use by a user, wherein said performance characteristic comprises a plurality of bending force measurements of said prosthetic foot over a period of time;storing data corresponding to said plurality of bending force measurements in a memory associated with said prosthetic foot; andprocessing said plurality of bending force measurements to identify at least one gait dynamic of the user. 2. The method of claim 1, wherein said measuring comprises detecting a resistance correlated to a bend radius of the flexible portion of said prosthetic foot. 3. The method of claim 1, wherein said measuring comprises measuring said performance characteristic on a step-by-step basis during locomotion by the user. 4. The method of claim 1, wherein the memory is provided on the prosthetic foot. 5. The method of claim 1, wherein said storing comprises storing the data on a memory residing on a device temporarily connected to the prosthetic foot. 6. The method of claim 1, further comprising connecting a computer to the prosthetic foot, the computer containing said memory to store said data. 7. The method of claim 1, further comprising analyzing the stored data to provide information indicative of at least one of alignment, usage statistics, safety, selection, and failure of the prosthetic foot. 8. The method of claim 7, wherein said analyzing is performed by a processor coupled to the prosthetic foot. 9. The method of claim 8, wherein the at least one gait dynamic comprises at least one of heel strike and toe load. 10. The method of claim 9, wherein said identifying the at least one gait dynamic further comprises determining a number of steps taken by a user over the period of time. 11. The method of claim 1, wherein the elongate deformable sensor comprises a resistive strip having a length-to-width ratio of at least 5:1. 12. The method of claim 1, additionally comprising dynamically aligning the prosthetic foot based on the at least one gait dynamic of the user. 13. The method of claim 12, wherein said dynamically aligning is performed during locomotion by the user. 14. The method of claim 12, wherein said dynamically aligning comprising adjusting a heel height of the prosthetic foot. 15. A prosthetic foot system, comprising: a prosthetic foot having at least one flexible portion configured to bend during locomotion by a user;at least one elongate deformable sensor adhered to a top surface of said flexible portion of t... | Össur hf | A61F002/48 A61F002/66 A61F002/68 A61B019/00 A61B005/00 G01L005/16 G01M019/00 G01M001/00
US07863892 | B2 | 2011-01-04 | Multiple SQUID magnetometer | 2006-10-06 | Multiple SQUID magnetometers that include at least two SQUID loops, each of which is composed of at least two Josephson Junctions connected in parallel with superconducting wires, are provided. The SQUID loops are fabricated such that they share a common Josephson Junction. Devices and application that employ the multiple SQUID magnetometers are also provided. | 1. A multiple SQUID magnetometer comprising: (a) a first SQUID loop comprising at least two Josephson Junctions connected in parallel by superconducting wires; and(b) a second SQUID loop comprising at least two Josephson Junctions connected in parallel by superconducting wires;wherein the first and second SQUID loops share a common Josephson Junction, and further comprising a sample coupled to the common Josephson Junction, wherein the sample is disposed in a pit in the common Josephson Junction. 2. The magnetometer of claim 1 wherein the first and second SQUID loops have diameters of no more than about 200 nm. 3. The magnetometer of claim 1 wherein the sample has dimensions smaller than the width of the superconducting wires. 4. The magnetometer of claim 3 wherein the width of the superconducting wires is no greater than about 100 nm. 5. The magnetometer of claim 1 wherein the sample comprises a magnetic nanoparticle. 6. The magnetometer of claim 5 wherein the magnetic nanoparticle comprises an endohedral fullerene. 7. The magnetometer of claim 6 wherein the endohedral fullerene is a N@C60 molecule. 8. The magnetometer of claim 1 wherein the sample is a biomolecule. 9. The magnetometer of claim 8 wherein the biomolecule is a protein. 10. A multiple SQUID magnetometer comprising: (a) a first SQUID loop comprising at least two Josephson Junctions connected in parallel by superconducting wires; and(b) a second SQUID loop comprising at least two Josephson Junctions connected in parallel by superconducting wires;wherein the first and second SQUID loops share a common Josephson Junction, and further comprising a third SQUID loop comprising two Josephson Junctions connected in parallel by superconducting wires, wherein the first, second and third SQUID loops share a common Josephson Junction, and wherein each Josephson Junction is configured to be separately biased. 11. The magnetometer of claim 10 further comprising a plurality of SQUID loops, the plurality of SQUID loops including at least four SQUID loops, wherein each of the SQUID loops in the plurality of SQUID loops comprises two Josephson Junctions connected in parallel by superconducting wires, and further wherein each of the SQUID loops in the plurality of SQUID loops share a common Josephson Junction. 12. The magnetometer of claim 10 wherein the first, second and third SQUID loops each comprise at least three Josephson Junctions. 13. A nuclear magnetic resonance apparatus comprising the magnetometer of claim 1 incorporated into a nuclear magnetic resonance instrument. 14. A multiple SQUID array comprising an array of SQUID loops, each SQUID loop comprising at least two Josephson Junctions connected in parallel by superconducting wires, wherein each SQUID loop in the array shares a common Josephson Junction with each of its neighboring SQUID loops, and further comprising a sample coupled to each common Josephson Junction. 15. The array of claim 14 wherein the array is a one dimen... | Florida State University Research Foundation | G01R033/035 A61B005/05
US08280477 | B2 | 2012-10-02 | Mono-phasic action potential electrogram recording catheter, and method | 2009-07-29 | Catheters and methods for obtaining monophasic action potential (“MAP”) electrograms include a flexible catheter body defining a longitudinal axis, and a distal assembly affixed to the catheter body distal end defining a distal tip. The distal assembly has at least three MAP recording electrodes, and at least one reference electrode for determining reference potential. The recording electrodes are each positioned a radial distance from the longitudinal axis in at least three different radial directions, defining a recording geometry substantially having radial symmetry. The reference electrode is a longitudinal distance from the recording geometry. Optional features include a steerable catheter shaft, one or more radio-frequency ablation electrodes, and a dedicated pacing electrode. Different possible embodiments include a dome housing having the recording electrodes in a fixed spatial arrangement, and a distal loop assembly having an array of electrodes on at least three flexible loop elements. | 1. A medical device, comprising: a catheter body having proximal and distal ends, and defining a longitudinal axis;a distal assembly affixed to the distal end of the catheter body, including at least two monophasic action potential electrogram recording electrodes, each positioned a radial distance from the longitudinal axis in different radial directions, and each having a spherical recording surface;wherein the recording electrodes substantially have radial symmetry around the longitudinal axis, and wherein the distal assembly further includes a dome housing having at least three sockets for receiving the recording electrodes in a fixed spatial arrangement;at least one reference electrode for determining reference potential, positioned a longitudinal distance from each recording electrode. 2. The medical device of claim 1, further comprising: at least one radio-frequency ablation electrode for ablating tissue; anda dedicated pacing electrode,wherein each of the at least one ablation electrode and pacing electrode are positioned a longitudinal distance from each recording electrode. 3. The medical device of claim 2, wherein the pacing and reference electrodes are recessed, to reduce direct contact with tissue by the pacing and reference electrodes. 4. The medical device of claim 2, wherein the reference and pacing electrodes have a cylindrical shape around the longitudinal axis; each recording electrode, reference electrode, and pacing electrode having an effective surface of substantially equivalent area. 5. The medical device of claim 2, further comprising an electrically insulating material positioned between each of the recording electrodes and any other electrode. 6. The medical device of claim 1, wherein the materials of the electrodes are selected from the group consisting of silver, silver chloride, platinum, iridium, gold, stainless steel, aluminum, and alloys and combinations thereof. 7. The medical device of claim 1, wherein the catheter shaft is resilient, for facilitating steady contact with tissue to improve signal fidelity. 8. The medical device of claim 1, wherein each recording electrode has a diameter of at most 2 millimeters. 9. The medical device of claim 1, further comprising an electrically insulating material positioned between each recording electrode and a corresponding socket in the dome housing, such that the dome housing is capable of radio-frequency ablation. 10. The medical device of claim 1, wherein each recording electrode extends an equal distance distally and radially beyond the dome housing. 11. The medical device of claim 1, wherein the recording electrodes are four in number, each positioned at 90 degrees around a circumference of the recording geometry, relative to adjacent recording electrodes. | Medtronic Cryocath LP | A61B005/042 A61B018/14 A61N001/05
US07865376 | B2 | 2011-01-04 | System and method for generating de-identified health care data | 2007-12-11 | A system and method for creating a unique alias associated with an individual identified in a health care database such that health care data, and particularly pharmaceutical-related data, can be efficiently gathered and analyzed. The system has a first data store for storing at least one record where each record includes a plurality of identification fields which when concatenated uniquely identify an individual, and at least one health care field corresponding to health care data associated with the individual. The system also has a second data store, and a processor. The processor selects a record of the first data store, then selects a subset of the plurality of identification fields within the selected record, and concatenates the selected subset of identification fields. Then the processor stores the concatenated identification fields in a record in the second data store with the at least one health care field from the selected record of the first data store. | 1. A system for creating a unique alias associated with an individual identified in a health care database, comprising: a first data store for storing at least one record, each record including a plurality of identification fields which when taken together uniquely identify an individual, and at least one health care field corresponding to health care data associated with the individual;a second data store; anda processor for performing the steps of: selecting a record of the first data store;selecting a subset of the plurality of identification fields within the selected record;concatenating the selected subset of identification fields;encrypting the concatenated identification fields; andstoring in a second record that does not include the plurality of identification fields in the second data store the encrypted concatenated identification fields as well as the at least one health care field from the selected record of the first data store. 2. The system of claim 1, wherein the health care data stored within the first data store corresponds to pharmaceutical claims data. 3. The system of claim 1, wherein the first data store and the second data store are both located within the same database. 4. The system of claim 1, wherein the first data store and the second data store are both located within different databases. 5. The system of claim 1, wherein the selected subset comprises last name, birthday and gender. 6. The system of claim 1, wherein the processor performs the further steps of: based on the concatenated identification fields and the at least one health care field of each record of the second data store, analyzing longitudinal, historical records of individuals using individual-level linking methodologies. 7. A system for creating a unique alias associated with an individual identified in a health care database, the database including a first data store for storing at least one record, each record including a plurality of identification fields which when taken together uniquely identify an individual, and at least one health care field corresponding to health care data associated with the individual, and a second data store, the system comprising: means for selecting a subset of the plurality of identification fields;means for concatenating the selected subset of identification fields;means for encrypting the concatenated identification fields; andmeans for storing in a second record that does not include the plurality of identification fields in the second data store the encrypted concatenated identification fields as well as the at least one health care field from the selected record of the first data store. 8. A method for creating a unique alias associated with an individual identified in a health care database, wherein the health care database stores at least one record, each record including a plurality of identification fields which when taken together uniquely identify an individual, and at least one health care field co... | SDI Health LLC | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07871598 | B1 | 2011-01-18 | Stable metal ion-lipid powdered pharmaceutical compositions for drug delivery and methods of use | 2000-05-10 | Microparticle compositions comprising metal ion-lipid complexes for drug delivery are described including methods of making the microparticle compositions and methods of treating certain conditions and disease states by administering the microparticle compositions. The metal ion-lipid complexes can be combined with various drugs or active agents for therapeutic administration. The microparticle compositions of the present invention have superior stability to other microparticle compositions resulting in a microparticle composition with longer shelf life and improved dispersability. The microparticle compositions of the present invention have a transition temperature Tm of at least 20° C. above the recommended storage temperature (Tst) for drug delivery. | 1. A powder composition comprising a plurality of microparticles, the microparticles comprising an active agent and a metal ion-lipid complex, and the powder composition has a metal ion concentration of greater than 0 and up to and including 25% w/w, and a lipid concentration of 25-90% w/w, and wherein the density of the microparticles as measured by He displacement is 0.5-2.0 g/ml and the Tg of the microparticles is at least 2° C. greater than that of microparticles without the metal ion. 2. The powder composition of claim 1 wherein the lipid component comprises a mixture of at least two lipids. 3. The powder composition of claim 1 wherein the lipid comprises a phospholipid. 4. The powder composition of claim 1 wherein the metal ions are chosen from the group consisting of lanthanide metals, actinide metals, group IIa and IIIb metals, transition metals or mixtures thereof. 5. The powder composition of claim 4 wherein the metal ion is chosen from the group consisting of calcium, zinc, aluminum, iron and magnesium, their water soluble salts and mixtures thereof. 6. The powder composition of claim 1 wherein the metal ion has more than a one positive charge. 7. The powder composition of claim 3 wherein the phospholipid is selected from the group consisting of dipalmitoylphosphatidylcholine, distearoylphosphatidylcholine, and dimyristoylphosphatidylcholine. 8. The powder composition of claim 7 wherein the microparticle has a phospholipid concentration from about 20-90% w/w. 9. The powder composition of claim 1 wherein the active agent is selected from the group consisting of antiallergics, antifungals, bronchodilators, pulmonary lung surfactants, analgesics, antibiotics, leukotriene inhibitors or antagonists, antihistamines, antiinflammatories, antineoplastics, anticholingerics, anesthetics, antituberculars, imaging agents, cardiovascular agents, enzymes, steroids, DNA, RNA, viral vectors, antisense agents, proteins, peptides, immunoglobulins, and combinations thereof. 10. The powder composition of claim 1 wherein the active agent is selected from the group consisting of albuterol, budesonide, fluticasone, salmeterol, formoterol, nicotine, triamcinolone, dexamethasone, beclomethasone, gentamicin, ciprofloxacin, paclitaxel, amphotericin, amikacin, tobramycin, insulin, human growth hormone, and salts thereof. 11. The powder composition of claim 1 further comprising a surfactant selected from the group consisting of nonionic detergents, nonionic block copolymers, ionic surfactants and combinations thereof. 12. The powder composition of claim 11 wherein the surfactant is selected from the group consisting of sorbitan esters, ethoxylated sorbitan esters, fatty acid salts, sugar esters, ethylene oxides, and combinations thereof. 13. The powder composition of claim 1 further comprising a polymer selected from the group consisting of polysaccharides, polyvinyl alcohol, polyvinyl pyrrolidone, polylactides, polyglycolides, polyethylene glycol, cycl... | Novartis AG | A61B005/055 A61K009/127 A61K009/50
US08600467 | B2 | 2013-12-03 | Optical sensor including disposable and reusable elements | 2010-07-01 | An embodiment of the present disclosure provides a noninvasive optical sensor or probe including disposable and reusable components. The assembly of the disposable and reusable components is straightforward, along with the disassembly thereof. During application to a measurement site, the assembled sensor is advantageously secured together while the componentry is advantageously properly positioned. | 1. A reusable component of a noninvasive optical sensor, the reusable component comprising: at least one energy emitter;at least one detector capable of detecting energy attenuated by body tissue and capable of outputting a signal usable to determine one or more physiological characteristics of the body tissue;a first casing housing one of the at least one energy emitter and the at least one detector, the first casing comprising at least one locking feature configured to interact with a corresponding locking feature of a disposable attachment of the noninvasive optical sensor upon application of the disposable attachment to the patient to fix movement of the first casing and the disposable attachment with respect to one another in at least one direction;a second casing housing the other of the at least one energy emitter and the at least one detector, comprising: a front portion having a bottom surface and a front surface extending from the bottom surface;at least one electrical component capable of electrically communicating with at least one corresponding electrical component in the disposable attachment; andat least one opening in the second casing extending along at least a portion of the bottom surface and at least a portion of front surface, the at least one electrical component of the reusable component accessible through the at least one opening. 2. The reusable component of claim 1, wherein the at least one locking feature of the first casing comprises at least one tab and the at least one locking feature of the disposable attachment comprises at least one groove, the at least one tab capable of sliding along groove upon application of the disposable attachment to the patient. 3. The reusable component of claim 1, further comprising a flexible connector connecting the first casing to the second casing. 4. A method of using the reusable component of claim 1, comprising: mating the first casing with the disposable attachment;mating the second casing with the disposable attachment;attaching the disposable attachment having the first and second casings mated thereto to the patient; andactivating the at least one energy emitter to emit energy into the body tissue. 5. A noninvasive optical sensor having a reusable portion and a disposable portion and capable of outputting a signal indicative of light attenuated by body tissue usable to determine one or more physiological characteristics of the body tissue, the optical sensor comprising: a reusable component including at least one energy emitter and at least one detector capable of detecting energy attenuated by body tissue and capable of outputting a signal usable to determine one or more physiological characteristics of the body tissue, the at least one energy emitter being housed in a first casing and the at least one detector being housed in a second casing; anda disposable component including a first receptacle mechanically matable with the first casing and a second receptacle mechanic... | Cercacor Laboratories, Inc. | A61B005/1455
US20110034798 | A1 | 2011-02-10 | SYSTEMS AND METHODS FOR GUIDING A MEDICAL INSTRUMENT | 2009-10-29 | This document discusses, among other things, a system for guiding a medical instrument within a body of a subject. The system includes at least one magnetic emitter coupled to the medical instrument. The system also includes a plurality of magnetic sensors configured to be placed on the subject and sense at least one magnetic field generated by the at least one magnetic emitter. The system associates a position of each of the plurality of magnetic sensors with an average anatomical representation of the subject. The system also determines a position of the medical instrument relative to the average anatomical representation and the plurality of magnetic sensors. Finally, the average anatomical representation and a representation of the position of the medical instrument relative to the average anatomical representation is displayed for a user. | 1. A system for guiding a medical instrument within a body of a subject, the system comprising: at least one magnetic emitter coupled to the medical instrument, the magnetic emitter configured to generate at least one magnetic field;a plurality of magnetic sensors configured to be placed on the subject, the plurality of magnetic sensors configured to sense the at least one magnetic field generated by the at least one magnetic emitter;an average anatomical representation of at least a portion of the subject; anda processing device communicatively coupled to the plurality of magnetic sensors, the processing device configured to: associate a position of each of the plurality of magnetic sensors with the average anatomical representation; anddetermine a position of the medical instrument relative to the average anatomical representation and the plurality of magnetic sensors using data from the plurality of magnetic sensors regarding the magnetic field generated by the magnetic emitter; anda display device communicatively coupled to the processing device, wherein the processing device is configured to direct the display device to display the average anatomical representation and to display a representation of the position of the medical instrument relative to the average anatomical representation. 2. The system of claim 1, wherein each of the plurality of magnetic sensors is attached to the subject at a known anatomical location such that the relationship between a position of each of the plurality of magnetic sensors and the average anatomical representation is known. 3. The system of claim 1, wherein the plurality of magnetic sensors comprise 3-axis magnetometers. 4. The system of claim 1, wherein the at least one magnetic emitter comprises a plurality of current-carrying coils. 5. The system of claim 1, wherein the at least one processing device is configured to direct a display device to: display a target position for the medical instrument relative to the average anatomical representation of the subject. 6. The system of claim 1, wherein the processing device is further configured to: determine an orientation of the medical instrument using data from the plurality of magnetic sensors regarding the magnetic field generated by the magnetic emitter. 7. The system of claim 6, wherein the processing device is further configured to direct the display device to: display a projected trajectory of the medical instrument using the position of the medical instrument and the orientation of the medical instrument. 8. The system of claim 7, wherein the processing device is further configured to direct the display device to: display a target position for the medical instrument relative to the projected trajectory of the medical instrument. 9. A method for guiding a medical instrument within a body of a subject, the method comprising: providing at least one magnetic emitter coupled to the medical instrument;providing an average anatomical representation ... | - | A61B005/05
US20110009731 | A1 | 2011-01-13 | METHOD AND SYSTEM FOR PERFORMING UPRIGHT MAGNETIC RESONANCE IMAGING OF VARIOUS ANATOMICAL AND PHYSIOLOGICAL CONDITIONS | 2010-07-08 | Vasculature or parenchyma is imaged using upright MRI techniques, on patients who may have conditions such as congestive heart failure, or otherwise be healthy. When an individual is horizontal, venous drainage is minimized, causing the vessels to remain engorged, also referred to herein as vascular congestion. Vascular congestion results in an enlarging of the vessels and surrounding tissue causing the vessels to be more visible on MRIs. The decrease in vascular visibility in upright subjects is in part, due to an increase in venous drainage. Patients suffering from coronary and/or pulmonary deficiencies (e.g. CHF) experience decreased rates and degrees of venous drainage. In one embodiment, the present invention uses upright imaging to visualize these enlarged vessels. | 1. A method for evaluating vasculature using magnetic resonance imaging, comprising the steps of: placing a patient into a magnetic resonance imaging apparatus with the patient in an upright position;capturing magnetic resonance images of the patient's vasculature; andevaluating the patient's vasculature based on or more magnetic resonance, signal characteristics of the captured images. 2. The method of claim 1, wherein the vasculature is lung vasculature. 3. The method of claim 1, wherein the vasculature is kidney vasculature. 4. The method of claim 1, wherein the vasculature is liver vasculature. 5. The method of claim 1, wherein the evaluation of the patient's vasculature is used to diagnose congestive heart failure. 7. The method of claim 1, wherein the signal characteristic is signal intensity. 8. The method of claim 1, wherein the signal characteristic is signal frequency. 9. The method of claim 1, wherein the signal characteristic is signal amplitude. 10. The method of claim 1, wherein the signal characteristic is signal phase. 11. A method for evaluating parenchyma from magnetic resonance images, comprising the steps of: placing a patient into a magnetic resonance imaging apparatus with the patient in an upright position;capturing magnetic resonance images of the parenchyma of a selected organ or organs; andusing one or more magnetic resonance signal characteristics to evaluate the imaged parenchyma. 12. The method of claim 11, wherein the parenchyma is lung parenchyma. 13. The method of claim 11, wherein the parenchyma is kidney parenchyma. 14. The method of claim 11, wherein the vasculature is liver parenchyma. 15. The method of claim 11, wherein the evaluation of the imaged parenchyma is used to diagnose congestive heart failure. 16. The method of claim 11, wherein the evaluation of the imaged parenchyma is used to evaluate gravitational effects on the human lungs. 17. The method of claim 11, wherein the signal characteristic is signal intensity. 18. The method of claim 11, wherein the signal characteristic is signal frequency. 19. The method of claim 11, wherein the signal characteristic is signal amplitude. 20. The method of claim 11, wherein the signal characteristic is signal phase. 21. A method for evaluating vasculature and parenchyma from magnetic resonance images, comprising the steps of: placing a patient into a magnetic resonance imaging apparatus with the patient in an upright position;capturing magnetic resonance images of the patient's vasculature and parenchyma of selected a selected organ or organs; andusing one or more magnetic resonance signal characteristics to evaluate the patient's vasculature and parenchyma. 22. The method of claim 21, wherein said evaluation of vasculature and parenchyma is used to diagnose congestive heart failure. 23. A magnetic resonance imaging magnet comprising: a magnet frame having a pair of pole faces spaced apart from one another along a horizontal pole axis and defin... | Fonar Corporation | A61B005/055
US20110046501 | A1 | 2011-02-24 | OPTIMIZATION OF A FLOW SENSOR | 2008-10-29 | A sensor for a flow measuring apparatus for lung function diagnostics or performance diagnostics, which sensor has an inner chamber, in particular an inner chamber which is oblong in the direction of flow, and has a resistor for producing a differential pressure between an inner chamber section located upstream from the resistor and between an inner chamber section located downstream from the resistor, which resistor consists of a planar component permeated by openings, which resistor is arranged at an angle between 20° and less than 90° to the direction of flow or to the longitudinal extension of the inner chamber. | 1. A sensor (11) for a flow measuring apparatus for lung function diagnostics or performance diagnostics, which sensor has an inner chamber which is oblong in the direction of flow (14), and has a resistor (13) for producing a differential pressure between an inner chamber section located upstream from the resistor and an inner chamber section located downstream from the resistor, which resistor (13) consists of a planar component permeated by openings, characterized by the resistor (13) being arranged at an angle (15) between 20° and less than 90° to the direction of flow (14) or to the longitudinal extension of the inner chamber. 2. The sensor (11) according to claim 1, wherein the resistor (13) is arranged at an angle (15) between 30° and 80°. 3. The sensor (11) according to claim 2, wherein the inner chamber has a cross-sectional surface vertical to the direction of flow (14) or to a longitudinal extension of the inner chamber of 1200 to 2400 square millimeters. 4. The sensor (11) according to claim 3 wherein the inner chamber has an elliptical cross section vertical to the direction of flow (14) or the longitudinal extension. 5. The sensor (11) according to claim 4, wherein the shorter main axis of the elliptical cross section is between 17 mm and 24 mm long, and the longer main axis of the elliptical cross section is between 22 mm and 32 mm long. 6. The sensor (11) according to claim 5, wherein the shorter main axis is 21 mm long and the longer main axis is 28 mm long. 7. The sensor (11) according to claim 4 wherein the resistor 13 has an angle (15) of greater than 20° and less than 90° relative to both main axes. 8. The sensor (11) according to claim 4, wherein the resistor (13) is arranged at an angle (15) between 40° and 50° to the direction of flow. 9. The sensor (11) according to claim 8, wherein the shorter main axis of the elliptical cross section is between 19 mm and 22 mm long, and the longer main axis of the elliptical cross section is between 25 mm and 29 mm long and the inner chamber has a cross-sectional surface vertical to the direction of flow (14) or to a longitudinal extension of the inner chamber of 1600 to 2000 square millimeters. 10. The sensor (11) according to claim 9, wherein the angle (15) at which the resistor (13) is arranged is 45° to the direction of flow. | - | A61B005/087
US20110046534 | A1 | 2011-02-24 | DEVICE AND METHOD FOR MONITORING AN ACCESS TO A PATIENT, IN PARTICULAR A VASCULAR ACCESS DURING AN EXTRACORPOREAL BLOOD TREATMENT | 2009-03-03 | The invention relates to the monitoring of an access to a patient at a puncture point, in particular the monitoring of the arterial and/or venous vascular access during an extracorporeal blood treatment. The device according to the invention is a magnetoelastic moisture sensor, which is embedded in an absorbent material and is laid on the puncture point. For the case in which infusion solution and/or blood exits from the puncture point, the incorrect seating of the puncture cannula can be recognized reliably using the device according to the invention. The absorbent material having the embedded magnetoelastic moisture sensor can be laid like conventional bandage material or a gauze bandage on the puncture point. | 1-15. (canceled) 16. A detection device for the detection of fluid, said detection device configured for use with a monitoring device for monitoring an access to a patient whereby fluid is provided to a patient or carried away from the patient via a tube line, the detection device comprising: a magneto-elastic moisture sensor embedded in an absorbent material. 17. The detecting device according to claim 16, wherein the device for the detection of fluid further comprises: a means for fixing the absorbent material to the patient. 18. The detecting device according to claim 17, wherein the means for fixing are designed such that the absorbent material can be fixed to the patient's skin in an adhesive manner. 19. A monitoring device for monitoring an access to a patient comprising: the detection device of claim 16;a measuring unit for measuring the resonance frequency of the magneto-elastic moisture sensor; andan evaluation unit for comparing the measured resonance frequency with a preset reference value, the evaluation unit configured to generate a control or alarm signal, which signals a defective patient access, if the amount of the difference between the measured resonance frequency and the preset reference value is greater than a preset threshold value. 20. The monitoring device according to claim 19, wherein the measuring unit for measuring the resonance frequency comprises: an exciter coil for generating an alternating magnetic field; anda detector coil for detecting an alternating magnetic field. 21. A blood treatment apparatus with an extracorporeal blood circuit, comprising: an arterial tube line with an arterial puncture cannula;a venous tube line with a venous puncture cannula; andthe monitoring device for monitoring the arterial or venous vascular access of claim 32. 22. The blood treatment apparatus according to claim 21, wherein the blood treatment apparatus further comprises: a control unit that cooperates with the monitoring device in such a way that the blood treatment is interrupted or an alarm is emitted when the control or alarm signal is generated. 23. The blood treatment apparatus according to claim 22, wherein the blood treatment apparatus further comprises: a blood pump for conveying blood through the extracorporeal blood circuit,wherein the control unit cooperates with the monitoring device in such a way that the blood pump is stopped when the evaluation unit generates the control or alarm signal. 24. The blood treatment apparatus according to claim 22, wherein the blood treatment apparatus further comprises: a shut-off element disposed in the extracorporeal blood circuit,wherein the control unit cooperates with the monitoring device in such a way that the shut-off element is closed when the evaluation unit generates the control or alarm signal. 25. The blood treatment apparatus according to claim 22, wherein the blood treatment apparatus further comprises: an alarm unit which emits an acoustic and/or optical alarm... | - | A61M001/00 A61B005/05
US07881768 | B2 | 2011-02-01 | Assessing the condition of a joint and devising treatment | 2007-04-24 | Methods are disclosed for assessing the condition of a cartilage in a joint, particularly in a human knee. The methods include converting an image such as an MRI to a three dimensional map of the cartilage. The cartilage map can be correlated to a movement pattern of the joint to assess the affect of movement on cartilage wear. Changes in the thickness of cartilage over time can be determined so that therapies can be provided. Information on thickness of cartilage and curvature of cartilage or subchondral bone can be used to plan therapy. Information on movement pattern can be used to plan therapy. | 1. A method of designing a patient-specific implant for treating an articular surface of a joint of a patient, the method comprising: using a computer to obtain electronic patient-specific image data of dimensions of at least a portion of the articular surface of said joint of said patient;deriving from said electronic patient-specific image data a patient-specific model of at least a portion of the articular surface of said joint; anddesigning a patient-specific implant configured to replace at least a portion of said articular surface of said joint based on said patient-specific model of the at least a portion of the articular surface;wherein said implant provides the articular surface derived at least in part from the dimensions of the at least a portion of the articular surface of the joint. 2. The method of claim 1, wherein designing the implant includes determining the shape of the implant. 3. The method of claim 1, wherein said image data includes both normal and diseased cartilage tissue. 4. The method of claim 1, wherein said articular surface includes bone. 5. The method of claim 1, wherein said articular surface includes cartilage. 6. The method of claim 1, wherein said patient specific model is derived, at least in part, from bone. 7. The method of claim 1, wherein said patient specific model is derived, at least in part, from cartilage. 8. The method of claim 1, wherein more than 30% of said implant has a thickness that is similar to normal cartilage. 9. The method of claim 1, wherein more than 50% of said implant has a thickness that is similar to normal cartilage. 10. The method of claim 1, wherein at least more than 70% of said implant has a thickness that is similar to normal cartilage. 11. The method of claim 1, wherein the implant is for one of a tibio-femoral articulation, a hip, and an ankle. 12. The method of claim 11, wherein more than 30% of said implant has a thickness that is similar to normal cartilage. 13. The method of claim 11, wherein more than 50% of said implant has a thickness that is similar to normal cartilage. 14. The method of claim 11, wherein at least more than 70% of said implant has a thickness that is similar to normal cartilage. 15. The method of claim 1, wherein at least a portion of said implant has a thickness that is similar to normal cartilage. 16. The method of claim 1, wherein substantially each portion of said implant has a thickness that is similar to normal cartilage. 17. The method of claim 1, wherein at least a portion of an undersurface of said implant substantially matches subchondral bone. 18. A method of creating a patient-specific implant, the method comprising: obtaining electronic cross-sectional or volumetric image data of at least a portion of an articular surface of a joint of a patient;deriving from the electronic image data a plurality of points of the at least a portion of the articular surface of said joint;deriving at least one or more patient-specific articul... | The Board of Trustees of the Leland Stanford Junior University | A61B005/00
US08061004 | B2 | 2011-11-22 | Method of manufacturing a test strip | 2009-08-20 | The test strip stabilizes skin prior to incision by a lancet tip. The innovative test strip includes a pair of skin contacting tabs that apply pressure linearly to the skin to stretch the skin taut. Beneficially an incision by a lancet tip to an accurate depth can be formed on the taut skin. Moreover, pressure from the tabs is applied linearly to the skin which results in a higher ratio of pressure to surface that can be reached as compared to applying pressure laminarily to skin. Some manufacturing techniques of the skin contacting tabs include cutting or stamping through all of the layers and test area of the test strip to form these tabs. These techniques increase the manufacturing efficiency and cost savings associated with test strips and integrated lancet testing devices. Contamination from other sampling events is eliminated since the test strips are disposed of after each use. | 1. A method of manufacturing a test strip, comprising: assembling said test strip having a test area including a test chemistry, said test area configured to collect a body fluid sample via capillary action and analyze said body fluid sample, andcutting through said test strip to form a first skin contacting tab and a second skin contacting tab after said assembling. 2. The method of claim 1, wherein said assembling includes: disposing said test area onto either a first panel or a second panel; andlaminating said first panel to said second panel. 3. The method of claim 1, wherein said cutting includes stamping the test strip. 4. The method of claim 1, wherein said cutting includes cutting through said test area. 5. The method of claim 1, further comprising: attaching a lancet to said test strip to form an integrated lancet test strip. 6. A method of manufacturing a plurality of test strips, comprising: providing a first panel and a second panel wherein each of said first panel and said second panel includes a continuous sheet of material;disposing a test area having a reagent onto either said first panel or said second panel;assembling said first panel with said second panel;cutting through said first panel and said second panel to form a plurality of test strips after said assembling; andcutting through said plurality of test strips to form a first skin contacting tab and a second skin contacting tab for each of said plurality of test strips. 7. The method of claim 6, wherein said cutting includes cutting through said test area. | Roche Diagnostics Operations, Inc. | B21D053/00 A61B005/00
US07865228 | B2 | 2011-01-04 | Multi-slice cerebral blood flow imaging with continuous arterial spin labeling MRI | 2008-03-13 | This invention is a method for multi-slice CBF imaging using continuous arterial spin labeling (CASL) with an amplitude modulated control which is both highly effective at controlling for off resonance effects, and efficient at doubly inverting inflow spins, thus retaining the signal advantages of CASL versus pulsed ASL techniques. | 1. A method for generating a blood flow image, the method comprising: perturbing arterial spins of blood flowing into a sample by applying a constant RF irradiation together with a magnetic field gradient;waiting a transit delay period before acquiring a first image of the sample, wherein the transit delay period is of a duration that allows for acquisition of the first image while the blood having perturbed arterial spins is exhibiting a desired flow characteristic within the sample;acquiring a first image of the sample;applying amplitude modulated RF irradiation with a magnetic field gradient which, together, mimic the effects of constant RF radiation;acquiring a second image of the sample;generating a blood flow image based on a difference between the first image and the second image. 2. The method of claim 1, wherein the sample comprises a blood vessel and a tissue portion, and the transit delay period is of a duration that allows for acquisition of the first image while the blood having perturbed arterial spins flows in the blood vessel. 3. The method of claim 1, wherein the sample comprises a blood vessel and a tissue portion, and the transit delay period is of a duration that allows for acquisition of the first image while the blood having perturbed arterial spins flows in the tissue portion of the sample. | The Trustees of the University of Pennsylvania | A61B005/055
US07894883 | B2 | 2011-02-22 | Method and apparatus for post-processing of episodes detected by a medical device | 2006-11-28 | A system and method for post-processing of sensing data generated by a medical device that includes transmitting a plurality of stored sensing data generated by a medical device to an access device, the stored sensing data including sensed atrial events and sensed ventricular events. The access device determines, in response to the transmitted data, instances where the medical device identified a cardiac event being detected in response to the sensing data, and verifies, in response to the transmitted data, whether the determined instances should have been identified by the medical device as a cardiac event being detected. | 1. A method of post-processing of sensing data generated by and stored within an implantable medical device, comprising: receiving the sensing data from the implantable medical device with an external access device, wherein the sensing data includes sensed atrial events and sensed ventricular events;determining, with the external access device, in response to the received data, instances where the implantable medical device identified a cardiac event being detected in response to the sensing data; andverifying, with the external access device, in response to the received data, whether the determined instances should have been identified by the implantable medical device as the cardiac event being detected,wherein verifying whether the determined instances should have been identified by the implantable medical device as the cardiac event being detected comprises classifying, with the external access device the cardiac event as one of ventricular tachycardia, ventricular fibrillation, supraventricular tachycardia, or unknown, wherein classifying the cardiac event is based on one or more of: determining, with the external access device, whether the sensed atrial events were both spontaneous and regular;determining, with the external access device, a ratio of sensed atrial events to sensed ventricular events in response to determining that the sensed atrial events were both spontaneous and regular;determining, with the external access device, whether the sensed atrial events are evenly distributed with the sensed ventricular events in response to determining that the ratio of sensed atrial events to sensed ventricular events is 1:1, wherein determining whether the sensed atrial events are evenly distributed with the sensed ventricular events comprises determining whether there is only one sensed atrial event located between each adjacent pair of sensed ventricular events;determining, with the external access device, whether there was an abrupt onset of the cardiac event in response to determining that the sensed atrial events were not evenly distributed with the sensed ventricular events;determining, with the external access device, whether one of the sensed ventricular events or the sensed atrial events initiated conduction of a heart associated with the sensing data during the cardiac event in response to determining that there was an abrupt onset of the cardiac event;comparing, with the external access device, morphologies of intervals within the cardiac event in response to determining that the ratio of sensed atrial events to sensed ventricular events is not 1:1. 2. The method of claim 1, wherein comparing morphologies of intervals further comprises generating a template in response to correlated morphologies of adjacent intervals prior to a detection interval corresponding to the cardiac being identified as the cardiac event and a morphology of the detection interval. 3. The method of claim 2, wherein generating a template comprises: determin... | Medtronic, Inc. | A61B005/04
US07892183 | B2 | 2011-02-22 | Method and apparatus for body fluid sampling and analyte sensing | 2003-07-03 | A body fluid sampling system is provided for use on a tissue site. In one embodiment, the system comprises a cartridge; a penetrating member driver; a plurality of penetrating members arranged in a radial configuration on the cartridge wherein sharpened distal tips of the penetrating members point radially outward; wherein an active one of the penetrating members may be operatively coupled to the penetrating member driver, the penetrating member driver moving the active one along a path out of a housing having a penetrating member exit, into the tissue site, stopping in the tissue site, and withdrawing out of the tissue site; and a plurality of analyte detecting members, wherein at least one of the analyte detecting members is positioned to receive fluid from a wound created by the active one of the penetrating members, wherein the detecting members are not pierced by the active one of the penetrating members. | 1. A body fluid sampling system for use on a tissue site, the system comprising: a disposable;a penetrating member driver;a plurality of penetrating members arranged in a radial configuration in the disposable, wherein sharpened distal tips of the penetrating members point radially outward, and wherein each of the plurality of penetrating members has a packing density of about 0.1 cm3;wherein an active one of said penetrating members is operatively coupled to said penetrating member driver, said penetrating member driver moving said active one of said penetrating members along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site;a processor coupled to the penetrating member driver and configured to provide instructions to the penetrating member driver for a fast-into of said active one penetrating members into the tissue site and slow-out velocity out of the tissue site; anda plurality of analyte detecting members positioned in the disposable, wherein at least one of said analyte detecting members is positioned to receive less than about 0.1 μL of body fluid from a wound created by said active one of said penetrating members, and wherein said plurality of analyte detecting members are not pierced by the active one of the penetrating members. 2. A system as in claim 1, further comprising a penetrating member sensor positioned to monitor the active one of said penetrating members coupled to said penetrating member driver, the penetrating member sensor configured to provide information relative to a depth of penetration of the active one of said penetrating members through a skin surface of the tissue site. 3. The system of claim 2, wherein the depth of penetration is about 100 to 2500 microns. 4. The system of claim 2, wherein the depth of penetration is 500 to 750 microns. 5. The system of claim 2, wherein the depth of penetration is no more than about 1000 microns beyond a stratum corneum thickness of the skin surface. 6. The system of claim 2, wherein the depth of penetration is no more than about 500 microns beyond a stratum corneum thickness of the skin surface. 7. The system of claim 1, wherein the penetrating member driver is selected from one of the following: a voice coil, a rotary voice coil, a solenoid, a motor and gear box, a nanomuscle, or a combination of any of the above. 8. The system of claim 1, wherein the processor is utilized to monitor position and speed of the active one of said penetrating members as the active one of said penetrating member moves in a first direction. 9. The system of claim 1, wherein the processor is utilized to monitor position and speed of the active one of said penetrating members as the active one of said penetrating members moves in a first direction toward the tissue site, wherein an application of a launching force to the active one of said penetrating members is controlled based on a position an... | Pelikan Technologies, Inc. | A61B005/00 A61B017/14 A61B017/32 B65D081/00
US07873407 | B2 | 2011-01-18 | Systems and methods for in-vivo optical imaging and measurement | 2005-12-06 | Disclosed are methods and systems for: (i) positioning a specimen inside an optical measurement system according to a reference image of the specimen indicative of its position and orientation during an earlier measurement using the optical measurement system; and (ii) measuring radiation emitted from the positioned specimen to provide information about an object embedded in the specimen. | 1. A method comprising: positioning a specimen inside an optical measurement system according to a reference image of the specimen indicative of the specimen's position and orientation during an earlier measurement using the optical measurement system; andmeasuring radiation emitted from the positioned specimen to provide information about an object embedded in the specimen. 2. The method of claim 1, wherein the reference image was recorded more than one hour prior to the present positioning of the specimen. 3. The method of claim 1, wherein the reference image was recorded more than one day prior to the present positioning of the specimen. 4. The method of claim 1, wherein the reference image was recorded based on light reflected or scattered from the specimen. 5. The method of claim 1, wherein the reference image was recorded based on fluorescent light emitted from the specimen. 6. The method of claim 1, wherein the emitted radiation is fluorescence or bioluminescence. 7. The method of claim 1, further comprising illuminating the positioned specimen to cause the emitted radiation. 8. The method of claim 1, wherein the emitted radiation comprises radiation emitted from the object embedded in the specimen. 9. The method of claim 1, wherein the specimen is a living animal. 10. The method of claim 9, wherein the object is a tumor embedded in the animal. 11. The method of claim 1, further comprising comparing the information about the object to information derived from the earlier measurement of the specimen. 12. The method of claim 1, further comprising at a later time, positioning the specimen inside the optical measurement system according to the reference image for a subsequent measurement using the optical measurement system to provide information about the object embedded in the specimen at a later time. 13. The method of claim 1, wherein the information about the object comprises information about the size, position, or shape of the object. 14. The method of claim 1, further comprising performing the earlier measurement and recording the reference image. 15. The method of claim 1, wherein the specimen is positioned to match the position and orientation of the reference image. 16. The method of claim 1, further comprising obtaining a live image of the specimen while positioning the specimen, and positioning the specimen based on a comparison between the live image and the reference image. 17. The method of claim 16, further comprising simultaneously displaying the live image and the reference image to aid in the positioning of the specimen. 18. The method of claim 17, further comprising superimposing the live image and the reference image to aid in the positioning of the specimen. 19. The method of claim 16, further comprising processing the live image to highlight edges or other features of the specimen to aid in the positioning. 20. The method of claim 16, wherein the comparison comprises providing a numerical index, based... | Cambridge Research & Instrumentation, Inc. | A61B006/00 A61B005/055 G01V008/00
US07865003 | B2 | 2011-01-04 | Image evaluation method for two-dimensional projection images and objects corresponding thereto | 2007-05-31 | 2-D projection images show the temporal profile of the distribution of a contrast medium in an examination object, which contains a vascular system and its surroundings. Each projection image comprises pixels with pixel values. The pixel values of pixels corresponding to one another in the projection images are defined by at least essentially locationally identical areas of the examination object. A computer assigns a uniform 2-D evaluation core that is uniform for all corresponding pixels at least in a sub-area of pixels corresponding to one another in the projection images that is uniform for the projection images. The computer defines at least one characteristic value for each pixel within each projection image based on the evaluation core assigned to the pixel and assigns it to the relevant pixel. Based on the temporal profile of the characteristic values, the computer defines parameters of at least one function of time, so that any deviation between the function parameterized with the parameters and the temporal profile of the characteristic values is minimized. Based on the parameters the computer defines a type and/or an extent and assigns them to a pixel of a 2-D evaluation image corresponding to the pixels of the projection images. The type indicates whether the respective pixel of the evaluation image corresponds to a vessel of the vascular system, a perfused part or a non-perfused part of the surroundings of a vessel of the vascular system. The extent is characteristic of perfusion. The computer outputs at least the sub-area of the evaluation image to a user via a display device. | 1. An image evaluation method for two-dimensional projection images to show a temporal profile of a distribution of a contrast medium in an examination object, wherein the examination object has a vascular system and a surrounding, wherein each projection image comprises a plurality of pixels with pixel values, wherein the pixel values of pixels corresponding to one another in the projection images are defined by at least essentially locationally identical areas of the examination object, comprising: assigning a uniform two-dimensional image evaluation core for all corresponding pixels to pixels corresponding to one another in the projection images at least in a sub-area of the projection images that is uniform for the projection images;determining at least one characteristic value for the pixel within the projection image based upon the evaluation core assigned to the pixel within the projection image;assigning the characteristic value to the pixel within the projection image;determining parameters of a function in time based upon a temporal profile of characteristic values assigned to corresponding pixels to minimize a deviation between the function of time parameterized with the parameters and the temporal profile of the characteristic values assigned to corresponding pixels;defining a type or an extent based upon the parameters and assigning the type or extent to a pixel of a two-dimensional evaluation image, wherein the type is distinctive of whether the respective pixel of the evaluation image corresponds to a vessel of the vascular system, a perfused part of the surroundings of a vessel of the vascular system or a non-perfused part of the surroundings of a vessel of the vascular system, and wherein the extent is distinctive for an extent of perfusion; andoutputting at least the sub-area of the evaluation image to a user via a display device. 2. The image evaluation method as claimed in claim 1, wherein a computer assigns a maximum of the pixel values in the evaluation core assigned to the respective pixel to the pixels as a characteristic value. 3. The image evaluation method as claimed in claim 2, wherein the computer defines parameters of a maximum function based upon a temporal profile of the maximums to minimize a deviation between the function of time parameterized with the parameters and the temporal profile of the maximum, and wherein the computer determines based upon the parameters whether the type “vessel” is assigned to the corresponding pixel of the evaluation image. 4. The image evaluation method as claimed in claim 1, wherein a mean value of the pixel values occurring in the evaluation core assigned to the respective pixel is assigned to the pixels as a distinctive value. 5. The image evaluation method as claimed in claim 4, wherein parameters of a mean value function are determined based upon the temporal profile of the mean value to minimize a deviation between the mean value function parameterized with the parameters a... | Siemens Aktiengesellschaft | G06K009/00 A61B005/02
US07878981 | B2 | 2011-02-01 | Systems and methods for intra-operative stimulation | 2007-01-09 | Improved assemblies, systems, and methods provide safeguarding against tissue injury during surgical procedures and/or identify nerve damage occurring prior to surgery and/or verify range of motion or attributes of muscle contraction during reconstructive surgery. A stimulation control device may incorporate a range of low and high intensity stimulation to provide a stimulation and evaluation of both nerves and muscles. A stimulation control device is removably coupled to a surgical device or is imbedded within the medical device to provide a stimulation and treatment medical device. A disposable hand held stimulation system includes an operative element extending from the housing, the housing includes an operative element adjustment portion and a visual indication to provide feedback or status to the user. | 1. A hand-held tissue stimulation system comprising: a tubular housing having a proximal end and a distal end, the housing being sized and configured to be held and controlled entirely within a surgical field and in a single hand by a user, the housing including a flexible probe adjustment portion including an elastic material extending from the distal end,a stimulation control device positioned entirely within the housing, the stimulation control device comprising stimulation signal generating circuitry for generation of an electrical stimulation signal having an amplitude and a duration, a microcontroller electrically coupled to the stimulation signal generating circuitry, and a source of power for the stimulation control device,a probe extending from the flexible probe adjustment of the housing and having an electrically conductive surface coupled to the signal generating circuitry, the electrically conductive surface being sized and configured for contact with a targeted tissue region and, when in contact, applying the electrical stimulation signal to the targeted tissue region, the flexible probe adjustment portion being sized and configured to elastically flex in response to pressure applied by a finger or a thumb of the hand holding the housing to adjust position of the electrically conductive surface at the targeted tissue region,a return electrode electrically coupled to the signal generating circuitry and being sized and configured for contact with tissue to provide an electrical flow path for the stimulation signal from the probe back to the stimulation control device,a first control device carried on the housing and electrically coupled to the microcontroller, the first control device being movable by a user from a first position to a second position for providing a power control input to the microcontroller to turn the stimulation signal generating circuitry on only for a predetermined total single use period, the first control device also being movable from the second position back to the first position for providing a control input to the microcontroller to place the stimulation signal generating circuitry in a power-on, standby condition during the predetermined total single use period, the first control device also being movable from the second position through a plurality of control positions for providing amplitude control inputs to the microcontroller to control selection of the amplitude within a predefined amplitude range during the total single use period,a second control device carried on the housing separate from the first control device and being electrically coupled to the microcontroller, the second control device comprising a potentiometer operable by a user during the total single use period independent of the first control device within a range of control positions for providing duration control inputs to the microcontroller to control adjustment of the duration from a predefined duration range, anda single visual i... | Checkpoint Surgical, LLC | A61B005/05 A61N001/00
US07877276 | B2 | 2011-01-25 | Messaging to remote patients in a networked health-monitoring system | 2006-02-27 | A networked health-monitoring system configured to collect and process patient health-related data. A plurality of remote patient sites each includes at least one display, a data management unit for collecting patient health-related data, a memory, and stored program instructions for generating health-monitoring related information on the display. A central server connects to the data management unit at each patient site. The system transmits messages for display at the patient sites and at least one of these messages is a message selected by a health care professional. | 1. A networked health-monitoring system, comprising: a plurality of remote sites, each site of said plurality of remote sites including a device having (i) at least one display, (ii) a data management unit configured to facilitate collection of health-related data of an individual, (iii) a memory circuit, and (iv) stored program instructions for collecting said health-related data and generating information in response to said health-related data; andat least one central server connectable for communication with the device at each remote site of said plurality of remote sites wherein(A) said at least one central server is configured to (i) receive said health-related data from said device, (ii) analyze said health-related data and transmit said analysis of said health-related data to one or more computers remotely located from and in signal communication with (a) said at least one central server and (b) said device, said one or more computers corresponding to one or more health care professionals associated with said individual, and (iii) transmit at least one message for presentation on said at least one display and (iv) said at least one message is sent by said one or more health care professionals in response to said analysis of said health-related data and(B) (i) the data management unit is physically separate from the at least one display, (ii) the memory circuit and the display form a part of the device, (iii) the device comprises a handheld device, and (iv) the memory circuit comprises a program cartridge. 2. The system of claim 1, wherein the at least one message includes step-by-step instructions. 3. The system of claim 1, wherein the at least one message includes results of a test. 4. The system of claim 1, wherein the at least one message is a multi-line message. 5. The system of claim 1 wherein the one or more health care professionals generate the at least one message. 6. The system of claim 1, wherein the at least one message is educational or motivational. 7. The system of claim 1, wherein the at least one central server is configured to transmit the at least one message to a specific individual. 8. The system of claim 7, wherein the at least one central server is configured to transmit said at least one message automatically to the specific individual. 9. The system of claim 7, wherein the at least one central server enables the specific individual to choose when to receive the at least one message. 10. The system of claim 7, wherein the at least one message can be stored before being transmitted to the specific individual. 11. The system of claim 1, wherein the device at each patient site is configured to monitor the health-related data of said individual. 12. The system of claim 11, wherein the data management unit facilitates collection of said health-related data by receiving said health-related data from the device. 13. The system of claim 11, wherein the device includes one or more of the set consisting of: (i) a... | Health Hero Network, Inc. | A61B005/00 G06Q010/00 G06F019/00
US07867172 | B1 | 2011-01-11 | Combination toothbrush and peak flow meter system | 2006-11-09 | A combination toothbrush and peak flow meter system for increasing the compliance of peak flow measurements in children and adults with asthma. The combination toothbrush and peak flow meter system includes a peak flow meter and a toothbrush head connected to an end of the peak flow meter. | 1. A combination toothbrush and peak flow meter system, comprising: a peak flow meter to measure a peak flow rate, wherein said peak flow meter includes an intake port and an exit port; anda toothbrush head connected to an end of said peak flow meter, wherein said toothbrush head includes a plurality of bristles. 2. The combination toothbrush and peak flow meter system of claim 1, wherein said toothbrush head is concentrically attached to said peak flow meter. 3. The combination toothbrush and peak flow meter system of claim 1, wherein said exit port extends rearwardly from a rear portion of said peak flow meter. 4. The combination toothbrush and peak flow meter system of claim 3, wherein said plurality of bristles of said toothbrush head extend forwardly from a front portion of said toothbrush head. 5. The combination toothbrush and peak flow meter system of claim 1, wherein said intake port is positioned opposite of said toothbrush head. 6. The combination toothbrush and peak flow meter system of claim 1, wherein said peak flow meter is comprised of a digital unit. 7. The combination toothbrush and peak flow meter system of claim 6, wherein said peak flow meter includes a display to show a measured peak flow rate. 8. The combination toothbrush and peak flow meter system of claim 1, wherein said peak flow meter is comprised of a mechanical analog unit. 9. The combination toothbrush and peak flow meter system of claim 1, wherein said toothbrush head includes a neck portion. 10. The combination toothbrush and peak flow meter system of claim 1, wherein said peak flow meter is comprised of an elongated structure. 11. A combination toothbrush and peak flow meter system, comprising: a peak flow meter to measure a peak flow rate, wherein said peak flow meter includes an intake port, an exit port and a second connecting end, and wherein said peak flow meter is comprised of an elongated structure; anda toothbrush head including a first connecting end, wherein said second connecting end is detachably connected to said first connecting end of said toothbrush head, wherein said toothbrush head includes a plurality of bristles. 12. The combination toothbrush and peak flow meter system of claim 11, wherein said toothbrush head is concentrically attached to said peak flow meter. 13. The combination toothbrush and peak flow meter system of claim 11, wherein said exit port extends rearwardly from a rear portion of said peak flow meter. 14. The combination toothbrush and peak flow meter system of claim 13, wherein said plurality of bristles of said toothbrush head extend forwardly from a front portion of said toothbrush head. 15. The combination toothbrush and peak flow meter system of claim 11, wherein said intake port is positioned opposite of said toothbrush head. 16. The combination toothbrush and peak flow meter system of claim 11, wherein said peak flow meter is comprised of a digital unit and wherein said peak flow meter includes a display t... | - | A61B005/08
US08799010 | B2 | 2014-08-05 | Telehealth scheduling and communications network | 2009-05-07 | A telehealth communications network accessible anywhere sufficient bandwidth connection is available facilitates a medical encounter between a patient at a patient site and a provider at a provider site remotely located from the patient, and includes an operating console comprising at least one memory and at least one processor coupled to the network, the operating console configured to execute at least one workflow for facilitating the medical encounter. | 1. A telehealth communications network accessible over the Internet for facilitating a medical encounter between a patient at a patient site and a provider at a provider site remotely located from the patient, the telehealth communications network comprising: an operating console comprising at least one memory and at least one processor coupled to the network, the operating console configured to execute a plurality of workflows, the plurality of workflows comprising a series of graphical user interfaces;wherein the workflows are instantiated based on a type of a user, the type of user determined based on an originating Internet protocol (“IP”) address and whether the IP address is associated with the patient site or the provider site;wherein one series of graphical user interfaces are configured to be viewable by the patient at the patient site and another series of the graphical user interfaces are configured to be viewable by the provider at the provider site remotely located from the patient, the graphical user interfaces for facilitating and guiding interactions between the patient and the provider during the medical encounter; andwherein the network is configured to command medical devices and network components coupled thereto, and at least one of the graphical user interfaces at the patient site is configured to be displayed simultaneously with a graphical user interface at the provider site and each comprises at least one image that guides use of at least one of the medical devices at the patient site. 2. The telehealth communications network of claim 1, further comprising a platform communicatively coupled to the operating console for providing an interface to operate the medical devices and the network components. 3. The telehealth communications network of claim 2, further comprising a clinic workstation for supporting the computational needs of the patient site and the provider site, wherein the clinic workstation comprises a device adaptor aggregator for communicatively coupling the medical devices. 4. The telehealth communications network of claim 1, further comprising a unified communication console configured to present data to the patient and the provider and to enable interactions between devices at the patient site and at the provider site. 5. The telehealth communications network of claim 4, wherein the unified communication console comprises a conference engine interface for managing a video conferencing connection. 6. The telehealth communications network of claim 4, wherein the unified communication console is communicatively coupled to a telemetry display for displaying telemetry data and for providing a user interface enabling the patient and provider to interact with the telemetry display settings and output. 7. The telehealth communications network of claim 6, further comprising a telemetry control for coupling the telemetry display to the medical devices. 8. The telehealth communications network of claim 7, furt... | UnitedHealth Group Incorporated | G06Q010/00 G06Q050/00
US08862214 | B2 | 2014-10-14 | Cardiac condition detection system | 2009-04-02 | A system is provided which processes cardiac condition data representative of heart electrical signals to identify a cardiac condition. An acquisition processor is conditioned for acquiring data representing a sequence of successive pulses of a type of electrical heart waveform of a patient. A computation processor is electrically coupled to the acquisition processor and is conditioned for calculating a pulse interval irregularity measure based on a sum of time interval differences occurring between pairs of successive pulses of the sequence of successive pulses and excluding time interval differences exceeding a predetermined maximum threshold from the sum. The computation processor is further conditioned to compare a calculated pulse interval irregularity measure with a predetermined irregularity measure threshold. An output processor is electrically coupled to the computation processor and conditioned for generating data representing an alarm message in response to the calculated pulse interval irregularity measure exceeding the predetermined threshold. | 1. A cardiac condition data processing system for processing data representative of heart electrical signals to identify a cardiac condition, comprising: an acquisition processor configured to acquire data representing a sequence of successive pulses of a type of electrical heart waveform of a patient;a computation processor configured to, calculate a pulse interval irregularity measure based on a sum of time interval differences occurring between pairs of successive pulses of said sequence of successive pulses, the time interval differences exceeding a predetermined maximum threshold are excluded from said sum prior to use in calculating the pulse interval irregularity measure andcompare a calculated pulse interval irregularity measure with a predetermined irregularity measure threshold; andan output processor configured to generate data representing an alert message in response to said calculated pulse interval irregularity measure exceeding said predetermined threshold. 2. The system according to claim 1, wherein each of the time interval differences comprises a difference between a first time interval occurring between a first pair of pulses and a second time interval occurring between a subsequent second pair of pulses, said first and second pair of pulses sharing a common pulse. 3. The system according to claim 1, including a defibrillator, for one of enabling a user to initiate defibrillation in response to said alert message and automatically initiating defibrillation in response to said alert message. 4. The system according to claim 1, including a heart pacing signal generator for automatically initiating generation of a heart pacing signal in response to said alert message. 5. The system according to claim 1, including a patient parameter monitoring device for automatically changing monitoring device operational settings in response to said alert message. 6. The system according to claim 1, including a treatment device for automatically changing operational settings and treatment provided to a patient in response to said alert message. 7. The system according to claim 1, wherein said predetermined maximum threshold from said sum comprises at least 290 milliseconds. 8. The system according to claim 1, wherein said sequence of successive pulses of said type of electrical heart waveform of a patient comprises R-wave pulses. 9. The system according to claim 8, wherein each pulse of successive pulses corresponds to a normal heart beat. 10. The system according to claim 9, wherein said computation processor inhibits calculation if the number of pairs of consecutive normal heart beats in a first predetermined number of successive heart beats is less than a second predetermined number. 11. The system according to claim 1, wherein said computation processor calculates said pulse interval irregularity measure based on a sum of the absolute value of said time interval differences. 12. The system according to claim 11, wherein said comp... | Draeger Medical Systems, Inc. | A61B005/04 A61N001/00 A61B005/024 A61N001/362 A61B005/0464 A61N001/39
US08790273 | B2 | 2014-07-29 | Noninvasive method and system for measuring pulmonary ventilation | 2008-09-05 | A pulmonary ventilation system comprising a method to determine at least one pulmonary ventilation parameter as a function of a plurality of measured anatomical distances and volume-motion coefficients, sensors for acquiring the anatomical distances, a device for determining the plurality of motion coefficients, and a device for determining the ventilation parameter based on the acquired anatomical distances and determined plurality of volume-motion coefficients. In one embodiment, the system further includes a method for acquiring base-line ventilation characteristics and a method for correlating the base-line ventilation characteristics to the ventilation parameter determined with the empirical relationship. | 1. A method for determining the tidal volume (VT) of a subject, comprising: determining a first anatomical characteristic representing a first linear displacement of the subject's rib cage in a first orientation using a first sensor system;determining a second anatomical characteristic representing a first linear displacement of the subject's abdomen in said first orientation using a second sensor system;determining a third anatomical characteristic representing a first axial displacement of the subject's chest wall in said first orientation using said first sensor system and said second sensor system;determining a first rib cage volume-motion coefficient representing said first orientation of the subject;determining a corrected abdominal volume-motion coefficient based on an upright abdominal volume-motion coefficient and a linear slope of an abdominal volume-motion coefficient as a function of an axial displacement of the subject's chest wall;determining a first chest well volume-motion coefficient representing said first orientation of the subject;providing a first mathematical relationship that is adapted to determine VT of the subject as a function of said first, second and third anatomical characteristics, said first rib cage volume-motion coefficient, said corrected abdominal volume-motion coefficient and said first chest wall volume-motion coefficient; anddetermining VT of the subject with said first mathematical relationship and said first, second and third anatomical characteristics, and said first rib cage volume-motion coefficient, said corrected volume-motion coefficient and first chest wall volume-motion coefficient, whereby said determined VT represents the subject's VT at said first orientation. 2. The method of claim 1, wherein said first axial displacement of said subject's chest wall is determined from a xiphi-umbilical distance. 3. The method of claim 1, wherein said first axial displacement of said subject's chest wall is determined from a sternal-umbilical distance. 4. The method of claim 1, wherein said first anatomical characteristic represents a second linear displacement of said rib cage during a first motion of the subject, said second anatomical characteristic represents a second linear displacement of said abdomen during said first motion of the subject, said third anatomical characteristic represents a second axial displacement of said chest wall during said first motion of the subject, said first rib cage volume-motion coefficient represents said first motion of the subject, and said first chest wall volume-motion coefficient represents said first motion of the subject, and wherein said determined VT represents the subject's VT during said first motion. 5. The method of claim 1, wherein said first orientation comprises an upright position. 6. The method of claim 5, wherein said first mathematical relationship comprises the following equation VT=α(ΔRC)+(βu+εXi)×(ΔAb)+γ(ΔXi)wherein ΔRC represents the linear displ... | adidas | A61B005/08 A61B005/04 A61B005/02 A61B005/103 G01B007/14 G01R033/02 A61B005/091
US08628477 | B2 | 2014-01-14 | Systems and methods for non-invasive determination of blood pressure | 2009-07-31 | Methods and systems for determining blood pressure from a pressure signal are disclosed. A patient's blood pressure may be determined by analyzing features of a wavelet transformation of a pressure signal obtained during an occlusion procedure. Ridges in a scalogram of the transformed signal may be identified and used to determine an envelope of a pressure oscillation signal, to which oscillometric blood pressure determination techniques may be applied. | 1. A method for determining a blood pressure measurement, comprising: receiving, from a sensor, an electronic signal responsive to an occluding device applied to a patient;using processor equipment for: transforming the electronic signal into a transformed signal based at least in part on a continuous wavelet transformation,generating a scalogram based at least in part on the transformed signal, wherein the scalogram comprises 3-dimensional information,identifying, within the scalogram, a ridge indicative of the patient's blood flow as a function of time, anddetermining a blood pressure measurement based at least in part on the identified ridge; andoutputting the blood pressure measurement to an output device. 2. The method of claim 1, further comprising applying a variable occluding pressure to a patient using the occluding device. 3. The method of claim 1, wherein the sensor is a pressure transducer associated with the occluding device. 4. The method of claim 1, wherein the blood pressure measurement comprises at least one of a systolic blood pressure, a diastolic blood pressure and a mean arterial pressure. 5. The method of claim 1, further comprising using processor equipment for determining an oscillation envelope based at least in part on the identified ridge, wherein determining a blood pressure measurement is based at least in part on the oscillation envelope. 6. The method of claim 5, wherein the oscillation envelope is based at least in part on a projection of the identified ridge. 7. The method of claim 1, wherein determining a blood pressure measurement is further based at least in part on a maximum amplitude of the identified ridge. 8. A system for determining a blood pressure measurement, comprising: an occluding device capable of being applied to a patient; a sensor, coupled to the occluding device, that outputs an electronic signal responsive to the occluding device when the occluding device is applied to the patient;processor equipment, coupled to the sensor, the processor equipment being capable of: transforming the electronic signal into a transformed signal based at least in part on a continuous wavelet transformation,generating a scalogram based at least in part on the transformed signal, wherein the scalogram comprises 3-dimensional information,identifying, within the scalogram, a ridge indicative of the patient's blood flow as a function of time,determining a blood pressure measurement based at least in part on the identified ridge; andan output device, coupled to the processor equipment, for outputting the blood pressure measurement. 9. The system of claim 8, wherein the occluding device is capable of applying a variable occluding pressure to a patient. 10. The system of claim 8, wherein the sensor is a pressure transducer associated with the occluding device. 11. The system of claim 8, wherein the blood pressure measurement comprises at least one of a systolic blood pressure, a diastolic blood pressure and a mea... | Nellcor Puritan Bennett Ireland | A61B005/02
US08509874 | B2 | 2013-08-13 | Method for non-contrast enhanced pulmonary vein magnetic resonance imaging | 2010-07-29 | A system and method for non-contrast enhanced pulmonary vein magnetic resonance imaging substantially suppresses the signal from cardiac tissue adjacent to the left atrium and pulmonary vein is provided. Significant conspicuity of the left atrium and pulmonary vein versus adjacent anatomical structures is produced. In this manner, more accurate measurements of pulmonary vein ostia size are facilitated, as well as more accurate registration of imaging volumes with a radiofrequency ablation catheter during pulmonary vein isolation procedures. In addition, more robust three-dimensional volume views of the left atrium and pulmonary vein are produced without the administration of contrast agents. | 1. A method for producing a non-contrast enhanced image of a subject with a magnetic resonance imaging (MRI) system, the steps comprising: a) identifying a desired imaging slab extending parallel to an imaging plane including at least a portion of the pulmonary vein;b) performing a pulse sequence with the MRI system that includes: i) applying an inversion recovery radiofrequency (RF) pulse in an inversion slab oriented in an inversion plane that extends not parallel to the imaging plane and not including the pulmonary vein of the subject;ii) allowing a selected inversion time (TI) to pass following the application of the inversion recovery RF pulse;iii) following the TI, acquiring image data from the desired imaging slab; andc) reconstructing an image of the subject including at least the portion of the pulmonary vein from the acquired image data. 2. The method of claim 1 wherein the desired inversion slab extends in a substantially sagittal orientation through the subject. 3. The method of claim 1 wherein the desired inversion slab extends in a substantially coronal orientation through the subject. 4. The method as recited in claim 1 wherein the image slab also extends through at least a portion of the left atrium. 5. The method of claim 1 wherein the TI is selected to be approximately 350 to 500 ms in duration. 6. The method of claim 1 wherein the desired inversion slab has a thickness of approximately 60 mm. 7. The method of claim 1 wherein step b) further includes acquiring an electrocardiogram (ECG) signal from the subject and identifying an R-wave in the acquired ECG signal and step b)i) includes applying the desired inversion RF pulse after the occurrence of an R-wave in the acquired ECG signal. 8. The method of claim 1 wherein step b)iii) includes acquiring navigator image data in addition to the image data and step c) includes reconstructing a navigator image from the navigator image data. 9. The method of claim 8 wherein step c) includes using the navigator image to compensate for respiratory motion when reconstructing the image of the subject from the acquired image data. 10. The method of claim 1 wherein the inversion slab extends through at least a portion of an aorta, right atrium, right ventricle, left ventricle, and left atrium of the subject. 11. A method for producing an non-contrast enhanced image of a subject's pulmonary vein with a magnetic resonance imaging (MRI) system, the steps comprising: a) acquiring an electrocardiogram (ECG) signal from the subject to at least identify an R-wave in the acquired ECG signal;b) performing a pulse sequence with the MRI system that causes the MRI system to: i) after identifying an occurrence of an R-wave in the acquired ECG signal, apply an inversion recovery radiofrequency (RF) pulse in an inversion slab that is oriented in a plane not parallel to an imaging plane and that does not include the pulmonary vein;ii) acquire image data in an imaging slab parallel to the imaging plane... | Beth Israel Deaconess Medical Center | A61B005/055
US08346333 | B2 | 2013-01-01 | Systems and methods for estimating values of a continuous wavelet transform | 2009-07-30 | According to embodiments, techniques for estimating scalogram energy values in a wedge region of a scalogram are disclosed. A pulse oximetry system including a sensor or probe may be used to receive a photoplethysmograph (PPG) signal from a patient or subject. A scalogram, corresponding to the obtained PPG signal, may be determined. In an arrangement, energy values in the wedge region of the scalogram may be estimated by calculating a set of estimation locations in the wedge region and estimating scalogram energy values at each location. In an arrangement, scalogram energy values may be estimated based on an estimation scheme and by combining scalogram values in a vicinity region. In an arrangement, the vicinity region may include energy values in a resolved region of the scalogram and previously estimated energy values in the wedge region of the scalogram. In an arrangement, one or more signal parameters may be determined based on the resolved and estimated values of the scalogram. | 1. A method for processing a scalogram to estimate one or more energy values in a wedge region of the scalogram, the method comprising: receiving a signal; andusing processing equipment for: transforming the signal based at least in part on a continuous wavelet transform;generating the scalogram based at least in part on the transformed signal;estimating the one or more energy values in the wedge region of the scalogram based at least in part on an estimation scheme;determining one or more signal parameters of the signal based at least in part on the one or more estimated energy values; andoutputting the determined one or more signal parameters to an output device. 2. The method of claim 1, wherein estimating the one or more energy values comprises: determining a plurality of estimation points and a set of locations corresponding to the plurality of estimation points in the wedge region of the scalogram;determining a set of estimation parameters corresponding to the estimation scheme; anddetermining a size and shape of a vicinity region. 3. The method of claim 2, wherein the determined size and shape of the vicinity region is based, at least in part, on one or more characteristics of the one or more signal parameters. 4. The method of claim 2, wherein the determined size and shape of the vicinity region is based, at least in part, on one or more characteristics of a patient. 5. The method of claim 1, wherein the signal comprises a photoplethysmograph (PPG) signal received from a sensor. 6. The method of claim 1, wherein the one or more signal parameters comprises a patient oxygen saturation level. 7. The method of claim 1, wherein the signal parameter comprises a patient respiration rate. 8. The method of claim 1, wherein the method further comprises: determining one or more confidence metrics based at least in part on the one or more estimated energy values; anddetermining if the one or more confidence metrics satisfy a constraint set. 9. The method of claim 8, wherein determining the one or more confidence metrics comprises: determining energy values and energy structure in a pulse band of the scalogram; anddetermining energy values and energy structure in a noise band of the scalogram. 10. The method of claim 1, wherein the estimate of at least one energy value in the one or more estimated energy values is based, at least in part, on a set of resolved energy values in a resolved region of the scalogram and a set of estimated energy values in the wedge region of the scalogram. 11. A system for processing a scalogram to estimate one or more energy values in a wedge region of the scalogram, the system comprising: a sensor capable of generating a signal; andone or more processors coupled to the sensor, wherein the one or more processors are capable of: receiving the signal from the sensor;transforming the signal based at least in part on a continuous wavelet transform;generating the scalogram based at least in part on the transformed sig... | Nellcor Puritan Bennett Ireland | A61B005/1455 A61B005/00 G01N033/48
US20110009714 | A1 | 2011-01-13 | HEMODYNAMIC MONITORS AND ALARMS | 2009-02-02 | A hemodynamic monitoring instrument includes a processor and an output device. The processor (30) is arranged to receive a physiological parameter indicative of heart rate and a physiological parameter indicative of arterial blood pressure and is configured to compute (50) a hemodynamic parameter correlating with systemic vascular resistance (SVR) based on the received physiological parameter indicative of heart rate and the received physiological parameter indicative of arterial blood pressure. The output device includes least one of: (i) a display (24) configured to display the computed hemodynamic parameter; and (ii) an alarm (32, 34) configured to generate a perceptible signal responsive to the computed hemodynamic parameter satisfying an alarm criterion (52). The processor in some embodiments computes the hemodynamic parameter using a fuzzy membership function representing the heuristic ‘quantitative ABP measure is low AND quantitative HR measure is slightly high or high’ OR ‘quantitative ABP measure is very low’. | 1. A hemodynamic monitoring instrument comprising: a processor arranged to receive a physiological parameter indicative of heart rate and a physiological parameter indicative of arterial blood pressure and configured to compute a hemodynamic parameter correlating with systemic vascular resistance (SVR) based on the received physiological parameter indicative of heart rate and the received physiological parameter indicative of arterial blood pressure, the hemodynamic parameter correlating with SVR quantifying the heuristic “ABP is low AND HR is slightly high or high” OR “ABP is very low” where APB denotes an arterial blood pressure indicated by the physiological parameter indicative of arterial blood pressure and HR denotes a heart rate indicated by the physiological parameter indicative of heart rate; andan output device including least one of (i) a display configured to display the computed hemodynamic parameter and (ii) an alarm configured to generate a perceptible signal responsive to the computed hemodynamic parameter satisfying an alarm criterion. 2. The hemodynamic monitoring instrument as set forth in claim 1, wherein the instrument is integral with a patient monitor configured to display at least one of (i) a heart rate derived from the physiological parameter indicative of heart rate and (ii) an arterial blood pressure derived from the physiological parameter indicative of arterial blood pressure. 3. The hemodynamic monitoring instrument as set forth in claim 1, further comprising: a memory storing a patient age, the processor configured to compute the hemodynamic parameter substantially correlating with SVR further based on the stored patient age. 4. The hemodynamic monitoring instrument as set forth in claim 1, further comprising: a memory storing a patient gender, the processor configured to compute the hemodynamic parameter substantially correlating with SVR further based on the stored patient gender. 5. The hemodynamic monitoring instrument as set forth in claim 1, further comprising: a memory storing a patient age and a patient gender, the processor configured to compute the hemodynamic parameter substantially correlating with SVR further based on the stored patient age and stored patient gender. 6. The hemodynamic monitoring instrument as set forth in claim 1, wherein the processor is configured to compute the hemodynamic parameter substantially correlating with SVR based on: a fuzzy membership function indicative of whether a heart rate indicated by the physiological parameter indicative of heart rate is slightly high or high,a fuzzy membership function indicative of whether an arterial blood pressure indicated by the physiological parameter indicative of arterial blood pressure is low, anda fuzzy membership function indicative of whether the arterial blood pressure indicated by the physiological parameter indicative of arterial blood pressure is very low. 7. The hemodynamic monitoring instrument as set forth in claim 1, wh... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/02
US20110009709 | A1 | 2011-01-13 | DETECTION OF CONGESTION FROM MONITORING PATIENT RESPONSE TO A RECUMBENT POSITION | 2010-09-22 | Systems and methods obtain information that allows detection of whether a patient suffers from congestion by monitoring a response of the patient to a recumbent position of the patient's body. The patient may be monitored to determine a respiration pattern for a non-recumbent position such as standing and a respiration pattern for a recumbent position such as lying down. The two patterns may be compared, either by a processing device or a physician, to determine a difference in the two respiration patterns. Furthermore, the congestion may be inferred from detecting an amount of time that the patient spends in a recumbent position or from detecting the recumbent angle that the patient obtains, either of which is presumed to indicate whether a recumbent position presents discomfort to the patient due to the presence of congestion. | 1. A system comprising: a reclination sensor configured to produce a reclination signal indicative of a reclination of a patient;reclination circuitry configured to produce reclination data from the reclination signal;a processing device configured to communicate with at least the reclination circuitry to determine, from at least the reclination data, an amount of respiratory distress, and to determine, from the amount of respiratory distress, whether the patient is suffering from congestion. 2. The system of claim 1, wherein the processing device is configured to determine a recumbent angle of the patient from the reclination data, the processing device being configured to compare the recumbent angle with a recumbent angle threshold to determine the amount of respiratory distress. 3. The system of claim 2, wherein the processing device is configured to limit the reclination data to a period when the patient is expected to be resting. 4. The system of claim 2, wherein the processing device is configured to limit the reclination data to a period when the patient is known to be resting. 5. The system of claim 1, comprising a clock configured to produce a timing signal, wherein the processor is configured to: determine, from the reclination data and the timing signal, an amount of recumbent time for which the patient is in a recumbent position; andcompare the amount of recumbent time to a threshold value to determine the amount of respiratory distress. 6. The system of claim 5, wherein the threshold value includes a time threshold related to a congestive condition. 7. The system of claim 1, comprising: a respiration sensor that produces a respiration signal indicative of respiration patterns of the patient; andrespiration circuitry that produces respiration pattern data from the respiration signal, wherein the processing device is configured to: determine, from the reclination data and the respiration pattern data, recumbent respiration pattern data and non-recumbent respiration pattern data;compare the recumbent respiration pattern data and non-recumbent respiration pattern data; anddetermine, using a result of the comparison, the amount of respiratory distress. 8. The system of claim 7, wherein the respiration sensor is configured to produce a respiration signal in response to a change in trans-thoracic impedance. 9. The system of claim 1, wherein the reclination sensor includes a tilt switch. 10. The system of claim 1, wherein the reclination sensor includes an accelerometer. 11. A method comprising: sensing reclination of a patient;determining an amount of respiratory distress using at least the sensed reclination; anddetermining, from the amount of respiratory distress, whether the patient is suffering from congestion. 12. The method of claim 11, comprising: determining a recumbent angle of the patient from the sensed reclination; andcomparing the recumbent angle with a recumbent angle threshold to determine the amount of respiratory... | - | A61B005/00
US20110046471 | A1 | 2011-02-24 | Electrode Lead Set for Measuring Physiologic Information | 2010-11-02 | An electrode lead set is provided for electrical connection to a body. The electrode lead set includes a flexible ribbonized cable core extending between a proximal end portion and a distal end portion. The distal end portion includes a plurality of branch end portions that are each configured to hold an electrode. The flexible ribbonized cable core is separable into a plurality of branches that are each joined to adjacent branches by a separable interface prior to separation. Each of the plurality of branches includes a corresponding one of the branch end portions. Each of the plurality of branches includes a generally planar ribbon conductor extending along the corresponding branch from the proximal end portion of the ribbonized cable core to the corresponding branch end portion. | 1. An electrode lead set for electrical connection to a body, said electrode lead set comprising: a flexible ribbonized cable core extending between a proximal end portion and a distal end portion, the distal end portion comprising a plurality of branch end portions that are each configured to hold an electrode, the flexible ribbonized cable core being separable into a plurality of branches that are each joined to adjacent branches by a separable interface prior to separation, each of the plurality of branches comprising a corresponding one of the branch end portions, wherein each of the plurality of branches comprises a generally planar ribbon conductor extending along the corresponding branch from the proximal end portion of the ribbonized cable core to the corresponding branch end portion, each of the generally planar ribbon conductors being shielded. 2. The electrode lead set according to claim 1, wherein the branch end portions each comprise an electrode housing that is configured to hold the corresponding electrode. 3. The electrode lead set according to claim 2, wherein the electrode housings each comprise a first opening for receiving a distal end portion of the corresponding generally planar ribbon conductor and a second opening for receiving an electrical contact of the corresponding electrode, and the first and second openings are relatively arranged such that the distal end portion of the corresponding generally planar ribbon conductor is configured to electrically connect to the electrical contact of the corresponding electrode. 4. The electrode lead set according to claim 3, wherein each of the electrode housings comprises a cover configured to substantially cover the first opening. 5. The electrode lead set according to claim 4, wherein each of the covers comprises an extension configured to engage the distal end portion of the corresponding generally planar ribbon conductor to facilitate electrical connection between the corresponding generally planar ribbon conductor and the electrical contact of the corresponding electrode. 6. The electrode lead set according to claim 2, wherein the electrode housings each comprise a first opening for connecting with an electrical contact of the corresponding electrode with a snap-fit connection. 7. The electrode lead set according to claim 1, wherein each of the plurality of branches comprises the corresponding generally planar ribbon conductor, an insulating material at least partially surrounding the generally planar ribbon conductor, an electrically conductive material at least partially surrounding the insulating material, and an insulating jacket at least partially surrounding the electrically conductive material. 8. The electrode lead set according to claim 1, wherein prior to separation the plurality of branches are configured such that the branch end portions are maintained in an array at the distal end portion of the substrate core, and each branch end portion within the array i... | - | A61B005/0408
US07885448 | B2 | 2011-02-08 | Diagnostic scanning microscope for information-enriched qualitative histopathology | 2007-02-27 | A microscope array with staggered rows of magnifying imaging systems is used to scan a biological tissue sample in a single linear pass to produce an image and corresponding optical-density data. A conventional computerized algorithm is used to identify, isolate and produce segmented images of nuclei contained in the image. The OD values corresponding to nuclear chromatin are used to identify numerical patterns known to have statistical significance in relation to the health condition of the biological tissue. These patterns are analyzed to detect pre-neoplastic changes in histologically normal-appearing tissue that suggest a risk for the development of a pre-malignant and a potentially malignant lesion. This information is then converted to a visually perceptible form incorporated into the image of the tissue sample and is displayed for qualitative analysis by a pathologist. | 1. A method of imaging a biological sample with a microscopic imaging system, comprising the following steps: providing a two-dimensional microscope array with a plurality of magnifying imaging systems disposed along a corresponding plurality of optical axes for imaging a sample onto a detector, said plurality of magnifying imaging systems being arranged in rows and configured to image respective sections of the sample;providing a scanning mechanism for producing a relative movement between the microscope array and the sample;operating the scanning mechanism along a linear direction of scan across the sample to produce an image of the sample from a plurality of image-forming signals corresponding to said plurality of magnifying imaging systems;analyzing said plurality of image-forming signals to produce a measure of image-forming features in said image, wherein said measure is a statistically significant indicator of pathology in portions of said image; andassigning a visually detectable marker to each of said portions of the image corresponding to image-forming signals that produced said measure;wherein said rows of imaging systems are staggered with respect to said direction of scan, such that each of the magnifying imaging systems acquires image data corresponding to a respective continuous strip of the sample along said linear direction of scan of the scanning mechanism. 2. The method of claim 1, wherein said image-forming signal is optical density. 3. The method of claim 2, further including the step of combining said marker with the image to produce an information-enriched image. 4. The method of claim 3, wherein said marker is color. 5. The method of claim 2, wherein said marker is color. 6. The method of claim 2, wherein said measure is a statistically significant combination of said optical-density features. 7. The method of claim 6, wherein said marker is color. 8. The method of claim 2, wherein said portions of the image are cell nuclei. 9. The method of claim 2, wherein said microscopic imaging system comprises a plurality of individual miniaturized microscopes in an array microscope. 10. An information-enriched image produced by the method of claim 9. 11. An information-enriched image produced by the method of claim 2. 12. Apparatus for imaging a biological sample with a microscopic imaging system, comprising: a two-dimensional microscope array with a plurality of magnifying imaging systems disposed along a corresponding plurality of optical axes for imaging a sample onto a detector, said plurality of magnifying imaging systems being arranged in rows and configured to image respective sections of the sample;a scanning mechanism for producing a relative movement between the microscope array and the sample along a linear direction of scan, so as to produce an image of the sample from a plurality of image-forming signals corresponding to said plurality of magnifying imaging systems;means for analyzing said plurality of image-... | DMetrix, Inc. | G06K009/00 G21K007/00 A61B005/05
US07993167 | B2 | 2011-08-09 | Shielded electrode connector | 2009-01-28 | An ECG lead set is described which is shielded against electrostatic charge hazards. An electrical shield is located at the end of each lead of the lead set and electrically shields the connection of the lead set to an ECG electrode. The electrical shield is covered by a nonconductive cover and is electrically connected to the shield of the coaxial cable of the lead set. | 1. A shielded connector for a body electrode comprising: a lead having a shielded signal conductor;a connector electrically attached to the signal conductor for attachment to a body electrode having a patient-contacting area which conducts received signals to the signal conductor; anda conductive electrical shield located above the connector and above the patient-contacting area when the connector is attached to the body electrode, the electrical shield being covered by a nonconductive cover and electrically connected to the shielding of the signal conductor, wherein the conductive electrical shield comprises a unitary disk-shaped shield having an outer periphery extending to a position between the periphery of the patient-contacting area and the outer periphery of the body electrode when the connector is attached to the body electrode. 2. The shielded connector of claim 1, wherein the shielded signal conductor further comprises a coaxial cable having a central signal conductor and an outer shielding conductor, wherein the conductive electrical shield is electrically connected to the outer shielding conductor. 3. The shielded connector of claim 2, wherein the connector is a snap connector. 4. The shielded connector of claim 2, wherein the connector is a clip-on connector. 5. The shielded connector of claim 2, wherein the body electrode further comprises an ECG electrode. 6. The shielded connector of claim 1, wherein the nonconductive cover presents a dielectric strength to an external conductor of at least 2000 volts/mm DC. 7. The shielded connector of claim 1, wherein the nonconductive cover presents a dielectric strength to an external conductor of at least 5000 volts/mm DC. 8. The shielded connector of claim 1, wherein the covered electrical shield extends outward and downward from the connector when the connector is attached to the body electrode. 9. The shielded connector of claim 1, wherein the conductive electrical shield comprises one of a conductive paint, a foil, or a conductive cloth. 10. The shielded connector of claim 9, wherein the conductive electrical shield exhibits a conductivity of 1.0 Ohm-cm or less. | Koninklijke Philips Electronics N.V. | H01R004/48
US07869850 | B2 | 2011-01-11 | Medical sensor for reducing motion artifacts and technique for using the same | 2005-09-29 | A sensor for pulse oximetry or other applications utilizing spectrophotometry may be adapted to reduce motion artifacts by fixing the optical distance between an emitter and detector. A flexible sensor is provided with a stiffening member to hold the emitter and detector of the sensor in a relatively fixed position when applied to a patient. Further, an annular or partially annular sensor is adapted to hold an emitter and detector of the sensor in a relatively fixed position when applied to a patient. A clip-style sensor is provided with a spacer that controls the distance between the emitter and detector. | 1. A sensor comprising: an emitter and a detector disposed on a conformable sensor body;a first cavity disposed on or in the sensor body comprising a first material; anda second cavity disposed on or in the sensor body comprising a second material, wherein the first material hardens upon exposure to the second material. 2. The sensor, as set forth in claim 1, comprising a breakable barrier between the first cavity and the second cavity. 3. The sensor, as set forth in claim 1, wherein the first material comprises cement or plaster and the second material comprises water. 4. The sensor, as set forth in claim 1, wherein the first material comprises one part of a two-part epoxy. | Nellcor Puritan Bennett LLC | A61B005/1455
US20110015503 | A1 | 2011-01-20 | MEDICAL APPARATUS FOR COLLECTING PATIENT ELECTROENCEPHALOGRAM (EEG) DATA | 2010-09-24 | The EEG Processing Unit comprises a semi-rigid framework which substantially conforms to the Patient's head and supports a set of electrodes in predetermined loci on the Patient's head to ensure proper electrode placement. The EEG Processing Unit includes automated connectivity determination apparatus which can use pressure-sensitive electrode placement ensuring proper contact with Patient's scalp and also automatically verifies electrode placement via measurements of electrode impedance through automated impedance checking. Voltages generated by the electrodes are amplified and filtered before being transmitted to an analysis platform, which can be a Physician's laptop computer system, either wirelessly or via a set of tethering wires. The EEG Processing Unit includes an automatic artifacting capability which identifies when there is sufficient clean data compiled in the testing session. This process automatically eliminates muscle- or other physical-artifact-related voltages. Clean data, which represents real brain voltages as opposed to muscle- or physical-artifact-related voltages, thereby are produced. | 1. An EEG Processing Unit for collecting EEG sensor data from a subject for forwarding to a data processing system, comprising: a plurality of EEG sensors, responsive to the presence of EEG signals for generating electrical signals representative of said EEG signals;frame having said plurality of EEG sensors mounted thereon in predetermined locations, for placement on a Patient's head thereby to site each of said plurality of EEG sensors at a corresponding predetermined location on a surface of the Patient's head to detect said EEG signals; andsignal processor, mounted on said frame, for processing said electrical signals representative of said EEG signals to remove artifacts therefrom. 2. The EEG Processing Unit of claim 1, further comprising: impedance measurement for automatically performing measurements indicative of contact of said EEG sensors to said head of said Patient. 3. The EEG Processing Unit of claim 1, further comprising: a test signal generator for executing a test to ascertain quality of signals being produced by said EEG Processing Unit. 4. The EEG Processing Unit of claim 1 wherein said signal processor comprises: artifacting processor for processing said electrical signals produced by said EEG sensors to eliminate artifacts from said electrical signals to produce processed electrical signals. 5. The EEG Processing Unit of claim 4 wherein said signal processor further comprises: data validation process for determining that said processed electrical signals are of sufficient quality to be output. 6. The EEG Processing Unit of claim 4, further comprising: an accelerometer, attached to said frame and responsive to movement of said frame in three dimensions, for generating a motion signal representative of a direction and magnitude of said three-dimensional motion. 7. The EEG Processing Unit of claim 6 wherein said signal processor further comprises: motion artifacting process, responsive to said motion signal, for processing said electrical signals to one of: remove motion artifacts from said electrical signals, and discard said electrical signals in the presence of movement of said frame. 8. The EEG Processing Unit of claim 1 wherein said frame comprises: a semi rigid framework formed to substantially conform to the head of the Patient. 9. The EEG Processing Unit of claim 8, further comprising: expansion mechanism connected to said frame for applying pressure to said plurality of EEG sensors to force said EEG sensors against the exterior surface of said Patient's head. 10. The EEG Processing Unit of claim 9, further comprising: a contact enhancer mechanism for providing movement and/or rotation of said EEG sensors to enhance contact between said EEG sensors and the scalp of said Patient. 11. A method of collecting EEG sensor data from a Patient for forwarding to a data processing system, comprising: generating, via a plurality of EEG sensors which are responsive to the presence of EEG signals, electrical signals represen... | WAVi | A61B005/0476 A61B005/11
US20110009724 | A1 | 2011-01-13 | PROVIDING CONTEXTUALLY RELEVANT ADVERTISEMENTS AND E-COMMERCE FEATURES IN A PERSONAL MEDICAL DEVICE SYSTEM | 2009-07-09 | A medical device system includes a portable (patient-worn or patient-carried) medical device that can automatically initiate certain e-commerce and/or advertisement presentation features in response to the detection of one or more triggering events. A triggering even may, for example, be associated with the monitoring of the operation or status of the medical device, the operation or status of the medical device system, a medical condition of the patient, treatment of the patient, and/or the preferences of the patient or user. In one application, the medical device is an insulin infusion pump and the user is automatically prompted to purchase, order, or otherwise obtain replaceable or replenishable items associated with the operation of the insulin infusion pump, such as a continuous glucose sensor, an infusion set, or an insulin reservoir. | 1. A method of controlling operation of a portable medical device having one or more consumables associated therewith, the method comprising: receiving status data corresponding to operating status of the portable medical device;determining, based upon analysis of the status data, when a consumable associated with the portable medical device requires replacement or replenishment;in response to the determining step, generating a message that indicates a need to replace or replenish the consumable; andcommunicating the message to the portable medical device. 2. The method of claim 1, wherein the message comprises an active link to a vendor of the consumable, the active link facilitating initiation of an electronic transaction for the consumable at the portable medical device. 3. The method of claim 1, wherein the message comprises an HTML document that facilitates initiation of an electronic transaction for the consumable at the portable medical device. 4. The method of claim 1, wherein the message comprises an email. 5. The method of claim 1, wherein the message comprises a voicemail. 6. The method of claim 1, wherein: the portable medical device comprises a medication delivery device that delivers a medication to a patient; andthe consumable comprises the medication. 7. The method of claim 1, wherein: the portable medical device comprises an infusion pump that cooperates with an infusion set for a patient; andthe consumable comprises a replacement infusion set for the infusion pump. 8. The method of claim 1, wherein: the portable medical device comprises an insulin infusion pump that responds to glucose measurements obtained from a glucose sensor worn by a patient; andthe consumable comprises a replacement glucose sensor. 9. The method of claim 1, wherein: the portable medical device comprises a controller configured to control delivery of a medication to a patient via a medication delivery device that is worn or carried by the patient; andthe consumable comprises the medication. 10. The method of claim 1, wherein: the portable medical device comprises a controller configured to control delivery of a medication to a patient via an infusion pump that cooperates with an infusion set worn by the patient; andthe consumable comprises a replacement infusion set for the infusion pump. 11. The method of claim 1, wherein: the portable medical device comprises a controller configured to control delivery of insulin to a patient via an insulin infusion pump that responds to glucose measurements obtained from a glucose sensor worn by the patient; andthe consumable comprises a replacement glucose sensor. 12. The method of claim 1, further comprising: obtaining geographic position data that indicates a location of the portable medical device; andidentifying, based upon the geographical position data, a provider of the consumable, the provider maintaining inventory of the consumable at a facility located within a predetermined distance from the locati... | MEDTRONIC MINIMED, INC. | A61M005/168 A61B005/15 A61M005/142 G06Q010/00 G06Q030/00
US20110015502 | A1 | 2011-01-20 | Analyte Monitoring and Management System and Methods Therefor | 2010-09-23 | Method and apparatus for providing multiple data receiver units in a data monitoring and management system such as analyte monitoring system where a first data receiver includes all of the functionalities for the data monitoring and management system receiver unit, and a second data receiver unit is configured with a limited functions to provide application specific convenience to the user or patient is disclosed. | 1. (canceled) 2. A method for processing sensor data from a continuous analyte sensor, the method comprising: receiving at a sensor electronics module sensor data from each of one or more sensors coupled to a host, wherein the sensor electronics module comprises a data structure storing delivery options each indicative of at least sensor data content respectively associated with display device attributes, alert condition attributes, and/or status attributes;determining at the sensor electronics module that at least some of the sensor data matches at least a first alert condition, wherein the first alert condition indicates at least one threshold level of an attribute of the sensor data;determining a first display device associated with the at least the first alert condition;selecting at least one delivery option in the data structure associated with one or more of a type of the first display device, one or more attributes of the first alert condition, and/or a status of the host;selecting data associated with the sensor data for transmission to the first display device, wherein the data selected comprises at least the sensor data content indicated in the selected delivery options;generating at the sensor electronics module a data package for transmission to the first display device, wherein the data package includes the selected data; andinitiating transmission of the data package to the first display device. 3. The method of claim 2, wherein the at least the first alert condition comprises a second alert condition that indicates at least one threshold level of an attribute of second sensor data from a second sensor. 4. The method of claim 2, wherein the second sensor comprises one or more of a temperature sensor. 5. The method of claim 2, wherein the status of the host is selected from the group comprising one or more of: resting, exercise, illness, mealtime, day, night, hyperglycemia, hypoglycemia, clinical risk, and noise. 6. The method of claim 2, wherein the data associated with the sensor data comprises one or more of filtered sensor data, calibrated sensor data, transformed sensor data, status data, event data, trend data, rate of change data, rate of acceleration data, sensor diagnostic data, alert condition data, graphical data, textual data, audible data, and/or tactile data. 7. The method of claim 2, wherein the sensor data content indicates portions of one or more of filtered sensor data, calibrated sensor data, transformed sensor data, trend data, status data, event data, rate of change data, rate of acceleration data, sensor diagnostic data, alert condition data, graphical data, textual data, audible data, and/or tactile data. 8. The method of claim 2, further comprising: determining a second display device associated with the at least the first alert condition;selecting at least one second delivery option in the data structure associated with one or more of a type of the second display device, one or more attributes of the f... | Abbott Diabetes Care Inc. | A61B005/145 A61B005/00 A61B005/01 A61M037/00
US07878986 | B2 | 2011-02-01 | Catheter deployment device | 2007-05-03 | An apparatus for deploying a needle within a lumen is provided. The apparatus includes a housing having a threaded bushing radially disposed therein. The bushing rigidly couples with a nose cone having a guide tip disposed at an end opposite the bushing for penetrating an arterial wall of a lumen. During operation, a user incrementally advances the bushing within the housing, thereby incrementally advancing the guide tip into the lumen. The nose cone also includes a flex guide having a slot configuration which couples with the guide tip which deploys into the lumen along with the guide tip. | 1. A device for advancing a guide tip through a lumen, the device comprising: an elongate catheter having an anchor operatively associated with the catheter at a distal portion of the catheter;a housing coupled with a proximal end of the catheter;a bushing disposed within the housing, the bushing rotationally associated with the housing such that rotational movement of the housing relative to the bushing incrementally advances the bushing within the housing; andan inner key coupled with the bushing at a proximal end of the inner key, the guide tip being coupled with the inner key such that as the bushing incrementally advances within the housing, the guide tip incrementally advances within the lumen. 2. The device as recited in claim 1, wherein the housing further includes a groove radially defined within an inner wall of the housing and the bushing further includes a bushing thread complementary to the groove. 3. The device as recited in claim 2, wherein the inner key further includes a flange, the bushing receiving the flange of the inner key within a cavity of the bushing. 4. The device as recited in claim 1, further comprising a nose cone disposed distally from the housing, the nose cone receiving the guide tip configured to penetrate tissue surrounding the lumen. 5. The device as recited in claim 4, the device further comprising: a flex guide coupled with the guide tip. 6. The device as recited in claim 1, wherein the housing further comprises a valve adapter integrating the housing and the inner key, the valve adapter providing a rotational interface between the inner key and the valve adapter. 7. The device as recited in claim 4, wherein the guide tip further includes a cam surface, a curved portion, and a tip for penetrating tissue surrounding the lumen. 8. The device as recited in claim 7, the device further comprising: the nose cone having a plurality of cam surfaces which engage with the guide tip cam surface, engagement of the plurality of cam surfaces and the guide tip cam surface pivoting the guide tip to deploy the guide tip from within the nose cone. 9. The device as recited in claim 8, wherein one of the plurality of cam surfaces is formed at an angle in a range from about 20 degrees to about 50degrees. 10. A device for advancing a guide tip through a lumen, the device comprising: a rotationally movable housing;a bushing disposed within the housing and being incrementally movable longitudinally within the housing upon rotational movement of the housing;an inner key coupled with the bushing, the inner key being advanced as the bushing advances; andan anchor disposed distally to a distal end of the housing, the anchor being deployable to anchor the device relative to the lumen during deployment of the guide tip. 11. The device as recited in claim 10, wherein the anchor comprises a balloon. 12. The device as recited in claim 10, wherein the anchor comprises a material capable of bowing. 13. The device as recited in claim... | Abbott Laboratories | A61B005/00 A61M025/00
US07883473 | B2 | 2011-02-08 | Blood sampling device | 2003-02-20 | A device (302) for sampling and/or analyzing blood or other body fluid of a subject. A housing (306) contains a plurality of lancets (308) and optionally includes test elements (324) to take up a sample of blood, an evaluation system and a display. A complete system that can be handled as a single device, for example in the form of a wristwatch, includes a multiplicity of test elements (324) and lancets (308), which can be brought successively to a working position to perform multiple measurements. A cassette or carrier (320) includes multiple lancets (308) and/or test elements (324), for insertion into the device. | 1. A lancing device comprising: a housing;a cassette removably mounted within said housing, said cassette comprising at least one lancet having a lancet body and a protective cap;a piston for propelling the lancet along a path of travel, said piston having a lancet coupling portion for releasably engaging and retracting the lancet to separate the lancet body and the protective cap along a retraction portion of the path of travel of the lancet, and thereafter advance the lancet along the path of travel into a lancing position;at least one guide member per lancet and in association therewith, the guide member being aligned transversely to the path of travel of the associated lancet within the cassette for engaging the separated protective cap of the associated lancet to guide the protective cap out of the path of travel; anda retainer position within the cassette adjacent each said at least one guide member to retain the separated protective cap. 2. The lancing device of claim 1, further comprising a biasing element for moving the protective cap out of the path of travel of the lancet after separation of the protective cap from the lancet body. 3. The lancing device of claim 2, wherein the biasing element moves the cap transversely out of the lancet travel path within a plane that is perpendicular to the lancet travel path. 4. The lancing device of claim 1, wherein the protective cap comprises a pair of recesses in opposed sides thereof, and wherein the at least one guide member comprises a pair of guide posts slidably engaged in said recesses to guide said protective cap out of the path of travel of the lancet. 5. The lancing device of claim 1, further comprising a cocking mechanism for arming the piston. 6. The lancing device of claim 5, wherein the cassette comprises a plurality of the lancets and the lancing device further comprises an advancing mechanism for sequentially advancing the lancets into an operating position, and wherein the cocking mechanism and the advancing mechanism are coupled together so that they operate together. 7. The lancing device of claim 6, wherein actuation of the cocking mechanism actuates the advancing mechanism to advance the cassette to sequentially engage each of said plurality of lancets with said piston. 8. The lancing device of claim 5, wherein actuation of the cocking mechanism retracts said piston to cause separation of the lancet body and the protective cap. 9. The lancing device of claim 1, wherein said cassette comprises a plurality of lancets radially arranged about an axis. 10. The lancing device of claim 9, wherein said cassette further comprises radially extending guides defining the path of travel of the lancet. 11. The lancing device of claim 9, wherein said plurality of lancets lie in a plane, and wherein the piston propels the lancet along a path of travel within that plane. 12. The lancing device of claim 9, further comprising a cocking mechanism for sequentially advancing the cassette... | Facet Technologies, LLC | A61B005/00
US08897869 | B2 | 2014-11-25 | Plaque type determination devices, systems, and methods | 2010-10-18 | Plaque type determination devices, systems, and methods. In at least one exemplary embodiment of a device for assessing composition of a plaque of the present disclosure, the device comprises an elongate body having a longitudinal axis and a distal end, a first excitation electrode and a second excitation electrode located along the longitudinal axis of the body near the distal end of the elongate body, and a first detection electrode and a second detection electrode along the longitudinal axis of the body and in between the first and second excitation electrodes, wherein when a current source in communication with at least one of the first excitation electrode and the second excitation electrode applies current thereto to facilitate measurement of two or more conductance values within a vessel containing at least part of the elongate body at or near a plaque site, a plaque type determination can be made based upon a calculation of tissue conductance using at least one of the two or more conductance values. | 1. A device for assessing composition of a plaque, the device comprising: an elongate body having a longitudinal axis and a distal end;a first excitation electrode and a second excitation electrode located along the longitudinal axis of the body; anda first detection electrode and a second detection electrode along the longitudinal axis of the body and in between the first and second excitation electrodes;wherein when a current source in communication with at least one of the first excitation electrode and the second excitation electrode applies current thereto to facilitate measurement of two or more conductance values within a vessel containing at least part of the elongate body at a plaque site, a plaque type determination from at least two types of plaques to distinguish one plaque type from another can be made based upon a calculation of tissue conductance using at least one of the two or more conductance values obtained by the device configured to obtain said values. 2. The device of claim 1, wherein the elongate body comprises a wire. 3. The device of claim 1, wherein the measurement of two or more conductance values is/are made using at least one of the first detection electrode and a second detection electrode. 4. The device of claim 1, wherein the plaque type determination identifies a plaque at the plaque site as being at least partially fatty if the value of % Gp as determined by%⁢⁢Gp=Gp[GA+GB2]×100is less than a threshold percentage, wherein A is a first concentration of saline and B is a second concentration of saline different from the first concentration of saline. 5. The device of claim 4, wherein A equals 0.5% saline, and wherein B equals 1.5% saline. 6. The device of claim 1, wherein a first distance between the first excitation electrode and the first detection electrode is equal to a second distance between the second excitation electrode and the second detection electrode. 7. The device of claim 1, wherein the measurement of two or more conductance values is/are made upon infusion of a first bolus into the vessel, wherein the first bolus contacts at least one of the first detection electrode and a second detection electrode. 8. The device of claim 1, wherein the measurement of two or more conductance values is/are made upon infusion of two boluses of differing conductive concentrations into the vessel, wherein the two boluses contact at least one of the first detection electrode and a second detection electrode. 9. The device of claim 7, wherein the first bolus comprises a first concentration of saline. 10. The device of claim 1, further comprising: a tube having a lumen, the tube surrounding at least part of the elongate body. 11. The device of claim 10, wherein the measurement of two or more conductance values is/are made upon infusion of a bolus through the tube into the vessel, wherein the bolus contacts at least one of the first detection electrode and a second detection electrode. 12. The device of claim 10,... | 3DT Holdings, LLC | A61B005/05 A61B005/053 A61B005/02 A61B017/00 A61B005/107
US20110009773 | A1 | 2011-01-13 | IMPLANTABLE SENSING MODULES AND METHODS OF USING | 2010-08-13 | Implantable sensing modules and methods for monitoring various physical parameters, including physical parameters of a living body and environmental parameters to which the living body may be subjected, for example, impacts. A method for monitoring impacts to which a living body is subjected entails the use of an implantable sensing module that has a rigid housing containing at least one energy storage device and at least one electromechanical sensing element that is responsive to impacts. The module generates data corresponding to impacts to which the electromechanical sensing element is subjected, and records the data in memory. The module is preferably implanted in a living body so that the module is connected to a rigid portion of the living body, in particular, a bone or tooth. | 1. A method of monitoring impacts to which a living body is subjected, the method comprising: providing an implantable sensing module that comprises a rigid housing containing at least one energy storage device, at least one electromechanical sensing element that is responsive to impacts, means for generating outputs corresponding to impacts to which the electromechanical sensing element is subjected, and means for recording data corresponding to the outputs;implanting the module in a living body so that the module is located internally within the living body and is connected to a rigid portion of the living body chosen from the group consisting of bones and teeth;monitoring impacts to the living body by monitoring a level of at least one impact to which the electromechanical sensing element is subjected within the living body;producing an output corresponding to the level of the at least one impact sensed by the electromechanical sensing element;storing data in the recording means within the module corresponding to the output of the electromechanical sensing element; and thenwirelessly retrieving the data stored in the recording means while the module remains implanted in the living body. 2. The method according to claim 1, wherein the recording means records the data without being supplied power external of the module. 3. The method according to claim 1, wherein the recording means comprises nonvolatile digital memory devices. 4. The method according to claim 1, wherein the energy storage device comprises a battery. 5. The method according to claim 1, wherein the energy storage device comprises an electrical capacitive element. 6. The method according to claim 1, wherein the at least one electromechanical sensing element comprises an accelerometer. 7. The method according to claim 1, wherein the at least one electromechanical sensing element comprises a plurality of integrated electromechanical switches, the electromechanical switches define open electrical paths when not subjected to an impact and are operable to define closed electrical paths that produce the outputs, the electromechanical switches define the closed electrical paths in response to impacts and produce the outputs while the electromechanical switches are subject to impacts that exceed thresholds of the electromechanical switches, and the electromechanical switches define the open electrical paths and do not produce the outputs when no longer subject to impacts that exceed the thresholds of the electromechanical switches. 8. The method according to claim 7, wherein the electromechanical switches have different thresholds so as to have different levels of sensitivity to impacts. 9. The method according to claim 8, wherein the outputs produced by the electromechanical switches are used to calculate an amplitude of an impact determined from the different thresholds of the electromechanical switches. 10. The method according to claim 8, wherein the outputs produced by the e... | EVIGIA SYSTEMS, INC. | A61B005/00
US20110034783 | A1 | 2011-02-10 | SYSTEMS AND METHODS FOR BALANCING POWER CONSUMPTION AND UTILITY OF WIRELESS MEDICAL SENSORS | 2009-08-10 | Systems, methods, and devices for balancing power consumption and utility of medical sensors are provided. For example, a wireless medical sensor device may include a sensor, data processing circuitry, and wireless transmission circuitry. The sensor may be capable of obtaining a raw measurement from a patient, and the data processing circuitry may be capable of sampling the raw measurement to obtain values. Further, the data processing circuitry also may be capable of determining an update interval based at least in part on an update factor associated with a status of the patient, and the wireless transmission circuitry may be capable of wirelessly transmitting one of the values to an external wireless receiver at the update interval. | 1. A medical device comprising: a medical sensor capable of obtaining a raw measurement from a patient;data processing circuitry capable of sampling the raw measurement to obtain a plurality of values and capable of determining an update interval based at least in part on an update factor associated with a status of the patient; andwireless transmission circuitry capable of wirelessly transmitting one of the plurality of values to an external wireless receiver at the update interval. 2. The device of claim 1, wherein the medical sensor comprises a photoplethysmographic sensor; a respirator band; a temperature sensor; a blood pressure sensor; an ECG sensor; or a pulse transit time sensor; or any combination thereof. 3. The device of claim 1, wherein the data processing circuitry is capable of determining the update interval based at least in part on the update factor associated with the status of the patient, wherein the update factor comprises a historical stability of the plurality of values; an absolute value of one of the plurality of values; a historical stability of a plurality of extraneous sensor values obtained from an extraneous medical sensor; an absolute value of an extraneous sensor value obtained from the extraneous medical sensor; an instruction from an external device; a button press or switch setting on the medical device; a current location of the patient; a current location of a clinician; a movement of the patient; an initialization status of the medical device; or a remaining battery life of the medical device; or any combination thereof. 4. The device of claim 1, comprising a button, wherein the wireless transmission circuitry is capable of transmitting the raw measurement to the external wireless receiver when the button is pressed. 5. The device of claim 1, comprising another medical sensor capable of obtaining another raw measurement of the patient, wherein the data processing circuitry is capable of sampling the other raw measurement to obtain a plurality of other values, wherein the data processing circuitry is capable of determining the update interval based at least in part on the update factor, and wherein the update factor comprises a historical stability of the plurality of other values or an absolute value of one of the plurality of extraneous sensor values. 6. The device of claim 1, wherein the wireless transmission circuitry is capable of transmitting the raw measurement to the external wireless receiver for a period of time when the determined update interval is beneath a threshold. 7. The device of claim 1, wherein the wireless transmission circuitry is capable transmitting a portion of the plurality of values to the external wireless receiver at the update interval. 8. The device of claim 1, comprising a memory device capable of storing at least a portion of the plurality of values, wherein the portion of the plurality of values have been obtained at an approximately constant sampling rate. 9. A method... | Nellcor Puritan Bennett LLC | A61B005/00
US07890153 | B2 | 2011-02-15 | System and method for mitigating interference in pulse oximetry | 2006-09-28 | A pulse oximetry sensor adapted to emit light from an emitter proximate to a patient's tissue and detect a portion of the emitted light on a detector proximate to the tissue. The pulse oximetry system is adapted to acquire position data for the emitter and for the detector with one or more position indicators. The pulse oximetry system is adapted to process the position data to obtain a change in position of the emitter and the detector; and to process pulse oximetry measurements using the change in position to obtain a motion-corrected pulse oximetry data. | 1. A pulse oximetry system comprising; a sensor body having, an emitter, a detector, and one or more position indicators;one or more position receivers separate and external from the sensor body and configured to determine the respective positions of the position indicators; anda processor-based monitor configured to determine changes in absolute position in space of the sensor body based on the positions determined by the one or more position receivers and to process pulse oximetry measurements from the detector using the changes in absolute position to obtain a motion-corrected pulse oximetry measurement. 2. The pulse oximetry system, as set forth in claim 1, wherein the processor-based monitor utilizes triangulation to determine the change in the absolute position of the sensor body. 3. The pulse oximetry system, as set forth in claim 1, wherein the processor-based pulse oximetry system comprises a display configured to display corrected pulse oximetry data. 4. The pulse oximetry system, as set forth in claim 1, wherein the absolute position comprises coordinates denoting the position in space of the sensor body. 5. A pulse oximetry system comprising; a sensor body comprising;an emitter;a detector;one or more position indicators adapted to detect a change in position of at least one of the sensor body, the emitter, or the detector;anda processor-based monitor configured to process pulse oximetry measurements from the detector using the change in position to obtain a corrected pulse oximetry measurement, wherein the processor-based monitor utilizes a data bank of light scatter/absorbance probability as a function of an optical path length to obtain the corrected pulse oximetry measurement. 6. The pulse oximetry system, as set forth in claim 5, wherein the change in position comprises a change in position of the emitter relative to the detector. 7. The pulse oximetry system, as set forth in claim 5, wherein the change in position comprises an absolute position change of the emitter or the detector. 8. The pulse oximetry system, as set forth in claim 5, wherein absolute position of the emitter or the detector is measured using a Hall effect device. 9. A tangible machine readable medium comprising: code for emitting light from an emitter proximate to a patient's tissue;code for detecting a portion of the light on a detector proximate to the tissue;code for acquiring three-dimensional data for the emitter or for the detector with one or more position indicators;code for processing the three-dimensional data to determine if there has been a change in position of the emitter or the detector;code for processing pulse oximetry measurements from the detector; andcode for correcting the pulse oximetry measurements based at least in part on the change in position of the emitter or detector utilizing a data bank of light scatter/absorbance probability as a function of the optical path length. 10. The tangible machine readable medium, as set forth ... | Nellcor Puritan Bennett LLC | A61B005/1455
US07894885 | B2 | 2011-02-22 | Coherent signal rejection in ECG | 2007-05-02 | A method for monitoring an electrocardiogram (ECG) signal of a subject, includes digitally sampling an average signal from at least a first ECG electrode, determining an average interference frequency, and digitally sampling and buffering a raw ECG signal from at least a second ECG electrode. The method further includes: filtering the raw ECG signal to generate a residual signal; calculating, based on the residual signal, a first amplitude and a first phase shift of a primary interference signal at the average interference frequency and a second amplitude and a second phase shift of one or more harmonic interference signals at respective multiples of the average interference frequency; and digitally subtracting the primary interference signal and the one or more harmonic interference signals from the raw ECG signal so as to generate and output a clean ECG signal. | 1. A method for monitoring an electrocardiogram (ECG) signal of a subject, comprising: digitally sampling an average signal from ECG electrodes attached to the subject, the average signal comprising a Wilson Central Terminal (WCT) signal;processing the WCT signal with a band-pass filter to establish a band-pass filter signal, the band-pass filter signal having a line interference component;processing the band-pass filter signal with a frequency meter and determining a frequency of the line interference component;determining an average interference frequency of the WCT signal based on the frequency of the line interference component;digitally sampling and buffering a raw ECG signal from another ECG electrode attached to the subject;calculating a moving average of the raw ECG signal;calculating a residual signal using the moving average of the raw ECG signal wherein the residual signal comprises frequency components at and above the average interference frequency;calculating, based on the residual signal, a first amplitude and a first phase shift of a primary interference signal at the average interference frequency and a second amplitude and a second phase shift of one or more harmonic interference signals at respective multiples of the average interference frequency; anddigitally subtracting the primary interference signal and the one or more harmonic interference signals from the raw ECG signal so as to generate and output a clean ECG signal. 2. The method according to claim 1, and comprising using an artifact processor for imposing a limit on a rate of change in the first amplitude over time prior to subtracting the primary interference signal from the raw ECG signal. 3. The method according to claim 2, wherein the limit on the rate of change in the first amplitude over time is a first limit and further comprising imposing a second limit on a rate of change in the second amplitude over time prior to subtracting the one or more harmonic interference signals from the raw ECG signal. 4. The method according to claim 1, wherein filtering the raw ECG signal comprises applying to the raw ECG signal a moving average filter, and wherein generating the residual signal comprises subtracting an output of the moving average filter from the raw ECG signal. 5. The method according to claim 1, wherein calculating the first amplitude and the first phase shift of the primary interference signal comprises finding a correlation between the residual signal and a periodic function having a frequency equal to the average interference frequency. 6. The method according to claim 1, wherein calculating the second amplitude and the second phase shift of the one or more harmonic interference signals comprises finding a correlation between the residual signal and a periodic function having a frequency equal to a multiple of the average interference frequency. | Biosense Webster, Inc. | A61B005/0402
US08802055 | B2 | 2014-08-12 | Use of the spin crossover complex [m | 2008-10-16 | A method for magnetic resonance imaging of a human or animal body includes providing a spin crossover complex having the general formula [Mx+(Ly)Hw]AN)z wherein, MX+ is a positively charged metal ion,Ly is a heteroaromatic ligand containing at least one nitrogen atom in its ring system,AN is a negatively charged inorganic anion,w is 0, 1 or 2,x is 2 or 3,y is 2 or 3, andz is 2, 3 or 4.;The spin crossover complex is applied as a contrast agent in a magnetic resonance imaging of the human or animal body. | 1. A method for magnetic resonance imaging of a human or animal body, the method comprising: introducing into the human or animal body a contrast agent comprising a spin crossover complex of formula [M(Ly)Hw](AN)z wherein,M is Fe2+,Fe3+, or Co2+, L is 4′-(4′″-pyridyl)-1,2′:6′1″-bispyrazolylpyridine,AN is a negatively charged inorganic anion,w is 0, 1, or 2,y is 2 or 3, andz is 2, 3, or 4; andconducting magnetic resonance imaging on the human or animal body. 2. The method as recited in claim 1, wherein M is Fe2+. 3. The method as recited in claim 1, wherein AN is BF4−, Cl−, SO42−, or ClO4−. 4. The method as recited in claim 1, wherein y is 2. 5. The method as recited in claim 1, wherein M is Fe3+. 6. The method as recited claim 1, wherein M is Co2+. 7. The method as recited in claim 2, wherein AN is BF4−. 8. The method as recited in claim 2, wherein AN is Cl−. 9. The method as recited in claim 2, wherein AN is SO42−. 10. The method as recited in claim 5, wherein AN is ClO4−. 11. The method as recited in claim 5, wherein AN is BF4−, SO42−, or Cl−. 12. The method as recited in claim 5, wherein AN is ClO4−. 13. The method as recited in claim 6, wherein AN is BF4−, SO42−, or Cl−, or ClO4−. 14. The method as recited in claim 1, wherein the contrast agent consists essentially of the spin crossover complex, and AN is BF4−, SO42−, Cl−, or ClO4−. 15. The method as recited in claim 4, wherein M is Fe2+. 16. The method as recited in claim 1, wherein the spin crossover complex has a formula: 17. The method as recited in claim 1, wherein the introducing is intravenous. 18. The method as recited in claim 1, wherein y is 3. 19. The method as recited in claim 4, wherein AN is BF4−, Cl−, SO42−, or ClO431 . | Karlsruher Institut fuer Technologie | A61K049/06 A61K049/00 A61K031/4439 A61K031/444
US08187183 | B2 | 2012-05-29 | Continuous glucose monitoring system and methods of use | 2010-10-11 | A continuous glucose monitoring system including a sensor configured to detect one or more glucose levels, a transmitter operatively coupled to the sensor, the transmitter configured to receive the detected one or more glucose levels, the transmitter further configured to transmit signals corresponding to the detected one or more glucose levels, and a receiver operatively coupled to the transmitter configured to receive transmitted signals corresponding to the detected one or more glucose levels, and methods thereof, are disclosed. In one aspect, the transmitter may be configured to transmit a current data point and at least one previous data point, the current data point and the at least one previous data point corresponding to the detected one or more glucose levels. | 1. A glucose monitoring system, comprising: a sensor configured to sample glucose levels of a user;a transmitter unit operatively coupled to the sensor, the transmitter unit comprising sensor electronics operatively coupled to the sensor to receive the sampled glucose levels and to transmit signals corresponding to the sampled glucose levels, a memory for storing a plurality of data points corresponding to the sampled glucose levels of the user, and a processor configured to generate one or more data packets for data transmission via the sensor electronics; anda receiver unit operatively coupled to the transmitter unit, the receiver unit configured to establish a wireless communication range with the sensor electronics of the transmitter unit and receive the one or more data packets from the sensor electronics of the transmitter unit when the receiver unit is placed within the established wireless communication range; wherein the sensor electronics include programming to transmit the one or more data packets with each data transmission, the one or more data packets including a current data point that corresponds to a glucose level of a current time period and at least one previous data point that corresponds to at least one glucose level of at least one previous time period, such that if the receiver unit does not successfully receive the current data point from the transmitter unit, at a subsequent data transmission, the current data point and the at least one previous data point are re-transmitted to the receiver unit with a new current data point, and the current data point is received at the receiver unit and ascertained from the subsequent data transmission. 2. The system of claim 1, wherein the one or more data packets are transmitted using a radio frequency (RF) communication link. 3. The system of claim 1, wherein the wireless communication range is established based on a wireless signal strength from the transmitter unit. 4. The system of claim 1, wherein the transmitter unit further includes a temperature detection section configured to monitor a temperature of an insertion site of the sensor to obtain a temperature reading, wherein the temperature reading is used to adjust the sampled one or more glucose levels. 5. The system of claim 4, wherein the receiver unit is further configured to receive at least one temperature reading related signal when the receiver unit is placed within the established wireless communication range. 6. The system of claim 1, wherein the transmitter unit is configured to encode the one or more data packets. 7. The system of claim 6, wherein the one or more data packets is encoded using Reed-Solomon encoding. 8. The system of claim 1, wherein the glucose level from the at least one previous time period corresponds to at least one glucose level preceding the current glucose level. 9. The system of claim 1, wherein the transmitter unit includes a switch configured to disable one or more functionalities o... | Abbott Diabetes Care Inc. | A61B005/00 G08B001/08 G06F007/00 G06F017/00
US08206299 | B2 | 2012-06-26 | Image guided high intensity focused ultrasound treatment of nerves | 2010-09-21 | A method for using high intensity focused ultrasound (HIFU) to treat neurological structures to achieve a desired therapeutic effect. Depending on the dosage of HIFU applied, it can have a reversible or irreversible effect on neural structures. For example, a relatively high dose of HIFU can be used to permanently block nerve function, to provide a non-invasive alternative to severing a nerve to treat severe spasticity. Relatively lower doses of HIFU can be used to reversibly block nerve function, to alleviate pain, to achieve an anesthetic effect, or to achieve a cosmetic effect. Where sensory nerves are not necessary for voluntary function, but are involved in pain associated with tumors or bone cancer, HIFU can be used to non-invasively destroy such sensory nerves to alleviate pain without drugs. Preferably, ultrasound imaging synchronized to the HIFU therapy is used to provide real-time ultrasound image guided HIFU therapy of neural structures. | 1. A method for using a high intensity focused ultrasound (HIFU) probe to provide therapy to a peripheral nerve at a treatment site in a patient, the method comprising the steps of: (a) positioning, external to the body, (i) a HIFU therapy probe having a focal point, and (ii) an ultrasound imaging device in proximity to the peripheral nerve;(b) creating an image of the treatment site with the imaging device by generating a sequence of ultrasonic imaging pulses that propagate to the treatment site;(c) verifying the position of the focal point of the HIFU probe relative to the treatment site in the image created in step (b) using imaging ultrasound;(d) after said verifying, administering a dose of HIFU to the peripheral nerve to achieve a therapeutic effect; and(e) during said administering, synchronizing and pulsing an HIFU wave produced by the HIFU probe with respect to the imaging pulses such that the HIFU probe's focal point can be viewed in the image during treatment. 2. The method of claim 1, wherein the imaging ultrasound comprises three dimensional (3D) ultrasound. 3. The method of claim 1, wherein the imaging ultrasound comprises two dimensional (2D) ultrasound. 4. The method of claim 1, wherein the step of verifying the position of the focal point of the HIFU probe relative to the treatment site comprises the steps of: (a) energizing the HIFU probe using a power level sufficient to change an echogenicity of tissue proximate the focal point, but insufficient to achieve a therapeutic effect; and(b) using the imaging ultrasound to detect the change in echogenicity, thereby verifying the position of the focal point. 5. The method of claim 1, further comprising the step of delivering a blood soluble agent to the patient prior to administering the dose of HIFU. 6. The method of claim 5, wherein the blood soluble agent is an anesthetic. 7. The method of claim 1, wherein the step of administering the dose of HIFU to the peripheral nerve comprises the step of using a HIFU power output level ranging from about 50 to about 2000 W/cm2. 8. The method of claim 1, wherein the step of administering the dose of HIFU to the peripheral nerve comprises the step of using a HIFU power output level ranging from about 1480 to about 1850 W/cm2. 9. The method of claim 1, wherein step (c) includes (i) generating a synchronization pulse that establishes a timing reference point for generating the sequence of ultrasonic imaging pulses in step (b), and (ii) pulsing the HIFU probe after a given time delay from the synchronization pulse sufficient to shift noise from the HIFU pulse away from the treatment site. | University of Washington | A61B008/00 A61H001/00
US07874999 | B2 | 2011-01-25 | Detecting access disconnect using needle sleeve | 2007-09-24 | An access needle for extracorporeal therapy is equipped with a mount and a sleeve having a metal or magnetic component at its distal end. The mount includes a sensor for detecting when the distal end is pressed against the mount. The sensor is a hall-effect sensor or a proximity sensor. When therapy is begun, the sleeve is rolled up and urged against the mount, exposing the needle for use with the patient. While the sleeve, and the piece of metal or magnet, remains in contact with the sensor, the needle has not been dislodged and therapy may continue. If the needle is dislodged, the sleeve pushes away from mount, moving the metal or magnet away from the sensor. The sensor notes the dislodgement and sends a signal to alert the patient or a caregiver. | 1. A method of detecting a needle disconnection, the method comprising: providing a needle having a mount with a sensor and a sleeve with a metallic target, the sensor configured for detecting the metallic target;placing the needle into an extracorporeal blood processing access site;urging the sleeve against the mount, thereby placing at least a portion of the metallic target near the sensor;taking a reading of the sensor after the metallic target is urged against the mount;monitoring the access site by taking additional readings of the sensor during an extracorporeal blood processing therapy; andsending a signal if a reading consistent with absence of the target from the sensor is detected, thereby indicating dislodgement of the needle. 2. The method of claim 1, further comprising taking a baseline reading of the sensor before the step of urging the sleeve against the mount. 3. The method of claim 1, wherein the signal is sent to a local output device to alert a person or to sound an alarm. 4. The method of claim 1, wherein the sensor is a hall-effect sensor or a proximity sensor. 5. A method for detecting needle disconnection at an access site, the method comprising: providing an access needle with an electroactive sleeve and a sensor on a mount of the access needle, wherein the sensor is configured to detect metal or a magnet on the sleeve;mounting the access needle on the access site of a patient and pressing the sleeve toward the mount;taking a reading of the sensor with the pressed sleeve;monitoring the access site during a therapy procedure by taking additional readings with the sensor; andsending a signal if a reading indicative of needle dislodgement is taken. 6. The method of claim 5, further comprising taking a first baseline reading of the sensor and the sleeve while the sleeve covers the needle, and saving data of the first baseline reading for comparison with the baseline reading. 7. The method of claim 5, further comprising sending a local alert as a result of sending the signal. 8. The method of claim 5, further comprising saving data of the reading. 9. A system for detecting needle dislodgement, the system comprising: an access needle suitable for mounting near an access site for extracorporeal therapeutic blood processing, the access needle including a mount;a sleeve having an electromagnetic target, the sleeve positioned over the needle and adjacent the mount;a sensor on the mount for detecting the electromagnetic element; anda communications circuit in communication with the sensor, wherein the needle is configured for placing into the access site, the sleeve is configured for pressing against the mount, the sensor is configured for detecting the electromagnetic element, and the communications circuit is configured for sending data indicative of a position of the electromagnetic element with respect to the sensor. 10. The system according to claim 9, wherein the electromagnetic target is selected from the group consis... | Baxter International, Inc. Baxter Healthcare S.A. | A61M037/00 A61M005/32 A61B005/00
US08233975 | B2 | 2012-07-31 | Method utilizing two wire electrode oscillator system for determining body impedance | 2010-09-24 | A method for measuring an impedance of a body includes the steps of electrically coupling first and second electrodes to the body and electrically biasing a variable frequency relaxation oscillator. The oscillator has both a positive resistive feedback network and a negative feedback network which includes the electrodes. The method further includes a step of measuring the period of oscillation of the variable frequency oscillator. The measured period of oscillation is indicative of the impedance of the body. | 1. A method for calibrating an apparatus for determining body impedance having first and second electrodes being suitable for being electrically coupled to the body, comprising: determining a first period of oscillation for the apparatus associated with an open circuit condition between the electrodes; anddetermining a second period of oscillation for the apparatus associated with a short-circuit condition between the electrodes;wherein, the first period of oscillation is indicative of a first calibration value, and the second period of oscillation in combination with a slope coefficient is indicative of a second calibration value. 2. The method of claim 1, wherein the slope coefficient is indicative of a calibration. 3. The method of claim 1, further comprising storing data indicative of the first period of oscillation. 4. A method for measuring the bioelectrical impedance of a body, comprising: electrically coupling first and second electrodes to the body;electrically biasing a variable frequency relaxation oscillator having both a negative feedback network comprising said electrodes, and a positive resistive feedback network;measuring the period of oscillation of said variable frequency oscillator, andcalculating the bioelectrical impedance of the body using the measured period of oscillation. 5. The method of claim 4, wherein the oscillator comprises at least one operational amplifier. 6. The method of claim 5, wherein the first of said electrodes is resistively coupled to an output terminal of said operational amplifier. 7. The method of claim 6, wherein the second of said electrodes is coupled to common through at least one capacitor and shorted to an inverting input terminal of said operational amplifier. 8. The method of claim 7, wherein said at least one capacitor charges and discharges exponentially through said electrodes with respect to said operational amplifier non-inverting input terminal. 9. The method of claim 4, further comprising the step of determining, by a microprocessor, the lean body mass of the body as a function of the weight of the body and the measured impedance of the body. 10. The method of claim 4, further comprising the step of determining, by a microprocessor, the lean body mass of the body as a function of the weight of the body, the measured impedance of the body and one or more of the height, a fitness factor, the sex and the age of the body. 11. The method of claim 10, wherein one or more coefficients for determining said lean body mass are stored in a memory accessible to the microprocessor and correspond to one or more of the height, the weight, the fitness factor, the measured impedance, the sex and the age of the body. 12. The method of claim 9, further comprising the step of determining, by the microprocessor, the percentage of body fat of the body, as a function of the lean body mass and the weight of the body. 13. The method of claim 9, further comprising the step of determining, by the micr... | Measurement Ltd. | A61B053/53 G01R035/00
US08670812 | B2 | 2014-03-11 | Physiological sensor having a waist | 2010-07-23 | An exemplary sensor includes a sensor pad having a first portion and a second portion separated by a waist portion. The waist portion is narrower than the first portion and the second portion. A light source that is disposed on the first portion is configured to generate near-infrared light and transmit the near-infrared light through part of a patient's body. A light detector that is disposed on the second portion is configured to detect near-infrared light that has traveled through the part of the patient's body. The near-infrared light detected indicates an amount of oxygen in the part of the patient's body through which the near-infrared light traveled. | 1. A sensor comprising: an integrated sensor pad having a first portion and a second portion separated by a waist portion, wherein the waist portion is narrower than the first portion and the second portion, and wherein said sensor pad is configured to be adhered to a patient's body by an adhesive disposed directly on an underside of said sensor pad;a light source at least partially housed within said sensor pad and configured to generate near-infrared light and transmit the near-infrared light through part of a patient's body; anda light detector at least partially housed within said sensor pad and configured to detect near-infrared light that has traveled through the part of the patient's body, wherein the near-infrared light detected indicates an amount of oxygen in the part of the patient's body through which the near-infrared light traveled. 2. A sensor as set forth in claim 1, wherein the second portion is configured to substantially prevent the light detector from receiving interfering light. 3. A sensor as set forth in claim 1, wherein at least one of the first portion, the waist portion, and the second portion includes a light blocking material. 4. A sensor as set forth in claim 1, further comprising a liner disposed on the adhesive. 5. A sensor as set forth in claim 4, wherein the liner defines a slit. 6. A sensor as set forth in claim 5, wherein the slit is located over the waist portion to facilitate removal of the liner. 7. A sensor as set forth in claim 1, wherein at least one of the first portion, the second portion, and the waist portion includes a flexible material. 8. A sensor as set forth in claim 1, wherein the light detector is disposed on the second portion and wherein the second portion has a larger surface area than the first portion. 9. A sensor as set forth in claim 1, wherein the sensor pad is formed from multiple layers having generally the same size and shape and forming the first portion, the second portion, and the waist portion. 10. A sensor pad comprising: a first portion configured to house a light source that generates near-infrared light and transmits the near-infrared light through a part of a patient's body;a second portion configured to house a light detector that detects near-infrared light after the near-infrared light has traveled through a part of the patient's body, wherein the light received indicates oxygen saturation of the part of the patient's body through which the light traveled; anda waist portion disposed between the first portion and the second portion, wherein the surface area of the second portion is larger than the surface area of the first portion and the waist portion is narrower than the first portion and the second portion. 11. A sensor pad as set forth in claim 10, wherein the second portion includes a light blocking material that substantially prevents the light detector from receiving interfering light. 12. A sensor pad as set forth in claim 10, wherein at least one of the... | Covidien LP | A61B005/00
US07878055 | B2 | 2011-02-01 | Sensor and analyzer for determining physiological limitations | 2006-05-22 | A sensing and analyzing device for determining when physical limits of a patient have been exceeded by the patient during exercise is provided. The device may comprise a sensing unit, analyzing unit and a warning unit. The sensing unit may be attached to critical locations on a body of the patient. The sensing unit may be operative to sense pressure. The analyzing unit may be attached to a convenient location on the patient's body. The analyzing unit may be operative to receive pressure data from the sensing unit to determine whether the patient has exceeded the physical limits of the patient. The warning unit may communicate with the analyzing unit and be activated by the analyzing unit when the physical limit of the patient has been exceeded for notifying the patient when the patient's physical limits have been exceeded. | 1. A sensing and analyzing device for determining when a patient has exceeded a range of motion during exercise, the device comprising: a. a sensing unit including a base pad having a plurality of first flex cables extending therefrom, and a plurality of spaced pressure pads connectable to the bottom of a patient's foot, each pressure pad being engaged to a respective first flex cable and configured to separately sense pressure applied by the patient proximate each pressure pad, the plurality of spaced pressure pads configured to each generate a responsive electrical pressure signal, the pressure signals being collectively representative of the a relative distribution of pressure on a patient's foot;b. an analyzing unit attachable to a convenient location on the patient's body, the analyzing unit includes adjustable threshold logic operative to receive the pressure signals from the sensing unit to determine whether the relative distribution of pressure on a patient's foot relates to a physiological orientation that exceeds the range of motion; andc. a warning unit being communicable with the analyzing unit and being activatable by the analyzing unit when the range of motion has been exceeded for notifying the patient when the range of motion has been exceeded. 2. The device of claim 1 wherein the plurality of pressure pads have a domed configuration which applies pressure to a conductor in the respective first flex cable when a force is applied to the pressure pads to change a resistance of the conductor. 3. The device of claim 1 further comprising a second flex cable for providing a communication pathway between the sensing unit and the analyzing unit to transfer data from the sensing unit to the analyzing unit. 4. The device of claim 3 wherein the second flex cable comprises: a. an upper portion communicable with the analyzing unit;b. a lower portion communicable with the sensing unit and the upper portion; andc. connector attached to the upper portion and the lower portion for detachably connecting the upper portion to the lower portion. 5. The device of claim 4 further comprising a remote device connectable to the connectors of the upper portion for downloading data stored in a memory of the analyzing unit. 6. The device of claim 3 wherein the second flex cable is conformable to the leg of a user. 7. The device of claim 3 wherein the second flex cable lower portion and the base pad are formed as an integral flex cable. 8. The device of claim 1 further comprising a remote device connectable to the analyzing unit for downloading data stored in a memory of the analyzing unit. 9. The device of claim 1 wherein the analyzing unit comprises: i. a memory for storing sensed data received from the sensing unit; andii. a clock and a sampler logic network for transferring data from the sensing unit to the memory in a timed sequence. 10. The device of claim 1 wherein the warning unit comprises a buzzer, audio alarm, a light, or combinations there... | - | A61B005/00 A61B005/103 G08B023/00
US20110040155 | A1 | 2011-02-17 | MULTIPLE SENSORY CHANNEL APPROACH FOR TRANSLATING HUMAN EMOTIONS IN A COMPUTING ENVIRONMENT | 2009-08-13 | Sensory inputs of a user can be received by a computing device. At least one of the sensory inputs can include a physiological input providing a physiological measurement from a body of the user. Each sensory input can be processed in a unique one of a set of standards-defined sensory channels, each corresponding to a specific emotion dimension. Processing the sensory inputs can transform the physiological measurement into an emotion dimension value. The emotion dimension values from each of the sensory channels can be aggregated to generate at least one emotion datum value, which is a standards-defined value for an emotional characteristic of the user. Historical data for a user can be optionally collected and used by a learning and calibration component to improve the accuracy of the generated emotion datum values for a specific individual. A programmatic action driven by the emotion datum value can be performed. | 1. A method for discerning human emotions in a computing environment comprising: receiving a plurality of sensory inputs of a user, wherein at least one of the sensory inputs comprises a physiological input providing a physiological measurement from a body of the user, said physiological measurement being obtained using a physiological sensor;processing each sensory input in a unique one of a plurality of standards-defined sensory channels, each standards-defined sensory channel corresponding to a specific emotion dimension, wherein said processing comprises transforming the physiological measurement into an emotion dimension value, said emotion dimension value abstracting the physiological measurement from specifics attributable to unique characteristics of the physiological sensor;aggregating the emotion dimension values from each of the sensory channels to generate at least one emotion datum value, which is a standards-defined value for an emotional characteristic of said user, wherein said emotion datum value is a value independent of any of said sensory devices and is a value independent of any single one of said standards-defined sensory channels; andperforming a programmatic action driven by the emotion datum value. 2. The method of claim 1, wherein said user is interacting with a graphical user interface to participate within a virtual environment in which said user is represented in said virtual environment as an avatar, which is presented to a plurality of other users of said virtual environment, wherein a presentation characteristic of said avatar is adjusted for the emotion datum value so that an emotion corresponding to the emotion datum value is presented to the other users. 3. The method of claim 2, wherein the presentation characteristic alters at least one of a facial expression of the avatar and a speech characteristic of the avatar in accordance with the emotion corresponding to the emotion datum value. 4. The method of claim 1, wherein the emotion datum value is generated without the user ever explicitly manually inputting semantic content for an emotion. 5. The method of claim 1, wherein said emotion datum value is an application independent value conforming to a defined standard, wherein a plurality of independent applications are able to consume the emotion datum value to drive application specific changes based upon emotions of the user. 6. The method of claim 1, further comprising: when aggregating the emotion dimension values, adjusting user specific factors specified within a stored user specific data set to normalize the generated emotion datum value to a baseline of an average person. 7. The method of claim 1, wherein said user participates in an application session of an application which performs said programmatic action driven by the emotion datum value, wherein during the application session, the plurality of sensory inputs from the user are continuously received, processed, and aggregated to continuously ge... | INTERNATIONAL BUSINESS MACHINES CORPORATION | A61B005/00 G06F003/048
US09144396 | B2 | 2015-09-29 | Apparatus and method of analyzing constituents of gas in oral cavity and alveolar gas | 2008-08-20 | An apparatus and a method of analyzing constituents of gas in an oral cavity and exhaled breath is disclosed. The apparatus for analyzing the constituents of gas in the oral cavity and exhaled breath according to the present invention includes a filter to filter the outside gas by adsorbing polar molecules and non-polar molecules in the outside air and by removing water in the outside air in order to use the outside gas as carrier gas. The apparatus can also include a plurality of solenoid valves for controlling the flow of a carrier gas; a sensor for detecting components of the exhaled breath; a pump to draw the gas in the oral cavity or the exhaled breath and the carrier gas and discharge the gases to the outside; a control unit for controlling the components of the apparatus; and a display device to display results calculated by the control unit. | 1. An apparatus for analyzing constituents of a gas in an oral cavity and an exhaled breath, the apparatus comprising: a filter, filled with a substance adsorbing a polar molecule and a non-polar molecule comprising silica gel and activated carbon, to filter an air;a first solenoid valve connected to the filter so as to provide the air passing through the filter thereinto and to control a flow of the gas in the oral cavity or the exhaled breath flowing through a connection port thereto;a second solenoid valve connected through a first connection tube to the first solenoid valve at a first port thereof, connected through a bypass tube to a third solenoid valve at a second port thereof so as to bypass the air, and connected through a sampling loop to the third solenoid valve at a third port thereof, the sampling loop filled with the gas in the oral cavity or the exhaled breath which is collected;a fourth solenoid valve connected through a second connection tube to the third solenoid valve;a fifth solenoid valve connected through an upper bypass tube provided with a speed controller to a port of the fourth solenoid valve;a column that is filled with a packing material for gas chromatography columns and that is connected to another port of the fourth solenoid valve to allow the gas in the oral cavity or the exhaled breath and the air to sequentially flow therethrough;a sensor chamber comprising a sensor connected to the column at an end thereof;a lower bypass tube connected to another end of the sensor chamber at an end thereof and to the fifth solenoid valve at another end thereof, the lower bypass tube provided with another speed controller;a suction pump configured at the rear portion of the column and connected to the fifth solenoid valve to draw the gas in the oral cavity or the exhaled breath and the air and then discharge the gas in the oral cavity or the exhaled breath and the air to the outside;a control unit connected to the respective solenoid valves, the sensor chamber, and the suction pump to control the respective solenoid valves, the sensor chamber, and the suction pump; anda display device to output results of concentrations of the gases measured by using signal and calculation treatments,wherein the control unit performs a sampling step, a measurement step, and a recovering step according to the following steps:when in the sampling step, the control unit operates:to close a first portion to the filter and to open a second portion to the connection port of the first solenoid valve, to close the second port which is connected to a bypass tube of the second solenoid valve, to open the third port which is connected to a sampling loop of the second solenoid valve, to open the third solenoid valve, to open a first port which is connected to an upper bypass tube of the fourth solenoid valve, and to close a second port which is connected to the column of the fourth solenoid valve;when in the measurement step, the control unit operates:to ope... | Yong Sahm Choe | A61B005/08 A61B005/097 G01N033/497 G01N001/22 G01N030/20