## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍 | 專利權人 | IPC
US07955278 | B1 | 2011-06-07 | Aspiration methods and devices for assessment of viscoelastic properties of soft tissues | 2009-11-30 | Methods for assessing viscoelastic properties of soft tissues are based on detecting an inflection point on a pressure-time plot when air is aspirated from a cavity placed over the tissue sample. A small diameter tube through which air aspiration is conducted is ultimately closed off by tissue being drawn into the cavity causing an abrupt change in pressure slope. First or second derivatives of the pressure-time plot can be used to detect the inflection point. Repeating the test with a different aspiration rates or after a predetermined relaxation time allows determining tissue viscosity and tissue creep in addition to tissue elasticity expressed as Young's modulus. | 1. An aspiration method for assessing viscoelastic properties of soft tissue, the method comprising the steps of: a) providing a test cavity in a sealed contact with said tissue, said test cavity includes an opening connected to an aspiration system, said opening located on a path of expected deformation of said tissue,b) activating said aspiration system to evacuate air through said opening from said cavity at a first rate of aspiration to draw said tissue into said cavity,c) conducting pressure measurements in said aspiration system and said test cavity as a function of time as said tissue is being drawn into said cavity, continue to draw said tissue into said cavity to cause closure of said opening by said tissue and separation of said test cavity from said aspiration system,d) detecting a first slope of said pressure measurements in said test cavity,e) detecting a first inflection point in said pressure measurements, said first inflection point defining a first pressure level and a first time point at which said pressure measurements deviate from said first slope indicating closure of said opening by said tissue and cessation of said tissue being drawn into said test cavity, said first inflection point is detected using a method selected from a group consisting of: i. detecting deviation of said first slope above a predetermined slope deviation threshold,ii. detecting a first derivative of said first slope exceeding a predetermined first derivative threshold, andiii. detecting a second derivative of said first slope exceeding a predetermined second derivative threshold,f) determining elasticity of said tissue using said first slope of said pressure measurements and said first pressure level at said first inflection point. 2. The method as in claim 1, wherein said step (d) further including constructing a first best-fit straight line through said measurements located prior to said inflection point and a second best-fit straight line through said measurements located after said inflection point, said method further including defining a corrected inflection point located at an intersection of said first best-fit straight line and said second best-fit straight line. 3. The method as in claim 1, wherein said test is stopped once said inflection point is determined and said cavity is immediately vented to release said tissue. 4. The method as in claim 1 further comprising venting said cavity to release said tissue, waiting during a first predetermined time to allow said tissue to relax, and repeating steps (b) through (f) wherein step (b) includes evacuating air from said cavity at a second rate of aspiration, said step (d) includes detecting a second slope of pressure measurements in said cavity, said step (e) includes detecting a second inflection point defining a second pressure level and a second time point at which said tissue is drawn into said test cavity, said step (f) further including determining tissue viscosity using said first press... | Artann Laboratories Inc. | A61B005/103 A61B005/117
US20110105871 | A1 | 2011-05-05 | MONITORING METHOD AND/OR APPARATUS | 2010-08-04 | A method and apparatus for substance monitoring. One application is an easy to handle continuous glucose monitor using a group of hollow out-of-plane silicon microneedles to sample substances in interstitial fluid from the epidermal skin layer. The glucose of the interstitial fluid permeates a dialysis membrane and reaches a sensor. Using MEMS technology, for example, allows well-established batch fabrication at low cost. | 1. A method of in vivo monitoring of an individual's interstitial fluid comprising: creating a plurality of fluid paths through a stratum corneum area of the individual's skin, the fluid paths each comprising a distal end in fluid communication with interstitial fluid of the individual, a proximal end in analyte communication with a sensing area located outside of the patient's body, and an interior space extending between the distal and proximal ends of the fluid path;allowing at least one analyte to passively diffuse from the patient's interstitial fluid through the fluid paths and into the sensing area; andsensing a concentration of the at least one analyte in the sensing fluid using a sensor located at least partially in the sensing area. 2. The method of claim 1 wherein fluid in the fluid paths is not moving when the at least one analyte is passively diffusing there-through. 3. The method of claim 1 wherein there is no fluid movement between the distal ends of the fluid paths and the sensor when the at least one analyte is passively diffusing through the fluid paths and into the sensing area. 4. The method of claim 1 wherein no interstitial fluid is extracted from the individual during the diffusion and sensing steps. 5. The method of claim 1 further comprising displaying analyte concentration information remote from the stratum corneum area of the individual's skin. 6. The method of claim 5 further comprising wirelessly communicating analyte concentration information to a display. 7. The method of claim 1 wherein the sensing step is performed by a sensor in fluid communication with the sensing area and the interior spaces of the fluid paths, the method further comprising calibrating the sensor by moving sensing fluid into the sensing area. 8. The method of claim 7 further comprising moving sensing fluid out of the sensing area as sensing fluid is moved into the sensing area. 9. The method of claim 1 further comprising attaching a device comprising the sensor area and the sensor to the individual with adhesive. 10. The method of claim 1 further comprising providing a membrane located between the distal ends of the fluid paths and the sensor. 11. The method of claim 1 wherein creating the plurality of fluid paths through the stratum corneum area of the individual's skin comprises applying a microneedle array to the stratum corneum area. 12. The method of claim 1 wherein the proximal ends of the fluid paths are in fluid communication with the sensing area. | The Regents of the University of California | A61B005/145
US20110150778 | A1 | 2011-06-23 | Assembly of Gas-Filled Microvesicle With Active Component For Contrast Imaging | 2011-01-31 | Assembly comprising a gas-filled microvesicle and a structural entity which is capable to associate through an electrostatic interaction to the outer surface of said microvesicle (microvesicle associated component—MAC), thereby modifying the physico-chemical properties thereof. Said MAC comprises a targeting ligand a diagnostic agent or any combination thereof. Optionally a bioactive agent can further be associated to the MAC. The assembly of the invention can be formed from gas-filled microbubbles or microballoons and a MAC having preferably nanometric dimensions, e.g. a micelle, and is used as an active component in diagnostically and/or therapeutically active formulations, in particular for enhancing the imaging in the field of ultrasound contrast imaging, including targeted ultrasound imaging, ultrasound-mediated drug delivery and other imaging techniques such as molecular resonance imaging (MRI) or nuclear imaging. | 1. An assembly comprising a gas-filled microvesicle bearing a first overall net charge and a component associated with said microvesicle wherein said component bears a second overall net charge opposite in sign to said first net charge and comprises a targeting ligand, a diagnostic agent or any combination thereof, and a biocompatible surface active agent. 2. An assembly according to claim 1 wherein said targeting ligand is selected from proteins, antibodies, antibody fragments, receptor molecules, receptor binding molecules, glycoproteins, lectins, peptides, oligopeptides, polypeptides, peptidomimetics, saccharides, polysaccharides, vitamins, steroids, steroid analogs, hormones, cofactors, bioactive agents, genetic material, nucleosides, nucleotides and polynucleotides. 3. An assembly according to claim 1 wherein said diagnostic agent is selected from magnetite nanoparticles, iodinated compounds and paramagnetic ion complexes. 4. An assembly according to claim 1 wherein said component associated with said gas-filled microvesicle further comprises a bioactive agent. 5. An assembly according to claim 1 further comprising at least a second component having an overall net charge opposite in sign to said first net charge and comprising a bioactive agent. 6. An assembly according to claim 1 further comprising at least a second component having an overall net charge equal in sign to the charge of the microvesicle. 7. An assembly according to claim 6 wherein said second component comprises a targeting ligand, a diagnostic agent, a bioactive agent or any combination thereof. 8. An assembly according to claim 1, wherein said component associated with said gas-filled microvesicle has a diameter of 300 nm or less. 9. An assembly according to claim 1 wherein said biocompatible surface active agent is an emulsifying agent, a dispersing agent or a mixture thereof. 10. An assembly according to claim 1 wherein said biocompatible surface active agent is an amphiphilic material. 11. An assembly according to claim 1 wherein said biocompatible surface active agent is selected among (C2-C10) organic acids, organic fatty acids comprising a (C12-C24) aliphatic chain, pharmaceutically acceptable salts thereof, esters thereof with polyoxyethylene; polyionic (alkali) salts; organic amines; amides; quaternary amine salts; aminoacids; phospholipids; esters of mono- or oligo-saccharides with (C12-C24), organic fatty acids; organic sulfonates; perfluoroorganic acids; polymeric surfactants; and mixtures thereof. 12. An assembly according to claim 1 wherein the ratio between the number of charges per mole of microvesicles and the number of charges per mole of the second component is from about 10:1 to about 1:10. 13. An assembly according to claim 12 wherein said ratio is of about 3:1 or less. 14. An assembly according to claim 12 wherein said ratio is of about 2:1 or less. 15. An assembly according to claim 12 wherein said ratio is of about 3:2 or less. 16. An a... | BRACCO RESEARCH S.A. | A61B005/055 A61B008/00 A61K039/395 A61K009/00 A61K038/00 A61K038/14 A61K031/70 A61K031/715 A61K031/56 A61K038/22 A61K031/7088 A61P043/00
US20110118594 | A1 | 2011-05-19 | MEDICAL IMPLANT TARGETING SYSTEM AND METHOD | 2009-08-19 | A system (10) and method for targeting a geometric feature (12) in a medical implant (14) located in a patient (16), such as a transverse bore (12) in an intramedullary nail (14) implanted in a femur (18) of a patient (16). The system (10) includes one or more emitters (20), one or more receptors (22), and a processor (24). The emitter(s) emits at least one of an acoustic field, a magnetic field, an electric field, or an electromagnetic field that will be altered by the geometric feature (12), and the receptor(s) (22) detects the altered field and generates a signal (26) responsive to the altered field. The processor receives the signal and generates a representation of the altered field. | 1. A system to target a geometric feature in a medical implant located in a patient, the system comprising: at least one emitter external to the implant to emit at least one of an acoustic field, a magnetic field, an electric field and an electromagnetic field that will be altered by the geometric feature;at least one receptor external to the implant to generate a signal responsive to the altered field; anda processor to receive the signal and generate a representation of the altered field. 2. The system of claim 1 further comprising a frame and a carrier mounted to the frame for translation and rotation along and about two axes relative to the medical implant, the at least one emitter and the at least one receptor mounted on the carrier for movement therewith along and about the two axes. 3. The system of claim 2 wherein the carrier is further mounted to the frame for translation and rotation along and about another axis relative to the medical implant. 4. The system of claim 2 wherein the at least one emitter and the at least one receptor are located on the carrier so that the medical implant can be positioned between the at least one receptor and the at least one emitter with the medical implant located in a patient. 5. The system of claim 2 wherein the at least one emitter and the at least one receptor are located on the carrier so that the medical implant can be positioned to one side of the at least one receptor and the at least one emitter with the medical implant located in a patient. 6. The system of claim 1 further comprising a comparator to compare at least a portion of the representation of the altered field to a shape of the known geometric feature. 7. The system of claim 2 further comprising a controller to position the carrier relative to the medical implant in response to the representation of the altered field. 8. The system of claim 7 wherein the controller comprises a comparator to compare at least a portion of the representation of the altered field to a shape of the known geometric feature, and the controller is responsive to the comparator to position the carrier relative to the implant. 9. The system of claim 7 wherein the controller is responsive to a representation of relative changes in an intensity of the altered field with respect to a reference coordinate system. 10. The system of claim 1 further comprising a surgical tool guide to be targeted at the geometric feature in response to the representation of the altered field. 11. The system of claim 2 further comprising a surgical tool guide mounted on the carrier to be targeted at the geometric feature in response to the representation of the altered field. 12. The system of claim 1 further comprising a user interface to transmit the representation of the altered field to a user. 13. The system of claim 12 wherein the user interface comprises at least one of a visual display and an audio signal. 14. The system of claim 1 wherein the at least one emitter com... | - | A61B005/05
US09138174 | B2 | 2015-09-22 | Displacement measurement in a fall detection system | 2009-05-05 | There is provided a personal fall detection system comprising a sensor module that is suitable for attachment to a user, the sensor module comprising an accelerometer for measuring accelerations acting on the sensor module; and a processor for analyzing the measured accelerations to determine if the sensor module has suffered an impact, and for estimating a vertical displacement of the sensor module during a time period immediately before the impact; wherein the processor estimates the vertical displacement of the sensor module by twice integrating the measurements from the accelerometer. | 1. A personal fall detection system configured to be attached to a user, the system comprising: an accelerometer configured to measure acceleration along at least two directions and output time stamped signals indicative of the measured acceleration; anda computer processor configured to: receive the time-stamped signals,from the time stamped signals, determine occurrence of an impact and an impact time at which the impact occurred,from the time stamped signals time stamped with times preceding the impact time, determine an onset time at which acceleration that ended with the detected impact started,calculate a gravitational acceleration at the onset time from the measured acceleration in the vertical direction at the onset time,from the time stamped signals between the onset time and the impact time, determine a vertical displacement of the accelerometer that occurred between the onset time and the impact time,classifying the impact as a fall and determining an alarm condition when the vertical displacement is between 0.5 and 1.5 m. 2. The personal fall detection system according to claim 1, wherein the processor is further is configured to: subtract the gravitational acceleration from the measured acceleration between the onset time and the impact time. 3. The personal fall detection system according to claim 1, further including: a transceiver and an antenna configured to transmit an alarm signal when the impact is classified as a fall. 4. The personal fall detection system according to claim 1, wherein the computer processor is configured to determine the occurrence of the impact in response to the measured accelerations being characteristic of a collision with a solid object. 5. The personal fall detection system according to claim 4, wherein the acceleration characteristic includes a change in the acceleration exceeding a predetermined threshold and wherein the processor is configured to determine the occurrence of the impact in response to changes in the measured acceleration exceeding the predetermined threshold. 6. The personal fall detection system according to claim 1, wherein the computer processor is configured to determine the onset time from the time stamped signals which are time stamped with times preceding the impact time by identifying at least one of: (a) a stable period in which the measured accelerations vary by less than a predetermined time threshold, the onset time being the time at which the stable period ends, and(b) a variation in a direction of the measured acceleration that is beyond a predetermined acceleration variation threshold, the onset time being the time at which the variation in direction exceeded the acceleration variation threshold. 7. The personal fall detection system according to claim 1, wherein the computer processor is configured to: determine a first onset time from a time at the end of a stable period; anddetermine a second onset time from a time at which a variation in direction exceeds an a... | KONINKLIJKE PHILIPS N.V. | G08B023/00 A61B005/117 A61B005/103 A61B005/11 A61B005/00 G08B021/04
US08936552 | B2 | 2015-01-20 | Physical activity monitor and data collection unit | 2009-05-13 | A physical activity data collection unit includes one or more infrared sensors configured to provide an output indicative of a pulse rate of a user of the physical activity data collection unit, at least one temperature sensor configured to provide an output indicative of at least a body temperature of the user, and at least one accelerometer configured to provide an output indicative of movements of the user. The physical activity data collection unit can also include a microcontroller configured to determine a pulse rate, a body temperature, and movement characteristics of the user of the data collection unit based on outputs from the one or more infrared sensors, the at least one temperature sensor, and the at least one accelerometer; determine a physical exertion level of the user based on one or more of the pulse rate, the body temperature, or the movement characteristics of the user; and store, in a memory, data indicative of the physical exertion level during a time period during which the physical exertion level exceeds a predetermined threshold. | 1. A physical activity data collection unit, including: two or more infrared sensors arranged on the data collection unit such that when the data collection unit is worn by a user, the two or more infrared sensors are positioned adjacent to a wrist of the user, the two or more infrared sensors being configured to provide an output indicative of a pulse rate of the user of the physical activity data collection unit;at least one temperature sensor configured to provide an output indicative of at least a body temperature of the user;at least one accelerometer configured to provide an output indicative of movements of the user; anda microcontroller configured to: receive, from the two or more infrared sensors, two or more output signals, including one output signal received from each one of the two or more infrared sensors;determine a noise level value for each one of the two or more received output signals;determine a weighting factor for each one of the two or more received output signals based on the noise level value determined for each one of the two or more received output signals, wherein a first weighting factor determined for an output of one of the two or more infrared sensors having a first noise level is greater than a second weighting factor determined for an output of another of the two or more infrared sensors having a second noise level higher than the first noise level;obtain weighted outputs by assigning the determined weighting factors to corresponding ones of the two or more received output signals;determine a pulse rate, a body temperature, and movement characteristics of the user of the data collection unit based on the weighted outputs, the output provided by the at least one temperature sensor, and the output provided by the at least one accelerometer, respectively;determine a physical exertion level of the user based on one or more of the pulse rate, the body temperature, or the movement characteristics of the user; andstore, in a memory, data indicative of the physical exertion level during a time period during which the physical exertion level exceeds a predetermined threshold. 2. The physical activity data collection unit of claim 1, further including a transceiver, and wherein the microcontroller is further configured to transmit data indicative of the exertion level of the user to an external data collection portal. 3. The physical activity data collection unit of claim 2, wherein the microcontroller is configured to determine the presence of stored physical exertion data and automatically initiate transmission of the stored data upon determining that the data collection unit has come within a predetermined range of the data collection portal. 4. The physical activity data collection unit of claim 2, wherein the microcontroller is configured to determine the presence of stored physical exertion data and initiate transmission of the stored data upon receiving an interrogation signal from the data collection portal. 5.... | HeartMiles, LLC | A61B005/00 A61B005/0205 A61B005/11 A61B005/024 A61B005/1455
US09020577 | B2 | 2015-04-28 | Systems, devices and methods for cartilage and bone grafting | 2009-05-29 | Disclosed are systems, methods, devices and products to identify suitable donor sites for harvesting bone-cartilage grafts and to implant such bone-cartilage grafts. In some embodiments, a method includes providing a computer having access to a donor database, the donor database comprising information on each of a plurality of donor joint sites of the body, receiving first data relating to a defect of a joint of a patient, the defect comprising an area of bone, a portion of which includes at least one of, for example, missing and/or damaged cartilage, and identifying, based on the first data, at least one donor site from the donor database of joints for harvesting a graft of bone and cartilage to repair the defect. | 1. A method to identify suitable donor sites for bone-cartilage grafts for repairing a defect in a joint of a patient, the method comprising: providing a computer having access to a donor database, the donor database comprising information related to each of a plurality of donor joint sites of the body;receiving first data relating to the defect of the joint of the patient, the defect comprising an area of bone, a portion of the defect including at least one of missing and damaged cartilage;identifying, using the first data, at least one donor site from the donor database of joints for harvesting a graft of at least one of bone and cartilage to repair the defect; andgenerating a harvesting template using the first data adapted for use at the at least one identified donor site,wherein the harvesting template guides, through physical interaction with a cutting implement, at least one of (i) removal of at least a portion of the defect, and (ii) creation of a graft from the at least one identified donor site. 2. The method of claim 1, wherein the first data comprises imaging data of an area including at least one of the defect and an area around the defect. 3. The method of claim 2, wherein the first data comprises data representative of a surface model of at least one of the defect and the area around the defect. 4. The method of claim 2, wherein the imaging data is captured via at least one of CT imaging, MRI imaging, X-ray imaging and laser imaging. 5. The method of claim 1, wherein receiving the first data comprises accessing from the donor database data representative of a non-damaged structure corresponding to the defect. 6. The method of claim 1, wherein identifying the at least one donor site comprises: performing comparisons of data representative of surface models of donor sites retrieved from the donor database to the first data, the first data comprising data representative of a surface model of at least one of the defect and an area around the defect; and determining, using the comparisons, the at least one donor site. 7. The method of claim 6, wherein the comparisons are performed using the surface model of the at least one of the defect and the area around the defect in different spatial orientations. 8. The method of claim 6, wherein performing the comparisons is performed according to a hierarchy of matching criteria. 9. The method of claim 8, wherein the matching criteria include one or more of: a) dimensions of the joint donor sites, the defect and the area around the defect, and b) impact of removal from the respective joint donor sites of one or more of cartilage and bone structure. 10. The method of claim 1, wherein the at least one donor site from the donor database includes at least one foot joint selected from the group consisting of a calcaneal-cuboid joint, intercuneiform joints, arsometatarsal joints, lesser metatasophalageal joints and interphalangeal joints. 11. The method of claim 1, further comprising: harves... | Yale University | A61B005/05 G06F019/00 A61F002/46 A61B017/00 A61B017/56
US08649835 | B2 | 2014-02-11 | Method of performing wavelength modulated differential laser photothermal radiometry with high sensitivity | 2010-11-17 | Methods are provided for the detection of an analyte in a sample using wavelength modulated differential photothermal radiometry with enhanced sensitivity. A wavelength modulated differential photothermal radiometry system, comprising two optical modulated beams, where each beam experiences different absorption by the analyte, is calibrated by controlling the relative phase difference between the modulated beams so that individual photothermal signals corresponding to each modulated beam are 180° out of phase, corresponding to peak sensitivity to analyte concentration. The system may be further calibrated by varying the relative intensities of the two modulated beams and measuring standards containing known analyte concentration in order to determine an optimal relative intensity for a given concentration range of interest. | 1. A method of operating a wavelength modulated differential photothermal radiometry system for detecting the presence of an analyte within a substance, said method comprising the steps of: providing a first optical beam and a second optical beam, wherein said first optical beam is characterized by a first wavelength and said second optical beam is characterized by a second wavelength, and wherein said first wavelength is selected so that analyte concentrations within a concentration range of interest will cause sufficient absorption of said first optical beam to produce detectable thermal-wave emission, and wherein an absorption coefficient of said analyte at said first wavelength exceeds that at said second wavelength;producing a first modulated beam and a second modulated beam by modulating an intensity of said first optical beam and an intensity of said second optical beam, respectively; wherein said first modulated beam and said second modulated beam are modulated at a substantially equal modulation frequency, and wherein a phase difference between said first modulated beam and said second modulated beam is approximately 180 degrees;providing a reference sample substantially free of analyte;directing and substantially overlapping said first modulated beam and said second modulated beams onto said reference sample;measuring a first photothermal signal produced by said first modulated beam;measuring a second photothermal signal produced by said second modulated beam;obtaining a relative photothermal phase between said first photothermal signal and said second photothermal signal;controlling a relative intensity of said first modulated beam and said second modulated beam such that a ratio of said first photothermal signal amplitude and said second photothermal signal amplitude is approximately unity; andcontrolling a relative phase of said first modulated beam and said second modulated beam so that said relative photothermal phase is approximately 180 degrees. 2. The method according to claim 1, further comprising the steps of: measuring a differential photothermal signal for one or more standards, each said standard comprising a known analyte concentration, while controlling a relative intensity of said first modulated beam and said second modulated beam to scan a photothermal amplitude ratio of said first photothermal signal and said second photothermal signal near unity; andselecting said photothermal amplitude ratio corresponding to a desired sensitivity over said concentration range of interest. 3. The method according to claim 2 further comprising the step of generating calibration data for said photothermal amplitude ratio corresponding to said desired sensitivity over said concentration range of interest. 4. The method according to claim 3 further comprising measuring a sample comprising an unknown analyte concentration within said substance, and determining said unknown analyte concentration by relating a differential photothermal s... | - | A61B005/1455
US20110137145 | A1 | 2011-06-09 | METHOD FOR COMBINING HIGHLY PERMEABLE PARTS OF A MAGNETIC SHIELD | 2010-12-13 | An inventive magnetic filling material for bridging interspaces in magnetic shields for shielding static or low-frequency magnetic fields. The magnetic filling material comprises at least one matrix material and at least one magnetic component embedded into the matrix material. The magnetic component has magnetically shielding properties. | 1. A magnetic filling material for bridging interspaces in magnetic shields and for shielding static or low-frequency magnetic fields, comprising: at least one matrix material; andat least one magnetic component embedded in the at least one matrix material, wherein the at least one magnetic component has magnetically shielding properties and comprises a material having a permeability of at least 100. 2. The magnetic filling material of claim 1, wherein the at least one matrix material has substantially diamagnetic properties. 3. The magnetic filling material of claim 1, wherein the at least one matrix material has a deformable state and a cured state, wherein the at least one matrix material is liquid or has at least partly plastic and/or elastic properties in the deformable state, and wherein the at least one matrix material is substantially rigid in the cured state. 4. The magnetic filling material of claim 1, wherein the at least one matrix material comprises an adhesive. 5. The magnetic filling material of claim 1, wherein the at least one matrix material comprises at a resin, a multi-component adhesive or at least one component of a multi-component adhesive. 6. The magnetic filling material of claim 1, wherein the at least one magnetic component comprises at least one ferromagnetic material having a permeability of at least 300. 7. The magnetic filling material of claim 1, wherein the at least one magnetic component comprises at least one ferromagnetic material having a permeability of at least 1000. 8. The magnetic filling material of claim 1, wherein the at least one magnetic component comprises a metallic material. 9. The magnetic filling material of claim 1, wherein the at least one magnetic component comprises at least one of the following materials: iron; nickel; an alloy comprising the elements iron and nickel comprising a nickel fraction of between 60% and 90%. 10. The magnetic filling material of claim 1, wherein the at least one magnetic component comprises an alloy comprising the elements iron and nickel and having a nickel fraction of between 75% and 80%. 11. The magnetic filling material of claim 1, wherein the at least one magnetic component comprises a powder or particles. 12. The magnetic filling material of claim 11, wherein the at least one magnetic component comprises particles having an average particle size of less than 500 micrometers. 13. The magnetic filling material of claim 12, wherein the particles have an average particle size of less than 90 micrometers. 14. The magnetic filling material of claim 12, wherein the particles have an average particle size of about 60 micrometers. 15. The magnetic filling material of claim 1, wherein the at least one magnetic component comprises a mass fraction of 5 to 50 weight percent of the magnetic filling material. 16. The magnetic filling material of claim 15, wherein the at least one magnetic component comprises a mass fraction of about 15 weight percent of the m... | - | A61B005/05 H01F007/00 B23P011/00
US20110152709 | A1 | 2011-06-23 | MOBILE BODY CONTROL DEVICE AND MOBILE BODY CONTROL METHOD | 2009-09-18 | A mobile body control device includes: a brain activity detecting unit that detects brain activity information of a user; a brain signal separating unit that separates an artifact component from the brain activity information detected by the brain activity detecting unit; a control signal generating unit that slides a sampling period for extracting brain data, at predetermined intervals in an overlapped manner in the brain activity information from which the artifact component is separated, successively calculates feature values for the brain data within each of the sampling periods obtained by sliding, and generates a control signal based on the feature values calculated; and a drive control unit that drives and controls a mobile body with a user riding thereon, based on the control signal generated. | 1. A mobile body control device comprising: a brain activity detecting unit that detects brain activity information of a user;a brain signal separating unit that separates an artifact component from the brain activity information detected by the brain activity detecting unit;a control signal generating unit for sliding a sampling period for extracting brain data, at predetermined intervals in an overlapped manner in the brain activity information from which the artifact component is separated by the brain signal separating unit, successively calculating feature values for the brain data within each of the sampling periods obtained by sliding, and generating a control signal based on the feature values calculated; anda drive control unit that drives and controls a mobile body with a user riding thereon, based on the control signal generated by the control signal generating unit. 2. The mobile body control device according to claim 1, further comprising a teacher signal generating unit that generates a teacher signal, wherein the brain signal separating unit performs learning using the teacher signal generated by the teacher signal generating unit, and separates the artifact component depending on a user. 3. The mobile body control device according to claim 1, further comprising a teacher signal generating unit that generates a teacher signal, wherein the control signal generating unit calculates a correlation between the feature values and the control signal by using the teacher signal generated by the teacher signal generating unit, and generates the control signal based on the feature values and the correlation calculated. 4. The mobile body control device according to claim 1, further comprising: a myoelectric potential detecting unit that detects a myoelectric potential of a user; anda stop determining unit that determines whether or not to stop the mobile body based on the myoelectric potential detected by the myoelectric potential detecting unit,wherein when the stop determining unit determines to stop the mobile body, the drive control unit controls the mobile body to be stopped. 5. The mobile body control device according to claim 1, wherein when the drive control unit successively receives a same control signal from the control signal generating unit a predetermined number of times or more, the drive control unit executes a control corresponding to the control signal. 6. The mobile body control device according to claim 1, wherein the brain activity detecting unit includes a plurality of sensors that detect a brain wave signal of a user, andthe control signal generating unit slides a sampling period for extracting brain data, at predetermined intervals in an overlapped manner in a plurality of the brain wave signals which are detected by the plurality of sensors and from which the artifact component is separated, successively calculates feature values for the brain data within each of the sampling periods obtained by sliding, and gen... | TOYOTA JIDOSHA KABUSHIKI KAISHA RIKEN | A61B005/0476
US09084820 | B2 | 2015-07-21 | Magnetic microstructures for magnetic resonance imaging | 2009-04-20 | A magnetic resonance contrast agent has a medium, and a contrast structure dispersed in the medium. The contrast structure comprises a magnetic material arranged to create a local region of a local magnetic field such that nuclear magnetic moments of a material when arranged within the local region precess at a characteristic Larmor frequency about a total magnetic field in the local region while in use, the characteristic Larmor frequency being identifiable with the contrast structure, and the total magnetic field in the local region being a substantially spatially uniform magnetic field. | 1. A magnetic resonance contrast agent, comprising: a medium; anda contrast structure dispersed in said medium,wherein said contrast structure comprises a magnetic material arranged as a pair of solid disk-shaped magnetic components defining a fixed space between a circular surface of each disk shape, wherein the centers of the circular surfaces of the disk shapes are arranged along a same axis normal to the circular surfaces, and wherein the circular surfaces are parallel to each other,wherein the fixed space creates a spatially extended region contained within a near-field region of the contrast structure over which an applied magnetic field results in a homogenous field, such that nuclear magnetic moments of a second material when arranged within said spatially extended region precess at a characteristic Larmor frequency, whereby the magnetic resonance contrast agent is adapted to emit a characteristic magnetic resonance signal of the magnetic material. 2. The magnetic resonance contrast agent according to claim 1, wherein a total magnetic field in said spatially extended region of said contrast structure is a local magnetic field created by said contrast structure. 3. The magnetic resonance contrast agent according to claim 1, wherein a total magnetic field in said spatially extended region of said contrast structure is a combination of a local magnetic field created by said contrast structure and a portion of a background magnetic field. 4. The magnetic resonance contrast agent according to claim 1, wherein said contrast structure has a structure that permits a fluid to at least one of flow through or diffuse through at least a portion of said spatially extended region of a local magnetic field created by said contrast structure. 5. The magnetic resonance contrast agent according to claim 4, wherein said second material when arranged within said spatially extended region is at least a portion of said fluid while said magnetic resonance contrast agent is in use. 6. The magnetic resonance contrast agent according to claim 1, wherein said second material arranged within said spatially extended region is at least one of a portion of said contrast structure or contained within said contrast structure. 7. The magnetic resonance contrast agent according to claim 1, wherein said magnetic resonance contrast agent provides a frequency-shifted nuclear magnetic resonance signal when used with at least one of a magnetic resonance imaging system or a nuclear magnetic resonance system. 8. The magnetic resonance contrast agent according to claim 1, wherein said contrast structure has a maximum dimension of less than about 5 mm. 9. The magnetic resonance contrast agent according to claim 1, wherein said contrast structure has a maximum dimension of at least about 10 nm and less than about 100 μm. 10. The magnetic resonance contrast agent according to claim 9, wherein said contrast structure has a maximum dimension of at least about 50 nm and less th... | The United States of America, as Represented by the Secretary, Department of Health and Human Services Office of Technology Transfer, National Institutes of Health The United States of America, Represented by the Secretary of Commerce | A61K049/18 A61B005/055 G01R033/12 G01R033/28 G01R033/56 G01R033/563 H01F010/14 G01R033/34
US08577433 | B2 | 2013-11-05 | Medical device alarm modeling | 2010-11-18 | Embodiments of the present disclosure relate to patient monitors with alarm modeling features that may be employed to set alarm limits. According to certain embodiments, the patient monitors may include a user interface for setting alarm limits that may be displayed on the patient monitors and/or on an external device, such as a central monitoring station. The user interface may allow a user to vary alarm limit settings and view how the settings change the alarm history for a representative data trend. | 1. A patient monitor, comprising: a medical device interface suitable for operable connection to a sensor;a display configured to display patient physiological data based on input received from the sensor and to display a graphical representation of representative physiological data;a graphical user interface configured to receive a user input that alters existing alarm limits for the patient physiological data; andan alarm modeling feature configured to analyze the representative physiological data to identify alarm conditions corresponding to the altered alarm limits and configured to display indicators that identify the alarm conditions associated with the altered alarm limits and that identify the alarm conditions associated with the previously existing alarm limits that no longer represent alarm conditions on the graphical representation. 2. The patient monitor of claim 1, wherein the patient physiological data and the representative physiological data comprise SpO2 values and wherein the graphical representation comprises a plethysmographic waveform. 3. The patient monitor of claim 1, wherein the representative physiological data comprises a historical trend of the patient physiological data. 4. The patient monitor of claim 1, wherein the representative physiological data comprises sample physiological data included within the alarm modeling feature. 5. The patient monitor of claim 1, wherein the graphical user interface includes a screen configured to display the graphical representation, the existing alarm limits, and highlighted sections of the graphical representation identifying alarm events corresponding to the existing alarm limits. 6. The patient monitor of claim 1, wherein the display comprises a touch screen and wherein the indicators comprise touch sensitive graphical elements selectable by a user to alter the alarm limits. 7. The patient monitor of claim 1, wherein the alarm limits comprise SpO2 limits, an oxygen saturation pattern detection threshold, or a value specifying a product of magnitude and duration of an oxygen desaturation event, or a combination thereof. 8. The patient monitor of claim 1, wherein the indicators comprise graphical elements selected by a user to adjust the alarm conditions for a corresponding portion of the graphical representation. 9. The patient monitor of claim 8, wherein the alarm model feature is configured to determine new alarm limits corresponding to the adjusted alarm conditions. 10. A patient monitoring system, comprising: one or more patient monitors; a central monitoring station coupled to the patient monitors and comprising: a display configured to display a graphical representation simulating patient physiological data received from the patient monitors; and a graphical user interface comprising alarm indicators configured to be displayed with the graphical representation to indicate alarm conditions for portions of the graphical representation, wherein each alarm indicator is d... | Covidien LP | A61B005/1455
US08747311 | B2 | 2014-06-10 | Method and apparatus for evaluating the effects of internal and external stress influences | 2010-11-18 | A measuring device and method is for enabling a user to measure and record changes in emotion during the day and translate it into a relative stress value per day. It may be used to accumulate the positive and negative emotional reactions that a user experiences during a predetermined time period. This method and device will provide a record of the associated distress and eustress and the cumulative effect therefrom on the individual. The user has a visual display of whether or not the emotions felt during a given day are positive or negative. The recognition of the stress level may remind or enable the user to take some action that will counterbalance excessive negative emotions known to contribute to major diseases and engage in activities that reduce or relieve the stress influences on the user before the end of each day. | 1. A computer-implemented method for evaluating a cumulative level of distress and eustress on an individual, the method implemented on a measuring device by software instructions contained in a memory coupled to a computer processor and including a user interface having a display means and an input means, comprising: providing a negative and a positive graded scale on the measuring device, the scale being moveable with respect to a cursor displayed on the measuring device;determining a time interval for recording emotional responses by the individual to life events and entering the time interval into the measuring device by the input means;judging and assigning positive and negative numeric measures of emotional responses perceived by the individual to each instance of life events during the determined time interval and entering the positive and negative numeric measures into the measuring device by the input means;accumulating the numeric measures of the emotional responses perceived by the individual during the determined time interval by the measuring device;displaying and evaluating the level of distress and eustress on the individual during the determined time interval based on the accumulated numeric measures of emotional responses perceived during the determined time interval by the individual; andusing the displayed level of distress and eustress as a reference to change and balance emotional reactions to life events by the individual. 2. The method of claim 1, further comprising evaluating long term trends over time of cumulative distress and eustress on the individual by comparing changes in the level of distress and eustress between successive time intervals. 3. The method of claim 1, wherein the determined time interval is selected from the group consisting of a day, a week and a month. 4. The method of claim 1, further comprising assigning positive numeric values based on an intensity of the positive measures of emotional responses and assigning negative numeric values based on an intensity of the negative measures of emotional responses. 5. The method of claim 4, further comprising: providing the measuring device with a displaying the positive and negative graded scale;setting the measuring device to a value of zero at the beginning of the time interval;using the input means, entering the assigned positive and negative numeric values based on the intensity of the measures of emotional responses into the measuring device during the time interval; andevaluating the level of distress and eustress on the individual during the time interval based on the final summed numeric values in the measuring device of the intensity of the measures of emotional responses perceived during the time interval by the individual. 6. The method of claim 5, wherein the measuring device is selected from the group consisting of an electronic measuring device, a personal digital assistant, a computer having non-transitory computer readable medium compris... | Make3Wishes, LLC | A61B005/00 G09B003/00 G09B001/30 G09B001/18
US08491504 | B2 | 2013-07-23 | Devices and methods for monitoring sit to stand transfers | 2010-08-04 | In certain embodiments of the present disclosure, a device for monitoring sit to stand transfers is provided. The device includes an electromechanical sensor and a computer, wherein the electromechanical sensor is configured to signal sit to stand transfers to the computer and the computer is configured to determine if the sit to stand transfers deviate from a normal sit to stand transfer pattern. | 1. A device for monitoring sit to stand transfers comprising: an electromechanical sensor and a computer,wherein the electromechanical sensor is configured to signal a number of sit to stand transfers to the computer and the computer is configured to determine if the number of sit to stand transfers deviate from a normal sit to stand transfer pattern and wherein the computer is configured to produce an alert if the number of sit to stand transfers deviate from a normal sit to stand transfer pattern. 2. The device of claim 1, wherein the electromechanical sensor comprises a film. 3. The device of claim 2, wherein the film comprises a polyolefin. 4. The device of claim 1, wherein the electromechanical sensor comprises a load cell. 5. The device of claim 1, wherein the alert is configured to be transmitted wirelessly. 6. The device of claim 1, further comprising an alert station, the alert station configured to receive the alert from the computer. 7. The device of claim 1, wherein the alert is configured to be transmitted via the internet. 8. The device of claim 1, wherein the device is configured to be positioned at a base of a seating surface so that the electromechanical sensor can signal sit to stand transfers. 9. The device of claim 1, wherein the device further comprises an RFID tag, the RFID tag configured to be placed on a user, the computer being capable of associating the sit to stand transfers with the user based upon the RFID tag. 10. The device of claim 1, wherein the device further comprises a plurality of RFID tags, the RFID tags configured to be placed on different users, the computer being capable of associating the sit to stand transfers with each different user based upon the RFID tag. 11. The device of claim 1, wherein the device is capable of signaling sit to stand transfers from a plurality of different seating surfaces. 12. The device of claim 1, wherein the electromechanical sensor and computer are not connected to one another by wire. 13. The device of claim 1 wherein the electromechanical sensor and computer are connected to one another by wire. 14. A method for monitoring sit to stand transfers comprising: utilizing a device to signal a sit to stand transfer, the device comprising an electromechanical sensor and a computer, the electromechanical sensor signaling a number of sit to stand transfers to the computer and the computer determining if the number of sit to stand transfers deviate from a normal sit to stand transfer pattern, the computer producing an alert when the number of sit to stand transfers deviate from a normal sit to stand transfer pattern. 15. The method of claim 14, wherein the electromechanical sensor comprises a film. 16. The method of claim 14, further comprising seeking medical attention for the user in response to a deviation from a normal sit to stand transfer pattern. 17. The method of claim 14, wherein the electromechanical sensor comprises a load cell. 18. The method of claim 14,... | University of South Carolina | A61B005/11
US08055722 | B2 | 2011-11-08 | Notification control through brain monitoring of end user concentration | 2009-12-02 | Embodiments of the present invention provide a method, system and computer program product for message notification management through brain monitoring of end user concentration. In an embodiment of the invention, a method for message notification management through brain monitoring of end user concentration can include receiving neurofeedback for an end user through an interface for a brain monitoring system, comparing the neurofeedback to a threshold level of brain activity indicating a degree of concentration of the end user, and suppressing message notifications for messages in a messaging system responsive to the neurofeedback exceeding the threshold level of brain activity, but otherwise permitting message notifications for messages in the messaging system. | 1. A method for message notification management through brain monitoring of end user concentration, the method comprising: receiving neurofeedback for an end user through an interface for a brain monitoring system;comparing the neurofeedback to a threshold level of brain activity indicating a degree of concentration of the end user; and,suppressing message notifications for messages in a messaging system responsive to the neurofeedback exceeding the threshold level of brain activity, but otherwise permitting message notifications for messages in the messaging system. 2. The method of claim 1, wherein the brain activity monitoring system is an electroencephalography (EEG) based brain activity monitoring system. 3. The method of claim 1, wherein the messaging system is an e-mail messaging system. 4. The method of claim 1, wherein the messaging system is an instant messaging system. 5. The method of claim 1, wherein the messaging system is a text messaging system. 6. The method of claim 1, wherein the messaging system is a voice mail messaging system. 7. A messaging data processing system comprising: a messaging system executing in memory by a processor of a host computing device;a brain activity monitoring system coupled to the messaging system; and,message suppression logic coupled to the messaging system and the brain activity monitoring system, the logic comprising program code that when executed by the processor compares neurofeedback received from the brain activity monitoring system to a threshold level of brain activity indicating a degree of concentration and suppresses message notifications for messages in a messaging system responsive to the neurofeedback exceeding the threshold level of brain activity, but otherwise permits message notifications for messages in the messaging system. 8. The system of claim 7, wherein the brain activity monitoring system is an electroencephalography (EEG) based brain activity monitoring system. 9. The system of claim 7, wherein the messaging system is an e-mail messaging system. 10. The system of claim 7, wherein the messaging system is an instant messaging system. 11. The system of claim 7, wherein the messaging system is a text messaging system. 12. The system of claim 7, wherein the messaging system is a voice mail messaging system. 13. The system of claim 7, wherein the messaging system is an incoming call alert system. 14. A computer program product comprising a computer usable storage medium storing computer usable program code for message notification management through brain monitoring of end user concentration, the computer program product comprising: computer usable program code for receiving neurofeedback for an end user through an interface for a brain monitoring system;computer usable program code for comparing the neurofeedback to a threshold level of brain activity indicating a degree of concentration of the end user; and,computer usable program code for suppressing message noti... | International Business Machines Corporation | G06F015/16
US08172768 | B2 | 2012-05-08 | Systems and methods for intra-operative stimulation | 2011-01-26 | Improved assemblies, systems, and methods provide safeguarding against tissue injury during surgical procedures and/or identify nerve damage occurring prior to surgery and/or verify range of motion or attributes of muscle contraction during reconstructive surgery. A stimulation control device may incorporate a range of low and high intensity stimulation to provide a stimulation and evaluation of both nerves and muscles. A stimulation control device is removably coupled to a surgical device or is imbedded within the medical device to provide a stimulation and treatment medical device. A disposable hand held stimulation system includes an operative element extending from the housing, the housing includes a visual indication to provide feedback or status to the user. | 1. A hand-held tissue stimulation system comprising: a tubular housing having a proximal end and a distal end, the housing being sized and configured to be held in and controlled by a single human hand,a stimulation control device positioned entirely within the housing, the stimulation control device comprising stimulation signal generating circuitry for generation of an electrical stimulation signal having an amplitude and a duration, a microcontroller electrically coupled to the stimulation signal generating circuitry, and a source of power for the stimulation control device,a probe extending from the housing and having an electrically conductive surface coupled to the signal generating circuitry, the electrically conductive surface being sized and configured for contact with a targeted tissue region and, when in contact, applying the electrical stimulation signal to the targeted tissue region,a return electrode electrically coupled to the signal generating circuitry and being sized and configured for contact with tissue to provide an electrical flow path for the stimulation signal from the probe back to the stimulation control device,a first control device carried on the housing and electrically coupled to the microcontroller, the first control device being movable by a user from a first position to a second position for providing a power control input to the microcontroller to turn the stimulation signal generating circuitry on only for a predetermined total single use period,a second control device carried on the housing separate from the first control device and being electrically coupled to the microcontroller, the second control device comprising a potentiometer operable by a user during the total single use period independent of the first control device within a range of control positions for providing duration control inputs to the microcontroller to control adjustment of the duration from a predefined duration range, anda visual indicator carried on the housing and being electrically coupled to the microcontroller for confirming status conditions during the total single use period including visual indications that vary according to different status conditions, the microcontroller configured to cause the visual indictor to display: (i) a first predetermined color and/or flash rate of light confirming, independent of contact between the probe and the targeted tissue region, that power to the stimulation signal generating circuitry is on due to operation of the first control device, (ii) a second predetermined color and/or flash rate of light different than the first predetermined color and/or flash rate of light confirming, when there is contact between the probe and the targeted tissue region, the delivery of the stimulation signal to the targeted tissue region through the probe and back to the stimulation control device through the return electrode due to operation of the first and/or second control device and (iii) a third predetermin... | Checkpoint Surgical, LLC | A61B005/06 A61N001/00
US07935062 | B2 | 2011-05-03 | Derivation of flow contour from pressure waveform | 2005-01-27 | The present invention provides a system and method for estimating a blood flow waveform contour from a pressure signal. An arterial or ventricular pressure signal is acquired from a pressure sensor. Landmark points are identified on the pressure waveform that correspond to features of a flow waveform. In one embodiment, the landmark pressure waveform points correspond to the onset of flow, the peak flow, and the end of the systolic ejection phase. The landmark pressure waveform points define a contour that approximates the flow contour. Beat-by-beat flow contour estimation can be performed to allow computation of flow-related hemodynamic parameters such as stroke volume or cardiac output for use in patient monitoring and/or therapy management. | 1. A method of hemodynamic monitoring in an implantable medical device, comprising: sensing a blood pressure signal using the implantable medical device;extracting a plurality of landmark points on the blood pressure signal corresponding in time to blood flow events using the implantable medical device; andderiving a flow contour estimate corresponding to the extracted points using the implantable medical device, wherein the blood pressure signal comprises a ventricular pressure signal. 2. The method of claim 1 wherein extracting a plurality of landmark points comprises identifying a point on the ventricular pressure signal corresponding in time to a maximum peak of a first time derivative of the pressure signal. 3. The method of claim 1 wherein extracting a plurality of landmark points comprises identifying a point on the ventricular pressure signal corresponding in time to a first peak of a third time derivative of the pressure signal. 4. The method of claim 1 wherein extracting a plurality of landmark points comprises identifying a point on the ventricular pressure signal corresponding in time to a minimum peak of the first time derivative of the pressure signal. 5. A method of hemodynamic monitoring in an implantable medical device, comprising: sensing a blood pressure signal using the implantable medical device;extracting a plurality of landmark points on the blood pressure signal corresponding in time to blood flow events using the implantable medical device; andderiving a flow contour estimate corresponding to the extracted points using the implantable medical device, wherein deriving the flow contour estimate comprises deriving a triangle defined by the extracted points. 6. A method of hemodynamic monitoring in an implantable medical device, comprising: sensing a blood pressure signal using the implantable medical device;extracting a plurality of landmark points on the blood pressure signal corresponding in time to blood flow events using the implantable medical device;deriving a flow contour estimate corresponding to the extracted points using the implantable medical device, wherein the flow contour estimate is an approximation of a flow waveform corresponding to the blood pressure signal; andcomputing a stroke volume using the derived flow contour estimate. 7. A system, comprising: means for sensing a blood pressure signal;means for extracting a plurality of landmark points on the blood pressure signal corresponding in time to blood flow events; andmeans for deriving a flow contour estimate corresponding to the extracted points, wherein the blood pressure signal comprises a ventricular pressure signal, wherein the means for extracting a plurality of landmark points comprises means for identifying a point on the ventricular pressure signal corresponding in time to a maximum peak of a first time derivative of the pressure signal. 8. The method of claim 6, wherein the stroke volume is computed as an area defined by the flow contour ... | Medtronic, Inc. | A61B005/02
US07960605 | B2 | 2011-06-14 | Methods for testing for caloric restriction (CR) mimetics | 2006-03-17 | Methods for treating neurological diseases and for testing Caloric Restriction (CR) mimetics or CR mimetic candidates. In one exemplary method, a CR mimetic candidate is administered to a transgenic animal and the effects of the administering are determined; the transgenic animal includes an added gene from another type of animal or a modified gene which is designed to produce a disease or ailment of the another type of animal, and the method seeks to determine whether the CR mimetic candidate improves the disease or ailment. Methods relating to neurological disease and other methods relating to CR mimetic testing are also described. | 1. A method of testing a Caloric Restriction (CR) mimetic or a CR mimetic candidate, the method comprising: selecting at least one of the CR mimetic or the CR mimetic candidate comprising a plant extract;administering the selected at least one of the CR mimetic or the CR mimetic candidate;determining antioxidant capabilities of the at least one of the CR mimetic or the CR mimetic candidate, wherein the determining includes determining: (i) a half maximal inhibitory concentration (IC50) of the CR mimetic or the CR mimetic candidate having a dose dependent scavenging effect on a hydroxyl radical is consistent with a scavenging effect of a radical scavenger made by combining vitamin C and vitamin E, (ii) an IC50 of the CR mimetic or the CR mimetic candidate having a dose dependent scavenging effect on a superoxide radical is consistent with a scavenging effect of a known superoxide radical scavenger and (iii) an IC50 of the CR mimetic or the CR mimetic candidate having a dose dependent scavenging effect on a lipid radical is consistent with a scavenging effect of a known lipid radical scavenger; anddetermining in a subject a level of a biological parameter of a known CR marker by comparing the level of the known CR marker to measurements of a corresponding CR marker from a subject having been administered at least one of the CR mimetic or the CR mimetic candidate, the biological parameter of the known CR marker includes an RNA transcript level of a gene known to be effected in gene expression by a CR diet or a known CR mimetic. 2. A method as in claim 1 further comprising: comparing the determined antioxidant capabilities to known antioxidants. 3. A method as in claim 1 wherein antioxidant capabilities are determined for only the CR mimetic candidate and wherein the known antioxidants comprise at least one of vitamin C, vitamin E, Acetyl-L-Carnitine, and α-lipoic acid. 4. A method as in claim 3 wherein the comparing of levels of biological parameters is part of a battery of screening tests designed to identify CR mimetic compounds and the screening tests include the determining of antioxidant capabilities. 5. The method as in claim 1 wherein determining the scavenging effect of the CR mimetic or the CR mimetic candidate on the hydroxyl radical, on the superoxide radical, and on the lipid radical comprises detecting the effect of the CR mimetic or the CR mimetic candidate on these radicals with a spectrometer. | BioMaker Pharmaceuticals, Inc. | A01K067/027 C12Q001/26 A61B005/00
US07959579 | B2 | 2011-06-14 | Apparatus for medical measurement | 2007-01-17 | A measuring device for medical applications in at least one example embodiment includes a first part provided with measuring means for measuring data of a patient, and a second part provided with communication means for sending the data. In at least one example embodiment, the second part is releasably connected to the first part. | 1. A cassette comprising: a housing including a bottom part, anda lid,said bottom part and said lid form an enclosure;at least one probe connector connected to said housing;for at least one of said at least one probe connector, at least one sensor in communication with said respective probe connector;a connector in communication with said at least one sensor; anda circuit located in said enclosure and in electrical communication between said at least one sensor and said connector, said circuit includes a timing circuit having a counter that prevents transmission of monitoring data being sent through said connector after a predetermined time has passed after connection to a communications module. 2. The cassette according to claim 1, wherein said circuit includes electrical components for controlling operation of said at least one sensor. 3. The cassette according to claim 1, wherein at least one of said at least one probe connector includes a catheter connector,a pressure chamber, andsaid catheter connector and said pressure chamber together have a passageway running from said catheter connector into said pressure chamber. 4. The cassette according to claim 3, wherein said sensor connected to said pressure chamber includes a fluid pressure sensor. 5. The cassette according to claim 3, wherein said pressure chamber includes at least one pressure relief opening having a one-way valve positioned such that fluid can be introduced past said valve into said pressure chamber, wherein when the pressure relief valve is brought into an opened state the pressure chamber is in free communication with an environment around said cassette. 6. The cassette according to claim 3, further comprising at least one catheter connected to said at least one catheter connector. 7. The cassette according to claim 3, further comprising at least one one way valve where said one way valve is connected to a pressure chamber. 8. The cassette according to claim 3, further comprising a cover over at least one pressure chamber. 9. The cassette according to claim 1, further comprising at least two electrode connectors on an outside of said housing, said electrode connectors electrically connected to said connector. 10. The cassette according to claim 1, further comprising a plate, said connector extending from said plate, said plate having at least one connecting pin configured to be attached to at least one module. 11. A cassette comprising: a housing including a bottom part, anda lid having a upper part and a lower part in interlocking fit such that said lower part is pivotable about said upper part to provide access to an area enclosed by said bottom part and said lid,said bottom part and said lid form an enclosure;at least one probe connector connected to said housing;for at least one of said at least one probe connector, at least one sensor in communication with said respective probe connector;a connector in communication with said at least one sensor, said connector... | Laborie Medical Technologies, Inc. | A61B005/00 A61B005/02 A61B005/117 A61B008/14
US07937140 | B2 | 2011-05-03 | Detection and diagnostic system and method | 2007-06-12 | A method and apparatus for detecting the presence of substances, compositions, constituents, proportionalities, of examined objects, and abnormalities and diseases associated with human tissue responsive to detected amplitudes and/or frequencies, such as resonant frequencies from the object/body being examined. Data in a database is utilized to identify the unbalanced condition, the foreign, toxic or harmful substance. The identified condition/substance may be used as an aid to select an appropriate treatment or corrective action. | 1. A method of examination, comprising: a) providing a database having categories of data associating characteristics of electrical signals with various substances, said database further containing data relating lack of one of a balanced state or harmonized state or a condition of an organ due to substances;b) scanning at least a portion of a body with an energy source;c) detecting signals from an organ of the scanned body portion;d) ascertaining presence of at least one characteristic from the detected signals;e) comparing the characteristic ascertained with data in at least one of said data categories in said database to ascertain presence of one of said substance; andf) employing the database for determining a lack of one of a balanced state or harmonized state or a condition of an organ due to the detected substance. 2. The method of claim 1 wherein step (b) further comprises: g) exposing said body portion to a selected one of sonic, electromagnetic, infra-red, light and laser energy. 3. The method of claim 1 wherein step (g) further comprises: varying an amplitude of the energy source over a given range. 4. The method of claim 1 wherein step (g) further comprises: varying a frequency of the energy source over a given range. 5. The method of claim 1 wherein step (g) further comprises: varying an amplitude and a frequency of the energy source over given respective ranges. 6. The method of claim 1 further comprising: selecting a treatment based on said second database responsive to step (b). 7. The method of claim 1, step (c) further comprising: employing a sensor to detect energy from said object. 8. The method of claim 7, step (c) further comprising: determining a characteristic of the received energy detected by said sensor. 9. A method of examination, comprising: a) creating a database having categories of data associating characteristics of electrical signals with various substances, said database further containing data relating lack of one of a balanced state or harmonized state or a condition of an organ due to substances;b) scanning at least a portion of an object with an energy source;c) detecting energy from an organ of the scanned portion;d) ascertaining at least one characteristic of the detected signal;e) comparing the characteristic ascertained with data from said database to identify at least one of said substances associated with the characteristic of said detected signal; andf) determining a lack of one of a balanced state or harmonized state of the organ due to the detected substance. 10. The method of claim 9 wherein step (b) further comprises: g) exposing said portion to a selected one of sonic, electromagnetic, infra-red, light and laser energy. 11. The method of claim 9 wherein step (g) further comprises: varying an amplitude of the energy source over a given range during scanning. 12. The method of claim 9 wherein step (g) further comprises: varying a frequency of the energy source over a given range during sc... | - | A61B005/05
US10154790 | B2 | 2018-12-18 | Method and system for monitoring sleep | 2008-08-11 | A method of monitoring sleep comprises simultaneously recording a person's electrocardiogram (ECG) and photoplethysmogram (PPG), deriving a plurality of parameters from the recorded data, and providing an output indicative of a sleep characteristic based upon an analysis of the parameters. The ECG and PPG may be recorded using an apparatus which is a combination of a Holter monitor and a pulse oximeter, which is wearable in ambulatory manner. | 1. A method of monitoring sleep stages or level of sleepiness, the method comprising using one or more processors to: acquire and record data of a person's electrocardiogram (ECG) and photoplethysmogram (PPG), the data being acquired and recorded simultaneously,derive a plurality of parameters from both the recorded ECG and PPG data, wherein the plurality of parameters includes (1) at least one parameter derived from the ECG data independently of the PPG data, and (2) at least one parameter derived from the PPG data independently of the ECG data,extract features from the plurality of parameters, andprovide an output indicative of a sleep stage or a quantitative level of sleepiness based upon an analysis of at least a subset of the extracted features from the derived plurality of parameters, wherein a trained statistical classifier in a processor of the one or more processors evaluates, on an epoch-by-epoch basis, the extracted features from both (1) the at least one parameter derived from the ECG data and (2) the at least one parameter derived from the PPG data, to generate the output indicative of the sleep stage or the quantitative level of sleepiness. 2. The method of claim 1, wherein the ECG and PPG data is acquired in an ambulatory manner. 3. The method of claim 2, wherein the ECG and PPG data is acquired using a combination of a Holter monitor and a pulse oximeter. 4. The method of claim 1, wherein the plurality of parameters include at least one of RR interval, pulse amplitude and respiratory rate. 5. The method of claim 4, wherein the plurality of parameters are analyzed by extracting the features, wherein the features are suitable for classifying sleep stages or quantitative level of sleepiness, and by feeding a subset of the features into a quantitative classifier to obtain an estimate of the sleep stage or quantitative level of sleepiness. 6. The method of claim 5, wherein the features include at least one of PPG spectral spread, PPG pulse amplitude, respiratory rate variability, detrended fluctuation analysis, time and frequency analysis and multi-scale entropy. 7. The method of claim 5, wherein the output is used to determine sleep stages comprising at least one of wakefulness, rapid-eye-movement sleep, and non-rapid-eye-movement sleep. 8. The method of claim 5, comprising assigning sleep stages or quantitative levels of sleepiness to specific epochs of time of greater than 10 seconds in duration. 9. The method of claim 8, comprising classifying a single epoch according to a classification of surrounding epochs. 10. The method of claim 5, further comprising estimating the output of a psychomotor vigilance test during wakefulness using the output of the quantitative classifier. 11. The method of claim 1, wherein the one or more processors derive an additional parameter from both the recorded ECG data and the recorded PPG data, and wherein the output indicative of a sleep stage or a quantitative level of sleepiness is based u... | University College Dublin, National University of Ireland | A61B005/00 A61B005/024 A61B005/04
US09797973 | B2 | 2017-10-24 | MR imaging with extended field of view | 2009-06-18 | Magnetic resonance examination system comprises displaceable carrier for supporting an object to be examined. The carrier can be moved over a two dimensional area. The magnetic resonance examination system is configured to acquire sets of magnetic resonance signals from the object for various positions of the carrier in the two dimensional area. | 1. A magnetic resonance examination system, comprising: an acquisition system configured to acquire magnetic resonance signals;a displaceable carrier configured to support an object to be examined; anda controller configured to control the acquisition system and the carrier, and further configured to: move the carrier along a continuous motion direction in a two dimensional area,shift the carrier along a direction transverse to the continuous motion direction to respective lateral positions, andfor respective lateral positions, acquire a set of magnetic resonance signals during continuous motion of the carrier along the continuous motion direction. 2. The magnetic resonance examination system of claim 1 and further having two magnetic poles between which a main magnetic field extends. 3. The magnetic resonance examination system of claim 1, wherein the controller is configured to apply oversampling of the acquired magnetic resonance signals in the direction transverse to the continuous motion direction of the carrier. 4. The magnetic resonance examination system of claim 1, wherein the controller is further configured to apply oversampling along at least one phase encoding direction. 5. The system of claim 1, wherein each set of magnetic resonance signals corresponds to a field of view, and wherein the controller is configured to acquire magnetic resonance signals corresponding to a plurality of fields of view including at least a first field of view and a second field of view which are displaced with respect to each other in the continuous motion direction, and at least a third field of view and a fourth field of view which are displaced with respect to each other in the direction transverse to the continuous motion direction. 6. The magnetic resonance examination system of claim 1, wherein the continuous motion direction is a first linear direction and the direction transverse to the continuous motion direction is a second linear direction orthogonal to the first linear direction, wherein the object to be examined has a width which extends in the second linear direction, and wherein the controller is configured to move the carrier, for each of the respective lateral positions, linearly along the first linear direction to acquire the set of magnetic resonance signals which span less than the entire width of the object to be examined. 7. The magnetic resonance examination system of claim 6, wherein the sets of magnetic resonance signals acquired for all of the respective lateral positions together span the entire width along the second linear direction of the examination area be imaged. 8. A method of operating a magnetic resonance examination system which employs a carrier configured to support an object to be examined, the method comprising: acquiring a plurality of sets of magnetic resonance signals from the object for various positions of the carrier in two dimensions, wherein at least a first one of the positions and a second one of t... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/00 G01R033/563
US20110160612 | A1 | 2011-06-30 | PAIN FREE HYPODERMIC NEEDLE | 2010-12-28 | A hypodermic needle having a hub with a distal surface extending at an obtuse angle from a short needle, and a bevel approximately parallel to the distal surface. | 1. A hypodermic needle comprising: a hub having a proximal opening for receiving a mating connector at a proximal end and a generally planar distal surface at a distal end; anda needle having a proximal end adjacent the distal surface of the hub and a distal end, the distal end comprising a bevel, the bevel having an anterior end and a posterior end,wherein the needle extends from the distal end of the hub at an obtuse angle with respect to the distal surface of the hub, and wherein the bevel is approximately parallel to the distal surface of the hub. 2. The hypodermic needle of claim 1, wherein a posterior portion of the distal end of the distal surface of the hub comprises a posterior surface extending inwardly with respect to the distal surface of the hub. 3. The hypodermic needle of claim 2, wherein the posterior surface is curved. 4. The hypodermic needle of claim 1, wherein the length of the needle between the distal surface of the hub and the anterior end of the bevel is between 1 millimeter and 3 millimeters. 5. The hypodermic needle of claim 1, wherein the distal surface of the hub has a surface area of between 0.5 cm2 and 15 cm2. 6. The hypodermic needle of claim 1, wherein the distal surface of the hub has a surface area of between 1 cm2 and 10 cm2. 7. The hypodermic needle of claim 1, wherein the bevel angle and the obtuse angle at which the needle extends from the distal end of the hub differ by less than 5%. 8. The hypodermic needle of claim 1, wherein the bevel angle and the obtuse angle at which the needle extends from the distal end of the hub differ by between 10% and 30%. 9. The hypodermic needle of claim 1, wherein the obtuse angle is approximately 135°. 10. The hypodermic needle of claim 1, wherein the bevel angle is between 135° and 170°. 11. The hypodermic needle of claim 1, further comprising a cap for enclosing the hypodermic needle. 12. The hypodermic needle of claim 1, further comprising a syringe attached to the proximal end of the hub. 13. The hypodermic needle of claim 1, further comprising an insulin pen attached to the proximal end of the hub. 14. A method of injecting a fluid into a subject with a hypodermic needle, comprising: (a) providing a hypodermic needle having: (i) a hub having a proximal opening for receiving a mating connector at a proximal end and a generally planar distal surface at a distal end; and(ii) a needle having a proximal end adjacent the distal surface of the hub and a distal end, the distal end comprising a bevel, the bevel having an anterior end comprising a tip and a posterior end,wherein the needle extends from the distal end of the hub at an obtuse angle with respect to the distal surface of the hub, and wherein the bevel is approximately parallel to the distal surface of the hub;(b) connecting the hypodermic needle to a device for retaining a fluid medicament;(c) puncturing the surface of the skin of the subject with the tip of the needle at an acute angle;(d) inserting the... | - | A61M005/00 A61M005/34 A61B005/15
US07966074 | B2 | 2011-06-21 | Apparatus and method enhancing muscular movement | 2007-05-07 | An apparatus to enhance muscular movement is provided. The apparatus includes at least one muscular movement sensor to sense a result of an attempt to move a muscle, a movement information controller to analyze the sensed muscular movement, and generate muscular movement information based on the analyzed muscular movements, and a muscular movement actuator to enhance the movement of the muscle according to the generated muscular movement information by actively controlling a deformation of the muscle movement actuator over the surface of the muscle. | 1. A wearable apparatus to enhance a muscular movement, the wearable apparatus comprising: at least one muscular movement sensor to sense a result of an initial muscular movement;a muscular movement actuator positioned external to a muscle; anda movement information controller to analyze the sensed result of the muscular movement and to generate muscular movement information configured to control the muscular movement actuator based on the analyzed muscular movement,wherein the muscular movement actuator is adapted to enhance the movement of the muscle, according to the muscular movement information generated by the movement information controller, by actively controlling a deformation of the muscle movement actuator over the surface of the muscle. 2. The apparatus of claim 1, wherein the muscular movement information comprises an intensity and a direction of the movement of the muscle. 3. The apparatus of claim 1, further comprising a storage module to store the muscular movement information as a database according to a set of standards, wherein the movement information controller controls the muscular movement actuator according to the muscular movement information stored in the storage module. 4. The apparatus of claim 3, wherein the standards comprise at least one of a gender of a user of the apparatus, an age of a user of the apparatus, a user body part to which the muscle corresponds, and a time when the user uses the apparatus. 5. The apparatus of claim 1, wherein the muscular movement actuator comprises: an electroactive polymer (EAP) configured as one of fibers and a pad;a pair of electrodes to contact opposite lateral sides of the EAP; andan electric circuit to apply a voltage to the electrodes. 6. The apparatus of claim 5, wherein the EAP is a wet polymer that causes movement of ions inside the wet polymer according to the intensity of the voltage, wherein the voltage varies according to the muscular movement information provided by the movement information controller. 7. The apparatus of claim 6, wherein the EAP contracts or relaxes the muscle by increasing or reducing curvature according to a movement of ions. 8. The apparatus of claim 7, wherein the EAP is formed as a single wavy layer having a uniform curvature, and the electrodes are formed by connecting one of an anode and a cathode to an upper curvature portion of the EAP and connecting whichever of the anode and the cathode is left unconnected to the upper curvature portion of the EAP to a lower curvature portion of the EAP. 9. The apparatus of claim 7, wherein the EAP is formed as a double wavy layer having a uniform curvature, and the electrodes are formed by connecting one of one of an anode and a cathode to the upper and lower curvature portions of the EAP and connecting whichever of the anode and the cathode is connected to none of the upper and lower curvature portions of the EAP to an interface between upper and lower parts of the EAP. 10. The apparatus of claim... | Samsung Electronics Co., Ltd. | A61N001/00 A61B005/04
US07966060 | B2 | 2011-06-21 | Method and apparatus for determining autofluorescence of skin tissue | 2004-01-29 | For determining autofluorescence of a clinically normal skin issue (7) of a patient, the tissue (7) is irradiated with electromagnetic radiation. An amount of fluorescent radiation emitted by the tissue (7) in response to the irradiation is measured and, in response thereto, a signal is generated which represents a determined autofluorescence in agreement with the measured amount of electromagnetic radiation. Because the tissue (7) is skin tissue (7) in vivo and irradiation is performed noninvasively, a method which is very simple to apply for determining autofluorescence is obtained. An apparatus especially adapted for use in this method is also described. | 1. A method for determining an advanced glycation/glycosylation end product content of non-locally anomalous, intact skin tissue of a human individual, comprising: using a radiation source to irradiate said skin tissue with electromagnetic excitation radiation;receiving and measuring an amount of electromagnetic, fluorescent radiation emitted by said skin tissue in response to said irradiation; andgenerating, in response to said measured amount of fluorescent radiation, a signal which represents said measured amount of fluorescent radiation; anddetermining an advanced glycation/glycosylation end product content of said skin tissue from said signal, and outputting said determined advanced glycation/glycosylation end product content;wherein said skin tissue is non-locally anomalous, intact skin tissue in vivo which is irradiated noninvasively and simultaneously in its entirety by directing said radiation from said radiation source to a portion of the outer surface of the skin, wherein fluorescent radiation emitted in response to said irradiation is simultaneously received from a surface area of the irradiated portions of the skin surface of at least 1 cm2, wherein said fluorescent radiation is received via a measuring window; and wherein said measuring window is oriented at an angle of 25-65° relative to the irradiated portion of the surface of the skin. 2. A method for determining an advanced glycation/glycosylation end product content of non-locally anomalous, intact skin tissue of a human individual, comprising: using a radiation source to irradiate said skin tissue with electromagnetic excitation radiation;receiving and measuring an amount of electromagnetic, fluorescent radiation emitted by said skin tissue in response to said irradiation; andgenerating, in response to said measured amount of fluorescent radiation, a signal which represents said measured amount of fluorescent radiation; anddetermining an advanced glycation/glycosylation end product content of said skin tissue from said signal, and outputting said determined advanced glycation/glycosylation end product content;wherein said skin tissue is non-locally anomalous, intact skin tissue in vivo which is irradiated noninvasively and simultaneously in its entirety by directing said radiation from said radiation source to a portion of the outer surface of the skin, wherein fluorescent radiation emitted in response to said irradiation is simultaneously received from a surface area of the irradiated portions of the skin surface of at least 1 cm2, wherein a supporting structure is held against the skin of the individual, wherein the irradiated skin tissue area is located behind an opening in the supporting structure, wherein the supporting structure supports a measuring window, wherein said fluorescent radiation is received via said measuring window and wherein said measuring window is held at a distance from the skin. 3. An apparatus for determining an advanced glycation/glycosylation end... | Diagnoptics Holding B.V. | A61B005/00 A61B005/1455
US20110137135 | A1 | 2011-06-09 | Context Aware Physiological Monitoring | 2009-12-09 | To make home patient care more successful, a machine may enable context aware physiological monitoring at home. For example, in one embodiment, the machine only allows the physiological monitor measurements to be made when the patient has proven that the patient has not undergone strenuous activity in a period preceding the measurement. That activity may skew the measurement and make the measurement unreliable. Without the machine information, the clinician may be unaware of the physical activity because the patient is not present in his or her office and could misdiagnose or misprescribe based on incorrect information. | 1. An apparatus comprising: a monitor interface to receive a physiological monitor measurement;a transmitter to transmit the measurement to a remote location; anda sensor interface to obtain information about kinematic activity from an activity sensor. 2. The apparatus of claim 1 including an activity sensor having an accelerometer to detect user motion, said activity sensor being mountable on the human body. 3. The apparatus of claim 2 wherein said activity sensor includes a gyroscope. 4. The apparatus of claim 2 wherein said activity sensor includes at least one of a temperature sensor, a moisture sensor, and a pulse meter. 5. The apparatus of claim 1 wherein said monitor interface to request kinematic activity information from the activity sensor in response to a request to provide a physiological monitor measurement. 6. The apparatus of claim 1, said sensor interface to receive information about user activity for a preset amount of time prior to the request for that information. 7. The apparatus of claim 1 wherein said sensor interface to determine the average activity level and to compare said average activity level to a threshold to determine whether excessive activity has occurred. 8. The apparatus of claim 1 including a controller to determine whether excessive activity has occurred and, if so, to provide an indication that the physiological monitor measurement should be delayed. 9. The apparatus of claim 1 to combine physiological monitor measurements and activity sensor information and to automatically transmit the combined data to a remote location. 10. The apparatus of claim 1 wherein said physiological monitor measurements include a pulse oximeter, a blood pressure cuff, and a peak flow meter. 11. A computer readable medium storing instructions executed by a computer to: in response to a request to provide a physiological monitor measurement, request information about user activity; andif said user activity is in excess of a threshold, providing an indication to delay the physiological monitor measurement. 12. The medium of claim 11 further storing instructions to combine the physiological monitor measurement and the activity information and to transmit the information to a remote location. 13. The medium of claim 11 further storing instructions to determine an average activity level, to compare that said average activity level to a threshold, and to determine whether excessive activity has occurred. 14. The medium of claim 11 further storing instructions to control the taking of physiological monitor measurements until an activity level is below a threshold. 15. A method comprising: receiving a request for a physiological monitor measurement;in response, obtaining information about the user's activity level from a body wearable sensor; anddetermining whether the physiological monitor measurement should be taken based on the activity level information. 16. The method of claim 15 including providing an indication that t... | - | A61B005/11 A61B005/01 A61B005/024 A61B005/022
US20110112414 | A1 | 2011-05-12 | SYSTEM FOR AUTOMATED EKG ANALYSIS | 2009-11-06 | Systems and methods for analyzing electronic cardiac signals for use in clinical diagnostics are described. System and method embodiments are configured to assist in the analysis of details of EKG signals and vector cardiograms to determine how patients should be categorized into specific cardiac risk categories, such as an acute coronary syndrome category. System configurations may comprise memory devices, computing systems, and EKG data sources positioned at various local or remote positions, and connected via various data connectivity modalities. Various parameters may be utilized to assist in the drawing of one or more conclusions regarding the cardiac condition of a patient. | 1. A system for cardiac condition detection, comprising: a. a source of EKG data pertinent to a patient; andb. a first computing system configured to: 1) receive a quantity of EKG data from the source;2) construct a 3-D representation of cardiac activity from the data;3) compute values for one or more parameters based upon at least one of said 3-D representation of cardiac activity and said EKG data;4) apply a multifactorial analysis protocol utilizing at least one of the values of the one or more parameters; and5) automatically draw one or more conclusions regarding the cardiac condition of the patient based at least in part upon the output of the multifactorial analysis protocol. 2. The system of claim 1, wherein the source comprises a plurality of EKG electrodes. 3. The system of claim 1, wherein the source comprises a memory device containing stored EKG data. 4. The system of claim 1, wherein the source comprises a reconstructed 12-lead EKG from a reduced cardiographic vector set. 5. The system of claim 1, further comprising an intermediate device interposed between the source and the first computing system, the intermediate device configured to capture data from the source and pass it to the first computing system, wherein the source is a plurality of EKG electrodes. 6. The system of claim 4, wherein the reduced vector set is provided by another device. 7. The system of claim 6, wherein the other device is selected from the group consisting of: a defibrillator, a pacemaker, a cardioverter, and a portable EKG system. 8. The system of claim 3, wherein the memory device is coupled to the first computing system with a connection selected from the group consisting of: a wired local connection; a wireless local connection; a wired remote location; and a wireless remote location. 9. The system of claim 1, wherein the source of the EKG data and first computing system are operatively coupled with a connection selected from the group consisting of: a wired local connection; a wireless local connection; a wired remote location; and a wireless remote location. 10. The system of claim 1, wherein the source of the EKG data and first computing system are physically integrated. 11. The system of claim 1, further comprising a storage device and an EKG system, the storage device operatively coupled between the first computing system and the EKG system and configured to capture data from the EKG system and allow selective access of such data from said storage device. 12. The system of claim 1, wherein the multifactorial analysis protocol further utilizes one or more parameters computed directly from the 12-lead EKG data. 13. The system of claim 1, wherein one of the one or more parameters is a ratio-metric parameter based upon two or more aspects of the EKG data. 14. The system of claim 13, wherein one of the one or more parameters is a ratio of Rmax relative to STshift or Tmax. 15. The system of claim 1, wherein one of the one or more parameters... | NEWCARDIO, INC | A61B005/0402
US20110112412 | A1 | 2011-05-12 | CUFF FOR BLOOD PRESSURE INFORMATION MEASURING DEVICE AND BLOOD PRESSURE INFORMATION MEASURING DEVICE EQUIPPED WITH THE SAME | 2009-06-18 | A cuff for a blood pressure information measuring device configured to be mounted on a measurement site for measuring blood pressure information has a fluid bag for compressing the measurement site, a flexible curved elastic plate located outside the fluid bag while the fluid bag is wrapped around the measurement site, and annular or arc-shaped to extend along the measurement site, an inflation and deflation mechanism secured to the curved elastic plate and inflating and deflating the fluid bag, and a bag-shaped cover body encapsulating the fluid bag, the inflation and deflation mechanism, and the curved elastic plate. | 1. A cuff for a blood pressure information measuring device configured to be mounted on a measurement site for measuring blood pressure information comprising: a fluid bag for compressing the measurement site;a flexible curved elastic plate located outside said fluid bag while said fluid bag is wrapped around the measurement site, and annular or arc-shaped to extend along the measurement site;an inflation and deflation mechanism secured to said curved elastic plate and inflating and deflating said fluid bag; anda bag-shaped cover body encapsulating said fluid bag, said inflation and deflation mechanism, and said curved elastic plate. 2. The cuff for a blood pressure information measuring device according to claim 1, wherein said inflation and deflation mechanism is disposed on a main surface of said curved elastic plate, andsaid main surface is opposite to a main surface of said curved elastic plate that faces said fluid bag. 3. The cuff for a blood pressure information measuring device according to claim 1, further comprising a protection member covering an exposed surface of said inflation and deflation mechanism. 4. The cuff for a blood pressure information measuring device according to claim 1, wherein said inflation and deflation mechanism includes a pressurizing pump and an exhaust valve. 5. The cuff for a blood pressure information measuring device according to claim 1, further comprising: a power supply unit supplying electric power for driving said inflation and deflation mechanism,wherein said power supply unit is encapsulated in said bag-shaped cover body. 6. The cuff for a blood pressure information measuring device according to claim 1, further comprising: a pressure detection unit for detecting an internal pressure of said fluid bag,wherein said pressure detection unit is encapsulated in said bag-shaped cover body (12). 7. A blood pressure information measuring device comprising: the cuff for the blood pressure information measuring device as recited in claim 1; andan associated unit provided separately and independently of said cuff for the blood pressure information measuring device,wherein said associated unit is provided with a display portion for displaying blood pressure information as a result of measurement. 8. The blood pressure information measuring device according to claim 7, wherein said associated unit is provided with an operation portion for entering an instruction to drive said inflation and deflation mechanism. 9. The blood pressure information measuring device according to claim 7, further comprising a communication unit enabling radio communication to be performed between said cuff for the blood pressure information measuring device and said associated unit. | OMRON HEALTHCARE CO., LTD. | A61B005/022
US07951357 | B2 | 2011-05-31 | Implantable power sources and sensors | 2005-07-13 | A protein is provided, including a glucose binding site, cyan fluorescent protein (CFP), and yellow fluorescent protein (YFP). The protein is configured such that binding of glucose to the glucose binding site causes a reduction in a distance between the CFP and the YFP. Substance monitoring apparatus (210) is also provided, including a semi-permeable barrier (212), adapted to be implanted in a body of a subject and to allow passage therethrough of a substance, while inhibiting passage therethrough of immune cells; and microorganisms (214), disposed within the semi-permeable barrier (212) so as to produce a measurable response to a level of the substance. A sensor (220) is adapted to measure the measurable response and not to measure a response of any mammalian cells that may be disposed within the semi-permeable barrier (212). | 1. Apparatus for detecting a concentration of a substance in a subject, the apparatus comprising: a fluorescence resonance energy transfer (FRET) measurement device; anda housing adapted to be implanted in the subject, the housing comprisingcells genetically engineered to produce, in situ, a FRET protein having a FRET complex comprising a fluorescent protein donor, a fluorescent protein acceptor, and a binding site for the substance. 2. The apparatus according to claim 1, wherein the protein that the cells are genetically engineered to produce comprises: a glucose binding site;cyan fluorescent protein (CFP); andyellow fluorescent protein (YFP),wherein the protein is configured such that binding of glucose to the glucose binding site changes a distance between the CFP and the YFP. 3. The apparatus of claim 2, wherein the protein is encoded by an isolated nucleic acid fragment having the nucleotide sequence in SEQ ID NO: 1. 4. The apparatus according to claim 1, comprising an inducer that is adminsterable to the subject and that regulates protein expression of the cells. 5. The apparatus according to claim 1, wherein the cells comprise beta or liver cells. 6. The apparatus according to claim 1, wherein the substance includes glucose, and wherein the binding site comprises a binding site for glucose. 7. The apparatus according to claim 1, wherein the protein comprises a protein selected from the group consisting of blue fluorescent protein (BFP), green fluorescent protein (GFP), and GFP/Rhodamine. 8. The apparatus according to claim 1, wherein the housing comprises a chamber which holds the cells. 9. The apparatus according to claim 1, wherein the cells are held directly within the housing, with no intervening chamber to hold the cells. 10. The apparatus according to claim 1, wherein the cells are disposed within a matrix. 11. The apparatus according to claim 1, wherein the cells are disposed in a suspension. 12. The apparatus according to claim 1, wherein the housing comprises a membrane that prevents passage therethrough of cells of the subject's body. 13. The apparatus according to claim 12, wherein the membrane is configured to prevent passage therethrough of the FRET protein. 14. The apparatus according to claim 1, wherein the cells are configured such that the protein produced by the genetically-engineered cells becomes positioned on the membranes of the cells. 15. The apparatus according to claim 14, wherein the protein comprises a targeting peptide that is configured to direct the protein to the membranes of the cells. 16. The apparatus according to claim 15, wherein the sequence of the targeting peptide is SEQ ID NO: 2. 17. The apparatus according to claim 1, comprising a drug capable of killing the cells, the drug being configured to be administered to the subject subsequently to the implantation of the housing in the subject. 18. The apparatus according to claim 1, wherein the fluorescence resonance energy transfer (FRET)... | Glusense Ltd. | A61B005/00 C07H021/04 C07K014/00 C12N015/85
US07963926 | B2 | 2011-06-21 | Trending of systolic murmur intensity for monitoring cardiac disease with implantable device | 2008-07-28 | Described is an implantable device configured to monitor for changes in the intensity and/or duration of a systolic murmur such as mitral regurgitation by means of an acoustic sensor. Such changes may be taken to indicate a change in a patient's heart failure status. Upon detection of a worsening in the patient's heart failure statue, the device may be programmed to alert clinical personnel over a patient management network and/or make appropriate adjustments to pacing therapy. | 1. An implantable cardiac device, comprising: a pressure sensor adapted for disposition so as to sense the acoustic vibrations produced by a systolic murmur and produce a pressure signal reflective thereof;circuitry for measuring the intensity of the systolic murmur from the pressure signal produced by the pressure sensor;an impedance sensor configured to measure transthoracic impedance reflective of cardiac volume;a controller programmed to record a baseline murmur intensity; and,wherein the controller is further programmed to measure a current murmur intensity as an average of intensity measurements taken over some period of time, compare the current murmur intensity to the baseline murmur intensity, and detect a change indicative of worsening cardiac status if the current murmur intensity exceeds the baseline murmur intensity by a first specified threshold amount and if the transthoracic impedance measurement indicates an increase in cardiac volume above a second specified threshold amount. 2. The device of claim 1 wherein the controller is programmed to measure a current duration of the murmur during the systolic phase, compare the current murmur duration to a baseline murmur duration, and set an alarm flag if the current murmur duration exceeds the baseline murmur duration by the first specified threshold amount. 3. The device of claim 1 further comprising a signal energy detector and an R-wave detector for enabling the signal energy detector to measure the intensity of the murmur during the systolic phase of the cardiac cycle. 4. The device of claim 1 further comprising a telemetry transceiver and wherein the controller is programmed to transmit an alarm message over a patient management network via telemetry if the current murmur intensity exceeds the baseline murmur intensity by the first specified threshold amount. 5. The device of claim 1 wherein the controller is further programmed to store the current murmur intensity measurements take over a specified period time and compute a trend in the murmur's intensity. 6. The device of claim 5 wherein the trend is computed by computing differences between consecutive current murmur intensity measurements over the specified period of time and summing the differences. 7. The device of claim 1 further comprising: a bandpass filter for filtering the pressure signal; and,a signal energy detector for measuring a parameter related to the signal energy of the filtered pressure signal to represent the murmur intensity. 8. The device of claim 1 further comprising circuitry for pacing the heart in a programmed mode and wherein the controller is programmed to adjust the programmed mode if the current murmur intensity exceeds the baseline murmur intensity by a specified threshold amount. 9. The device of claim 8 wherein the device is configured to pace the heart in an atrial triggered mode and wherein the controller is programmed to adjust a programmed AV delay if the current murmur intensity excee... | Cardiac Pacemakers, Inc. | A61B005/02
US08343062 | B2 | 2013-01-01 | Digital control method for measuring blood pressure | 2010-10-29 | A system and method of digital control for a blood pressure measurement system is provided. According to at least one embodiment, a photo-plethysmographic (PPG) system produces a frequency signal that corresponds to the measured light in the PPG system. Such light may be indicative of blood volume in a vein or artery. The frequency signal may be used to control one or more pressure valves of the system in order to measure blood pressure and hold the frequency signal constant. | 1. A device for continuous blood pressure measurement comprising: a pressure cuff adapted to be placed over an artery in a human finger, the cuff including a PPG system having at least one light source and at least one light detector, the at least one light detector being associated with a light-to-frequency (LFC) conversion device that is configured for producing an AC signal based on light intensity measured from the one or more light detectors, wherein a frequency of the AC signal is representative of blood volume in the artery;a pressure system comprising:at least one pump,at least one valve or valve system, andat least one pressure sensor; anda controller configured for controlling pressure in the cuff by altering the valve or valve system and holding the AC signal substantially constant by altering the pressure in the cuff. 2. The blood pressure measurement device according to claim 1, wherein the controller is a computing device. 3. The blood pressure measurement device according to claim 2, wherein the LFC device is electrically connected to a digital input of the computing device. 4. The blood pressure measurement device according to claim 2, wherein the at least one valve or valve system is controlled via a digital output from the computing device. 5. The blood pressure measurement device according to claim 4, wherein the digital output is a pulse width modulated (PWM) signal. 6. The blood pressure measurement device according to claim 2, wherein the at least one light source and the at least one light detector are controlled via a digital output of the computing device. 7. The blood pressure measurement device according to claim 6, wherein a blanking-interval of the light source is used for ambient light detection. 8. The blood pressure measurement device according to claim 2, wherein the computing device receives information from a calibration device. 9. The blood pressure measurement device according to claim 1, wherein the blood pressure measurement device receives information from a hydrostatic correction system. 10. The blood pressure measurement device according to claim 1, wherein the blood pressure measurement device receives scaling information from at least one other device. 11. The blood pressure measurement device according to claim 10, wherein the blood pressure measurement device calculates a scaled blood pressure curve and transmits the scaled blood pressure curve to the other device. 12. A method for continuously measuring blood pressure comprising: placing a photo-plethysmographic (PPG) system over an artery or vein in a human finger, the PPG system producing a PPG signal based on volume of the artery or vein, the PPG system including at least one light source and at least one light detector;utilizing a computing device to alter a pressure inside the cuff by altering a valve or valve system that is connected to a pump and pressure sensor,wherein, based on the measured blood volume of the artery or vein, a fr... | CNSystems Medizintechnik AG | A61B005/02
US07955261 | B2 | 2011-06-07 | System and methods for processing analyte sensor data | 2010-03-23 | Systems and methods for processing sensor analyte data, including initiating calibration, updating calibration, evaluating clinical acceptability of reference and sensor analyte data, and evaluating the quality of sensor calibration. During initial calibration, the analyte sensor data is evaluated over a period of time to determine stability of the sensor. The sensor may be calibrated using a calibration set of one or more matched sensor and reference analyte data pairs. The calibration may be updated after evaluating the calibration set for best calibration based on inclusion criteria with newly received reference analyte data. Fail-safe mechanisms are provided based on clinical acceptability of reference and analyte data and quality of sensor calibration. Algorithms provide for optimized prospective and retrospective analysis of estimated blood analyte data from an analyte sensor. | 1. A method for calibrating data from a glucose sensor, the method comprising: receiving sensor data from a glucose sensor, including one or more sensor data points; anddetermining, using an electronic system, a stability of the glucose sensor by comparing the sensor data with a waveform template, wherein an amount of agreement between the template and the sensor data is indicative of a level of sensor stability. 2. The method of claim 1, further comprising calibrating the sensor data only if a predetermined level of sensor stability has been determined. 3. The method of claim 1, further comprising receiving reference data, including one or more reference data points, wherein determining a stability further comprises evaluating a correlation between reference and sensor data points. 4. The method of claim 3, wherein receiving reference data comprises receiving data obtained from a blood glucose test. 5. The method of claim 3, wherein receiving reference data comprises receiving an internal wired communication from a reference glucose monitor. 6. The method of claim 1, wherein determining a stability further comprises evaluating a frequency content of the sensor data. 7. A system for calibrating data from a glucose sensor, the system comprising: a sensor data module configured to receive sensor data from the glucose sensor, including one or more sensor data points; anda processor module configured to determine a stability of the glucose sensor by comparing the sensor data with a waveform template, wherein a predetermined amount of agreement between the template and the sensor data is indicative of sensor stability. 8. The system of claim 7, wherein processor module is configured to calibrate the glucose sensor only after the glucose sensor reaches a predetermined level of sensor stability. 9. The system of claim 7, further comprising a reference input module configured to receive one or more reference glucose points from a reference glucose monitor, wherein the processor module is further configured to determine a level of sensor stability by evaluating a correlation between reference and sensor data points. 10. The system of claim 7, further comprising a reference input module configured to receive one or more reference glucose points from a reference glucose monitor, wherein the one or more reference data points are obtained from a blood glucose test. 11. The system of claim 7, further comprising a reference input module configured to receive one or more reference glucose points from a reference glucose monitor, wherein the one or more reference data points are obtained from a reference glucose monitor integral with said system. 12. The system of claim 7, wherein processor module is further configured to determine a stability of the glucose sensor by evaluating a frequency content of the sensor data. 13. The method of claim 1, further comprising prompting a user for reference glucose data responsive to a predetermined level of sensor... | DexCom, Inc. | A61B005/00
US08185181 | B2 | 2012-05-22 | Method and apparatus for detecting false hypoglycemic conditions | 2010-10-29 | Embodiments of the present disclosure include detecting a concurrent occurrence of a decrease in monitored analyte level and a corresponding decrease in monitored on-skin temperature, confirming a presence of an impending hypoglycemic condition, and asserting a notification corresponding to the confirmed impending hypoglycemic condition. Devices, methods, systems and kits incorporating the same are also provided. | 1. A method, comprising: receiving, using one or more processors, a plurality of time spaced analyte related data monitored by an analyte sensor in fluid contact with interstitial fluid during a first time period;detecting, using the one or more processors, when one or more of the received plurality of time spaced analyte related data crosses a predetermined analyte threshold level during the first time period;receiving, using the one or more processors, a plurality of time spaced temperature data during the first time period;determining, using the one or more processors, a rate of change of the received plurality of time spaced temperature data and detecting when the determined rate of change crosses a predetermined rate of temperature change;determining, using the one or more processors, a rate of change of the received plurality of time spaced analyte related data and comparing a slope of the determined rate of change of the received plurality of time spaced analyte related data to a slope of the rate of change of the received plurality of time spaced temperature data; andasserting, using the one or more processors, a notification when the determined rate of change of the received plurality of the time spaced temperature data reaches the predetermined rate of temperature change and when the one or more of the received plurality of time spaced analyte related data reaches the predetermined threshold analyte level during the first time period. 2. The method of claim 1 further including determining, using the one or more processors, when the monitored analyte level based on the received plurality of time spaced analyte related data indicates approaching the predetermined analyte threshold level during a second time period. 3. The method of claim 2 wherein the first time period precedes the second time period. 4. The method of claim 1 wherein the asserted notification includes one or more of an audible alert, a vibratory alert, a visual alert, or one or more combinations thereof. 5. The method of claim 1 wherein the predetermined analyte threshold level is associated with one of an impending hypoglycemic condition or an onset of hypoglycemic condition. 6. The method of claim 1 wherein the slope of the determined rate of change of the received plurality of analyte related data and the slope of the rate of change of the received plurality of time spaced temperature data are coincident. 7. The method of claim 1 wherein the asserted notification includes an impending hypoglycemic condition. 8. The method of claim 1 wherein when the determined rate of change of the received plurality of the time spaced temperature data does not exceed the predetermined rate of temperature change, deactivating a notification function. 9. The method of claim 8 wherein the deactivated notification function includes a hypoglycemic alarm. 10. An apparatus, comprising: one or more processors; anda memory storing instructions which, when executed by the one or more p... | Abbott Diabetes Care Inc. | A61B005/05
US07961326 | B2 | 2011-06-14 | Method of detecting the concentration of an analyte | 2008-08-21 | A system and method is provided for detecting concentration of an analyte in a fluid. The method comprises detecting an optical property of a first region of two or more regions in a system, the first region located in a container having a reservoir for one or more modifiers of one or more optical properties of the first region. The movement of the one or more modifiers is responsive to changes in concentration of the analyte. A next step detects an optical property of a second region of the two or more regions in the system, the second region located in a container having a reservoir for one or more modifiers of one or more optical properties of the second region. The movement of the one or more modifiers is responsive to changes in concentration of a compound, where the compound is something other than the analyte. The detected optical property of the first region and the detected optical property of the second region are used in embodiments to separate the effect of the analyte on the detected optical property of the first region from the effect of the compound, where the compound is an interfering compound. | 1. A method for detecting concentration of an analyte in a fluid, the method comprising: providing a reflective optical cavity system having two or more light-transmissive regions at least partially bounded by components having light-transmissive and light-reflective properties, at least two of the components having substantially parallel surfaces;detecting an optical property of a first region of the two or more regions in the system, the first region located in a container having a reservoir for one or more modifiers of one or more optical properties of the first region, the movement of said one or more modifiers being responsive to changes in concentration of said analyte;detecting an optical property of a second region of the two or more regions in the system, the second region located in a container having a reservoir for one or more modifiers of one or more optical properties of the second region; the movement of said one or more modifiers being responsive to changes in concentration of a compound other than said analyte in the fluid;using said detected optical property of the first region and said detected optical property of the second region to separate the effect of said analyte on said detected optical property of the first region from the effect of said compound other than said analyte, said compound being an interfering compound; anddetermining the concentration of the analyte based on the difference in the effect of the analyte and the effect of the compound other than the analyte. 2. A method according to claim 1 wherein one or more optical property detected is the intensity of light transmitted through or reflected from an optical cavity detection region and further including determining the shift in the spectral position of one or more intensity peak. 3. A method according to claim 2 further including determining the shift in the spectral position of one or more intensity peak divided by the free spectral range. 4. A method according to claim 1 further including determining changes in refractive index in the first and second containers of a system with a first container in which the first region is located and a second container in which the second region is located. 5. A method according to claim 1 further including determining changes in optical absorbance in the first and second containers of a system with a first container in which the first region is located and a second container in which the second region is located. 6. A method according to claim 4 further including determining a concentration of the first analyte by the determined refractive index. 7. A method according to claim 5 further including determining a concentration of the first analyte by the determined absorption. 8. A method according to claim 1 further including, detecting an optical property of an optical cavity detection region in at least a third container which is preloaded with a known reference fluid, the preloaded third container being sealed ... | Palo Alto Research Center Incorporated | A61B005/00 G01J003/46
US07955271 | B2 | 2011-06-07 | Tape transport lance sampler | 2006-10-13 | A lancet-sampler system is configured to automatically remove a protective cover from a lancet and automatically unpack a test pad just prior to use. This minimizes the risk of injury and reduces the chance of cross-contamination between the lancet and the test pad. The lancet defines a capillary groove for drawing body fluid from the incision via capillary action and a sample transfer opening for collecting the fluid from the groove. A carrier tape is coupled to the lancet. The carrier tape includes a test pad for analyzing the fluid. The tape is folded around the test pad to form an airtight package. The test pad is located at a position to align with the sample transfer opening when the tape is unfolded. The protective cover covers a portion of the lancet, and when the tape is pulled, the protective cover is automatically pulled from the lancet. | 1. A body fluid sampling device, comprising: a lancet including a lancet tip configured to lance tissue;a protective cover covering at least a portion of the lancet tip; anda tape coupled to the lancet and the protective cover, the tape having a slackened section between the lancet and the protective cover for allowing removal of the protective cover from the lancet tip when the tape is pulled, wherein the slackened section is longer than the distance between where the tape is coupled to the lancet and where the tape is coupled to the protective cover. 2. The device of claim 1, further comprising: a test pad for analyzing a body fluid sample disposed on the tape at the slackened section;the tape at the slackened section being folded around the test pad to form a packet for protecting the test pad prior to use; andthe packet being configured to unfold when the tape is pulled. 3. The device of claim 2, wherein: the lancet includes a capillary groove extending along the lancet tip for drawing the body fluid sample and a sample transfer opening fluidly coupled to the capillary groove; andthe test pad is disposed on the tape at a position to align with the sample transfer opening when the packet is unfolded for transferring fluid from the sample transfer opening to the test pad. 4. The device of claim 3, further comprising a cover foil covering at least a portion of the capillary groove to form a closed capillary channel. 5. The device of claim 4, wherein the sample transfer opening has a vent slot that is left uncovered by the cover foil for venting air as the body fluid sample fills the sample transfer opening. 6. The device of claim 3, further comprising: the capillary groove being partially etched;the sample transfer opening being fully etched;the sample transfer opening being wider than the capillary groove; andthe lancet having a fully etched section having the same width as the capillary groove, the fully etched section being located between the capillary groove and the sample transfer opening to provide a gradual transition between the capillary groove and the sample transfer opening to allow momentum of the body fluid sample to carry the body fluid sample to the sample transfer opening. 7. The device of claim 1, further comprising: a first tape connector coupling the protective cover to the tape; anda second tape connector coupling the lancet to the tape. 8. The device of claim 1, wherein the lancet defines an actuator engagement opening configured to engage an actuator member of a lancing mechanism to fire the lancet. 9. The device of claim 1, wherein the protective cover includes at least a pair of protective foils between which the lancet tip is sandwiched. 10. The device of claim 1, wherein the slackened section includes a section of the tape that is folded. 11. The device of claim 10, wherein the section of the tape that is folded is folded in a fanfold manner. 12. The device of claim 1, wherein the slackened section is at lea... | Roche Diagnostics Operations, Inc. | A61B005/00 C12Q001/00
US07970451 | B2 | 2011-06-28 | Textile-based electrode | 2008-12-31 | Textile-based electrodes include a fabric portion having stretch-recovery non-conductive yarns and an electrically conductive region having stretch-recovery electrically conductive yarn filaments. The electrodes can further include float yarns and can be configured in a textured or ribbed construction. When incorporated into a garment, the electrodes can be used to monitor biophysical characteristics, such as the garment wearer's heart rate. | 1. A textile-based electrode system, comprising: at least a first fabric portion and a second fabric portion;wherein: the first fabric portion comprises at least a first electrically conductive region;the second fabric portion comprises at least a second electrically conductive region;the first and second fabric portions at least partially overly one another;the electrically conductive region of the first fabric portion and the electrically conductive region of the second fabric portion cooperate to provide a region of at least partial physical contact; andthe region of at least partial physical contact establishes electrical conduction between the first electrically conductive region and the second electrically conductive region. 2. The textile-based electrode system according to claim 1, wherein at least one of: (i) the electrically conductive region of the first fabric portion and (ii) the electrically conductive region of the second fabric portion, comprises a float yarn. 3. The textile-based electrode system according to claim 2, wherein: (i) the electrically conductive region of the first fabric portion and (ii) the electrically conductive region of the second fabric portion both comprise float yarns that are stitched together. 4. The textile-based electrode system according to claim 1, wherein at least one of: (i) the electrically conductive region of the first fabric portion and (ii) the electrically conductive region of the second fabric portion, comprises at least a portion of an elastified electrically conductive yarn. 5. The textile-based electrode system according to claim 1, wherein at least one of: (i) the electrically conductive region of the first fabric portion and (ii) the electrically conductive region of the second fabric portion, comprises an elastic yarn at least partially plated with a conductive yarn. 6. The textile-based electrode system according to claim 1, wherein at least one of: (i) the electrically conductive region of the first fabric portion and (ii) the electrically conductive region of the second fabric portion, comprises a fabric having a textured or ribbed construction. 7. The textile-based electrode system according to claim 1, wherein at least one of: (i) the electrically conductive region of the first fabric portion and (ii) the electrically conductive region of the second fabric portion, further comprises at least a portion including a hydrophilic material. 8. The textile-based electrode system according to claim 1, wherein at least one hydrophobic material is incorporated into an electrically conductive region of either the first fabric portion or the second fabric portion or both fabric portions. 9. The textile-based electrode system according to claim 1, further comprising means for transferring signals from at least one of the electrically conductive regions to a measuring device. 10. The textile-based electrode system according to claim 9, wherein the measuring device is attached to at least... | Textronics, Inc. | A61B005/04
US08880162 | B2 | 2014-11-04 | Device, method, and computer-readable storage medium for detecting events in cardiac signals | 2010-11-08 | The present disclosure generally relates to a method, a device, and a computer-readable storage medium for detecting heart beats from cardiac signals whose quality, expressed in terms of signal amplitude and signal-to-noise ratio, varies dynamically in time.;Hence, a method, a device, and a computer-readable storage medium for detecting electrical signals originating from a human or animal heart is proposed. The method includes the following steps: a) identifying an initial indication of the event in at least one of the signal channels,b) deciding whether or not the identified initial indication confirms the event depending on the quality of the signal channels in which initial indications are identified; andc) determining a point-of-detection for the event depending on the quality of the signal channels and depending from the shape of the signal. | 1. A method for detecting events in cardiac signals originating from at least two signal channels in a cardiac medical device, the method comprising the following steps: a. identifying an initial indication of an event in signal samples from at least one of the signal channels,b. deciding whether or not the identified initial indication confirms the event depending on the quality of the signal channels in which initial indications are identified, each signal channel's quality being dependent on: (1) the number of consecutive samples in the signal channel exceeding a settable detection threshold, and(2) the number of times a noise timer for the signal channel is restarted, the noise timer being with the cardiac medical device, wherein the channel's noise timer is restarted if: i. a sample magnitude for the signal channel rises above a settable noise threshold, orii. a sample polarity for the signal channel changes while the sample magnitude remains above the settable noise threshold; andc. determining a point-of-detection for the event depending on: (1) the quality of the signal channels, and(2) the shape of the signal. 2. The method according to claim 1, wherein step b comprises analyzing the quality of the signal channels, starting after a settable number of indications has been identified. 3. The method according to claim 1, wherein step b comprises, for confirmation of the event, selecting the channel that has the best quality. 4. The method according to claim 1, wherein for determining the point-of-detection, the channel having the best quality is selected. 5. The method according to claim 4, wherein the point-of-detection is determined to be the location of a peak from the selected channel. 6. The method according to claim 1, wherein the initial indication is identified depending on a number of consecutive samples having a magnitude that exceeds a settable detection threshold. 7. The method according to claim 6, wherein the initial indication is identified if an accumulated count of the consecutive samples having a magnitude that exceeds the settable detection threshold reaches a settable sample count limit. 8. The method according to claim 6, wherein, for determining the number of consecutive samples, such samples are considered which exceed the settable detection threshold while a noise timer is inactive. 9. The method according to claim 1, wherein the noise timer is started following the initial indication of an event. 10. The method according to claim 1, wherein at least one of the noise timer and the settable detection threshold are set to be channel-specific. 11. The method according to claim 1, wherein determining the point-of-detection depends on a slope of the signal. 12. The method according to claim 1, where after the initial indication of the event in at least one of the signal channels a detection holdoff timer is started, having an associated holdoff timer interval, and a search for a peak value is initialised in the... | BIOTRONIK CRM Patent AG | A61B005/04 A61B005/0456 A61B005/00
US08838226 | B2 | 2014-09-16 | Multi-channel brain or cortical activity monitoring and method | 2009-12-01 | The present invention relates to a quantitative electroencephalogram (QEEG) monitor and system capable of monitoring and displaying simultaneously neuropathological characteristic and activity of both sides of a subject's brain. The methods include various indices and examination of differences in these indices by which neurophysiological conditions or problems can be identified and treated. These methods, and the systems and devices using these methods preferably can be used for identifying these neurophysiological conditions or brain dysfunction with monitors and methods for seizure detection, for sedation monitoring, for anesthesia monitoring, and the like. These bilateral brain monitoring methods and systems, and the devices using these methods can be used by individuals or clinicians with little or no training in signal analysis or processing. These bilateral monitoring methods can also be used in a range of applications. | 1. A method of determining electroencephalogram (EEG) signal quality in a device for quantifying brain or cortical activity as a function of depth of sedation or anesthesia comprising steps of: a. anesthetizing or sedating a subject;b. monitoring the subject with a brain having a left hemisphere and a right hemisphere under anesthesia or sedation device for quantifying brain or cortical activity as a function of depth of sedation or anesthesia, the device with at least two measurement electrodes, and at least one reference electrode, the at least two measurement electrodes comprising at least one electroencephalogram (EEG) electrode, having a signal, positioned to monitor left hemisphere brain or cortical activity and at least one EEG electrode, having a signal, positioned to monitor right hemisphere brain or cortical activity of the subject's brain, the reference electrode comprising at least one EEG electrode, each electrode providing an EEG analog signal which is subsequently converted to a digital signal;c. measuring the brain or cortical activity of both the subject's left and right brain hemispheres essentially simultaneously over a period of time part of the period of time over which the subject is under sedation or anesthesia;d. calculating based in part on the digital signals with the processor at least one numerical quantitative electroencephalogram (QEEG) time-series index corresponding to the brain or cortical activity as a function of depth of anesthesia or sedation of each of the left and right hemispheres of the subject's brain over the period of time part of the period of time over which the subject is under sedation or anesthesia;e. displaying the at least two numerical time-series indices, with at least one numerical time-series index corresponding to the brain or cortical activity of each of the left and the right hemispheres over the period of time, part of the period of time over which the subject is under sedation or anesthesia, on a monitor simultaneously;f. comparing the numerical time-series indices of each hemisphere's cortical or brain activity as a function of depth of anesthesia or sedation; andg. determining, based at least in part on differences between the at least two numerical QEEG time-series indices corresponding to the brain or cortical activity as a function of depth of anesthesia or sedation, whether the signal of at least one of the EEG electrodes has poor quality. 2. The method of claim 1 wherein an additional step is comparing the calculated numerical brain or cortical activity time-series indices corresponding to the brain or cortical activity as a function of depth of anesthesia or sedation for each of the subject's brain hemispheres and determining based on differences between these indices whether the subject has suffered, or is suffering some neuropathological activity. 3. The method of claim 1 wherein an additional step is comparing the calculated numerical brain or cortical activity time-series i... | Neuro Wave Systems Inc | A61B005/0476 A61B005/00
US08467843 | B2 | 2013-06-18 | Optical sensor configuration for ratiometric correction of blood glucose measurement | 2009-11-04 | Embodiments of the invention are directed to an optical sensor for detecting blood glucose. The sensor comprises a chemical indicator system disposed within a gap between the distal end of an optical fiber and an atraumatic tip portion, wherein the optical fiber and atraumatic tip portion are coupled by a coupling member, such as a rod or hypotube or cage that traverses the gap. The sensor further comprises a means for generating and detecting an optical reference signal unrelated to the blood glucose, such that ratiometric correction of blood glucose measurements for artifacts in the optical system is enabled. | 1. A sensor for detecting an analyte concentration in a blood vessel, comprising: an optical fiber with proximal and distal ends;an atraumatic tip portion with proximal and distal ends, wherein the proximal end of the atraumatic tip portion is separated from the distal end of the optical fiber, such that a gap exists between the atraumatic tip portion and the optical fiber;a rod with proximal and distal ends, wherein the proximal end of the rod is attached to the distal end of the optical fiber, and wherein the distal end of the rod is attached to the proximal end of the atraumatic tip portion, such that the rod traverses the gap and couples the optical fiber to the atraumatic tip portion;a chemical indicator system disposed within the gap and capable of generating an emission light signal in response to an excitation light signal, wherein the intensity of the emission light signal is related to the analyte concentration; anda selectively permeable membrane disposed over the gap, wherein the sensor is sized for deployment within the blood vessel. 2. The sensor of claim 1, wherein the chemical indicator system is further immobilized within the gap by a hydrogel. 3. The sensor of claim 1, further comprising a temperature sensor. 4. The sensor of claim 1, wherein the optical fiber has a diameter of between about 0.005 inches and about 0.020 inches. 5. The sensor of claim 1, further comprising a reflective region. 6. The sensor of claim 5, wherein the reflective region comprises a reflective surface of the proximal end of the rod. 7. The sensor of claim 1, wherein the rod is attached to the optical fiber by heating. 8. The sensor of claim 1, wherein the rod is attached to the optical fiber by a reflective adhesive. 9. The sensor of claim 8, wherein the shape of the distal end of the atraumatic tip portion is configured to reduce trauma within the blood vessel and selected from the group consisting of hemispherical, parabolic, and elliptical. 10. The sensor of claim 1, wherein the distal end of the atraumatic tip portion is flexible. 11. The sensor of claim 1, wherein the distal end of the atraumatic tip portion is deformable. 12. The sensor of claim 1, wherein the distal end of the atraumatic tip portion is formed from at least one material selected from the group consisting of plastics, polymers, gels, metals and composites. 13. The sensor of claim 1, wherein the rod is formed from at least one material selected from the group consisting of metal, metal alloy, plastic, polymer, ceramic, and composite material. 14. The sensor of claim 13, wherein the rod is formed from stainless steel, titanium, or Nitinol. 15. The sensor of claim 1, wherein the rod is cylindrical. 16. The sensor of claim 15, wherein the rod diameter is between about 0.002 inches and about 0.010 inches. 17. The sensor of claim 1, wherein the rod is flexible. 18. The sensor of claim 1, wherein the rod is stiffer than the optical fiber. 19. The sensor of claim 1, wher... | Glumetrics, Inc. | A61B005/00
US08532769 | B2 | 2013-09-10 | Heart rate variability distinction | 2009-10-30 | Systems, methods and devices for monitoring, analyzing, and processing a patient's heart rate signal for HRV characteristics are described herein. A first heart rate signal is acquired. The first heart rate signal includes at least one indication of an interval duration of cardiac activity. At least one accelerating portion and at least one decelerating portion of the first heart rate signal are identified. An average heart rate signal is acquired. The accelerating portion or the decelerating portion of the first heart rate signal is replaced with the average heart rate signal to produce a second heart rate signal. A frequency spectrum of the second heart rate signal may be obtained and utilized to predict or detect one or more autonomic conditions of a patient. Therapy may initiated or titrated in response to prediction or detection of the autonomic condition. | 1. A method, comprising: acquiring a first heart rate signal that includes an indication of at least one interval duration of a heart rate of a patient;identifying at least one accelerating portion of the first heart rate signal for which the heart rate is increasing;identifying at least one decelerating portion of the first heart rate signal for which the heart rate is decreasing;determining an average heart rate signal; andreplacing, in the first heart rate signal, one of only the accelerating portion or only the decelerating portion with the average heart rate signal to produce a second heart rate signal such that the second heart rate signal is one of an accelerating signal when only the decelerating portion is replaced and a decelerating signal when only the accelerating portion is replaced. 2. The method of claim 1, wherein acquiring the first heart rate signal comprises: acquiring a cardiac electrical signal; andanalyzing the cardiac electrical signal to determine the at least one interval duration. 3. The method of claim 1, wherein acquiring the first heart rate signal comprises: acquiring the first heart rate signal based on at least one signal selected from the group consisting of: a pressure signal;an audio signal;a force signal; anda motion signal. 4. The method of claim 1, further comprising acquiring a frequency response of the second heart rate signal. 5. The method of claim 4, further comprising filtering the frequency response of the second heart rate signal to remove one or more frequencies associated with the average heart rate signal. 6. The method of claim 4, further comprising diagnosing or predicting at least one autonomic condition of a patient based on the frequency response of the second heart rate signal. 7. The method of claim 6, further comprising responding to the diagnosis or prediction of at least one autonomic condition of a patient by initiating or modifying at least one therapy selected from the group consisting of: overdrive pacing;anti-tachycardia pacing;spinal stimulation;vagal stimulation;vagal branch stimulation;AV nodal stimulation;fat pad stimulation; anddrug delivery therapy. 8. The method of claim 6, wherein diagnosing or predicting at least one autonomic condition of a patient based on the frequency response of the second heart rate signal includes estimating sympathetic or parasympathetic modulation. 9. The method of claim 6, wherein diagnosing or predicting at least one autonomic condition of a patient based on the frequency response of the second heart rate signal includes diagnosing or predicting a tachyarrhythmia in the patient. 10. The method of claim 6, wherein diagnosing or predicting at least one autonomic condition of a patient based on the frequency response of the second heart rate signal includes diagnosing or predicting the development of cardiovascular disease in the patient. 11. The method of claim 9, wherein replacing, in the first heart rate signal, one of the accelerating p... | Medtronic, Inc. | A61N001/00
US08467863 | B2 | 2013-06-18 | Sensing apparatus for sensing an object | 2009-08-20 | A sensing apparatus for sensing an object includes an arrangement of sensing elements for sensing a property of the object. The sensing elements are operable in a contact mode, in which a sensing is performable, while the sensing elements are in contact with the object, and in a non-contact mode, in which a sensing is performable, while the sensing elements are not in contact with the object. The sensing apparatus further includes a mode determination unit for determining whether or not a sensing element is in contact with the object. | 1. A sensing apparatus for sensing an object, comprising: an arrangement of sensing elements for sensing a property of the object and acquiring electrical signals, wherein the sensing elements are operable in a contact mode, in which a sensing is performable, while the sensing elements are in contact with the object, and in a non-contact mode, in which a sensing is performable, while the sensing elements are not in contact with the object; andan electrical signal processing unit configured to perform both a non-contact mode filtering of the electrical signals to retain only frequencies in a first range if the sensing elements are in the non-contact mode, and to perform a contact mode filtering of the electrical signals to retain only frequencies in a second range if the sensing elements are in the contact mode, wherein the first range is different from the second range. 2. The sensing apparatus as claimed in claim 1, further comprising a mode determination unit for determining whether a sensing element is in contact with the object not. 3. The sensing apparatus as claimed in claim 1 further comprising a detection unit, wherein the sensing elements are adapted to acquire an electrical signal, which depends on the property to be determined, by performing a unipolar acquisition by the detection unit, if the sensing elements are in the non-contact mode, and to acquire an electrical signal, which depends on the electrical potential, by performing a unipolar or bipolar acquisition by the detection unit, if the sensing elements are in the contact mode. 4. The sensing apparatus as claimed in claim 1, further comprising a tube for containing at least the arrangement of sensing elements for guiding the sensing elements in the interior of the object for sensing the interior of the object. 5. The sensing apparatus as claimed in claim 1, further comprising a holding structure for holding the arrangement of sensing elements, wherein the sensing elements are arranged on the holding structure, wherein the holding structure is adjustable between a folded condition and at least one unfolded condition and wherein the holding structure comprises an elongated shape in the folded condition. 6. The sensing apparatus as claimed in claim 5, wherein the holding structure is adapted such and the sensing elements are arranged such that the spacing between the sensing elements in a first unfolded condition is different to the spacing between the sensing elements in a second unfolded condition. 7. The sensing apparatus as claimed in claim 1, wherein the sensing elements are adapted to apply energy and to receive energy. 8. The sensing apparatus as claimed in claim 1, further comprising markers for being visualized in an imaging apparatus. 9. The sensing apparatus as claimed in claim 8, further comprising a registering unit for registering the sensing elements and a model of the object. 10. The sensing apparatus of claim 1, wherein the first range is from 0.1 to 300 H... | Koninklijke Philips N.V. | A61B005/0428
US08758257 | B2 | 2014-06-24 | Instrument including a movement sensor for positioning an effective portion and method of using same | 2010-12-17 | An instrument positionable with respect to a body valve in the body of a living subject for performing a predetermined procedure on the living subject at a predetermined location relative to the body valve. The instrument includes: a movement sensor for sensing movements of the body valve; an effective portion for performing the predetermined procedure when the effective portion is at the predetermined location; and an effective portion positioner for positioning the effective portion in the living subject. The effective portion is positionable using the effective portion positioner in response to the movements sensed by the movement sensor for positioning the effective portion at the predetermined location in order to perform the predetermined procedure. | 1. An instrument positionable with respect to a body valve in the body of a living subject for implanting anchors into living tissue at a predetermined location relative to said body valve, said instrument comprising: a movement sensor for sensing movements of said body valve;an effective portion for implanting anchors when said effective portion is at said predetermined location, said effective portion including an anchor implanting mechanism for implanting said anchors at said predetermined location; andan effective portion positioner for positioning said effective portion in said living subject;whereby said effective portion is positionable using said effective portion positioner in response to said movements sensed by said movement sensor for positioning said effective portion at said predetermined location in order to use said anchor implanting mechanism for implanting said anchors at said predetermined location;wherein said movement sensor includes a deformable element having a configuration, a rigidity and dimensions such that said deformable element is substantially freely deformable by said valve when abutting against said valve; anda deformation sensor for sensing deformations of said deformable element. 2. An instrument as defined in claim 1, wherein said deformable element is deformable in a predetermined plane. 3. An instrument as defined in claim 1, wherein said deformable element is substantially elongated, said deformable element defining a deformable element proximal end and a substantially opposed deformable element distal end, said deformable element proximal end being fixed with respect to said effective portion and located substantially adjacent to said effective portion and said deformable element distal end being substantially freely movable relative to said effective portion. 4. An instrument as defined in claim 3, wherein said deformable element includes a series of deformable element segments articulated to each other and said deformation sensor includes segments movement sensors for sensing movements of said deformable element segments relative to each other. 5. An instrument as defined in claim 1, wherein said deformation sensor includes a strain sensor for sensing strain in said deformable element. 6. An instrument as defined in claim 1, wherein said deformation sensor includes an electrical sensor for sensing electric activity in adjacent tissue. 7. An instrument as defined in claim 1, wherein said deformable element is hingedly attached to the remainder of said instrument. 8. An instrument as defined in claim 1, wherein said instrument is substantially elongated and defines an instrument proximal end section and a substantially longitudinally opposed instrument distal end section, said effective portion and said movement sensor being provided substantially adjacent to said instrument distal end section, and said effective portion positioner being provided, at least in part, substantially adjacent to said inst... | - | A61B005/02
US08500637 | B2 | 2013-08-06 | Viscosimetric biosensor for monitoring analyte levels | 2009-07-14 | The invention relates to a viscosimetric sensor for monitoring an analyte level in a patient or in body fluids ex vivo. In particular, the invention relates to a sensor for monitoring analytes, such as glucose levels in a patient. The sensor comprises at least two hermetically closed chambers (2a, 2b), including an activation chamber and a measuring chamber, containing a sensitive fluid (3), the chambers communicating with each other through at least one microchannel (5), and a semi-permeable barrier (4) allowing the analyte to enter or leave the chambers. Each said chamber comprises at least one flexible membrane (6, 7) configured for antagonistic variations of the volume of the chambers. | 1. A viscosimetric biosensor suitable for the measurement of an analyte in a fluid surrounding or on the sensor, including at least two hermetically closed chambers, including an actuation chamber and a sensing chamber, containing a sensitive fluid responsive to the analyte to be measured, the chambers communicating with each other through at least one channel,semi-permeable membrane allowing the analyte to enter or leave the chambers,each said chamber comprising a flexible membrane configured for antagonistic variations of the volume of the chambers,at least one displacement sensing element configured to measure the displacement of at least one of said flexible membranes, andat least one actuating element configured to actuate displacement of at least one of said flexible membranes. 2. A viscosimetric biosensor according to claim 1 wherein the two chambers are hermetically sealed by the semi-permeable membrane on a side of the chambers opposite the flexible membranes. 3. A viscosimetric biosensor according to claim 1 wherein the at least one channel is located directly adjacent the semi-permeable membrane. 4. A viscosimetric biosensor according to claim 1, wherein the actuating element for actuating the flexible membrane of the actuation chamber is a piezoelectric element. 5. A viscosimetric biosensor according to claim 1, wherein the sensing element for sensing the movement of the flexible membrane of the sensing chamber is a piezoelectric element. 6. A viscosimetric biosensor according to claim 1, wherein the actuating element for actuating the flexible membrane of the actuation chamber is a MEMS device. 7. A viscosimetric biosensor according to claim 1, wherein the sensing element for sensing the movement of the flexible membrane of the sensing chamber induced by the variation of the volume of the sensitive fluid in the actuation chamber is a MEMS device. 8. A viscosimetric biosensor according to claim 6, wherein the MEMS device comprises a suspended gate MOSFET having an electrode on the flexible membrane acting as the suspended gate. 9. A unitary biosensing device for the measurement of a plurality of analytes in a fluid surrounding or on the sensor, comprising a plurality of viscosimetric biosensors portions, each biosensor portion being configured to respond to a specific respective one of the plurality of analytes to be measured, each biosensor portion comprising: at least two hermetically closed chambers each defining a variable volume, including an activation chamber and a measuring chamber, containing a sensitive fluid responsive to the specific analyte to be measured, the chambers communicating with each other through at least one channel;a semi-permeable membrane allowing the analyte to enter or leave the chambers;each said chamber comprising a flexible membrane configured for antagonistic variations of the volume of the chambers;at least one displacement sensing element configured to measure the displacement of at least one... | Ecole Polytechnique Federale de Lausanne (EPFL) | A61B005/00 G01N011/00 G01N033/48
US08753286 | B2 | 2014-06-17 | Spirometer with replaceable flow tube | 2010-12-23 | The invention relates to a spirometer for measuring pulmonary respiration of a test subject, comprising a body and a sensor unit, wherein the sensor unit can be interchanged with further sensor units. In a corresponding method, such a spirometer is operated with a sensor unit and the sensor unit is subsequently interchangeably replaced by a further sensor unit. | 1. A spirometer for measuring pulmonary respiration of a test subject, comprising: a sensor unit, the sensor unit comprising a flow tube,a mass flow sensor disposed within the flow tube, the mass flow sensor being adapted to sense a mass flow of air flowing through the flow tube, anda sensor identification code; anda body, adapted to removably attach the sensor unit thereto, the body comprising a plurality of ambient air sensors adapted to sense a plurality of ambient air characteristics, the ambient air sensors being operatively connected to an analysis unit,wherein the mass flow sensor becomes operatively connected to the analysis unit when the sensor unit is attached to the body, andwherein the mass flow sensor comprises a first temperature element that is exposed to the mass flow of air flowing through the flow tube, andwherein during operation of the spirometer the first temperature element is heated to a target temperature, andwherein the mass flow of air flowing through the flow tube is derived from an amount of power that is required to maintain the target temperature of the first temperature element while being exposed to the mass flow of air flowing through the flow tube, andwherein the sensor unit is identified by reading the sensor identification code through a data connection between the sensor unit and the body when the sensor unit is attached to the body, andwherein the pulmonary respiration of the test subject is determined in the analysis unit by evaluating the mass flow of air flowing through the flow tube sensed by the mass flow sensor and the plurality of ambient air characteristics sensed by the plurality of ambient air sensors. 2. The spirometer as in claim 1, wherein the plurality of ambient air sensors comprises a temperature sensor adapted to sense ambient air temperature, a humidity sensor adapted to sense ambient air humidity, and a pressure sensor adapted to sense ambient air pressure. 3. The spirometer as in claim 1, further comprising a second temperature element adapted to sense a temperature of air in the flow tube, wherein the target temperature of the first temperature element is a predetermined delta value above the temperature sensed by the second temperature element. 4. The spirometer as in claim 1, wherein the sensor unit is disposable and the body is reusable. 5. The spirometer as in claim 1, wherein the sensor unit is standardized with respect to a set measurement standard. 6. The spirometer as in claim 1, wherein the sensor unit is standardized with respect to a lung volume of the test subject or to a set mass flow. 7. The spirometer as in claim 2, wherein the pulmonary respiration of the test subject determined in the analysis unit is a volume flow of air flowing through the flow tube, or a volume of air flowing through the flow tube over time, or both as a function of the mass flow of air flowing through the flow tube, ambient air temperature, ambient air humidity, and ambient air pressure. 8. The... | sendsor GmbH | A61B005/085 A61B005/09 A61B005/097
US08423122 | B2 | 2013-04-16 | Localization of capsule with a synthetic source of quadrupoles and dipoles | 2009-07-06 | This invention relates to methods and apparatus for localizing an in vivo imaging device by means of a single magnetic source coil assembly and a single magnetic detector coil assembly. This invention also relates to methods and apparatus to enable the use of a single magnetic source coil assembly to also transmit and receive information, images, and controls signals, as well as for the coil assembly to be used in DC-DC voltage conversion. A user interface with a display provides the user with the option of viewing selected images captured by the in vivo imaging device. | 1. A system comprising: an encapsulated in-vivo imaging device, comprising a single electromagnetic sensor coil for determining localization of said in vivo imaging device in a vicinity of an electromagnetic field, said localization computed based on determining an electrical signal across said electromagnetic sensor coil, said signal being induced by said electromagnetic field;a set of quadrupole electromagnetic radiator source coils to radiate said electromagnetic field, said electromagnetic field being sensed by said electromagnetic sensor coil, and said set of electromagnetic radiator source coils being located outside of said encapsulated in-vivo imaging device;wherein said set of quadrupole electromagnetic radiator source coils comprises three mutually orthogonal pairs of parallel coils, and wherein a distance between said parallel coils within said pair is substantially smaller than a distance between said electromagnetic radiator source coils and said electromagnetic sensor coil. 2. The system of claim 1, further comprising an analog-to-digital converter, wherein said converter is used for detecting an electrical signal generated by said radiated electromagnetic field across said electromagnetic sensor coil within said in-vivo imaging device for purposes of localization of said device. 3. The system of claim 1, further comprising a DC-DC converter, wherein said converter is used for delivering power to said in-vivo device. 4. The system of claim 3, wherein said electromagnetic sensor coil is part of said DC-DC converter. 5. The system of claim 1, further comprising a transceiver, wherein said transceiver is to transmit information from and receive information by said in-vivo imaging device. 6. The system of claim 5, wherein said electromagnetic sensor coil is used as an antenna by said transceiver for transmitting information from and receiving information by said in-vivo imaging device. 7. The system of claim 1, wherein said encapsulated in-vivo imaging device further comprising an analog-to-digital converter, a DC-DC converter, a transceiver and means for switching between said coil, and said analog-to-digital converter, said DC-DC converter, and said transceiver within said in-vivo imaging device. 8. The system of claim 1, further comprising a power source to deliver power to said in-vivo device. 9. The system of claim 1, further comprising means for processing information generated in said in-vivo device. 10. The system of claim 1, further comprising a processor and interface circuitry means for controlling said in-vivo device. 11. The system of claim 1, further comprising illumination means for illuminating a body lumen being imaged by said in-vivo device. 12. The system of claim 1, further comprising an imager. 13. The system of claim 1, further comprising a transceiver to transmit data to and receive data from said in-vivo imaging device, said transceiver located outside a body lumen or a cavity being imaged. 14. The s... | Given Imaging Ltd. | A61B005/05 G01S005/04
US08635053 | B2 | 2014-01-21 | Impedance analysis for esophageal maladies | 2010-10-19 | A system and method for automatically analyzing impedance and pH data from an esophageal probe includes a data collection system that collects and stores the output from the sensors for a certain period of time to locate reflux episodes in the waveforms. The data analysis system uses wavelet analysis to assist in locating bolus entry and exit points in the waveforms and to smooth waveforms for additional analysis. It also distinguishes between candidate acid reflux episodes and candidate non-acid reflux episodes for determining actual acid reflux episodes and non-acid reflux episodes, and it distinguishes between, and provides different signal processing for, signals from patients that have healthy esophageal tissue and those that have diseased esophageal tissue. | 1. An automated method of identifying and classifying reflux episodes in data comprising a plurality of impedance versus time measurements obtained simultaneously at a plurality of different locations spaced apart from each other along a length of an esophagus between the pharynx and the stomach collected and stored in a collection device, wherein the data is indicative of impedance versus time waveforms, including proximal waveforms from locations closest to the pharynx, distal waveforms from locations closest to the stomach, and mid-channel waveforms between the proximal waveforms and the distal waveforms, said method comprising: providing the impedance versus time data to a data analysis device, which performs wavelet transformation to the impedance versus time waveforms to obtain smoothed curves of the impedance versus time waveforms and detail curves showing local positive and negative maxima peaks corresponding to positive and negative slopes in the smoothed impedance versus time waveforms;determining with the data analysis device whether the smoothed impedance versus time waveforms exhibit either normal baseline impedance levels that are characteristic of conductivity of healthy esophageal tissue or low baseline impedance levels that are characteristic of conductivity of diseased esophageal tissue;determining with the data analysis device whether change in impedance in the smoothed impedance versus time waveforms during gastrointestinal tract events are either: (i) large as indicated by the ratio of the mean distribution of minimum impedance to baseline impedance levels being not greater than a value; or (ii) low as indicated by the ratio of the mean distribution of minimum impedance to baseline impedance levels being greater than the value;if it is determined by the data analysis device that the smoothed impedance versus time waveforms exhibit low baseline levels and that the change in impedance during gastrointestinal tract events are low, then applying an adaptive filter with the data analysis device to at least one of the smoothed impedance versus time waveforms, wherein a primary input for the adaptive filter comprises a delayed sequence of the said at least one smoothed impedance versus time waveform, and determining a candidate episode by transition of impedance level in the adaptive filter output to a gastric contents impedance level; andif it is determined by the data analysis device that either the smoothed impedance versus time waveforms do not exhibit low baseline levels or the change in impedance during gastrointestinal tract events are not low, then determining a candidate episode by transition of impedance level in the smoothed waveforms to a gastric contents impedance level. 2. The method of claim 1, wherein the value is 0.5. 3. The method of claim 1, wherein the adaptive filter is applied to at least one mid-channel waveform. 4. The method of claim 1, wherein the data in the collection device also comprises data that is... | Sandhill Scientific, Inc. | G06G007/48
US09392966 | B2 | 2016-07-19 | Fall prevention system | 2009-08-27 | There is provided a system for fall prevention for a user, comprising one or more sensors for attachment to respective portions of the body, each sensor being adapted to measure movement of the respective portion of the body and to translate the movement into a signal; and a controller adapted to receive the signal or signals and to determine a risk of falling by estimating a trajectory of the center of mass of the user relative to a trajectory of a lower portion of the body from the signal or signals. | 1. A system for fall prevention for a user, comprising: one or more sensors configured to be attached to respective portions of the body of a walking user to measure movements of a moving center of mass of the user and a moving base of support of the body and to translate the measured movements into signals, the base of support of the user being a surface defined by a step width and a step length of the user when walking; anda controller configured to receive the signals and to determine a risk of falling of the walking user by estimating a difference between a trajectory of the moving center of mass of the user relative to a trajectory of the moving base of support of the body from the signals. 2. The system as claimed in claim 1, wherein the controller is further configured to: determine whether the signals have a reoccurring pattern over time in multi-dimensional space; andbased on periodicity of the reoccurring pattern, determining whether the user is in a walking steady state. 3. The system as claimed in claim 2, wherein the controller is further configured to: determine periodicity of the signals when the user is in the walking steady state by detecting peaks in at least one of the movement signals to delineate individual steps. 4. The system as claimed in claim 3, wherein the controller is further configured to: from the signals for each step, determining (1) the step size and the step width and (2) a location of the center of mass; anddetermining whether the center of mass is aligned within the base of support. 5. The system as claimed in claim 4, wherein the controller is further configured to: from the signals, determine a velocity of the center of mass and whether the center of mass is within the base of support; anddetermining the risk of falling in response to one of: the center of mass being outside the base of support and having a velocity which is insufficient to bring the center of mass back within the base of support, orthe center of mass is within the base of support and the velocity is outwards from the base of support. 6. The system as claimed in claim 1, wherein the controller is further configured to: from the signals, determine a velocity of the center of mass and whether the center of mass is within the base of support; anddetermining the risk of falling in response to one of: the center of mass being outside the base of support and having a velocity which is insufficient to bring the center of mass back within the base of support, andthe center of mass is within the base of support and the velocity is outwards from the base of support. 7. The system as claimed in claim 1, wherein the one or more sensors include accelerometers, and the trajectories are estimated by integrating the signals from the accelerometers. 8. The system as claimed in claim 1, wherein the one or more sensors include accelerometers, and the controller is configured to estimate the trajectory of the center of mass of the user relative to the tr... | KONINKLIJKE PHILIPS N.V. | A61B005/103 A61B005/117 A61B005/11 A61B005/00 G08B021/04 A63B024/00
US09907962 | B2 | 2018-03-06 | Arrhythmia prediction based on heart rate turbulence | 2009-10-29 | In general, the disclosure describes techniques for predicting the occurrence of an arrhythmia based on an indication of heart rate turbulence. An example method comprises sensing a parameter indicative of heart rate turbulence, measuring heart rate turbulence based on the sensed parameter, and predicting an occurrence of an arrhythmia based on the measured heart rate turbulence. | 1. A method comprising: sensing cardiac contractility of a heart;deriving a plurality of values of an interval between cardiac contractions of the heart from the sensed cardiac contractility;measuring heart rate turbulence based on the interval values derived from the sensed cardiac contractility;predicting an occurrence of a tachyarrhythmia based on the measured heart rate turbulence; anddelivering a cardiac stimulation therapy to the heart in response to predicting the occurrence of the tachyarrhythmia, wherein the cardiac stimulation therapy is configured to prevent the occurrence of the tachyarrhythmia. 2. The method of claim 1, further comprising charging a therapy delivery circuit in response to predicting the occurrence of the tachyarrhythmia. 3. The method of claim 2, wherein the therapy delivery circuit comprises a circuit for delivery of a cardioversion. 4. The method of claim 3, wherein charging of the therapy delivery circuit occurs in response to the predicting of the occurrence of the tachyarrhythmia. 5. The method of claim 1, further comprising identifying an abnormal heartbeat, wherein measuring heart rate turbulence comprises measuring heart rate turbulence resulting from the abnormal heartbeat. 6. The method of claim 5, further comprising delivering a signal configured to induce the abnormal heartbeat, wherein identifying the abnormal heart beat comprises identifying the abnormal heartbeat induced by the signal. 7. The method of claim 5, wherein the abnormal heartbeat comprises at least one of a premature atrial contraction or a premature ventricular contraction. 8. The method of claim 5, wherein measuring heart rate turbulence comprises calculating at least one of turbulence onset or turbulence slope. 9. The method of claim 5, further comprising comparing the measured heart rate turbulence to a threshold value, wherein predicting the occurrence of the tachyarrhythmia comprises predicting the occurrence of the tachyarrhythmia based on the comparison. 10. The method of claim 5, wherein identification of the abnormal heartbeat comprises identification of a naturally occurring heartbeat. 11. The method of claim 1, wherein the sensed cardiac contractility is sensed using one or more of: a sensor intracardiac or intravascular pressure, an accelerometer or other sensor capable of detecting heart or blood sounds, vibrations, or motion, an intra-cardiac or intra-arterial impedance, optical or ultrasonic sensor capable of detecting changes in blood flow associated with cardiac contractions, or an optical sensor capable of detecting oxygen saturation changes associated with cardiac contractions. 12. The method of claim 1 wherein the cardiac contractility is sensed using a sensor responsive to physical contractions of the heart. 13. The method of claim 1, wherein the therapy is configured to reduce an effect of the tachyarrhythmia or terminate the tachyarrhythmia if the tachyarrhythmia occurs. 14. A system comprising: a sensin... | Medtronic, Inc. | A61N001/365 A61B005/02 A61N001/362 A61B005/024 A61B005/0402 A61B005/046 A61B005/00
US20110144527 | A1 | 2011-06-16 | Artificial Airway with Integrated Core Temperature Monitor | 2009-12-10 | This is a device for use in clinical, surgical intro- and post-operative patient core temperature monitoring, which utilizes an artificial airway to integrate a temperature probe, removing the need for an external cord for connecting to an external display unit, or the traditional way of an esophageal tube that necessitates more medical procedures. Consequently, for patients under general anesthesia with ET Tube or LMA, accurate monitoring of core temperature becomes feasible. As a result, the overall cost for surgical operation, including that of anesthesia, is reduced and the with reduced risk of complication associated with extra oral procedural for inserting an esophageal tube to measure temperature. | 1. An artificial airway with core temperature monitor, comprising: An in-body portion having an air tube containing an inside end and an outside end;A temperature probe located near the tip of said inside end of said air tube;A signal wire inside said air tube, connecting said temperature probe to a connecting socket near said outside end;An outside body portion having a temperature display unit having a connector for attaching to the connecting socket on said air tube; and,An air sac located near the inside end of said air tube and an air pump connect to said air sac, whereby said air pump can be operated to inflate said air sac by a pumping tube connected to said air sac. 2. The device of claim 1, wherein said signal wire further having a probing end connected to said temperature display unit and an external end that contains a wire connecting socket with a connector for connecting to either the connector of said temperature display unit or an interface for an external display. 3. The device of claim 1, wherein said temperature probe is embedded in the inside wall of said air tube. 4. The device of claim 1, wherein said temperature probe is embedded on the sac wall of said air sac. 5. The device of claim 1, wherein said temperature display unit further having a digital display readout, battery compartment and a switch in association with a printed circuit board. 6. An artificial airway with core temperature monitor, comprising: An in-body portion having an air tube containing an inside end and an outside end;A temperature probe located near the tip of said inside end of said air tube;An air sac located near the inside end of said air tube and an air pump connected to said air sac, whereby said air pump can be operated to inflate said air sac by a pumping tube connected to said air sac; and,A signal wire inside said air tube, connecting said temperature probe to a connecting socket near said outside end, whereby said connecting socket allows the temperature sensing and reading to be sent to an external monitor by an external cord. | - | A61B005/01 A61M016/04
US20110152695 | A1 | 2011-06-23 | System for Processing Exercise-Related Data | 2010-12-07 | A method, apparatuses, and a system for carrying out exercise-related processing are described. The apparatus includes an exercise-measurement circuitry configured to measure exercise-related measurement data related to a user carrying out an exercise. The apparatus further comprises a communication circuitry configured to provide the portable apparatus with wireless communication capability. The apparatus further comprises a processing circuitry configured to receive the exercise-related measurement data from the exercise-measurement circuitry, to process the exercise-related measurement data, and to communicate with an external user interface apparatus over a bidirectional wireless communication connection through the communication circuitry so as to cause transmission of the processed exercise-related measurement data to the user interface apparatus and to receive configuration data from the user interface apparatus. | 1. A portable apparatus comprising: an exercise-measurement circuitry configured to measure exercise-related measurement data related to a user carrying out an exercise;a communication circuitry configured to provide the portable apparatus with wireless communication capability; anda processing circuitry configured to:receive the exercise-related measurement data from the exercise-measurement circuitry,receive configuration data from an external user interface apparatus over a bidirectional wireless communication connection established through the communication circuitry, wherein the received configuration data includes exercise-related parameters;process the exercise-related measurement data according to the received exercise-related parameters in order to obtain advanced exercise-related data, andcommunicate the advanced exercise-related measurement data to the user interface apparatus over the bidirectional wireless communication connection. 2. The apparatus of claim 1, wherein the advanced exercise-related data processed by the processing circuitry comprises energy expenditure data. 3. The apparatus of claim 1, wherein the advanced exercise-related data processed by the processing circuitry comprises performance capacity data. 4. The apparatus claim 1, wherein the processing circuitry is further configured to process the exercise-related measurement data on the basis of the configuration data received from the user interface apparatus, wherein the configuration data comprises user-related parameters describing characteristics of the user. 5. The apparatus claim 1, wherein the processing circuitry is further configured to process the exercise-related measurement data on the basis of the configuration data received from the user interface apparatus, wherein the configuration data comprises exercise-related parameters defining instructions or guidance for an exercise to be conducted. 6. The apparatus of preceding claim 1, further comprising at least one of the following sensors: a thermodynamic sensor, a motion sensor, wherein the received configuration data includes exercise-related parameters and wherein the processing circuitry is further configured to process measurement data received from the at least one of the thermodynamic sensor and motion sensor by calculating the advanced exercise-related data from the received measurement data on the basis of the received configuration data. 7. The apparatus of preceding claim 1, wherein the communication circuitry is further configured to establish a first wireless communication link and a second wireless communication link to the user interface apparatus, and tooperate the first wireless communication link and the second wireless communication link in parallel by using the first wireless communication link for controlling the operation of the second wireless communication link. 8. The apparatus of claim 1, further comprising: a memory unit configured to store audio tracks;an audio decoder co... | POLAR ELECTRO OY | A61B005/02 A61B005/00 A61B005/11
US07955276 | B2 | 2011-06-07 | Capsule endoscope | 2008-12-12 | A capsule endoscope is disclosed that includes a capsule and a transparent cover having a central axis and that is located in front of an objective optical system within the capsule. The objective optical system has a depth of field, a field of view, and an image surface that has curvature of field, and is arranged on the central axis of the transparent cover and the capsule. The field of view of the objective optical system includes a spherical surface that is tangent to the transparent cover at a periphery of the field of view, and the near point of the depth of field on the optical axis of the objective optical system is positioned farther from the objective optical system than is the transparent cover. Various conditions are satisfied so as to provide high quality imaging. | 1. A capsule endoscope comprising: a capsule having a central axis;an illumination light source;an objective optical system that forms images of a region illuminated by the illumination light emitted from the illumination light source, the objective optical system having a depth of field, a field of view, and an image surface that has curvature of field such that, the higher the image height becomes, the greater the image surface tilts toward the object side;an image pickup device that receives the images formed by the objective optical system; anda transparent cover having a central axis and located in front of the objective optical system; whereinthe objective optical system is arranged on the central axis of the transparent cover and the capsule;at the periphery of the field of view, the transparent cover has a radius of curvature substantially equal to one-half the outside diameter of the capsule; andthe near point of the depth of field along the optical axis of the objective optical system is positioned farther from the transparent cover than is the distance along the optical axis of the objective optical system from the transparent cover to a hypothetical hemispherical surface having a radius equal to one-half the outside diameter of the capsule and that is positioned to enclose the capsule in front of the objective optical system. 2. The capsule endoscope according to claim 1, wherein the image pickup device is arranged so that an image receiving surface thereof is positioned at a point that is conjugate, with regard to the objective optical system, to the intersection point of the hypothetical hemispherical surface and the optical axis of the objective optical system. | Olympus Corporation | A61B005/103
US20110160554 | A1 | 2011-06-30 | DEVICE AND METHOD FOR DETERMINING AT LEAST ONE CHARACTERIZING PARAMETER OF MULTILAYER BODY TISSUE | 2008-06-18 | A device for the non-invasive measurement of a glucose level, body hydration or another characterizing parameter of body tissue comprises at least two coplanar waveguides arranged on a common support. An AC signal is applied to the first ends of the coplanar waveguides, and the signal arriving at the second end is measured. The coplanar waveguides have differing gap widths, such that their electric fields have different reach into the body tissue. This allows obtain depth resolved information about the permittivities of individual tissue layers and to obtain more accurate results. | 1. A device for determining at least one characterizing parameter P of living body tissue, in particular a glucose level or water content, comprising a number N>1 of coplanar waveguides, each coplanar waveguide comprising a center strip electrode between ground electrodes, wherein at least some of said coplanar waveguides have different gap widths between their center strip electrode and their ground electrodes for generating electrical fields of different reach,a signal generator generating at least one AC signal, wherein first ends of said coplanar waveguides are connected to said signal generator,a measuring unit, wherein a second ends of said coplanar waveguides are connected to said measuring unit for measuring N measured parameters mi,a control unit for determining said characterizing parameter P from at least part of said measured parameters mi. 2. The device of claim 1, comprising at least one coplanar waveguide having a gap width of 100 μm or less. 3. The device of claim 1, comprising at least one co-planar waveguide having a gap width of at least 1 mm, in particular between 1 and 4 mm. 4. The device of claim 1, comprising at least one co-planar waveguide having a gap width of at least 4 mm. 5. The device of claim 1, wherein said signal generator generates an AC signal having a frequency of at least 50 MHz. 6. The device of claim 1, wherein said ground electrodes are wider than said center strip electrodes. 7. The device of claim 1, further comprising a non-conductive cover layer covering said electrodes. 8. Use of the device of claim 1, for measuring a glucose level. 9. Use of the device of claim 1, for measuring a water content. 10. A method for determining at least one characterizing parameter P of body tissue, in particular a glucose level or water content, comprising, applying a number N>1 of coplanar waveguides to a skin region of said body tissue, each coplanar waveguide comprising a center strip electrode between ground electrodes, wherein at least some of said coplanar waveguides have different distances between their center strip electrode and their ground electrodes generating electrical fields of different reach by means of said coplanar waveguides by applying an AC signal to a first end of each coplanar waveguide,measuring N measured parameters mi depending on a signal exiting from a second end of each coplanar waveguide,determining said characterizing parameter P from at least part of said measured parameters mi. 11. The method of claim 10, wherein at least a first of the measured parameters mi1 of a layer i1 depends on the characterizing parameter P as well as on a further parameter Q and at least a second of the measured parameters mi2 of a layer i2 above said layer i1 depends on said further parameter Q and to a lesser degree on the characterizing parameter P, said method further comprising the step of combining the measured parameters mi1 and mi2 for obtaining a calculated value (∈i1) that depends less on the ... | - | A61B005/145
US20110137181 | A1 | 2011-06-09 | Heart pulse detector | 2009-12-08 | A heart pulse detector includes a light emitting diode for emitting a short pulse light toward body tissue and a photo sensor for receiving light signal modulated by blood pulse flow through body tissue. The output current of photo sensor is charging a capacitor during pulse light and is discharging the same capacitor during no pulse light for same period of pulse light. The ambient light signal is then cancelled out from the detected light signal. The light current signal is transferred to a voltage signal after sample and hold procession. Finally heart pulse signal is detected out from the voltage signal without the interference of ambient light. | 1. A heart pulse detector for detecting the heart pulse signal from blood pulse signal of human body, comprising: timing generation means for providing periodically a first pulse with period in the range between 0.5 milliseconds to 20 milliseconds, and a second pulse with active interval of duration t2 in the range between 20 microseconds to 200 microseconds during non-active interval of said first pulse, and a third pulse with active interval of duration t2 during non-active interval of said first pulse and the start of active interval of said third pulse is right after the end of active interval of said second pulse, and a fourth pulse with active interval of duration t4 in the range between 20 microseconds to 200 microseconds during non-active interval of said first pulse and non-active interval of said second pulse and the start of active interval of said fourth pulse is after the end of active interval of said third pulse;a light emitting diode for emitting a pulsed light toward body tissue during the active intervals of said second pulse;a photo sensor for providing a sensor current signal in response to a received light signal which sourced from said pulsed light and ambient light and modulated by blood pulse flowing through body tissue;a first capacitor;a first switch connected to the two terminals of said first capacitor for discharging said first capacitor during the active intervals of said first pulse to a zero charge;signal detection means for charging charge corresponding to said sensor current signal into said first capacitor during the active intervals of said second pulse and for discharging charge corresponding to said sensor current signal from said first capacitor during the active intervals of said third pulse to cancel the ambient light signal related charge previously stored in said first capacitor and for providing a first voltage signal corresponding to the blood pulse signal related charge stored in said first capacitor after the ambient light signal is cancelled;a second capacitor;signal sampling means for storing said first voltage signal on said second capacitor during the active intervals of said fourth pulse and for providing a second voltage signal proportional to the voltage containing blood pulse signal stored on said second capacitor;means for providing a heart pulse signal in response to blood pulse signal contained in said second voltage signal to the output of said heart pulse detector;a power regulator for providing a constant voltage source to the circuits of said heart pulse detector. 2. A heart pulse detector as claimed in claim 1, wherein said signal detection means is comprised of a first operational amplifier having negative input terminal connected to a voltage reference source;a first p-channel field-effect transistor having source terminal connected to said constant voltage source and gate terminal connected to the output terminal of said first operational amplifier and drain terminal connected to ... | Holylite Microelectronics Corp. | A61B005/024
US20110152648 | A1 | 2011-06-23 | PULSE OXIMETER SENSOR | 2009-12-21 | A method of sensing or measuring various characteristics of blood is disclosed. The method includes providing a sensor having an emitter and a receiver coupled to external processing apparatus. The sensor has a generally elongated configuration designed to be formed into a generally U-shaped orientation so that the sensor can be positioned over the skin of a patient adjacent the patient's mouth with the emitter and receiver in an aligned orientation on opposite sides of the skin. One of the emitter and receiver is positioned against the skin outside of the mouth and the other is positioned against the skin on the inside of the mouth. The sensor is used to sense or measure at least one selected characteristic of the blood of the patient. | 1. A method of sensing or measuring various characteristics of blood including the steps of: providing a sensor including an emitter and a receiver coupled to external processing apparatus, the sensor having a generally elongated configuration designed to be formed into a generally U-shaped orientation;positioning the sensor over the skin of a patient adjacent the patient's mouth with the emitter and receiver in an aligned orientation on opposite sides of the skin, one of the emitter and receiver against the skin outside of the mouth and the other of the receiver and the emitter against the skin on the inside of the mouth; andusing the sensor to sense or measure at least one selected characteristic of the blood of the patient. 2. A method as claimed in claim 1 wherein the step of providing a sensor includes providing a sensor with a generally elongated thin body designed to be formed into the generally U-shaped orientation. 3. A method as claimed in claim 2 wherein the step of providing the sensor with the generally elongated thin body designed to be formed into the generally U-shaped orientation includes providing a body including a thin formable base layer, a pliable padding material positioned on a surface of the formable base layer, and a smooth layer of material fixed to a surface of the pliable padding material. 4. A method as claimed in claim 1 wherein the step of providing a sensor includes providing a sensor constructed to form a spring clip and the step of positioning the sensor over the skin of the patient adjacent the patient's mouth includes grasping the skin of the patient adjacent the patient's mouth with the spring clip. 5. A method as claimed in claim 4 wherein the step of providing the sensor constructed to form a spring clip grasping the skin of a patient adjacent the patient's mouth and positioning the emitter and receiver aligned on opposite sides of the skin includes positioning the emitter against the skin outside of the mouth and the receiver against the skin on the inside of the mouth. 6. A method as claimed in claim 5 wherein the step of positioning the emitter and receiver in an aligned orientation on opposite sides of the skin includes applying a conductive material to the smooth layer of material overlying the emitter. 7. A method as claimed in claim 1 wherein the step of positioning the sensor over the skin of the patient adjacent the patient's mouth includes taping the sensor to the skin of the patient adjacent the patient's mouth. 8. A method of sensing or measuring various characteristics of blood including the steps of: providing a sensor including an emitter and a receiver coupled to external processing apparatus, providing the sensor with a generally elongated thin body including a thin formable base layer, a pliable padding material positioned on a surface of the formable base layer, and a smooth layer of material fixed to a surface of the pliable padding material, the body being designed to be formed i... | - | A61B005/1455
US20110137134 | A1 | 2011-06-09 | METHOD AND SYSTEM FOR ADMINISTERING AN ANAESTHETIC | 2008-01-17 | A method and system for objectively scoring intra-operative pain during general anaesthesia based on the patient's mean arterial pressure and heart rate. The index is used for closed-loop control of the intra-operative analgesia through adjustment of the drug infusion level according to fuzzy logic. It is further displayed along with other components of anaesthesia and important patient data on a monitoring display for presentation to medical staff. | 1. A method for displaying an indicator of a current pain level of a patient being administered an analgesic, the method comprising: providing a display device;measuring a current mean arterial pressure and heart rate of the patient;deriving the indicator from said measured current mean arterial pressure and heart rate; anddisplaying said derived indicator on said display device. 2. The method of claim 1, wherein said derived indicator is displayed on said display device as numerical data, graphical data, colour coding and combinations thereof. 3. The method of claim 1, wherein desired target values of said mean arterial pressure and said heart rate are determined prior to administering said analgesic and said deriving an indicator comprises comparing said current mean arterial pressure and heart rate to said target values using fuzzy logic rules. 4. The method of claim 1, wherein said derived indicator is defined in a range from a first level to a second level, said first level representing an excessive analgesia level and said second level representing an insufficient analgesia level. 5. The method of claim 4, wherein said range comprises a plurality of predetermined regions, at least a first one of said predetermined regions representing inadequate pain control, at least a second one of said predetermined regions representing good pain control, and at least a third one of said predetermined regions representing excellent pain control. 6. The method of claim 5, further comprising calculating a change in rate of infusion based on said derived indicator, recalculating said change in rate of infusion based on an average change in said derived indicator over time and a current value of said derived indicator, and adjusting a rate of infusion of the analgesic according to said recalculated change in rate of infusion. 7. The method of claim 6, wherein said calculating a change in rate of infusion comprises maintaining said infusion, stopping said infusion, or increasing said infusion according to said predetermined region said derived indicator lies in. 8. The method of claim 6, wherein said recalculating said change in rate of infusion comprises computing a first correction factor representative of a temporal variation of said derived indicator and a second correction factor representative of a current physiological state of the patient and applying said first and second correction factors to said calculated change in rate of infusion. 9. A system for displaying an indicator representative of a current pain level of a patient being administered an analgesic, the system comprising: a monitoring subsystem for measuring a current mean arterial pressure and heart rate of the patient and deriving the indicator from said measured mean arterial pressure and heart rate; anda display device coupled to said monitoring subsystem for displaying said derived indicator. 10. The method of claim 9, wherein said display device displays said derived indicato... | - | A61B005/00
US20110152696 | A1 | 2011-06-23 | AUDIBLE BIOFEEDBACK HEART RATE MONITOR WITH VIRTUAL COACH | 2010-12-21 | A system and associated method that provides a virtual coach to a person during an exercise routine includes a sensor used to sense the level of a physiological parameter of the exerciser, such as heart rate, a processor programmed to compare the sensed physiological parameter to a target level of the physiological parameter, and, depending on the comparison, provide coaching prompts to the exerciser. Coaching prompts include audible and visual prompts. In another feature, a visual image of a virtual coach may be made available for display to the exerciser, and in yet another feature, a video of the virtual coach saying the audio prompt may be displayed. The system allows for selection from a plurality of virtual coaches. The audio and video of virtual coaches are recordings of real people and may be produced or obtained from other sources. In one embodiment, audio, visual, and video of the virtual coach may be downloaded from the Internet. | 1. A system for monitoring a physiological parameter and coaching an exerciser during exercise, comprising: a sensor configured to sense a physiological parameter and provide a sensor signal representative of the actual level of the sensed parameter;an input device configured to provide a target level of the physiological parameter;a processor programmed to receive and compare the target level of the physiological parameter with the actual sensed level and in response to the comparison, play an audible coaching prompt indicative of the difference between the actual and the target; andan audio transducer configured to receive the played audible coaching prompt and transduce it into audio energy for listening by the exerciser. 2. The system for monitoring and coaching of claim 1 wherein the physiological parameter sensed comprises heart rate. 3. The system for monitoring and coaching of claim 1 wherein additional coaching prompts are used that vary in intensity of coaching. 4. The system for monitoring and coaching of claim 1 wherein at least one prompt comprises a recording of a voice message from an actual person. 5. The system for monitoring and coaching of claim 1 wherein the processor is further programmed to select a prompt based on whether the actual sensed parameter is higher or lower than the target rate. 6. The system for monitoring and coaching of claim 1 further comprising a memory, wherein the processor is further programmed to store in the memory the last prompt provided and in the event that a subsequent prompt is needed, the processor is further programmed to refer to the stored prompt in selecting the next prompt. 7. The system for monitoring and coaching of claim 1 further comprising a memory, wherein the memory has stored a plurality of prerecorded audio prompts from which the processor is programmed to select in playing an audio coaching prompt, the stored prompts having a predetermined sequence for playing, the processor being programmed so as to retrieve and play the prompts in the order of that predetermined sequence. 8. The system for monitoring and coaching of claim 1 further comprising a memory and a display, wherein the memory has stored a visual image of a coach face, and wherein the processor is further programmed to retrieve that visual image from the memory and output that image to the display when playing an audio prompt. 9. The system for monitoring and coaching of claim 1 further comprising a memory and a display, wherein the memory has stored a video of a coach providing the audio prompt played by the processor, the processor being further programmed to retrieve that video from the memory and output that image to the display when playing an audio prompt. 10. The system for monitoring and coaching of claim 9 further comprising a memory and an input port, wherein the processor is further programmed to receive a live video prompt at the input port, record the live prompt as a video file, and store that recor... | Hall Ryan Laboratories, Inc. | G09B019/00 A61B005/024
US07957781 | B2 | 2011-06-07 | Method and circuit for storing and providing historical physiological data | 2006-06-02 | Embodiments of the present invention include systems and methods that relate to pulse oximetry. Specifically, one embodiment includes an oximeter sensor comprising a light emitting element configured to emit light, a light detector configured to receive the light, and an accelerometer configured to detect motion of the oximeter sensor. | 1. An oximeter sensor, comprising: a light emitting element configured to emit light;a light detector configured to receive the light;an accelerometer configured to detect motion of the oximeter sensor; anda memory storing an algorithm configured to correct physio-optic effects from motion detected by the accelerometer. 2. The oximeter sensor of claim 1, wherein the accelerometer comprises a MEMS accelerometer incorporated into a chip within the oximeter sensor. 3. An oximeter system, comprising: an oximeter sensor comprising: a light emitting element configured to emit light;a light detector configured to receive the light;an accelerometer configured to detect motion of the oximeter sensor; anda sensor memory storing an algorithm configured to correct physio-optic effects from motion detected by the accelerometer; anda monitor communicatively coupleable with the sensor, the monitor configured to receive light detector and accelerometer information from the sensor and to implement the algorithm. 4. The oximeter system of claim 3, wherein the oximeter sensor comprises an alarm mechanism configured to indicate the presence of a designated amount of motion detected by the accelerometer. 5. The oximeter system of claim 3, wherein the monitor comprises a suspend mechanism configured to suspend operation of the oximeter system when a designated amount of motion is detected by the accelerometer. 6. The oximeter system of claim 3, wherein the accelerometer comprises a MEMS accelerometer incorporated into a chip within the oximeter sensor. 7. An oximeter monitor, comprising: a drive circuit configured to provide signals to circuitry of an oximeter sensor;a circuit configured to receive optical data from a detector of the oximeter sensor and information from an accelerometer of the oximeter sensor, the information relating to motion of the oximeter sensor; anda circuit configured to receive and implement an algorithm from memory of the oximeter sensor, wherein the algorithm is configured to correct physio-optic effects from motion detected by the accelerometer in the optical data. 8. The monitor of claim 7, comprising an alarm mechanism configured to indicate the presence of a designated amount of motion detected by the accelerometer. 9. A method of monitor operation, comprising: providing signals to circuitry of an oximeter sensor;receiving information from an accelerometer of the oximeter sensor, the information relating to motion of the oximeter sensor; andreceiving and implementing an algorithm from a memory of the oximeter sensor, wherein the algorithm is configured to correct physio-optic effects from motion detected by the accelerometer. 10. The method of claim 9, comprising indicating the presence of a designated amount of motion detected by the accelerometer with accelerometer data. 11. The method of claim 9, comprising suspending operation of the oximeter sensor when a designated amount of motion is detected by the accelerometer. 12. A... | Mallinckrodt Inc. | A61B005/1455
US20110137189 | A1 | 2011-06-09 | PHYSIOLOGICAL SIGNAL SENSING SYSTEM WITHOUT TIME AND PLACE CONTRAINT AND ITS METHOD | 2010-01-28 | A physiological signal sensing system and a physiological signal sensing method applied in any time and place are disclosed. The present invention is used by applying wireless transmitting electrocardiogram (ECG) data and contactless charging so as to achieve the purpose of sensing physiological signals without any time and place constraints. | 1. A physiological signal sensing system without time and place constraint, comprising: an electrocardiogram collector;a neural activity analyzer, analyzing neural activity from electrocardiogram data collected by the electrocardiogram collector;a wireless transceiver; anda remote data processing device;wherein the wireless transceiver is used for transmitting electrocardiogram signal collected by the electrocardiogram collector or the processed neural activity data to the remote data processing device or receiving data from the remote data processing device, and the remote data processing device is used for receiving and analyzing the obtained signal and returning results to the wireless transceiver or sending them to a medical organization. 2. The physiological signal sensing system of claim 1, wherein the electrocardiogram collector comprises a non-invasive measuring device for measuring the electrocardiogram data. 3. The physiological signal sensing system of claim 2, wherein the non-invasive measuring device comprises a detecting electrode and a reference electrode, and the detecting electrode and a reference electrode are mounted on the surface of the skin. 4. The physiological signal sensing system of claim 1 further comprising a radio stimulating device to give a user an electrical stimulation treatment. 5. The physiological signal sensing system of claim 1 further comprising a feedback device for suggesting or warning a user. 6. The physiological signal sensing system of claim 1, wherein the wireless transceiver and the remote data processing device respectively include a transmission interface which is a wireless communication protocol selected from a group of radio, wireless network, infrared, Bluetooth, radio frequency, GSM, PHS and CDMA. 7. The physiological signal sensing system of claim 1 further comprising a battery for maintenance operation of the electrocardiogram collector and the wireless transceiver and charging with a contactless mode 8. The physiological signal sensing system of claim 7, wherein the contactless mode is a charging process selected from a group of induction, optoelectronics, piezoelectricity, acousto-electricity, wind power and thermal electricity. 9. The physiological signal sensing system of claim 1 further comprising: a acceleration sensing component, collecting signals and signal processing results and sending them to the remote data processing device via the wireless transceiver;a physical activity calculating device for calculating a physical activity from the signal collected by the acceleration sensing component. 10. The physiological signal sensing system of claim 1, wherein the neural activity analyzer is used for carrying on spectral analysis to the heartbeat data. 11. A physiological signal recording method without time and place constraint, comprising: measurement of electrocardiogram signals;wireless transmission of data; andcontactless charging of equipment. 12. The physiological sig... | NATIONAL YANG-MING UNIVERSITY | A61B005/0408
US07967759 | B2 | 2011-06-28 | Endoscopic system with integrated patient respiratory status indicator | 2006-01-19 | The system 10 includes an endoscopic probe 20 and an expired air sampling device 24 functionally connected to a control console 28. In use, the endoscopic probe 20 is routed through a body lumen of a patient to, for example, visualize a selected region of a patient's body. As the endoscopic probe 20 is routed through the body lumens, the expired air sampling device 24 collects expired air from the patient, generates signals indicative of patient respiratory status, and outputs the generated signals to the control console 28. The control console 28 processes the signals and monitors the respiratory status of the patient. If the respiratory status of the patient changes based on the processed signals of the expired air sampling device 24, the control console 28 may output an audible or visual alert signal. | 1. An endoscope system, comprising: an endoscopic probe, the endoscopic probe having a proximal end, a distal end, and an elongate passageway extending between the proximal and distal ends; the elongate passageway configured to route at least one of an illumination system or a visualization system, wherein at least a distal end portion of the endoscopic probe is configured for insertion into a natural orifice of a body and routable through a tortuous natural lumen of the body; anda sampling device attached to the endoscopic probe at a position proximate the distal end portion of the endoscopic probe such that the sampling device is configured for insertion into the body, wherein the sampling device includes a capnographic sensor capable of generating signals indicative of CO2 gas concentration during endoscopic use, and wherein the capnographic sensor includes an infrared LED light associated with an infrared detector, and a reference light emitter associated with a reference light detector. 2. The endoscopic system of claim 1, wherein the capnographic sensor measures CO2 concentration by IR spectroscopy. 3. The endoscopic system of claim 1, wherein the distal end portion of the endoscopic probe includes a distal-most end face, and wherein the sampling device is positioned adjacent the distal-most end face of the endoscopic probe. 4. The endoscopic system of claim 3, wherein the capnographic sensor includes an elongate sampling chamber, and wherein the distal-most end face of the elongate sampling chamber is a aligned with the distal-most end face of the endoscopic probe. 5. The endoscopic system of claim 1, further comprising a control console, the control console receiving signals from the capnographic sensor, processing the received signals, and displaying results of the processed signals. 6. The endoscopic system of claim 5, wherein the control console includes an alert device for generating an alert signal indicative of patient distress based on the received signals from the capnographic sensor. 7. The endoscopic system of claim 6, wherein the alert signal is either visual or audible. 8. The endoscopic system of claim 1, further comprising a respiration sensor. 9. The endoscopic system of claim 8, wherein the respiration sensor generates signals indicative of breath cycles. 10. An endoscopic system, comprising: an endoscopic probe routable through a passageway of a body, the endoscopic probe having a proximal end, a distal end, and an elongate working channel therebetween, wherein the elongate working channel includes at least one of an illumination system or a visualization system, and wherein at least a distal end portion of the endoscopic probe is configured for insertion into a natural orifice of a body and routable through a tortuous natural lumen of the body; anda sampling device coupled to the endoscopic probe and having a capnographic sensor capable of generating signals indicative of respiratory gas concentrations during en... | Boston Scientific Scimed, Inc. | A61B005/08
US07945314 | B1 | 2011-05-17 | System and method for emulating a surface EKG for use with transtelephonic monitoring of an implantable medical device | 2006-05-10 | A surface electrocardiogram (EKG) is emulated using signals detected by internal leads of an implanted device. In one example, emulation is performed using a technique that concatenates portions of signals sensed using different electrodes, such as by combining far-field ventricular signals sensed in the atria with far-field atrial signals sensed in the ventricles. In another example, emulation is performed using a technique that selectively amplifies or attenuates portions of a single signal, such as by attenuating near-field portions of an atrial unipolar signal relative to far-field portions of the same signal. The surface EKG emulation may be performed by the implanted device itself or by an external programmer based on cardiac signals transmitted thereto. A transtelephonic monitoring network is also described, wherein the emulated surface EKG (or raw data used to emulate the EKG) is relayed from an implanted device to a remote monitor, typically installed in a physician's office. | 1. A method for emulating a surface electrocardiogram (EKG) of a patient in which an implantable cardiac stimulation device is implanted, the method comprising: sensing first and second cardiac signals using electrodes implanted within the patient;identifying a pair of baseline points in each of the first and second cardiac signals;selectively concatenating that portion of the first cardiac signal that lies between the pair of baseline points with that portion of the second cardiac signal that lies outside the pair of baseline points, using the baseline points as concatenation points to generate an emulated surface EKG using the implantable device;transmitting the emulated surface EKG to an external relay device; andrelaying the emulated surface EKG to a remote monitor using the external relay device. 2. The method of claim 1 wherein the step of relaying the emulated surface EKG to a remote monitor using the external relay device is performed by transmitting emulated surface EKG data via a communication channel comprising one or more of the public switched telephone network (PSTN), a wireless cellular communication network, a satellite communication network, a local area network (LAN), a wide area network (WAN) and the Internet. 3. The method of claim 1 wherein the step of transmitting the emulated surface EKG to an external relay device is performed by transmitting the data via a telemetry wand placed over the chest of the patient in proximity to the implanted device. 4. The method of claim 1 wherein external relay device is part of a bedside monitor and wherein the step of transmitting the emulated surface EKG to an external relay device is performed to transmit the emulated surface EKG directly to the bedside monitor. 5. The method of claim 1 wherein the step of sensing separate cardiac signals includes the steps of: sensing atrial signals using at least one atrial electrode; andsensing ventricular signals using at least one ventricular electrode. 6. The method of claim 5 wherein the step of selectively concatenating portions of the cardiac signals includes the steps of: identifying far-field ventricular signals within the atrial channel signals;identifying far-field atrial signals within the ventricular channel signals; andconcatenating the far-field atrial signals and the far-field ventricular signals to generate the emulated surface EKG. 7. The method of claim 1 wherein the step of sensing cardiac signals using electrodes implanted within the patient includes the steps of: sensing atrial signals using unipolar sensing by employing an atrial electrode in combination with a housing of the device; andsensing ventricular signals using unipolar sensing by employing a ventricular electrode in combination with a housing of the device. 8. The method of claim 1 wherein the step of selectively concatenating portions of the cardiac signals includes the additional step of: smoothing the emulated surface EKG at concatenation points. 9. The metho... | Pacesetter, Inc. | A61B005/04
US08285355 | B2 | 2012-10-09 | Biomedical electric wave sensor | 2010-05-12 | A biomedical electric wave sensor includes a base, a central pole, a dry electrode, a case, and a plurality of ribs. When the central pole lowers down, the ribs radiate and expand outward to push aside the hair of a subject, and the dry electrode exposes from the case and contacts the skin of the subject to measure a physiological electric wave signal from the subject. The present invention may overcome the intervention problem caused by hair and achieve the measurement of biomedical electric wave signal. | 1. A biomedical electric wave sensor, comprising: a base comprising a plurality of holes at a rim of the base, wherein a lower surface of the base is configured as a dry electrode;a case housing the base and exposing the lower surface of the base;a central pole connected to an upper surface of the base, penetrating through the case and configured for lifting and lowering relative to the case; anda plurality of ribs configured at an inner side of the case corresponding to the holes, the ribs radiating inward relative to the central pole, penetrating through the holes and protruding from the lower surface of the base, wherein as the central pole lowers down, the ribs protruding from the lower surface of the base radiate and expand outward to push aside hair of a subject, and the dry electrode is exposed from the case and contacts skin of the subject to measure a physiological electric wave signal from the subject. 2. The biomedical electric wave sensor as claimed in claim 1, wherein the base has a round shape. 3. The biomedical electric wave sensor as claimed in claim 2, wherein the case has a round shape. 4. The biomedical electric wave sensor as claimed in claim 1, wherein the base has a regular polygon shape. 5. The biomedical electric wave sensor as claimed in claim 4, wherein the case has a regular polygon shape. 6. The biomedical electric wave sensor as claimed in claim 1, wherein the central pole comprises an internal circuit configured therewithin and electrically connected to the dry electrode for outputting the physiological electric wave signal. 7. The biomedical electric wave sensor as claimed in claim 1, wherein the case further comprises a round disk, and the central pole penetrates through the round disk. 8. The biomedical electric wave sensor as claimed in claim 7, wherein the central pole further comprises a tenon configured for retaining the round disk as the central pole lowers down. 9. The biomedical electric wave sensor as claimed in claim 8, wherein the tenon is made of plastic material. 10. The biomedical electric wave sensor as claimed in claim 1, further comprising a top disk connected with the central pole. 11. The biomedical electric wave sensor as claimed in claim 10, wherein the top disk further comprises a grip configuration on a top surface of the top disk. 12. The biomedical electric wave sensor as claimed in claim 1, further comprising a hair dividing mechanism extending from a bottom of each of the ribs. 13. The biomedical electric wave sensor as claimed in claim 12, wherein the hair-dividing mechanism has a V shape or a rake shape. 14. The biomedical electric wave sensor as claimed in claim 1, wherein the dry electrode comprises a plurality of microstructure probes. 15. The biomedical electric wave sensor as claimed in claim 14, wherein the microstructure probe comprises a plurality of MEMS (Micro Electro Mechanical Systems) probes or a plurality of CNT (carbon nanotubes) probes. | National Chiao Tung University | A61B005/04
US08224423 | B2 | 2012-07-17 | Means of tracking movement of bodies during medical treatment | 2010-06-18 | During preprogrammed medical treatment and remote controlled surgery tracking the movements of the body being treated and integrating those tracked movements with the preprogrammed/remote controlled treatment to arrive at an integrated modified treatment track and following that modified track during the treatment. The treating instrument may be a solid scalpel or high-energy radiation, such as X-rays, ultra sound, laser beams or the like. | 1. A system for improving the accuracy of preprogrammed treatment on a body of a patient having an inside portion that is in need of the treatment and an outside portion that is moveable during the treatment, comprising: at least one emission source;a plurality of markers configured for independent attachment to a plurality of locations adjacent a surface of the moveable outside portion of the body proximate to the inside portion in need of the treatment to emit a corresponding plurality of signals, wherein each of the markers comprise:a support element comprising a base for securing the marker adjacent the surface of the moveable outside portion of the body;a signal emitter operatively associated with the support element; andan optical fiber coupling the signal emitter and the emission source;a timing device configured to cause the signal emitters to respectively emit signals under conditions sufficient to differentiate which signal emitter is sending each signal, respectively;at least one receiver disposed remote from the body and positioned to receive the plurality of signals from the signal emitters, wherein each of the support elements is configured to position a respective signal emitter at an angular offset from a perpendicular direction relative to the base to direct emissions from the respective signal emitter towards the at least one receiver;a remotely-controlled surgical device;a computing device comprising a processor and wherein the processor is configured with software executable instructions to cause the computing device to perform operations comprising: tracking the movement of the outside portion of the body based on the received plurality of signals from the signal emitters;identifying and mapping the inside portion of the body that is intended to be subjected to treatment;preprogramming a treatment path adapted to be followed by the remotely-controlled surgical device; andintegrating the tracked movements with the treatment path to form a modified treatment path; anda guidance apparatus configured for causing the remotely-controlled surgical device to treat the inside portion of the body along the modified treatment path while substantially preventing the remotely-controlled surgical device from departing from the modified treatment path to any substantial extent. 2. The system of claim 1, wherein the remotely-controlled surgical device is configured to be operated without benefit of a surgeon and to perform a treatment selected from the group consisting of a high energy radiation source sufficient to render the inside portion of the body necrotic and a high energy ultrasound radiation sufficient to break up stones. 3. The system of claim 1, wherein the at least one emission source comprises a plurality of light emitting diodes (LEDs). 4. The system of claim 3, wherein the timing device is configured to cause the signal emitters to respectively emit signals under conditions sufficient to differentiate which emitter is sendi... | Boulder Innovation Group, Inc. | A61B005/05 A61N007/00
US08821416 | B2 | 2014-09-02 | Patient management device, system and method | 2009-05-29 | A portable apparatus for managing a plurality of events related to a patient and for recording and storing input data related to said events is disclosed. The portable apparatus comprises a man machine user interface; a timer unit; a control unit, arranged to provide an invitation to the patient as an instruction via the interface to perform a motion exercise, upon a trigger from the timer unit or upon a patient self initiation. In an embodiment the apparatus comprises a sensor to record measurements related to said motion exercise. Further, a storage unit is arranged to retrievably store patient input data and said measurements. This enables diagnosing a degree of a neurological disease or an effect of a medication on said neurological disease. | 1. A method of diagnosing a degree of a neurological disease or an effect of a medication on said neurological disease, said method comprising: transmitting a questionnaire via a communication link from an application on a local computer system to a portable apparatus, wherein said portable apparatus comprises an integrated accelerometer and said questionnaire comprises questions and an instruction to initiate a predefined motion exercise;registering movement of said portable apparatus during said predefined motion exercise in response to said instruction by registration of an output signal from said accelerometer and storing said output signal as stored measurement data elements;correlating said questions and said instruction to said measurement data elements; andanalyzing said measurement data elements for determining said degree of said neurological disease or an effect of said medication on said neurological disease. 2. The method of claim 1, further comprising comparing values of parameters analyzed from said measurement data with corresponding values from a standard population for said diagnosing. 3. The method of claim 1, further comprising: registering said movement and storing said output signal as time stamped measurement data elements at repeated fixed or arbitrary time intervals or upon user initiation, anddetermining a trend for a parameter from a plurality of said time stamped measurement data elements. 4. The method of claim 3, wherein said neurological disease is Parkinson's disease, and said method further comprises identifying OFF periods and/or ON periods of a Parkinson's disease patient from said stored measurement data elements. 5. The method of claim 1, wherein said portable apparatus is an integrated apparatus comprising: a user interface comprising an input unit for registering patient input data elements and an output unit for providing information elements to a patient;a timer unit arranged to trigger at least a first event of said plurality of events and to tag said input data with a time stamp;a control unit, arranged to provide an instruction via said output unit to said patient to perform a desired movement in a motion exercise, upon said trigger from said timer unit or upon a patient self-initiation;at least one sensor unit arranged to record measurement data elements related to said motion exercise of said patient, wherein said at least one sensor unit comprises said accelerometer adapted to measure said desired motion of said portable apparatus; anda storage unit, arranged to retrievably store said patient input data elements and said measurement data elements, together with said time stampwherein:said portable apparatus is adapted to be held by said patient, or attached in a fixed relation to a body portion of said patient, and thus to serve as a mass to be moved in a desired manner by said patient during said motion exercise in such a manner that said measurement data elements are provided that are related t... | Cunctus AB | A61B005/103 A61B005/117
US07941206 | B2 | 2011-05-10 | Automatic orientation determination for ECG measurements using multiple electrodes | 2010-01-27 | Cardiac monitoring and/or stimulation methods and systems provide monitoring, defibrillation and/or pacing therapies. A signal processor receives a plurality of composite signals associated with a plurality of sources, separates a signal using a source separation algorithm, and identifies a cardiac signal using a selected vector. The signal processor may iteratively separate signals from the plurality of composite signals until the cardiac signal is identified. The selected vector may be updated if desired or necessary. A method of signal separation involves detecting a plurality of composite signals at a plurality of locations, separating a signal using source separation, and selecting a vector that provides a cardiac signal. The separation may include a principal component analysis and/or an independent component analysis. Vectors may be selected and updated based on changes of position and/or orientation of implanted components and changes in patient parameters such as patient condition, cardiac signal-to-noise ratio, and disease progression. | 1. A system, comprising: a plurality of implantable electrodes configured for sensing a plurality of composite signals associated with an unknown number of sources, the plurality of implantable electrodes comprising at least three spatially separated implantable electrodes;a housing configured for implantation in a patient;a control circuitry provided in the housing;a memory provided in the housing;a processing circuitry provided in the housing and coupled to the memory, the processing circuitry configured to perform a source separation that separates signal components of the sensed plurality of composite signals according to their sources to produce a set of signal vectors, the processing circuitry configured to: select, from the set of signal vectors, a target vector associated with a target signal, the target vector representative of a signal vector of the set of signal vectors associated with higher quality cardiac data relative to other signal vectors of the set of signal vectors;detect a change in a predetermined condition or an event occurrence indicating that changing which signal vector of the set of signal vectors is selected as the target vector may be useful in sensing cardiac activity; andupdate, in response to the change in the predetermined condition or event occurrence, selection of the target vector by performing a subsequent source separation on the detected plurality of composite signals and reselecting the same or a different signal vector associated with higher quality cardiac data relative to other signal vectors of the set of signal vectors as the target vector based on the subsequent source separation; andcommunications circuitry provided in the housing and configured to facilitate communication of information between the housing and a device external of the housing. 2. The system according to claim 1, wherein the at least three spatially separated implantable electrodes are provided at the housing. 3. The system according to claim 1, wherein the at least three spatially separated implantable electrodes comprise at least four electrodes, and wherein one of the at least four electrodes does not lie in the same plane as the other of the at least four electrodes. 4. The system according to claim 1, wherein at least one of the at least three spatially separated implantable electrodes is provided at the housing and at least two of the at least three spatially separated implantable electrodes are provided on a cardiac lead arrangement. 5. The system according to claim 1, wherein the processing circuitry is configured to perform blind source separation that separates signal components of the sensed plurality of composite signals according to their sources to produce the set of signal vectors. 6. The system according to claim 1, wherein the processing circuitry is configured to update selection of the target vector in response to at least one of a change in orientation of the housing and a change of a signal to noise ratio of... | Cardiac Pacemakers, Inc. | A61B005/04
US08783102 | B2 | 2014-07-22 | Protective container for holding reusable diagnostic components | 2010-12-17 | A protective container for holding a reusable control part of a transcutaneous sensor system for detecting at least one analyte in a bodily fluid is disclosed. The control part includes at least one coupling, which has at least one sensor coupling for connection to at least one transcutaneous sensor. The protective container has at least one container housing. The control part can be held in the container housing. The container housing is adapted to shield the control part from environmental influences. The container housing also has at least one connector which can be connected to the coupling and seals the latter in a media-tight fashion. | 1. A protective container for holding a reusable control part of a transcutaneous sensor system for detecting at least one analyte in a bodily fluid, wherein the control part includes at least one coupling having at least one sensor coupling for connection to at least one transcutaneous sensor, wherein the protective container comprises: at least one container housing, wherein the control part is positionable in the container housing, wherein the container housing is adapted to shield the control part from environmental influences, and wherein the container housing has at least one connector, the connector being connectable to the coupling and sealing the coupling in a media-tight fashion when connected thereto; and wherein the protective container is adapted to perform a function test of the control part. 2. The protective container as claimed in claim 1 wherein the coupling has at least one control coupling for connecting the control part to at least one control component, and wherein the connector comprises at least one sensor connector for connection to the sensor coupling and at least one control connector for connection to the control coupling. 3. The protective container as claimed in claim 1 wherein the container housing has an electrical shielding property. 4. The protective container as claimed in claim 1 wherein the connector is adapted to establish an electrical connection to the coupling. 5. The protective container as claimed in claim 1 wherein the connector provides ESD protection for the control part. 6. The protective container as claimed in claim 1 wherein the protective container is adapted to supply at least one test current and/or at least one test voltage to the coupling to thereby perform the function test of the control part. 7. The protective container as claimed in claim 6 wherein the protective container is adapted to undertake a leakage-current measurement between at least two contacts on the coupling to thereby perform the function test of the control part. 8. The protective container of claim 6 wherein the protective container supplies a test current to the coupling to thereby perform the function test of the control part. 9. The protective container of claim 6 wherein the protective container supplies a test voltage to the coupling to thereby perform the function test of the control part. 10. The protective container as claimed in claim 1 wherein the protective container is adapted to transmit at least one result of the function test to at least one further instrument and/or one user. 11. The protective container as claimed in claim 1 wherein the protective container is adapted to transmit an alert to a user and/or a further instrument if a fault is established in the control part during the function test. 12. The protective container as claimed in claim 1 wherein the protective container has at least one electrical energy supply, the at least one electrical energy supply comprising one or more of the fol... | Roche Diagnostics Operations, Inc. | A61B005/145 A61B005/00 A61B005/1495 B65D085/00 B65D085/86 A61B005/1486
US09462959 | B2 | 2016-10-11 | Methods and systems that use implanted posture sensor to monitor left atrial pressure and/or inter-thoracic fluid volume | 2009-12-30 | In specific embodiments, a method to monitor left atrial pressure and/or intra-thoracic fluid volume of a patient, comprises (a) monitoring posture of the patient using a posture sensor implanted within the patient, and (b) using portions of an impedance signal, obtained using implanted electrodes, to monitor the left atrial pressure and/or intra-thoracic fluid volume of the patient. Each portion of the impedance signal used to monitor the left atrial pressure and/or intra-thoracic fluid volume of the patient corresponds to a period after which the patient has maintained a predetermined posture for at least a predetermined period of time, and during which the patient has remained in the predetermined posture. | 1. A method to monitor left atrial pressure (LAP) of a patient comprising: (a) storing data indicative of an impedance signal for at least a first specified period of time, obtained using implanted electrodes, and a corresponding posture of the patient, as determined using an implanted posture sensor;(b) identifying, based on the stored data, one or more snapshots of the impedance signal where the patient is in a predetermined posture and where prior to each snapshot the patient substantially maintained the predetermined posture for at least a second specified period of time, and during the snapshot the patient maintained the predetermined posture for at least a third specified period of time;(c) calculating, for each of the identified one or more snapshots of the impedance signal, a separate zLAP value indicative of the patient's left atrial pressure (LAP) while the patient is in the predetermined posture;(d) identifying an undesirably abnormal change in the patient's left atrial pressure (LAP) in response to at least one of (d.1) detecting a change in the calculated zLAP values by a predefined amount, or(d.2) detecting a trend in the calculated zLAP values that are symptomatic of organ damage and/or failure; and(e) one or both of triggering an alert or adjusting therapy administered to the patient in response to identifying the abnormality. 2. The method of claim 1, further comprising, after step (b), writing over or otherwise erasing the stored data, indicative of the impedance signal and the corresponding posture of the patient, not associated with the identified snapshots. 3. The method of claim 1, wherein after the first specified period of time steps (a), (b) and (c) are repeated for further periods of time of the same or different length until actively interrupted. 4. The method of claim 1, wherein steps (a), (b), (c), (d) and (e) are performed by an implanted device including and/or connected to the implanted electrodes; and wherein when an alert is triggered, the method further comprises: (f) transmitting the alert from the implantable device to a non-implanted device; and(g) displaying a visual warning and/or issuing an audible warning by way of the non-implanted device in response to the alert. 5. The method of claim 1, wherein steps (a), (b), (c), (d) and (e) are performed by an implanted system including and/or connected to the implanted electrodes; and wherein when therapy administered to the patient is adjusted, the method comprises: adjusting medication based on a physician pre-determined formula; anddisplaying a visual indication of the medication adjustment to the patient. 6. The method of claim 1, wherein steps (a), (b), (c), (d) and (e) are performed by an implanted system including and/or connected to the implanted electrodes; and wherein when therapy administered to the patient is adjusted, the method comprises adjusting electrical therapy administered by an implantable cardioverter-defibrillator (ICD) or cardiac resyn... | Pacesetter, Inc. | A61N001/00 A61B005/053 A61B005/021 A61B005/11 A61B005/00 A61N001/362 A61N001/365 A61N001/372 A61N001/368 A61N001/37 A61N001/39
US09883799 | B2 | 2018-02-06 | Method of identifying when a patient undergoing hemodialysis is at increased risk of death | 2010-12-02 | The invention is directed to a method of identifying a patient undergoing periodic hemodialysis treatments at increased risk for death that includes determining at least one of the patient's clinical or biochemical parameters, including systolic blood pressure, serum albumin concentration level, body weight, body temperature, serum bicarbonate concentration level, serum potassium concentration level, serum calcium concentration level, hemoglobin concentration level, serum phosphorus concentration level, neutrophil to lymphocyte ratio, equilibrated normalized protein catabolic rate (enPCR), equilibrated fractional clearance of total body water by dialysis and residual kidney function (eKdrt/V), EPO resistance index, transferrin saturation index, serum ferritin concentration level, serum creatinine concentration level, platelet count, Aspartat-Aminotransferase level, and Alanin-Aminotransferase level at periodic hemodialysis treatments, and identifying a patient as having an increased risk for death if the patient has a significant change in the rate of change of at least one of the patient's clinical or biochemical parameters. The invention is also directed to a method of identifying an increased mortality risk factor for a patient undergoing periodic hemodialysis treatment. The method includes analyzing data of deceased patients that were previously undergoing periodic hemodialysis treatments by performing a longitudinal analysis backwards in time of changes in a clinical or biochemical parameter the patients, and identifying a significant change in the rate of decline or the rate of increase in a clinical or biochemical parameter before death of the patients. | 1. A method of identifying and treating a patient undergoing periodic hemodialysis treatments at increased risk for death, comprising: a) determining at least one clinical or biochemical parameter associated with an increased risk of death of the patient and monitoring said parameter periodically before and/or after the patient is undergoing hemodialysis treatments;b) determining a significant change in the rate of change of the at least one clinical or biochemical parameter from a retrospective record review of parameter values of the patient determined at prior hemodialysis treatments;c) identifying the patient as having an increased risk for death because the patient has the significant change in the rate of change of the at least one clinical or biochemical parameter; andd) treating the patient having an increased risk for death within a sufficient lead time to decrease the patient's risk of death. 2. The method of claim 1, wherein the at least one clinical or biochemical parameter of the patient includes systolic blood pressure, serum albumin concentration level, body weight, body temperature, serum bicarbonate concentration level, serum potassium concentration level, serum calcium concentration level, hemoglobin concentration level, serum phosphorus concentration level, neutrophil to lymphocyte ratio, equilibrated normalized protein catabolic rate (enPCR), equilibrated fractional clearance of total body water by dialysis and residual kidney function (eKdrt/V), Erythropoietin (EPO) resistance index, transferrin saturation index (TSAT), serum ferritin concentration level, serum creatinine concentration level, platelet count, Aspartat-Aminotransferase level, and Alanin-Aminotransferase level. 3. The method of claim 2, wherein a significant change is determined by using a statistical method. 4. The method of claim 2, wherein a significant change is a change from a steady level to an increase or a decrease in the rate of change of the at least one clinical or biochemical parameter. 5. The method of claim 2, wherein a significant change is a change in character of the rate of change of the at least one clinical or biochemical parameter. 6. The method of claim 1, further including initiating at least one diagnostic procedure prior to the step of treating the patient. | Fresenius Medical Care Holdings, Inc. | A61B005/00 G06Q050/24 A61B005/0205 A61B005/021
US09974485 | B2 | 2018-05-22 | Method and system for providing remote healthcare services | 2009-12-11 | An approach is provided for healthcare information management and providing remote healthcare examination. A platform receives position data corresponding to a mark to be projected onto a human subject by a user equipment. The mark specifies placement of a medical sensor onto the subject. The platform collects measurement data from the medical sensor coupled to the subject at the mark. | 1. A method comprising: receiving, by one or more devices and from a user device of a medical professional, position data corresponding to a mark drawn or designated, by the medical professional and via a user interface of the user device, on an image of a subject, the user device and the medical professional being located remotely from the one or more devices;initiating, by the one or more devices and based on the position data, projection of the mark onto the subject, the mark specifying placement of a medical sensor by the subject or a party that is assisting the subject, andthe medical sensor being connected to the one or more devices;collecting, by the one or more devices, measurement data from the medical sensor;receiving, by the one or more devices and from the user device, a request to stop the projection of the mark onto the subject when the medical professional eliminates, via the user interface of the user device, the mark from the image of the subject; andstopping, by the one or more devices and based on the request, the projection of the mark onto the subject. 2. The method of claim 1, further comprising: detecting movement of the subject;determining a new position for the mark based on the movement; andinitiating projection of the mark at the new position onto the subject. 3. The method of claim 1, further comprising: capturing the image of the subject; andinitiating transmission of the image to the user device. 4. The method of claim 1, further comprising: initiating transmission of the collected measurement data for storage in a database that is accessible by the user device. 5. The method of claim 4, wherein the database is maintained by a social network service provider or a telecommunication service provider. 6. An apparatus comprising: a memory; anda processor configured to: receive, from a user device of a medical professional, position data corresponding to a mark drawn or designated, by the medical professional and via a user interface of the user device, on an image of a subject, the user device and the medical professional being located remotely from the apparatus;initiate projection of the mark onto the subject, the mark specifying placement of a medical sensor by the subject or a party that is assisting the subject, andthe medical sensor being connected to the apparatus;collect measurement data from the medical sensor;receive, from the user device, a request to stop the projection of the mark onto the subject when the medical professional eliminates, via the user interface of the user device, the mark from the image of the subject; andstop, based on the request, the projection of the mark onto the subject. 7. The apparatus of claim 6, wherein the processor is further configured to: detect movement of the subject,determine a new position for the mark based on the movement, andinitiate projection of the mark at the new position onto the subject. 8. The apparatus of claim 6, wherein the processor is further configur... | Verizon Patent and Licensing Inc. | G06Q010/00 A61B005/00 G06F019/00
US09320571 | B2 | 2016-04-26 | Lead tracking and positioning system and method | 2009-12-30 | Image data can be obtained with an imaging device. A location of the imaging device relative to a subject can be determined. A location of an instrument can be tracked relative to the subject using a tracking system. Also, the tracked location of the instrument can be illustrated relative to the image data. | 1. A lead placement system for placing a lead in a subject, the lead placement system comprising: the lead;a stylet comprising a body, a curved portion and an angled portion, wherein the stylet is tubular shaped,the body extends along a first axis between a proximal end of the stylet and the curved portion,the angled portion extends along a second axis from the curved portion to a distal end of the stylet, andthe second axis is at a fixed predetermined angle relative to the first axis;a first tracking device disposed within the angled portion;a second tracking device disposed (i) within the stylet, (ii) within or proximal the curved portion, and (iii) proximal to the first tracking device,wherein the stylet is configured to be inserted in and extend along at least a portion of the lead; anda plurality of electrodes (i) configured to supply an electrical current to the subject to stimulate the subject, and (ii) disposed on the lead and surrounding portions of at least one of the angled portion, the curved portion and the body,wherein the lead has an inner diameter and an outer diameter,each of the plurality of electrodes has an inner diameter and an outer diameter, andat least one of (i) the inner diameter of each of the plurality of electrodes is equal to the inner diameter of the lead, or (ii) the outer diameter of each of the plurality of electrodes is equal to the outer diameter of the lead. 2. The lead placement system of claim 1, further comprising a third tracking device disposed within the body and proximal to the second tracking device, wherein the second tracking device and the third tracking device are configured to define the first axis relative to the second axis. 3. The lead placement system of claim 2, further comprising: the lead; andan electrode,wherein at least a portion of the third tracking device is disposed within a volume surrounded by the electrode when the stylet is within the lead. 4. The lead placement system of claim 3, further comprising at least one processor comprising hardware configured to: track locations of the first tracking device, the second tracking device, and the third tracking device; andbased on the locations, determine a direction of the angled portion relative to the body. 5. The lead placement system of claim 4, further comprising a display device operable to display: at least one icon representing at least one of the first tracking device, the second tracking device, the third tracking device, and the lead; anda second icon representing the orientation of the angled portion relative to the body. 6. The lead placement system of claim 2, wherein: the first tracking device, the second tracking device, and the third tracking device are at predetermined fixed positions within the stylet;the stylet has an inner diameter; andthe first tracking device, the second tracking device, and the third tracking device are all disposed within the stylet. 7. The lead placement system of claim 6, wherein the first ... | Medtronic, Inc. | A61B005/05 A61B019/00
US20110151432 | A1 | 2011-06-23 | METHODS AND SYSTEMS TO COLLECT AND PREPARE SAMPLES, TO IMPLEMENT, INITIATE AND PERFORM ASSAYS, AND TO CONTROL AND MANAGE FLUID FLOW | 2010-10-20 | Methods and systems to related to sample collection, assays, and fluid control and management, including methods and systems to implement hand-held portable assays, to activate an assay system, to collect and prepare samples, to capture antibodies, and to trap or capture gas bubbles. Methods and systems disclosed herein, and portions thereof, may be implemented alone and/or in various combinations with one another. | 1. A system, comprising: a portable housing having a sample region, the sample region having a fluid inlet and a fluid outlet, the portable housing further having a fluid chamber and a fluid passage from the fluid chamber to the sample region fluid inlet;a movable cover to enclose the sample region; anda mechanically actuated fluid controller to force fluid from the fluid chamber, through the fluid passage, the sample region fluid inlet, and the sample region fluid outlet, to move at least a portion of a sample from the sample region through the fluid outlet. 2. The system of claim 1, wherein the movable cover is configured to prevent the mechanically actuated fluid controller from forcing fluid flow when the movable cover is in an open position. 3. The system of claim 1, wherein the sample region is located on an enclosable portion of a surface of the portable housing. 4. The system of claim 1, wherein the moveable cover includes a mechanically slideable cover. 5. The system of claim 4, wherein an outer surface of the portable housing has a substantially cylindrical shape, and wherein the moveable cover includes a slideable sleeve surrounding a portion of the portable housing. 6. The system of claim 1, wherein the mechanically actuated fluid controller includes a mechanical actuator to close the cover and to actuate fluid flow when the movable cover is closed. 7. The system of claim 6, wherein: the movable cover and the mechanical actuator are releasably coupled to one another; andthe mechanical actuator is configured to close the cover and to thereafter disengage from the cover and actuate fluid flow when the cover is closed. 8. The system of claim 7, further including: a breakable tab to releasably couple the mechanical actuator to the moveable cover. 9. The system of claim 8, wherein the moveable cover includes a slideable sleeve surrounding a portion of the portable housing. 10. The system of claim 1, wherein the moveable cover includes a hinged cover. 11. The system of claim 1, wherein the moveable cover includes a rotating cover. 12. The system of claim 1, wherein the moveable cover includes a movable window. 13. The system of claim 1, wherein the moveable cover includes a removable window. 14. The system of claim 1, further including: a sample collection pad disposed within the sample region to receive a biological sample. 15. The system of claim 1, further including: a sample collection pad disposed within the sample region to receive an environmental sample. 16. The system of claim 1, further including: a sample collection pad disposed within the sample region to absorb a sample. 17. The system of claim 15, wherein the sample region and the sample collection pad are dimensioned to receive less than approximately 250 micro-liters of the biological sample. 18. The system of claim 1, further including: a lance device to pierce a membrane to release a biological sample therefrom. 19. The system of claim 1, further includin... | BOSTON MICROFLUIDICS | C12Q001/70 G01N001/00 B01D035/00 G01N033/53 C12M001/00 G01N033/566 A61B005/157
US20110144471 | A1 | 2011-06-16 | FLEXIBLE PROBE STRUCTURE AND METHOD FOR FABRICATING THE SAME | 2009-12-16 | The present invention discloses a flexible probe structure comprises at least one electrode using a CNT layer as the electrode interface. The CNT layer disposed on the electrode surface is processed with an UV-ozone treatment to form a great number of carbon-oxygen functional groups on the surface of CNT. The carbon-oxygen functional groups can greatly reduce the interface impedance of the electrode and the biological tissue fluid. Thereby, the measurement can achieve better quality. The present invention also discloses a method for fabricating a flexible probe structure, which comprises steps: preparing a flexible substrate; forming a conductive layer on the flexible substrate, and defining an electrode, a wire and a metal pad on the conductive layer; forming a CNT layer on the electrode; forming an insulating layer on the conductive layer to insulate the wire from the environment; and processing the CNT layer with an UV-ozone treatment. | 1. A flexible probe structure comprising a base and at least one probe connected to said base, said probe having at least one electrode, wherein each said electrode is electrically connected to a metal pad on said base via a wire, and wherein said base and said probe are made of a flexible polymeric material, and wherein said electrode uses a CNT (carbon nanotube) layer as an electrode interface, and wherein said CNT layer is processed with an UV (ultraviolet ray)-ozone treatment to form carbon-oxygen functional groups on an outmost layer thereof. 2. The flexible probe structure according to claim 1, wherein said base and said probe are made of a material selected from a group consisting of polyimide, poly-para-xylylene, a thick photoresist SU-8, polydimethylsiloxane and benzocyclobutene. 3. A method for fabricating a flexible probe structure comprising the steps of: preparing a flexible substrate;forming a conductive layer on said flexible substrate, and defining an electrode, a wire and a metal pad on said conductive layer;forming a CNT (carbon nanotube) layer on said electrode;forming an insulating layer on said conductive layer; andprocessing said CNT layer with an UV (ultraviolet ray)-ozone treatment to form carbon-oxygen functional groups on an outmost layer of said CNT layer. 4. The method for fabricating a flexible probe structure according to claim 3, wherein said flexible substrate is made of a material selected from a group consisting of polyimide, poly-para-xylylene, a thick photoresist SU-8, polydimethylsiloxane and benzocyclobutene. 5. The method for fabricating a flexible probe structure according to claim 3, wherein said conductive layer is made of a metallic material selected from a group consisting of gold, silver, aluminum, copper, platinum, and an alloy thereof. 6. The method for fabricating a flexible probe structure according to claim 3, wherein in said UV-ozone treatment, said CNT layer is illuminated with an ultraviolet ray having an illumination intensity of 25-35 mW/cm2. 7. The method for fabricating a flexible probe structure according to claim 6, wherein said ultraviolet ray has a wavelength of 254 nm. 8. The method for fabricating a flexible probe structure according to claim 3, wherein said insulating layer is made of a material selected from a group consisting of polyimide, poly-para-xylylene, a thick photoresist SU-8, polydimethylsiloxane and benzocyclobutene. 9. The method for fabricating a flexible probe structure according to claim 3, wherein said CNT layer is formed on said electrode with a chemical vapor deposition method. 10. The method for fabricating a flexible probe structure according to claim 9, wherein said CNT layer is formed on said electrode via a catalytic layer; said catalytic layer is made of a material selected from a group consisting of iron, cobalt, nickel, and an alloy thereof. 11. The method for fabricating a flexible probe structure according to claim 9, wherein said CNT layer is s... | - | A61B005/0478 H05K003/10
US07953474 | B2 | 2011-05-31 | Apparatus, method and optical sensor module using a tilter for body fat measurement | 2007-06-12 | An apparatus, method and optical sensor module for body fat measurement includes at least one light source, a tilter to control an irradiation angle of the light source to radiate light from the light source into biological tissue at a predetermined angle, and an optical detector to detect an optical signal scattered from the biological tissue, and transform the scattered optical signal into an electric signal to acquire more accurate body fat information about a deep body fat layer. | 1. An optical sensor module for a body fat measurement apparatus, said apparatus being arranged to measure a thickness of a subcutaneous fat layer between skin and muscle in a biological tissue of an organism, the module comprising: at least one light source configured for use in a reflection scattering-absorption process;a tilter to control an irradiation angle of the at least one light source to radiate light from the at least one light source into the biological tissue at a predetermined angle other than an angle perpendicular to the biological tissue so that the radiated light penetrates deeper into the biological tissue, the tilter being placed between the at least one light source and the biological tissue whereby the light incident upon the biological tissue from the at least one light source falls substantially at an angle other than one perpendicular to the biological tissue;an optical detector to detect a scattered optical signal and to transform the scattered optical signal into an electric signal, the scattered optical signal being scattered from the light radiated into the biological tissue; andan optical power control unit to control a driving current applied to the at least one light source,wherein the optical detector further comprises a transformer unit to transform the scattered optical signal into an electrical signal, andwherein the irradiation angle is controlled such that the radiated light comprisesa component directed away from the optical detector. 2. The module of claim 1, wherein the tilter is comprised of a material having a predetermined refraction index to radiate light from the at least one light source into the biological tissue at the predetermined angle. 3. The module of claim 1, wherein the tilter comprises a solid water epoxy or a polymer compound. 4. The module of claim 1, wherein the tilter is a prism and the irradiation angle ranges from 2 degrees to 45 degrees. 5. The module of claim 1, wherein: the at least one light source comprises at least two light sources, each being a different distance from the optical detector; andthe tilter is located at one of the at least two light sources located farthest from the optical detector. 6. The module of claim 1, wherein: the electric signal is applied to an electric signal calculation unit; andthe electric signal calculation unit receives the electric signal and calculates the thickness of the fat layer of a corresponding measurement point. 7. The module of claim 6, wherein the electric signal calculation unit comprises: a signal amplification unit to amplify the electric signal transformed by the transformer unit;a body fat information calculation unit to calculate body fat information comprising the thickness of the fat layer from the amplified electric signal; andan output control unit to control the body fat information to be output to a predetermined display unit or sound output unit. 8. A body fat measurement apparatus, said apparatus being arranged to ... | Samsung Electronics Co., Ltd. | A61B006/00 A61B005/00
US08296982 | B2 | 2012-10-30 | Form for wristband with adjacent labels | 2009-11-11 | An improved identification band such as a wristband is provided with preprinted adhesive labels having a correlated unique identification code on the identification band and each of the labels. A label portion is separable from the wristband portion such that the wristband may be applied to a person or object to be identified. The label portion may be affixed to a specimen vial or other object to be identified as having a correlation with the person or object wearing the wristband. Individual labels are separable from the label portion, and the band portion in certain embodiments, so as to be applicable to other objects to be correlated to the person or object. | 1. An identification band, comprising: an elongated flexible strap consisting of a wristband portion and a label portion;the wristband portion consisting of a wristband, said wristband consisting of a primary identification zone, a secondary identification zone and a band portion, wherein the primary and secondary identification zones are each substantially co-extensive with the width of the flexible strap and the band portion is substantially less than the width of the flexible strap, and wherein the primary identification zone is configured to receive identification indicia and the secondary identification zone displays a unique identification code; andthe label portion being disposed immediately adjacent to and substantially co-extensive with the length of the band portion and consisting of a plurality of detachable labels, wherein the label portion is substantially less than the width of the flexible strap and each of the labels displays the unique identification code. 2. The identification band of claim 1, wherein the band and label portions are offset from a centerline of the flexible strap such that a junction between the band portion and the label portion approximates the centerline. 3. The identification band of claim 2, wherein the adjacent label and band portions together are substantially co-extensive with the width of the flexible strap. 4. The identification band of claim 2, wherein the junction between the band portion and the label portion comprises a line of weakness such that the label portion is separable from the band portion. 5. The identification band of claim 1, wherein the wristband portion is configured to form a closed loop. 6. The identification band of claim 5, wherein the wristband portion has a closure mechanism adjacent to the primary identification zone, disposed at a first end of the flexible strap and configured to receive the band portion. 7. The identification band of claim 6, wherein the closure mechanism comprises an adhesive closure or a snap closure. 8. The identification band of claim 6, wherein the label portion has an attachment mechanism disposed at a second end of the flexible strap and configured to attach the label portion to an article related to a specific wearer or object. 9. The identification band of claim 1, wherein the primary identification zone is configured to receive the identification indicia by hand-writing, printing or a pre-printed tag. 10. The identification band of claim 9, wherein the pre-printed tag is configured to adhere to the primary identification zone, be disposed under an adhesive transparent cover film on the primary identification zone, or inserted into a clear window on the primary identification zone. 11. The identification band of claim 1, wherein the elongated flexible strap defines a substantially rectangular structure. 12. An identification band, comprising: an elongated flexible strap defining a substantially rectangular structure consisting of a wristban... | Precision Dynamics Corporation | A44C005/00
US07959567 | B2 | 2011-06-14 | Device to enable quick entry of caloric content | 2007-10-22 | A monitoring apparatus that includes a sensor device and an I/O device in communication with the sensor device that generates derived data using the data from the sensor device. The derived data cannot be directly detected by the associated sensors. Alternatively, an apparatus that includes a wearable sensor device and an I/O device in communication with the sensor device that includes means for displaying information and a dial for entering information. Alternatively, an apparatus for tracking caloric consumption and caloric expenditure data that includes a sensor device and an I/O device in communication with the sensor device. The sensor device includes a processor programmed to generate data relating to caloric expenditure from sensor data. Alternatively, an apparatus for tracking caloric information for an individual that utilizes a plurality of classification identifiers for classifying meals consumed by the individual, each of the classification identifiers having a corresponding caloric amount. | 1. An apparatus for tracking caloric consumption and caloric expenditure of an individual, comprising: an I/O device adapted to enable a user to enter information relating to foods consumed by said individual, said information comprising a user's estimate of the relative size of a meal in relation to said user's normal-sized meal; anda processor in communication with said I/O device, said processor for receiving said user's estimate and programmed to generate data corresponding to the caloric content of said meal based on at least one of (a) aggregated data of the caloric content associated with other users' estimates of the relative size of meals in relation to said other users' normal-sized meals, and (b) a previous estimate of said user of the relative size of a previous meal in relation to said user's normal-sized meals. 2. The apparatus of claim 1, wherein processor is further programmed to cause said I/O device to provide the user with an interface enabling said user to enter said information relating to foods consumed by said individual. 3. The apparatus of claim 2, wherein said interface provides the user with a list of selectable identifiers corresponding to user's estimate. 4. The apparatus of claim 3, wherein said selectable identifiers comprise small, medium, large, and extra large. 5. The apparatus of claim 3, wherein said selectable identifiers relate to the user's estimate of the relative density of a meal in relation to said user's normal meal. 6. The apparatus of claim 3, wherein said selectable identifiers comprise a point value corresponding to said user's estimate. 7. The apparatus of claim 1, wherein said meal comprises a snack. 8. The apparatus of claim 1, wherein said processor is programmed to adjust data corresponding to the caloric content of said meal based on said user's weight and information relating to calories expended by said individual. | BodyMedia, Inc. | A61B005/00
US07970455 | B2 | 2011-06-28 | Ingestible device platform for the colon | 2005-05-19 | An ingestible pill platform for colon imaging is provided, designed to recognize its entry to the colon and expand in the colon, for improved imaging of the colon walls. On approaching the external anal sphincter muscle, the ingestible pill may contract or deform, for elimination. Colon recognition may be based on a structural image, based on the differences in diameters between the small intestine and the colon, and particularly, based on the semilunar fold structure, which is unique to the colon. Additionally or alternatively, colon recognition may be based on a functional image, based on the generally inflammatory state of the vermiform appendix. Additionally or alternatively, pH, flora, enzymes and (or) chemical analyses may be used to recognize the colon. The imaging of the colon walls may be functional, by nuclear-radiation imaging of radionuclide-labeled antibodies, or by optical-fluorescence-spectroscopy imaging of fluorescence-labeled antibodies. Additionally or alternatively, it may be structural, for example, by visual, ultrasound or MRI means. Due to the proximity to the colon walls, the imaging in accordance with the present invention is advantageous to colonoscopy or virtual colonoscopy, as it is designed to distinguish malignant from benign tumors and detect tumors even at their incipient stage, and overcome blood-pool background radioactivity. | 1. An ingestible pill, adapted to recognize entry to a colon, comprising: an ingestible-pill body, adapted to be swallowed by a living being and adapted for traveling down a gastrointestinal tract of the living being, the body including an expansion mechanism and an outer portion, wherein at least part of the outer portion is adapted to be moved by said expansion mechanism relative to an opposite part of the outer portion and wherein the part of the outer portion includes at least one radioactive-emission detector for detecting pathologies in the colon; anda colon-recognition trigger mechanism, associated with the ingestible-pill body and configured to produce a colon-entry signal, when the ingestible-pill body enters the colon of the living being,wherein the outer portion comprises a skin, andwherein the at least one radioactive-emission detector contacts the skin. 2. The ingestible pill of claim 1, wherein the colon-recognition trigger mechanism is selected from the group consisting of: an ultrasound sensor, configured to produce the colon-entry signal, based on the difference in structure between the small intestine and the colon,an MRI sensor, configured to produce the colon-entry signal, based on the difference in structure between the small intestine and the colon,a video camera, configured to produce the colon-entry signal, based on the difference in structure between the small intestine and the colon,an ultrasound sensor, configured to produce the colon-entry signal, upon imaging the semilunar fold structure, which is unique to the colon,an MRI sensor, configured to produce the colon-entry signal, upon imaging the semilunar fold structure, which is unique to the colon,a video camera, configured to produce the colon-entry signal, upon imaging the semilunar fold structure, which is unique to the colon,a radioactive-emission sensor, configured to produce the colon-entry signal, upon functional sensing the generally inflammatory state of the vermiform appendix,an optical-fluorescence-spectroscopy sensor, configured to produce the colon-entry signal, upon functional sensing the generally inflammatory state of the vermiform appendix,a pH sensor, configured to produce the colon-entry signal, upon sensing the pH environment of the colon, which is greater than 7,an enzyme sensor, configured to produce the colon-entry signal, upon sensing the enzyme environment of the colon,a chemical sensor, configured to produce the colon-entry signal, upon sensing the chemical environment of the colon,a bacterial sensor, configured to produce the colon-entry signal, upon sensing the flora environment of the colon, anda combination of the above. 3. The ingestible pill of claim 1, wherein the expansion mechanism is configured to expand the ingestible pill body into an expanded structure, having expanded length and width parameters, which are between a factor of two and a factor of five greater than the initial length and width parameters, so as to push the outsi... | Spectrum Dynamics LLC | A61B006/00 A61B005/05 A61B008/00
US07942816 | B2 | 2011-05-17 | Psychotic manifestation and mental state evaluation apparatus and evaluation method | 2006-08-09 | A psychotic manifestation and mental state evaluation apparatus and method which can individually discriminate among symptoms by taking advantage of stimulating Noh mask images and also evaluate with a high probability whether a person is suffering from a specific symptom at the time of diagnosing, clinical examination, assessment, or counseling. | 1. A psychotic manifestation and mental state evaluation apparatus comprising: a database which stores data of stimulating Noh mask images that express various kinds of human emotions according to changes of the angles at which the masks are photographed from tilting up to tilting down;an image display apparatus for displaying the stimulating images;means for displaying questions, referred to as emotion items, about the feeling elicited by the stimulating images that have been one by one shown on the image display apparatus as well as displaying a response from a subject being evaluated to each question based on response data which includes affirmative response data, reaction data and a reaction-time value;means for gathering statistical data of the response to each emotion item with respect to the emotion elicited by each stimulating image; andprocessing means for evaluating psychotic manifestation and mental state based on the gathered statistical data including the number of affirmative responses, a number of expression evaluation variations, a balance coefficient of expression and a deviation score from healthy subjects' reference data that have been obtained by recalculating affirmative responses from each emotion item,wherein said database stores predetermined combinations of stimulating images photographed at different angles that have been selected from said stimulating images and plural emotion items that have been selected from said emotion items with respect to more than one psychotic manifestation and mental state,wherein said processing means collects a reaction value and a reaction-time value from the response data of the response to each emotion item that corresponds to each stimulating image, collects reaction data by comparing the reaction and reaction-time values with reference values, and collects attention data based on the reaction data with respect to each stimulating image, and evaluating psychotic manifestation and mental state based on the collected data, thereby creating attention data in regard to a specific emotion item with respect to a specific stimulating image,wherein the attention data, when used, permits said psychotic manifestation to be easily discriminated, said psychotic manifestation being one of organic brain damage, schizophrenia and depressive tendency,wherein said attention data with respect to specified stimulating image and specified emotion item has a three-dimensional structure in which plus values and minus values coexist and is displayed on said image display apparatus to permit a visual understanding of the condition of the subject,wherein after the collection, said processing means identifies a specific symptom that corresponds to said psychotic manifestation based on the attention data and a result of comparing of the selected stimulating images and the emotion items to the stimulating image and emotion item stored in the database, andwherein individual data is created for each specific symptom ... | Shinji Satoh Seiko Minoshita | A61B005/16
US08060193 | B2 | 2011-11-15 | Systems and methods for graphic display of ST-segment deviation | 2011-03-09 | Systems and methods are provided for monitoring a patient and graphically representing ST-segment deviations. A receiver acquires, from a plurality of leads, electrocardiogram (ECG) signals that each includes an ST-segment. A processor processes the ECG signals to determine baseline and current ST-segment deviations relative to an isoelectric value. A display module displays a graph that includes a first axis representing ST-segment deviation values and a second axis representing the plurality of leads. At each location along the second axis representing a respective lead, the graph displays a first set of graphic indicia representing the baseline ST-segment deviations and a second set of graphic indicia representing the current ST-segment deviations. In certain embodiments, the graphic indicia are represented by bar graphs. In other embodiments, the graphic indicia are represented by symbols that may be connected by line segments. | 1. A method for monitoring a patient and graphically representing ST-segment deviations, the method comprising: acquiring electrocardiogram (ECG) signals from a plurality of leads, each ECG signal including a sequence of ST-segments between a QRS complex and a T wave;determining a baseline ST-segment deviation from an isoelectric value for each of the plurality of leads;determining a current ST-segment deviation from the isoelectric value for each of the plurality of leads;displaying a graph comprising: a first linear axis representing ST-segment deviation values; anda second linear axis representing the plurality of leads, wherein each of the plurality of leads is associated with a respective location along the second linear axis;displaying a first set of bar graphs representing the baseline ST-segment deviations, wherein each first bar graph is displayed with respect to one of the locations corresponding to a respective lead and is aligned relative to the first linear axis based on the baseline ST-segment deviation for the respective lead; anddisplaying a second set of bar graphs representing the current ST-segment deviations, wherein each second bar graph is displayed with respect to one of the locations corresponding to a respective lead and is aligned relative to the first linear axis based on the current ST-segment deviation for the respective lead;wherein each second bar graph has a width in the direction of the second linear axis that is narrower than a width in the direction of the second linear axis of the corresponding first bar graph for a respective lead;wherein each second bar graph has an origin with respect to the first axis at the baseline ST-segment deviation for a respective lead and extends along the first axis in a positive or negative direction depending on whether the current ST-segment deviation is respectively greater or less than the baseline ST-segment deviation; andwherein, if the current ST-segment deviation is positive and is less than the baseline ST-segment deviation for a particular lead, the second bar graph for the particular lead extends in the negative direction into the corresponding first bar graph. 2. The method of claim 1, wherein displaying each first bar graph comprises: displaying the first bar graph between a first value relative to the first linear axis and a second value relative to the first linear axis, wherein the first value corresponds to the isoelectric value, and wherein the second value corresponds to the baseline ST-segment deviation for the respective lead. 3. The method of claim 2, wherein displaying each second bar graph comprises: displaying the second bar graph between the second value relative to the first linear axis and a third value relative to the first linear axis, wherein the third value corresponds to the current ST-segment deviation for the respective lead. 4. The method of claim 1, further comprising: highlighting, at each of the locations along the second linear axis, at l... | Mindray DS USA, Inc. | A61B005/04
US07943161 | B2 | 2011-05-17 | Contrast agent coated medical device | 2005-12-08 | A magnetically opaque medical device is disclosed wherein a contrast agent is incorporated into the actual device. The medical device is generally comprised of a base material forming the structure of the device and a contrast agent, such as tocopherol and tocopherol derivative solutions or suspensions, gadolinium, or nickel sulfate integrated into the base material itself or posited on a substantial portion of an exterior surface of the device. The device may include other additional functional agents and layers. | 1. An implantable medical device comprising: a base material forming a structure for introduction into a vessel of a patient, the base material comprising silicone or a silicone derivative, the structure having at least one surface,at least one contrast agent layer posited atop the base material and over at least one surface of the structure, the contrast agent comprising tocopherol; and;a non-porous poly(p-xylylene)polymer layer posited over the contrast agent layer, wherein the non-porous layer discourage the contrast agent being dissipation into surrounding tissue. 2. The implantable medical device of claim 1, wherein the contrast agent is tocopherol, wherein the contrast agent layer comprises about 0.01 mg to about 10 mg of the contrast agent per cm2 of the gross surface area of the structure, wherein the medical device further comprises a therapeutic agent coating layer posited on the non-porous layer, the therapeutic agent coating layer comprising about 100 μg to about 300 μg of a therapeutic agent per 0.001 inch of the total coating thickness on the device surface. 3. The implantable medical device of claim 1, wherein the contrast agent is applied to the structure by one of dipping, spraying, bulk distribution, plasma deposition or vapor phase deposition. 4. An implantable medical device comprising: a base material forming a structure for implantation into a patient, a magnetically opaque contrast agent intermingled in the base material, and a non-porous poly(p-xylylene)polymer coating layer posited over a surface of the base material, wherein the base material comprises silicone or a silicone derivative and forms the main structure of the device, wherein the non-porous layer discourage the contrast agent being dissipation into surrounding tissue. 5. The implantable medical device of claim 1, wherein the contrast agent is tocopherol. 6. The implantable medical device of claim 1, wherein the contrast agent is attached to the surface of an adhesion promotion coating layer formed on an outer surface of the base material. 7. The implantable medical device of claim 6, wherein the coating layer has a thickness between 2 to 50 Å. 8. The implantable medical device of claim 1, wherein the contrast agent is injected or otherwise included in wells, grooves or channel formations integrated into the shape of the device as formed by the base material. 9. The implantable medical device of claim 1, wherein the device comprises a stent, catheter, cannulae, wire guide, balloon or bladder. 10. The implantable medical device of claim 1, wherein the device is configured for vascular system implantation. 11. The implantable medical device of claim 1, wherein the device is a stent. 12. The implantable medical device of claim 4, wherein upon implantation in a patient the non-porous layer is configured to directly face tissue or blood. 13. The implantable medical device of claim 4, wherein the device comprises a stent, catheter, cannulae, wire guide, b... | Cook Incorporated | A61B005/055 A61F002/06
US07967757 | B2 | 2011-06-28 | Arterial pressure-based, automatic determination of a cardiovascular parameter | 2008-10-23 | One or more cardiovascular parameters is estimated as a function of the arterial pressure waveform, in particular, using at least one statistical moment of a discrete representation pressure waveform having an order greater than one. Arterial compliance, the exponential pressure decay constant, vascular resistance, cardiac output, and stroke volume are examples of cardiovascular parameters that can be estimated using various aspects of the invention. In one embodiment of the invention, not only are the first four moments (mean, standard deviation, skewness, and kurtosis) of the pressure waveform used to estimate the cardiovascular parameter(s) of interest, but also heart rate, statistical moments of a set of pressure-weighted time values, and certain anthropometric patient measurements such as age, sex, body surface area, etc. | 1. A method for determining a cardiovascular parameter of a subject comprising: sensing, over a period corresponding to at least one cardiac cycle, an input signal that indicates arterial blood pressure;computing, using at least a portion of a processing system, a set of pressure-weighted time values, each corresponding to the product of a sensing time, relative to an initial time, and arterial pressure at the sensing time; andestimating the cardiovascular parameter as a function of the set of pressure-weighted time values. 2. A method as in claim 1, in which the cardiovascular parameter is arterial compliance. 3. A method as in claim 2, further comprising: determining an approximating function relating a set of clinically determined reference measurements to arterial compliance, in which the approximating function is a function of the statistical moment(s) of the pressure-weighted time values and of a set of anthropometric values;computing the statistical moment(s) of the pressure-weighted time values, as well as measuring the set of anthropometric values of the subject;estimating the arterial compliance of the subject by evaluating the approximating function with the computed statistical moment(s) of the pressure-weighted time values, as well as with the measure set of anthropometric values of the subject. 4. A method as in claim 1, in which the cardiovascular parameter is vascular resistance. 5. A method as in claim 1, in which the cardiovascular parameter is cardiac output. 6. A method as in claim 1, in which the cardiovascular parameter is stroke volume. 7. A method as in claim 1, in which the cardiovascular parameter is a pressure decay constant. 8. A method as in claim 1, further comprising estimating the cardiovascular parameter as a function of the kurtosis of the pressure-weighted time values. 9. A method as in claim 1, further comprising estimating the cardiovascular parameter as a function of the standard deviation of the pressure-weighted time values. 10. A method as in claim 1, further comprising estimating the cardiovascular parameter as a function of the skewness of the pressure-weighted time values. 11. A method as in claim 1, further comprising: measuring a predetermined set of anthropometric parameters of the subject; andestimating the cardiovascular parameter as a function of the measured anthropometric parameters. 12. A method as in claim 1, further comprising: computing at least two statistical moments not only of the input signal, but also of the set of pressure-weighted time values;in which the approximating function is a function of the computed statistical moments of both the input signal and of the set of pressure-weighted time values. 13. A method as in claim 1, further comprising: measuring the subject's heart rate;estimating stroke volume as a function of the product of the estimated arterial compliance and the standard deviation the input signal; andcomputing cardiac output value as a function of the pro... | Edwards Lifesciences Corporation | A61B005/02
US07962204 | B2 | 2011-06-14 | Method of recommending neurophysiological therapies | 2008-11-18 | A method and system for utilizing neurophysiologic information obtained by techniques such as quantitative electroencephalography (QEEG), electrode recordings, MRI in appropriately matching patients with therapeutic entities is disclosed. The present invention enables utilization of neurophysiologic information, notwithstanding its weak correlation with extant diagnostic schemes for mental disorders, for safer and expeditious treatment for mental disorders, discovering new applications for therapeutic entities, improved testing of candidate therapeutic entities, inferring the presence or absence of a desirable response to a treatment, and deducing the mode of action of one or more therapeutic entities. In particular, methods for effectively comparing neurophysiologic information relative to a reference set are disclosed along with database-based tools for deducing therapeutic entity actions on particular patients such that these tools are readily accessible to remote users. | 1. A method for recommending treatment, comprising: a) providing; i) a first set of quantified neurophysiologic information obtained from a patient having a behaviorally diagnosed psychiatric disorder;ii) a second set of quantified neurophysiologic information obtained from at least one reference population of individuals, wherein said second information set comprises at least one therapy responsivity profile, wherein said profile includes a first treatment modality, wherein said first treatment modality is recommended independently of said diagnosed psychiatric disorder and associated symptoms;b) comparing said first set of quantified neurophysiologic information with said second set of quantified neurophysiologic information by cluster analysis, under conditions such that a correlation between said patient and said at least one therapy responsivity profile is indicated, wherein said comparing is performed by a computer;c) recommending said first treatment modality for said patient based upon said indicated correlation. 2. The method of claim 1, wherein said therapy responsivity profile comprises determined effects of said at least one treatment modality. 3. The method of claim 1, wherein said indicated correlation comprises a relatively high correlation such that said at least one reference population was relatively effectively treated with said at least one treatment modality. 4. The method of claim 1, wherein said indicated correlation comprises a relatively low correlation such that said at least one reference population was relatively ineffectively treated with said at least one treatment modality. 5. The method of claim 1, further comprising repeating step (b) and step (c) after administering said first treatment modality to said patient, wherein said repeated step (c) recommends a second treatment modality. 6. The method of claim 1, wherein said first and second sets of quantified neurophysiological information comprise electroencephalographic information. 7. The method of claim 6, wherein said electroencephalographic information was recorded by electrodes placed in accordance with the International 10/20 system. 8. The method of claim 1, wherein said behaviorally diagnosed psychiatric disorder is selected from the group consisting of a major depressive disorder, a psychosomatic disorder, an atypical asthma disorder, an anxiety disorder, infancy disorders, childhood disorders, adolescence disorders, eating disorders, delirium disorders, dementia disorders, amnestic disorders, general cognitive disorders, impulse control disorders, mood disorders, personality disorders, hypoactive sexual desire disorders, sleep disorders, somatoform disorders, and substance-related disorders. 9. The method of claim 1, wherein said treatment modality is selected from the group consisting of psychiatric therapy, including for example therapeutic entity therapy, medicinal therapy, psychotherapy, verbal therapy, convulsive therapy, neuromodulation ther... | CNS Response, Inc. | A61B005/04
US20110118578 | A1 | 2011-05-19 | HYPOGLYCEMIC TREATMENT METHODS AND SYSTEMS | 2009-11-17 | A system for treating hypoglycemia includes a continuous blood glucose monitor that monitors blood glucose levels of a patient on a continuous basis over a period of time. An alarm and injector apparatus is configured to be worn by the patient and includes a glucagon delivery mechanism for automatically delivering glucagon to the patient if a hypoglycemic condition is detected using the continuous blood glucose monitor. | 1. A method of treating hypoglycemia, comprising: detecting a hypoglycemic condition using a continuous blood glucose monitor that monitors blood glucose levels of a patient on a continuous basis over a period of time;providing an alarm alerting of the hypoglycemic condition; anddelivering glucagon to the patient using an alarm and injector apparatus comprising a glucagon delivery mechanism for delivering glucagon to the patient. 2. The method of claim 1 further comprising providing an alarm signal to the alarm and injector apparatus in response to detecting the hypoglycemic condition. 3. The method of claim 2 further comprising determining when the blood glucose level of the patient falls below a predetermined level. 4. The method of claim 3, wherein the alarm signal is provided from the continuous blood glucose monitor to the alarm and injector apparatus when the blood glucose level of the patient falls below the predetermined level. 5. The method of claim 3, wherein the alarm signal is provided from a blood glucose meter to the alarm and injector apparatus when the blood glucose level of the patient falls below the predetermined level. 6. The method of claim 2, wherein the alarm and injector apparatus includes a controller controlling an alarm mechanism providing the alarm. 7. The method of claim 6, wherein, after the step of providing the alarm, the controller looking for a patient-initiated input signal provided using an input mechanism within a predetermined time period. 8. The method of claim 7, wherein if the patient-initiated input signal is not received by the controller within the predetermined time period, the controller actuating the glucagon delivery mechanism. 9. A method of treating hypoglycemia, comprising: detecting a hypoglycemic condition using a continuous blood glucose monitor that monitors blood glucose levels of a patient on a continuous basis over a period of time;providing a first alarm signal to an alarm and injector apparatus worn by the patient if the hypoglycemic condition is above a predetermined blood glucose level;providing a second alarm signal to the alarm and injector apparatus worn by the patient if the hypoglycemic condition is below the predetermined blood glucose level; andproviding an alarm using the alarm and injector apparatus in response to one or both of the first alarm signal and the second alarm signal, the alarm and injector apparatus comprising a glucagon delivery mechanism for delivering glucagon to the patient. 10. The method of claim 9 further comprising delivering glucagon using the glucagon delivery mechanism upon receipt of the second alarm signal by the alarm and injector apparatus. 11. The method of claim 10, wherein the step of delivering the glucagon is provided automatically if the alarm is not acknowledged within a preselected period of time. 12. The method of claim 10, wherein the step of delivering the glucagon is initiated manually. 13. The method of claim 9, wherein the ... | ROCHE DIAGNOSTICS OPERATIONS, INC. | A61B005/145 A61M005/00
US07937276 | B2 | 2011-05-03 | Computerized system and method for generating an immunization schedule in a healthcare environment | 2009-01-08 | A computerized method and system for generation an immunization schedule for a person in healthcare environment is provided. The system receives a request for an immunization schedule for a person and obtains information from an electronic medical record of the person stored within a comprehensive healthcare system. The system obtains one or more immunization schedules and utilizes the information from the electronic medical record of the person and the immunization schedule to generate a customized immunization schedule for the person. In another embodiment, the system receives immunization information for a person and stores the immunization information in the electronic medical record of the person within a comprehensive healthcare system. | 1. A computer-implemented method in a computer system having a processor and memory for preventing an inappropriate live virus combination from being administered to a person, the method comprising: receiving, from a clinician utilizing a first computer process, an indication of two or more live virus immunizations to be administered to a person;generating, utilizing a second computer process, a custom immunization schedule for the person, wherein the custom immunization schedule is based on at least one immunization administered to the person during a former clinical visit, the two or more live virus immunizations to be administered to the person as received from the clinician, and a standard schedule of recommended immunizations;accessing, utilizing a third computer process, a database comprised of information identifying one or more combinations of live virus immunizations that create a potential adverse reaction when administered in combination;determining, with a fourth computer process, the two or more live virus immunizations constitute an inappropriate live virus combination based, at least in part, on the database comprised of information identifying one or more combinations of live virus immunizations that create a potential adverse reaction;updating, with a fifth computer process, the custom immunization schedule for the person in response to receiving an input regarding an immunization to be administered during a present clinical visit;generating an output, with a sixth computer process, indicating that the two or more live virus immunizations constitute an inappropriate live virus combination;the first, the second, the third, the fourth, the fifth, and the sixth computer processes are executed utilizing one or more computing devices. 2. The method of claim 1, further comprising: comparing the two or more live virus immunizations to the information regarding inappropriate virus combinations in the database to determine the two or more viruses constitute the inappropriate live virus combination. 3. A system in a computerized environment for preventing an inappropriate live virus combination from being administered to a person, the system comprising: a receiving module using a processor and memory for receiving two or more live virus immunizations to be administered to a person;a determining module using a processor and memory for determining whether the two or more live virus immunizations constitute an inappropriate live virus combination, wherein an inappropriate live virus combination is determined by accessing a database comprised of information identifying one or more combinations of live virus immunizations that create a potential adverse reaction when administered in combination;an updating module for updating a custom immunization schedule for the person in response to receiving an input regarding the immunization to be administered during a present clinical visit, wherein the custom immunization schedule is based on at least... | Cerner Innovation, Inc. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US08597185 | B2 | 2013-12-03 | Residual-based monitoring of human health | 2011-01-31 | Improved human health monitoring is provided in the context of sensor measurements of typical vital signs and other biological parameters, by a system and method using an empirical model of the parameters and disposed to estimate values of the parameters in response to actual measurements. Residuals resulting from the difference between the estimates and actual measurements are analyzed for robust indications of incipient health issues. Residual analysis is both more robust and more sensitive than conventional univariate range checking on vital signs. | 1. A system for monitoring the health status of a human, comprising: an ambulatory device disposed to receive multiple biological parameters from a plurality of sensors configured to be attached to a monitored human;a computer disposed to obtain monitored observations of multiple biological parameters from said ambulatory device; anda computer-accessible memory for storing a set of exemplary observations of said multiple biological parameters characteristic of the interrelationships between said multiple biological parameters under normal health conditions for the monitored human;said computer being configured to generate an estimate of at least one of said multiple biological parameters in a monitored observation obtained from said ambulatory device, as a linear combination of at least some of said set of exemplary observations, determined from a kernel-based comparison of the monitored observation and at least some exemplary observations; andsaid computer being further configured to compare said estimate to the corresponding value of the at least one of said multiple biological parameters in said monitored observation, to generate a residual, and make a determination of the health status of the human there from. 2. The system according to claim 1, wherein said ambulatory device communicates with said sensors by means of local wireless communications. 3. The system according to claim 2, wherein said local wireless communications is IEEE 802.15.1. 4. The system according to claim 1, wherein said plurality of sensors includes an oximeter. 5. The system according to claim 1, wherein said plurality of sensors includes an optical blood glucose meter. 6. The system according to claim 1, wherein said plurality of sensors includes an accelerometer. 7. The system according to claim 1, wherein said plurality of sensors includes at least a pair of temperature sensors disposed to measure ambient temperature and a temperature of the monitored human. 8. The system according to claim 1, wherein said ambulatory device comprises a cell radio and said observations of multiple biological parameters are communicated to said computer at least in part over a cellular communications network. 9. The system according to claim 1, wherein said ambulatory device records said multiple biological parameters and transfers them periodically upon being attached by a cable to an internet-connected appliance, and said appliance is configured to upload said multiple biological parameters to said computer over an internet connection. 10. The system according to claim 1, wherein said set of exemplary observations of said multiple biological parameters comprises at least some observations previously obtained from the monitored human under normal health conditions using such an ambulatory device. 11. The system according to claim 1, wherein said computer is further configured to use said monitored observation for localization so that a subset of said exemplary observations i... | Ventura Gain LLC | G06N003/063 G06N003/02
US08380295 | B2 | 2013-02-19 | Detection of waveform artifact | 2009-10-30 | An implantable medical device system including a physiological sensor detects signal artifact in a signal waveform acquired by the sensor. Features of individual waveforms in the sensor signal are extracted. Sample waveforms are classified by expert observation into at least two classes including an artifact class. A distribution range for each of the extracted features from the sample waveforms is determined for each of the classes. Waveform classification criteria are established in response to the determined distribution ranges. | 1. A medical device system for monitoring a physiological signal in a patient and detecting artifact in the physiological signal, the system comprising: an implantable physiological sensor for sensing the physiological signal;a display for displaying the physiological signal to enable an expert to classify individual waveforms of the physiological signal based on visual observation, the individual waveforms classified into at least two classes comprising an artifact class;a user interface for receiving user input to establish the classification of each of the individual waveforms;a memory for storing the user input classifications of the individual waveforms;a processor configured to:receive the implantable physiological sensor signal;separate the signal into a plurality of individual waveforms;extract a plurality of features from each individual waveform;determine a distribution range for the plurality of features for each of the classes, andestablish waveform classification criteria in response to the determined distribution ranges; andan implantable processor adapted to be coupled to an implantable physiological sensor, the implantable processor configured to:acquire unknown waveforms from the implantable physiological sensor,extract the plurality of features from the unknown waveforms,classify the unknown waveforms according to the established waveform classification criteria, anddetect signal artifact in response to a waveform being classified in the artifact class, wherein the at least two classes further comprise an atypical class and a typical class, the implantable processor further configured to: determine a percentage of the unknown waveforms classified as each of the artifact class, atypical class, and typical class;compute a monitoring metric using the unknown waveforms; andreport the monitoring metric with the percentage of one of the artifact, atypical and typical classes to establish a confidence level of the reported metric. 2. The system of claim 1 wherein determining the distribution range comprises computing a morphology score as a function of the extracted features and determining the distribution range of the morphology score for each of the classes. 3. The system of claim 1 wherein the implantable processor is further configured to compute a metric for monitoring a physiological condition of the patient, the metric computed as a function of at least one of the features extracted from the unknown waveforms in response to an unknown waveform not being classified in the artifact class. 4. The system of claim 1 wherein the plurality of features extracted comprises a non-physiological feature of the waveform, the non-physiological feature used for classifying the unknown waveforms; and the implantable processor further configured to extract a physiological feature from an unknown waveform in response to the unknown waveform not being classified in the artifact class, andcompute a metric as a function of the physiological feat... | Medtronic, Inc. | A61B005/04
US08364247 | B2 | 2013-01-29 | Methods for determining the sex of birds' eggs | 2010-12-06 | The invention relates to methods for determining the sex of birds' eggs. It can be used for an early determination of the sex, in particular in poultry farming, and in this respect also with non-incubated eggs. In the method in accordance with the invention, electromagnetic radiation is emitted onto the blastodisk of an egg by a radiation source and, after a switching off of the radiation source, the decay behavior of the autofluorescence intensity excited by the electromagnetic radiation is detected by a detector at the irradiated region of the blastodisk with time resolution and spectral resolution for at least one wavelength of the autofluorescence. The fractal dimension DF is calculated using the determined measured intensity values and the value of the fractal dimension DF is compared with a species-specific and sex-specific limit value, wherein, when the limit value is exceeded, the respective egg is classified as female and, when it is not reached, the egg is classified as male. | 1. A method for determining the sex of birds' eggs, said method comprising the steps of: (a) emitting electromagnetic radiation to the blastodisk of an egg by a radiation source,(b) switching off of the radiation source,(c) detecting the decay behavior of the autofluorescence intensity excited by the electromagnetic radiation with a detector at the irradiated region of the blastodisk with time resolution and spectral resolution for at least one wavelength of the autofluorescence,(d) calculating, using a microprocessor, fractal dimension DF using the determined measured intensity values,(e) comparing the value of the fractal dimension DF with a species-specific and sex-specific limit value, and(f) classifying the respective egg as female when the limit value is exceeded and classifying the egg as male when the limit value is not reached. 2. A method in accordance with claim 1, characterized in that the fractal dimension DF1 is determined by determining the smallest error squares. 3. A method in accordance with claim 1, characterized in that the fractal dimension DF2 is determined by determining the difference-autocorrelation function C(t) of the autofluorescence intensity decay behavior. 4. A method in accordance with claim 1, wherein the blastodisk is irradiated with monochromatic electromagnetic radiation and a fluorescence spectrometer is used for determining the autofluorescence, with the irradiation taking place so that the beam for the autofluorescence excitation is incident onto the blastodisk so that no radiation is reflected on the same axis. 5. A method in accordance with claim 1, wherein the detection is carried out with a known, constant sampling rate. 6. A method in accordance with claim 5, wherein the detection is carried out with a constant sampling rate in the range of 50 ps to 1000 ps. 7. A method in accordance with claim 1, wherein electromagnetic radiation from the wavelength range of UV radiation is used for the fluorescence excitation. 8. A method in accordance with claim 1, wherein the electromagnetic radiation used for the fluorescence excitation is directed to the blastodisk by means of a first optical fiber and the electromagnetic radiation occurring as a consequence of autofluorescence is directed to the detector via a second optical fiber and in this respect the two optical fibers are aligned at an obliquely inclined angle with respect to one another. 9. A method in accordance with claim 1, wherein the determination is carried out in ovo. | Fraunhofer-Gesellschaft zur Foerderung der Angwandten Forschung E.V. | A61B006/00
US08801622 | B2 | 2014-08-12 | System and method for pulse rate calculation using a scheme for alternate weighting | 2011-03-07 | Embodiments of the present invention relate to a method for analyzing pulse data. In one embodiment, the method comprises receiving a signal containing data representing a plurality of pulses, the signal generated in response to detecting light scattered from blood perfused tissue. Further, one embodiment includes performing a pulse identification or qualification algorithm on at least a portion of the data, the pulse identification or qualification algorithm comprising at least one constant, and modifying the at least one constant based on results obtained from performing the pulse identification or qualification algorithm, wherein the results indicate that a designated number of rejected pulses has been reached. | 1. A method for analyzing pulse data, comprising: receiving a signal at a pulse oximeter, the signal representative of light detected after being scattered by tissue and containing data representing a plurality of pulses;performing a pulse qualification algorithm on at least a portion of the data using a processor of the pulse oximeter, the pulse qualification algorithm utilizing at least one constant; andmodifying the at least one constant using the processor when results obtained from performing the pulse qualification algorithm meet a condition. 2. The method of claim 1, comprising: emitting the light from a light source of a sensor into the tissue;detecting the light with a detector of the sensor after the light has been scattered by the tissue; andgenerating the signal with the detector. 3. The method of claim 1, wherein the pulse qualification algorithm detects potential pulse amplitudes in the portion of the data. 4. The method of claim 3, wherein the condition comprises the potential pulse amplitudes being generally below a threshold. 5. The method of claim 1, wherein modifying the at least one constant comprises swapping a primary constant value with an alternate constant value. 6. The method of claim 1, comprising keeping a count of consecutive pulses rejected by the pulse qualification algorithm. 7. The method of claim 6, wherein the condition comprises the count of consecutive pulses rejected by the pulse qualification algorithm reaching a value. 8. The method of claim 1, comprising keeping a count of consecutive pulses qualified by the pulse qualification algorithm. 9. The method of claim 8, wherein the condition comprises the count of consecutive pulses qualified by the pulse qualification algorithm reaching a value. 10. A system for analyzing pulse data, comprising: a receiver configured to receive a signal representative of light detected after being scattered by tissue, the signal containing data representing a plurality of pulses;a filter configured to run a pulse qualification algorithm on at least a portion of the data, the pulse qualification algorithm comprising at least one constant; anda subsystem configured to modify the at least one constant when results obtained from running the pulse qualification algorithm meet a condition. 11. The system of claim 10, wherein the subsystem is configured to change a value of the at least one constant to an alternate value when the results obtained from running the pulse qualification algorithm meet the condition by indicating that a designated number of rejected pulses has been reached. 12. The system of claim 10, wherein the subsystem is configured to change a value of the at least one constant to an alternate value when the results obtained from running the pulse qualification algorithm meet the condition by indicating that a designated number of qualified pulses has been reached. 13. The system of claim 10, comprising a sensor, wherein the sensor comprises: an emitter co... | Covidien LP | A61B005/02
US08398568 | B2 | 2013-03-19 | Methods and instruments for assessing bone fracture risk | 2011-01-14 | Methods and instruments for assessing bone, for example fracture risk, in a subject in which a test probe is inserted through the skin of the subject so that the test probe contacts the subject's bone and the resistance of the test bone to microscopic fracture by the test probe is determined. Macroscopic bone fracture risk is assessed by measuring the resistance of the bone to microscopic fractures caused by the test probe. The microscopic fractures are so small that they pose negligible health risks. The instrument may also be useful in characterizing other materials, especially if it is necessary to penetrate a layer to get to the material to be characterized. | 1. A method for assessing bone fracture risk in a subject without surgically exposing the bone surface, the method comprising: inserting a probe assembly through the skin and soft tissue down to a bone, the probe assembly comprising a reference probe having a tubular section and a test probe slidably disposed within the tubular section;establishing a stationary contact point directly on the surface of the bone with the reference probe;inserting the test probe a microscopic distance into the bone to create one or more microscopic fractures in the bone; anddetermining one or more of the following parameters: the extent of insertion of the penetrating end of the test probe into the bone;the resistance of the bone to penetration of the test probe;the resistance of bone to removal of the test probe after it penetrates the bone;thereby assessing bone fracture risk in the subject. 2. The method according to claim 1, wherein the microscopic fractures are on the order of 0.01 cubic millimeters or smaller. 3. The method according to claim 1, wherein the test probe is a single use test probe, and wherein the test probe is discarded after use. 4. The method according to claim 1, wherein determination of the one or more parameters further comprises determining a distance the test probe inserts into the bone under a predetermined force. 5. The method according to claim 1, wherein the method does not comprise exposing or visualizing the bone surface. 6. The method according to claim 1, wherein the method is applied to a volume of bone allowing the one or more parameters to be determined over at least several osteons. 7. The method according to claim 1, wherein the test probe is symmetrically sharpened. 8. The method according to claim 1, wherein the microscopic fractures are on the order of 0.01 cubic millimeters or smaller, and wherein the test probe is a single use test probe discarded after use, and wherein determination of the one or more parameters further comprises determining a distance the test probe inserts into the bone under a predetermined force, and wherein the method does not comprise exposing or visualizing the bone surface, and wherein the method is applied to a volume of bone allowing the one or more parameters to be determined over at least several osteons. 9. A method for assessing bone fracture risk in a subject without surgically exposing the bone surface, the method comprising: inserting a probe assembly through the skin and soft tissue down to a bone, the probe assembly comprising a reference probe having a tubular section and a test probe slidably disposed within the tubular section;establishing a stationary contact point directly on the surface of the bone with the reference probe;inserting the test probe into the bone to create a yield in a small probed volume of the bone; anddetecting irreversible changes in force v. distance curves by probing the yield to determine bone toughness,thereby assessing bone fracture risk in the su... | The Regents of the University of California | A61B005/103 A61B005/117 G01N003/48 G01N003/42 G01N003/32 G01N003/44
US08758262 | B2 | 2014-06-24 | Respiratory disease monitoring system | 2010-11-24 | An automated system for monitoring respiratory diseases, such as asthma, provides noninvasive, multimodal monitoring of respiratory signs and symptoms that can include wheeze and cough. Some embodiments employ a mobile device, such as a cell phone, in which raw data from a microphone and an accelerometer are processed, analyzed, and stored. Data can be collected continuously. Time domain and frequency domain analyses of signals to determine, e.g., energy, duration, and spectral content of candidate sounds can be employed to discriminate symptoms of interest from background sounds and to establish significance. Accelerometer signals are analyzed to determine activity levels. Analyses of a user's symptoms and activity level prior to, during, and after an event can provide meaningful determinations of disease severity and predict future respiratory events. The system can provide a summary of data, as well as an alarm when symptom severity reaches a threshold. | 1. An asthma monitoring system, comprising: a sound sensing device that detects sound over at least a portion of the audible spectrum and converts the detected sound to a sound signal;an acceleration sensing device that detects acceleration and converts the detected acceleration to an acceleration signal;an output device; anda computation module, operably connected to the sound sensing device, the acceleration sensing device, and the output device, the computation module comprising:a memory, containing instructions for processing the sound signal and the acceleration signal, wherein the instructions comprise at least one algorithm that converts a time waveform of the sound signal into a frequency representation and performs time-domain analysis and frequency-domain analysis of at least one of the sound signal and the acceleration signal, and wherein the instructions further comprise an algorithm that measures a spectral-peak-to-noise ratio of the sound signal;a communication module configured to transmit signals to an external device; anda processor module, operably connected to the memory, the communication module, and the output device, and configured to receive the sound signal and the acceleration signal;wherein the processor module retrieves the instructions from the memory, processes the sound signal and the acceleration signal according to the instructions, stores data generated according to the instructions, sends at least one signal generated according to the instructions to the output device, and sends data through the communication module to the external device according to the instructions; wherein the at least one signal is generated based on the frequency representation, the time-domain analysis, the frequency-domain analysis, and the spectral-peak-to-noise ratio. 2. The system of claim 1, wherein the instructions further comprise at least one algorithm capable of discriminating an asthma symptom from another sonic event in the sound signal; wherein the at least one signal is generated further based on a result of the discriminating. 3. The system of claim 1, wherein the instructions further comprise an algorithm that measures peaks of the frequency representation at frequencies in about the 200-800 Hz range, and compares these to a predetermined value associated with wheezing that is stored in the memory; wherein the at least one signal is generated further based on the peaks of the frequency representation. 4. The system of claim 1, wherein the instructions further comprise an algorithm that measures an amplitude and duration of a sonic event of the sound signal and measures a power spectral density over a selected band of frequencies, and compares the power spectral density to a predetermined value associated with coughing that is stored in the memory; wherein the at least one signal is generated further based on the power spectral density. 5. The system of claim 1, wherein the instructions further comprise an algorithm that ... | University of Rochester | A61B005/08
US09028258 | B2 | 2015-05-12 | Combined cognitive and physical therapy | 2010-11-09 | The present invention provides method and apparatus to perform combined cognitive and motor rehabilitation on a computerized non-portable system or on single portable device. A patient can play a variety of games that require the patient to perform a variety of memory exercises which involve physical exertion. The activities of the patient are monitored with pattern analysis software which provides feedback to the patient. The feedback can include voice synthesis, video guidance, progression messages etc. Patient data obtained while the patient is performing each of the memory exercises is stored locally on a database module and then uploaded to a cloud server. A remote psychologist/psychiatrist monitors the patient by logging into the same cloud, and updating cognition exercises. The same therapist can have live chats with the patient for further interaction and coaching. | 1. A method of providing cognitive and physical rehabilitation to a patient with a computer, comprising: having the patient play a virtual reality game through an avatar that simultaneously elicits cognitive mental effort and physical exertion on a limb of the patient, operating on a computerized non-portable system or on single portable device;measuring cognitive responses of the patient by measuring results of the patient playing the virtual reality game, wherein the measured cognitive responses includes measurements of reaction speed, short term memory, long term memory, pairing memory, focusing and executive function;measuring motor responses of the patient with one or more devices while the patient plays the virtual reality game and reporting the motor response to the computer, wherein the measured motor responses include limb strength, limb speed, limb range of movement, motor control, endurance, hand-eye coordination, and dual-tasking;the computer reporting the measured cognitive responses and the measured motor responses through a network to a server that stores the measured cognitive responses and the measured motor responses;reviewing the measured cognitive responses and the measured motor responses and adjusting the virtual reality game difficulty or the virtual reality game according to whether the patient is progressing in the cognitive and physical rehabilitation;a first medical services provider accessing the measured cognitive responses of the patient by accessing the server over the network;a second medical services provider accessing the measured motor responses of the patient by accessing the server over the network; andthe first medical services provider and the second medical services provider providing instructions to the computer based on the measured cognitive responses and the measured motor responses of the patient. 2. The method of claim 1, comprising repeating the steps of claim 1 with the computer directing the patient in accordance with the instructions from the first and second medical services providers. 3. The method of claim 1, comprising the computer analyzing the cognitive responses and the measured motor responses and providing feedback to the patient. 4. The method of claim 1, wherein the one or more devices is selected from the group consisting of: one or more cameras, one or more accelerometers, one or more weights, one or more dynamometers, one or more timers, one or more counters, one or more goniometers, one or more tilt tables and combinations thereof. 5. The method of claim 1, wherein the network is the Internet to mediate cloud rehabilitation. 6. A system for providing simultaneous cognitive and physical rehabilitation to a patient with a computer, comprising: a computer that provides instructions to the patient to play a virtual reality game through an avatar;a first set of one or more devices that measure cognitive responses of the patient that are connected to the computer and that report the ... | Bright Cloud International Corp. | A63B021/06 A63F013/00 G09B019/00 A61B005/04 A61B005/22 A61H001/02 A63B069/00 A61B005/11 A61B005/0482 A63F009/24 A63B023/12 A63B071/06 G06F019/00 A61B005/00 A61F005/37 A63B024/00 A63B022/00 A63B059/00
US10041876 | B2 | 2018-08-07 | Photoacoustic measurement apparatus | 2009-07-09 | A measurement apparatus capable of measuring a position and a size of an absorber with high accuracy, which includes: a light source unit for emitting a pulse beam; an illumination optical unit for leading the pulse beam emitted by the light source unit to an inside of an inspection object; and an acoustic signal detection unit for detecting a photoacoustic signal generated by the pulse beam in which the illumination optical unit includes a first and second illumination optical units that are arranged so that the inspection object is irradiated with the pulse beam from both sides thereof opposingly; and the acoustic signal detection unit is provided so that a detection surface of the acoustic signal detection unit is positioned on the same side as that of one of irradiation surfaces of the inspection object which the first and second illumination optical units irradiate with the pulse beam. | 1. A measurement apparatus, comprising: a light source configured to emit pulsed light of a particular wavelength;an optical system having a pair of optical irradiation ends each of which is optically coupling with the light source and configured to irradiate an inspection object with the pulsed light;an acoustic detection unit comprising a transducer and configured to face the inspection object to detect an acoustic wave generated in the inspection object and propagated from the inspection object and to convert the acoustic wave into an electric signal; anda driving mechanism configured to move the pair of optical irradiation ends and the acoustic detection unit with respect to the inspection object simultaneously,wherein the pair of optical irradiation ends are configured to face each other when the inspection object is disposed therebetween so that the inspection object is irradiated with the pulsed light from both sides of the inspection object,wherein the acoustic detection unit is disposed on the same side as one of the pair of optical irradiation ends,wherein the driving mechanism moves the acoustic detection unit intermittently, andwherein the acoustic detection unit detects a first acoustic wave generated by a first light irradiation in a first wavelength and a second acoustic wave generated by a second light irradiation in a second wavelength during a time period when the driving mechanism does not move the acoustic detection unit. 2. The measurement apparatus according to claim 1, wherein the acoustic detection unit is arranged on a same side as that of a first optical irradiation end, and wherein the first optical irradiation end and the acoustic detection unit are disposed so that a full width at half maximum of an irradiation light intensity distribution of axial symmetry of a region illuminated by the first optical irradiation end overlaps with a full width at half maximum of a region for detecting the acoustic wave by the acoustic detection unit. 3. The measurement apparatus according to claim 2, wherein the first optical irradiation end and a second optical irradiation end are disposed so that a full width at half maximum of an irradiation light intensity distribution of axial symmetry of a region irradiated by the first optical irradiation end overlaps with a full width at half maximum of an irradiation light intensity distribution of axial symmetry of a region irradiated by the second optical irradiation end. 4. The measurement apparatus according to claim 3, wherein the driving mechanism maintains a position relationship among the first optical irradiation end, the second optical irradiation end, and the acoustic detection unit. 5. The measurement apparatus according to claim 1, wherein at least one end of the pair of optical irradiation ends comprises an end of an optical fiber. 6. The measurement apparatus according to claim 5, wherein the light source includes a pulsed laser light source having a wavelength within a ra... | Canon Kabushiki Kaisha | A61B005/05 G01N021/17 A61B005/00 A61B005/145
US09326718 | B2 | 2016-05-03 | Blood glucose measurement for small blood volume | 2010-07-28 | A system is proposed for detection of at least one analyte in a body fluid, in particular for detection of blood glucose. The system is designed to generate a sample of the body fluid and to transfer at least some of the sample to at least one test element, in particular a test panel. The system is designed such that a time period between the generation of the sample and the application to the test element is less than 1 second, preferably less than 500 ms. | 1. A system for detecting at least one analyte in a body fluid, the system comprising: at least one test element comprising a test panel configured to detect the at least one analyte;at least one lancet configured to generate a body fluid sample of less than 1 μl;an actuator configured to incite the at least one lancet;at least one transfer device configured to transfer a whole or partial volume of the body fluid sample generated by the at least one lancet to the at least one test element in a transfer time of less than 1 second; anda measurement device configured to detect the at least one analyte, configured to detect an actual sample volume of the body fluid sample generated by the at least one lancet including a volume of the body fluid sample transferred to the at least one test element, configured to correct a measured value of the at least one analyte taking into account the actual sample volume, and configured to at least partially compensate changes in concentration to the measured value of the at least one analyte taking into account evaporation effects dependent on sample volume, wherein the measurement device comprises at least one sensor for detecting the actual sample volume. 2. The system of claim 1, wherein the transfer device comprises at least one capillary, the at least one lancet comprising at least partially a capillary, the at least one lancet comprising a partially opened capillary, or the at least one lancet comprising a wholly or partially integrated capillary and a wholly or partially integrated at least one test element. 3. The system of claim 2, wherein the transfer device is the at least one lancet comprising the wholly or partially integrated capillary and the wholly or partially integrated at least one test element. 4. The system of claim 2, wherein the transfer device comprises at least one hydrophilic coating. 5. A system for detecting at least one analyte in a body fluid, the system comprising: at least one test element comprising a test panel configured to detect the at least one analyte;at least one lancet configured to generate a body fluid sample of less than 1 μl;an actuator configured to incite the at least one lancet;at least one capillary configured to transfer a whole or partial volume of the body fluid sample generated by the at least one lancet to the at least one test element in a transfer time of less than 1 second, wherein the at least one capillary has a ratio of fillable capillary length to capillary diameter of less than 100; anda measurement device configured to detect the at least one analyte, configured to detect an actual sample volume of the body fluid sample generated by the at least one lancet including a volume of the body fluid sample transferred to the at least one test element, configured to correct a measured value of the at least one analyte taking into account the actual sample volume, and configured to at least partially compensate changes in concentration to the measured value... | Roche Diabetes Care, Inc. | A61B005/157 B65D081/00 A61B005/15 A61B005/145
US20110144466 | A1 | 2011-06-16 | Catheter-Free Implantable Needle Biosensor | 2011-02-22 | Electrochemical biosensors consisting of two or more needle-shaped electrodes for monitoring of molecules such as glucose in the subcutaneous tissue are provided. The two electrodes are parallel to each other, and are fixed perpendicularly to the planar side of a sensor base which provides skin fixation and electric connection. The needle electrodes are made of rigid conducting materials, capable of inserting into the skin without the need of catheters. One electrode is the sensing electrode and the other, the reference electrode. A compound sensing membrane is attached onto the sensing electrode surface to detect analyzed molecules. | 1. An catheter-free subcutaneously implantable apparatus for producing a signal reflecting the concentration of molecules, said apparatus comprising: at least one needle shaped anode, at least one needle shaped cathode, a supporting base for mounting the anodes and the cathodes, and a contact means for connecting the anodes and the cathodes with a signal reading system. 2. The apparatus of claim 1, wherein the molecules are glucose. 3. The apparatus of claim 2, wherein the anode comprises a anode core, a first anode cover layer, a second anode cover layer, a third anode cover layer; wherein the cathode comprises a cathode core, a first cathode cover layer, a second cathode cover layer, a third cathode cover layer. 4. The apparatus of claim 3, wherein the anode core is a metal core; the first anode cover layer is a noble metal layer; the second anode cover layer is a platinum layer, the third anode cover layer is a sensing layer; wherein the cathode core is a metal core, the first cathode cover layer is a silver layer, the second cathode cover layer is a silver chloride layer, and the third cathode cover layer is a polymer layer. 5. The apparatus of claim 4, wherein the sensing layer is a compound membrane, including enzyme and biocompatible polymer material. 6. The apparatus of claim 5, wherein the polymer material is medical grade polydimethylsiloxane copolymer. 7. The apparatus of claim 1, wherein the supporting base is a molded insulating plastic base. 8. The apparatus of claim 7, further comprises a mold plastic cover. 9. The apparatus of claim 8, further comprises conducting circuits. 10. The apparatus of claim 8, further comprises a circular medical adhesive tape. 11. The apparatus of claim 8, wherein the contact means is an electric receptacle. 12. The apparatus of claim 1, wherein the molecules are Oxygen. 13. The apparatus of claim 12, wherein the anode comprises a anode core, a first anode cover layer, a second anode cover layer; wherein the cathode comprises a cathode core, a first cathode cover layer, and a second cathode cover layer. 14. The apparatus of claim 13, wherein the anode core is a metal core; the first anode cover layer is a silver layer; the second anode cover layer is a polymer layer; wherein the cathode core is a metal core, the first cathode cover layer is a platinum layer, the second cathode cover layer is a biosensing layer. 15. The apparatus of claim 14, wherein the biosensing layer is a complex membrane, including enzyme and biocompatible polymer materials. 16. A percutaneously implanted miniature sensor assembly, comprising a base, an anode, a cathode, a first contact plate, and a second contact plate; wherein the base further comprises a first mounting hole surface, a second mounting hole surface, a base top surface and a base bottom surface, wherein the first mounting hole surface defines a first mounting hole, and the second mounting hole surface defines a second mounting hole; wherein the first ... | - | A61B005/05
US20110152712 | A1 | 2011-06-23 | Impedance Measurement Tissue Identification in Blood Vessels | 2010-12-20 | A system for radio frequency (RF) wire guidance during RF occlusion canalization including a RF wire configured to travel through a blood vessel, an impedance measure unit configured to measure an electrical impedance between the RF wire and a reference electrode, a RF power source, an activation controller configured to connect the RF power source to the RF wire when the impedance measure unit indicates that the RF wire is contacting an occlusion material, a RF wire steering system; and a controller configured to activate the steering system to maintain the RF wire within the occlusion material based on impedance measured by the impedance measure unit. | 1. A method of guiding a vascular electrode through a vascular occlusion comprising: contacting the vascular occlusion with the vascular electrode;measuring a first impedance between the vascular electrode and a reference electrode;moving the vascular electrode;measuring a second impedance between the vascular electrode and the reference electrode; andif the difference in impedance between the first measured impedance and the second measured impedance is large enough to indicate that the vascular electrode is no longer contacting the vascular occlusion, then moving the vascular electrode until the vascular electrode contacts the vascular occlusion. 2. The method of claim 1 further comprising energizing the vascular electrode with radio frequency electricity sufficient to vaporize part of the vascular occlusion. 3. The method of claim 2 further comprising energizing the vascular electrode with radio frequency electricity if the second measured impedance indicates that the vascular electrode is in contact with the vascular occlusion. 4. The method of claim 1 wherein the second measured impedance is a measurement of resistivity. 5. The method of claim 1 wherein the second measured impedance is a measurement of relative permittivity. 6. The method of claim 1 wherein the second measured impedance is a measurement of complex impedance. 7. The method of claim 1 wherein the second measured impedance is a measurement relative to a baseline impedance measured in blood. 8. The method of claim 1 wherein the second measured impedance is measured by energizing the vascular electrode with an excitation voltage of between 20 kHz and 500 kHz. 9. A system for radio frequency (RF) wire guidance during RF occlusion canalization comprising: a RF wire configured to travel through a blood vessel;an impedance measure unit configured to measure an electrical impedance between the RF wire and a reference electrode;a RF power source;an activation controller configured to connect the RF power source to the RF wire when the impedance measure unit indicates that the RF wire is contacting an occlusion material;a RF wire steering system; anda controller configured to activate the steering system to maintain the RF wire within the occlusion material based on impedance measured by the impedance measure unit. 10. The system of claim 9 wherein the activation controller is further configured to alternate connecting the RF power source and the impedance measure unit to the RF wire. 11. The system of claim 9 wherein the RF power source and the impedance measure unit are simultaneously connected to the RF wire. 12. A method of characterizing tissue within a blood vessel comprising: contacting a first blood vessel site with a vascular electrode;measuring the impedance between the vascular electrode and a reference electrode;moving the vascular electrode to a second blood vessel site;repeating the measuring and the moving until the impedance of a plurality of vessel sites have... | - | A61B005/053 A61B018/14
US20110144513 | A1 | 2011-06-16 | METHODS FOR IMPROVED ANALYSIS OF HEART RATE VARIABILITY | 2010-12-15 | Methods for analysis of heart rate variability can be used in gaming and other software applications as disclosed herein; wherein improvements include a faster analysis of heart rate variability, prediction of future values, and assessment of value deviations, among others. The described methods interpolate variations in heart rate variability using one or more of the factors: LF amplitude factor (kALF), HF amplitude factor (kAHF), and deviation factor (kdev). Other features of the improved methods are described in detail herein. | 1. A method for analysis of heart rate variability, comprising: (i) providing an electrocardiographic device;(ii) measuring a users heart rate over a period of time;(iii) interpolating the value of low frequency (LF) variations and high frequency (HF) variations in the heart rate using one or more normalization factors selected from:LF amplitude factor (kALF), HF amplitude factor (kAHF), and deviation factor (kdev);(iv) representing the proportion of adjusted LF variations to adjusted HF variationsin the form of HRV=LFadj/HFadj. 2. The method of claim 1, wherein said period of time is sec−1. | - | A61B005/0402
US20110105952 | A1 | 2011-05-05 | RELATIVELY SMALL DEVICES APPLIED TO THE SKIN, MODULAR SYSTEMS, AND METHODS OF USE THEREOF | 2010-10-29 | The present invention generally relates, in certain aspects, to relatively small devices applied to the skin, modular systems, and methods of use thereof. In some aspects, the device is constructed and arranged to have more than one module. For instance, the device may have a module for delivering to and/or withdrawing fluid from the skin and/or beneath the skin of a subject and a module for transmitting a signal indicative of the fluid delivered to and/or withdrawn from the skin and/or beneath the skin of the subject, a module for analyzing a fluid withdrawn from the skin and/or beneath the skin of the subject, or the like. In some embodiments, the modules are connectable and/or detachable from each other, and in some cases, the connections and/or detachments may be performed while the device is in contact with the subject, e.g., while affixed to the subject. In some embodiments, the device may be repeatedly applicable to the skin of the subject to deliver to and/or withdraw fluid from the skin and/or beneath the skin of a subject, e.g., at the same location, or at different locations on the skin of the subject. In some aspects, the devices may be self-contained and/or have a relatively small size, and in some cases, the device may be sized such that it is wearable and/or able to be carried by a subject. For example, the device may have a mass and/or dimensions that allow the device to be carried or worn by a subject for various periods of time, e.g., at least about an hour, at least about a day, at least about a week, etc., or no more than about an hour, no more than about 10 min, etc. | 1. A device for analysis of an extractable fluid from the skin and/or from beneath the skin of a subject, the device comprising: a support structure for application to the skin of the subject;a fluid transporter associated with the support structure; anda signal structure for generating a signal indicative of the determination of at least one analyte contained within the fluid,wherein the support structure and the signal structure are constructed and arranged to be readily connectable and/or detachable from each other. 2. The device of claim 1, wherein the support structure further comprises an adhesive for adhesion of the support structure to the skin. 3-5. (canceled) 6. The device of claim 1, wherein the fluid transporter is able to be activated using an actuator. 7. The device of claim 6, wherein the actuator comprises a button. 8. The device of claim 1, wherein the fluid transporter is able to be activated remotely. 9. The device of claim 1, wherein the device comprises a timing mechanism for timed activation of the fluid transporter after removal from packaging containing the device. 10-11. (canceled) 12. The device of claim 1, wherein the support structure further comprises a microfluidic channel in fluidic communication with the fluid transporter. 13. The device of claim 1, wherein the fluid comprises blood. 14. (canceled) 15. The device of claim 1, wherein the device further comprises a storage chamber in fluidic communication with the fluid transporter. 16. The device of claim 1, wherein the device is able to determine an analyte contained within the fluid withdrawn from the skin and/or from beneath the skin of the subject. 17. The device of claim 1, wherein the device is able to determine the analyte without subsequent removal of the device from the skin of the subject. 18. The device of claim 1, wherein the device is able to automatically determine the analyte. 19. The device of claim 1, wherein the device is able to display a result of the determination of the analyte. 20. The device of claim 1, wherein the fluid transporter comprises a needle. 21. The device of claim 20, wherein the needle is a microneedle. 22-23. (canceled) 24. The device of claim 1, wherein the fluid transporter is immobilized relative to the support structure. 25. The device of claim 1, wherein the device is able to withdraw the fluid transporter from the skin of the subject after exposure of the fluid transporter to the skin of the subject. 26. (canceled) 27. The device of claim 1, wherein the support structure is able to apply a pressure less than atmospheric pressure to the skin of the subject. 28. The device of claim 1, wherein the support structure comprises a self-contained vacuum chamber having a pressure less than atmospheric pressure. 29. The device of claim 1, wherein the fluid transporter is able to move relative to the support structure. 30. The device of claim 1, wherein the fluid transporter is able to be activated by the sub... | Seventh Sense Biosystems, Inc. | A61B005/00
US07946983 | B2 | 2011-05-24 | Body-insertable apparatus and power supplying method therefor | 2006-12-14 | An object of the present invention is to easily start to operate a body-insertable apparatus which is introduced into a subject and executes a predetermined function. In a capsule endoscope 3 of the invention, a magnetic switch 14 is arranged in a tubular capsule casing 16 of the capsule endoscope 3 in a direction substantially perpendicular to a longitudinal axis t direction of the capsule casing 16. A movable electrode of the magnetic switch 14 is brought into contact with a fixed electrode of the magnetic switch by a magnetic induction action by a magnetic field generated by a magnet 6, which is applied to the magnetic switch 14 from the outside of the capsule endoscope 3 in a direction substantially parallel to the longitudinal axis t direction of the capsule casing 16. As a result, electric power can be supplied from a power source unit to a function executing unit. | 1. A body-insertable apparatus comprising: a function executing unit which executes a predetermined function;a power source unit which supplies electric power to the function executing unit;a capsule body which contains the function executing unit and the power source unit; anda magnetic switch which is provided in the capsule body, and controls the power supply from the power source unit to the function executing unit by a pair of contacts, the pair of contacts being brought into contact with or separated from each other by a magnetic field, the magnetic field being applied from the outside of the capsule body in a direction substantially parallel to a longitudinal axis direction of the capsule body, whereinthe function executing unit comprises an illuminating unit, an imaging unit and a radio transmitting unit; andthe magnetic switch is mounted on a board on which one of the illuminating unit and the radio transmitting unit is mounted, the board being perpendicular to a central axis of the capsule body. 2. The body-insertable apparatus according to claim 1, wherein the magnetic switch is provided on a board on which the illuminating unit is mounted. 3. The body-insertable apparatus according to claim 1, wherein the magnetic switch is provided on a board on which the radio transmitting unit is mounted. 4. A body-insertable apparatus comprising: a function executing unit which executes a predetermined function and includes at least an imaging unit for imaging a body cavity;a power source unit which supplies electric power to the function executing unit;a switch unit which disconnectively connects the function executing unit and the power source unit; anda capsule body which is tubular and rotationally symmetric with respect to a longitudinal axis direction, and contains the function executing unit, the power source unit, the switch unit, and an imaging board on which the imaging unit is mounted;wherein the switch unit comprises a base seat, a fixed electrode, a movable electrode and a pair of electrodes, the fixed electrode and the movable electrode each being made of a conductive and magnetic material and each being provided on a surface of the base seat, the pair of electrodes being provided on a reverse side of the base seat and electrically connected to the fixed electrode and the movable electrode, respectively, the fixed electrode and the movable electrode being brought into contact with or separated from each other by a magnetic field applied in a contact and separation direction of the fixed electrode and the movable electrode; andthe switch unit controls the power supply from the power source unit to the function executing unit in accordance with an action of a magnetic field, the magnetic field being applied from the outside of the capsule body in a direction substantially parallel to the longitudinal axis direction. 5. The body-insertable apparatus according to claim 4, wherein each of said movable and fixed electrodes are magnetiz... | Olympus Medical Systems Corp. Olympus Corporation | A61B001/00 A61B005/07
US07967771 | B2 | 2011-06-28 | System for use in draining fluid from a brain or spinal fluid cavity unto another body cavity of a human being | 2008-10-24 | This invention describes a method for processing pressure signals derivable from locations inside or outside a human or animal body or body cavity. Different aspects of the invention relate to a method for optimal differentiating between cardiac beat- and artifact-induced pressure waves and a method for obtaining new and improved information from said pressure signals. In particular, this invention describes a system for draining fluid from a brain or spinal fluid cavity according to the inventive method of processing pressure signals. | 1. A system for use in draining fluid from a brain or spinal fluid cavity unto another body cavity of a human being, comprising: a first drainage tube having an inlet thereof configured to be located in said brain or spinal fluid cavity, said first drainage tube connected to the inlet of a fluid flow controllable valve,one or more sensor(s) configured to measure fluid pressure in one of: the brain cavity, the spinal cavity and said another body cavity, when located in such cavity or attached to said first drainage tube inside or outside such cavity,a pressure transducer linked to said sensor(s) and capable of transforming continuous pressure-related signals derivable from said sensor(s) into processable signals,a processing unit configured to receive said processable signals,a valve-opening regulator having an input linked to output means of the processing unit via a regulator unit and a control unit and having its output operatively connected to said valve to control fluid flow mass rate through said valve,a second drainage tube from an outlet of said valve, said second drainage tube at a downstream end having a distal outlet configured to be located in said another human body cavity, anda power supply to power said pressure transducer, said processing unit, said valve opening regulator, said control unit, and said regulator unit,wherein said processing unit comprises: sampling means configured to sample said pressure-related signals at specific intervals,converter means configured to convert the sampled signals into pressure related digital data with a time reference,identifying means capable of identifying from said digital data single pressure waves related to cardiac beat-induced pressure waves, related to artifacts, or a combination of cardiac beat-induced waves and artifacts,computing means configured to compute time sequence (TS.x)-related parameters of said single pressure waves during individual time sequence windows,analyzing means configured to establish an analysis output of said time sequence (TS.x)-related parameters for a selectable number of said time sequence windows, andoutput means for delivering said analysis output as parameter signals to said regulator unit. 2. A system according to claim 1, wherein said system further comprises an information transfer means linked to the processor and configured to communicate with external processing equipment, and wherein an assembly of said fluid flow controllable valve, said valve-opening regulator, said control unit, said regulator unit, said processing unit, said pressure transducer, said power supply, said information transfer means, and said first and second drainage tubes are located below a skin surface of said human body. 3. A system according to claim 2, wherein said assembly configured to be located below a skin surface of said human body in a thoracic or abdominal body area thereof and with part of said first and second drainage tubes thereat. 4. A system according to clai... | DPCOM AS | A61M005/00 A61B005/00
US20110152663 | A1 | 2011-06-23 | MEDICAL IMAGE DIAGNOSTIC APPARATUS, MEDICAL IMAGE DISPLAY DEVICE, PERSONAL INFORMATION MANAGEMENT SYSTEM | 2010-12-21 | A medical image diagnostic apparatus includes an input unit to input personal information of a subject, a communication unit to communicate with a personal information managing server for storing the personal information, and a judge unit to judge whether the personal information can be transmitted to a personal information managing server. The communication unit transmits the personal information to the personal information managing server and receives personally-identifying information for identifying the personal information in response to the personal information transmitted, when the judge unit judges that the personal information can be transmitted. A memory for examination information stores the personally-identifying information. | 1. A medical image diagnostic apparatus producing image data of an examination subject, comprising: an input unit to input personal information of the subject;a judge unit to judge whether the personal information can be transmitted to a personal information managing server that stores the personal information;a communication unit to transmit the personal information to the personal information managing server and to receive personally-identifying information that identifies the personal information, when the judge unit judges that the personal information can be transmitted to the personal information managing server; andan examination information memory to store the personally-identifying information. 2. The medical image diagnostic apparatus according to claim 1, wherein the judge unit is configured to judge whether the personal information can be transmitted based on position information obtained by a position information obtaining unit. 3. The medical image diagnostic apparatus according to claim 2, wherein the judge unit is configured to judge that the personal information can be transmitted when the medical image diagnostic apparatus is in a healthcare center, and is configured to judge that the personal information can not be transmitted when the medical image diagnostic apparatus is outside the healthcare center. 4. The medical image diagnostic apparatus according to claim 1, further comprising an encrypt unit to encrypt the personal information, wherein the communication unit is configured to transmit the encrypted personal information encrypted by the encrypt unit to the personal information managing server. 5. The medical image diagnostic apparatus according to claim 4, further comprising: a personal information memory to store the personal information encrypted by the encrypt unit; anda delete unit configured to delete the personal information,wherein the personally-identifying information is stored by the examination information memory. 6. The medical image diagnostic apparatus according to claim 4, further comprising: a personally-identifying managing unit to read the examination information related to the personally-identifying information to identify the personal information returned by the personal information managing server, in response to the personal information transmitted to the personal information managing server; anda display unit to display the examination information. 7. The medical image diagnostic apparatus according to claim 1, wherein the input unit is configured to input the personally-identifying information stored by the examination information memory, the medical image diagnostic apparatus further comprising:a decrypt unit to decrypt the encrypted personal information identified by the personally-identifying information returned from the personal information managing server in response to the personally-identifying information transmitted;a personally-identifying information managing unit to read the exa... | KABUSHIKI KAISHA TOSHIBA TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/05
US07962202 | B2 | 2011-06-14 | Method and apparatus for verifying a determined cardiac event in a medical device based on detected variation in hemodynamic status | 2006-04-28 | A method and apparatus for verifying a determined cardiac event in a medical device based on detected variation in hemodynamic status that includes a physiologic sensor sensing physiologic signals to generate a plurality of variation index samples corresponding to the sensed signals, and a microprocessor computing a variation index trend associated with a predetermined number of variation index samples of the plurality of variation index samples, determining whether deviations of the predetermined number of variation index samples over predetermined sampling windows are less than a deviation threshold, and determining, in response to the deviations being less than the deviation threshold, a corrected variation index trend in response to the changes in the variation index trend during the predetermined sampling windows. | 1. A medical device, comprising: a physiologic sensor sensing physiologic signals to generate a plurality of variation index samples corresponding to the sensed signals; anda microprocessor computing a variation index trend associated with a predetermined number of variation index samples of the plurality of variation index samples, determining whether deviations of the predetermined number of variation index samples over predetermined sampling windows are less than a deviation threshold, adjusting, in response to the deviations being less than the deviation threshold, a corrected variation index trend in response to the changes in the variation index trend during the predetermined sampling windows, and not adjusting the corrected variation index in response to the deviations not being less than the deviation threshold. 2. The device of claim 1, wherein the variation index trend corresponds to a measure of change in tissue oxygenation. 3. The device of claim 1, wherein the determined corrected index trend includes an offset determined in response to the deviations not being less than the deviation threshold. 4. The device of claim 1, wherein the microprocessor determines, in response to the deviations being less than the deviation threshold, whether the sensed signals are associated with an unstable rhythm. 5. The device of claim 4, wherein the microprocessor determines whether the computed variation index trend is less than a predetermined index trend threshold. 6. The device of claim 1, wherein the variation index trend corresponds to relationships between proportions of red light intensities and infrared light intensities to corresponding baseline intensities. 7. The device of claim 1, wherein the variation index trend corresponds toii0-i*i0*,wherein i is an intensity of red light, i* is an intensity of infrared light, io is an intensity of red light baseline, and io* is an intensity of infrared light baseline. 8. The device of claim 1, further comprising a plurality of sensors sensing cardiac signals, wherein the microprocessor detects a cardiac event in response to the sensed cardiac signals and confirms the detected cardiac event in response to the corrected variation index trend. 9. A method of determining a cardiac event in a medical device, comprising: sensing physiologic signals to generate a plurality of variation index samples corresponding to the sensed signals;computing a variation index trend associated with a predetermined number of variation index samples of the plurality of variation index samples;determining whether deviations of the predetermined number of variation index samples over predetermined sampling windows are less than a deviation threshold;adjusting, in response to the deviations being less than the deviation threshold, a corrected variation index trend in response to the changes in the variation index trend during the predetermined sampling windows; andnot adjusting the corrected variation index in response... | Medtronic, Inc. | A61B005/0432
US20110141016 | A1 | 2011-06-16 | MOUSE FOR MEASURING CONSISTENCY OF BLOOD OXYGEN | 2009-12-15 | A mouse for measuring consistency of blood oxygen includes a mouse case, on which two transmitters, a transmitting circuit, a receiver, a receiving feedback circuit, a filtrate amplifying circuit, and a calculating unit are disposed. The transmitting circuit serves to control the transmitters to alternately radiate a red glow and an infrared light to the receiver. The receiving feedback circuit applies to convert them into an electrical signal, and the filtrate amplifying circuit obviates any promiscuous frequency existing in the electrical signal and amplifies the electrical signal. The calculating unit accordingly computes the signal to obtain the value of consistency of blood oxygen, which would thence be outputted from a communication interface for informing users about the consistency of blood oxygen, so as to prevent the discomfort suffered by the occurrence of oxygen deficiency in blood. | 1. A mouse for measuring consistency of blood oxygen, comprising: a mouse case having a controlling circuit therein, said mouse case including a communication interface and at least one controlling button, said communication interface and said controlling button being electrically connected to said controlling circuit; characterized in that said mouse further including:two transmitters disposed on said mouse case;a transmitting circuit disposed inside said mouse case and electrically connected to said transmitters;a receiver disposed on said mouse case;a receiving feedback circuit disposed inside said mouse case and electrically connected to said receiver and said transmitting circuit;a filtrate amplifying circuit disposed inside said mouse case and electrically connected to said receiving feedback circuit; anda calculating unit disposed inside said mouse case and electrically connected to said filtrate amplifying circuit and said communication interface, respectively. 2. The mouse as claimed in claim 1, wherein said communication interface is either a wire cable or a wireless transceiver. 3. The mouse as claimed in claim 1, wherein said mouse case is provided with a screen which is electrically connected to said calculating unit. 4. The mouse as claimed in claim 1, wherein said transmitters and said receiver are disposed either at a top portion or a side portion of said mouse case. | - | G09G005/08 A61B005/05
US20110144463 | A1 | 2011-06-16 | DEVICE, SYSTEM AND METHOD FOR MODULAR ANALYTE MONITORING | 2009-02-26 | The present invention relates to a device, system and a method for measuring an analyte and in particular, to such a device, system and method in which analyte from a bodily fluid is measured and/or monitored. | 1-85. (canceled) 86. A system for measuring a blood analyte concentration from a blood sample, the system comprising: a. a disposable housing having an opening for receiving said blood sample for measuring said analyte and at least two disposable electrical contacts for performing said measurement; and wherein said disposable housing comprises a single skin piercing element, a single measuring element for measuring said analyte concentration in blood and a coupling mechanism for attaching said housing over said measuring area; andb. a reusable housing having at least two reusable electrical contacts corresponding to said at least two disposable electrical contacts; and comprising a power source, a controller, a transponder and at least one mechanical actuator for performing and communicating said measurement; andc. wherein said reusable housing is associated with said disposable housing over said at least two corresponding reusable electrical contacts and said at least two disposable electrical contacts; and wherein said at least one mechanical actuator is coupled to and manipulates said single skin piercing element and said single measuring element; andd. wherein upon a triggering signal said at least one mechanical actuator manipulates the movement of said single skin piercing element to pierce the skin and then after a given period at least one mechanical actuator manipulates the movement of said single measuring element, from a first position of a piercing area to a second position of said piercing area, such that said measuring element opening wipes said piercing area and samples the blood drawn from said skin piercing area; ande. therein triggering the measurement of said analyte concentration in said blood sample within said reusable housing and communicating said measurement to said remote controller;f. and wherein said remote controller is further adapted to receive said measurement and to display it the in the form of a concentration of said blood analyte. 87. The system of claim 86 wherein said given period is a period that allows the blood to pool about said piercing area forming a blood pooling area. 88. The system of claim 87 wherein said given period is a predetermined period that allow the blood to pool in said blood pooling area. 89. The system of claim 87 wherein said given period is a determined by measuring the amount of accumulated blood within said blood pooling area using an optical element. 90. The system of claim 87 wherein said piercing element is adapted for piercing only the dermis layer of the skin. 91. The system of claim 87 wherein said housing further comprises at least one tissue treatment element for improving blood extraction, wherein said treatment element is chosen from the group consisting of electrical energy, ultrasound energy, optical energy, acoustic energy, vasodilatation drugs, applied heating, massage, mechanical energy and any combination thereof. 92. The system of claim 91 wherein a remote con... | - | A61B005/1468
US07998574 | B2 | 2011-08-16 | Biomedical sensing methods employing hydro-insensitive alternating current responsive composites | 2010-12-14 | A method is disclosed of detecting a bioelectrical signal from a subject. The method includes the steps of applying a composite material to a subject wherein the composite material includes a polymeric material and a polar material that is substantially dispersed within the polymeric material; coupling monitoring equipment to the second side of the composite material; permitting the polar material within the polymeric material to respond to the bioelectrical signal within the subject; and detecting a responsive electrical signal from the composite material that is representative of the bioelectrical signal. The polar material exhibits molecular compatibility with the polymeric material such that the polar material neither blooms to a surface of the polymeric material nor crystallizes within the polymeric material, and the composite material has a first side for contacting the subject and a second side. | 1. A method of detecting a bioelectrical signal from a subject, said method comprising the steps of: applying a composite material to a subject wherein the composite material includes a polymeric material and a polar material that is substantially dispersed within the polymeric material, wherein said polar material exhibits molecular compatibility with the polymeric material such that the polar material neither blooms to a surface of the polymeric material nor crystallizes within the polymeric material, said composite material having a first side for contacting the subject and a second side;coupling monitoring equipment to the second side of the composite material;permitting the polar material within the polymeric material to respond to the bioelectrical signal within the subject; anddetecting a responsive electrical signal from the composite material that is representative of the bioelectrical signal. 2. The method as claimed in claim 1, wherein said step of permitting the polar material within the polymeric material to respond to the bioelectric signal includes permitting the polar material to become aligned with the bioelectrical signal as the bioelectric signal increases within the subject. 3. The method as claimed in claim 2, wherein said method further includes the step of permitting the polar material within the polymeric material to respond to the bioelectrical signal as the bioelectric signal decreases within the subject by becoming non-aligned with the bioelectrical signal, thereby providing a discharge electric field as the polar material becomes non-aligned. 4. The method as claimed in claim 3, wherein said step of detecting a responsive electrical signal from the composite material that is representative of the bioelectrical signal involves detecting the discharge electric field using the monitoring equipment. 5. The method as claimed in claim 1, wherein said polymeric material is an adhesive. 6. The method as claimed in claim 1, wherein said polymeric material is an acrylic adhesive. 7. The method as claimed in claim 1, wherein said polymeric material is a pressure sensitive adhesive. 8. The method as claimed in claim 1, wherein said polar material is a salt. 9. The method as claimed in claim 1, wherein said polar material is an organic salt. 10. The method as claimed in claim 1, wherein said composite includes a concentration of said polar material within said polymeric material of between about 5% and about 65% based on weight. 11. The method as claimed in claim 1, wherein said composite includes a concentration of said polar material within said polymeric material of between about 10% and about 35% based on weight. 12. The method as claimed in claim 1, wherein said composite has a DC resistance of at least about 100 k Ohms. 13. The method as claimed in claim 1, wherein said composite has a dielectric constant that changes in the presence of and changes with alternating electric field. 14. A method of detecting an alt... | FLEXcon Company Inc. | H01B001/12 A61B001/12 A61B005/0402
US07949384 | B2 | 2011-05-24 | Method to correct magnetic field/phase variations in proton resonance frequency shift thermometry in magnetic resonance imaging | 2008-02-22 | PRF shift MRI data is acquired. The PRF shift MRI data may include signals affected by both a desired PRF shift and an undesired PRF shift. Thus, example systems and methods describe manipulating the PRF shift MRI data to make it substantially free of the effects of the undesired PRF shift, which facilitates displaying certain MRI images based on the desired PRF shift. | 1. A system, comprising: a data store for storing an input signal data received from an MRI or NMR system, where the input signal data includes two or more related sets of image data, where a set of image data is a set of k-space data that contains a peak k-space signal location;a first logic for processing the input signal data into a processed signal data, where the effects of one or more undesired phase shifting events on the input signal data are reduced in the processed signal data, where the undesired phase shifting events are one or more of, a variation in a main magnetic field in the MRI or NMR system and a misaligning of echoes in the input signal data,where the variation in the main magnetic field is one or more of, a temporal variation, a spatial variation, and a field strength variation,where the first logic identifies the peak k-space signal location for two or more of the related sets of image data and aligns the two or more related sets of image data based on peak k-space signal locations, andwhere the first logic processes the input signal by computing a phase shift due to a variation in B0 by processing k-space data centered around and local to a peak k-space signal location; anda second logic for processing the processed signal data to quantify the effects of one or more desired phase shifting events. 2. The system of claim 1, where a desired phase shifting event is a temperature change. 3. The system of claim 1, where the desired phase shifting event is one or more of a motion change, a velocity change, and an acceleration change. 4. The system of claim 1, where the first logic employs an indirect interpolation algorithm to identify the peak k-space signal location. 5. The system of claim 1, where the second logic establishes a first set of processed signal data as a reference signal. 6. The system of claim 5, where the second logic compares one or more second sets of processed signal data to the reference signal to create one or more sets of difference data. 7. The system of claim 6, comprising a second data store for storing the reference signal and one or more sets of difference data. 8. The system of claim 6, comprising a display for displaying an image developed from the reference signal and one or more sets of difference data. 9. The system of claim 6, where the second logic creates the one or more sets of difference data by determining one or more phase shifts between one or more elements of the processed signal data according to Δφ=Δφref−φ2. 10. The system of claim 9, where the second logic computes φref and φ2 according to φ=tan−1(I(x,y)/R(x,y)). 11. The system of claim 6, comprising the second logic quantifying the effects of temperature change as a desired phase shifting event according to: Δφ(x,y,t)=γ*B0(x,y,t)*δ*TEeff*ΔT(x,y,t). 12. The system of claim 11, where Δ is approximately −0.01 ppm/C.°. 13. The system of claim 11, where γ is approximately 2π*43 MHz/T. 14. The system of claim 1, where the secon... | Case Western Reserve University | A61B005/05
US07947508 | B2 | 2011-05-24 | Systems and methods for intra-oral diagnosis | 2007-10-31 | Devices, systems, and methods are disclosed for intra-orally diagnosing a body condition. | 1. A dental appliance, comprising: an appliance body;a sensor attached to the appliance body configured for intra-orally diagnosing a body condition while the dental appliance is within an oral cavity, where the sensor detects the body condition based on an interaction between a first chemical compound and a second chemical compound; anda processor located on the appliance body configured for providing a diagnosis that is based on the body condition. 2. The appliance of claim 1, where at least a portion of the sensor is made from the first chemical compound. 3. The appliance of claim 1, where the sensor is designed to detect a body condition that is a chemical reaction. 4. The appliance of claim 1, where the sensor is designed to detect a body condition that is a chemical reaction, where at least a portion of the sensor is formed from a first chemical compound, and where the chemical reaction is caused by a reaction to a second chemical compound. 5. The appliance of claim 1, where the sensor is designed to detect a body condition that is a chemical reaction, where at least a portion of the sensor is formed from a first chemical compound, and where the chemical reaction is caused by a reaction to a second chemical compound that is produced when a particular malady is present in the body. 6. The appliance of claim 1, where the sensor is designed to detect an interaction between two chemicals. 7. The appliance of claim 1, where a sensitivity of the sensor is focused on a characteristic that may change in the body. 8. The appliance of claim 1, where a sensitivity of the sensor is focused on identifying whether a quantity of a byproduct that may be produced in the body is above a threshold level. 9. The appliance of claim 1, where the sensor can change color. 10. A dental system, comprising: an appliance body;a processor provided on the appliance body;a sensor attached to the appliance body configured for intra-orally diagnosing a body condition while the dental appliance is within an oral cavity, providing a diagnosis using the processor; andwhere the sensor detects the body condition based on a chemical reaction between a first chemical compound and a second chemical compound and where the chemical reaction can be sensed by the sensor and a signal can be generated that can be provided to the processor. 11. The system of claim 10, where the processor is on the appliance. 12. The system of claim 10, where the signal can be processed by the processor to develop a diagnosis of a medical condition on the appliance. 13. The system of claim 12, where the appliance includes a delivery mechanism for administering medication based upon the diagnosed condition. 14. The system of claim 12, where the appliance includes a transmitting mechanism for transmitting the diagnosis to a remote computing device for review by one or more treatment professionals. 15. The system of claim 12, where the appliance includes a transmitting mechanism for transmittin... | Align Technology, Inc. | A61B005/00
US07935307 | B2 | 2011-05-03 | Disposable, refillable glucometer with cell phone interface for transmission of results | 2008-05-30 | A disposable glucometer is disclosed which includes a testing region for receiving a test strip and electronic components to perform blood glucose monitoring, a cell phone interface for relaying test results to a monitoring station or server (which can be through a Bluetooth radio), and a multi-strip test strip holder. The glucometer includes a mechanism to automatically load test strips, one at a time, from the multi-strip test strip holder into the testing region, and a mechanism to eject the strip from the testing region after the test is complete. The test strip holder is preferably encased in a transparent material or otherwise includes a visual indicator so that the number of strips remaining strips and available for testing can be determined. In one embodiment, the entire unit is disposable, and is discarded when all of the strips have been used. In another embodiment, the test strip holder portion of the unit is removed and replaced with a new test strip holder which is loaded with test strips. | 1. A glucometer holding multiple glucose test strips which can be loaded for testing in succession, and each used test strip can be ejected following test completion, comprising: a body outlining a chamber holding multiple glucose test strips, said chamber including a platform with glucose test strips placed thereon in a stack on one side of the platform, and springs positioned on the other side of the platform so as to advance the glucose test strips in the direction of the stack;a cam with an edge portion designed to fit behind the test strip in the stack which is furthest from the platform to allow its selection from the other test strips in the stack, wherein the cam and an edge portion reside in a slot in the body such that the cam can move in a direction transverse to the direction of the stack and then stop at a particular location so that the selected test strip extends a predetermined amount through a port in the side of the glucometer, such that a sample can be applied to the exposed portion of said selected test strip for testing, when said selected test strip is in such testing position;electronic testing components in electrical connection with the selected test strip when it is said testing position, wherein said electronic testing components include a microcontroller which identifies when the selected test strip is at the testing position and initiates testing for glucose in a sample in contact with said selected test strip:a wire or wireless link for connecting the glucometer to a wireless device which can transmit test results to a monitoring station; andsaid cam and said bore are designed such that the cam can further advance the selected test strip following testing, such that, said selected test strip is ejected through the port and from the glucometer, after which, the cam returns to a position where it can advance the next test strip in the stack. 2. The glucometer of claim 1 wherein the chamber is detachable and can be replaced with a different test strip holder filled with unused test strips. 3. The glucometer of claim 1 wherein the number of test strips in the holder can be viewed. 4. The glucometer of claim 1 wherein the cam and said bore are designed such that the cam's forward travel terminates at the end of the bore whereby the strip is ejected. 5. The glucometer of claim 1 wherein the electronic testing components maintain a consistent reference voltage across the strip and determine the variation in current flowing through the strip, over time, to determine the glucose in the sample. 6. The glucometer of claim 1 wherein the wireless device is a Bluetooth radio, cellular phone or PDA. 7. The glucometer of claim 6 wherein glucometer is direct is connected to a cellular phone or PDA. 8. The glucometer of claim 6 wherein test results are transmitted from the cell phone or PDA to a monitoring station or a server which stores the results. | EOS Health, DVC. | G01N027/00 A61B005/00 B65D081/00 C12M001/34
US07956756 | B2 | 2011-06-07 | REM-sleep directed visual alarm system and method | 2009-02-19 | The present apparatus and method for using the same detects Rapid Eye Movement (REM) in a sleeping mammal and awakens the mammal after the cessation of a specific episode of REM. The system comprises an alarm setting unit for setting a predetermined wakeup time, a data collection unit for collecting physiological data from the mammal over time, a processing unit for determining the occurrence and cessation of REM and for providing a stimulation signal for awakening the mammal within a prescribed window of time before the predetermined wakeup time and after the cessation of REM and after the mammal's nadir in body temperature with respect to time. | 1. An apparatus for detecting Rapid Eye Movement (REM) in a sleeping mammal and awakening the mammal after the cessation of REM, the apparatus comprising: a. an alarm time setting means for setting a predetermined time;b. data collection means for collecting physiological data from the mammal over time;c. processing means for processing said physiological data in near real time to detect the occurrence of REM in said mammal; andd. a means for providing a stimulation signal for awakening said mammal after cessation of REM wherein said cessation of REM occurs within a maximum number of minutes before said predetermined time. 2. The apparatus of claim 1 wherein said maximum number of minutes is between 60 and 120. 3. The apparatus of claim 2 wherein said physiological data comprises an amount of eye movement over time. 4. The apparatus of claim 3 wherein said stimulation signal comprises auditory stimulation. 5. The apparatus of claim 3 wherein said processing means compares an amount of eye movement during a second time interval relative to an amount of eye movement during a first time interval. 6. The apparatus of claim 3 wherein said physiological data further comprises the body temperature of said mammal, wherein said processing means further comprises detecting the nadir of said temperature with respect to time, and wherein said stimulation signal is provided not before said nadir. 7. The apparatus of claim 6 wherein said stimulation signal comprises photic stimulation. 8. The apparatus of claim 7 wherein said photic stimulation is a light source having a wavelength of between 420-500 nanometers and wherein said light source increases in illumination from a value greater than 0 LUX to a value of not more than 10,000 LUX. 9. The apparatus of claim 7 further comprising a means for adjusting the ambient lighting of a room. 10. The apparatus of claim 9 wherein said photic stimulation is delivered by said ambient lighting. 11. A method of detecting Rapid Eye Movement (REM) sleep in a mammal and awakening the mammal after the cessation of REM sleep, the method comprising the steps of: a. setting a predetermined time;b. collecting physiological data from the mammal over time;c. processing said physiological data in near real time to detect the occurrence of REM in said mammal; andd. providing a stimulation signal for awakening said mammal after cessation of REM wherein said cessation of REM occurs within a maximum number of minutes before said predetermined time. 12. The method according to claim 11 wherein said physiological data comprises an amount of eye movement over time, and wherein said processing step determines a relative amount of eye movement based on said physiological data during a second time interval compared to said physiological data during a first time interval, and wherein said relative amount of eye movement falls below a minimum predetermined threshold before said step of providing a stimulation signal. 13. The method ... | - | G08B023/00 A61B005/08
US07941210 | B2 | 2011-05-10 | Electroimpedance tomograph with common-mode signal suppression | 2006-06-29 | An electroimpedance tomograph with a plurality of electrodes (1) is provided, which can be placed on the body of a patient and are connected via a selector switch (60) with a control and evaluating unit (20). The control and evaluating unit (20) cooperates with the selector switch (60) such that two electrodes each are supplied with an alternating current from an AC power source (22) and the detected analog voltage signals of the other electrodes are processed in order to reconstruct therefrom the impedance distribution of the body in the plane of the electrodes, wherein a symmetrical AC power source is used to reduce common-mode signals. To further suppress interferences due to common-mode signals, provisions are made for the control and evaluating unit (20) to be set up, furthermore, for detuning the common-mode signal of the alternating current on the body against the ground by means of a common-mode signal measuring electrode (4) and, based on this, the symmetry of the symmetrical AC power source such that the common-mode signal on the body is minimized, and the corresponding detuning parameters are stored for each electrode pair. | 1. An electroimpedance tomography comprising: a plurality of electrodes for placement on a body of a patient;a control and evaluating unit;a selector switch;a measuring amplifier;an AD converter;a symmetrical AC power source used to reduce common-mode signals, said electrodes being connected to said control and evaluating unit via said selector switch, wherein said control and evaluating unit cooperates with said selector switch such that two of said electrodes as an electrode pair are each supplied with an alternating current from said symmetrical AC power source and the detected analog voltage signals of the other said electrodes are processed in order to reconstruct therefrom the impedance distribution of the body in the plane of the electrodes; anda common-mode signal measuring electrode, said control and evaluating unit detuning the common-mode signal of the alternating current applied on the body with respect to an electrical ground by means of said common-mode signal measuring electrode and, based on this, the symmetry of said symmetrical AC power source such that the common-mode signal on the body is minimized, and the corresponding detuning parameters are stored for each electrode pair. 2. An electroimpedance tomography in accordance with claim 1, wherein said control and evaluating unit is set up in a measuring mode to poll the detuning parameters determined for minimizing the common-mode signal for a particular electrode pair being connected and to detune the symmetrical AC power source according to these parameters. 3. An electroimpedance tomography in accordance with claim 1, further comprising a compensation AC power source having an output that is sent to the symmetrical AC power source, and said control and evaluating unit is set up to control an amplitude and phase of said output of said compensation AC power source such as to detune the symmetry of said symmetrical AC power source in order to minimize the common-mode signal on the body. 4. An electroimpedance tomograph in accordance with claim 3, further comprising an isolation transformer inserted between said symmetrical AC power source and said selector switch, wherein an output of a secondary winding of said isolation transformer forms the output of said symmetrical AC power source. 5. An electroimpedance tomography in accordance with claim 4, wherein said compensation AC power source is connected to a center tap of said secondary winding of said isolation transformer. 6. An electroimpedance tomograph in accordance with claim 1, further comprising passive compensation members connected to a first output of said symmetrical AC power source, wherein said compensation members have at least one resistor, at least one capacitor, and at least one control transistor, and said control and evaluating unit is set up to control said compensation members via at least one control transistor such that the symmetry of said symmetrical AC power source is detuned such as to minimize the... | Dräger Medical GmbH | A61B005/05 G01R027/00 G01R027/08
US07945318 | B2 | 2011-05-17 | Peripheral impedance plethysmography electrode and system with detection of electrode spacing | 2009-03-16 | In a system and method for peripheral impedance plethysmography, an electrode for application to the patient's limb includes two outer current electrodes and two inner voltage electrodes. A distance between the two inner electrodes is automatically input into an analyzing device, either as a pre-stored value or as determined automatically from the electrode. Peripheral blood flow is calculated in accordance with that distance. | 1. A method for peripheral impedance plethysmography, the method comprising: (a) attaching an electrode to an extremity of a patient, the electrode comprising two outer current electrodes and two inner voltage electrodes;(b) applying an alternating current to the two outer current electrodes;(c) measuring a voltage in each of the two inner voltage electrodes induced by the alternating current;(d) automatically receiving, in a processing device, the voltage measured in step (c) and a value of a distance between the two inner voltage electrodes, wherein the value of the distance is automatically received in the processing device by automatically generating a signal representing the value of the distance and automatically receiving the signal into the processing device; and(e) calculating, in the processing device, a volume change in the extremity in accordance with the voltage measured in step (c) and the value of the distance received in step (d) such that a result of said calculating depends on the value of the distance. 2. The method of claim 1, wherein the electrode is structured such that the distance between the inner voltage electrodes has a predetermined value, the predetermined value is previously stored in a medium accessible to the processing device, and step (d) comprises reading the predetermined value from the medium to produce said signal. 3. The method of claim 2, wherein the medium is part of the electrode, such that different electrodes having different values of the distance are usable with the processing device. 4. The method of claim 3, wherein the electrode comprises an electrical connector by which the electrode is connected to the processing device, and wherein the predetermined value is coded into the electrical connector such that the electrical connector constitutes the medium. 5. The method of claim 3, wherein the electrode comprises a resistor whose resistance corresponds to the predetermined value, and wherein step (d) comprises measuring the resistance and determining the predetermined value in accordance with the resistance. 6. The method of claim 5, wherein the predetermined value is determined in accordance with the resistance through a look-up table. 7. The method of claim 2, wherein the medium is part of the processing device. 8. The method of claim 1, wherein the electrode is structured such that the distance between the inner voltage electrodes has a variable value and step (d) comprises automatically determining the variable value. 9. The method of claim 8, wherein at least one of the electrode and the processing device comprises a distance measuring device for measuring the distance. 10. The method of claim 9, wherein the distance measuring device comprises a distance transducer in the electrode. 11. The method of claim 9, wherein the electrode further comprises a connecting portion connects the two inner voltage electrodes and which is structured to be unfolded to provide any of a plurality of valu... | Smithmarks, Inc. | A61B005/05 A61B005/04
US08442622 | B2 | 2013-05-14 | Vertebra segmentation apparatus, vertebra segmentation method, and recording medium with program for vertebra segmentation | 2010-11-26 | Based on generated tomographic images, a first feature quantity representative of a clearness of sectional shapes crossing the Z′ axis at the vertebrae is calculated for each of respective points on the Z′ axis. Further, based on the generated tomographic images, a second feature quantity representative of a clearness of sectional shapes parallel to the Z′ axis at the vertebrae is calculated for each of the respective points on the Z′ axis. A third feature quantity representative of a regularity of an array of the vertebrae is calculated for each of the respective points on the Z′ axis, based on the calculated first feature quantity and the calculated second feature quantity. Positions of the vertebrae on the Z′ axis are calculated based on the third feature quantity, which is calculated for each of respective points on the Z′ axis. | 1. A vertebra segmentation apparatus comprising: an image generator constructed and arranged to generate tomographic images at intervals along central axes of respective vertebrae;a first feature quantity calculator constructed and arranged to calculate a first feature quantity representative of a clearness of sectional shapes crossing the central axis at the vertebrae, for each of respective points on the central axis, based on the tomographic images generated by the image generator;a second feature quantity calculator constructed and arranged to calculate a second feature quantity representative of a clearness of sectional shapes parallel to the central axis at the vertebrae, for each of the respective points on the central axis, based on the tomographic images generated by the image generator;a third feature quantity calculator constructed and arranged to calculate a third feature quantity representative of a regularity of an array of the vertebrae, for each of the respective points on the central axis, based on the first feature quantity calculated by the first feature quantity calculator and the second feature quantity calculated by the second feature quantity calculator; anda vertebral position estimator constructed and arranged to estimate positions of the vertebrae on the central axis based on the third feature quantity, which is calculated by the third feature quantity calculator, for each of the respective points on the central axis. 2. A vertebra segmentation apparatus according to claim 1, wherein the first feature quantity comprises a feature quantity for extracting an annular pattern around each of the respective points on the central axis. 3. A vertebra segmentation apparatus according to claim 1, wherein the second feature quantity comprises a feature quantity for extracting a tubular pattern extending along the central axis. 4. A vertebra segmentation apparatus according to claim 1, further comprising: a function matcher constructed and arranged to match the third feature quantity calculated for each of the respective points on the central axis to a prescribed periodic function or a quasi-periodic function,wherein the vertebral position estimator is constructed and arranged to estimate the positions of the vertebrae on the central axis based on the prescribed periodic function or the quasi-periodic function matched by the function matcher. 5. A vertebra segmentation apparatus according to claim 4, wherein the prescribed periodic function or the quasi-periodic function includes a phase as a variable, and further comprising: a phase determiner constructed and arranged to determine the phase successively for the points on the central axis. 6. A vertebra segmentation apparatus according to claim 1, further comprising: an identifying information allocator constructed and arranged to allocate predefined identifying information to each of the vertebrae depending on the position thereof on the central axis, which is estimated by the... | Fujifilm Corporation | A61B005/00
US08620416 | B2 | 2013-12-31 | System and method for ischemia classification with implantable medical device | 2011-01-18 | An implantable medical device monitors ST segment data collected from EGM. ST trends are established and monitored over time. The IMD is able to discern whether the data indicate supply ischemia, demand ischemia, or other physiological causes. The IMD is then able to provide appropriate information and alerts. | 1. A method comprising: obtaining an electrogram representing cardiac data from an implantable medical device for a plurality of cardiac cycles;determining an ST segment data point for each of the cardiac cycles;initiating a timer interval when one of the ST segment data points exceeds a threshold value;calculating a slope between two of the ST segment data points that occur during the timer interval; anddetermining whether the ST segment data points indicate ischemia based on the calculated slope. 2. The method of claim 1, further comprising determining that the ST segment data points do not indicate supply ischemia if the calculated slope is less than a first slope value. 3. The method of claim 1, further comprising determining that the ST segment data points indicate demand ischemia if the calculated slope is less than a first slope value. 4. The method of claim 1, wherein the two of the ST segment data points are separated by a plurality of the ST segment data points. 5. The method of claim 1, further comprising determining a maximum ST segment data point value during the timer interval, wherein the two of the ST segment data points include the maximum ST segment data point value. 6. The method of claim 1, further comprising determining a maximum ST segment data point value during the timer interval, wherein the two of the ST segment data points occur prior to the maximum ST segment data point value. 7. The method of claim 1, further comprising: determining a maximum ST segment data point value during the timer interval;determining an amount of time from initiation of the timer interval until the maximum ST segment data point value occurs; anddetermining whether the ST segment data points indicate ischemia based on the amount of time. 8. The method of claim 7, further comprising determining that the ST segment data points do not indicate supply ischemia if the amount of time is less than a first short value or greater than a first long value. 9. The method of claim 7, further comprising determining that the ST segment data points do not indicate supply ischemia if an ST segment data point occurring at a termination of the timer interval is less than a predetermined percentage of the maximum ST segment data point value. 10. The method of claim 1, wherein determining an ST segment data point further comprises: identifying a peak of an R wave;adding a predetermined time interval to the identified peak; andidentifying a recorded value at a time identified by the predetermined time interval after the peak of the R wave as the ST segment data point. 11. The method of claim 10, further comprising: determining a PR isoelectric baseline value; andsubtracting the PR isoelectric baseline value from the recorded value to identify the ST segment data point. 12. The method of claim 1, further comprising determining that the ST segment data points do not indicate supply ischemia if less than a predetermined percentage of the ST segment data point... | Medtronic, Inc. | A61B005/0452
US08983619 | B2 | 2015-03-17 | Testing communication during implantation | 2010-03-26 | A system and method are described for testing communication through a patient during implantation using telemetry coupling electrodes on a delivery catheter. In one example, at least two telemetry coupling electrodes may be placed on or within a delivery catheter to test conductive communication with external body electrodes during implantation. In some instances, the telemetry coupling electrodes of the delivery catheter may approximate the spacing of telemetry electrodes on an IMD. In this manner, testing conductively coupled communication with telemetry coupling electrodes of the catheter may be used to mimic the telemetry electrodes on the IMD and determine a target position and/or orientation of an electrode or electrodes of the IMD for successful conductive communication through the body. | 1. An implantable medical device delivery system comprising: a catheter configured to be located at least partially within a patient, the catheter comprising: an elongate catheter body having a proximal end and a distal end, the distal end configured to be located within the patient;at least two telemetry coupling electrodes located on the distal end of the catheter body; anda feed line coupled to the at least two telemetry coupling electrodes;at least one body electrode configured for placement on the patient; andan external device, located outside of the patient, coupled to the feed line at the proximal end of the catheter and to the body electrode,wherein the external device is configured to: generate and receive a telemetry communication signal that is communicated through the patient between the body electrode and the telemetry coupling electrodes;measure at least one parameter of the telemetry communication signal indicative of a quality of the telemetry communication signal for communication of data;determine the quality of the telemetry communication signal for communication of data based on the measurement of the at least one parameter; andreport the quality of the telemetry communication signal to a user of the external device. 2. The system of claim 1, wherein the external device is further configured to predict a future telemetry communication signal quality between an implanted medical device and the external device based on the quality of the telemetry communication signal between the body electrode and the telemetry coupling electrodes. 3. The system of claim 1, wherein the at least one body electrode is one of a plurality of body electrodes configured for placement on the patient, each of which is coupled to the external device, the system further comprising: a localization module that generates position information defining positions of the telemetry coupling electrodes within the patient using a plurality of signals communicated between the plurality of body electrodes and the at least two telemetry coupling electrodes,wherein the external device is configured to measure the at least one parameter of the telemetry communication signal indicative of the quality of the telemetry communication signal between the plurality of body electrodes and the telemetry coupling electrodes for communication of data at various positions of the telemetry coupling electrodes and to store the position information generated by the localization module and the measurement of the at least one parameter indicative of the quality of the telemetry communication signal at the various positions. 4. The system of claim 3, wherein the localization module generates position information using signals communicated between at least one of: the plurality of body electrodes and the at least two telemetry coupling electrodes;the plurality of body electrodes and an implantable electrode coupled to an implantable medical device; ora sensor at the distal end of the... | Medtronic, Inc. | A61N001/08 A61B005/00 A61N001/372 A61B019/00 A61N001/375
US09174186 | B2 | 2015-11-03 | Supercrystalline colloidal particles and method of production | 2008-08-15 | The present invention concerns size- and shape-controlled, colloidal superparticles (SPs) and methods for synthesizing the same. Ligand-functionalized nanoparticles such as nonpolar-solvent-dispersible nanoparticles, are used, and the solvophobic interactions can be controlled. Advantageously, supercrystalline SPs having a superlattice structure, such as a face-centered cubic structure, can be produced. Further, the methods of the invention can provide SPs that self-assemble and are monodisperse. The SPs can be doped with organic dyes and further assembled into more complex structures. | 1. A colloidal superparticle (SP) of controlled shape and size, comprising an aggregate of nanoparticles, wherein the nanoparticles have a diameter of less than 100 nm and wherein adjacent nanoparticles are bound together through non-covalent interactions, wherein the SP comprises a superlattice structure and the SP is dispersible in a polar solvent. 2. The colloidal SP of claim 1, wherein the superlattice structure is a face-centered cubic (fcc) superlattice structure. 3. The colloidal SP of claim 1, wherein the SP comprises one type of nanoparticle. 4. The colloidal SP of claim 1, wherein the SP comprises two or more types of nanoparticle. 5. The colloidal SP of claim 1, wherein the nanoparticles are doped. 6. The colloidal SP of claim 1, wherein the SP comprises a dye inside or outside the nanoparticles. 7. The colloidal SP of claim 1, wherein the SP is synthesized by mixing a first solution comprising ligand-functionalized nanoparticles with a surfactant under aqueous conditions to form a second solution comprising nanoparticle-micelles; and mixing the second solution with a solvent to form a third solution comprising colloidal SPs. 8. The colloidal SP of claim 1, wherein the SP is spherical or partially spherical in shape. 9. The colloidal SP of claim 1, comprising one or more drugs or biologically active molecules. 10. The colloidal SP of claim 9, wherein the one or more drugs are analgesics, anesthetics, anti-inflammatory agents, anthelmintics, anti-arrhythmic agents, antiasthma agents, antibiotics, anticancer agents, anticoagulants, antidepressants, antidiabetic agents, antiepileptics, antihistamines, antitussives, antihypertensive agents, antimuscarinic agents, antimycobacterial agents, antineoplastic agents, antioxidant agents, antipyretics, immunosuppressants, immunostimulants, antithyroid agents, antiviral agents, anxiolytic sedatives, astringents, bacteriostatic agents, beta-adrenoceptor blocking agents, blood products and substitutes, bronchodilators, buffering agents, cardiac inotropic agents, chemotherapeutics, contrast media, corticosteroids, cough suppressants, diagnostic agents, diagnostic imaging agents, diuretics, dopaminergics, free radical scavenging agents, growth factors, haemostatics, immunological agents, lipid regulating agents, muscle relaxants, proteins, peptides and polypeptides, parasympathomimetics, parathyroid calcitonin and biphosphonates, prostaglandins, radio-pharmaceuticals, hormones, sex hormones, time release binders, anti-allergic agents, stimulants and anoretics, steroids, sympathomimetics, thyroid agents, vaccines, vasodilators, or xanthines. 11. The colloidal SP of claim 9, wherein the one or more biologically active molecules are chemotherapeutics, toxins, radiotherapeutics, radiosensitizing agents, gene therapy vectors, antisense nucleic acid constructs, transcription factor decoys, imaging agents, diagnostic agents, agents known to interact with an intracellular protein, polypeptides, or po... | UNIVERSITY OF FLORIDA RESEARCH FOUNDATION, INC. | B01J013/00 A61K009/16 A61K047/02
US10117721 | B2 | 2018-11-06 | Real-time surgical reference guides and methods for surgical applications | 2011-02-10 | Apparatus and methods are described which guide a surgeon in performing a reconstructive or cosmetic procedure. The apparatus and methods utilize three dimensional presentations of the target surgical site incorporating one or more virtual surgical guides which help attain proper alignment and orientation of the post surgical outcome. | 1. An apparatus for guiding a surgical reconstructive or cosmetic procedure on a patient, said apparatus comprising: a real-time, multidimensional visualization module including a photosensor and an image processor for producing real-time multidimensional visualizations of a portion of a target surgical site, the portion including an identifiable anatomical feature of the patient;a data processor communicatively coupled to the real-time, multidimensional visualization module and configured to generate in real-time: (i) at a first instant of time a first virtual surgical guide for display in conjunction with a first real-time multidimensional visualization of the portion of the target surgical site such that the first virtual surgical guide is aligned with respect to the identifiable anatomical feature, and(ii) at subsequent instances of time after the first instance of time subsequent virtual surgical guides for display in conjunction with respective subsequent real-time multidimensional visualizations of the portion of the target surgical site such that the subsequent virtual surgical guides are aligned with respect to the identifiable anatomical feature; anda display for presenting the real-time multidimensional visualizations in conjunction with the virtual surgical guides including the first real-time multidimensional visualization with the first virtual surgical guide and the subsequent real-time multidimensional visualizations with the respective subsequent virtual surgical guides,wherein the data processor is configured to: receive, from a user, an indication that (a) a first portion of the subsequent virtual surgical guides is to remain fixed regardless of movement of the identifiable anatomical feature, and (b) a second portion of the subsequent virtual surgical guides is to track movement of the identifiable anatomical feature,determine movement of the identifiable anatomical feature, using image processing, by determining a first location of the identifiable anatomical feature within the first real-time multidimensional visualization, determining subsequent locations of the identifiable anatomical feature displayed within the respective subsequent real-time multidimensional visualizations, and determining a difference between the first location and the subsequent locations,generate the subsequent virtual surgical guides in real time after the determined movement of the identifiable anatomical feature, andautomatically position the subsequent virtual surgical guides such that the first portion of the subsequent virtual surgical guides remains fixed independent of the determined movement of identifiable anatomical feature while the second portion of the subsequent virtual surgical guides is moved by the determined difference so as to be in the same position as the first virtual surgical guide relative to the identifiable anatomical feature. 2. The apparatus according to claim 1, further comprising, a user control input in communication ... | TrueVision Systems, Inc. | A01B001/00 A61F009/007 A61B034/00 A61B090/00 A61F002/00 A61F002/12
US09683835 | B2 | 2017-06-20 | System and method for scanning an intraoral cavity | 2010-12-29 | According to the invention, a method and system are provided for scanning, and for facilitating scanning of, an intraoral cavity. The target parts of the intraoral cavity that it is desired to have scanned are identified, and the spatial relationships between a scanning device and the target parts of the intraoral cavity suitable for enabling said target parts to be scanned by said scanning device, are also identified or otherwise determined. These relationships are then displayed, and the displayed relationships are used as a guide for scanning the intraoral cavity. | 1. A method for providing scanning guidance for a scanning device, comprising: (a) receiving, at a computer system, user input selecting target parts of an intraoral cavity to be subsequently scanned by said scanning device;(b) determining, using the computer system, spatial relationships between the scanning device and said target parts of the intraoral cavity suitable for enabling said target parts to be subsequently scanned by said scanning device, wherein visual guide instructions for positioning and orienting said scanning device are generated by the computer system based on said spatial relationships; and(c) displaying, with the computer system on a display, said visual guide instructions for positioning and orienting said scanning device relative to the intraoral cavity in a manner substantially conforming to said spatial relationships, thereby providing the scanning guidance for subsequent scanning of said target parts by said scanning device;wherein said scanning device scans the target parts of the intraoral cavity in a manner substantially corresponding to said spatial relationships;wherein said visual guide instructions comprise a symbol representative of a profile corresponding to an expected view of said target parts or of ancillary parts in said intraoral cavity associated with said target parts, the expected view corresponding to the determined spatial relationships, and the symbol comprising a line defining an interface between a plurality of image zones. 2. The method according to claim 1, wherein said scanning device provides 3D data of said target parts for use in a predetermined dental procedure. 3. The method according to claim 2, wherein step (a) includes receiving user input identifying said ancillary parts in said intraoral cavity associated with said target parts, wherein 3D data of said ancillary parts is also required for use in said predetermined dental procedure; andsaid step (b) comprises, for each said target part and each said ancillary part, determining for said scanning device a series of spatial parameters, each comprising scanning station data for enabling said scanning device to fully scan said corresponding target part or ancillary part;wherein said scanning station data of said series of spatial parameters includes a proximity and a relative orientation of said scanning device with respect to said target part or ancillary part such as to enable said scanning device to obtain 3D topographical data of an area of said target part or ancillary part;wherein said series provides 3D topographical data for a corresponding plurality of said areas, wherein said spatial parameters of said series are determined such that at least some adjacent said areas overlap one another. 4. The method according to claim 3, wherein step (c) comprises displaying a nominal image comprising an image of at least a portion of a nominal intraoral cavity comprising corresponding said target parts and said ancillary parts, and an image o... | ALIGN TECHNOLOGY, INC. | A61B005/05 G01B011/24 A61B005/107 A61B005/00 A61C013/00 G06F019/00 A61C009/00
US20110160434 | A1 | 2011-06-30 | FIBROMYALGIA TEST METHOD | 2009-07-06 | A method for diagnosing or testing for fibromyalgia with a specific peptide in the blood as an indicator, as well as a method for effectively evaluating or assessing a fibromyalgia drug with the peptide as an indicator. A method for diagnosing or testing for fibromyalgia or for evaluating or assessing a fibromyalgia drug, by subjecting a patient's serum to peptide analysis using as an indicator (biomarker) a peptide that demonstrates a specific expression amount in the blood of a fibromyalgia patient. | 1. A method for diagnosing or testing for fibromyalgia, comprising: measuring peptides in a blood by mass spectrometry; and using, as an indicator, a peptide having a mass-to-charge ratio (m/z) of 650.0, 656.0, 657.0, 666.0, 672.0, 679.4, 685.4, 737.5, 795.6, 810.6, 825.5, 853.7, 854.1, 861.2, 861.5, 868.4, 890.0, 911.8, 912.1, 921.4, 942.6, 949.3, 964.6, 969.8, 970.5, 1027.9, 1028.7, 1042.5, 1055.9, 1085.9, 1086.7, 1144.8, 1741.1, 1789.60, 1898.8, 1944.7, 2012.2, 2013.7, 2030.8, 2034.1, 2046.6, 2053.4, 2056.1, 2071.2, 2073.9, 2083.6, 2118.0, 2130.2, 2187.2, 2212.0, 2253.9, 2268.5, 2270.7, 2359.5, 2534.1, 2570.5, 2600.8, 2604.1, 2625.4, 2726.2, 2819.4, 3325.9, 3364.2 or 3380.1. 2. A method for diagnosing or testing for fibromyalgia, comprising: measuring peptides in a blood by mass spectrometry; and using, as an indicator, a peptide having a mass-to-charge ratio (m/z) of 615.6, 637.4, 1020.6, 1021.3, 1061.6, 1077.8, 1206.8, 1208.0, 1262.0, 1350.7, 1350.8, 1419.6, 1420.4, 1465.9, 1467.6, 1521.0, 1547.6, 1562.9, 1563.0, 1564.5, 1618.9, 1972.2, 2452.7, 2452.8, 2454.6, 2646.9, 2770.5, 2934.0, 2981.7, 3193.1 or 3263.9. 3. A method for diagnosing or testing for fibromyalgia using, as an indicator, a peptide in blood derived from high molecular weight kininogen, fibrinogen, inter-alpha-trypsin inhibitor H4, apolipoprotein E, complement C3f or transthyretin A chain. 4. A method for diagnosing or testing for fibromyalgia using, as an indicator, a peptide in blood corresponding to an amino acid sequence from position 438 to 456, position 439 to 456, position 440 to 456 or position 458 to 477 of high molecular weight kininogen. 5. A method for diagnosing or testing for fibromyalgia using, as an indicator, a peptide in blood corresponding to an amino acid sequence from position 19 to 35, position 20 to 35, position 21 to 35, position 22 to 35, position 24 to 35, position 26 to 35 or position 28 to 35 of fibrinogen alpha chain. 6. A method for diagnosing or testing for fibromyalgia using, as an indicator, a peptide in blood corresponding to an amino acid sequence from position 626 to 642 or position 669 to 687 of inter-alpha-trypsin inhibitor H4. 7. A method for testing for fibromyalgia using, as an indicator, a peptide in blood corresponding to an amino acid sequence from position 212 to 232 of apolipoprotein E. 8. A method for diagnosing or testing for fibromyalgia using, as an indicator, a peptide in blood corresponding to an amino acid sequence from position 9 to 16 of complement C3f. 9. A method for diagnosing or testing for fibromyalgia using, as an indicator, a peptide in blood corresponding to an amino acid sequence from position 101 to 123 of transthyretin A chain. 10. A method for diagnosing or testing for fibromyalgia using, as an indicator, a peptide in blood the amino acid sequence of which is any one of SEQ ID NO. 1 to 15. 11. A method for evaluating or assessing a fibromyalgia drug using, as an indicator, the peptide according to clai... | NIPPON ZOKI PHARMACEUTICAL CO., LTD. | C07K014/00 H01J049/26 A61B005/00
US07941201 | B2 | 2011-05-10 | Microprobe array structure and method for manufacturing the same | 2006-11-15 | A microprobe array structure with self-stabilization capability and a method for manufacturing the same are proposed. The microprobe array structure is used to sense various biopotential signals, and is characterized in that the etching parameters are controlled during etching to manufacture a microprobe structure with a reduced bottom cross section so that the microprobe can be firmly stabilized in the skin tissue. Moreover, a conducting layer is formed on the microprobe to sense signals. A design of electric isolation between microprobes is also proposed. The microprobe array can therefore be used for the measurement of various biopotential signals, and can also be used as a stimulus. | 1. A microprobe structure used to measure biopotential signals, said microprobe structure comprising: an elongated tapering needle member having an increasing width in a first section integrally formed with a second section, said second section having an acute contour terminating at a point to define a tip portion; anda conducting layer covering said needle member including said tip portion. 2. The microprobe structure as claimed in claim 1, wherein said tip portion is conical. 3. The microprobe structure as claimed in claim 1, wherein said tip portion can puncture the Stratum Corneum and Stratum Germinativum to allow said microprobe to get into said Stratum Germinativum and measure said biopotential signals. 4. The microprobe structure as claimed in claim 1, wherein said biopotential signals include electroencephalographic (EEG) signals, electrocardiographic (ECG) signals and electromyographic (EMG) signals. 5. The microprobe structure as claimed in claim 1, wherein the total length of said tip portion and said needle member is 250˜450 micrometers. 6. The microprobe structure as claimed in claim 1, wherein a connection between said tip portion and a remaining portion of said needle member includes edges and corners to enhance a stabilization capability of said microprobe. 7. The microprobe structure as claimed in claim 1, wherein a ratio of a cross-sectional area at an interface between said tip portion and a remaining portion of said needle member and a bottom of said needle member is larger than 1. 8. The microprobe structure as claimed in claim 1, wherein the material of said conducting layer is polymer, thin film material or metal. 9. The microprobe structure as claimed in claim 1, wherein a bottom of said needle member is connected to a substrate and used as a sensing electrode that is connected to an external circuit and used to receive said measured biopotential signals. 10. The microprobe structure as claimed in claim 9, wherein the material of said substrate is silicon. 11. The microprobe structure as claimed in claim 9, wherein a plurality of said microprobes is disposed on said substrate to form a microprobe array. 12. The microprobe structure as claimed in claim 11, wherein said conducting layers of said microprobes are connected together to measure one of said biopotential signals. 13. The microprobe structure as claimed in claim 11, wherein said conducting layers of said microprobes are mutually isolated, and said microprobe array is used to measure a plurality of said biopotential signals. 14. The microprobe structure as claimed in claim 13, wherein each of said microprobes with mutually isolated conducting layers is used as a stimulus. 15. The microprobe structure as claimed in claim 14, wherein when said microprobe is used as a stimulus, a voltage or a current is conducted into said microprobe to electrify said microprobe or render pressure to the punctured location. 16. The microprobe structure as claimed in clai... | National Chiao Tung University | A61B005/04
US07953476 | B2 | 2011-05-31 | Discrimination of supraventricular tachycardia and ventricular tachycardia events | 2006-02-16 | A method and system for discrimination of supraventricular tachycardia and ventricular tachycardia events. Morphological features points are extracted from normal sinus rhythm (NSR) complexes and used to generate a NSR template. A numerical convolution is performed using the NSR template and the feature points for each sensed NSR to give a NSR filter output. Using a plurality of NSR complexes, a median NSR filter output template is determined, where the median NSR filter output template has a median value for each value in the NSR filter output. The median NSR filter output template is then used during a tachycardia event to distinguish tachycardia events as either ventricular tachycardia events or supraventricular tachycardia events. | 1. A system, comprising: means for storing a normal sinus rhythm output template;means for sensing cardiac signals, including tachycardia complexes;means for generating a tachycardia complex output;means for comparing the tachycardia complex output to the normal sinus rhythm output template and classifying the tachycardia complexes as either ventricular tachycardia (VT) complexes or supraventricular tachycardia (SVT) complexes based on the comparison of the tachycardia complex output to the normal sinus rhythm output template; andmeans for extracting tachycardia feature points from a number of the tachycardia complexes, wherein the means for generating a tachycardia complex output includes means for generating the tachycardia complex output from the normal sinus rhythm output template and the tachycardia feature points for each of the number of tachycardia complexes. 2. The system of claim 1, wherein the means for generating a tachycardia complex output includes means for generating a tachycardia complex output representative of the tachycardia complexes. 3. The system of claim 1, further comprising means for determining whether a number of VT complexes exceeds a predetermined threshold, and means for declaring a VT event when the number of VT complexes exceeds the predetermined threshold. 4. The system of claim 1, further comprising means for receiving a number of cardiac rate signals, and means for detecting a tachycardia event, wherein the tachycardia complexes are associated with the tachycardia event. 5. The system of claim 1, further comprising means for delivering appropriate cardioversion and defibrillation therapy for the classified tachycardia complexes. 6. The system of claim 1, wherein the means for generating a tachycardia complex output includes means for convolving the tachycardia feature points with the NSR filter output template. 7. A system, comprising: means for sensing cardiac signals, including tachycardia complexes and normal sinus rhythm (NSR) complexes;means for determining a NSR output template;means for generating a tachycardia complex output representative of the tachycardia complexes;means for comparing the tachycardia complex output to the NSR output template;means for classifying the sensed tachycardia complexes as either ventricular tachycardia (VT) complexes or supraventricular tachycardia (SVT) complexes based on the comparison of the tachycardia complex output to the NSR output template;means for determining whether a number of the VT complexes exceeds a predetermined threshold; andmeans for declaring a VT event when the number of the VT complexes exceeds the predetermined threshold,wherein the means for determining the NSR output template includes: means for sampling the NSR complexes;means for locating feature points on sampled NSR complexes and forming NSR complex vectors;means for determining a NSR template based on the NSR complex vectors;means for creating a filter impulse response (h(t)) using the NS... | Cardiac Pacemakers, Inc. | A61B005/0464
US08215186 | B2 | 2012-07-10 | Lower-limb prosthesis force and moment transducer | 2011-01-06 | A computerized prosthesis alignment system includes a transducer that can measure socket reactions in the anterior/posterior plane and the right/left planes, while canceling or reducing the transverse forces on the measurements of these socket reactions. In addition, the transducer is also capable of determining the axial load or weight experienced by the prosthesis. The computerized prosthesis alignment system is in communication with a host computer. The moment data from the transducer is interpreted by the user via a computer interface. The host computer includes memory for storing one or more applications. These applications receive data from the transducer, interpret the data with discrete algebraic or fuzzy logic algorithms, and display the output numerically and graphically. Applications may also interpret the data to provide analyses to the user for aligning the prosthesis. | 1. A transducer, comprising: a transducer body;an adaptor on the transducer body for coupling to a prosthesis;a beam disposed on each of anterior, posterior, right and left sides of the adaptor, wherein the beams support the adaptor on the transducer body;a first strain-measuring device on the transducer to measure forces in a first plane; anda second strain-measuring device on the transducer to measure forces in a second plane substantially orthogonal to the first plane;wherein the first and the second strain-measuring devices are arranged to reduce measuring forces in a third plane orthogonal to the first and second planes. 2. The transducer of claim 1, wherein the first plane is the anterior/posterior plane, the second plane is the right/left plane, and the third plane is the transverse plane. 3. The transducer of claim 1, comprising third, fourth, fifth, sixth, seventh and eighth strain-measuring devices, wherein a pair of strain-measuring devices is provided on each of the beams on the anterior, posterior, right, and left sides of the adaptor. 4. The transducer of claim 1, further comprising memory to store data. 5. The transducer of claim 1, comprising a first and a second strain gage on a beam on the anterior side of the adaptor, a third and fourth strain gage on a beam on the posterior side of the adaptor, a fifth and sixth strain gage on a beam on the right side of the adaptor, and a seventh and eighth strain gage on a beam on the left side of the adaptor. 6. The transducer of claim 5, wherein two sets of four strain gages is each arranged in a balanced bridge, and each balanced bridge produces a voltage output representative of a moment in a plane. 7. The transducer of claim 6, further comprising a temperature-compensating resistor connected to each bridge, wherein voltage is supplied through the temperature-compensating resistor to each bridge. 8. The transducer of claim 7, further comprising amplification circuits for amplifying the output from the balanced bridges and an analog-to-digital conversion circuit that receives the amplified output from the balanced bridges. 9. The transducer of claim 1, further comprising circuitry to measure the axial load on the transducer. 10. A transducer, comprising: a transducer body;an adaptor on the transducer body for coupling to a prosthesis;a beam disposed on each of the anterior, posterior, right and left sides of the adaptor, wherein the beams support the adaptor to the transducer body, each beam having a strain-measuring device positioned on each of at least two opposite side surfaces of each beam;a first set of strain-measuring devices configured to measure forces in a first plane; anda second set of strain-measuring devices configured to measure forces in a second plane substantially orthogonal to the first plane;wherein the first and the second sets of strain-measuring devices are arranged to reduce measuring forces in a third plane orthogonal to the first and second planes. 11. ... | Orthocare Innovations LLC | G01L001/22
US20110152713 | A1 | 2011-06-23 | System for Recording Measured Values in or on an Organism, and Method for Producing a Componet of this System | 2009-07-17 | A system for detecting measured values in or on an organism (10) comprising at least one secondary device (12) attachable to the organism or implantable in the organism, the impedance of which depends on a state of the surroundings of the implantable secondary device, a coil arrangement (14, 16) placeable outside the organism that acts as a primary device for generating an electromagnetic alternating field in the vicinity of the secondary device in an implanted state, and an evaluation device (20) placeable outside of the organism for detecting and analysing measured values depending on the impedance of the secondary device, wherein the secondary device comprises electrically conductive means (22) which are applied using thin film technology.;The invention further relates to a method for producing a system component. | 1. A system for detecting measured values in or on an organism, comprising: at least one secondary device attachable to the organism or implantable in the organism, the impedance of which depends on a state of the surroundings of the implantable secondary device,a coil arrangement placeable outside the organism as a primary device for generating an electromagnetic alternating field in the vicinity of the secondary device in an implanted state, andan evaluation device placeable outside of the organism for detecting and analysing measured values depending on the impedance of the secondary device,wherein the secondary device comprises an electrically conductive element applied using thin film technology. 2. The system according to claim 1, wherein the secondary device comprises an electric oscillating circuit the impedance and resonant frequency of which depend on the state of the surroundings of the secondary device. 3. The system according to claim 1, wherein the electrically conductive element comprises a first electrically conductive means (26) portion applied to the implantable secondary device which constitutes a component of an oscillating circuit,an electrically insulating layer is applied on the first electrically conductive portion,second electrically conductive portion which constitutes another component of the oscillating circuit applied on the electrically insulating layer, andthe first electrically conductive portion is in contact with the second electrically conductive portion so that the oscillating circuit is formed. 4. The system according to claim 3, wherein the first electrically conductive portion comprises an exterior with windings and an interior with capacitive properties. 5. The system according to claim 4, wherein the secondary device includes a measuring device detecting measured values depending on the state of tissue surrounding the secondary device,the secondary device includes a transmitter device transmitting signals depending on the measured values, anda receiver device placeable outside the organism receives signals transmitted by the transmitter device and supplies the signals to the evaluation device. 6. The system according to claim 5, wherein the transmitter device comprises at least one RFID transponder. 7. The system according to claim 6, wherein a readable information content of the RFID transponder is changeable depending on measured values provided by the measuring device. 8. The system according to claim 5, wherein a plurality of RFID transponders is provided which are activatable or deactivatable depending on measured values provided by the measuring device. 9. A method for producing a secondary device in which an electric voltage is inducible by a primary device, comprising the steps of: providing an implant;applying first electrically conductive portion of an oscillating circuit on the implant;applying an electrically insulating layer on the electrically conductive means;applying second electric... | - | A61B005/053
US20110137141 | A1 | 2011-06-09 | Wireless Monitoring of Multiple Vital Signs | 2009-12-03 | Devices, systems and methods are presented for monitoring a plurality of vital statistics, sending the statistics over a communications network, and generating real-time feedback for the user. A device includes a microcontroller, a transceiver, a plurality of logic units, a plurality of transducers, and a plurality of sensors. The sensors measure a plurality of vital statistics for a user. Alerts are communicated to the user via a speaker. Audio signals are detected via a microphone and processed by the logic. The transceiver enables wireless communication, directly or across a network. The transceiver further communicates with a wireless communication device. The plurality of health sensors includes a thermometer, a pulse oximeter, and a blood glucose meter. | 1. A wireless in-ear health monitoring device, comprising: a processor;a memory in communication with the processor;a transceiver in communication with the processor, the transceiver being configured to communicate over a wireless network;a plurality of health sensors coupled to the device, the plurality of health sensors measuring a corresponding plurality of vital statistics;a sensing logic on the memory to operate the plurality of health sensors and to receive the corresponding plurality of vital statistics from the plurality of health sensors;a communication logic on the memory to operate a microphone and a speaker coupled to the device; anda hearing aid logic on the memory to process an incoming audio signal. 2. The device of claim 1, wherein the plurality of health sensors includes a thermometer, a pulse oximeter, and a blood glucose meter. 3. The device of claim 2, wherein the blood glucose meter and the pulse oximeter further comprise a Near Infra-Red (NIR) detector adapted to be clipped to an earlobe. 4. The device of claim 2, wherein the thermometer and the speaker are adapted to be inserted into an ear canal. 5. The device of claim 1, further comprising alerting logic to generate an alert when one of the plurality of vital statistics exceeds a maximum or minimum threshold. 6. The device of claim 5, wherein the alert is transmitted to a user of the device via the speaker. 7. The device of claim 5, wherein the alert is transmitted to a remote server on a network via the transceiver. 8. The device of claim 1, wherein the communications logic operates the transceiver to communicate with a wireless communication device. 9. The device of claim 8, wherein the transceiver is one of a BLUETOOTH and a ZIGBEE transceiver. 10. The device of claim 1, wherein the hearing aid logic processes an incoming audio signal based upon a hearing profile for a user of the device. 11. The device of claim 10, further comprising a second microphone to generate the audio signal that is processed by the hearing aid logic. 12. A system for monitoring health of a user, the system comprising: a microcontroller including a processor, a memory, and a transceiver;a plurality of health sensors in communication with the microcontroller, the plurality of health sensors measuring a corresponding plurality of vital statistics for the user;a headset in communication with the microcontroller, the headset including a speaker and a microphone;a hearing aid logic to process an incoming audio signal based upon a hearing profile for the user; andan alerting logic in communication with the microcontroller to generate an alert via the speaker when one of the plurality of vital statistics exceeds a maximum or minimum threshold;wherein the transceiver is configured to allow communication with a wireless telecommunication device. 13. The system of claim 12, wherein the plurality of health sensors includes a thermometer, a pulse oximeter, and a blood glucose meter. 14. The sy... | AT&T INTELLECTUAL PROPERTY I, L.P. | A61B005/1455
US20110118568 | A1 | 2011-05-19 | ASSISTANCE DEVICE | 2008-08-29 | Disclosed is an assistance device which can be used for the puncture into the skin with a puncture tool having a puncture needle and the measurement of a desired component in the blood with a blood glucose monitoring system equipped with a measurement tool. The assistance device comprises a main body which contacts with the skin and a lid which is connected to the main body in a freely openable and closable manner, wherein one of the main body and the lid has a puncture guide for positioning the puncture needle toward a position on the skin to be punctured with the puncture needle, and the other has a blood collection guide for positioning the measurement tool toward the position on the skin punctured with the puncture needle. | 1. An assistance device for puncturing skin with a puncture device having a puncture needle, and measuring a desired blood component with a measurement device on which a measurement tool having a blood sampler is mounted, comprising: a main body; anda lid openably and closably joined to said main body,wherein one of said main body and said lid has a puncture guide for positioning said puncture needle in alignment with a puncture position on the skin, and the other of said main body and said lid has a blood sampling guide for positioning said measurement tool in alignment with a region of the skin which is punctured by said puncture needle. 2. An assistance device according to claim 1, wherein said main body has a bleeding accelerator for pressing the skin upon contact therewith in order to cause blood to flow out from the skin, said bleeding accelerator being of a tapered shape that is progressively greater in diameter in a direction toward the skin. 3. An assistance device according to claim 1, further comprising: fixing means for fixing said main body to a blood sampling region. 4. An assistance device according to claim 1, further comprising: biasing means for biasing said lid in an opening direction; andholding means for holding said lid in a closed state. 5. An assistance device according to claim 1, wherein when said lid is closed, a centerline of a hole of said puncture guide and a centerline of a hole of said blood sampling guide are held in alignment with each other. 6. An assistance device according to claim 1, wherein the cross section of a hole of one of said puncture guide and said blood sampling guide, which is provided on said lid, is included within the cross section of a hole of the other of said puncture guide and said blood sampling guide, which is provided on said main body, in planar projection. 7. An assistance device according to claim 3, wherein said fixing means comprises two curved arcuate ring members which form an annular shape when the ring members engage with each other, and one of the two ring members has sawtooth engaging recesses for adjusting the diameter of said fixing means. 8. An assistance device according to claim 3, wherein said main body is detachably mounted on said fixing means, said assistance device further comprising: positioning means for positioning said main body with respect to said fixing means in a plurality of orientations. | TERUMO KABUSHIKI KAISHA | A61B005/157
US20110160560 | A1 | 2011-06-30 | Pressure sensor apparatus, system and method | 2009-12-29 | An implantable pressure sensor system having a sensor assembly configured and adapted to measure pressure in a volume, the sensor assembly including at least a first MEMS pressure sensor, an application-specific integrated circuit (ASIC) having memory means, temperature compensation system, drift compensation system, and power supply means for powering the sensor assembly, the first MEMS pressure sensor having a pressure sensing element that is responsive to exposed pressure, the pressure sensing element being adapted to generate a pressure sensor signal representative of the exposed pressure, the temperature compensation system being adapted to correct for temperature induced variations in the pressure sensor signal, the drift compensation system being adapted to correct for pressure and temperature induced pressure sensor signal drift. | 1. An implantable pressure sensor system, comprising: a sensor assembly configured and adapted to measure pressure in a volume, said sensor assembly including at least a first MEMS pressure sensor, an application-specific integrated circuit (ASIC) having memory means, a temperature compensation system, a drift compensation system, and power supply means for powering said sensor assembly,said first MEMS pressure sensor having a first pressure sensing element that is responsive to first exposed pressure, said first pressure sensing element being adapted to generate a first pressure sensor signal representative of said first exposed pressure,said temperature compensation system being adapted to correct for at least one temperature induced variation in said first pressure sensor signal,said drift compensation system being adapted to correct for pressure induced first pressure sensing element signal drift. 2. The sensor system of claim 1, wherein said sensor assembly includes a memory module. 3. The sensor system of claim 2, wherein said temperature compensation system includes a first temperature sensor that is responsive to a first temperature, said first temperature sensor being adapted to generate a first temperature signal representative of said first temperature. 4. The sensor system of claim 3, wherein said first temperature sensor is disposed proximate said first MEMS pressure sensor. 5. The sensor system of claim 3, wherein said temperature compensation system further includes at least one calibrated temperature coefficient, said calibrated temperature coefficient being stored in said ASIC memory means. 6. The sensor system of claim 3, wherein said ASIC is adapted to correct for said temperature induced variation in said first pressure signal as a function of said first temperature signal and said calibrated temperature coefficient. 7. The sensor system of claim 2, wherein said drift compensation system includes at least one pressure induced drift characterization. 8. The sensor system of claim 7, wherein said drift compensation system is further adapted to correct for temperature induced first pressure sensing element signal drift. 9. The sensor system of claim 8, wherein said drift compensation system includes at least one temperature induced drift characterization. 10. The sensor system of claim 9, wherein said drift compensation system includes memory means adapted to store said pressure and temperature induced drift characterizations and at least one algorithm adapted to correct said first capacitance variation output from said first sensing element with said pressure and temperature induced drift characterizations. 11. The sensor system of claim 1, wherein said sensor assembly includes a sensor housing having an internal chamber. 12. The sensor assembly of claim 11, wherein said first MEMS pressure sensor is disposed in said sensor housing internal chamber. 13. The sensor system of claim 12, wherein said sensor assembly incl... | - | A61B003/16 G01L019/04 G01L009/12 A61B005/00
US20110152664 | A1 | 2011-06-23 | MEASURING DEVICE AND A METHOD FOR DETERMINING MOVEMENT IN A TISSUE | 2009-06-02 | A measuring device contains a microwave transmitter, a microwave receiver and a control device. The control device controls the microwave transmitter in such a manner that the latter transmits a microwave signal into a tissue. The tissue contains moving constituents. The tissue controls the microwave signal. The moving constituents of the tissue change the frequency of the microwave signal. The control device controls the microwave receiver in such a manner that the latter receives the scattered and/or frequency-changed microwave signal. The control device determines from the received microwave signal a movement of the moving constituents of the tissue. | 1. A measuring device with a microwave transmitter, a microwave receiver and a control device, wherein the control device controls the microwave transmitter in such a manner that the latter transmits a microwave signal into a tissue,wherein the tissue contains moving constituents,wherein the tissue scatters the microwave signal,wherein the moving constituents of the tissue change the frequency of the microwave signal,wherein the control device controls the microwave receiver in such a manner that the latter receives the scattered and/or frequency-changed microwave signal,wherein the control device is formed in such a manner that it determines from the received microwave signal a movement of the moving constituents of the tissue and detects tissues changes in an automated manner. 2. The measuring device according to claim 1, characterized in that the measuring device contains at least one antenna,that, by means of the at least one antenna, the microwave signal transmitted by the microwave transmitter is directed successively to given locations within the tissue, and/or the microwave signal received by the microwave receiver is directed to given locations within the tissue,and that the control device determines from the received microwave signals of the given locations a microwave tomography of the tissue. 3. The measuring device according to claim 1, characterized in that the measuring device contains a display device,that the control device controls the display device, andthat the control device displays a direction and strength of the movement on the display device. 4. The measuring device according to claim 2, characterized in that the control device displays the microwave tomography on the display device. 5. The measuring device according to claim 4, characterized in that the control device displays the microwave tomography and the movement in a superimposed manner on the display device, andthat the control device displays the movement in a color-coded manner on the display device. 6. The measuring device according to claim 1, characterized in that the moving constituents of the tissue are blood, that the control device detects blood vessels and/or tumors. 7. The measuring device according to claim 1, characterized in that the transmitted microwave signal is a frequency sweep through different frequencies, andthat the microwave receiver receives at a fixed frequency. 8. The measuring device according to claim 1, characterized in that the transmitted microwave signal is largely a mono-frequency, andthat the microwave receiver receives at different frequencies by means of a frequency sweep. 9. A method for determining movement within a tissue with a microwave transmitter, a microwave receiver and a control device, wherein the control device controls the microwave transmitter and the microwave receiver,wherein the tissue contains moving constituents,the method comprising: transmitting a microwave signal into the tissue by the microwave tr... | ROHDE & SCHWARZ GMBH & CO. KG | A61B005/0265
US07952475 | B2 | 2011-05-31 | Communication system for monitoring the health status of a patient, communication device, sensor device and method | 2007-05-08 | A communication system for monitoring the health status of a patient includes a communication device and a sensor device, the communication device comprising a first communication interface for communication with the sensor device and comprising a second communication interface for communication with a health care center, the communication device being able to generate a warning message upon failure of either the first communication interface or the second communication interface, wherein the generation of a warning message is prevented if the failure of the first communication interface begins and ends during a first time interval and/or the generation of a warning message is prevented if the failure of the second communication interface begins and ends during a third time interval and/or the communication device comprises a detection device for detecting a critical or non-critical health status, and the generation of a warning message is prevented if both the duration of failure of the second communication interface is shorter than a second time interval and a non-critical health status is detected. | 1. A communication device for monitoring the status of a user, comprising a first communication interface for communication with a sensor device and a second communication interface for communication with a health care center, the communication device generating a warning message upon failure of either the first communication interface or the second communication interface, wherein the communication device comprises a detection device for detecting a critical or a non-critical status, and generating the warning message is prevented if boththe duration of failure of the second communication interface is shorter than a second time interval anda non-critical status is detected. 2. Communication device according to claim 1, further comprising a positioning device for determining the position of the communication device. 3. The communication device according to claim 1, wherein the generation of a warning message is prevented if the failure of the first communication interface begins and ends during a first time interval or if the failure of the second communication interface begins and ends during a third time interval. 4. The communication device according to claim 1, wherein the second time interval is dynamically adjustable. 5. The communication device according to claim 3, wherein the first time interval or the third time interval is dynamically adjustable. 6. Method of monitoring the status of a user (60), comprising: linking a communication device and a sensor device via a first communication interface, the communication device having a second communication interface for communication with a health care center,transferring a sensor signal to the communication device via the first communication interface, the sensor signal comprising health information as to whether the status is critical or not, andgenerating a warning message upon failure of either the first communication interface or the second communication interface, whereinthe status is detected via the sensor device and the generation of the warning message is prevented if boththe duration of the failure of the second communication interface isshorter than a second time interval andnon-critical status is detected. 7. Method according to claim 6, wherein the sensor device senses at least one of: a heart rate of the user, a blood pressure of the user, a temperature of the user, a breathing rate of the user, a ECG (electrocardiogram) of the user, a motion of the user, a galvanic skin response of the user, an oxygen level of the user, a blood flow of the user, an acidity of the user, a glucose level of the user as an indication of the status of the user. 8. Method according to claim 6, wherein the health information is implicit and the detection device detects the status of the user from the implicit health information by applying a detection algorithm. 9. Method according to claim 6, wherein the health information is explicit and the explicit health information is generated inside th... | Koninklijke Philips Electronics N.V. | G08B001/08 G08B023/00 A61B005/00
US07957788 | B2 | 2011-06-07 | Interventional magnetic resonance imaging based on global coherent free precession | 2005-02-02 | Methods and systems for obtaining intravascular magnetic resonance images of blood flow are disclosed. In preferred forms, a train of radio frequency (RF) pulses is produced by an intravascularly introduced RF transmitter positioned in proximate location to the blood flow so as to create a continuous stream of coherently excited protons of the blood flow. The coherently excited protons of the blood flow are sampled as the protons freely precess while flowing through a region of three dimensional space unaffected by the ongoing intravascular RF excitation. An image of the sampled coherently excited protons may then be constructed. | 1. A method for obtaining magnetic resonance images of flowing blood in a vascular vessel comprising: (a) positioning a radio frequency (RF) transmitter in the vascular vessel,(b) conducting global coherent free precession (GCFP) magnetic resonance imaging (MRI) of blood flowing intravascularly in the vessel by operating the RF transmitter within the vascular vessel; and(c) obtaining images of the flowing blood. 2. A method as in claim 1, wherein (b) comprises operating the RF transmitter at a frequency which is exactly on resonance and at a zero phase increment. 3. A method as in claim 1 or 2, wherein (b) comprises operating a MRI timing sequence which prevents magnetic field gradients required for imaging from causing excited protons to accumulate phase shifts as the protons move through three dimensional space. 4. A method of obtaining intravascular magnetic resonance images (MRI) of blood flow comprising: (a) intravascularly positioning a radio frequency (RF) transmitter in proximity to the blood flow,(b) causing the RF transmitter to produce a train of RF pulses so as to form a continuous stream of coherently excited protons of the blood flow,(c) performing global coherent free precession (GCFP) magnetic resonance imaging (MRI) by continuously sampling the coherently excited protons of the blood flow as the protons freely precess while flowing through a region of three dimensional space unaffected by the ongoing RF excitation, and(d) generating an image of the sampled coherently excited protons. 5. A method as in claim 4, wherein step (a) further comprises providing an intravascular catheter having a RF transmitter at a distal end thereof. 6. A method as in claim 4 or 5, which comprises operating a MRI timing sequence which prevents magnetic field gradients required for imaging from causing excited protons to accumulate phase shifts as the protons move through three dimensional space. 7. A system for obtaining magnetic resonance images of flowing blood in a vascular vessel comprising: a catheter having a radio frequency (RF) transmitter at a distal end thereof configured to be positioned in the vascular vessel,a magnetic resonance data acquisition and imaging system configured to conduct global coherent free precession (GCFP) magnetic resonance imaging (MRI) of blood flowing intravascularly in the vessel and obtain images thereof, whereinthe RF transmitter is configured to produce a train of RF pulses so as to form a continuous stream of coherently excited protons of the blood flow within the vascular vessel during operation of the data acquisition and imaging system so that images of the flowing blood therein are obtained. 8. A system as in claim 7, wherein the RF transmitter is configured to operate at a frequency which is exactly on resonance and at a zero phase increment. 9. A system as in claim 7 or 8, wherein the RF transmitter and the data acquisition and imaging system are configured to operate to achieve a MRI timing sequenc... | Duke University | A61B005/05 G01V003/00
US07957985 | B2 | 2011-06-07 | Personalized information discovery and presentation system | 2008-09-02 | A personalized health information discovery and presentation system is presented. The system has an input device, a subject-profile module, an information database, an alignment module, a solutions module, an integration engine, a plan module, a graphical user interface, and a processor. The input device allows a user to enter a personalized subject-profile, the subject-profile module receives the subject-profile and converts it into a recognized health profile. The health information database comprises a library of health information, and the alignment module aligns the recognized health profile with relevant health information. The solution's module parses the relevant information into information subsets, and the integration engine integrates the information; identifies conflicting health information, contraindications, or health warnings; creates notifications; and compiles data for presentation to the user. The plan module converts the integrated health information into a personalized health maintenance plan for display in a form selected by the user. | 1. An iterative, health information discovery and presentation system for producing an optimized-personalized health maintenance plan for a user, the system comprising: a processor;an input device operable to allow a user to enter a personalized subject-profile into a computing system, wherein the personalized subject-profile includes information from answers to a template questionnaire designed to to produce a first personalized health maintenance plan for optimization into an optimized-personalized health maintenance plan;a subject-profile module embodied in a non-transitory computer readable storage medium for receiving the personalized subject-profile and converting the personalized subject profile into a recognized health profile for the subject;a health information database embodied in a non-transitory computer readable storage medium and comprising a library of health information;an alignment module embodied in a non-transitory computer readable storage medium for aligning the recognized health profile with health information from the health information database that is determined by the system as relevant to the recognized health profile;a solutions module embodied in a non-transitory computer readable storage medium for parsing the relevant health information into information subsets that include a subset selected from the group consisting of activity, nutrition, stress, and sleep;an integration engine embodied in a non-transitory computer readable storage medium for (i) integrating rating information from the recognized health profile with the relevant health information; (ii) identifying any conflicting health information, contraindications, and health warnings obtained from the health information database; and (iii) compiling data comprising the integrated information and any conflicting health information, contraindications, and health warnings for presentation to the user;a plan module embodied in a non-transitory computer readable storage medium for converting the compiled data into a presentation format for the first personalized health maintenance plan, or an optimized-personalized health maintenance plan; andan output device operable for displaying the first personalized health maintenance plan, or the optimized-personalized health maintenance plan, to the user;wherein,the system functions to produce the optimized personalized health maintenance plan based on an “nth” iteration, n being an integer greater than or equal to 1;where n=1, a first iterated questionnaire is provided to the user and is based on the answers to the template questionnaire;where n>1, an additional nth questionnaire is provided to the user and is based on the answers to the (n−1)th questionnaire;and,the personalized health maintenance plan is optimized to the nth iteration to increase the level of specificity of the personalized health maintenance plan for the user. 2. The computing system of claim 1, wherein the subject-profile comprises one or more of th... | Dynamic Health Innovations, LLC | A61B005/00
US08506500 | B2 | 2013-08-13 | Method and apparatus for morphology-based arrhythmia classification using cardiac and other physiological signals | 2011-03-08 | A tachyarrhythmia detection and classification system classifies tachyarrhythmias based on an analysis of morphological features of a cardiac signal enhanced by using one or more physiological parameters indicative of hemodynamic stability and/or activity level. The tachyarrhythmia detection and classification system computes a measure of similarity between an arrhythmic waveform of the cardiac signal, a template waveform for that cardiac signal, such as a correlation coefficient representative of the correlation between morphological features of the arrhythmic waveform and morphological features of the template waveform. A detected tachyarrhythmia episode is classified by comparing the measure of similarity to a threshold that is dynamically adjusted using the one or more physiological parameters. | 1. A tachyarrhythmia detection and classification system, comprising: a cardiac sensing circuit configured to sense a cardiac signal;a hemodynamic sensor configured to sense a hemodynamic signal;a tachyarrhythmia detector coupled to the cardiac sensing circuit, the tachyarrhythmia detector configured to detect a tachyarrhythmia episode using the cardiac signal;a physiological parameter generator coupled to the physiological sensor, the physiological parameter generator configured to measure an arrhythmic hemodynamic parameter from the hemodynamic signal during the detected tachyarrhythmia episode; anda morphology-based tachyarrhythmia classification module coupled to the cardiac sensing circuit, the physiological parameter generator, and the tachyarrhythmia detector, the morphology-based tachyarrhythmia classification module configured to extract arrhythmic morphological features of an arrhythmic heart beat of the detected tachyarrhythmia episode from the cardiac signal, compute a measure of similarity between the extracted arrhythmic morphological features of the arrhythmic heart beat and template morphological features of a template heart beat, and compare the computed measure of similarity to a beat classification threshold to classify the arrhythmic heart beat, the morphology-based tachyarrhythmia classification module including a threshold generator configured to dynamically adjust the beat classification threshold using the arrhythmic hemodynamic parameter and a template hemodynamic parameter,wherein the hemodynamic signal comprises at least one of an impedance signal indicative of a transthoracic impedance or a cardiac impedance, a pressure signal indicative of an arterial pressure, a signal indicative of blood flow, a signal indicative of heart sounds, an oximetry signal indicative of blood oxygenation, a plethysmographic signal indicative of cardiac output, a respiratory signal, or a cardiac contractility signal, and the hemodynamic sensor comprises at least one of an impedance sensor configured to sense the impedance signal, a pressure sensor configured to sense the pressure signal, a flow sensor configured to sense the signal indicative of blood flow, a heart sound sensor configured to sense the signal indicative of heart sounds, a pulse oximetry sensor configured to sense the oximetry signal, a photoplethysmography sensor configured to sense the plethysmo graphic signal, a respiratory sensor configured to sense the respiratory signal, or a cardiac contractility sensor configured to sense the cardiac contractility signal. 2. The system of claim 1, wherein the hemodynamic signal comprises the impedance signal, and the hemodynamic sensor comprises the impedance sensor. 3. The system of claim 2, wherein the physiological parameter generator is configured to measure arrhythmic stroke impedance from the impedance signal during the detected tachyarrhythmia episode, the stroke impedance being a difference between a maximum impedance and a m... | Cardiac Pacemakers, Inc. | A61B005/0472 A61B005/02
US08852114 | B2 | 2014-10-07 | Heart pulse rate monitor | 2010-12-21 | Disclosed herein are methods and devices for monitoring a heartbeat. In one embodiment, the device may comprise a sensor package mountable over a pulse location of a user. The sensor package may include a first sensor element configured to sense at least one signal at the pulse location and to provide a first output signal comprising a heart pulse signal and a first set of noise artifacts, a second sensor element configured to sense at least one signal at the pulse location and to provide a second output signal indicative of a second set of noise artifacts, and a mechanically isolating material located between the first sensor element and the second sensor element. The device may further comprise processing circuitry connected to the sensor package and configured to extract the heart pulse signal from the first output signal based on the first output signal and the second output signal. | 1. A heartbeat monitoring device comprising: at least one sensor package mountable over a pulse location of a user, wherein each sensor package includes:a first sensor configured to sense at least one signal at the pulse location and to provide a first output signal comprising a heartbeat signal and a first set of noise artefacts;(ii) a layer of isolating material located above the first sensor; and(iii) a second sensor located above the first sensor and configured to sense at least one signal at the pulse location and to provide a second output signal comprising a second set of noise artefacts; andprocessing circuitry connected to each of the at least one sensor package, the processing circuitry being configured to extract the heartbeat signal from the first output signal based on at least the first output signal and the second output signal. 2. The heartbeat monitoring device of claim 1, wherein at least one of the first sensor and the second sensor comprises a piezoelectric sensor. 3. The heartbeat monitoring device of claim 2, wherein each piezoelectric sensor comprises a polyvinylidene fluoride film sensor. 4. The heartbeat monitoring device of claim 1, wherein the first sensor comprises a first material and the second sensor comprises a second material. 5. The heartbeat monitoring device of claim 4, wherein the first material and the second material each have an acoustic impedance that is matched to an acoustic impedance of the user's tissue. 6. The heartbeat monitoring device of claim 4, wherein the first material is different from the second material. 7. The heartbeat monitoring device of claim 6, wherein one of the first material or the second material has an acoustic impedance that is matched to an acoustic impedance of the user's tissue, and the other of the first material or the second material has an acoustic impedance that is mismatched to the acoustic impedance of the user's tissue. 8. The heartbeat monitoring device of claim 1, wherein the processing circuitry comprises a microprocessor in which at least one of the first output signal, the second output signal, or the extracted heartbeat signal is stored. 9. The heartbeat monitoring device of claim 1, wherein the processing circuitry comprises filtering circuitry configured to filter at least one of the first output signal or the second output signal. 10. The heartbeat monitoring device of claim 1, wherein the processing circuitry comprises analogue-to-digital circuitry for converting the extracted heartbeat signal into a digital signal. 11. The heartbeat monitoring device of claim 1, wherein the processing circuitry comprises wireless transmission equipment for transmitting at least the extracted heartbeat signal to an external device. 12. The heartbeat monitoring device of claim 1, further comprising a strap for attaching the device to the user. 13. The heartbeat monitoring device of claim 1, further comprising an oximeter configured to determine a level of oxygen in... | Stichting IMEC Nederland | A61B005/02 A61B005/024
US08613702 | B2 | 2013-12-24 | Device, apparatus and method for obtaining physiological signals by way of a feeding tube | 2009-08-11 | A neonatal feeding tube (10) includes electronics and instrumentation for monitoring a neonate and for provides nourishment to the neonate. The tube (10) includes electrodes (20) for sensing ECG signals of the neonate. Thermistors (22, 24, 28, 30) are placed at various points along the tube (10) to measure the neonate's temperature at those points. Breathing effort is measured by calculating a pressure differential at two pressure ports (32, 34). Pulse and SpO2 are measured at a fiber optic window (35). The electrodes (20), a distal electrode (64) and a light source (66) aid in helping a caregiver position the tip (12) of the tube (10) correctly in the stomach of the neonate. | 1. An orogastric or nasogastric feeding tube including: a tubular construction defining at least one lumen that provides a pathway for nourishment from outside of a subject into a stomach or small intestine of the subject;an oropharynx thermistor for monitoring a temperature of the subject in a region of the subject's pharynx;a hypopharynx thermistor for monitoring a temperature of the subject inferior to the pharynx of the subject; anda monitor which measures respiration flow rate from a relative temperature change between the oropharynx thermistor and the hypopharynx thermistor. 2. The orogastric or nasogastric feeding tube as set forth in claim 1, wherein the tubular construction includes at least one of: a molded or jacketed build-up of discrete parts including lumen, wires, electromechanical components;wires layered onto a center lumen that is over-molded, cast, or encapsulated after layering;molded in-wall wires in a single lumen or a multi-lumen extrusion; ora discrete bundle of wires installed into one of multiple lumens, the multiple lumens physically separate a feeding path in one of the lumens from the thermistors;wherein the wires are connected with the oropharynx thermistor and the hypopharynx thermistor. 3. The orogastric or nasogastric feeding tube as set forth in claim 1, further including: a plurality of lumens, a wiring bundle, and/or optical fiber bundles substantially in a center of the lumens. 4. The orogastric or nasogastric feeding tube as set forth in claim 1, further including: a supra-diaphragmatic pressure port for monitoring a pressure of the subject superior to the diaphragm of the subject; anda sub-diphragmatic pressure port for monitoring a pressure of the subject inferior to the diaphragm of the subject. 5. The orogastric or nasogastric feeding tube as set forth in claim 4, wherein the monitor further implements an algorithm to: from the pressure of the subject inferior to the diaphragm of the subject monitored with the sub-diphragmatic pressure port, calculate a change in pressure to generate a respiration effort signal. 6. The orogastric or nasogastric feeding tube as set forth in claim 1, wherein the monitor further implements the algorithm to: measure cardiac activity waveforms measured by the electrodes; andanalyze the waveforms to determine when the feeding tube is located at a selected location, such as when the active electrode is adjacent an SA node to position the tip based on biometric or demographic information related to the patient, e.g. biometric information, age, gender, head circumference, or the like, and confirm that the tube is properly placed, should be further advanced, or should be retracted. 7. The orogastric or nasogastric feeding tube as set forth in claim 6, the monitor further implements the algorithm to: select a subset of the electrodes that detect equipotential cardiac signals to continuously monitor cardiac activity of the subject. 8. The orogastric or nasogastric feeding tube... | Koniklijke Philips N.V. | A61B005/00 A61J015/00 A61B005/02 A61B005/042 A61B005/087 A61B005/04 A61B005/0402 A61B005/0205 A61B005/1459 G06F019/00
US09326692 | B2 | 2016-05-03 | Blood pressure measurement device and blood pressure measurement method | 2010-11-16 | A blood pressure measurement device is characterized in that the said device is designated to arithmetically apply to the computed blood pressure value a correction value that is in accordance with the difference in atmospheric pressure between the position of the blood pressure transmitting part and the position of the heart of the subject wearing the blood pressure transmitting part as obtained from the sensor signal, on the basis of the internal pressure variation of the blood pressure transmitting part. | 1. A blood pressure measurement device to measure blood pressure of a subject, the blood pressure measurement device comprising: a blood pressure transmitting part;a pressure transmission channel configured to be selectively switched between an air pressure state in which atmospheric pressure is introduced to the pressure transmission channel and an internal pressure state in which an internal pressure of the blood pressure transmitting part is introduced to the pressure transmission channel;a sensor coupled to the pressure transmission channel to measure air pressure in the pressure transmission channel in the air pressure state at first and second positions and to measure an internal pressure variation in the pressure transmission channel in the internal pressure state at the second position;a counting unit configured to count a sensor signal of the sensor with a first frequency in the air pressure state, and to count the sensor signal of the sensor with a second frequency in the internal pressure state, the second frequency being higher than the first frequency;an instruction component configured to instruct the subject to move the blood pressure transmitting part to the first position adjacent to a heart of the subject, in order to measure the air pressure by the sensor at the first position;a measurement switching unit configured to switch a pressure state for measurement between the internal pressure state and the air pressure state;a blood pressure measuring part configured to compute a blood pressure value on the basis of the internal pressure variation obtained from the sensor signal counted in the counting unit; anda blood pressure value correction unit configured to arithmetically apply to the computed blood pressure value a correction value that is in accordance with the difference in the air pressure between the second position and the first position as obtained from the sensor signal of the sensor counted in the counting unit. 2. The blood pressure measurement device according to claim 1, wherein a precision of measuring the air pressure at the second position and the first position through the use of the sensor in the air pressure state is more precise than a precision of measuring the internal pressure variation through the use of the sensor in the internal pressure state. 3. A blood pressure measurement device to measure blood pressure of a subject, the blood pressure measurement device comprising: a sensor configured to measure air pressure at first and second positions while a state for measurement is in an air pressure state in which atmospheric pressure is measured by the sensor, and to measure an internal pressure variation at the second position while a state for measurement is in an internal pressure state in which an internal pressure of a blood pressure transmitting part is measured by the sensor;a counting unit configured to count a sensor signal of the sensor with a first frequency in the air pressure state, and to c... | Seiko Epson Corporation | A61B005/0215 A61B005/022
US20110137188 | A1 | 2011-06-09 | SLEEP ASSISTANT SYSTEM | 2010-01-15 | The present invention relates to a system for sleeping assistance which comprises an electrocardiogram collector, a heartbeat recognition device, a frequency domain analysis device and a nerve active judgment device. Frequency spectrum and heartbeat rate analysis are made by the measured electrocardiogram to find out what nerve actives effect sleeping and to give corresponding feedback. | 1. A sleep assistant system, which comprises: (a) an electrocardiogram collector;(b) a heartbeat recognition device, which obtains cardiac cycle sequences from electrocardio-signals collected by said electrocardiogram collector;(c) a frequency domain analysis device, which obtains power of different frequency from said cardiac cycle sequences; and(d) a nerve active judgment device, which determines whether said cardiac cycle sequences and said power of different frequency have any effect on disrupted sleep. 2. The sleep assistant system of claim 1, wherein said electrocardiogram collector includes a detect electrode and a reference electrode for attaching on the skin surface. 3. The sleep assistant system of claim 1, wherein said heartbeat recognition device includes a peck detection procedure and a noise separation procedure. 4. The sleep assistant system of claim 1, wherein frequency domain analysis device comprises: (a) sampling the heart cycle;(b) Hamming operation for prevent heart cycle leak each other of individual frequency component in sampling;(c) fast Fourier transform to transfer sampling results of heart cycle into power density spectrum; and(d) calculates the power of low frequency (0.04˜0.15 Hz) and high frequency(0.15˜0.4 Hz) from said power density spectrum by integral quantification to obtain total power and low-high frequency power ratio. 5. The sleep assistant system of claim 1, wherein said nerve active judgment device includes: determine whether said cardiac cycle sequences or said low-high power ratio are higher than warning value for the diagnosis of sleep disrupted effect. 6. The sleep assistant system of claim 1, which further includes an automatic shut-off element. 7. The sleep assistant system of claim 1, which further comprises: (a) an acceleration sensor;(b) a motility computing device, which computes body motilities from said accelerator's signals of acceleration sensor; and(c) a hyperergasia judging device, which interprets said computed body motilities and estimates its sleep disrupted effect. 8. The sleep assistant system of claim 7, wherein said acceleration sensor is an acceleration induction integrated circuit. 9. The sleep assistant system of claim 7, wherein said motility computing device measures acceleration to obtain vector sum of triaxial directions (x, y, z). 10. The sleep assistant system of claim 1, which further comprises a wireless transceiver and a remote data processing device, wherein wherein said wireless transceiver responsible for transmit recorded electrocardio-signals of said electrocardiogram collector or recorded acceleration signals of said acceleration sensor to said remote data processing device, or receive data of said remote data processing device; said remote data processing device responsible for receive and analyse said recorded signals and return results to said wireless transceiver or to health agencies. 11. A sleep states estimation method, which includes: (a) measurin... | NATIONAL YANG-MING UNIVERSITY | A61B005/0402 A61B005/11
US20110152630 | A1 | 2011-06-23 | METHOD FOR RECORDING AND FORWARDING VITAL SIGNS, AS WELL AS DEVICE FOR THIS PURPOSE | 2010-12-03 | In a method for recording and forwarding vital signs of a user, vital signs are measured, possibly linked, and displayed to the user. Once the user has accepted the displayed vital signs or their linkages, the user releases them for transmission by way of an authentification. | 1. A method for recording and forwarding vital signs of a user, comprising: measuring the vital signs;displaying the vital signs to the user; andonce the user has accepted the displayed vital signs or a linkage of the vital signs, releasing the vital signs by the user for transmission by way of an authentification. 2. The method according to claim 1, further comprising linking the vital signs. 3. The method according to claim 1, further comprising buffer-storing the vital signs until the user releases the vital signs for transmission via an authentification, or until, the user deletes the vital signs. 4. The method according to claim 1, further comprising transmitting the authentification to a telemedical center together with the vital signs. 5. The method according to claim 1, further comprising transmitting a device identification of an authentification unit at which the user has authentificated himself, but not the authentification itself. 6. The method according to claim 1, wherein the authentification takes place by checking biometrical data via a key or a code. 7. A device for recording and forwarding vital signs of a user, comprising: an authentification unit to which medical measuring devices for measuring vital signs are able to be connected, and from which the vital signs are able to be transmitted to a medical center only following an authentification by the user. 8. The device according to claim 7, wherein the authentification unit communicates with the medical center through a base station. 9. The device according to claim 7, wherein the authentification unit includes a display unit, which allows a user to perform the authentification only following the display of the vital signs and their linkage. 10. The device according to claim 7, wherein the authentification unit includes a memory for determined vital signs, and a linking device for the determined vital signs. 11. The device according to claim 7, wherein a device identification, which is able to be transmitted following the authentification, is stored in the authentification unit. 12. The device according to claim 7, wherein for authentification purposes, the authentification unit includes at least one sensor for an input of at least one of (a) biometrical data, (b) a code and (c) a key. | - | A61B005/00
US08343078 | B2 | 2013-01-01 | Methods for evaluating the integrity of a uterine cavity | 2009-11-11 | Methods, systems and devices for evaluating the integrity of a uterine cavity. A method comprises introducing transcervically a probe into a patient's uterine cavity, providing a flow of a fluid (e.g., CO2) through the probe into the uterine cavity and monitoring the rate of the flow to characterize the uterine cavity as perforated or non-perforated based on a change in the flow rate. If the flow rate drops to zero or close to zero, this indicates that the uterine cavity is intact and not perforated. If the flow rate does not drop to zero or close to zero, this indicates that a fluid flow is leaking through a perforation in the uterine cavity into the uterine cavity or escaping around an occlusion balloon that occludes the cervical canal. | 1. A method for detecting perforations in a patient's uterus, said method comprising: transcervically introducing a probe into a uterine cavity of the patient;expanding an expandable-collapsible frame to expand a fluid-tight wall on the probe in the uterine cavity;introducing a flow of a fluid into the uterine cavity in a region between an external surface of the fluid-tight wall and an interior wall of the uterus, wherein the fluid comprises a gas; andmeasuring a flow rate of the fluid as it is being introduced into the uterine cavity, wherein a measured flow rate which drops to a level below a predetermined threshold minimum level indicates that the uterus is not perforated. 2. The method of claim 1, wherein the flow is introduced through the probe and the probe provides no outlet for flow from the uterine cavity. 3. The method of claim 1, further comprising generating a signal, responsive to the flow rate, that characterizes the uterine cavity as at least one of perforated or non-perforated. 4. The method of claim 1, further comprising generating a signal responsive to the flow rate not dropping below a predetermined level, that characterizes the uterine cavity as perforated. 5. The method of claim 4, wherein the predetermined level is 0.05 slpm. 6. The method of claim 1, further comprising generating a signal responsive to the flow rate dropping below a predetermined level, that characterizes the uterine cavity as non-perforated. 7. The method of claim 6, wherein the predetermined level is 0.05 slpm. 8. The method of claim 1, wherein measuring comprises measuring the flow rate over a predetermined first time interval after initiation of the flow. 9. The method of claim 8, wherein the first time interval is at least 5 seconds. 10. The method of claim 8, wherein the first time interval is in a range from 1 second to 60 seconds. 11. The method of claim 8, wherein the first time interval is in a range from 10 seconds to 30 seconds. 12. The method of claim 8, wherein measuring comprises measuring the flow rate over a second predetermined time interval after the first time interval. 13. The method of claim 12 wherein the second interval is at least 1 second. 14. The method of claim 12 wherein the second interval is at least 5 second. 15. The method of claim 12 wherein the second is at least 10 seconds. 16. The method of claim 1, wherein measuring comprises providing a signal, responsive to the flow rate, that characterizes the uterine cavity as at least one of perforated or non-perforated, and wherein the signal is at least one of visual, aural and tactile. 17. The method of claim 1 further comprising, prior to introducing the flow, positioning a member within the cervical canal that prevents a flow of the fluid out of the uterine cavity. 18. The method of claim 17, wherein the member is positioned about an exterior of the probe. 19. The method of claim 18 wherein the member is expanded in the cervical canal. 20. The method of cl... | Minerva Surgical, Inc. | A61B005/103 A61B005/117 A61M031/00
US08165675 | B2 | 2012-04-24 | Systems and methods for use with an implantable medical device for discriminating VT and SVT based on ventricular depolarization event timing | 2009-10-29 | Techniques are described for discriminating ventricular tachycardia (VT) from supraventricular tachycardia (SVT) using an implantable medical device capable of multi-site ventricular sensing. In one example, ventricular depolarization events are detected within a patient by the implantable device during a tachyarrhythmia, at both a left ventricular sensing site and a right ventricular sensing site. Ventricular event timing differences are then ascertained. The device compares the ventricular event timing differences detected during the tachyarrhythmia with predetermined supraventricular event timing differences for the patient, such as event timing differences previously detected within the patient during sinus rhythm or extrapolated from sinus rhythm values. The device then distinguishes VT from SVT based on the comparison of the event timing differences detected during the tachyarrhythmia with the predetermined supraventricular event timing differences. Morphological waveform analysis can also be performed, when needed, to further distinguish VT from SVT. | 1. A method for use by an implantable medical device capable of multi-site ventricular sensing within a patient, the method comprising: detecting intrinsic ventricular electrical events within the patient during a tachyarrhythmia at each of a plurality of different sites within the ventricles;detecting ventricular event timing differences between the ventricular events as detected at the plurality of different sites;comparing the ventricular event timing differences detected during the tachyarrhythmia with predetermined supraventricular event timing differences for the patient; anddistinguishing a ventricular tachycardia (VT) from a supraventricular tachycardia (SVT) based on the comparison of the ventricular event timing differences detected during the tachyarrhythmia with the predetermined supraventricular event timing differences. 2. The method of claim 1 wherein the device is equipped with left ventricular (LV) and right ventricular (RV) leads and wherein detecting intrinsic ventricular electrical events during the tachyarrhythmia at each of a plurality of different sites within the ventricles includes separately detecting ventricular events within the RV and within the LV. 3. The method of claim 2 wherein detecting ventricular timing differences LV includes detecting RV−LV timing differences (TRV−LV) between the ventricular events as detected in the RV and in the LV. 4. The method of claim 3 wherein comparing the ventricular timing differences detected during the tachyarrhythmia with the predetermined supraventricular event timing differences includes: comparing ventricular timing differences (TRV−LV) detected during the tachyarrhythmia with predetermined supraventricular timing differences (TRV−LV/SV) detected within the patient during a prior sinus rhythm. 5. The method of claim 4 wherein distinguishing VT from SVT includes: identifying the tachyarrhythmia as VT if a magnitude of the difference between TRV−LV and TRV−LV/SV exceeds a predetermined threshold. 6. The method of claim 4 wherein distinguishing VT from SVT includes: identifying the tachyarrhythmia as VT if TRV−LV and TRV−LV/SV have a different sign. 7. The method of claim 4 wherein average values for TRV−LV and TRV−LV/SV are employed in the comparison. 8. The method of claim 3 further including detecting a current ventricular rate during the tachyarrhythmia. 9. The method of claim 8 wherein comparing the ventricular timing differences detected during the tachyarrhythmia with the predetermined supraventricular event timing differences includes: comparing ventricular timing differences (TRV−LV) detected during the tachyarrhythmia at the current ventricular rate with predetermined supraventricular timing differences (TRV−LV/SV(rate)) determined for the patient for substantially the same ventricular rate. 10. The method of claim 9 further including the prior steps of: detecting ventricular timing differences (TRV−LV/SV(rate)) within the patient for a range of ventricular rat... | Pacesetter, Inc. | A61N001/36
US07949386 | B2 | 2011-05-24 | Computer-aided osteoplasty surgery system | 2007-03-20 | A method for performing computer-assisted orthopaedic surgery includes the steps of: (1) producing and displaying three-dimensional geometrical models of first and second bones, the first and second bones forming a joint; (2) identifying a zone of impingement between the first bone and the second bone on at least one of the bones; and (3) generating and displaying a color map of at least one surface of at least one bone, the at least one surface being within the zone of impingement, the color map including different colors representing different depths of bone to be removed in order to achieve an increased range of motion between the first and second bones. | 1. A method for performing computer-assisted orthopaedic surgery comprising the steps of: using a device to acquire first coordinate points on a surface of a first bone;using the device to acquire second coordinate points on a surface of a second bone that forms a joint with the first bone;producing and displaying a three-dimensional geometrical surface model of the first bone based at least initially on the acquired first coordinate points;producing and displaying a three-dimensional geometrical surface model of the second bone based at least initially on the acquired second coordinate points;moving the first and second bones relative to one another and detecting maximum amplitudes of rotation between the first and second bones;identifying a zone of impingement between the first and second bones on at least one of the bones;generating and displaying as a color map at least one surface of at least one bone model, the at least one surface being within the zone of impingement, the color map including different colors representing different depths of bone to be removed in order to achieve an increased range of motion between the two bones, the colors of the color map indicating a distance between a current bone surface and an objective bone surface for each point of the at least one surface, the objective bone surface being the bone surface that permits a target increased range of motion between the two bones to be achieved;using a tracked tool to remove bone in the zone of impingement based on real time information provided on the color map; andmoving again the first and second bones relative to one another and detecting an increased amount of rotation between the first and second bones. 2. The method of claim 1, further including the step of: representing the different colors by a continuous or discrete spectrum of varying brightness of a single color. 3. The method of claim 1, wherein the step of acquiring points on each of the surfaces of the first and second bones comprises the step of: projecting a current pointer position overtop of registered fluoroscopic images of the bones. 4. The method of claim 3, further including the step of: extracting bone contours from the registered fluoroscopic images to assist in generating the three-dimensional geometrical surface models. 5. The method of claim 1, further including the step of: superimposing the generated three dimensional geometrical surface models over registered fluoroscopic images of the bones to allow a surgeon to verify model accuracy. 6. The method of claim 1, further including the steps of: continuously, in real time, updating the three-dimensional geometrical surface models of the bones as bone is removed; and superimposing the updated three dimensional geometrical surface models over registered fluoroscopic images of the bones to allow a surgeon to compare the updated surface of the three dimensional geometrical models with original surface models in the fluoroscopic images. 7. A... | A2 Surgical | A61B005/05
US07951079 | B1 | 2011-05-31 | Apparatus and method for correlating consumable intake with physiological parameters | 2007-12-21 | A computer-assisted method and apparatus for providing information pertaining to consumptions of a consumer. | 1. A computer-assisted method for providing information pertaining to the effect of consumptions on physiological parameters of a consumer comprising: receiving at a first processor a plurality of first information pertaining to an encoded product identifier identifying a consumable, time data with respect to the time of each said consumption, and quantity data with respect to amount consumed of each consumable;for each of said identified consumables, identifying in an electronic database, consumption data using encoded product identifiers and retrieving the consumption data;receiving at a second processor second information on the status of one or more physiological parameters at a plurality of points in time associated with the consumer of said consumable;computationally generating from electronic data from said processors third information, using said first information, said second information, and the consumption data, related to the effect of the consumption of said consumable(s) upon said one or more physiological parameters of said consumer over time; anddisplaying said third information on a display to said consumer. 2. The method as recited in claim 1 wherein the said first and second processor are the same processor and the processor receives both said first information, said second information, and the consumption data and processes such information to automatically generate said third information. 3. The method as recited in claim 1 wherein an additional processor processes said first and said second information to generate said third information. 4. The method as recited in claim 1 wherein said physiological parameter relates to at least one of the group consisting of blood sugar, pulse rate and blood pressure. 5. The method as recited in claim 1 wherein the consumable is a foodstuff. 6. The method as recited in claim 1 wherein the consumable is a drug. 7. The method as recited in claim 1 wherein the consumable is a pre-packaged good. 8. The method as recited in claim 1 wherein the third information is advice pertaining to reducing the consumption of particular consumables by said consumer. 9. The method of claim 1 wherein said third information is advice pertaining to increasing consumption of one or more consumables by said consumer. 10. The method of claim 1 wherein said third information is advice pertaining to undertaking or altering a treatment regimen for the individual. 11. A computerized system for acquiring and processing information relative to the consumptions and physiological status of a consumer comprising: an apparatus comprising: an input module for inputting information pertaining to the amount of a consumable consumed by a consumer and the status of a physiological parameter of the consumer at points in time;a data storage unit configured to store information pertaining to consumptions by a consumer over a plurality of time points, and status information over a plurality of time points pertaining to at le... | - | A61B005/00 A61B019/00
US09125635 | B2 | 2015-09-08 | Sampling device for vaginal specimen | 2009-02-03 | A sampling device for vaginal specimens is revealed. The sampling device for vaginal specimens includes a tubular unit whose one end is disposed with a sampling unit, and at least one hole arranged on side walls of the same end of the tubular unit. The other end of the tubular unit is disposed with a saclike body or is inserted with an injection unit so as to send solution through the tubular unit to the hole for flushing vagina. By the sampling unit, a flushed vaginal specimen is collected. Thus the collection procedures of the specimen are easy and convenient to be performed and users have no uncomfortable feelings due to the lubricant solution. | 1. A sampling device for vaginal specimens comprising: a tubular unit whose front end is disposed with at least one hole and a flow channel is mounted in the tubular unit while the flow channel is connected with the hole, said at least one hole disposed on a side wall of the front end of the tubular unit;a sampling unit arranged on the front end of the tubular unit; anda saclike body connected with a rear end of the tubular unit and having a pour opening that connects with the flow channel;wherein the tubular unit comprises:a first tubular member for being disposed with the sampling unit and side walls of one end of the first tubular member is mounted with the hole, a first flow channel of the flow channel mounted in the first tubular member; anda second tubular member threaded with a rear end of the first tubular member, a second flow channel of the flow channel mounted in the second tubular member, the first flow channel and the second flow channel are configured to be in a staggered arrangement before in use, and the first flow channel and the second flow channel are configured to be connected with each other correspondingly while the tubular unit is rotated in use. 2. The device as claimed in claim 1, wherein a drawing string whose one end is fastened and connected with the sampling unit is arranged in the tubular unit. 3. The device as claimed in claim 1, wherein said at least one hole is covered by a film. 4. The device as claimed in claim 1, wherein at least one scraping member is disposed on surface of the sampling unit. 5. The device as claimed in claim 4, wherein said at least one scraping member is a convex surface, a protrusion or a projecting object. 6. The device as claimed in claim 1, wherein the sampling unit is an absorptive member. 7. The device as claimed in claim 1, wherein a membrane or a stop valve is disposed on the pour opening or between the pour opening and the flow channel. 8. The device as claimed in claim 1, wherein the saclike body is configured to be pivoted to or threaded with the tubular unit. 9. The device as claimed in claim 1, wherein the saclike body is with hardness, support strength or elasticity. |  | A61B005/00 A61B010/02 A61B010/00 A61B017/22
US08525386 | B2 | 2013-09-03 | Dynamically adjustable orthotic device | 2009-12-09 | An orthotic device comprises a flexible support structure comprising at least one surface for contacting a body part of a user, a plurality of pressure sensors configured for coupling to a microcontroller, and a plurality of displacement regions. Each region defines a portion of said flexible support structure, wherein each portion includes at least one sensor disposed on or below the at least one surface and at least one electrically deformable unit. Each unit comprises at least one electroactive material and is configured for coupling to the microcontroller and to a power source. The device is dynamically adjustable to change its shape and support properties, when an electrical voltage is applied to the electroactive material under the control of a microcontroller. | 1. An orthotic device comprising: a flexible support structure comprising at least one surface for contacting a body part of a user;a plurality of pressure sensors configured for coupling to a microcontroller; anda plurality of displacement regions, each said region defining a portion of a flexible support structure, wherein each said portion includes at least one said sensor disposed on or below said at least one surface, and at least one electrically deformable unit,wherein each said unit comprises at least one electroactive material and is configured for coupling to said microcontroller wherein:a microcontroller coupled to the plurality of pressure sensors and said flexible support structures determines the shape of the flexible support structure based on input from the pressure sensors; andwherein the plurality of displacement regions displace the plurality of pressure sensors of the orthotic device when the displacement regions deform. 2. The device of claim 1, wherein at least one said deformable unit comprises a unimorphic structure. 3. The device of claim 1, wherein at least one said deformable unit comprises a multi-morphic structure. 4. The device of claim 1, wherein each said displacement region has, respectively, a first physical configuration when an electrical voltage is not applied to said deformable unit, and has a second physical configuration when an electrical voltage is applied to said deformable unit. 5. The device of claim 4, wherein said second physical configuration increases physical support for a portion of the body part contacting the device compared to said first physical configuration. 6. The device of claim 1 wherein at least one said pressure sensor comprises a piezoelectric material capable of emitting an electrical voltage when depressed. 7. An orthotic module comprising: the device of claim 1;a power source electrically coupled to said orthotic device and to said microcontroller. 8. The module of claim 7 further comprising: an analog to digital converter coupled between said orthotic device and said microcontroller; anda digital to analog converter coupled between said microcontroller and said orthotic device. 9. A method of adjusting a dynamically adjustable orthotic device, comprising: a) coupling the device of claim 1 to a microcontroller and a power source;b) causing an individual in need of an orthotic device to be positioned so that an external body part of the individual contacts said at least one surface of the device;c) causing the body part to move and exert pressure on said at least one surface sufficient to depress at least one said sensor, causing each depressed sensor to emit a respective electrical signal indicative of said pressure;d) operating the microcontroller to measure each said electrical signal, wherein each said measurement is correlated to a physical address of a respective sensor and deformable unit in a said region and includes electrical properties and elapsed time between mea... | Texas Instruments Incorporated | H01L041/09 H01L041/107 H01L041/113 H01L041/18
US20110105919 | A1 | 2011-05-05 | MEDICAL DEVICE | 2009-10-30 | A medical device in accordance with the present invention comprises a connection device adapted to be wrapped around and securely attached to a patient, a data storage device attached to or integral with the connection device, wherein the data storage device comprises an input port adapted to receive medical data pertaining to the patient, a memory adapted to store the medical data, and an output port adapted to produce the medical data. A medical device in accordance with the present invention may also comprise a connection device adapted to be wrapped around and securely attached to a patient, a data storage device attached to or integral with the connection device, wherein the data storage device comprises a memory adapted to receive medial data pertinent to a patient, and a transceiver adapted to transmit and receive the medical data. | 1. A medical device, comprising: a connection device adapted to be wrapped around and securely attached to a patient;a data storage device attached to or integral with the connection device, wherein the data storage device comprises: an input port adapted to receive medical data pertaining to the patient;a memory adapted to store the medical data; andan output port adapted to produce the medical data. 2. The medical device of claim 1, wherein the input port comprises: a microphone adapted to convert speech into an audio signal for audio input;a signal conditioner adapted to process the audio signal; andan analog to digital converter adapted to digitize the processed audio signal to generate the medical data. 3. The medical device of claim 1, wherein the input port comprises: a biometric sensor adapted to generate a biometric signal;a signal conditioner adapted to process the biometric signal; andan analog to digital converter adapted to digitize the processed biometric signal to generate the medical data. 4. The medical device of claim 3, wherein the biometric sensor is adapted to measure the temperature of the patient. 5. The medical device of claim 3, wherein the biometric sensor is adapted to measure the pulse rate of the patient. 6. The medical device of claim 3, wherein the biometric sensor is adapted to measure movement of the patient's limb. 7. The medical device of claim 1, wherein the output port comprises: a digital to analog converter adapted to convert digitized audio data based on the medical data into an analog audio signal;a signal conditioner adapted to process the analog audio signal; anda speaker for audio output adapted to produce sound based on the processed analog audio signal. 8. The medical device of claim 1, wherein the data storage device is comprises a receiver adapted to connect to a corresponding connector on the connection device. 9. The medical device of claim 1, wherein medical data comprises one or more of the following: information related to a patient;information related to the patient's waiver;information related to the patient's consent;information related to patient allergies;information related to special surgical instructions; orinformation related to message to medical personnel; 10. The medical device of claim 1, further comprising a processor for receiving medical data and storing in memory. 11. The medical device of claim 1, wherein the input port comprises a wireless peer-to-peer connection interface. 12. A medical device, comprising: a connection device adapted to be wrapped around and securely attached to a patient;a data storage device attached to or integral with the connection device, wherein the data storage device comprises: a memory adapted to receive medial data pertinent to a patient; anda transceiver adapted to transmit and receive the medical data. 13. The medical device of claim 12, wherein the transceiver comprises a Bluetooth device adapted to transmit and receive medical data... | - | A61B005/024 A61B005/00 A61B005/11 G10L011/00
US20110160567 | A1 | 2011-06-30 | FUNCTIONAL MRI CARDIAC OPTIMIZATION | 2010-12-21 | An implantable or other ambulatory device, such as a pacer, defibrillator, or other cardiac function management device, can use imaging information, such as one or more of cardiac functional magnetic resonance imaging (fMRI) information or cardiac magnetic resonance imaging (MRI) information, such as for helping optimize one or more parameters of the implantable or other ambulatory device. | 1. An apparatus comprising: a processor circuit, configured to provide at least one cardiac function management device parameter, of an ambulatory medical device, having a value established at least in part using cardiac functional imaging information obtained from a functional magnetic resonance imaging (fMRI) device. 2. The apparatus of claim 1, comprising: a port, configured to be communicatively coupled to the fMRI device and to the processor circuit;wherein the processor circuit is configured to establish the value of the at least one cardiac function management device parameter at least in part using the cardiac functional imaging information obtained from the fMRI device; andwherein the cardiac functional imaging information obtained from the fMRI device is obtained using the port. 3. The apparatus of claim 1, comprising: a therapy circuit, configured to be communicatively coupled to the processor circuit and to provide a therapy to a subject;wherein the cardiac function management device parameter comprises a cardiac function management therapy control parameter configured to control at least one of a therapy timing, a therapy energy level, or a therapy location; andwherein the processor circuit is configured to control operation of the therapy circuit using the at least one cardiac function management therapy control parameter. 4. The apparatus of claim 1, wherein the at least one cardiac function management device parameter has a value established using information obtained from an imaging apparatus including at least one of an echocardiogram, a computed tomography (CT) scan, a positron emission tomography (PET) scan, magnetic resonance imaging (MRI) scan, or an X-ray image. 5. The apparatus of claim 1, wherein the processor circuit comprises a mode configured to: establish operative functionality of the ambulatory medical device that is compatible for use during an MRI procedure; andallow cardiac signal sensing during the MRI procedure. 6. The apparatus of claim 1, wherein the processor circuit comprises a mode configured to vary a value of the at least one cardiac function management device parameter when a subject associated with the ambulatory medical device is undergoing an imaging procedure. 7. The apparatus of claim 1, wherein the at least one cardiac function management device parameter has a value established using cardiac functional imaging information obtained from the fMRI device including information about at least one of a deoxyhemoglobin level or an amount of blood flow. 8. The apparatus of claim 7, wherein the establishing the value of the at least one cardiac function management device parameter comprises establishing the at least one cardiac function management device parameter associated with at least one of a decreased deoxyhemoglobin level indicated by the cardiac functional imaging information obtained from the fMRI device or an increased amount of blood flow indicated by the cardiac functional imaging infor... | - | A61B005/055
US07955279 | B2 | 2011-06-07 | Device for evaluating a position of balance for the human body | 2001-01-10 | According to the invention, the patient (4) is placed upon a patient support (3) between an X-ray source (1) and a support for the target plate (2) sensitive to the X-rays. The patient support (3) is connected to a device for detecting the horizontal position of the global axis of gravity (5) of the patient when the X-ray source (1) is in operation. The patient (4) can, for example, remain standing. The inventive device can superimpose an image (Hg) of the global axis of gravity (5) on the digital radiographic image of the patient (4). | 1. A system for evaluating an equilibrium position of the human body, including a source of X-rays, means for supporting a target plate sensitive to X-rays to produce on the target plate a radiographic image of a patient, and patient support means for supporting the patient in a fixed position between the source of X-rays and the support means for the target plate and for generating position signals imaging the horizontal position of the global axis of gravity of the patient relative to the source of X-rays and the target plate, the system including: a) scanning means which digitize the radiographic image of the patient produced on the target plate, and which thereby generate a digitized radiographic image,b) means for generating, as a function of said position signals, a digitized image of the cast shadow of the global axis of gravity in the plane of the target plate,c) means for combining with said digitized radiographic image of the patient the digitized image of the cast shadow of the global axis of gravity for generating a combined digitized image, andd) means for viewing said combined digitized imagewherein the means for generating a digitized image of the cast shadow of the global axis of gravity include:e) a calculation unit and an associated memory in which geometrical data are stored corresponding to the relative positions of the source of X-rays, the plane of the target plate, and the patient support means in a fixed system of axes, andf) a program stored in a program memory area of the memory, the program adapted for instructing the calculation unit to store said position signals at the time of taking the radiographic image, to calculate, as a function of said position signals, the straight line segment intersecting the plane of the target plate and the vertical plane passing through the source of X-rays and the horizontal position of the global axis of gravity, and to store the result of this calculation, which constitutes the image of the global axis of gravity. 2. A system according to claim 1, wherein: the scanning means includes a scanner for scanning the radiographic image produced on the target plate sensitive to X-rays and producing a series of digital signals constituting the digitized radiographic image,the program is adapted to instruct the calculation unit to receive said digital signals and stores them in the memory,the program is adapted to provide the means for combining by instructing the calculation unit to modify the stored digitized radiographic image by substituting the image of the global axis of gravity, thereby generating said combined digitized image,and the program is adapted to provide the means for viewing by instructing the calculation unit to display said combined digitized image on a monitor, or print it out on a support. 3. A system according to claim 1, wherein the program includes a calibration sequence, adapted for instructing the calculation unit to scan the radiographic image of a calibration obje... | SDGI Holdings, Inc. | A61B005/103
US07942818 | B2 | 2011-05-17 | Obstetric analgesia system | 2007-01-26 | The present invention relates to systems and methods for providing a short-acting analgesic agent in the management of pain during labor, wherein the system enables efficient, real-time prediction of contractions for the coordinated administration of analgesia such that the peak effectiveness of the analgesic coincides with the intermittent pain of labor. | 1. A maternal labor pain management system comprising: a) at least one analgesic pain management means;b) a system for administering the pain management means; andc) a system for analyzing uterine activity in the prediction of labor contractions to predict time of contraction onset relevant for use in coordinating the administration of the pain management means, wherein the system for analyzing uterine activity provides at least one signal regarding the predicted time of contraction onset and wherein the system for analyzing uterine activity is operably connected to the system for administering the pain management means, wherein the system for analyzing uterine activity comprises: i) at least two abdominal electrode sensors to receive a mixture of maternal and fetal vital signals; andii) a computing means for analyzing the mixture of maternal and fetal vital signals from the sensors, said computing means including hardware and software, wherein said software comprises a function to extract EHG data from the maternal and fetal vital signals for use in determining the predicted time of contraction onset, andwherein the system for administering the pain management means is responsive to the signal regarding the predicted time of contraction onset to automatically initiate and cease the delivery of the analgesic pain management means. 2. The maternal labor pain management system of claim 1, wherein the analgesic pain management means is an intravenous analgesic that is highly titratable with a rapid and predictable onset and short duration of bioactivity. 3. The maternal labor pain management system of claim 2, wherein the analgesic is Remifentanil. 4. The maternal labor pain management system of claim 1, wherein the system for analyzing uterine activity provides the at least one signal regarding the predicted time of contraction onset up to ten to twenty seconds before significant pressure increases in the uterus. 5. The maternal labor pain management system of claim 1, further comprising a pulse oximeter and/or a respiratory rate monitoring apparatus. 6. The maternal labor pain management system of claim 1, further comprising an automated shut-off feature for shutting off the delivery of the analgesic pain management means should detrimental effects from the delivery of the analgesic pain management means be exhibited. 7. The maternal labor pain management system of claim 1, further comprising an automated feature for increasing inspired oxygen concentration following exhibited detrimental effects from the delivery of the analgesic pain management means. 8. The maternal labor pain management system of claim 1, wherein the system for analyzing uterine activity comprises at least 8 abdominal electrode sensors. | University of Florida Research Foundation, Inc. Convergent Engineering, Inc. | A61B005/00
US07959582 | B2 | 2011-06-14 | Method and apparatus for penetrating tissue | 2007-03-21 | These and other objects of the present invention are achieved in a body fluid sampling system for use on a tissue site that includes a single drive force generator. A plurality of penetrating members are operatively coupled to the force generator. The force generator moves each of the members along a path out of a housing with a penetrating member exit, into the tissue site, stops in the tissue site, and withdraws out of the tissue site. A flexible support member couples the penetrating members to define a linear array. The support member is movable and configured to move each of the penetrating members to a launch position associated with the force generator. A user interface is configured to relay at least one of, penetrating member performance or a penetrating member setting. | 1. A body fluid sampling system for use on a tissue site, the system comprising: a drive force generator;a plurality of penetrating members operatively coupled to said force generator, said force generator moving each of said members along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site;a plurality of analyte sensors, each of an analyte sensor of the plurality of analyte sensors being associated with a penetrating member;a flexible support member carrying said penetrating members parallel to each other on its top to define an array, said support member being movable and configured to move each of said penetrating members to a launch position associated with said force generator; anda position sensor configured to provide an indication of a position of the penetrating member during actuation; anda user interface coupled to a processor, wherein in response to an input at the user interface by a user, the processor provides an input to the drive force generator of a lancing penetration parameter selected from at least one of, penetrating member depth of penetration, velocity, braking, or retraction. 2. A body fluid sampling system for use on a tissue site, the system comprising: a drive force generator;a plurality of penetrating members operatively coupled to said force generator, said force generator moving each of said members along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site;a plurality of analyte sensors, each of an analyte sensor of the plurality of analyte sensors being associated with a penetrating member;a flexible support member carrying said penetrating members parallel to each other on its top to define an array, said support member being movable and configured to move each of said penetrating members to a launch position associated with said force generator. 3. A body fluid sampling system for use on a tissue site, the system comprising: a drive force generator;a plurality of penetrating members operatively coupled to said force generator, said force generator moving each of said members along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site;a plurality of analyte sensors, each of an analyte sensor of the plurality of analyte sensors being associated with a penetrating member;a flexible support member carrying said penetrating members parallel to each other on its top to define an array, said support member being movable and configured to move each of said penetrating members to a launch position associated with said force generator; anda user interface coupled to a processor, wherein in response to an input at the user interface by a user, the processor provides an input to the drive force generator of a lancing penetration para... | Pelikan Technologies, Inc. | A61B005/00
US07949381 | B2 | 2011-05-24 | Transcutaneous analyte sensor | 2008-04-11 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A transcutaneous analyte sensor system comprising: an electrode configured to measure a level of glucose in a host;a membrane disposed over the electrode and configured to limit transport of glucose to the electrode; andsensor electronics operably connected to the electrode and configured to measure a current flow produced by the electrode to generate sensor data, wherein the sensor electronics are configured to measure the current flow in at least a picoAmp range, wherein the sensor electronics are further configured to calibrate the sensor data based on reference data from a reference lucose monitor and further configured to guide or validate the calibration based on predetermined sensitivity information between about 5 pA/mg/dL and about 25 pA/mg/dL, wherein the predetermined sensitivity information is derived from in vivo testing of at least one other continuous glucose sensor or in vitro testing, and wherein the electrode comprises an exposed electroactive working electrode with a surface area between about 0.000084 cm2 and about 0.016 cm2. 2. The transcutaneous analyte sensor system of claim 1, wherein the sensor electronics are configured to directly measure the level of current flow produced by the electrode. 3. The transcutaneous analyte sensor of claim 1, further comprising an analog-to-digital converter configured to translate the current flow measurement to a digital signal. 4. The transcutaneous analyte sensor of claim 1, wherein the membrane comprises a resistance domain configured with a permeability ratio of at least about 50:1 analyte to co-analyte concentration. 5. The transcutaneous analyte sensor of claim 1, wherein the permeability ratio is at least about 200:1. 6. The system of claim 1, wherein the sensor electronics are further configured to guide or validate the calibration based on predetermined baseline information, and wherein the baseline information is derived from in vivo testing of at least one other continuous glucose sensor or in vitro testing. 7. The method of claim 6, wherein the sensor electronics are further configured to adjust a sensitivity value or a baseline value determined from the calibration based on the predetermined sensitivity information or the predetermined baseline information. | DexCom, Inc. | A61B005/05 A61B005/00
US07953264 | B2 | 2011-05-31 | Classifying image features | 2009-06-03 | Methods are disclosed for classifying different parts of a sample into respective classes based on an image stack that includes one or more images. | 1. A method comprising using one or more processors to perform the following steps: decomposing a set of spectral images of a sample into an unmixed image set, wherein each member of the unmixed image set corresponds to a spectral contribution from a different component in the sample; andclassifying different parts of the sample into respective classes based on an image stack comprising one or more of the unmixed images. 2. The method of claim 1, wherein the images in the image stack used for classification comprise two or more of the unmixed images. 3. The method of claim 1, wherein classification involves both spectral and spatial information in the image stack. 4. The method of claim 1, further comprising: generating a composite image based on the set of spectral images of the sample,wherein the spatial intensities of two or more different members of a group consisting of the spectral images and the unmixed images are weighted differently and combined to produce the composite image, andwherein the one or more images in the image stack comprise the composite image. 5. The method of claim 4, wherein the set of spectral images comprises n images, and wherein the one or more images in the image stack used for classification comprises fewer than n images. 6. The method of claim 4, wherein the composite image is generated according to a weighting function that changes monotonically with a spectral wavelength. 7. The method of claim 4, wherein the spatial intensities of two or more different ones of the unmixed images are weighted differently and combined to produce the composite image. 8. The method of claim 4, wherein the weighting function is selected to enhance a contrast between features contributed to the composite image from the two or more different members of the group consisting of the spectral images and the unmixed images. 9. The method of claim 1, wherein a neural network is used for the classifying. 10. The method of claim 9, wherein classifying different regions of the sample into the different classes comprises identifying selected regions of the image stack that correspond to each of the individual classes, training the neural network to recognize the classes based on the selected regions, and applying the trained neural network to the additional regions of the image stack. 11. The method of claim 1, further comprising generating an output image showing classified regions of the sample. 12. The method of claim 1, further comprising obtaining the one or more images in the image stack. 13. The method of claim 1, wherein the sample is a tissue section. 14. Apparatus comprising a computer readable medium storing a program that causes a processor to carry out the method of claim 1. 15. The method of claim 1, further comprising outputting information about the classified different parts of the sample to a user. 16. The method of claim 1, wherein the set of spectral images comprises n images, and wherein the one or more image... | Cambridge Research & Instrumentation, Inc. | G06K009/00 G01N001/00 A61B005/05
US08427150 | B2 | 2013-04-23 | Magnetic resonance imaging apparatus and method of determining execution order | 2009-06-22 | It has an object to suppress reduction of an examination efficiency caused by limitation of SAR (specific absorption rate) to the minimum level in an examination that plural imaging sequences containing different imaging sequences are consecutively executed. In consideration of the time variation of time average SAR value in the overall examination, the order of executing the imaging sequences is determined so that the SAR value falls within a predetermined range and the waiting time between imaging sequences is minimum. The determination is performed by calculating a waiting time of each of all conceivable execution orders and selecting an execution order having the minimum waiting time. | 1. A magnetic resonance imaging apparatus for executing an examination comprising a plurality of imaging sequences, characterized in that the plurality of imaging sequences contain imaging sequences having different time variations in average specific absorption rate of electromagnetic waves within a predetermined time, and the magnetic resonance imaging apparatus comprises: an execution order determining unit that determines an execution order of the plurality of imaging sequences so that the average specific absorption rate within the predetermined time falls within a predetermined range and the examination time is shortest; andan imaging unit that executes the plurality of imaging sequences according to the execution order determined by the execution order determining unit. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the execution order determining unit further comprises: a time variation calculator that calculates the time variation of the average specific absorption rate within the predetermined time with respect to each of the plurality of imaging sequences by using imaging parameters;a measurement pattern extracting unit that extracts all conceivable measurement patterns each of which is provided as a conceivable execution order of the plurality of imaging sequences; andan examination time calculator that calculates the examination time with respect to each of all the extracted measurement patterns when the examination is executed in the order of the measurement pattern concerned while the average specific absorption rate within the predetermined time falls within a predetermined range, and a measurement pattern having the shortest examination time is determined as an execution order on the basis of the examination time of each measurement pattern calculated by the examination time calculator. 3. The magnetic resonance imaging apparatus according to claim 2, wherein the examination time calculator calculates a waiting time with respect to each of imaging sequences to be executed secondly and subsequently, the waiting time being defined as a time required from a time when a just-before imaging sequence is finished till a time when the imaging sequence concerned is started so that the average specific absorption rate within the predetermined time falls within a predetermined range, and calculates the examination time on the basis of the calculated waiting time. 4. The magnetic resonance imaging apparatus according to claim 2, further comprising: a display unit; andan input unit, wherein the execution order determining unit displays on the display unit, the examination time of each of the measurement patterns calculated by the examination time calculator, and determines a measurement pattern selected through the input unit as the execution order. 5. The magnetic resonance imaging apparatus according to claim 2, wherein the imaging unit comprises: a comparator that actually measures the time variation of the a... | Hitachi Medical Corporation | G01V003/00
US09072436 | B2 | 2015-07-07 | Device for measuring information regarding blood pressure | 2011-03-09 | A measurement device includes an air bag on a peripheral side, an air bag on a central side, and an air bag positioned therebetween. During blood pressure measurement, the air bags integrally behave, and the blood pressure is measured from the change in inner pressure thereof. During pulse wave measurement, the inner pressure of the air bag on the peripheral side is maintained to higher than or equal to the systolic blood pressure, and the measurement site is avascularized. The inner pressure of the air bag on the central side is maintained to the pressure near the blood pressure value, and the pulse wave is measured from the change in inner pressure thereof. The inner pressure of the air bag positioned therebetween is opened to atmospheric pressure, and the vibration of the air bag on the central side is prevented from propagating to the air bag on the peripheral side. | 1. A blood pressure information measurement device comprising: a fluid bad including a first air bag, a second air bag, and a third air bag, said fluid bag arranged such that the three air bags are independently pressure controlled and are closely attached to each other in a direction from a peripheral side, which is closest to a wrist being measured, towards a central side, which is furthest from the wrist being measured, when a cuff including the fluid bag is attached to a measurement site;an inner pressure adjustment unit including an air pump, a first air valve connected to the first air bag and a second air valve connected to the second air bag for adjusting an inner pressure of the fluid bag;a connection unit including a first two-port electromagnetic valve connected between the first air bag and the second air bag, and a second two-port electromagnetic valve connected between the second air bag and the third air bag, each of said first and second two-port electromagnetic valves being configured to switch between a connection state and a disconnection state;a control unit for controlling the inner pressure of each of the air bags by controlling the inner pressure adjustment unit and the connection unit; anda measurement unit for acquiring blood pressure information based on a change in the inner pressure of the fluid bag,wherein the control unit performs, as a first control, a control for changing the respective inner pressure of the air bags at a same pressure such that the first air valve and the second air valve are closed and, at a same time, the three air bags are connected to each other through the first and second two-port electromagnetic valves, which are both in the connection state, andwherein the control unit performs, as a second control, a control for changing the respective inner pressure of the three air bags at different pressures such that the first air valve is maintained closed and the first two-port electromagnetic valve is turned to the disconnection state to set the first air bag at a first predetermined pressure, the second air valve is opened to set the second air bag at a second predetermined pressure that is less than the first predetermined pressure, and the second two-port electromagnetic valve is controlled between the disconnection state and the connection state to set the third air bag at a third predetermined pressure that is between the first and second predetermined pressures, andwherein the measurement unit calculates a blood pressure value as the blood pressure information based on the change in the inner pressure of the fluid bag by the first control in the control unit, and acquires a pulse wave waveform as the blood pressure information based on the change in the inner pressure of the fluid bag by the second control in the control unit. | OMRON HEALTHCARE Co., Ltd. | A61B005/02 A61B005/0225 A61B005/021 A61B008/04
US08644920 | B2 | 2014-02-04 | Implantable medical device having magnetic resonance tomography antenna | 2010-12-02 | The disclosure relates to a device and a method for detecting electromagnetic fields, in particular fields occurring in imaging magnetic resonance tomography. | 1. An implantable medical device (IMD), comprising: a hermetically sealed housing,at least one control unit,at least one power supply,at least one RF antenna, andat least one RF communication unit, wherein the RF communication unit in conjunction with the at least one RF antenna are configured to allow communication between an external programming device and the implantable medical device;at least one frequency diplexer;a magnetic resonance tomography (MRT) interference detection unit having a demodulation unit, wherein the MRT interference detection unit is connected to said at least one RF communication unit orsaid at least one RF antenna orsaid at least one RF communication unit and said at least one RF antenna, by said at least one frequency diplexer;wherein said demodulation unit is configured to detect RF rotary fields typical of MRT and transmit an MRT detection signal from the MRT detection unit to the at least one control unit; and,wherein the MRT interference detection unit is additionally connected or connectable via one of said at least one frequency diplexer to an electrode line that is present, and wherein the demodulation unit is configured to detect the MRT only when RF rotary fields are simultaneously detected via the at least one RF antenna and via the electrode line. 2. The IMD according to claim 1, wherein the at least one RF antenna is configured to communicate in a medical implant communication service (MICS) frequency band. 3. The IMD according to claim 1, wherein the at least one RF antenna is configured to communicate in at least one of an industrial, scientific, and medical (ISM) band. 4. The IMD according to claim 1, wherein the demodulation unit is configured to detect the MRT based on an evaluation of a frequency spectrum typical for that MRT, that occurs at the at least one RF antenna. 5. The IMD according to claim 1, wherein the at least one RF antenna has maxima in three spatial directions, preferably in three orthogonal spatial directions. 6. The IMD according to claim 1, wherein the IMD is switched to an MRT safe state when the MRT interference detection unit has detected an MRT field, and the MRT-safe state has ventricular fibrillation (VF) detection which may be prolonged by a predetermined period of time. 7. The IMD according to claim 1, wherein the MRT interference detection unit is further connected to indicators configured to indicate MRT interference fields or to at least one MRT sensor. 8. The IMD according to claim 1, wherein the at least one control unit is configured, in case of MRT electromagnetic field detection, to: change to an MRI-safe state;remain for a prolonged period of time in an MRI-safe state;remain for a prolonged period of time in a state that is otherwise insensitive to electromagnetic interference fields;synchronize electrical measurements with field intensity minimum values that occur with periodic or pulsed electromagnetic fields;synchronize a stimulation using the minimum valu... | Biotronik CRM Patent AG | A61B005/00
US08565866 | B2 | 2013-10-22 | Method and device for monitoring acute decompensated heart failure | 2007-12-21 | An implantable medical device has an impedance determiner for determining a cardiogenic impedance signal based on electric signals sensed by connected electrodes. A parameter calculator processes the impedance signal to calculate an impedance parameter representative of the cardiogenic impedance in connection with the diastolic phase of a heart cycle. This parameter is then employed by the device for monitoring acute decompensated heart failure status of a subject. | 1. An implantable medical device comprising: an electrode connecting arrangement connectable to at least two electrodes;an impedance determiner that determines an impedance signal associated with a heart of a subject over at least a diastolic phase of a heart beat based on electric signals sensed by said at least two electrodes;a signal analyzer configured to identify a global maximum in said impedance signal during said heart beat, and to identify first local maximum in said impedance signal during the diastolic phase following the identified global maximum in the impedance signal;a parameter calculator that calculates, based on said impedance signal, an impedance parameter by applying a non-zero first weight to a diastolic portion of said impedance signal corresponding to a portion of said diastolic phase immediately preceding said identified first local maximum and up to directly following said identified first local maximum and applying a second weight to a non-diastolic portion of said impedance signal corresponding to a non-diastolic phase of said heart beat, said first weight being larger than said second weight; and an acute decompensated heart failure, ADHF, monitor that monitors an ADHF status based on said impedance parameter. 2. The implantable medical device according to claim 1, wherein said electrode connecting arrangement is connectable to at least one implantable electric lead comprising said at least two electrodes and is attachable to said heart. 3. The implantable medical device according to claim 1, wherein said parameter calculator calculates, based on said impedance signal, a diastolic impedance parameter representative of said impedance associated with said heart during at least a portion of said diastolic phase. 4. The implantable medical device according to claim 1, wherein said parameter calculator performs an isolated impedance analysis of said impedance signal during said at least a portion of said diastolic phase to calculate said impedance parameter. 5. The implantable medical device according to claim 1, wherein said second weight is zero. 6. The implantable medical device according to claim 1, wherein said impedance determiner determines said impedance signal as an average of multiple cardiogenic impedances measured during different heart beats. 7. The implantable medical device according to claim 1, further comprising: a signal applier that applies, through said at least two electrodes,a current signal or a voltage signal to at least a portion of said heart during at least said diastolic phase of said heart beat; anda signal measurer that measures a resulting voltage signal or a resulting current signal sensed by said at least two electrodes over at least a portion of said heart during at least said diastolic phase of said heart beat, said impedance determiner calculating said impedance signal based on said current signal and said resulting voltage signal or said voltage signal and said resulting current si... | St. Jude Medical AB | A61B005/0295
US08814801 | B2 | 2014-08-26 | Vessel wall monitoring apparatus | 2010-11-24 | A vessel wall monitoring apparatus includes: a first detecting unit which detects vessel diameter information based on first biological information obtained from a subject; a first producing unit which differentiates the vessel diameter information detected by the first detecting unit, to produce a vessel diameter function; a second detecting unit which detects blood pressure based on second biological information obtained from the subject; a second producing unit which performs a logarithmic operation on the blood pressure detected by the second detecting unit, to produce a logarithmic blood pressure function; and an outputting unit which produces an impedance model expression by using the vessel diameter function, the logarithmic blood pressure function, and mechanical characteristic values including a stiffness, viscosity, and inertia, and which calculates and outputs at least one of the stiffness, the viscosity, and the inertia based on the impedance model expression. | 1. A vessel wall monitoring apparatus comprising: a first detecting unit which detects vessel diameter information based on first biological information obtained from a subject;a first producing unit which differentiates the vessel diameter information detected by the first detecting unit, to produce a vessel diameter function;a second detecting unit which detects blood pressure based on second biological information obtained from the subject;a second producing unit which performs a logarithmic operation on the blood pressure detected by the second detecting unit, to produce a logarithmic blood pressure function; andan outputting unit which produces an impedance model expression by using: the vessel diameter function; the logarithmic blood pressure function; and mechanical characteristic values including a stiffness, viscosity, and inertia in a mechanical impedance model of a vessel wall, and which calculates and outputs at least one of a stiffness, viscosity, and inertia of a vessel wall of the subject based on the produced impedance model expression. 2. The vessel wall monitoring apparatus according to claim 1, wherein the first detecting unit detects the vessel diameter information based on the first biological information obtained by an ultrasonic diagnostic apparatus using an ultrasonic echo. 3. The vessel wall monitoring apparatus according to claim 1, wherein the vessel diameter information includes a plethysmogram, andthe first detecting unit detects the plethysmogram based on the first biological information obtained by a photoelectric sensor or a strain gauge. 4. The vessel wall monitoring apparatus according to claim 1, wherein the first producing unit uses the vessel diameter information detected by the first detecting unit based on the first biological information obtained at a reference appearance timing and a next reference appearance timing of a blood pressure waveform,the second producing unit uses the blood pressure detected by the second detecting unit based on the second biological information obtained at the reference appearance timing and the next reference appearance timing of the blood pressure waveform, andeach of the reference appearance timing and the next reference appearance timing is obtained from an inflection point of the blood pressure waveform. 5. The vessel wall monitoring apparatus according to claim 4, wherein the logarithmic blood pressure function includes dPb(t)=ln(Pb(t))−ln(Pb(t0)), andthe vessel diameter function includes: ε(t)=(r(t)−r(t0))/r(t0);ε′(t) which is a first derivative of ε(t); andε″(t) which is a second derivative of ε(t),where t0 indicates the reference appearance timing, t indicates a time period from the reference appearance timing to the next reference appearance timing, Pb(t) indicates a blood pressure value at t, and r(t) indicates a vessel diameter at t, andthe impedance model expression includesdPb(t)=ln(Pb(t))−ln(Pb(t0))=βε(t)+ηε′(t)+με″(t), where β indicates the stiffness, η indi... | Nihon Kohden Corporation Hiroshima University | A61B005/02 A61B005/00
US08543181 | B2 | 2013-09-24 | Sensor holder for medical sensor | 2010-10-06 | A sensor holder for a medical sensor. The sensor holder includes a housing surrounding a hollow for a subject appendage, the housing having a first channel for a gas volume following the hollow, a first aperture, a second aperture and a third aperture for a ventilation of the gas volume of the first channel. | 1. A sensor holder for a medical sensor comprising: a housing;a hollow defined by said housing and configured to receive a subject appendage;said housing defining a first channel for a gas volume adjacent said hollow, said housing further defining each of a first aperture, a second aperture and a third aperture in communication with said first channel for a ventilation of the gas volume of said first channel. 2. The sensor holder according to claim 1, wherein said first aperture and second aperture are configured to locate on opposite ends said first channel, said third aperture being therebetween and said first, second and third apertures are configured to allow airflow outside said housing to improve cooling of the appendage when the appendage is in place inside the hollow. 3. The sensor holder according to claim 2, wherein said third aperture between said opposite ends of said first channel is configured to locate on an edge of said first channel which is apart from a border between said hollow and said first channel. 4. The sensor holder according to claim 1, wherein said housing having a proximal end and a distal end comprises a second channel for a gas volume adjacent said hollow, a fourth aperture, a fifth aperture and a sixth aperture for a ventilation of the gas volume of said second channel, said fourth and fifth apertures being on opposite ends of said second channel, said sixth aperture being therebetween. 5. The sensor holder according to claim 4, wherein said first channel and second channel are on different sides of the hollow, typically on opposite sides of said hollow. 6. The sensor holder according to claim 4, wherein said third aperture and sixth aperture are perpendicular to said hollow and a bending direction of the appendage when the appendage is in place inside said hollow. 7. The sensor holder according to claim 5, wherein said first channel is substantially parallel with said second channel and which first and second channels are substantially parallel with the appendage when the appendage is inserted into said hollow. 8. The sensor holder according to claim 4, wherein said sixth aperture between said opposite ends of said second channel is configured to locate on an edge of said second channel which is apart from a border between said hollow and said second channel. 9. The sensor holder according to claim 4, wherein a cross-sectional area of at least one of said third aperture and sixth aperture is at least 32 mm2. 10. The sensor holder according to claim 4, wherein an inner width W1 of the housing including a width of the first channel, the second channel and the hollow therebetween at the proximal end of said housing is at least 30.5 mm. 11. The sensor holder according to claim 4, wherein an inner width W2 of the housing including a width of the first channel, the second channel and the hollow therebetween at the distal end of said housing is at least 25 mm. 12. The sensor holder according to claim 1, wherei... | General Electric Company | A61B005/1455
US09295426 | B2 | 2016-03-29 | Detection of a physiologic parameter with a probe | 2011-02-28 | The invention provides a device for contacting a surface of a patient's body to determine a physiologic parameter in a measurement region of a tissue of the patient. The device typically comprises a sensor responsive to the physiologic parameter and a probe housing the sensor. The probe is constructed to allow the sensor to be secured at a sensing site adjacent to the measurement region, without disturbing the blood flow within the measurement region of the tissue. The device may also include a means for reducing interference in the sensing area. Preferably, the device further comprises an indicating means operably connected to the sensor for indicating an analyte quantity and/or concentration associated with the physiologic parameter. | 1. A device for contacting a tissue sensing site of a patient to determine a physiologic parameter of the patient, comprising: a probe body having a longitudinal axis, a top covered portion parallel to said longitudinal axis, said top portion comprising a material that is impermeable and a bottom open portion defining a perimeter structured to be positioned against a surface of the tissue sensing site, said perimeter defining a recess therewithin, said probe body and said perimeter structured to prevent application of excessive pressure to the sensing site to allow for substantially undisturbed blood flow in the tissue; anda sensor responsive to a concentration of an analyte through a chemical reaction with a reactant, said analyte concentration representative of the physiologic parameter, said sensor having a first end and a sensor body, said first end coupled to said probe body and said sensor body unobstructedly suspended with the recess;wherein upon placement of the probe body against the sensing site, said impermeable top portion and said perimeter are structured to create a covered, isolated measurement area within the recess such that the sensor is exposed to the tissue sensing site and the covered, isolated measurement area prevents the analyte from dissipating from the measurement area. 2. The device of claim 1, wherein the probe body further comprises a means for reducing interference at the sensing site. 3. The device of claim 2, wherein the means for reducing interference at the sensing site is an integral portion of the probe. 4. The device of claim 1, wherein the probe is constructed to allow the sensor to be secured at the sensing site. 5. The device of claim 4, wherein a surface of the probe is secured to the sensing site with an adhesive. 6. The device of claim 1, wherein the sensor is detachable from the probe body. 7. The device of claim 2, wherein the means for reducing interference is adapted to reduce interference from electromagnetic radiation, air flow or both. 8. The device of claim 3, wherein the means for reducing interference comprises an isolating means for isolating and preventing said analyte from dissipating from the measurement area and for allowing the analyte to come into equilibrium within the measurement area. 9. The device of claim 8, wherein the isolating means for isolating and preventing an analyte from dissipating from the measurement area represents an integral portion of the probe. 10. The device of claim 1, wherein the device is sized to fit into the nares, mouth, or cheek of the patient. 11. The device of claim 1 wherein the analyte is selected from the group consisting of hydrogen ions, CO2, pCO2, O2, glucose, pyruvate, acetyl-CoA, citrate, isocitrate, alpha-ketoglutarate, succinyl-CoA, ADP, ATP, succinate, fumarate, malate oxaloacetate, NAD, NADH, and combinations thereof. 12. The device of claim 1, further comprising an indicating means for indicating analyte concentration in the patient... | ExoStat Medical, Inc. | A61B005/05 A61B005/00 A61B005/01 A61B005/1459 A61B005/145 A61B005/1473 A61B008/06 A61B008/12
US20110152653 | A1 | 2011-06-23 | SYSTEMS AND METHODS FOR IMPLEMENTING RAPID RESPONSE MONITORING OF BLOOD CONCENTRATION OF A METABOLITE | 2009-06-07 | Systems and methods for monitoring the concentration of glucose or other metabolites by way of a low-volume microdialysis-probe (10) operative in accordance with a pulsed mode of flow through a flow path (12) defined by a layered structure. Also described are a system and a method for deployment of the microdialysis probe (10) within the body and a technique for deriving a metabolite concentration of the basis of a rate of change of an optical parameter. | 1. A method for monitoring the concentration of a metabolite in vivo within a body fluid, the method comprising the steps of: (a) providing a monitoring device comprising: (i) a microdialysis probe formed at least in part from a membrane permeable to the metabolite, the probe defining a probe volume containing dialysate,(ii) a measuring cell having a measuring volume, and(iii) a flow path connecting from the probe volume to the measuring volume;(b) bringing the probe into contact with the body fluid;(c) maintaining substantially zero flow conditions during a diffusion period to allow diffusion of the metabolite into the dialysate within the microdialysis probe; and(d) generating a dialysate flow to carry a quantity of dialysate from the probe volume along the flow path to the measuring volume,wherein said generating a dialysate flow is performed such that said quantity of dialysate passes from the probe volume to the measuring volume in a given transport time, said transport time being no more than 25 percent of said diffusion period. 2. A device for monitoring the concentration of a metabolite in vivo within a body fluid, the device comprising: (a) a microdialysis probe for bringing into contact with the body fluid, said probe being formed at least in part from a membrane permeable to the metabolite, said probe defining a probe volume containing dialysate;(b) a measuring cell having a measuring volume;(c) a flow path connecting from the probe volume to the measuring volume; and(d) a fluid flow controller deployed to control flow of dialysate through said probe volume and said flow path to said measuring volume, said fluid flow controller being configured to generate a pulsed flow pattern including: (i) a diffusion period of substantially zero flow conditions to allow diffusion of the metabolite into the dialysate within said probe; and(ii) a dialysate flow to carry a quantity of dialysate from the probe volume along said flow path to the measuring volume,wherein said flow controller generates said dialysate flow such that said quantity of dialysate passes from the probe volume to the measuring volume in a given transport time, said transport time being no more than 25 percent of said diffusion period. 3. The method of claim 1, wherein said transport time is no more than 10 percent of said diffusion period. 4. The method of claim 1, wherein said substantially zero flow conditions and said dialysate flow are generated repeatedly as part of a pulsed flow pattern, and wherein a single fluid pulse of said pulsed flow pattern transports said quantity of dialysate from the probe volume to the measuring volume. 5. The method of claim 1, wherein at least part of the flow path from the probe volume to the measuring volume is formed as a channel enclosed between at least two layers of a layered structure. 6. The method of claim 5, wherein said layered structure further comprises a reagent inlet in flow connection with said flow path for introducing a ... | G-Sense Ltd. | A61B005/1468
US20110130637 | A1 | 2011-06-02 | PROTEIN PRODUCT FOR MODIFYING CARDIOVASCULAR HEALTH | 2009-03-19 | A composition for use in the treatment or prophylaxis of conditions mediated by endothelial function in mammals is provided. The composition includes a milk protein hydrolysate generated by treating milk or milk whey protein with a proteolytic enzyme having subtilisin or subtilisin-like activity and/or glutamyl endopeptidase or glutamyl endopeptidase-like activity. The composition is useful in the management of vascular conditions. | 1. A composition for use in the treatment or prophylaxis of conditions mediated by endothelial function in mammals, comprising a milk protein hydrolysate prepared by treating milk or milk whey protein with a proteolytic enzyme having subtilisin or subtilisin-like activity and/or glutamyl endopeptidase or glutamyl endopeptidase-like activity. 2. The composition of claim 1 wherein the proteolytic enzyme is derived from Bacillus species. 3. The composition of claim 1 wherein the proteolytic enzyme is derived from Bacillus licheniformis. 4. The composition of claim 1 wherein the proteolytic enzyme comprises Alcalase™ from Bacillus licheniformis. 5. The composition of claim 1 wherein the hydrolysate comprises a fractionated whey protein hydrolysate. 6. The composition of claim 5 wherein the whey fraction comprises a molecular weight fraction of 5 kDa or less. 7. The composition of claim 6 wherein the whey fraction contains greater than 60% of peptide material having a molecular weight of less than 2 kDa. 8. The composition of claim 7 wherein the whey fraction comprises greater than 70% of peptide material having a molecular weight of less than 2 kDa. 9. The composition of claim 8 wherein the whey fraction comprises greater than 80% of peptide material having a molecular weight of less than 2 kDa. 10. The composition of claim 9 wherein the whey fraction comprises greater than 55% of peptide material having a molecular weight of less than 1 kDa. 11. The composition of claim 10 wherein the whey fraction comprises greater than 65% of peptide material having a molecular weight of less than 1 kDa. 12. The composition of claim 11 wherein the whey fraction comprises greater than 40% of peptide material having a molecular weight of less than 500 Daltons. 13. The composition of claim 1 wherein the milk protein hydrolysate has a degree of hydrolysis of greater than 10%, preferably from about 15% to about 25%. 14. The composition of claim 13 wherein the degree of hydrolysis is approximately 19%. 15. The composition as claimed in claim 1 comprising between 5 g and 18 g of the milk protein hydrolysate. 16. The composition of claim 15 comprising about 14 g of the milk protein hydrolysate. 17. The composition of claim 1 wherein the milk protein hydrolysate has greater than 80% solubility between pH 2.0 to 8.0. 18. The composition of claim 17 wherein the milk protein hydrolysate has greater than 95% solubility between pH 2.0 to 8.0. 19. The composition of claim 1 wherein the whey protein hydrolysate has a foam stability of less than 10% after 15 minutes standing following foam formation. 20. The composition of claim 19 having a foam stability of less than 5% after 15 minutes standing following foam formation. 21. The composition of claim 1 in which the whey protein hydrolysate has an amino acid composition as outlined for the permeate fraction in the examples presented herein. 22. The composition of claim 1 for use in the treatment, prophylaxis ... | University of Dundee | A61B005/145 A61K038/02 A61P009/00 A61P009/12 A61P003/10 A61P003/00
US20110124992 | A1 | 2011-05-26 | OPTIMIZED SENSOR GEOMETRY FOR AN IMPLANTABLE GLUCOSE SENSOR | 2011-01-28 | An implantable sensor for use in measuring a concentration of an analyte such as glucose in a bodily fluid, including a body with a sensing region adapted for transport of analytes between the sensor and the bodily fluid, wherein the sensing region is located on a curved portion of the body such that when a foreign body capsule forms around the sensor, a contractile force is exerted by the foreign body capsule toward the sensing region. The body is partially or entirely curved, partially or entirely covered with an anchoring material for supporting tissue ingrowth, and designed for subcutaneous tissue implantation. The geometric design, including curvature, shape, and other factors minimize chronic inflammatory response at the sensing region and contribute to improved performance of the sensor in vivo. | 1. An implantable sensor for use in measuring a concentration of an analyte in a bodily fluid, the sensor comprising: a substantially cylindrical body comprising a sensing region and adapted for transport of at least one analyte from a bodily fluid to the sensing region, wherein the substantially cylindrical body is defined by a curved lateral surface and two ends, and wherein the sensing region is located on the curved lateral surface such that when a foreign body capsule forms around the sensor in vivo, a contractile force is exerted by the foreign body capsule toward the sensing region. 2. The sensor of claim 1, wherein the substantially cylindrical body comprises a thermoplastic plastic or a thermoset plastic. 3. The sensor of claim 2, wherein the substantially cylindrical body comprises epoxy. 4. The sensor of claim 1, further comprising a battery and a circuit board housed within the substantially cylindrical body. 5. The sensor of claim 1, further comprising an application specific integrated circuit housed within the substantially cylindrical body. 6. The sensor of claim 1, wherein the sensing region comprises a plurality of sensors. 7. The sensor of claim 6, wherein the plurality of sensors are configured to sense a same analyte. 8. The sensor of claim 6, wherein the plurality of sensors are configured to sense at least two different analytes. 9. The sensor of claim 1, wherein the sensing region is configured for measurement of a concentration of glucose in the bodily fluid. 10. The sensor of claim 1, wherein the sensing region comprises an electrode. 11. The sensor of claim 1, wherein the sensor is configured for implantation in a soft tissue of a body. 12. The sensor of claim 1, further comprising a porous biointerface material that covers at least a portion of the sensing region. 13. The sensor of claim 12, wherein the biointerface material comprises interconnected cavities dimensioned and arranged to create contractile forces that counteract a generally uniform downward fibrous tissue contracture caused by the foreign body capsule in vivo so as to interfere with formation of occlusive cells. 14. The sensor of claim 1, wherein the sensing region encircles the curved lateral surface. | DexCom, Inc. | A61B005/1473
US20110137183 | A1 | 2011-06-09 | EVALUATE AORTIC BLOOD PRESSURE WAVEFORM USING AN ADAPTIVE PERIPHERAL PRESSURE TRANSFER FUNCTION | 2009-06-30 | The invention relates to method for reconstructing an aortic blood pressure waveform of a person from a peripheral blood pressure waveform of the person comprising the steps of determining at least one pre-selected parameter of the peripheral blood pressure waveform, reconstructing the aortic blood pressure waveform from the peripheral blood pressure waveform using a pressure transfer function having at least one adjustable characteristics, wherein said adjustable characteristics is determined using the at least one pre-selected parameter of the peripheral blood pressure waveform. The invention further relates to a device for reconstructing an aortic blood pressure waveform from a peripheral blood pressure waveform and a computer program product. | 1. A method for reconstructing an aortic blood pressure waveform of a person from a peripheral blood pressure waveform of the person comprising: determining at least one pre-selected parameter of the peripheral blood pressure waveform; andreconstructing the aortic blood pressure waveform from the peripheral blood pressure waveform using a pressure transfer function having at least one adjustable characteristic, wherein said adjustable characteristics determined using the at least one pre-selected parameter of the peripheral blood pressure waveform. 2. The method according to claim 1, wherein the at least one pre-selected parameter of the blood pressure waveform is selected from a group comprising: cardiac output (CO), heart rate (HR), stroke volume (SV), total peripheral resistance (TPR), pressure wave delay (Delay), peripheral mean arterial pressure (MAP), and diastolic arterial pressure (DBP). 3. The method according to claim 2, further comprising selecting a resonance peak of the pressure transfer function for the adjustable characteristic. 4. A method according to claim 3, wherein the resonance peak of the pressure transfer function (Fpeak) is adjusted according to the following equation: Fpeak=k*CO+const,where k is an assigned coefficient value, and const is an assigned constant value. 5. A method according to claim 3, wherein the resonance peak of the pressure transfer function (Fpeak) is adjusted according to one of the following equations: Fpeak=k1*(MAP−DBP)+k2*HR−k3*Delay+const, orFpeak=k4*(MAP−DBP)+k5*HR+constwhere k1, k2, k3, k4, and k5 are assigned coefficient values, and const is an assigned constant value. 6. The method according to claim 1, further comprising measuring the peripheral blood pressure waveform. 7. A method according to claim 6, wherein the peripheral blood pressure waveform is measured during an exercise of the person. 8. The method according to claim 6, wherein the peripheral blood pressure waveform is measured non-invasively. 9. The method according to claim 1, further comprising calculating a physiological parameter characteristic of a health condition of the person from at least aortic blood pressure waveform data. 10. A device for reconstructing an aortic blood pressure waveform from a peripheral blood pressure waveform comprising: a memory unit for storing a pressure transfer function having at least one adjustable characteristic and at least the peripheral blood pressure waveform; anda processor arranged for: determining at least one pre-selected parameter of the peripheral blood pressure waveform;determining the at least one adjustable characteristic of the pressure transfer function from the at least one pre-selected parameter of the peripheral blood pressure waveform; andreconstructing the aortic blood pressure waveform from the peripheral blood pressure waveform using the pressure transfer function with the at least one adjustable characteristic. 11. A device according to claim 10, further compris... | BMEYE B.V. | A61B005/021
US20110125058 | A1 | 2011-05-26 | PATIENT-ENACTED SAMPLING TECHNIQUE | 2010-11-24 | The present invention is generally directed to systems and methods for sampling fluids from subjects. The fluid may be any suitable bodily fluid, e.g., blood or interstitial fluid. In some cases, the subject is a patient. The subject may have a device that can be applied to the skin of the subject (e.g., by the subject, or another person), and the device is able to obtain a sample of fluid. The fluid may be stored within a reservoir in the device, and the fluid may be obtained from the subject at any convenient time, e.g., at home, away from a healthcare setting, etc. In some embodiments, the device, or a portion thereof, may be returned to a clinical and/or laboratory setting to analyze the fluid stored within the device. | 1. A method for obtaining a fluid sample from the skin and/or from beneath the skin of a subject, comprising: providing a fluid access and storage device, comprising a fluid storage reservoir, to a non-healthcare-professional person;directing the non-healthcare-professional person to use the device whereby, in the absence of a healthcare professional, the device is applied to the skin of the subject to obtain a fluid sample from the skin and/or from beneath the skin of the subject into the fluid storage reservoir of the device, and removed from the skin of the subject thereby defining a stored sample of fluid within the device; andtransporting the fluid storage reservoir including the stored sample of fluid to a clinical and/or laboratory setting. 2-5. (canceled) 6. The method of claim 1, wherein the device further comprises a support structure. 7-10. (canceled) 11. The method of claim 6, wherein the device further comprises a transport means for obtaining the sample from the skin and/or from beneath the skin of a subject, the transport means in fluidic communication with the reservoir. 12-14. (canceled) 15. The method of claim 11, wherein the transport means comprises a microneedle. 16-19. (canceled) 20. The method of claim 11, wherein the device is able to withdraw the transport means from the skin of the subject after exposure to the skin. 21. (canceled) 22. The method of claim 11, wherein the transport means is able to move within the device. 23-24. (canceled) 25. The method of claim 1, wherein the device is able to apply a vacuum to the skin of the subject. 26. The method of claim 1, wherein the device comprises a self-contained vacuum chamber. 27. (canceled) 28. The method of claim 1, wherein the subject activates the device to obtain the sample using an activator. 29-38. (canceled) 39. A device for obtaining fluid from a subject, comprising: a fluid transporter;a fluid storage reservoir; andan indicator of one or more conditions associated with the introduction of fluid into the storage reservoir and/or one or more conditions associated with the storage of fluid in the storage reservoir, wherein the indicator is activated upon the accessing of fluid by the fluid transporter and/or the introduction of fluid into the storage reservoir. 40-77. (canceled) 78. A method, comprising: providing a non-healthcare-professional person with a fluid access device;directing the non-healthcare-professional person to apply the fluid access device to the skin of a subject to withdraw fluid from the skin and/or from beneath the skin of the subject into the device; anddirecting the non-healthcare-professional person to cause transport at least a portion of the device containing the withdrawn fluid to a separate location for analysis. 79. (canceled) 80. The method of claim 78, wherein the fluid comprises blood. 81. The method of claim 78, wherein the non-healthcare-professional person is the subject. 82. The method of claim 78, comprisin... | Seven Sense Biosystems, Inc. | A61B005/15 G09B019/00
US20110125208 | A1 | 2011-05-26 | METHODS AND SYSTEMS TO MONITOR CARDIAC CONTRACTILITY | 2009-12-30 | An implanted sensor produces a signal that is indicative of changes in arterial blood volume, such as a photoplethysmography signal or an impedance plethysmography signal. A metric is determined from the signal for each of the plurality of periods. Changes in cardiac contractility are monitored based on changes in the determined metric. | 1. For use with an implanted system, a method for monitoring cardiac contractility, comprising: (a) for each of a plurality of periods, using an implanted sensor to produce a signal that is indicative of changes in arterial blood volume;(b) determining a metric from the signal for each of the plurality of periods; and(c) monitoring changes in cardiac contractility based on changes in the determined metric. 2. The method of claim 1, wherein: step (a) comprises using an implanted photoplethysmography (PPG) sensor to produce a photoplethysmography signal that is indicative of changes in arterial blood volume. 3. The method of claim 1, wherein: step (a) comprises using an implanted impedance plethysmography sensor to produce an impedance plethysmography (IPG) signal that is indicative of changes in arterial blood volume. 4. The method of claim 1, wherein: the metric determined at step (b) is indicative of an upward slope of the signal; anddetecting an increase in the metric at step (c) is indicative of an increase in the cardiac contractility; anddetecting a decrease in the metric of step (c) is indicative of a decrease of the cardiac contractility. 5. The method of claim 1, wherein: the metric determined at step (b) is indicative of a time between two features of a cycle of the signal, the two features occurring between and inclusive of a foot indicative of beginning of systole and a dicrotic notch indicating closing of the aortic valve during the cycle of the signal;detecting a decrease in the metric at step (c) is indicative of an increase in the cardiac contractility; anddetecting an increase in the metric at step (c) is indicative of a decrease in the cardiac contractility. 6. The method of claim 1, wherein: the metric determined at step (b) is indicative of an amplitude of the signal where a dicrotic notch of the signal occurs;detecting an increase in the metric at step (c) is indicative of an increase in the cardiac contractility; anddetecting a decrease in the metric at step (c) is indicative of a decrease of cardiac contractility. 7. The method of claim 1, wherein: the metric determined at step (b) is determined from a derivative of the signal. 8. The method of claim 1, further comprising: (d) adjusting at least one pacing parameter based on the monitored cardiac contractility. 9. The method of claim 8, wherein: step (d) comprises adjusting the at least one pacing parameter to modify the monitored cardiac contractility be closer to a desired cardiac contractility. 10. The method of claim 1, further comprising: (d) adjusting at least one pacing parameter and determining a change in the monitored cardiac contractility; and(e) using the change in monitored cardiac contractility to determine a presence of dyssynchrony or a change in dyssynchrony. 11. The method of claim 1, further comprising: (d) providing therapy based on the determined cardiac contractility. 12. An implantable system for monitoring cardiac contractility comprising: ... | - | A61N001/362 A61B005/02
US20110118586 | A1 | 2011-05-19 | MEDICAL IMAGING APPARATUS CONTROLLER AND METHOD THAT ACCOMMODATE DIFFERENT CONTROL PROGRAM VERSIONS | 2010-11-18 | In a method, a system and a computer program product to control a medical imaging apparatus or to extend a control method for such an apparatus, the divided into different segments. A physical segment includes all control modules that are dependent on the platform of the medical apparatus. A logical segment includes all control modules that are independent of a platform. The logical segment of the control method is swapped out to an additionally provided auxiliary control computer while the physical segment remains on the (previous) control computer. At the previous control computer a data transfer module is provided that receives data of the auxiliary control computer and relays it to the medical apparatus for control. The auxiliary control computer advantageously possesses standardized interfaces. | 1. A control method for operating a medical imaging apparatus comprising the steps of: using a control platform to execute a control procedure to control a medical imaging apparatus via a plurality of control modules;among said plurality of control modules, designating platform-dependent control modules that are implemented by a control computer of the medical imaging apparatus and designating platform-independent control modules that are implemented by an auxiliary control computer of the medical imaging apparatus;placing the auxiliary control computer in data exchange communication with the control computer via a standardized interface; andin said control computer, providing a data transfer module and, via said data transfer module, receiving control data from the auxiliary control computer via said standardized interface and relaying said control data from the auxiliary control computer to the medical imaging apparatus via the control computer for, at least in part, controlling operation of said medical imaging apparatus dependent on the control data from the auxiliary control computer. 2. A method as claimed in claim 1 comprising controlling said medical imaging apparatus with said platform via at least one driver that is implemented at the control computer and interfacing said driver with said control computer via a driver platform. 3. A method as claimed in claim 2 wherein said medical imaging apparatus comprises a hardware component and an imaging component and comprising operating at least one of said hardware component and said imaging apparatus via said driver. 4. A method as claimed in claim 2 comprising embodying said driver in an operating system of the medical imaging apparatus. 5. A method as claimed in claim 1 comprising loading a modified version of said platform into the auxiliary control computer for use in the auxiliary control computer with no physical modifications to said control computer. 6. A method as claimed in claim 1 comprising configuring a plurality of different versions of said control platform for use with different console computers having different platforms and being configured for operating different types of imaging apparatuses. 7. A method as claimed in claim 1 comprising operating said control procedure through a console computer, and placing said console computer in a further data exchange communication with said auxiliary control computer via a further standardized interface. 8. A method as claimed in claim 7 comprising also placing said console computer in data exchange communication with said control computer. 9. A method as claimed in claim 7 comprising also placing said control computer in direct data exchange communication with said medical imaging apparatus. 10. A method as claimed in claim 1 comprising operating said control computer and said auxiliary control computer with a real-time operating system. 11. A method as claimed in claim 1 wherein said imaging apparatus comprises hardware c... | - | A61B005/055
US20110144461 | A1 | 2011-06-16 | PULSE FREQUENCY MEASURING METHOD AND APPARATUS | 2010-12-14 | A method and an apparatus are provided for measuring a pulse frequency in a bio-signal measurement device. A bio-signal collected by a sensor is applied as an input signal of a notch filter. A filter coefficient of the notch filter is adaptively changed according to a result of tracking the bio-signal in the notch filter and calculating a pulse frequency corresponding to the filter coefficient of the notch filter. | 1. A method for measuring a pulse frequency in a bio-signal measurement device, the method comprising the steps of: applying a bio-signal collected by a sensor as an input signal of a notch filter; andadaptively changing a filter coefficient of the notch filter according to a result of tracking the bio-signal in the notch filter and calculating a pulse frequency corresponding to the filter coefficient of the notch filter. 2. The method of claim 1, wherein calculating a pulse frequency comprises: determining a degeneration period of the bio-signal by using the input signal and an output signal of the notch filter;determining an impulse noise period in which an impulse noise signal is introduced, by measuring a power of the bio-signal tracked by the notch filter;delaying a tracking speed of the notch filter in the degeneration period and the impulse noise period; andcalculating the pulse frequency corresponding to the filter coefficient of the notch filter. 3. The method of claim 2, wherein determining a degeneration period of the bio-signal comprises determining a corresponding period as a degeneration period when a difference between a power of the input signal and a power of the output signal is less than a predetermined power reference value and continuously maintaining a filter coefficient from before the degeneration period during the degeneration period. 4. The method of claim 3, wherein the filter coefficient of the notch filter detected in a normal period of the bio-signal tracked by the notch filter is stored and updated during a predetermined period. 5. The method of claim 2, wherein the impulse noise period is determined by comparing two signals that are envelope-estimated by applying different attack time constants to the power of the bio-signal tracked by the notch filter. 6. The method of claim 5, wherein determining an impulse noise period comprises: detecting a first envelope envelope-estimated by applying a first attack time constant to the power of the bio-signal tracked by the notch filter;detecting a second envelope envelope-estimated by applying a second attack time constant, which is slower than the first attack time constant, to the power of the bio-signal tracked by the notch filter; anddetermining the impulse noise period, when a ratio of the second envelope to the first envelope is less than an envelope reference value. 7. The method of claim 1, wherein, before the bio-signal collected by the sensor is applied to the notch filter, motion artifacts are removed from the bio-signal by a High Pass Filter (HPF). 8. The method of claims 1, wherein the bio-signal is one of an ElectroCardioGram (ECG) and a PhotoPlethysmoGraphy (PPG). 9. An apparatus for measuring a pulse frequency in a bio-signal measurement system, the apparatus comprising: a bio-signal processor for adaptively changing a filter coefficient of a notch filter according to a result of tracking a bio-signal in the notch filter when the bio-signal collected ... | Samsung Electronics Co., Ltd. Industry-Academic Cooperation Foundation Yonsei University | A61B005/0402 A61B005/1455 G06F019/00
US20110125004 | A1 | 2011-05-26 | Analysis by photo acoustic displacement and interferometryl | 2009-07-24 | A device for analyzing a material, the device having a photo acoustic generator (40) for generating pressure waves in the material by pulsed illumination by a light source, and a sensor (20, 30, 50) for producing a signal using self mixing interferometry, the signal representing displacements of the material due to the pressure waves, and a signal processor (60) for processing the signal to analyze the material. This combination of photo acoustic stimulation and self mixing interferometry sensing helps reduce or avoid errors or distortions caused by sensor contact on the material, and can provide a better ratio of sensitivity for a given degree of compactness of the device. | 1. A device for analyzing a material, the device having a photo acoustic generator (40) for generating pressure waves in the material by pulsed illumination by a light source, and a sensor (30, 50) for producing a signal using self mixing interferometry, of displacements of the material due to the pressure waves, and a signal processor (60) for processing the signal to analyze the material. 2. The device of claim 1, the sensor comprising a laser diode (30) for outputting light which is directed towards the material so that a proportion is reflected so as to mix in a cavity (20) of the laser diode to form a modulation, the sensor also having an analyzer (50) for generating a signal representing the displacements by analyzing the modulation. 3. The device of claim 1 having an opto mechanical arrangement for scanning the sensor over an area of the material. 4. The device of claim 1, having an array of sensors for simultaneously generating signals for an area of the material. 5. The device of claim 1 having a near infra red spectroscope sensor for generating a further measurement of the material. 6. The device of claim 1, the sensor being arranged to sense parts of the material below a surface of the material. 7. The device of claim 1, the sensor being arranged to sense parts of the material on a surface of the material. 8. The device of claim 1, arranged as a non invasive monitor for monitoring concentration of constituents of blood in a human or animal body. 9. A method of analyzing a material, having the steps of photo acoustically generating pressure waves in the material by pulsed illumination by a light source, producing a signal using self mixing interferometry, of displacements of the material due to the pressure waves, and processing the signal to analyze the material. 10. The method of claim 9, having the step of generating an image of the material. 11. The method of claim 9 for non invasive monitoring of blood in a human or animal body, having the step of processing the signal to determine the concentration of given constituents of the blood. 12. The method of claim 9 having the steps of processing the signal to provide alignment information, and aligning the device on the material on the basis of the information before processing the signal to analyze the material. 13. A controller for controlling a device for analyzing a material, the controller having means for controlling photo acoustic generation of pressure waves in the material by pulsed illumination by a light source, and means for processing a signal to analyze the material, the signal being of displacements of the material due to the pressure waves obtained from self mixing interferometry. 14. The controller of claim 13 for use with a device for non invasive monitoring of blood in a human or animal body, the means for processing the signal outputting a value related the concentration of given constituents of the blood. 15. The controller, claim 13 having means for p... | KONINKLIKJKE PHILIPS ELECTRONICS N.V. | A61B005/05
US20110136076 | A1 | 2011-06-09 | IDENTIFICATION DEVICE | 2010-11-09 | The invention provides an identification device, which comprises a dental implantation carrier and an information storage element therein. The identification device has good privacy and bio-compatibility. The invention also provides a method for identification by using the device. | 1. An identification device, comprising: a dental implantation carrier; andan information storage element positioned in said implantation carrier. 2. The identification device according to claim 1, wherein said implantation carrier is a dental crown, a dental post, a dental implant or a bone screw. 3. The identification device according to claim 1, wherein said information comprises user identification, user personal information or a combination thereof. 4. The identification device according to claim 3, wherein said information further comprises a remaining sum of electronic money. 5. The identification device according to claim 1, wherein said information storage element is a chip, a flash memory, a metal-etched code or a combination thereof. 6. The identification device according to claim 5, wherein said chip is an identification chip. 7. The identification device according to claim 5, wherein said chip is a contactless chip, a contact chip or a combination thereof. 8. The identification device according to claim 1, wherein said information storage element is coated by a metal prosthesis. 9. The identification device according to claim 8, wherein a metal of said metal prosthesis is titanium. 10. A method for identification, comprising the following steps: (a) positioning at least one identification device in accordance with claim 1 within an oral cavity of a subject to be identified; and(b) reading an information storage element of said identification device. 11. The method according to claim 10, wherein said identification device is positioned within a tooth or a bone inside said oral cavity. 12. The method according to claim 10, further comprising a step of removing said identification device from said oral cavity before said step (b). 13. The method according to claim 10, wherein a type of said reading is contactless-reading or contact-reading. 14. The method according to claim 10, wherein said subject comprises human beings or animals. | Yu-Jung Li Sheng-Yi Chen | A61C019/00 A61B005/117
US08554300 | B2 | 2013-10-08 | MRI gradient field detector | 2010-12-02 | A device and a method are disclosed for detecting electromagnetic fields, in particular, fields occurring in magnetic resonance tomography (MRT) or magnetic resonance imaging (MRI) tests. An implantable medical device (IMD), contained in a hermetically sealed housing, includes a control unit, a programming coil, and a communication unit, wherein the communication unit, together with the programming coil, is designed to allow communication between an external programming device and the IMD by utilizing alternating electromagnetic fields, and further comprising a detection unit for MRT interference fields, characterized in that the detection unit is designed in such a way that voltage profiles induced in the programming coil and originating from a pulsed alternating electromagnetic field of the MRT (gradient field) are detected, and a corresponding MRT detection signal is transmitted from the detection unit to the control unit, if communication with a programming device is not detected at the same time. | 1. An implantable medical device (IMD) contained in a hermetically sealed housing, comprising: a programming coil enclosed by said hermetically sealed housing;a communication unit coupled with the programming coil, wherein the communication unit is configured to enable communication between an external programming device and the IMD by utilizing alternating electromagnetic fields;a detection unit configured to sense MRT interference fields, wherein voltage pulses induced in the programming coil and originating from a pulsed alternating electromagnetic field of the MRT are detected; anda control unit that configured to receive an MRT detection signal from the detection unit, wherein communication with the external programming device via the programming coil is not detected at the same time;wherein the IMD is connected to at least one electrode which is connected to the detection unit, andwherein the MRT detection signal is transmitted only when the sensed MRT interference fields is simultaneously detected via the programming coil enclosed by said hermetically sealed housing and via the at least one electrode. 2. The IMD according to claim 1, wherein the detection unit identifies the pulsed alternating electromagnetic fields by detecting a threshold value and counting the induced voltage pulses per unit time. 3. The IMD according to claim 1, wherein the MRT interference detection is based on evaluation and comparison of characteristics of the induced voltage pulses with voltage pulse characteristics typically associated with MRT. 4. The IMD according to claim 3, wherein said voltage pulse characteristics include frequencies. 5. The IMD according to claim 1, wherein the MRT detection signal is configured to cause a change in an operating state of the IMD. 6. The IMD according to claim 1, wherein the detection of the pulsed alternating electromagnetic fields is accomplished using a programmable filter and a trigger unit connected to the programming coil. 7. The IMD according to claim 1, wherein an MRT state may be set either permanently, until a next reprogramming, or for a predetermined or settable first time period, and VF detection is extended by a predetermined second time period in order to ensure detection of the pulsed alternating electromagnetic fields, even for temporary zero crossings of gradient fields in the plane of the implant. 8. The IMD according to claim 7, wherein the extension of the VF detection by the second time period is up to 30 s, preferably up to 10 s. 9. The IMD according to claim 1, wherein the detection unit is connected to an indicator of MRT interference fields and an MRT sensor configured to sense MRT interference fields, wherein the MRT detection occurs in response to detection by the sensor. 10. The IMD according to claim 1, wherein at least one of the following measures is introduced in response to MRT detection: changing to an MRI-safe state;remaining for a prolonged period of time in an MRI-safe state tha... | Biotronik CRM Patent AG | A61B005/055
US08255042 | B2 | 2012-08-28 | Monitoring physiological condition and detecting abnormalities | 2010-11-30 | A system and method for monitoring an individual's physiological condition and detecting abnormalities therein, comprising concurrently receiving an electrocardiograph signal and a ballistocardiograph signal. The electrocardiograph and ballistocardiograph signals are conditioned to minimize background extraneous noise after which, each signal is concurrently processed and analyzed to detect repeating cyclical patterns and further characterized to identify individual components of the repeating cycles. At least one individual component in one signal is selected as a reference marker or a selected component in the other signal. The two signals are then synchronized, outputs produced therefrom and stored in a database. | 1. A system for monitoring an individual's physiological condition comprising: a first device for coupling to said individual, said first device for detecting and transmitting an electrocardiograph signal of said individual;a second device for coupling to said individual, said second device for detecting and transmitting a ballistocardiograph signal of said individual;a computer including a microprocessor in communication with said first device and said second device, said computer for receiving said electrocardiograph signal and said ballistocardiograph signal, converting said electrocardiograph signal into electrocardiograph data, converting said ballistocardiograph signal into ballistocardiograph data and synchronizing said electrocardiograph data with said ballistocardiograph data, said computer generating an output indicative of said individual's physiological condition. 2. The system according to claim 1, wherein said electrocardiograph signal and said ballistocardiograph signal are converted by an analog-digital converter. 3. The system according to claim 1, wherein said first device comprises an electrocardiograph. 4. The system according to claim 1, wherein said second device comprises an accelerometer. 5. The system according to claim 4, wherein the accelerometer is configured to be positioned on the sternum of said individual. 6. The system according to claim 1, wherein said computer includes a database in communication with said microprocessor, said database for storing said electrocardiograph data and said ballistocardiograph data. 7. The system according to claim 6, wherein said output comprises a comparison of said electrocardiograph data and said ballistocardiograph data of said individual with reference electrocardiograph and ballistocardiograph wave pattern data, said reference electrocardiograph and ballistocardiograph wave pattern data being stored in said database. 8. The system according to claim 6, wherein said computer includes a graphical user interface in communication with said microprocessor, said graphical user interface for allowing an operator to input data into said database. 9. The system according to claim 1, wherein synchronizing said electrocardiograph data with said ballistocardiograph data includes aligning an R peak of an electrocardiograph waveform corresponding to said electrocardiograph data with an H peak of a ballistocardiograph waveform corresponding to said ballistocardiograph data. 10. The system according to claim 1, wherein said output is provided on a device selected from the group consisting of monitors, screens, printers and plotters. 11. The system according to claim 1, wherein said output comprises a synchronized electrocardiograph-ballistocardiograph waveform. 12. The system according to claim 1, wherein said output comprises a comparison of at least one resting-stage electrocardiograph-ballistocardiograph signal set and at least one post-exercise electrocardiograph-ballistocardiog... | Heart Force Medical Inc. | A61B005/02
US08221352 | B2 | 2012-07-17 | Disposable safety needle system and safety enclosure means | 2010-01-29 | Techniques disclosed herein include a disposable safety needle assembly and safety enclosure for internally shielding a used or contaminated needle. The needle assembly has a needle cannula slidably disposed within a safety enclosure. The safety enclosure includes an internally disposed flexible safety member. The flexible safety member has face portion transversely situated and having a safety feature disposed to the needle and that slides thereon. Retracting the needle within the safety enclosure to a point wherein the sharp tip of the needle is beyond the needle protection section of the flexible member causes the protection section to travel transversely to the direction of motion of the needle. As a result, the safety feature receives the sharp tip of the needle and prevents the needle from exiting the enclosure, thereby maintaining the needle safely after use. | 1. An apparatus for securing and protecting the tip of a needle, the apparatus comprising: a needle having a body portion and a diameter and a tip of the needle;a safety enclosure having an interior space, a first base, a pair of second bases, and a longitudinal axis, the needle being slidably disposed within said interior space of said safety enclosure;a base portion;a flexible member having a length, said base portion being flexibly connected to said safety enclosure by said flexible member;a flexible safety member having a face portion and rear portion, said flexible safety member being disposed substantially longitudinally within said interior space, said face portion being disposed substantially transversely within said interior space;said face portion including a safety feature sized and dimensioned to receive the tip of the needle;said safety enclosure having two positions, a use position and a safe position;said use position in which the tip of the needle extends beyond the interior space of said safety enclosure and said face portion is prevented from transverse movement by the needle;said safe position in which said safety enclosure is moved such that the tip of the needle is brought into said interior space of said safety enclosure to a safe position where the tip of the needle is brought further into the interior space than said face portion of said flexible safety member, wherein said face portion is able to move transversely and to retain the tip of the needle, wherein the tip of the needle is received by the safety feature;wherein the safety feature comprises an orifice in said face portion of said flexible safety member that is sized and dimensioned to receive the tip of the needle;wherein said orifice of said safety feature has a diameter that is less than the diameter of the needle body; andwherein the flexible safety member is constructed of material that has sufficient strength to prevent the tip of the needle from further expanding said diameter of said orifice of said safety feature. 2. The apparatus of claim 1, further comprising: a latch that removably couples the safety enclosure to the base portion. 3. The apparatus of claim 1 wherein the safety enclosure is substantially a rectangular parallelepiped in shape, and wherein said rectangular parallelepiped is a rectangular cuboid. 4. The apparatus of claim 1, wherein said safety enclosure is substantially cylindrical in shape. 5. The apparatus of claim 1 wherein said safety enclosure is constructed of a material having sufficient strength to resist being punctured by the tip of the needle. 6. The apparatus of claim 1 further including a piece of hollow tubing having an outer surface, said hollow tubing being mechanically attached to the needle to be in fluid communication with the tip of the needle, said base portion being fixedly disposed about the outer surface of the hollow tubing, and wherein said hollow tubing is flexible. 7. The apparatus of claim 1 further incl... |  | A61M005/00
US07942822 | B1 | 2011-05-17 | Implantable medical device with sleep apnea detection control and method | 2006-07-31 | A method for use in an implantable medical device comprises the steps of monitoring respiration with an amplifier having a gain, generating a moving apneic threshold based on recent respiration cycles, accumulating differences between amplitudes of respiration cycles and the moving apnea detection threshold and comparing the accumulated differences against an apnea detection threshold to detect the onset of an episode of apnea. The method further comprises measuring respiration levels upon detecting the onset of apnea, confirming the episode of apnea based upon the respiration levels measured upon detecting the onset of apnea; and adjusting one of the gain of the amplifier and the apnea detection threshold so that the time from the detection of onset of apnea to the time of confirmation of the episode of apnea is within a predetermined time range following the detection of the onset of apnea. | 1. A method for use in an implantable medical device, the method comprising; monitoring respiration with a variable gain amplifier;generating a moving apneic threshold based on recent respiration cycles;accumulating differences between amplitudes of respiration cycles and the moving apnea detection threshold;detecting the onset of an episode of apnea based upon the accumulated differences;measuring respiration levels upon detecting the onset of apnea;confirming the episode of apnea based upon the respiration levels measured upon detecting the onset of apnea; andadjusting the gain of the variable gain amplifier so that the time from the detection of onset of apnea to the time of confirmation of the episode of apnea is within a predetermined time range. 2. The method of claim 1, wherein the predetermined time range is between ten seconds and twenty seconds. 3. The method of claim 1, wherein detecting the onset of an episode of apnea based upon the accumulated differences is performed by comparing the accumulated differences against a fixed apnea detection threshold. 4. The method of claim 1, wherein measuring respiration levels is performed by making respiration measurements at preset time intervals upon and after detecting the onset of apnea. 5. The method of claim 4, wherein confirming the episode of apnea is based upon making a set number of consecutive respiration measurements of apnea respiration levels within the predetermined time range. 6. The method of claim 1, further comprising delivering apnea therapy upon confirmation of the onset of an episode of apnea. 7. A method for use in an implantable medical device, the method comprising: monitoring respiration with an amplifier having a gain;generating a moving apneic threshold based on recent respiration cycles;accumulating differences between amplitudes of respiration cycles and the moving apnea detection threshold;detecting the onset of an episode of apnea based upon the accumulated differences and an apnea detection threshold;measuring respiration levels upon detecting the onset of apnea;confirming the episode of apnea based upon the respiration levels measured upon detecting the onset of apnea; andadjusting one of the gain of the amplifier and the apnea detection threshold so that the time from the detection of onset of apnea to the time of confirmation of the episode of apnea is within a predetermined time range. 8. The method of claim 7, further comprising delivering apnea therapy upon confirmation of the onset of an episode of apnea. 9. The method of claim 7, wherein the predetermined time range is between ten seconds and twenty seconds. 10. The method of claim 7, wherein adjusting comprises adjusting the gain of the amplifier and wherein detecting the onset of an episode of apnea comprises comparing the accumulated differences against a fixed apnea detection threshold. 11. The method of claim 7 wherein measuring respiration levels is performed by making respiration measureme... | Pacesetter, Inc. | A61B005/08
US20110105862 | A1 | 2011-05-05 | DEVICE FOR ACQUIRING AND PROCESSING PHYSIOLOGICAL DATA OF AN ANIMAL OR OF A HUMAN IN THE COURSE OF A PHYSICAL OR MENTAL ACTIVITY | 2009-04-28 | The invention relates to a system allowing the acquisition and processing of data representative of the physical or mental activity and/or of the physiological state of human or animal individuals. This system includes, for each individual, a unique individual electronic box encasing several sensors capable of measuring physical and/or biological quantities related to the physical and/or biological activity of the wearer of the box and of returning information. Each electronic box moreover is provided with an interface for radio communication with a device allowing the management of the data gathered from the individual boxes and management. | 1. A system for the acquisition and processing of data representing the physical or mental activity and/or the physiological state of a plurality of human or animal individuals, comprising, for each individual, a single individual electronic box containing several sensors able to measure physical and/or biological values relating to the physical activity of the carrier of the box, each electronic box also being provided with means of radio communication with a remote management device for managing the data collected from the individual boxes. 2. A system according to claim 1, wherein each individual electronic box contains movement sensors able to deliver in real time a signal representing the spatial movements of the individual, and at least one physiological sensor able to deliver a signal representing at least one physiological value of the individual, as well as a computer receiving the sensor signals as an input and delivering synchronised and digitised signals as an output, and a memory for storing said synchronised and digitised signals. 3. A system according to claim 1, wherein the communication between the individual electronic boxes and the remote management device is bidirectional and also makes it possible to send information or instructions to each individual electronic box. 4. A system according to claim 3, wherein the information and/or instructions received are reproduced by the individual electronic box in the form of voice messages or light, audible or electrical signals. 5. A system according to claim 3, wherein the remote management device is provided with means for adjusting in real time the bandwidth of the radio communication with each individual electronic box. 6. A system according to claim 3, wherein the radio communication between the individual electronic boxes and the remote management device takes place by means of a mobile telephony communication network. 7. A system according to claim 3, wherein the radio communication between the individual electronic boxes and the remote management device takes place by means of a dedicated self-contained communication network. 8. A system according to claim 7, further comprising positioning beacons distributed around the area of activity of the individuals, with respect to which each electronic box determines its instantaneous position and transmits it to the remote management device. 9. A system according to claim 8, wherein the positioning beacons comprise a function of communication relay between the individual electronic boxes and the remote management device, so as to ensure continuity of the radio communication between the individual electronic boxes and the remote management device, when the latter is beyond the direct range of communication of the individual electronic boxes. 10. A system according to claim 7, wherein the individual electronic boxes and the remote management device are configured so as to function as an auto-routing and auto-adaptive radio commu... | UNIVERSITE DU SUD TOULON-VAR | A61B005/00
US07970174 | B2 | 2011-06-28 | Medical marker tracking with marker property determination | 2007-06-25 | A method for tracking at least one medical marker is provided, wherein actual properties of the at least one marker are compared with nominal properties of the at least one marker. A basis for subsequent use of information obtained from the at least one marker is formed based on the comparison. | 1. A method for tracking at least one medical marker, comprising: comparing, using a processor, actual properties of the at least one marker with nominal properties of the at least one marker; andforming a basis for subsequent use of information obtained from the at least one marker from the comparison,wherein comparing includes determining there is a deviation in quality of the at least one marker if the actual properties of the at least one marker do not correspond to the nominal properties of the at least one marker. 2. The method according to claim 1, further comprising assigning a unique identifier to the at least one marker. 3. The method according to claim 1, further comprising storing base properties of the at least one marker as nominal properties of the at least one marker. 4. The method according to claim 1, further comprising continuously or periodically storing actual properties of the at least one marker as nominal properties of the at least one marker. 5. The method according to claim 4, further comprising concluding properties of the at least one marker have changed when newly acquired actual properties of the at least one marker deviate from the nominal properties of the at least one marker. 6. The method according to claim 1, further comprising moving the at least one marker into a detection range of an image detection unit in order to determine properties of the at least one marker. 7. The method according to claim 1, wherein determining includes ascertaining there is a deviation in marker quality if the correspondence between the actual properties and the nominal properties deviates by a threshold value. 8. The method according to claim 7, wherein the threshold value is pre-defined or user settable. 9. The method according to claim 7, wherein the at least one marker is a plurality of markers, further comprising setting the threshold value based on a measuring procedure of the plurality of markers and/or a comparison of different markers of the plurality of markers. 10. The method according to claim 1, wherein the at least one marker is a plurality of markers, and once it is determined that a marker of the plurality of markers deviates in quality, then performing at least one of: indicating the deviation in marker quality;removing any marker that deviates in quality from tracking calculations;requesting replacement of the marker that deviates in quality; orarithmetically correcting positional details defined by the marker that deviates in quality. 11. The method according to claim 10, wherein arithmetically correcting includes recalculating the image center of the deviating marker. 12. The method according to claim 11, wherein recalculating the image center includes using a part of the deviating marker's outer contour that matches the actual and/or nominal properties of the deviating marker, and disregarding parts deviating marker's outer contour that do not match the actual and/or nominal properties. 13. A method f... | Brainlab AG | G06K009/00 A61B005/05
US07937132 | B2 | 2011-05-03 | Hybrid imaging method to monitor medical device delivery and patient support for use in method | 2008-02-14 | This invention discloses a method and apparatus to deliver medical devices to targeted locations within human tissues using imaging data. The method enables the target location to be obtained from one imaging system, followed by the use of a second imaging system to verify the final position of the device. In particular, the invention discloses a method based on the initial identification of tissue targets using MR imaging, followed by the use of ultrasound imaging to verify and monitor accurate needle positioning. The invention can be used for acquiring biopsy samples to determine the grade and stage of cancer in various tissues including the brain, breast, abdomen, spine, liver, and kidney. The method is also useful for delivery of markers to a specific site to facilitate surgical removal of diseased tissue, or for the targeted delivery of applicators that destroy diseased tissues in-situ. | 1. An apparatus for substantially immobilizing an anatomy of interest of a patient during MRI imaging of the anatomy of interest using a MRI scanner having a bore and during ultrasound imaging of the anatomy of interest, the apparatus comprising: a patient support structure including a superior section and an inferior section, the inferior section being offset horizontally a distance from the superior section to define an opening therebetween, the inferior section connected to the superior section by a connection member lateral to the opening;a transporter having a base and a pedestal, the pedestal supporting the patient support structure and the patient support structure horizontally translatable relative to the transporter between an extended and retracted position to allow the patient support structure to be positioned into the bore of the MRI scanner in the extended position, at least a portion of the opening extending beyond the pedestal in a horizontal direction providing medial, lateral and anterior access to the anatomy of interest in the retracted position;an immobilization device coupled to the patient support structure, the immobilization device having a membrane support frame defining an open area and a deformable membrane configured to constrain the anatomy of interest within the opening in a selected position and conformation during both of a MRI imaging exam and an ultrasound imaging exam; andthe deformable membrane being coupled to the membrane support frame and covering at least a portion of the open area, and the deformable membrane constructed from a non-magnetic and substantially ultrasound-transparent material, wherein a portion of the first deformable membrane is adjacent to the anatomy of interest when the anatomy of interest is constrained by the immobilization device. 2. The apparatus of claim 1, wherein the deformable membrane is of a thickness selected to allow puncturing or cutting of the deformable membrane by a standard medical instrument. 3. The apparatus of claim 1, further comprising a plurality of fiducial markers coupled to the immobilization device, wherein the fiducial markers are constructed from a material visible on images reconstructed from data acquired during the MRI imaging. 4. The apparatus of claim 1, wherein the immobilization device further comprises a second membrane support frame defining an open area and a second deformable membrane, wherein the membrane support frames are translatably positionable on opposite sides of a breast, and are configured to compress the breast by decreasing the distance between the membrane support frames, the second deformable membrane being coupled to the second membrane support frame and covering at least a portion of the open area and constructed from a non-magnetic and substantially ultrasound-transparent material. 5. The apparatus of claim 1, further comprising: a mechanical stage coupled to the patient support structure; anda holder coupled to the mechanical ... | Sunnybrook Health Sciences Center | A61B005/05
US07951087 | B2 | 2011-05-31 | Method for correction of posture dependence on heart sounds | 2009-12-23 | A system to monitor heart sounds. The system comprises an implantable heart sound sensor configured to produce an electrical signal representative of at least one heart sound, an implantable posture sensor operable to produce an electrical signal representative of a patient's posture, and a controller circuit. The controller circuit is configured to determine a patient posture, measure at least one heart sound in correspondence with at least one corresponding patient posture, adjust the heart sound measurement by using the corresponding determined patient posture to reduce or remove variation in the heart sound measurement due to patient posture, detect a change in the adjusted heart sound measurement, and provide an indication of congestive heart failure to a user or an automated process in response to the detected change. | 1. An apparatus comprising: an implantable heart sound sensor configured to produce an electrical signal representative of at least one heart sound, the heart sound associated with mechanical activity of a patient's heart;an implantable posture sensor operable to produce an electrical signal representative of a patient's posture; anda controller circuit, communicatively coupled to the heart sound sensor interface circuit and the posture circuit, wherein the controller circuit is configured to: determine a patient posture using the posture sensor signal;measure the at least one heart sound in correspondence with at least one corresponding determined patient posture;adjust the heart sound measurement by using the corresponding determined patient posture to reduce or remove variation in the heart sound measurement due to patient posture;detect a change in the adjusted heart sound measurement; andprovide an indication of congestive heart failure (CHF) to a user or an automated process in response to the detected change. 2. The apparatus of claim 1, wherein the implantable heart sound sensor is configured to produce an electrical signal representative of an S3 heart sound, andwherein the controller circuit is configured to: measure the S3 heart sound in correspondence with the determined patient posture;detect a change in an adjusted S3 heart sound measurement; andgenerate an alarm indicative of progression of CHF in response to the detected change in the adjusted S3 heart sound measurement. 3. The apparatus of claim 1, including: a therapy circuit communicatively coupled to the controller circuit and configured to provide electrical pacing therapy to the patient, andwherein the controller circuit is configured to initiate delivery of pacing resynchronization therapy according to the detected change in the adjusted heart sound measurement. 4. The apparatus of claim 1, including: a trans-thoracic impedance sensor communicatively coupled to the controller circuit and configured to provide an electrical signal representative of trans-thoracic impedance of the patient, andwherein the controller circuit is configured to: measure trans-thoracic impedance of the patient in correspondence with the determined patient posture; andprovide the indication of CHF according to the adjusted heart sound measurement and the measured trans-thoracic impedance. 5. The apparatus of claim 4, wherein the trans-thoracic impedance sensor is configured to provide an electrical signal representative of near-DC trans-thoracic impedance, andwherein the controller circuit is configured to provide the indication of CHF according to the adjusted heart sound measurement and the measured near-DC trans-thoracic impedance. 6. The apparatus of claim 1, including: an intra-cardiac impedance sensor communicatively coupled to the controller circuit and configured to provide an electrical signal representative of trans-thoracic impedance of the patient, andwherein the controller circuit ... | Cardiac Pacemakers, Inc. | A61B005/02
US09020585 | B2 | 2015-04-28 | Electronic identity card | 2008-06-17 | A method comprises the steps of obtaining a baseline measurement corresponding to a representation of brain activity from a living body and storing the baseline measurements on a storage device configured to be wearable on the living body in combination with obtaining, at a time subsequent to the baseline measurement, a second measurement representing brain activity from the living body, reading the baseline measurement from the storage device and evaluating the living body based on a comparison of the second measurement to the baseline measurement. A system for analyzing electrical brain activity, the system comprises a component for reading, from a storage medium, data corresponding to a baseline measurement of brain activity of a living body and a component for determining data representing a current brain activity of the living body in combination with a component for evaluating the living body by comparing the baseline measurement data and the current brain activity data. | 1. A method comprising: receiving, using electrodes, a baseline measurement corresponding to a representation of brain activity from a living body;storing the baseline measurements on a storage device configured to be wearable on the living body;obtaining, using electrodes connected to a portable device, at a time subsequent to the baseline measurement, a second measurement representing brain activity from the living body;evaluating, via a processing unit of the portable device, the living body based on a comparison of the second measurement to the baseline measurement stored in the storage device, wherein the storage device is portable independently of the portable device;determining, via the processing unit of the portable device, whether the second measurement is indicative of a persistent change in the brain activity of the living body;replacing, via one of the portable device and a docking station, the baseline measurement with a current measurement if the second measurement is indicative of a persistent change in the brain activity of the living body; anddetermining, via the processing unit, if the living body is experiencing an impairment based on the comparison of the baseline data to the current brain activity data, wherein if an impairment is detected, the processing unit prevents operation of a machine operably associated with the system. 2. The method of claim 1, wherein the brain activity measurement is at least one of a qEEG, a BAER and an ERP. 3. The method of claim 1, further comprising filtering noise from the baseline measurement via a filtering arrangement. 4. The method of claim 1, further comprising amplifying the baseline measurement via an amplifier. 5. The method of claim 1, further comprising storing personal data about the living body on the portable storage device. 6. The method of claim 5, wherein the personal data includes age, sex, date, and time of measurement. 7. The method of claim 1, further comprising obtaining the measurements by using one of a montage including leads on a forehead of the living body and any selected array from the International 10-20 System with any selected reference. 8. The method of claim 1, wherein the baseline measurement is performed at the docking station. 9. The method according to claim 8, wherein the docking station is mobile. 10. The method of claim 8, wherein the baseline measurement is transmitted to a database via one of a wired and a wireless network, the docking station writing the baseline measurement from the database to the storage device. 11. The method according to claim 1, wherein the storage device includes one of a USB drive, a memory card, a memory stick, a memory chip, a CD-ROM, a floppy disk, a zip disk and a flash card. 12. The method according to claim 1, wherein the evaluating includes performing a substance abuse check on the living body based on the comparison. 13. The method according to claim 1, wherein the evaluating includes determining an extent... | New York University | A61B005/04 A61B005/117 A61B005/0476 A61B005/0484 A61B005/00
US08543180 | B2 | 2013-09-24 | Method and apparatus for total hemoglobin measurement | 2010-09-03 | The present invention pertains to a method and apparatus for total hemoglobin measurement. A modulated optical signal based on a digital code sequence is transmitted to human tissue. A temporal transfer characteristic is derived from the modulated optical signal. Total hemoglobin is determined based on the temporal transfer characteristic. | 1. A total hemoglobin measurement system for measuring total hemoglobin in human tissue comprising: a signal generator for generating a digital modulation signal representing a code sequence;an optical illumination source coupled to said signal generator for receiving said digital modulation signal and configured for transmitting a modulated optical signal along an optical transmission path to said human tissue in response to said digital modulation signal;a detector configured to be optically coupled to said human tissue for receiving said modulated optical signal after transmission through the optical transmission path through said human tissue;a processor coupled to said detector for deriving a temporal transfer characteristic for said optical signal received by said detector and for determining said total hemoglobin based on said temporal transfer characteristic;a second signal generator for generating a second digital modulation signal representing a second code sequence, wherein said second code sequence and said code sequence have the same ordering of code elements and begin at different code elements; anda second optical illumination source coupled to said second signal generator for receiving said second digital modulation signal and configured for transmitting a second modulated optical signal to said human tissue in response to said second digital modulation signal. 2. The total hemoglobin measurement system of claim 1 further comprising: a delay component optically coupled to said optical illumination source for increasing length of said optical transmission path. 3. The total hemoglobin measurement system of claim 2 wherein said delay component further comprises a reflective surface. 4. The total hemoglobin measurement system of claim 1 wherein said code sequence further comprises multiple identical code sequences of a single code sequence. 5. The total hemoglobin measurement system of claim 1 wherein said code sequence is circular. 6. The total hemoglobin measurement system of claim 1 wherein said code sequence is orthogonal. 7. The total hemoglobin measurement system of claim 1 wherein said code sequence is unipolar. 8. The total hemoglobin measurement system of claim 1 wherein said code sequence is a Galois code. 9. The total hemoglobin measurement system of claim 1 wherein said code sequence is a maximal-length sequence. 10. A method for measuring total hemoglobin in a human tissue comprising: generating a digital modulation signal associated with a code sequence;generating a modulated optical signal based on said digital modulation signal;transmitting said modulated optical signal to said human tissue;receiving a modified version of said modulated optical signal after transmission through said human tissue;deriving a temporal transfer characteristic for said modified version of said modulated optical signal;calculating said total hemoglobin based on said temporal transfer characteristic;delaying said code sequence to ge... | Covidien LP | A61B005/1455
US08799013 | B2 | 2014-08-05 | Mammography information system | 2010-11-23 | A method and system for analyzing and retrieving tissue abnormality tracking data, providing a tool for a radiologist that includes a report summarizing the statistical frequency of diagnosed patients, both locally and nationally, with tissue region-of-interest classifications similar to the tissue images taken of the anatomy of an individual patient. In one embodiment, a computer aided diagnostic program can suggest region-of-interest attributes and classifications for user review. In one embodiment, user can select a seed area for computer aided diagnostic program analysis, the CAD analysis can suggest region-of-interest attributes and classifications. In yet another embodiment, the region-of-interest abnormality attributes suggested by the computer aided diagnostic program can be stored and compared with the user selected abnormality attributes and classifications. In one embodiment the images are of human breast tissue. | 1. A configurable mammography diagnostic system, comprising: a plurality of electronic displays, at least one of the plurality of electronic displays configured to display a breast tissue image;a graphical user interface presented on at least one of the plurality of electronic displays and including: a breast tissue image on which at least one region of interest in the breast tissue image can be indicated,a plurality of possible characteristics from a predetermined medical lexicon according to which an automated categorization of a region of interest in the breast tissue image can be indicated, wherein each of the plurality of possible characteristics in the predetermined medical lexicon is associated with a unique alphanumeric value, anda profiler display configured to present statistical information related to a comparison of the automated categorization with a database of existing categorizations of a plurality of existing breast tissue images, each of the plurality of existing breast tissue images having been diagnosed;anda processing engine including a computer aided diagnosis (CAD) tool configured to identify the at least one region of interest in the breast tissue image, to provide the automated categorization of the region of interest, to present the automated categorization of the region of interest via the graphical user interface, and configured to dynamically update the profiler display based on a change in one or more of the plurality of possible characteristics and the unique alphanumeric value associated with the changed characteristic of the region of interest. 2. The system of claim 1, wherein the breast tissue image is selected from the group consisting of an X-ray image, a CT image, an MRI image, an ultrasound image, and a pathology image. 3. The system of claim 1, wherein each of the plurality of existing breast tissue images is cataloged in the database according to a characteristic of the breast tissue image. 4. The system of claim 1, wherein the plurality of possible characteristics from the predetermined medical lexicon includes an abnormality type, size, and profile. 5. The system of claim 1, wherein the computer aided diagnosis (CAD) tool is further configured to identify the plurality of possible characteristics from the predetermined medical lexicon of an abnormality depicted in the at least one region of interest. 6. The system of claim 1, wherein the computer aided diagnosis (CAD) tool is further configured to identify the at least one region of interest in the breast tissue image based on a user selection of a portion of the breast tissue image within the region of interest depicted in the breast tissue image. 7. The system of claim 1, wherein the computer aided diagnosis (CAD) tool is further configured to identify the at least one region of interest in the breast tissue image based on a user selection of a portion of the breast tissue image containing the region of interest depicted in the breast tissue imag... | PenRad Technologies, Inc. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00 A61B005/08 A61B005/06
US08657756 | B2 | 2014-02-25 | Implantable device employing movement sensing for detecting sleep-related disorders | 2011-02-15 | Methods and systems for evaluating a pathological condition include acquiring movement information, such as electromyogram (EMG) information, and sleep disordered breathing (SDB) information, and detecting the presence of a pathological condition using both movement and SDB information. Methods may involve sensing physiological signals including at least muscle movement signals. Sleep-related disorders are detected using the sensed physiological signals, the sleep-related disorders including at least an involuntary muscle movement disorder and sleep-disordered breathing. Methods and systems also provide for detecting and treating a sleep-related disorder using movement and SDB information. Cardiac, respiratory, nerve stimulation, drug, or a combination of such therapies may be delivered to treat a detected or diagnosed pathological condition. | 1. A system, comprising: a plurality of sensors configured to respectively sense a plurality of physiological signals including at least muscle movement signals comprising electromyogram signals;a detector configured to detect sleep-related disorder events using the sensed physiological signals, the sleep-related disorder events comprising involuntary muscle movement events of non-respiratory muscles and sleep-disordered breathing events, the detection of the sleep-disordered breathing events and the involuntary muscle movement events based at least in part on the electromyogram signals; anda processor configured to calculate a composite sleep disorder index, the composite sleep disorder index quantifying combined effect of the involuntary muscle movement events and the sleep-disordered breathing events, wherein at least one of calculating and detecting is configured to be implemented at least in part patient-internally. 2. The system of claim 1, further comprising communications circuitry configured to communicate one or both of the plurality of physiological signals and information associated with the detected sleep-related disorder events to a patient-external processing system. 3. The system of claim 1, further comprising communications circuitry configured to communicate one or more of the plurality of physiological signals, information associated with the detected sleep-related disorder events, and the composite sleep disorder index to an implantable medical device. 4. The system of claim 1, further comprising at least one of a controller configured to control therapy to treat a condition associated with one or more of the detected sleep-related disorder events based on the sleep movement disorder index and a therapy delivery unit configured to deliver the therapy to treat the condition associated with one or more of the detected sleep-related disorder events based on the composite sleep disorder index. 5. The system of claim 1, wherein the processor is configured to calculate the composite sleep disorder index including calculating an apnea/hypopnea index and a movement disorder index, the apnea/hypopnea index representative of an amount of apnea and hypopnea events relative to time and the movement disorder index representative of an amount of the involuntary muscle movement events relative to time. 6. The system of claim 1, wherein the sleep-disordered breathing events comprise one or more of sleep apnea, hypopnea, and Cheyne-Stokes respiration. 7. The system of claim 1, wherein the involuntary muscle movement events comprise one or more of bruxism, periodic limb movement disorder, and restless leg syndrome. 8. The system of claim 1, wherein the plurality of sensors includes a transthoracic impendence sensor. 9. The system of claim 1, wherein the processor is further configured to determine one or more sleep stages based at least in part on one or more of the physiological signals. 10. The system of claim 1, wherein the processo... | Cardiac Pacemakers, Inc. | A61B005/08 A61B005/04 A61B005/05
US08500649 | B2 | 2013-08-06 | Blood pressure measuring apparatus and method for measuring blood vessel elasticity | 2010-12-21 | A blood pressure measuring method which can measure blood vessel elasticity, so that that the blood vessel elasticity can be commonly applied to all subjects, calculating pulse transit time at the time when no pressure is applied, calculating pulse transit time in an area where the pressure is changed by the cuff, and then calculating the elasticity of a blood vessel, using the pulse transit times. | 1. A blood pressure measuring apparatus for measuring blood vessel elasticity, comprising: a Pulse Wave Velocity (PWV) measuring module for measuring PWV, using an electrocardiogram signal and a pulse wave signal;a blood pressure measuring module for measuring blood pressure by pressurizing or decompressing a cuff; anda control unit, wherein the control unit calculates an initial pulse transit time, using the PWV at a time when the cuff is not pressurized, wherein the initial pulse transit time is determined by dividing an arterial length from the heart to a pulse wave measuring point by the PWV, the control unit calculates pulse transit time in a pressure applied cuff area when the pressure is applied by the cuff, and the control unit calculates blood vessel elasticity, using the initial pulse transit time according to mean blood pressure in an area corresponding to the arterial length, a length of an area where pressure is applied by the cuff, the arterial length, a transmural pressure in the cuff area, and an increment of pulse transmit time in the cuff area according to the transmural pressure in the cuff area. 2. The apparatus of claim 1, wherein the PWV measuring module comprises: an electrocardiogram measuring unit for measuring the electrocardiogram signal; anda pulse wave measuring unit for measuring the pulse wave signal. 3. The apparatus of claim 1, wherein the blood pressure measuring module comprises: a cuff for pressurizing or decompressing an artery, the cuff being adapted to be wound on the upper part of an arm of a subject;an air pump for pressurizing the cuff;a pressure sensor for detecting pressure for pressurizing the cuff;an analog-digital converter for converting a signal from the pressure sensor to an electric signal; anda valve control unit for reducing the cuff pressure. 4. The apparatus of claim 1, wherein the initial pulse transit time is determined by dividing the arterial length from the heart to the pulse wave measuring point by the PWV, and Young's modulus of the artery is substituted for the PWV. 5. The apparatus of claim 4, wherein the blood pressure in the cuff area corresponds to mean blood pressure minus the pressure applied by the cuff. 6. The apparatus of claim 4, wherein the arterial length is determined by measuring the length from the heart to the pulse wave measuring point. 7. The apparatus of claim 4, wherein the arterial length is determined through a regression equation, using the sex and height of a subject. 8. The apparatus of claim 5, wherein the control unit calculates a blood vessel elastic modulus as the Young's modulus, using the initial pulse transmit time according to the mean blood pressure in an area corresponding to the arterial length, and the increment of pulse transmit time in the cuff area according to the blood pressure in the cuff area. 9. The apparatus of claim 8, wherein the blood vessel elastic modulus is expressed by,α=slope⁡(-2⁢⁢ln⁡(LcL⁢PTTo+Δ),Pt)wherein the slope functi... | Samsung Electronics Co., Ltd | A61B005/02
US08588878 | B2 | 2013-11-19 | Simultaneous measurement of pulse and regional blood oxygen saturation | 2010-11-12 | Methods and systems are provided that allow for the simultaneous calculation of pulse and regional blood oxygen saturation. An oximeter system that includes a sensor with a plurality of emitters and detectors may be used to calculate a pulse and/or regional blood oxygen saturation. A plurality of light signals may be emitted from light emitters. A first light signal may be received at a first light detector and a second light signal may be received at a second light detector. A pulse and/or regional blood oxygen saturation value may be calculated based on the received first and/or second light signals. The pulse and regional blood oxygen saturation values may be calculated substantially simultaneously. The calculated pulse and regional blood oxygen saturation values as well as other blood oxygen saturation values may be displayed simultaneously in a preconfigured portion of a display. | 1. A method for calculating blood oxygen saturations, the method comprising: emitting a plurality of light signals from light emitters;receiving a first light signal at a first light detector, wherein the received first light signal comprises a first set of impulses corresponding to a first wavelength of emitted light emitted from the light emitters and a second set of impulses corresponding to a second wavelength of light emitted from the light emitters;receiving a second light signal at a second light detector;calculating a pulse blood oxygen saturation based on a first subset of the first and second sets of impulses of the received first light signal, using a processor, wherein the calculation of pulse blood oxygen saturation is not based on a second subset of the first and second sets of impulses; andcalculating a regional blood oxygen saturation based on the second subset of the first and second sets of impulses and the received second light signal, using a processor, wherein the calculating a pulse blood oxygen saturation and the calculating a regional blood oxygen saturation are done substantially simultaneously. 2. The method of claim 1, wherein the first light detector is positioned at a first distance from the light emitters and the second light detector is positioned at a second distance, greater than the first distance, from the light emitters. 3. The method of claim 1, wherein the first light detector is associated with a first mean path length and wherein the second light detector is associated with a second mean path length, greater than the first mean path length. 4. The method of claim 1, wherein the emitting comprises emitting a first wavelength of light from light emitters for a first duration and emitting a second wavelength of light from light emitters for a second duration. 5. The method of claim 4, wherein the first wavelength is smaller than the second wavelength and the first duration is greater than the second duration. 6. The method of claim 1, wherein the plurality of light signals comprises a RED light signal and an IR light signal. 7. The method of claim 1, wherein the plurality of light signals comprises broadband light signals. 8. The method of claim 1, wherein the calculating the pulse blood oxygen saturation comprises a ratio-of-ratios technique based on a light intensity and wherein the calculating the regional blood oxygen saturation involves a differencing technique based on the light intensity. 9. The method of claim 1, wherein the calculating the pulse blood oxygen saturation comprises pre-processing the received first light signal using a first technique and the calculating the regional blood oxygen saturation comprises pre-processing the received first light signal using a second technique. 10. The method of claim 1, wherein the calculated pulse blood oxygen saturation and the calculated regional blood oxygen saturation are used to calculate a venous/capillary blood oxygen saturation. 11. A syste... | Covidien LP | A61B005/1455
US10094818 | B2 | 2018-10-09 | Bacterial/cellular recognition impedance algorithm | 2009-05-07 | A method for characterising cells or cell structures in a sample comprising: obtaining at least one normalized impedance response of the sample over at least one frequency range; and characterising at least one cell using at least one characteristic of the normalized impedance response. | 1. A method for characterizing unidentified cells or cell structures in a sample, said method comprising: implanting at least a portion of a device within a subject containing the sample, the device comprising a controller and at least two electrodes;monitoring, via the controller, at least one characteristic of the sample located in an electrical path between the electrodes;obtaining, via the controller and utilizing the at least two electrodes, at least one normalized impedance response of the sample over at least one frequency range, wherein the normalized impedance is a measured impedance divided by a baseline impedance, wherein the measured impedance comprises at least one of a real or an imaginary part, and wherein the baseline impedance is an impedance response of the sample over the at least one frequency range with substantially no cells in an electrical path between at least two electrodes;identifying, via the controller and within a set of data containing a plurality of normalized impedance responses associated with a corresponding plurality of cell types, at least one characteristic of the normalized impedance response over the at least one frequency range that is indicative of a single cell type within the plurality of cell types; andcharacterizing, via the controller, at least one unidentified cell to determine cell identity by using the identified at least one characteristic of the normalized impedance response over the at least one frequency range,wherein: the determined cell identity is characteristic of at least a cell type and a stage of growth of the at least one unidentified cell in the sample; andthe at least one characteristic of the normalized impedance response used for characterizing the at least one unidentified cell is at least one of a number of peaks in the normalized impedance. 2. A method according to claim 1, wherein the frequency range is a continuous or quasi-continuous frequency range or spectrum. 3. A method according to claim 1, wherein the frequency range comprises a series of discrete frequency measurements. 4. A method according to claim 1, wherein the frequency range is any frequency range lying between 0.1 Hz and 33 MHz. 5. A method according to claim 1, wherein at least one of: the at least one cell is a bacteria or other single cell organism;the at least one cell is a plant or animal cell such as a plant or animal tissue cell;at least one cell is arranged in a structure, such as plant or animal tissue or a multi-cell organism. 6. A method according to claim 1, wherein the baseline impedance response is at least one of an initial or a calculated or estimated or standard impedance response of the system. 7. A method according to claim 6, wherein the baseline impedance response of the system is further based at least in part upon measurement of an impedance response of at least one of a similar or a standardised system over the at least one frequency range with substantially no cells in an electrica... | University of Strathclyde | G01N033/487 A61B005/053 A61B005/1468 A61B005/1473 A61B005/00 C12M001/34 G01N027/02
US09439591 | B2 | 2016-09-13 | Analysis system and method for determining an analyte in a body fluid | 2010-12-07 | A system for determining an analyte in a body fluid includes a magazine comprising two partial magazines, one including chambers containing analysis elements having a sample contact zone and a reagent system containing a reagent, whose reaction with the body fluid results in a change of a variable, and another including chambers containing puncturing elements having a tip having a capillary channel, which forms a fluid connection between the tip and a sample transfer zone of the puncturing element. The system may also include an analysis instrument having a puncturing drive for driving a puncturing element on a movement path, a mounting for receiving the magazine such that one chamber of the magazine at a time is located in a functional position, in which a puncturing element in the chamber can be moved by the puncturing drive, and a measuring and evaluation apparatus for measuring the change of the variable. | 1. An analysis system for determining an analyte in a body fluid sample acquired through a prick in the skin, including: a multi-part magazine comprising at least a puncturing element magazine and an analysis element magazine, each having a plurality of chambers, wherein in the analysis element magazine, the plurality of chambers are analysis element chambers, which each contain one analysis element, each analysis element having a sample contact zone and a reagent system containing at least one reagent, whose reaction with the sample results in a measurable change of a measuring variable, andin the puncturing element magazine, the plurality of chambers are puncturing element chambers, which each contain one puncturing element, each puncturing element having a tip for piercing into the skin, the tip having a capillary channel which forms a fluid connection between the tip and a sample transfer zone of the puncturing element,a reusable analysis instrument having a puncturing drive for driving a puncture movement of a puncturing element on a movement path, which comprises a propulsion phase in the puncturing direction and, after reaching a reversal point of the puncture movement, a retraction phase opposite to the puncturing direction,a mounting for receiving the multi-part magazine in such a manner that one puncturing element chamber of the puncturing element magazine at a time is located in a functional position, in which a puncturing element in the puncturing element chamber can be moved by the puncturing drive, anda measuring and evaluation apparatus for measuring the measurable change of the measuring variable and for determining a desired analysis result,whereinin the functional position of the multi-part magazine, both the puncturing element magazine and the analysis element magazine can be brought into a relative position, in which a puncturing element exit opening of a puncturing element chamber and a puncturing element entry opening of an analysis element chamber are adjacent to one another aligned in such a manner that a puncturing element can be moved from the puncturing element chamber through the puncturing element exit opening of the puncturing element chamber and through the puncturing element entry opening of the adjacent analysis element chamber into the analysis element chamber,the puncturing element entry opening of the analysis element chamber is sealed by means of a sealing film which is opened before or during the puncture movement of the puncturing element from the puncturing element chamber into the analysis element chamber, andthe movement path of the puncturing element includes a transfer position in which the sample transfer zone of the puncturing element is positioned adjacent to the sample contact zone of the analysis element, to thereby produce a fluid connection by a relative movement, transverse to the puncturing direction, between the puncturing element and the analysis element to transfer the sample from the punctu... | Roche Diabetes Care, Inc. | A61B005/157 A61B005/15 A61B005/145 A61B005/1455 A61B005/151
US20110160544 | A1 | 2011-06-30 | SYSTEM AND METHOD FOR ANALYSIS OF MEDICAL DATA TO ENCOURAGE HEALTH CARE MANAGEMENT | 2010-12-29 | System and method analyze medical data of a patient having a disease afflicted health condition and action is taken to encourage the patient to perform wellness-enhancing activities and to take and report medical data more frequently. Consideration in the form of reduced insurance costs, medical supply costs, and medical equipment costs are given to patients who comply. Health care providers are advised that analysis of the patient's medical data is reimbursable and are also encouraged to perform such analyses through rewards. Patient data may be stored in a remote memory site, accessed by HCPs for analysis, and proprietary software may be used to communicate directly to insurance companies or other medical benefit entities. | 1. A medical data analysis system to encourage more rigorous management of a health condition, the system comprising: a medical sensor configured to sense a physiological parameter relevant to a particular disease-afflicted health condition and to provide medical data representative of the sensed parameter;a processor programmed to receive the medical data that is representative of the sensed parameter and relevant to the afflicted health condition, process the received medical data, process an analysis of the afflicted health condition that is based on the medical data, and automatically determine the status of the afflicted health condition based on the analysis; andthe processor further programmed to automatically provide a level of encouragement to more rigorously manage the afflicted health condition, wherein the level of encouragement is based on the determined status of the afflicted health condition. 2. The medical data analysis system of claim 1, wherein the encouragement comprises a reward and wherein the processor is further programmed to automatically provide the level of reward based on the determined status of the afflicted health condition. 3. The medical data analysis system of claim 2, wherein the reward comprises an automatically granted discount on a health insurance cost. 4. The medical data analysis system of claim 2, wherein the reward comprises an automatically granted discount on co-payment for office visits. 5. The medical data analysis system of claim 1, wherein the encouragement comprises a penalty and wherein the processor is further programmed to provide the level of penalty based on the determined status of the afflicted health condition. 6. The medical data analysis system of claim 1, wherein the encouragement comprises a reward and wherein the processor is further programmed to a provide a reward for the more frequent sensing of a physiological parameter relevant to the particular disease-afflicted health condition and for more frequently providing medical data representative of the sensed parameter to the processor for processing. 7. The medical data analysis system of claim 1, wherein the processor is further programmed to provide a report of the processed medical data, the analysis, and the determined status of the afflicted health condition. 8. The medical data analysis system of claim 1, further comprising a memory located remotely from the processor, and wherein the processor comprises a plurality of processors, one of which is programmed to receive the medical data that is representative of the sensed parameter and relevant to the afflicted health condition, process the received medical data, and store the processed medical data in the memory. 9. The medical data analysis system of claim 8, wherein another of the plurality of processors is programmed to retrieve the stored medical data from the memory, process an analysis of the afflicted health condition that is based on the medical data, automatica... | ABBOTT DIABETES CARE INC. | A61B005/00
US20110137180 | A1 | 2011-06-09 | Systems and Methods for Fitting a Cochlear Implant System to a Patient Based on Stapedius Displacement | 2010-12-02 | An exemplary system includes an implantable cochlear stimulator configured to be implanted within a patient, a sound processor communicatively coupled to the implantable cochlear stimulator and configured to direct the implantable cochlear stimulator to generate and apply electrical stimulation to one or more stimulation sites within a cochlea of the patient, and a displacement sensor assembly configured to measure a displacement of the stapedius of the patient that occurs in response to the application of the electrical stimulation and transmit displacement data representative of the measured displacement to the sound processor. Corresponding systems and methods are also disclosed. | 1. A system comprising: an implantable cochlear stimulator configured to be implanted within a patient;a sound processor communicatively coupled to the implantable cochlear stimulator and configured to direct the implantable cochlear stimulator to generate and apply electrical stimulation to one or more stimulation sites within a cochlea of the patient; anda displacement sensor assembly configured to measure a displacement of the stapedius of the patient that occurs in response to the application of the electrical stimulation and transmit displacement data representative of the measured displacement to the sound processor. 2. The system of claim 1, wherein the displacement sensor assembly is configured to be coupled to at least one of a stapedius tendon, a stapedius muscle, and a stapes of the patient. 3. The system of claim 1, wherein the sound processor is further configured to automatically fit the implantable cochlear stimulator to the patient in accordance with the displacement data. 4. The system of claim 1, wherein the sound processor is further configured to adjust one or more control parameters governing an operation of at least one of the sound processor and the implantable cochlear stimulator in accordance with the displacement data. 5. The system of claim 4, wherein the one or more control parameters comprise a most comfortable current level. 6. The system of claim 1, wherein the sound processor is further configured to determine a range of intensities of the electrical stimulation that are most comfortable for the patient in accordance with the displacement data. 7. The system of claim 1, further comprising a fitting device communicatively coupled to the sound processor, wherein the sound processor is configured to transmit the displacement data to the fitting device, and wherein the fitting device is configured to fit the implantable cochlear stimulator and the sound processor to the patient in accordance with the displacement data. 8. The system of claim 1, wherein the displacement sensor assembly is electrically coupled to the implantable cochlear stimulator and configured to transmit the displacement data to the sound processor by way of the implantable cochlear stimulator. 9. The system of claim 1, wherein the displacement sensor assembly is configured to wirelessly transmit the displacement data to the sound processor. 10. The system of claim 1, wherein the displacement sensor assembly is configured to measure the displacement of the stapedius by measuring a strain of a stapedius tendon of the patient. 11. The system of claim 10, wherein the displacement sensor assembly comprises a strain gauge assembly comprising: a buckle frame configured to be coupled to an outer portion of the stapedius tendon and deform in response to the displacement of the stapedius tendon; anda strain gauge coupled to the buckle frame and configured to measure the strain by detecting the deformation of the buckle frame. 12. The system of clai... | Advanced Bionics AG, c/o Froriep Renggli | A61B005/05 A61B006/00
US20110118559 | A1 | 2011-05-19 | NEUROLOGICAL AND/OR PSYCHOLOGICAL TESTER | 2011-01-19 | A method to be performed on a computer for neuropsychological evaluation of a person includes calculating complexities of a plurality of stimuli by a pattern description length algorithm for a pattern of each the stimulus, presenting stimuli to the user in reaction to which the user has to perform a cognitive test, the stimuli being presented in order of complexity and adjusting a complexity level for a next cognitive test based on the user's response. | 1. A method to be performed on a computer for neuropsychological evaluation of a person, the method comprising: calculating complexities of a plurality of stimuli by a pattern description length algorithm for a pattern of each said stimulus;presenting stimuli to said user in reaction to which said user has to perform a cognitive test, said stimuli being presented in order of complexity, andadjusting a complexity level for a next cognitive test based on said user's response. 2. A method according to claim 1 wherein said pattern description length algorithm for a stimulus j is: PDL(j)mini[nj(i)+log pi(i)]wherein n is a word description length, n=L−S where L is a Lempel-Ziv code and S is a description-facilitating sequence; wherein i is a tracing through a pattern of said stimulus; wherein nj(i) is the number of bits required to describe a word created from pattern j using tracing i and wherein pj(i) is the a-priori probability for compression using each tracing i. 3. A method according to claim 1 wherein calculating said complexity includes a measure reflecting a ratio between complexities of stimuli. 4. A method according to claim 3, wherein a task complexity ratio for stimuli patterns A and B is RPC=PC(A)*PC(A)/PC(B). 5. A method to be performed on a computer for neuropsychological evaluation of a person, the method comprising: presenting tests containing stimuli in reaction to which said person has to perform a cognitive task, said tests having multiple tasks of different complexity levels and wherein said complexity levels are associated with said stimuli;during prior training of the person on a test and during said test: defining an expected response to said stimuli,measuring the response of said person to the stimuli associated with the test,analyzing said response,reducing a complexity of stimuli for a next task if said person incorrectly responds to stimuli of said task, andincreasing a complexity of stimuli for a next task if said person correctly responds to stimuli of said task;calculating said complexity for each next task prior to determining said next task;changing an expected response time of a particular stimuli as a function of said person's responses;storing results from said tests; andevaluating performance of said person in reaction to stimuli of said tests. 6. The method according to claim 5 and also comprising: determining a score S for each person, whereinS=∑iwifi&lt;Tand wherein T is a threshold defining a disorder, fi are a set of features from the received responses in a test and wi are per feature weights based on population statistics. 7. The method according to claim 5 and also comprising: calculating complexities of a plurality of stimuli by a pattern description length algorithm for a pattern of each said stimulus;presenting stimuli to said user in reaction to which said user has to perform a cognitive test, said stimuli being presented in order of complexity, andadjusting a complexity level for a next cognit... | - | A61B005/00
US20110160607 | A1 | 2011-06-30 | QEED-Guided Selection and Titration of Psychotropic Medications | 2008-12-12 | Data corresponding to brain electrical activity of a patient from a QEEG is processed to determine which brain activity data deviates from that of a normative profile. The deviant brain activity data is used to determine a brain state vector (BSV) in a multidimensional brain electrical signal space, wherein the orientation of the BSV indicates the nature of the deviation from a normal state and may be used to prescribe a medicine to the patient and wherein the length of the BSV quantifies the degree of abnormality exhibited by the patient and may indicate the dosage of the medicine to be prescribed to the patient. | 1. A method of determining a pharmacological treatment, comprising: determining whether a difference exists between brain activity data corresponding to electrical brain activity of a living body and control data corresponding to normative brain activity data;generating a first brain state vector (BSV) for the living body having an orientation and a length corresponding, respectively, to a quality and a magnitude of any determined difference; anddetermining a pharmacological treatment based on at least one of the length and the orientation of the first BSV. 2. The method according to claim 1, wherein the determining of the pharmacological treatment includes: selecting at least one drug associated with a second BSV orientation that is opposite that of the first BSV. 3. The method according to claim 2, father comprising: determining a dosage of the drug based on the length of the first BSV. 4. The method according to claim 2, wherein the determining of the pharmacological treatment includes: selecting a first drug associated with a third BSV orientation and a second drug associated with a fourth BSV orientation so that a vector summation of the third BSV and the fourth BSV produces a resultant BSV that is substantially opposite in orientation with respect to the orientation of the first BSV of the living body. 5. The method according to claim 4, wherein the pharmacological treatment is one of manually administered and automatically administered by a titration unit. 6. The method according to claim 4, further comprising: monitoring the first BSV after a beginning of the pharmacological treatment to determine whether the first BSV is being reduced to a point of origin of a brain state vector space. 7. The method according to claim 6, further comprising: if the magnitude of the first BSV one of remains constant and becomes larger after the beginning of the pharmacological treatment, selecting a new pharmacological treatment. 8. A system for determining a pharmacological treatment, comprising: an arrangement for determining whether a difference exists between brain activity data corresponding to electrical brain activity of a living body and control data corresponding to normative brain activity data;an arrangement for generating a first brain state vector (BSV) for the living body having an orientation and a length corresponding, respectively, to a quality and a magnitude of any determined difference; andan arrangement for determining a pharmacological treatment based on at least one of the length and the orientation of the first BSV. 9. The system according to claim 8, wherein the arrangement for determining the pharmacological treatment includes: an arrangement for selecting at least one drug associated with a second BSV orientation that is opposite that of the first BSV. 10. The system according to claim 8, wherein the arrangement for determining the pharmacological treatment includes: an arrangement for selecting a first drug associated w... | NEW YORK UNIVERSITY | A61B005/0476
US20110144517 | A1 | 2011-06-16 | Video Based Automated Detection of Respiratory Events | 2011-02-23 | A computer implemented method for automated sleep monitoring, including recording live images of a subject sleeping, transmitting the recorded images to a computing device, receiving the transmitted images at the computing device, performing motion analysis of the subject based on the received images, computing an air flow signal that represents the amount of air flowing into the lungs of the subject over time from the results of said motion analysis, automatically detecting respiratory events based on said airflow signal, wherein a respiratory event is a breathing disturbance, and displaying the respiratory events experienced by the subject on a monitor. An apparatus is also described. | 1. Apparatus for automatically monitoring sleep, comprising: a video recorder for recording live images of a subject sleeping, comprising a transmitter for transmitting the recorded images to a computing device; anda computing device communicating with said video recorder transmitter, comprising: a receiver for receiving the transmitted images;a motion analyzer for performing motion analysis of the subject based on the received images;a respiratory event detector for (1) computing an air flow signal that represents the amount of air flowing into the lungs of the subject over time from the results of said motion analysis, and (2) automatically detecting respiratory events based on said airflow signal, wherein a respiratory event is a breathing disturbance; anda monitor for displaying the respiratory events experienced by the subject, detected by said computing device. 2. The apparatus of claim 1 wherein said monitor also displays the received images. 3. The apparatus of claim 1 wherein said respiratory event detector identifies candidate respiratory events, wherein a candidate respiratory event is indicated by the derivative of the airflow signal being uniformly negative during a first interval followed by a second interval during which the derivative of the airflow signal is uniformly positive. 4. The apparatus of claim 3 wherein a respiratory event is determined from a candidate respiratory event by applying a threshold to the first interval, the period of time during which the airflow signal is uniformly negative 5. The apparatus of claim 3 wherein a respiratory event is determined from a candidate respiratory event by applying a threshold to the average derivative of the airflow signal during said first time interval or to the average derivative of the airflow signal during the second time interval. 6. The apparatus of claim 3 wherein said monitor also displays derived respiratory information, wherein said derived respiratory information is a member of the group consisting of respiratory events, candidate respiratory events, and the air flow signal. 7. The apparatus of claim 1 wherein said computing device further comprises: a non-volatile memory; anda log manager for selectively storing the received images and derived respiratory information in the non-volatile memory for subsequent post-analysis or display, wherein said derived respiratory information is a member of the group consisting of the airflow signal and the detected respiratory events. 8. A computer implemented method for automated sleep monitoring, comprising: recording live images of a subject sleeping;transmitting the recorded images to a computing device;receiving the transmitted images at the computing device;performing motion analysis of the subject based on the received images;computing an air flow signal that represents the amount of air flowing into the lungs of the subject over time from the results of said motion analysis;automatically detecting respiratory events ... | - | A61B005/087
US20110112413 | A1 | 2011-05-12 | METHOD FOR AUTOMATED EKG ANALYSIS | 2009-11-06 | Systems and methods for analyzing electronic cardiac signals for use in clinical diagnostics are described. System and method embodiments are configured to assist in the analysis of details of EKG signals and vector cardiograms to determine how patients should be categorized into specific cardiac risk categories, such as an acute coronary syndrome category. System configurations may comprise memory devices, computing systems, and EKG data sources positioned at various local or remote positions, and connected via various data connectivity modalities. Various parameters may be utilized to assist in the drawing of one or more conclusions regarding the cardiac condition of a patient. | 1. A method for detecting a first cardiac condition, comprising: a. providing a quantity of 12-lead EKG data from a patient;b. constructing a 3-D representation of cardiac activity from the data;c. computing values for one or more parameters based upon at least one of said 3-D representation of cardiac activity and said EKG data;d. applying a multifactorial analysis protocol utilizing at least one of the values of the one or more parameters; ande. automatically drawing one or more conclusions regarding a first cardiac condition of the patient based at least in part upon the output of the multifactorial analysis protocol. 2. The method of claim 1, further comprising acquiring the quantity of 12-lead EKG data from the patient. 3. The method of claim 1, wherein providing comprises reconstructing a 12-lead EKG recording from a reduced cardiographic vector set from the patient. 4. The method of claim 3, further comprising receiving the vector set from another device. 5. The method of claim 4, wherein the other device is selected from the group consisting of an implantable defibrillator, an implantable pacemaker, an implantable cardioverter, a portable EKG monitoring system, and a desktop EKG system. 6. The method of claim 2, wherein providing further comprises transmitting the EKG data from a plurality of electrodes to a computing system configured to construct the 3-D representation and compute values for the one or more parameters. 7. The method of claim 6, wherein transmitting the EKG data from the electrodes to the computing system occurs in real or near-real time. 8. The method of claim 6, wherein transmitting the EKG data from the electrodes to the computing system is delayed with utilization of a memory device interposed between the electrodes and computing system. 9. The method of claim 1, wherein automatically drawing one or more conclusions regarding the first cardiac condition of the patient further comprises categorizing the patient into a specific category of acute coronary syndrome relative to other patients. 10. The method of claim 1, wherein the electrodes and computing system are interfaced with a connection selected from the group consisting of a local wired connection, a local wireless connection, a remote wired connection, and a remote wireless connection. 11. The method of claim 1, wherein the multifactorial analysis protocol further utilizes one or more parameters computed directly from the 12-lead EKG data. 12. The method of claim 1, further comprising providing feedback to a healthcare provider regarding the conclusions. 13. The method of claim 1, further comprising computing one or more confidence indices regarding the one or more conclusions based at least in part upon the one or more parameter values. 14. The method of claim 1, wherein the multifactorial analysis protocol comprises quantitative analysis selected from the group consisting of discriminatory analysis, regression analysis, and hierarchical analysis.... | NEWCARDIO, INC | A61B005/02
US20110118617 | A1 | 2011-05-19 | Heart Rate Measuring Device and Method for Measuring Heart Rate Thereof | 2009-11-25 | This invention discloses a heart rate measuring device and a method for measuring heart rate thereof. The heart rate measuring device can be put on a user body. The heart rate measuring device includes a first electrode, a second electrode, a plurality of optical indicating units, a signal processing circuit and a control unit. When the user's both hands touch the first electrode and the second electrode, the control unit provides a heart signal by using the first electrode, the second electrode and the signal processing circuit. The control unit controls the optical indicating units to output electrocardiographic signals. | 1. A heart rate measuring device, worn by a subject, comprising: a first electrode;a second electrode;a plurality of optical indicating units;a signal processing circuit, coupled to the first electrode and the second electrode; anda control unit, coupled to the signal processing circuit and the optical indicating units, wherein when the first electrode and the second electrode are both in contact with the subject, the control unit provides electrocardiographic signals by using the first electrode, the second electrode and the signal processing circuit, the control unit further controlling the optical indicating units to output electrocardiographic signals. 2. The heart rate measuring device of claim 1, wherein the first electrode and the second electrode are used to detect the electrocardiographic signals. 3. The heart rate measuring device of claim 1, wherein the first electrode and the second electrode are used to detect electrocardiographic signals, the signal processing unit used to process the detected electrocardiographic signals. 4. The heart rate measuring device of claim 1, wherein the heart rate measuring device is off when the first electrode and the second electrode are not both in contact with both hands of the subject. 5. The heart rate measuring device of claim 1, wherein the optical indicating units include a first optical indicating unit, a second optical indicating unit, a third optical indicating unit and a fourth optical indicating unit, indicating that the subject is detected with cardiac arrhythmia as all four lights are on or flashing. 6. The heart rate measuring device of claim 1, wherein the optical indicating units include a first optical indicating unit, a second optical indicating unit, a third optical indicating unit and a fourth optical indicating unit, indicating that the heart rate of the subject is too slow as the first optical indicating unit is on. 7. The heart rate measuring device of claim 6, wherein when the second optical indicating unit is on, the subject's heart rate is normal. 8. The heart rate measuring device of claim 6, wherein when the third optical indicating unit is on, the subject's heart rate is a little fast. 9. The heart rate measuring device of claim 6, wherein when the third optical indicating unit is on, the subject's heart rate is too fast. 10. The heart rate measuring device of claim 1, further including an audio indicating unit coupled to the control unit. 11. The heart rate measuring device of claim 10, wherein the audio indicating unit makes a sound with each heart beat along the electrocardiographic signals. 12. A heart rate measuring method, used to measure a subject's heart rate, comprising: providing a first electrode and a second electrode;providing a plurality of optical indicating units;when the first electrode and the second electrode are both in contact with the subject, the device starts measuring the heart rate of the subject to generate electrocardiographic signals... | NATIONAL YANG MING UNIVERSITY | A61B005/0404 A61B005/0468 A61B005/024
US08197425 | B2 | 2012-06-12 | Ankle-foot orthosis prescription assistive device | 2009-11-13 | A measurement device (10) for quantitatively measuring stiffness, torque or range of motion (ROM) around the joint axis of the ankle-foot complex of a lower limb in the sagittal and coronal planes, the device (10) comprising: a plate (13, 21) for placement of the ankle-foot complex, the plate (13, 21) operatively connected to an actuator (14, 16) for manual actuation; an angular measurement device (12, 17) to measure an angular position of the ankle-foot complex; a torquemeter (11) to measure resistance torque of the ankle-foot complex; an electromyograph (48) to monitor muscular activity of muscles of the lower limb; and computer software to record the measured angular position, resistance torque of the ankle-foot complex and EMG of the lower limb muscles. | 1. A measurement device for quantitatively measuring stiffness, torque or range of motion (ROM) around the joint axis of the ankle-foot complex of a lower limb of a seated subject in the sagittal and coronal planes, the device comprising: a plate for placement of the ankle-foot complex, the plate operatively connected to an actuator for manual actuation;an angular measurement device operatively connected to the plate to measure an angular position in the sagittal and coronal planes of the ankle-foot complex of the seated subject around the rotational center of the ankle-foot complex;a torquemeter operatively connected to the plate to measure resistance torque in the sagittal and coronal planes of the ankle-foot complex of the seated subject around the joint rotational center of the ankle-foot complex;an electromyograph (EMG) to monitor muscular activity of muscles of the lower limb; andcomputer software to record the measured angular position in the sagittal and coronal planes of the ankle-foot complex, the measured resistance torque in the sagittal and coronal planes of the ankle-foot complex, and the measured EMG of the muscles of the lower limbwherein the computer measures the stiffness, torque, or range of motion based on the measured angular position, the measured resistance torque, or the measured EMG. 2. The device according to claim 1, wherein the plate is a foot plate or a rotary plate. 3. The device according to claim 2, wherein the angular measurement device is a potentiometer or a protractor. 4. The device according to claim 2, wherein the foot plate has a range of motion in the sagittal and coronal planes, and the range of motion of the foot plate is adjustable with stoppers positioned under a handle. 5. The device according to claim 1, further comprising a base and base frames to support the device, wherein the angular position of the base relative to the base frame is adjustable in the sagittal plane. 6. The device according to claim 1, further comprising a metronome to monitor angular velocity of the foot plate. 7. The device according to claim 3, further comprising a rotary plate to fix the ankle-foot complex via the foot plate, the foot plate being freely movable on the rotary plate and its height is adjustable to enable correct positioning of the axis of the rotary plate to the estimated rotational center of the ankle-foot complex with reference to the extended axis of the torquemeter and the potentiometer. 8. The device according to claim 7, wherein the actuator is a handle or a steering wheel to manually rotate the rotary plate around the axis of the torquemeter and the potentiometer. 9. The device according to claim 1, wherein the EMG comprises electrodes attached to dorsiflexor and plantarflexor muscles. 10. The device according to claim 1, further comprising a hand-held dynamometer and associated software to measure muscle strength of each joint to determine an ankle-foot orthosis (AFO) prescription. 11. The devi... | Hong Kong Polytechnic University | A61B005/103 A61B005/117 A61B005/04
US08323194 | B2 | 2012-12-04 | Detection of bubbles during hemodynamic monitoring when performing automated measurement of blood constituents | 2009-12-18 | Example embodiments of the present invention provide methods and apparatuses that enable the detection of bubbles so that hemodynamic performance can be assured following an automated blood analyte measurement. An example apparatus according to the present invention comprises a blood access system, adapted to remove blood from a body and infuse at least a portion of the blood back into the body. The infusion of at least a portion of the blood back in to the body can be done in a manner to assure that no bubbles of clinical significance are injected into the patient. Additionally an example embodiment can assess for the presence of bubbles in the fluid column that can affect hemodynamic monitoring performance. If a condition exists where hemodynamic monitoring performance cannot be assured, an example embodiment can provide appropriate warning or corrective actions. | 1. A method of determining the presence of a bubble in a blood access system comprising at least one pressure detector, comprising: a. Using the pressure detector to determine a first frequency response of the system at a first time;b. Using the pressure detector to determine a second frequency response of the system at a second time;c. Determining if a bubble is present in the system by comparing the first and second frequency responses. 2. A method as in claim 1, wherein the blood access system is in fluid communication with the circulatory system of a patient. 3. A method as in claim 2, wherein a. determining a first frequency response comprises determining a frequency-dependent characteristic of the output of the pressure detector responsive to changes in blood pressure in the circulatory system, andb. determining a second frequency response comprises determining a frequency-dependent characteristic of the output of the pressure detector responsive to changes in blood pressure in the circulatory system. 4. A method as in claim 1, wherein comparing the first and second frequency responses comprises subtraction, division, Fourier transform analysis, wavelet analysis, vector comparison, derivative processing, or a combination thereof. 5. A method as in claim 1, wherein a. determining a first frequency response comprises determining the power spectral density of the output of the pressure detector during a first time interval, andb. determining a second frequency response comprises determining the power spectral density of the output of the pressure detector during a second time interval, andc. determining if a bubble is present comprises determining whether the second frequency response indicates more power at frequencies corresponding to presence of a bubble than does the first frequency response. 6. A method of determining the presence, concentration, or both of an analyte in the blood of a patient, comprising: a. Placing a blood access system in fluid communication with the circulatory patient, wherein the blood access system comprises at least one pressure sensor, at least one analye sensor, and at least one pump;b. Using the pressure sensor to determine the frequency response of the blood access system at a first time before step c;c. Operating the pump to withdraw blood from the patient to the analyte sensor;d. Operating the analyte sensor to determine the presence, concentration, or both of an analyte in the withdrawn blood;e. Using the pressure sensor to determine the frequency response of the blood access system at a second time after step c;f. Determining if a bubble is present in the blood access system by comparing the frequency response determined at the first time with the frequency response determined at the second time. 7. A method as in claim 6, further comprising operating the pump to move the bubble out of portions of the blood access system where the bubble has an effect on the frequency response of the blood access sys... | InLight Solutions, Inc. | A61B005/00
US07962196 | B2 | 2011-06-14 | Method and system for determining the location of a medical instrument relative to a body structure | 2007-08-02 | A method for determining a location of an object relative to a body structure includes: providing a location of a first part of the object relative to a second part of the object, said second part different from said first part; attaching a marker device to the body structure; detecting a location of the marker device relative to the body structure; positioning the object relative to the body structure; detecting a location of the first part of the object relative to the marker device; and calculating a relative location of the second part of the object relative to the location of the body structure based on the location of the first part relative to the second part, the location of the marker device relative to the body structure, and the location of the first part relative to the marker array. | 1. A method for determining a location of an object relative to a body structure, comprising: a) providing a location of a first point or first plane corresponding to a first part of the object relative to a second point or second plane corresponding to a second part of the object, said second part different from said first part;b) attaching a body marker device to the body structure;c) detecting a location of the body marker device relative to the body structure;d) positioning the object relative to the body structure;e) detecting a location of the first point or first plane relative to the body marker device, said detecting performed without using a marker array attached to the object; andf) calculating a relative location of the second part of the object relative to the location of the body structure based on the location of the first point or first plane relative to the second point or second plane, the location of the body marker device relative to the body structure, and the location of the first point or first plane relative to the body marker device. 2. The method according to claim 1, wherein the object is a medical instrument at least partially situated inside the body structure. 3. The method according to claim 1, wherein providing the location of the first point or first plane relative to the second point or second plane includes measuring the location of the first point or first plane relative to the second point or second plane. 4. The method according to claim 1, wherein the object comprises at least a first element and a second element, wherein the first part is a part of the first element, and the second part is a part of the second element. 5. The method according to claim 4, wherein the second element is removably attachable to the first element such that a relative location between the first part and the second part is known prior to attachment of the second element to the first element. 6. The method according to claim 1, wherein detecting a location of the first point or first plane relative to the body marker device includes optically detecting a location of the first point or first plane relative to the body marker device. 7. A system for determining a location of an object relative to a body structure, comprising: a) a device for providing a location of a first point or first plane corresponding to a first part of the object relative to a second point or second plane corresponding to a second part of the object, said second part different from the first part;b) a body marker device;c) a first detection device for detecting a location of the body marker device relative to the body structure when the body marker device is stationary relative to the body structure;d) a second detection device for detecting a location of the first point or first plane relative to the body marker device; ande) a computational unit configured to calculate, without using a marker device attached to the object, a location of the second part... | Brainlab AG | A61B005/055
US08249709 | B2 | 2012-08-21 | Methods and apparatus for automatically tracking heart failure status | 2011-03-01 | Assessing symptomatic and asymptomatic physiologic changes due to chronic heart failure involves apparatus and methods for gauging degradation and possible improvement using automated measurement of inter-ventricular conduction time, both alone and in combination with other automated physiologic tests. Conduction times increase due to the greater distance a wavefront must traverse as a heart enlarges. Analysis of conduction time can be used to verify the occurrence of cardiac remodeling due to heart failure as well as beneficial reverse remodeling due to successful heart failure therapy delivery. Patient activity level(s) and presence/increase in pulmonary fluids can also be used to automatically determine changes in heart failure status and/or predict hospitalization. Conduction time is monitored between electrodes positioned in the left and right ventricles of the heart via endocardial or epicardial electrodes. | 1. A method of assessing heart failure status in a patient, comprising: measuring a current inter-ventricular cardiac conduction time (IVCT);comparing the measured current IVCT to a previously measured IVCT;determining a rate of change of IVCT from the comparison;comparing the rate of change of the IVCT to a defined range of change; andidentifying a possible change in cardiac status responsive to the rate of change of the IVCT being within the range of change; andresponsive to the identification of possible change in cardiac status, performing one of:storing the current IVCT in a memory structure,notifying a patient of a possible change in cardiac status,notifying a clinician of a possible change in cardiac status,generating an alarm;the method further comprising:launching a stimulation wave front from a first electrode positioned proximate a first ventricular chamber; andsensing a local cardiac depolarization at a second electrode positioned proximate a second ventricular chamber,wherein detecting cardiac conduction time includes detecting a time between launching the wave front and sensing the local cardiac depolarization. 2. A method according to claim 1, wherein the first electrode includes an endocardial electrode positioned within the first ventricular chamber, and the second electrode includes an epicardial electrode positioned on a surface of the second ventricular chamber. 3. A method according to claim 1, wherein the first electrode includes an endocardial electrode positioned within the first ventricular chamber, and the second electrode includes an endocardial electrode positioned within the second ventricular chamber. 4. A method according to claim 1, wherein the first electrode includes an epicardial electrode positioned on a surface of the first ventricular chamber, and the second electrode includes an epicardial electrode positioned on a surface of the second ventricular chamber. 5. A method according to claim 1, wherein the first ventricular chamber is the right ventricular chamber and the second ventricular chamber is the left ventricular chamber. 6. A method according to claim 1, wherein the first ventricular chamber is the left ventricular chamber and the second ventricular chamber is the right ventricular chamber. 7. A method according to claim 1, further comprising one of: assessing an activity level of the patient and assessing an intra-thoracic fluid status of the patient; andcomparing said activity level and said fluid status to respective threshold values, and in the event that the activity level is lower than said threshold and said fluid status exceeds said threshold then:notifying a patient or a clinician of a treble-confirmed heart failure progression situation. 8. A method according to claim 7, further comprising one of: assessing an activity level of the patient and assessing an intra-thoracic fluid status of the patient; andcomparing said activity level and said fluid status to respective threshold values, ... | Medtronic, Inc. | A61N001/37
US08660628 | B2 | 2014-02-25 | Analyte sensors comprising blended membrane compositions and methods for making and using them | 2009-12-21 | Embodiments of the invention provide analyte sensors having elements designed to modulate their chemical reactions as well as methods for making and using such sensors. In certain embodiments of the invention, the sensor includes an analyte modulating membrane that comprises a blended mixture of a linear polyurethane/polyurea polymer, and a branched acrylate polymer. | 1. An amperometric analyte sensor apparatus comprising: a base layer comprising a working electrode;an analyte sensing layer coating the working electrode; andan analyte modulating layer coating the analyte sensing layer, wherein:the analyte modulating layer comprises a blended mixture of a linear polyurethane/polyurea polymer and a branched acrylate polymer;the linear polyurethane/polyurea polymer exhibits a permeability to glucose that decreases by between 1% and 8% per degree centigrade as temperature is increased from 22 to 40 degrees centigrade;the branched acrylate polymer exhibits a permeability to glucose that increases by between 1% and 8% per degree centigrade as temperature is increased from 22 to 40 degrees centigrade;the linear polyurethane/polyurea polymer and the branched acrylate polymer are blended at a ratio of between 1:1 and 1:20 by weight %; andthe analyte modulating layer comprises a blend of:(1) a polyurethane/polyurea polymer formed from a mixture comprising: (a) a diisocyanate;(b) a hydrophilic polymer comprising a hydrophilic diol or hydrophilic diamine; and(c) a siloxane having an amino, hydroxyl or carboxylic acid functional group at a terminus; and(2) a branched acrylate polymer formed from a mixture comprising: (a) a butyl, propyl, ethyl or methyl-acrylate;(b) an amino-acrylate;(c) a siloxane-acrylate; and(d) a poly(ethylene oxide)-acrylate. 2. The analyte sensor apparatus of claim 1, wherein the blended analyte modulating layer exhibits a permeability to glucose that changes less than 2% per degree centigrade over a temperature range of 22 to 40 degrees centigrade. 3. The analyte sensor apparatus of claim 1, wherein the analyte modulating layer facilitates in vivo hydration of the sensor so that levels of an in vivo analyte can be sensed less than 45 minutes after sensor implantation into an in vivo environment. 4. The analyte sensor apparatus of claim 1, wherein the sensor does not include a separate layer of an adhesion promoting material disposed between the analyte sensing layer and the analyte modulating layer. 5. The analyte sensor apparatus of claim 1, further comprising at least one of: a protein layer coating the analyte sensing layer; ora cover layer coating the analyte sensor apparatus, wherein the cover layer comprises an aperture positioned on the cover layer so as to facilitate an analyte present in an in vivo environment from contacting and diffusing through an analyte modulating layer; and contacting the analyte sensing layer. 6. The analyte sensor apparatus of claim 1, wherein the analyte sensing layer comprises glucose oxidase. 7. The analyte sensor apparatus of claim 1, wherein the base layer comprises a plurality of electrodes including the working electrode, a counter electrode and a reference electrode. 8. The analyte sensor of claim 7, wherein the base layer comprises a plurality of working electrodes, counter electrodes and reference electrodes. 9. The analyte sensor of claim 1, where... | Medtronic MiniMed, Inc. | A61B005/05
US08579830 | B2 | 2013-11-12 | Neuropathy diagnostic device | 2009-07-17 | A neuropathy diagnostic device includes a body having a neutral temperature surface disposed on an exterior of the body, a variable temperature surface disposed on the exterior of the body, and a vibrating surface having a variable vibration amplitude. The vibrating surface extends outwardly from the body. A touch sensor extends outwardly from the body. A processor is disposed within the body. The processor is adapted to set a temperature of the neutral temperature surface, vary a temperature of the variable temperature surface, and vary the vibration amplitude of the vibrating surface. | 1. A neuropathy diagnostic device comprising: a body having a first side and a second side opposite the first side, the body further comprising:a neutral temperature surface extending outwardly from the first side;a variable temperature surface extending outwardly from the second side; anda vibration surface having a variable vibration amplitude and extending outwardly from the body between the first side and the second side; anda processor disposed within the body, wherein the processor is adapted to set a temperature of the neutral temperature surface, vary a temperature of the variable temperature surface, and vary the vibration amplitude of the vibration surface. 2. The neuropathy diagnostic device according to claim 1, wherein the variable temperature surface is variable between a first temperature less than the neutral temperature surface and a second temperature greater than the neutral temperature surface. 3. The neuropathy diagnostic device according to claim 2, wherein the variable temperature surface is incrementally variable by about two degrees below the neutral temperature surface. 4. The neuropathy diagnostic device according to claim 2, wherein the variable temperature surface is incrementally variable by about two degrees above the neutral temperature surface. 5. The neuropathy diagnostic device according to claim 1, wherein the vibration surface comprises a pressure sensor. 6. The neuropathy diagnostic device according to claim 5, wherein the pressure sensor is coupled to an indicator for indicating sufficient pressure applied by the pressure sensor to a surface. 7. The neuropathy diagnostic device according to claim 1, further comprising an output for recording data generated by each of the neutral temperature surface, the variable temperature surface, and the vibration surface. 8. The neuropathy diagnostic device according to claim 1, further comprising a touch sensor extending outwardly from the body. 9. The neuropathy diagnostic device according to claim 1, wherein the device is handheld. 10. The neuropathy diagnostic device according claim 1, wherein the body comprises a single handheld body. 11. A neuropathy diagnostic device comprising: a body having:a neutral temperature surface extending outwardly from the body;a variable temperature surface extending outwardly from the body; anda vibration surface having a variable vibration amplitude, the vibration surface extending outwardly from the body;a processor disposed within the body, wherein the processor is adapted to set a temperature of the neutral temperature surface, vary a temperature of the variable temperature surface, and vary the vibration amplitude of the vibration surface; anda first button operatively coupled to the processor to alternatively vary one of the temperature of the variable temperature surface and the vibration of the vibration surface in only a first direction; anda second button operatively coupled to the processor to alternatively vary t... | Prosenex, LLC | A61B005/00 A61B019/00 A61F002/00 A61F007/00 G01K001/08 G01K001/14 G01K013/00
US08630695 | B2 | 2014-01-14 | Locator, device and method for electronically locating and reading the setting of an adjustable valve | 2007-12-14 | The invention is aimed at improving valve pinpointing and reading devices by proposing to activate a magnetic detection source allowing measurements regularly-distributed in a circle around an axis at predetermined measurement positions. An exemplary locator comprises a casing (26) provided with a mechanism for detecting and analyzing the magnetic field of the magnetic dipole of the valve. It comprises:—a magnetic detection source (36) able to detect and measure the magnetic field of the magnetic center of the valve at predetermined measurement positions distributed in a circle around an axis (X-X′), linked to—a microprocessor (32) for analyzing the measurements and for generating detection signals. | 1. A method for electronically locating a magnetic center of an implanted magnetic adjustable valve exhibiting a fluid flow reference axis (Ax) and a center furnished with a magnetic dipole (24), wherein said method comprises of the following steps: pinpointing the fluid flow reference axis (Ax) for flow of fluid through the valve (10),activating a mechanism for electronically detecting a magnetic field (36, 74, 82) at predetermined measurement positions (α1-α12) angularly and uniformly distributed in a circle around an axis (X-X′),electronically measuring the magnetic field (CM) emitted by the valve at the predetermined measurement positions (α1-α12),displacing the mechanism for electronically detecting a magnetic field until the magnetic field measured is substantially the same at substantially diametrically opposite positions, the magnetic center of the valve being in the center between said opposite positions. 2. The locating method according to claim 1, furthermore comprising the step of pinpointing an orientation of the dipole (24) of the valve with respect to the flow axis (Ax) on the basis of the magnetic field measurements, the orientation of the dipole determining a setting of the valve. 3. The locating method according to claim 2, in which the orientation of the dipole is pinpointed by implementing the following steps: a) searching for a measurement position Imax1 having a strongest negative signal;b) searching, on the basis of the position Imax1 determined in step a) and according to a direction of rotation of the mechanism, among the subsequent measurement positions, for a position Imax2 having a strongest positive signal;c) searching for a measurement position Imin1 of a sensor having a weakest signal in terms of absolute value situated between the two positions Imax1 and Imax2 determined in steps a) and b);d) determining, from among the measurement positions preceding and succeeding the position Imin1 determined in step c), the position Imin2 of the sensor having the weakest signal in terms of absolute value;a setting position of the valve being equal to: Imin1+Imin2, if Imin1 is less than Imin2, or to Imin1+Imin2+1, if Imin2 is less than Imin1. 4. The locating method according to claim 1, comprising, furthermore, the steps of calculating a depth of implantation of the valve on the basis of a mean of the absolute values of the measurements taken at each position, and of indicating the depth of the valve by comparing the mean obtained with predetermined calibration values. 5. The locating method according to claim 1, in which the mechanism for electronically detecting a magnetic field (36, 74, 82) is rotated about said axis (X-X′) and in a determined direction of rotation (R), the measurements of the magnetic field (CM) emitted by the valve being made at the predetermined measurement positions (α1−α12). 6. The locating method according to claim 5, in which the mechanism for electronically detecting a magnetic field (36) moves i... | Sophysa | A61B005/05 A61B019/00
US08805533 | B2 | 2014-08-12 | Systems and methods of neuromodulation stimulation for the restoration of sexual function | 2011-02-21 | Systems and methods use an external and/or implantable pulse generator system for neuromodulation stimulation to treat sexual dysfunction by the unilateral or bilateral stimulation of a target nerve A and/or a target nerve B using one or more leads and electrodes implanted in tissue in the pelvic region. The electrical stimulation waveform may be conveyed to the target nerve A electrode for a first predetermined amount of time, and conveyed to the target nerve B electrode for a second predetermined amount of time. | 1. A stimulating catheter system comprising a catheter, the catheter comprising: an elongated flexible body having a proximal end and a distal end and at least one lumen extending at least a portion of the body;an inflatable balloon disposed at or near the distal end;an electrode array disposed along the flexible body, wherein respective electrodes of the electrode array are regularly spaced along the flexible body; andan electrical lead coupled to each electrode in the electrode array and extending through the body, wherein the electrical lead is configured to be connected to an electrical source via an external connection,wherein the catheter is configured to deliver neuromodulation stimulation to a patient from the electrical source via the electrical lead and the electrode array,wherein the at least one lumen includes a first, andwherein the first lumen exits the flexible body at or near the distal end and serves as a path for fluid flow in and out of the flexible body at or near the distal end. 2. The catheter system according to claim 1, wherein the electrode array is disposed proximal to the inflatable balloon along the flexible body. 3. The catheter system according to claim 1, wherein the flexible body is configured to be positioned within the urethra of a male and/or female animal or human. 4. The catheter system according to claim 1, wherein the flexible body is configured to be positioned within the rectum of a male and/or female animal or human. 5. The catheter system according to claim 1, wherein the flexible body is configured to be positioned within the vagina of a female animal or human. 6. The catheter system according to claim 1, wherein at least one electrode is positioned on or near the distal tip of the flexible body. 7. The catheter system according to claim 1, wherein the electrode array is coupled to the outside of the flexible body. 8. The catheter system according to claim 1, wherein the flexible body is malleable. 9. The catheter system according to claim 1, wherein the flexible body is adapted to be bendable while positioned within a body cavity, and is configured to be bendable during removal without causing damage to the body cavity or internal structures. 10. The catheter system according to claim 1, wherein the flexible body is pre-bent in a desired configuration so as to maintain its shape during insertion and stimulation. 11. The catheter system according to claim 1, wherein the electrode array comprises a proximal portion and a distal portion, the electrodes of the proximal portion are spaced apart a first distance, and the electrodes of the distal portion are spaced apart at a second distance less than the first distance. 12. The catheter system according to claim 1, wherein the electrode array comprises a proximal portion and a distal portion, the electrodes of the proximal portion are spaced about 1 cm to about 5 cm apart, and the electrodes of the distal portion are spaced apart about 0.1 mm to a... | Medtronic, Inc. | A61N001/36
US07967817 | B2 | 2011-06-28 | Multi-electrode lead | 2008-07-15 | A multi-electrode lead comprises an elongate carrier having a longitudinal axis. A plurality of electrical conductors is carried by the carrier, a plurality of electrodes being connected in spaced relationship to each conductor. The electrodes are arranged at axially spaced intervals along the carrier so that, along the length of the carrier, any one electrode associated with any one of the conductors only once has an electrode associated with another one of the conductors adjacent to that any one electrode. In another aspect of the invention, an ablating device comprises an elongate tubular sleeve for effecting spot ablation of tissue, a sensing electrode catheter being receivable in a passage of a sleeve for assisting in positioning of an ablating electrode, in use. Preferably, the sensing electrode catheter is the multi-electrode lead of the first aspect of the invention. | 1. An ablating device, the ablating device comprising: an elongate, tubular sleeve defining an open passage; andan ablating electrode carried at a distal end of the sleeve for effecting spot ablation of tissue, a sensing electrode catheter being receivable in the open passage of the sleeve for assisting in positioning of the ablating electrode, in use; anda sensing electrode catheter received in the open passage of the sleeve, the sensing electrode catheter being in the form of a multi-electrode lead comprising: an elongate carrier having a longitudinal axis; anda plurality of electrical conductors carried by the carrier, a plurality of electrodes being connected in spaced relationship to each conductor, the plurality of electrodes being arranged at axially spaced intervals along the carrier so that, along the length of the carrier, any one electrode associated with any one of the conductors only once has an electrode associated with another one of the conductors adjacent to the any one electrode. 2. The ablating device of claim 1 in which the carrier comprises a mandrel about which the conductors are carried. 3. The ablating device of claim 2 in which the mandrel is a tubular mandrel to form a lumen of the lead. 4. The ablating device of claim 1 in which a covering of a non-conductive material is applied about the plurality of electrical conductors to cover the plurality of electrical conductors. 5. The ablating device of claim 4 in which the plurality of electrodes is formed by elements of electrically conductive material applied to an outer surface of the covering. 6. The ablating device of claim 5 in which the electrically conductive material from which the elements are made is a biocompatible metal. 7. The ablating device of claim 6 in which each element is in the form of a band or annulus arranged about the carrier. 8. The ablating device of claim 7 in which, at least at a region of the carrier having the plurality of electrodes, the covering is of a porous material. 9. The ablating device of claim 8 in which, where the covering is of a porous material, a metal is applied about the covering to permeate through the pores of the covering to make contact with the plurality of electrical conductors. 10. The ablating device of claim 6 in which the covering is removed at regions where it is desired to form electrodes to expose the plurality of electrical conductors at a region with the metal forming the elements being in direct electrical contact with their associated conductors. 11. The ablating device of claim 10 in which each electrode is defined by a region of each conductor exposed by the removal of the covering at that region. 12. The ablating device of claim 1 in which the ablating electrode is a ring electrode arranged at the distal end of the sleeve. 13. A method of ablating tissue at a site in a patient's body, the method comprising: inserting an ablating device, as claimed in claim 1, through a body vessel of the patient;pl... | Cathrx Ltd. | A61B018/14 A61B005/04
US07969117 | B2 | 2011-06-28 | Method and device for charging peripherals | 2009-11-06 | The invention includes a peripheral charging system for a computer. The peripheral charging system comprises a retention mechanism configured to retain a peripheral having a battery, the retention mechanism comprising a charging mechanism, and a system battery electrically connected to the charging mechanism. When the peripheral is retained by the retention mechanism, the peripheral contacts the charging mechanism of the retention mechanism to charge the peripheral battery. The invention also comprises a method for charging a peripheral for a mobile computing device. The method comprises retaining the peripheral in a retention mechanism of the mobile computing device, and charging a battery of the peripheral device from a battery of the mobile computing device. The retention mechanism of the mobile computing device includes a charging mechanism configured to supply charge to a charging mechanism of the peripheral. | 1. A system comprising: a peripheral charging module arranged within a mobile computing device and configured to charge a first battery of a peripheral device through an electrical connection with a second battery of the mobile computing device; anda controller configured to manage power consumption of the second battery and determine if the second battery is capable of providing power to the first battery,wherein the mobile computing device contains firmware and software that supports wireless use of the peripheral device by the mobile computing device,wherein the mobile computing device is adapted to communicate wirelessly with the peripheral device and the peripheral device is adapted to communicate wirelessly with the mobile computing device, andwherein the peripheral device comprises a wireless medical device,wherein the controller is configured to determine if the second battery has a power capacity level above a predetermined charge threshold to charge the first battery, to determine an operating time remaining to the mobile computing device provided by the second battery and to determine the predetermined charge threshold is based on the determined operating time. 2. The system of claim 1, further comprising a programmable voltage regulator. 3. The system of claim 2, further comprising voltage control interface. 4. The system of claim 1, wherein the system is configured to provide power at a default voltage and current. 5. The system of claim 1, wherein the controller is configured to manage charging power. 6. The system of claim 5, wherein the controller is configured to prioritize charging of one or more peripheral devices. 7. A mobile computing device comprising: a peripheral device receiving member arranged within the mobile computing device and dimensioned such that a peripheral device is physically retained within the peripheral device receiving member;a peripheral charging module arranged within the mobile computing device and electrically connected to the peripheral device receiving member and configured to charge a first battery of the peripheral device through an electrical connection with a second battery of the mobile computing device;a controller configured to manage power consumption of the second battery and determine if the second battery is capable of providing power to the first battery; andfirmware and software that supports wireless use of the peripheral device by the mobile computing device andwherein the mobile computing device is adapted to communicate wirelessly with the peripheral device and the peripheral device is adapted to communicate wirelessly with the mobile computing device,wherein the peripheral device comprises a wireless medical device,wherein the controller is configured to determine if the second battery has a power capacity level above a predetermined charge threshold to charge the first battery, to determine an operating time remaining to the mobile computing device provided by the second batt... | Intel Corporation | H02J007/00 A61B005/00
US08260409 | B2 | 2012-09-04 | Self-adjusting ECG morphological feature correlation threshold | 2011-02-01 | An apparatus comprises first and second sensing circuits, a template generator circuit, and a correlation circuit. The correlation circuit is configured to identify a first fiducial position in a third cardiac signal sensed using the first sensing circuit during a detected rhythm with elevated ventricular rate, align the template correlation features and the correlation features of a fourth cardiac signal using the first fiducial position, calculate a correlation using the correlation features of the template and the correlation features of the fourth cardiac signal, and iteratively searching for a replacement to the first fiducial position in the third cardiac signal according to the calculated correlation. | 1. A method comprising: sensing a first intrinsic cardiac signal during a normal sinus rhythm (NSR) using at least one heart rate sensing electrode and sensing a second intrinsic cardiac signal during the NSR using at least a second sensing electrode;generating a template using the first and second sensed cardiac signals, the template including a fiducial point on the first cardiac signal and correlation features on the second cardiac signal;sensing a third intrinsic cardiac signal during a detected rhythm with elevated ventricular rate using the rate sensing electrode and sensing a fourth intrinsic cardiac signal during the detected rhythm using the second sensing electrode;identifying a first fiducial position in the third cardiac signal and aligning the template correlation features and the correlation features of the fourth cardiac signal using the first fiducial position;calculating a correlation using the correlation features of the template and the correlation features of the fourth cardiac signal;iteratively searching for a replacement to the first fiducial position in the third cardiac signal according to the calculated correlation; anddetecting tachyarrhythmia using the calculated correlation. 2. The method of claim 1, wherein the iteratively searching for the replacement to the first fiducial position includes: identifying a central fiducial position, an earlier fiducial position and a right fiducial position in the third cardiac signal, wherein the central fiducial position is the first fiducial position;calculating a correlation at each of the fiducial positions; andreplacing the central fiducial position with a fiducial position having the maximum correlation. 3. The method of claim 2, wherein replacing the central fiducial position includes: incrementing a max count for each of the fiducial positions when the corresponding fiducial position has the maximum correlation for that cardiac cycle;replacing the central fiducial position with the fiducial position having a highest max count; anditeratively restarting the search for a replacement fiducial position every N cardiac cycles, where N is an integer. 4. The method of claim 3, including: calculating the correlation for the next beat when the beat count is less than N; andreplacing the central fiducial position with the fiducial position having a highest max count when the beat count is equal to N. 5. The method of claim 1, including limiting the search for a replacement to the first fiducial position to within a time interval of an original first fiducial position. 6. The method f claim 1, wherein aligning the template correlation features includes aligning the four sensed cardiac signals. 7. The method of claim 1, wherein the iteratively searching for the replacement to the first fiducial position includes: choosing a second fiducial position in the third signal that is shifted from the first fiducial position by at least one sample;re-calculating the correlation; anddetermi... | Cardiac Pacemakers, Inc. | A61B005/0402
US07937131 | B2 | 2011-05-03 | PET—MRI hybrid apparatus and method of implementing the same | 2004-12-28 | There is provided a PET-MRI hybrid apparatus and method for integrating a PET image and an MRI image so that anatomical, hemodynamical and molecular information on human tissues are simultaneously presented in a single image. The PET-MRI hybrid system comprises a first scanner for obtaining anatomical and hemodynamical information, and a second scanner for obtaining molecular and functional information on the human tissues. Along a path between the first scanner and the second scanner, a transferring railway system which includes runs, and a movable bed for supporting a subject installed on the railway. The PET-MRI hybrid system also comprises a “RF+ magnetic” shield and a “magnetic” shield between path between the first scanner and the second scanner, which switch between an open status and a close status in a completely synchronized manner to assure a complete magnetic shield for the PET system at any given time. The subject is fastened on the bed and transferred along the railway between the first and second scanner to provide accurately fused MRI and PET images. | 1. An imaging apparatus for providing anatomical, hemodynamical, molecular and functional information on a subject, the imaging apparatus comprising: a magnetic resonance imaging scanner configured to obtain said anatomical information as well as hemodynamical information on said subject in an integrated form, said magnetic resonance imaging scanner generating magnetic fields external to said magnetic resonance imaging scanner, said magnetic resonance imaging scanner vulnerable to external RF signals whereby the magnetic resonance imaging scanner fails to correctly obtain anatomical information;a PET scanner configured to obtain said molecular and functional information, said PET scanner disposed in close proximity to the first scanner, said PET scanner vulnerable to magnetic fields generated by said magnetic resonance imaging scanner whereby the PET scanner is damaged by said magnetic fields generated by said magnetic resonance imaging scanner;an RF and magnetic shield enclosing a first predetermined space including said magnetic resonance imaging scanner but not the PET scanner, the RF and magnetic shield inhibiting external RF fields sufficiently that the magnetic resonance imaging scanner does not fail to correctly obtain anatomical information caused by said external RF fields, and preventing magnetic fields of said magnetic resonance imaging scanner from leaking outside the first predetermined space, a first opening defined in the RF and magnetic shield, the first opening selectively closable with a first shutter;a magnetic shield enclosing a second predetermined space including said PET scanner but not the magnetic resonance image scanner, and shielding the PET scanner from said magnetic fields out of said magnetic resonance imaging scanner, whereby the RF and magnetic shield enclosing the first predetermined space and the magnetic shield enclosing the second predetermined space together providing provide sufficient shielding such that the PET scanner is not reached by the magnetic fields generated by said magnetic resonance imaging scanner, a second opening defined in the magnetic shield enclosing the second predetermined space, the second opening selectively closable with a second shutter;an imaging processor configured to construct a fused image by processing anatomical and hemodynamical information obtained by the magnetic resonance imaging scanner, and molecular and functional information obtained by the PET scanner;a transferring railway system oriented in a substantially horizontal plane and extending between the magnetic resonance imaging scanner and the PET scanner through the first opening and the second opening, the railway system including a selectively rotatable portion;a bed movable along said transferring railway and for supporting said subject; anda controller coupled to the bed, the first shutter, and the second shutter, configured to control the movement of the bed along the railway, and coordinating the opening and closi... | Gachon University of Medicine & Science Industry-Academic Cooperation Foundation | A61B005/05
US09155496 | B2 | 2015-10-13 | Low oxygen in vivo analyte sensor | 2011-02-18 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous and subcutaneous measurement of glucose in a host. | 1. A transcutaneous glucose monitoring system comprising: an in vivo portion and an ex vivo portion;wherein the in vivo portion comprises:an implantable body comprising an electrode configured to measure a glucose level in a host;a membrane disposed over the electrode and configured to limit transport of glucose to the electrode, wherein the membrane comprises polyurethaneurea, wherein the membrane is configured to provide the system with substantial linearity at glucose concentration of up to 400 mg/dL at an oxygen concentration of 0.45 mg/L;a layer comprising an enzyme configured to catalyze a reaction of glucose and oxygen from a biological fluid surrounding the membrane; andwherein the ex vivo portion comprises a sensor electronics unit operably connected to the electrode and configured to measure a current produced by the electrode. 2. The transcutaneous glucose monitoring system of claim 1, wherein the sensor electronics unit is configured to directly measure the current produced by the electrode. 3. The transcutaneous glucose monitoring system of claim 1, further comprising a biointerface membrane configured to support tissue ingrowth. 4. The transcutaneous glucose monitoring system of claim 1, further comprising an analog-to-digital converter configured to translate the current into a digital signal. 5. The transcutaneous glucose monitoring system of claim 1, wherein the electrode comprises an exposed electroactive working electrode surface with a surface area of from about 0.00002 in2 to about 0.0079 in2. 6. The transcutaneous glucose monitoring system of claim 1, wherein the membrane comprises a resistance domain configured to have a permeability ratio of at least about 50:1 of glucose to an interferant. 7. The transcutaneous glucose monitoring system of claim 1, wherein the membrane comprises a resistance domain configured to have a permeability ratio of at least about 200:1 of glucose to an interferant. 8. The transcutaneous glucose monitoring system of claim 1, wherein the sensor system is configured to have, in operation, a sensitivity of from about 1 pA/mg/dL to about 100 pA/mg/dL. 9. The transcutaneous glucose monitoring system of claim 1, wherein the sensor system is configured to have, in operation, a sensitivity of from about 5 pA/mg/dL to about 25 pA/mg/dL. 10. The transcutaneous glucose monitoring system of claim 1, wherein the sensor system is configured to have, in operation, a sensitivity of from about 3.5 to about 7.5 pA/mg/dL. 11. The transcutaneous glucose monitoring system of claim 1, wherein the sensor electronics unit is configured to measure glucose at an oxygen concentration of less than about 0.3 mg/L. 12. The transcutaneous glucose monitoring system of claim 1, wherein the sensor electronics unit is configured to measure glucose at an oxygen concentration of less than about 0.15 mg/L. 13. The transcutaneous glucose monitoring system of claim 1, wherein the sensor electronics unit is configured to meas... | DexCom, Inc. | A61B005/145 A61B005/00 A61B005/15 A61B005/1486 A61B005/1495
US08515532 | B2 | 2013-08-20 | Compensation of motion artifacts in capacitive measurement of electrophysiological signals | 2009-08-24 | The invention relates to a system and a method in which an electrophysiological signal is sensed capacitively with at least two closely spaced electrodes such that the electrodes experience strongly correlated skin-electrode distance variations. To be able to derive a motion artifact signal, the capacitive coupling between the electrodes and skin is made intentionally different. With a signal processing means the motion artifact signal can be removed from the measured signal to leave only the desired electrophysiological signal. Since the measured quantity is dependant on the electrode-skin distance itself, the system and method do not need to rely on the constancy of a transfer function. Hereby, they give reliable motion artifact free output signals. | 1. A system for providing motion compensated capacitive measurement of electrophysiological signals, the system comprising: a sensor probe comprising a first capacitive plate and a second capacitive plate arranged to be placed in a proximity of a skin of a human or animal body, wherein a first capacitance between the first capacitive plate and the skin is different from a second capacitance between the second capacitive plate and the skin, and wherein the first capacitive plate is connected to an input of a first amplifier for providing a first measurement signal and the second capacitive plate is connected to an input of a second amplifier for providing a second measurement signal as a measure of an electrophysiological signal of the human or animal body; anda processor connected to the sensor probe and configured for performing the acts of: receiving said first and second measurement signals from the first and second capacitive plates;determining a motion modulated signal from at least one signal of the first measurement signal and the second measurement signal; andproviding a motion compensated output signal from the motion modulated signal and the at least one signal. 2. The system according to claim 1, wherein the first and second capacitive plates of the sensor probe have mutually different areas. 3. The system according to claim 1, wherein the sensor probe moreover comprises an active shield. 4. The system according to claim 1, wherein the determining act comprises calculating a difference between the first measurement signal and a factor applied to the second measurement signal, and wherein the factor is a ratio between the first and second measurement signals integrated by respective integrators. 5. The system according to claim 1, wherein the determination act comprises averaging or low pass filtering of the first and second measurement signals over a predetermined time interval. 6. The system according to claim 1, wherein the providing act comprises using a least-mean-square algorithm to adaptively remove the effect of the motion modulated signal from the at least one signal. 7. The system according to claim 1, wherein the first and second capacitive plates are placed so close that they effectively measure the same signal of the user. 8. The system according to claim 1, wherein the first and second capacitive plates of the sensor probe have different distances to a measurement surface of the sensor probe. 9. The system according to claim 1, wherein the first and second capacitive plates of the sensor probe have different distances to a measurement surface of the sensor probe and the first and second capacitive plates have mutually different areas. 10. A signal processing unit for compensating for motion, said signal processing unit comprising: a first amplifier configured to receive a first measurement signal from a first capacitive plate of a sensor probe arranged to be placed in a proximity of a skin of a human or animal bod... | Koninklijke Philips N.V. | A61B005/04 A61B005/0408 A61B005/05 A61B005/053
US08498689 | B2 | 2013-07-30 | Real time motion information capture in an MRI environment | 2009-11-04 | Systems, methods and articles of manufacture are disclosed for capturing motion information in a magnetic resonance imaging (MRI) environment. A light sink in the MRI environment may detect light emitted from a plurality of light sources. Each of the plurality of light sources may emit light of a different frequency. Further, each of the plurality of light sources may be located at a different spatial position in the MRI environment. The detected light may be analyzed. A change in spatial position of the light sink may be computed based on the analysis. | 1. A computer-implemented method for capturing motion information in a magnetic resonance imaging (MRI) environment, based on light sources having different frequencies and different spatial positions, the method comprising: detecting, by a light sink in the MRI environment, light emitted from a plurality of light sources, wherein each of the plurality of light sources emit light of a different frequency of a plurality of frequencies, wherein each of the plurality of light sources is located at a different spatial position of a plurality of spatial positions in the MRI environment, and wherein the light sink is detachably fastened to a subject;analyzing the detected light to determine, by operation of one or more computer processors, a change in spatial position of the light sink; anddetermining a change in spatial position of the subject during an MRI scan, based on the determined change in spatial position of the light sink. 2. The computer-implemented method of claim 1, wherein the method further comprises: receiving an MRI image of the subject during the MRI scan; andcorrecting a motion in the received MRI image, based on the determined change in spatial position of the subject. 3. The computer-implemented method of claim 1, wherein analyzing the detected light comprises determining a change in peak frequency of the detected light. 4. The computer-implemented method of claim 3, wherein determining a change in peak frequency of the detected light includes decomposing, via a fast Fourier transform, the detected light into frequency components of the detected light. 5. The computer-implemented method of claim 1, wherein the method further comprises: determining the plurality of frequencies for the plurality of light sources; anddetermining the plurality of spatial positions for the plurality of light sources. 6. The computer-implemented method of claim 1, wherein the light sink is an open-ended tip of an optical fiber cable, wherein each of the plurality of light sources is an open-ended tip of an optical fiber cable, and wherein the location of each of the plurality of light sources is fixed in the MRI environment. 7. A computer-readable memory containing a program which, when executed, performs an operation for capturing motion information in a magnetic resonance imaging (MRI) environment, based on light sources having different frequencies and different spatial positions, the operation comprising: detecting, by a light sink in the MRI environment, light emitted from a plurality of light sources, wherein each of the plurality of light sources emit light of a different frequency of a plurality of frequencies, wherein each of the plurality of light sources is located at a different spatial position of a plurality of spatial positions in the MRI environment, and wherein the light sink is detachably fastened to a subject;analyzing the detected light to determine, by operation of the one or more computer processors, a change in spatial positi... | International Business Machines Corporation | A61B005/05
US08460218 | B2 | 2013-06-11 | Patient weight bearing monitor | 2010-11-30 | A weight sensor includes a pad of electrically insulative material having a through aperture and electrode plates secured to its top and bottom surfaces at least partially covering the aperture. An electrical contactor in the aperture responds to a compressive force applied across the pad thickness to electrically interconnect the electrodes and close a circuit that permits an alarm to be actuated. The sensor is disposed in footwear worn by a patient being rehabilitated from a lower limb surgery or injury. The stiffness/compressibility of the pad determines the applied patient weight or compressive force required to permit the contactor to close the circuit between the electrodes. | 1. A method of actuating an alarm when the weight of a patient applied on a lower limb of the patient exceeds predetermined weights, said method comprising the steps of: (a) providing a first sensor as a first electrically insulative compressible pad between two first pad electrodes that are spaced across the thickness of the first pad, the first pad being comprised of a material having a first density;(b) selecting said first density to calibrate the first sensor such that, in response to application of a force exceeding a first predetermined force across the first pad thickness, compression of the first pad effects electrical contact between the two first pad electrodes;(c) providing at least a second sensor as a second electrically insulative compressible pad between two second pad electrodes that are spaced across the thickness of the second pad, the second pad being comprised of a material having a second density;(d) selecting said second density different from said first density to calibrate the second sensor such that, in response to application of a force exceeding a second predetermined force across the second pad thickness, compression of the second pad effects electrical contact between the two second pad electrodes, wherein said second predetermined force is different from said first predetermined force;(e) alternatively positioning said first and second sensors beneath the foot of the patient to enable weight of the patient to be applied through the lower limb across the sensor thickness;(f) providing an electrical signal in response to electrical contact between the two first pad electrodes and in response to electrical contact between the two second pad electrodes; and(g) actuating said alarm with said electrical signal. 2. The method of claim 1 wherein step (b) includes: (b.1) providing a first pad aperture extending through the entire thickness of said first pad; and(b.2) securing a first pad electrical contactor to one of said first pad electrodes and extending through said aperture toward the other of said first pad electrodes, said first pad contactor being positioned to be spaced from said other of said first pad electrodes in the absence of force applied across the first pad thickness and in contact with said other of said first pad electrodes in response to application of a force exceeding said first predetermined force across the first pad thickness;and wherein step (d) includes:(d.1) providing a second pad aperture extending through the entire thickness of said second pad; and(d.2) securing a second pad electrical contactor to one of said second pad electrodes and extending through said second pad aperture toward the other of said second pad electrodes, said contactor being positioned to be spaced from said other of said second pad electrodes in the absence of force applied across the second pad thickness and in contact with said other of said second pad electrodes in response to application of a force exceeding said sec... | - | A61B005/00
US08940527 | B2 | 2015-01-27 | Integrated device for analyte testing, confirmation, and donor identity verification | 2010-12-22 | The present invention provides an apparatus for fluid sample collection and analyte testing, including a sample receiving member and at least one membrane test strip, and optionally a sample retention member, fingerprint acquisition pad, and/or fluid collector. It also provides a fluid collection apparatus having an absorbent material, compression element, and closure element, and optionally a lid that allows the apparatus to be used in conjunction with a fluid container. Also provided are methods of collecting, testing, and retaining a fluid sample and verifying the identity of one or more individuals associated with the sample, such as the test subject, test administrator, and/or witnesses. | 1. An apparatus comprising: a sample receiving member having an opening for receiving a fluid sample;a sample retention member in fluid communication with the sample receiving member to retain a portion of the fluid sample;at least one membrane test strip in fluid communication with the sample receiving member to indicate the presence or absence of at least one analyte in the fluid sample; anda fluid collector to collect the fluid sample and convey the fluid sample into the sample receiving member, wherein the fluid collector comprises an absorbent material to absorb the fluid sample; a compression member operatively associated with the absorbent material; and a closure member capable of sealing the open end of the sample receiving member when the fluid collector is inserted in the sample receiving member. 2. The apparatus of claim 1, wherein the closure member is capable of creating an airtight seal in the open end of the sample receiving member. 3. The apparatus of claim 1, further comprising means for securing the fluid collector within the sample receiving member after the fluid collector has been inserted into the sample receiving member. 4. The apparatus of claim 3, wherein the means for securing the fluid collector within the sample receiving member comprise at least one annular projection within the sample receiving member. 5. The apparatus of claim 1, wherein the fluid collector further comprises a handle removably attached to the upper surface of the closure member. 6. The apparatus of claim 1, further comprising a saliva-producing substance. 7. The apparatus of claim 6, wherein the saliva-producing substance is sugar, salt, acid, or any combination thereof. 8. The apparatus of claim 7, wherein the saliva-producing substance is sodium chloride, citric acid, or any combination thereof. 9. The apparatus of claim 1, further comprising a fingerprint acquisition pad. 10. The apparatus of claim 9, wherein the fingerprint acquisition pad is enclosed by a cover attached to the housing. 11. The apparatus of claim 9, wherein the fingerprint acquisition pad is an ink-based fingerprint acquisition pad. 12. The apparatus of claim 11, further comprising a dispenser able to dispense an ink that can elicit a signal in the ink-based fingerprint acquisition pad. 13. The apparatus of claim 9, wherein the fingerprint acquisition pad is an inkless fingerprint acquisition pad. 14. The apparatus of claim 13, further comprising a dispenser able to dispense an activator that can elicit a signal in the inkless fingerprint acquisition pad. 15. The apparatus of claim 13, wherein the inkless fingerprint acquisition pad is an immunoassay-based fingerprint acquisition pad. 16. The apparatus of claim 15, wherein the immunoassay-based fingerprint acquisition pad is in fluid communication with the sample receiving member. 17. The apparatus of claim 16, wherein a channel provides the fluid communication between the sample receiving member and the at least... | Lamina Equities Corp. | C12M001/34 A61B010/00 G01N033/558 G01N033/94 A61B019/00
US08374683 | B2 | 2013-02-12 | Medical instrument with probe, probe cover, and methods of using the same | 2009-12-30 | A medical instrument including a probe adapted to be inserted into an orifice of an animal's body. The instrument includes an emitter of electromagnetic radiation that is sensed by a sensor mounted on the probe. The sensor can then determine, by variations in the amount of radiation received from the emitter, whether a probe cover is mounted over the probe. A sensor can also be used to determine whether the walls of the body orifice block the radiation received from the emitter, thereby indicating position of the sensor, and thus, the probe within the orifice. A special probe cover is disclosed, and methods of using the medical instrument are also disclosed. | 1. A medical instrument including a support structure;a probe attached on said support structure, said probe possessing an outer peripheral surface and adapted to be inserted into an orifice in an animal's body;an emitter that emits electro-magnetic radiation toward said outer peripheral surface of said probe, said emitter mounted on said support structure;a sensor adapted to sense said radiation, said sensor mounted on said outer peripheral surface of said probe in a position so as to receive said radiation;a probe cover adapted to be selectively mounted over and to substantially cover said probe and said sensor, said probe cover being partially transparent to said radiation such that when said probe cover substantially covers said probe, said sensor senses reduced transmittance of said radiation; anda microprocessor operatively connected to said sensor, said microprocessor configured to monitor the magnitude of said radiation sensed by said sensor and to evaluate the degree of any reduction of said monitored magnitude as an indication of the insertion of said probe into the body orifice. 2. A medical instrument according to claim 1 wherein said emitter comprises a blue LED and wherein said radiation possesses a strong intensity in the bandwidth substantially in the range of about 500-550 nanometers. 3. A medical instrument according to claim 1 wherein said probe cover transmits between about 40-60 percent of said radiation impinging thereon within a selected bandwidth. 4. A medical instrument according to claim 1 wherein said probe cover is coated with a material that blocks a substantial portion of said radiation within a selected bandwidth from passing therethrough. 5. A medical instrument according to claim 1 wherein said probe cover possesses a substantially frusto-conical configuration. 6. A medical instrument according to claim 1 wherein said medical instrument consists essentially of an ear instrument. 7. A medical instrument according to claim 1 wherein said medical instrument consists essentially of an infrared thermometer. 8. A medical instrument according to claim 1 wherein said probe extends in a substantially longitudinal direction and possesses an outer exterior surface, wherein said sensor extends in a substantially elongate, longitudinal direction along said outer peripheral surface of said probe and whereby when said probe is inserted into the orifice, said radiation is partially blocked by the walls of the orifice from reaching said sensor and as said probe is inserted deeper into the orifice even more of said radiation is blocked by the walls of the orifice from reaching said sensor. 9. A medical instrument according to claim 8 wherein said microprocessor is configured to calculate a depth of insertion of said probe into the orifice in a substantially linear relationship with the decrease in the magnitude of said radiation sensed by said sensor. 10. A medical instrument according to claim 9 wherein said microprocesso... | - | A61B006/00
US09848898 | B2 | 2017-12-26 | Navigation and tissue capture systems and methods | 2009-03-27 | Navigation and tissue capture systems and methods for navigation to and/or capture of selected tissue using the innate electrical activity of the selected tissue and/or other tissue are described. In the context of left atrial appendage closure, the systems and methods can be used to navigate to the left atrial appendage and capture/control the appendage while a closure instrument (suture, clip, ring) is placed over the appendage and tightened down or a closure method (ablation, cryogenic procedures, stapling, etc.) is performed to close the left atrial appendage. The use of innate electrical activity for navigating devices may be used in connection with other tissues and/or areas of the body. | 1. A method of navigating a device to a left atrial appendage, the method comprising: delivering a capture device on a capture shaft into a pericardial sac, wherein the capture device comprises a first jaw and a second jaw, wherein each of the first jaw and the second jaw comprises an interior surface, wherein the first jaw and the second jaw comprise an open configuration in which the first jaw and the second jaw are open and a closed configuration in which the first jaw and the second jaw are closed, wherein the interior surface of the first jaw is located closer to the interior surface of the second jaw when the first jaw and the second jaw are in the closed configuration than when the first jaw and the second jaw are in the open configuration, and wherein a first electrode is located on the internal surface of the first jaw and a second electrode is located on the internal surface of the second jaw;moving the first jaw and the second jaw from the open configuration to the closed configuration after delivering the capture device into the pericardial sac, wherein moving the first jaw and the second jaw captures a left atrial appendage between the first jaw and the second jaw;detecting an EGM signal using the first electrode on the first jaw and the second electrode on the second jaw of the capture device after moving the first jaw and the second jaw from the open configuration to the closed configuration to capture the left atrial appendage within the pericardial sac and confirming capture of the left atrial appendage by the capture device by determining if the EGM signal is associated with tissue of the left atrial appendage;wherein delivering the capture device on the capture shaft into the pericardial sac comprises advancing the capture device over ventricular myocardium tissue before the capture device reaches the left atrial appendage, detecting a capture shaft EGM signal using a capture shaft electrode located on the capture shaft proximal from the capture device, and confirming that the capture shaft electrode is contacting ventricular tissue using the capture shaft EGM signal. 2. A method of navigating a device to the left atrial appendage according to claim 1, the method further comprising confirming capture of atrial tissue by the capture device by electrically stimulating the atrial tissue and confirming that the atrial tissue is being paced. 3. A method of navigating a device to the left atrial appendage according to claim 1, the method further comprising determining that the left atrial appendage is not captured between the first jaw and the second jaw if the EGM signal is not associated with tissue of the left atrial appendage. 4. A method of navigating a device to the left atrial appendage according to claim 1, wherein the first jaw and the second jaw comprise atraumatic nesting retention structure located at distal ends of each of the first jaw and the second jaw, wherein proximal ends of each of the first jaw and the second ... | MAYO FOUNDATION FOR MEDICAL EDUCATION AND RESEARCH AEGIS MEDICAL INNOVATIONS INC. | A61B017/00 A61B017/29 A61B005/042 A61B005/06 A61B017/02 A61B017/122 A61B017/128 A61B018/02 A61B018/14 A61B018/20 A61B018/00
US20110137143 | A1 | 2011-06-09 | SENSOR WITH FINE NEEDLE HAVING CHANNEL FORMED THEREIN | 2009-07-16 | This sensor with a fine needle having a channel formed therein includes a fine needle portion of resin provided with a channel for sucking body fluid, a liquid sump portion, connected with the channel, pooling the body fluid, a sensor portion for detecting a specific component contained in the body fluid pooled in the liquid sump portion, a cover portion, in which the sensor portion is arranged, arranged to cover the liquid sump portion and an air hole connected with the channel formed in the fine needle portion through the liquid sump portion. | 1. A sensor with a fine needle having a channel formed therein, comprising: a fine needle portion of resin provided with a channel for sucking body fluid;a liquid sump portion, connected with said channel, pooling said body fluid;a sensor portion for detecting a specific component contained in the body fluid pooled in said liquid sump portion;a cover portion, in which said sensor portion is arranged, arranged to cover said liquid sump portion; andan air hole connected with the channel formed in said fine needle portion through said liquid sump portion. 2. The sensor with a fine needle having a channel formed therein according to claim 1, so formed as to introduce said body fluid into said liquid sump portion through said channel by capillarity. 3. The sensor with a fine needle having a channel formed therein according to claim 1 or 2, wherein said channel is so formed that the side surface of at least part of said channel is opened to the atmosphere. 4. The sensor with a fine needle having a channel formed therein according to claim 1, wherein the resin of said fine needle portion is biocompatible resin. 5. The sensor with a fine needle having a channel formed therein according to claim 4, wherein said biocompatible resin is polylactic acid. 6. The sensor with a fine needle having a channel formed therein according to claim 1, wherein said sensor portion is that for detecting glucose. 7. The sensor with a fine needle having a channel formed therein according to claim 1, wherein an end portion of said air hole on the side opened to the atmosphere is arranged in the vicinity of said liquid sump portion. 8. The sensor with a fine needle having a channel formed therein according to claim 1, wherein said liquid sump portion is provided in an outer shape having no angular portion on a suction side in plan view. 9. The sensor with a fine needle having a channel formed therein according to claim 1, wherein said sensor portion includes a detection electrode for detecting the specific component contained in said body fluid and an introduction confirmation electrode provided in the vicinity of an end portion of said liquid sump portion opposite to the side connected with said channel for confirming that said body fluid has been introduced into said liquid sump portion. 10. The sensor with a fine needle having a channel formed therein according to claim 1, further comprising a body portion provided with said liquid sump portion, wherein said body portion and said cover portion are made of the same type of resin material, and compression-bonded to each other without employing an adhesive. 11. The sensor with a fine needle having a channel formed therein according to claim 1, further comprising a body portion provided with said liquid sump portion, wherein said body portion and said cover portion are bonded to each other with a double-faced adhesive tape. 12. The sensor with a fine needle having a channel formed therein according to claim 1, further... | LIGHTNIX, iNC. Sysmex Corporation | A61B005/157
US20110125038 | A1 | 2011-05-26 | SYSTEM AND METHOD FOR HEART FAILURE PREDICTION | 2010-05-18 | Monitoring a health status of a human subject at defined intervals using a questionnaire. The questionnaire provides a standard script for capture, including data from a patient entered on a visual analog scale (VAS). A tool accurately assesses entries on the VAS for use by an algorithm. The captured data is input into a computer together with a physician assessment of health symptoms of the subject and provided to the algorithm for an assessment of a risk of acute heart failure of the subject. The risk is computed using the algorithm and captured data from a plurality of defined intervals. In one method, the health status of the subject as being either improved or worsening is output to the physician as a function of a value of the computed risk. In another method, a survival function outcome for the subject is predicted using the output of the algorithm. | 1. A method for monitoring a health status of a human subject, comprising the steps of: capturing medical data concerning the health of the subject at defined intervals using a questionnaire, wherein the questionnaire provides a standard script for data capture, wherein a first part of the questionnaire has a first portion of the captured data on a Likert scale and a second portion on a visual analog scale, and wherein a second part of the questionnaire has the captured data in the form of an assessment by a physician of health symptoms of the subject;inputting the captured data into a computer;extracting a calibrated value corresponding to the second portion of the captured data from the visual analog scale using a tool;providing the captured data and the calibrated value to an algorithm configured to assess a risk of acute heart failure;computing the risk of acute heart failure by executing the algorithm in the computer using the captured data input from a plurality of the defined intervals and using calibrated values associated with each of the plurality of the defined intervals; andoutputting to the physician the health status as being either improved or worsening as a function of a value of the computed risk. 2. The method of claim 1, wherein the intervals are spaced by unequal amounts of time. 3. The method of claim 2, wherein the first interval is about six hours from a first data capture, a second interval is about six hours from the first interval, a third interval is about twelve hours the second interval, a fourth interval is about twenty-four hours from the third interval, and wherein a plurality of successive intervals are twenty-four hours apart. 4. The method of claim 1, wherein the standard script causes a sequential capture of medical data. 5. The method of claim 1, wherein the capturing and inputting steps are performed concurrently through an electronic data form at or after each defined interval. 6. The method of claim 5, wherein the extracting step is performed by code executing in the computer that configures the computer to determine directly or indirectly a value on the visual analog scale that corresponds to a location marked by the human subject. 7. A method of predicting a survival function outcome of a human subject, comprising the steps of: capturing medical data concerning the health of the subject at defined intervals using a questionnaire, wherein the questionnaire provides a standard script for data capture, wherein a first part of the questionnaire has a first portion of the captured data on a Likert scale and a second portion on a visual analog scale, and wherein a second part of the questionnaire has the captured data in the form of an assessment by a physician of health symptoms of the subject;inputting the captured data into a computer;extracting a calibrated value corresponding to the second portion of the captured data from the visual analog scale using a tool;providing the captured data and the calib... | Momentum Research Inc. | A61B005/02
US20110152718 | A1 | 2011-06-23 | DEVICE AND METHOD FOR EVALUATING THE HYDRIC LOSS OF A PERSON OR AN ANIMAL BY SWEATING | 2010-12-21 | A device for evaluating the hydric loss of a person or an animal by sweating is provided with: a support comprising a measuring face and a contact face designed to be applied on the skin of the person or animal,a swelling element by absorption of sweat, securedly affixed to the support, said swelling element being arranged to be in direct contact or in fluidic communication with the skin and,a measuring instrument to measure the increase of at least one dimension of the swelling element. Increase of the dimension of the swelling element is caused by absorption of said sweat.;A method for evaluating the hydric loss of a person or an animal using such a device is also described. | 1. A device for evaluating the hydric loss of a person or an animal by sweating, provided with: a support comprising a measuring face and a contact face designed to be applied to the skin of the person or animal,a swelling element, secure affixed to the support, that swells by absorption of sweat, said swelling element being arranged to be in direct contact or in fluidic communication with the skin and,a measuring instrument for measuring the increase of at least one dimension of the swelling element, said increase being caused by absorption of said sweat. 2. The device according to claim 1, wherein the swelling element comprises a hydrophilic polymer. 3. The device according to claim 2, wherein the hydrophilic polymer is chosen from synthetic hydrophilic polymers, natural polysaccharides, semi-synthetic polysaccharides and derivatives of cellulose and their salts. 4. The device according to claim 2, wherein the hydrophilic polymer is chosen from polyvinylpyrrolidones, polymers derived from acrylic and methylacrylic acids and their salts, and amino acid polymers. 5. The device according to claim 1, wherein the swelling element comprises swelling stimulation agents of said element. 6. The device according to claim 1, wherein the measuring instrument is chosen from photodetectors, capacitive sensors, inductive sensors, mechanical position sensors and sensors of strain gage type. 7. The device according to claim 1, comprising an armature situated on the measuring face and wherein at least a part of the measuring instrument is supported by said armature. 8. The device according to claim 1, wherein the contact face comprises a drainage layer of the sweat to the swelling element. 9. The device according to claim 1, wherein the support comprises a through hole opening out onto the contact face and onto the measuring face in which the swelling element is housed and wherein the measuring instrument is arranged at the level of the measuring face. 10. The device according to claim 9, wherein the swelling element is housed and stressed in the through hole so that an increase of a dimension of the swelling element makes said element protrude out from the through hole and wherein the measuring instrument is arranged above the through hole. 11. The device according to claim 10, wherein the measuring face comprises a confinement element of the swelling element salient from the through hole. 12. The device according to claim 11, wherein the confinement element is formed by a deformable membrane fixed to the measuring face and covering at least a part of the swelling element. 13. The device according to claim 11, wherein the confinement element is formed by a rigid or semi-rigid layer arranged above the swelling element which operates in conjunction with guiding elements salient from the surface of the measuring face and wherein said layer is mobile along the guiding elements. 14. A method for evaluating the hydric loss of a person or an animal by swe... | COMMISSARIAT A L'ENERGIE ATOMIQUE ET AUX ENERGIES ALTERNATIVES | A61B005/00
US20110137185 | A1 | 2011-06-09 | Saddle Faced Small Animal Sensor Clip | 2010-05-21 | A noninvasive photoplethysmographic sensor platform for small animals provides a spring biased sensor clip, wherein at least one side of the sensor clip is provided with a saddle faced clip face member. The saddle shape of the clip face member may have a hinge end, toward the hinge of the clip, which is longer in the longitudinal direction of the clip and shorter in depth than the distal end. The shorter distal end side (measured longitudinally) of this saddle shape facilitates the ability to align the transmitted and received light with the bone, while the overall saddle shape of the facing provides a physical grip to capture enough tissue to prevent the clip from relocating over time while it is attached to the limb. The shorter hinge end side (depth-wise) also allows the clip to close when it is fully assembled. The edges of the saddle shaped clip are preferably rounded off to prevent the contusions of the tissue that may result from long-term contact. This saddle faced feature works well on only one side of the clip. | 1. A noninvasive sensor platform for small animals includes sensor clip, wherein at least one side of the sensor clip is provided with a saddle faced clip member having a sensor coupled thereto and having laterally extending projections at distal ends of the clip member, wherein each laterally extending projection has a generally concave face. 2. The noninvasive sensor platform according to claim 1 wherein the sensor clip includes a pivot. 3. The noninvasive sensor platform according to claim 2 wherein saddle faced clip face member is provided only on only one side of the clip. 4. The noninvasive sensor platform according to claim 3 further including an integral diffuser into the clip face member which is aligned with one of the transmission or receiving source on the clip, wherein the diffuser includes a substantially encapsulated diffuser material provided within a diffuser pocket formed in the clip face member. 5. The noninvasive sensor platform according to claim 4 wherein the laterally extending projection at the pivot end of the clip face member has a shorter height than the laterally extending projection at the opposed end of the clip face member. 6. The noninvasive sensor platform according to claim 5 wherein the laterally extending projection at the pivot end of the clip face member has a longer longitudinal length from the sensor than the laterally extending projection at the opposed end of the clip face member. 7. The noninvasive sensor platform according to claim 2 further including an integral diffuser into the clip face member which is aligned with one of the transmission or receiving source on the clip, wherein the diffuser includes a substantially encapsulated diffuser material provided within a diffuser pocket formed in the clip face member. 8. The noninvasive sensor platform according to claim 7 wherein the laterally extending projection at the pivot end of the clip face member has a shorter height than the laterally extending projection at the opposed end of the clip face member. 9. The noninvasive sensor platform according to claim 8 wherein the laterally extending projection at the pivot end of the clip face member has a longer longitudinal length from the sensor than the laterally extending projection at the opposed end of the clip face member. 10. The noninvasive sensor platform according to claim 2 wherein the laterally extending projection at the pivot end of the clip face member has a shorter height than the laterally extending projection at the opposed end of the clip face member. 11. The noninvasive sensor platform according to claim 10 wherein the laterally extending projection at the pivot end of the clip face member has a longer longitudinal length from the sensor than the laterally extending projection at the opposed end of the clip face member. 12. The noninvasive sensor platform according to claim 2 wherein the laterally extending projection at the pivot end of the clip face member has a longer longitu... | STARR LIFE SCIENCES CORP. | A61B005/02 A61B005/00
US20110144510 | A1 | 2011-06-16 | METHODS TO IDENTIFY DAMAGED OR SCARRED TISSUE BASED ON POSITION INFORMATION AND PHYSIOLOGICAL INFORMATION | 2009-12-16 | An exemplary system includes one or more processors; memory; and control logic, of one or more modules operable in conjunction with the one or more processors and the memory, to acquire myocardial potential data associated with position information, acquire myocardial electrical activation data associated with position information, acquire myocardial position data with respect to time, generate isopotential contours based on the potential data, generate isochronal contours based on the electrical activation data, generate isomotion contours based on the position data with respect to time, and overlay the generated isopotential contours, isochronal contours and isomotion contours on a display to indicate a region of myocardial damage or myocardial scarring with respect to a map that comprises anatomical markers. Various other methods, devices, systems, etc., are also disclosed. | 1. A system comprising: one or more processors;memory; andcontrol logic to acquire myocardial potential data associated with position information,acquire myocardial electrical activation data associated with position information,acquire myocardial position data with respect to time,generate isopotential contours based on the potential data,generate isochronal contours based on the electrical activation data,generate isomotion contours based on the position data with respect to time, andoverlay the generated isopotential contours, isochronal contours and isomotion contours on a display to indicate a region of myocardial damage with respect to a map that comprises anatomical markers. 2. The system of claim 1 wherein the control logic to overlay comprises control logic to relatively weight the isopotential contours, the isochronal contours and the isomotion contours to increase or decrease their respective contribution to the overlay on the display to more accurately indicate a region of myocardial damage with respect to a map that comprises anatomical markers. 3. The system of claim 1 further comprising control logic to render adjustable controls to a display to individually weight the isopotential contours, the isochronal contours and the isomotion contours to increase or decrease their respective contribution to the overlay on the display to more accurately indicate a region of myocardial damage. 4. The system of claim 1 further comprising control logic to select only isopotential contours that comprise values less than a predetermined value to more accurately indicate a region of myocardial damage. 5. The system of claim 1 further comprising control logic to select only isochronal contours that comprise values greater than a predetermined value to more accurately indicate a region of myocardial damage. 6. The system of claim 1 further comprising control logic to select only isomotion contours that comprise values less than a predetermined value to more accurately indicate a region of myocardial damage. 7. The system of claim 1 further comprising control logic to select only some of the contours based on one or more predetermined values andto weight the selected contours to increase or decrease their respective contribution to the overlay on the display to more accurately indicate a region of myocardial damage. 8. The system of claim 1 further comprising control logic to indicate myocardial damage by outlining a scar region with respect to a map that comprises anatomical markers. 9. The system of claim 1 further comprising control logic to acquire fractionation data associated with position information, generate isofractionation metric contours based on the fractionation data and overlay the generated isofractionation metric contours and the generated isomotion contours on a display to indicate a region of myocardial damage with respect to a map that comprises anatomical markers wherein the fractionation data comprises data selected from ... | PACESETTER, INC. | A61B005/0452
US07967752 | B2 | 2011-06-28 | Sensor system as well as an arrangement and method for monitoring a constituent and in particular glucose in body tissue | 2006-06-19 | The invention generally relates to a sensor system. In particular to a sensor for glucose monitoring. The invention also also relates to an arrangement and a method for monitoring a constituent and in particular glucose in body tissue using a sensor system. | 1. A sensor system especially for glucose monitoring, the sensor system comprising: a support that can be at least partially inserted into body tissue;a fluid channel integrated on the support for passing through an interstitial fluid;a sensor that responds to a constituent of the body tissue, wherein the sensor is integrated on the support and configured for direct contact with the body tissue separately from the fluid channel;a signal processor electrically coupled to the sensor for generating a measurement result;a calibration device configured to be located outside the body tissue to calibrate an output signal of the sensor, the calibration device having a reference sensor that is fluidly coupled to the fluid channel for passing through the interstitial fluid and electrically coupled to the signal processor; anda pumping unit for removing interstitial fluid not in contact with the sensor through the fluid channel and into the calibration device for use in calibrating the sensor. 2. The sensor system according to claim 1, wherein the sensor is arranged on a surface of the support which, when inserted, is in contact with the body tissue. 3. The sensor system according to claim 2, wherein the sensor in a multielectrode configuration has several electrodes arranged on the surface of the support which engages with the body tissue. 4. The sensor system according to claim 1, wherein the sensor responds directly to the constituent independently of the amount of the locally present tissue fluid. 5. The sensor system according to claim 1, wherein the sensor is in direct engagement with the body tissue at a distance axial from an opening of the fluid channel on an active sensor region side. 6. The sensor system according to claim 1, wherein an interstitial fluid and/or a calibration solution can be fed via the fluid channel to the calibration device for the sensor that is located outside of the body tissue. 7. The sensor system according to claim 1, wherein the fluid channel is connected to a metering unit to feed a medication liquid. 8. The sensor system according to claim 7, wherein the medication liquid is insulin. 9. The sensor system according to claim 1, wherein the sensor has an active sensor region that can be brought into contact with the body tissue. 10. The sensor system according to claim 1, wherein the sensor is designed to continuously detect glucose as a constituent in the interstitial fluid. 11. The sensor system according to claim 1, wherein the support is formed by a flexible flat substrate wherein the fluid channel and the sensor are designed as a thin layer structure integrated on the substrate as a composite part. 12. The sensor system according to claim 11, wherein the flexible support can be inserted into the body tissue by means of an insertion aid. 13. The sensor system according to claim 1, wherein the fluid channel has at least one opening on an active sensor region side, at least one opening located outside of the... | Roche Diagnostics Operations, Inc. | A61B005/00
US20110144525 | A1 | 2011-06-16 | METHOD AND DEVICE FOR CHARACTERIZING THE EFFECT OF A SKIN TREATMENT AGENT ON SKIN | 2008-06-18 | In a method for characterizing skin treatment agent, a device having several sets of electrodes is applied to the skin. The electrode sets have differing electrode distances, such that fields having different reach can be generated. Inverse profiling is used to calculate the dielectric permittivities of individual skin layers, which in turn allows to observe the water transport mechanism in the skin. These transport mechanisms can be used to assess the effect of the agent on the skin. An advantageous device for implementing this method comprises coplanar waveguides for generating the fields. | 1. A method for characterizing an effect of a skin treatment agent on skin comprising the steps of (a) applying the skin treatment agent to a skin region,(b) applying a measuring device to said skin region, said measuring device having several sets of electrodes, wherein each set comprises at least two electrodes and wherein the electrodes of each set of electrodes have a distance Wi from each other and wherein there are N>1 sets having different distances Wi,(c) generating, by means of said different sets of said electrodes, at least a N electrical fields within said skin region, said electrical fields having differing penetration depths into said skin region,(d) measuring at least N measured parameters mi wherein each measured parameter mi depends on an effective permittivity seen a different one of said N electrical fields(e) calculating, from said measured parameters mi at least one characterizing parameter descriptive of a permittivity of said skin region for a given depth. 2. The method of claim wherein step (e) further comprises calculating, from said measured parameters mi, several characterizing parameters pi, in particular N characterizing parameters pi, for differing depths of said skin region. 3. The method of claim 2, wherein each characterizing parameter pi, depends on a permittivity ∈i of a single layer i of said skin region. 4. The method of claim 2, wherein each characterizing parameter pi is the permittivity ∈i of a single layer i of said skin region. 5. The method of claim 2, wherein each characterizing parameter pi is the water content of a single layer i of said skin region. 6. The method of claim 1, wherein said measuring device comprises at least one set of electrodes having a mutual distance 0.1 mm or less. 7. The method of claim 1, wherein said measuring device comprises at least one set of electrodes having a mutual distance of at least 0.1 mm. 8. The method of claim 7, wherein said measuring device comprises at least one set of electrodes having a mutual distance between 0.1 mm and 0.2 mm. 9. The method of claim 1, wherein said measuring device comprises at least one set of electrodes having a mutual distance of at least 1 mm. 10. The method of claim 1, wherein the applied electrical fields have a frequency between 50 MHz to 100 GHz. 11. The method of claim 1, comprising the steps of repetitively determining, at different times after application of the skin treatment agent, the at least one characterizing parameter. 12. The method of claim 1, comprising the step of measuring an optical reflection or transmission of the skin region. 13. The method of claim 1, comprising the step of measuring a temperature of the skin region, and in particular where a moisture content of a layer i of said skin region is calculated from a permittivity ∈i of said layer i and said temperature. 14. The method of claim 1, comprising the step of measuring a conductivity of said skin region for a frequency below 1 MHz. 15. The meth... | - | A61B005/053
US07959583 | B2 | 2011-06-14 | Integrated lancet and bodily fluid sensor | 2007-02-09 | A lancing device integrated with a sensing strip and method of making and using the same are provided. A lance and analyte sensing strip are combined in a single, preferably disposable, device. In a preferred embodiment, a slot is provided in a support member so the lance can be engaged by a lance driver and driven into the skin producing a bodily fluid sample from a wound. The sensing area of the sensing strip is adjacent to the wound location reducing dexterity requirements for patients. | 1. A cartridge containing two or more bodily fluid sample acquisition and testing devices, each of said acquisition and testing devices comprising: a lancet adapted to be engaged by a lancing device;a bodily fluid chemistry sensing strip comprising a sensing area; anda primary support member adapted to fixedly engage said sensing strip, and to slidably engage said lancet, said primary support member further comprising a slot adapted to allow contact between said lancet and said lancing device, said lancing device being adapted to slide said lancet in the direction of said slot between an unexposed position and an exposed position. 2. A cartridge of claim 1, wherein the sensing area is in electrical contact with two electrodes. 3. A cartridge as in claim 1, wherein said sensing area is adjacent to said exposed position of said lancet. 4. A cartridge as in claim 1, wherein said sensing strip comprises a wick. 5. A cartridge as in claim 1, wherein said sensing strip is adapted to sense bodily fluid having a volume less than one microliter. 6. A cartridge as in claim 1, wherein said sensing strip is adapted to sense bodily fluid having a volume less than one half microliter. 7. A cartridge as in claim 1, further comprising a spring adapted to retain said plurality of acquisition and testing devices against a proximal end of said cartridge. 8. A cartridge as in claim 7, further comprising a first slot aligned with a proximal one of said plurality of acquisition and testing devices. 9. A cartridge as in claim 8, further comprising a second slot aligned with a proximal one of said plurality of acquisition and testing devices, wherein said first slot is further aligned with said slot is said proximal one of said plurality of acquisition and testing devices, and said second slot is adapted to permit said proximal one of said plurality of acquisition and testing devices to be ejected from said cartridge. 10. A cartridge containing two or more bodily fluid sample acquisition and testing devices, each of said devices comprising: a primary support member having a bodily fluid chemistry sensing strip attached thereto, said sensing strip comprising a sensing area in proximity to a sensing end of said primary support member;a cantilever member comprising a first end hinged to said support member;a lancing member comprising a lancet support member and a lancet having a pointed end attached to said lancet support member, said lancet support member being hinged to a second end of said cantilever member;said primary support member, cantilever member, and lancing member adapted to adopt a first orientation being substantially coplanar, and to adopt a second orientation with said lancing member defining an angle with said primary support member such that said pointed end of said lance is adjacent to said sensing area of said sensing strip. 11. The cartridge of claim 10, wherein said sensing area is in electrical communication with two electrodes. 12. The ca... | Becton, Dickinson and Company | A61B005/00
US07966058 | B2 | 2011-06-21 | System and method for registering an image with a representation of a probe | 2003-12-31 | A system and method is provided for registering a representation of a probe with an image. One embodiment of a method comprises acquiring an image of or pertaining to a heart and registering a representation of a probe which is in or adjacent to the heart with the image using a heart vector of the heart. | 1. A method comprising: acquiring an image of or pertaining to a heart;acquiring a first data set pertaining to one or more locations of a heart vector of the heart, the first data set being spatially correlated with the image;acquiring a second data set pertaining to one or more locations of the heart vector of the heart;receiving first and second data sets at a processor; andregistering, via the processor, a representation of a probe with the image by registering the location of the heart vector from the first data set with the location of the heart vector from the second data set, wherein the second data set is acquired using at least one lead positioned on a skin surface, wherein the location of the heart vector from the second data set can be determined relative to the lead, and wherein the location of the probe can also be determined relative to the lead. 2. The method of claim 1, wherein the image comprises one or more images obtained using computed tomography, magnetic resonance, or ultrasound. 3. The method of claim 1, wherein the acquiring the second data step and the registering step are performed on a repeating basis. 4. A method comprising: acquiring an image of or pertaining to a heart;acquiring a first and a second data set using a lead system;receiving the first and second data sets at a processor;registering, via the processor, a location of a first heart vector from the first data set relative the lead system at a skin surface of an imaged subject, wherein the first heart vector represents a summation of electrical currents at a particular time, the summation having a direction and an amplitude;registering, via the processor, a location of a second heart vector from the second data set relative to the lead system; andadjusting the size or position of the image dependent on a change in the location between the first and second heart vector generated from the first and second data sets, respectively. 5. The method of claim 4, further comprising registering a representation of a probe with an image, the probe being located in or adjacent to a heart. 6. The method of claim 4, wherein the image is correlated to a first heart vector data set and the image is adjusted by comparing the first heart vector data set to a second heart vector data set. 7. A system comprising: a lead system located at a skin surface of an imaged subject and operable to acquire a first data set and a second data set pertaining to one or more locations of a first and second heart vector, respectively, of the heart;a processor configured to be communicatively coupled to a probe and further configured to register the first heart vector from the first data set with the second heart vector from the second data set, the probe being configured to be located in or adjacent to a heart;memory configured to store:an image of at least a portion of the heart;the first data set pertaining to one or more locations of the first heart vector of the heart, the first data ... | General Electric Company | A61B005/055 A61B008/00 G06K009/32
US07946985 | B2 | 2011-05-24 | Method and system for providing sensor redundancy | 2006-12-29 | A closed loop system or semi-closed loop system for infusing insulin using sensor values applies a redundant sensor system as a fail-safe method against sensor failure. The redundant glucose sensors are used corroborate each other and a failing sensor is detected if the sensors no longer corroborate each other. The use of redundant sensors has the additional benefit of producing better sensor signals compared to the use of a single sensor. | 1. A method of sensing blood glucose values using redundant sensors, comprising: obtaining a first glucose reading from a first glucose sensor located at a first site;obtaining a second glucose reading from a second glucose sensor located at a second site;deriving a first predictive value for the first glucose sensor corresponding to the first glucose reading using the second glucose reading as an input;deriving a second predictive value for the second glucose sensor corresponding to the second glucose reading using the first glucose reading as an input;determining a first error between the first predictive value and the first glucose reading;determining a second error between the second predictive value and the second glucose reading;comparing a sum of the absolute error values of the first and second errors to a threshold;determining a failing sensor exists if the sum of the absolute error values exceed the threshold;determining whether the first glucose sensor or second glucose sensor has the least error in the sensor signal if the sum of the absolute error values does not exceed the threshold;calculating a blood glucose value based on the glucose sensor having the least error in the sensor signal; andreporting the calculated blood glucose value. 2. The method of claim 1, further comprising: requesting a meter glucose value if the sum of the absolute values exceed the threshold; andcomparing the meter glucose value to the first and second glucose readings to determine if the first or second glucose sensor is failing. 3. The method of claim 1, wherein the first glucose sensor and the second glucose sensor use different types of technology. 4. The method of claim 1, wherein the first site is subcutaneous tissue, and the second site is blood plasma. 5. The method of claim 1, wherein the first site and the second site is subcutaneous tissue. 6. The method of claim 1, wherein the second sensor acts as a watchdog for the first sensor. 7. The method of claim 1, wherein the steps of deriving a first predictive value and deriving a second predictive value uses a first and a second adaptive filter. 8. The method of claim 1, wherein the reported blood glucose value is calculated based on the predictive value for the glucose sensor having the least error in the sensor signal. 9. The method of claim 1, wherein the reported blood glucose value is calculated based on the glucose reading for the glucose sensor having the least error in the sensor signal. 10. The method of claim 1, wherein the reported blood glucose value is calculated based on the glucose sensor having the least error in the sensor signal and a calibration factor for calibrating the sensor signal from the glucose sensor having the least error in the sensor signal. 11. A closed loop system or semi-closed loop system for infusing insulin using sensor values comprising: an infusion pump delivering insulin based on a reported blood glucose value derived by a controller configured to per... | Medtronic MiniMed, Inc. | A61B005/00
US07966063 | B2 | 2011-06-21 | Body composition measuring apparatus | 2006-03-09 | There are provided a hold portion that can be held in a hand of a subject, a cable having one end connected to the hold portion, and a rest portion on which a foot of the subject can be rested. The rest portion includes a reel portion connected to the other end of the cable for winding the cable, a first detection unit for detecting a draw-out extent to which the cable is drawn out from the reel portion, and a second detection unit for detecting an angle formed by the rest portion and the cable. Based on the detected draw-out extent and angle, the part length of a body part of the subject is calculated. | 1. A body composition measuring apparatus comprising: a plurality of electrodes to be in contact with a plurality of predetermined parts, associated respectively with said electrodes, of the body of a subject;a hold portion that can be held in a hand of said subject, the hold portion comprising a display unit to display body composition data;a cable having one end connected to said hold portion;a rest portion on which a foot of said subject can be rested,said rest portion comprising,a receptacle for receiving the hold portion, such that when the hold portion is received in the receptacle the hold portion is substantially encased in the rest portion,a reel portion connected to the other end of said cable for winding said cable,a first detection unit for detecting a draw-out extent to which said cable is drawn out from said reel portion, anda second detection unit for detecting an angle formed by said rest portion and said cable proximate to a top surface of the rest portion;said body composition measuring apparatus further comprisinga first part-length calculation unit for calculating the part length of a body part of said subject based on said draw-out extent detected by said first detection unit and said angle detected by said second detection unit;a first measurement unit for measuring impedance of said body part by using said plurality of electrodes; anda first body composition calculation unit for calculating body composition of said body part based on said calculated part length and said measured impedance. 2. The body composition measuring apparatus according to claim 1, further comprising: an operation unit operated by the subject; anda sense unit for sensing an instruction signal from said operation unit, whereinsaid first part-length calculation unit calculates said part length of the body part when said sense unit senses the instruction signal for instructing to start calculation of the part length. 3. The body composition measuring apparatus according to claim 1, wherein said plurality of electrodes includeelectrodes for hand provided to said hold portion and including at least a pair of an electrode for current and an electrode for voltage, andelectrodes for foot provided to said rest portion and including at least a pair of said electrode for current and said electrode for voltage. 4. The body composition measuring apparatus according to claim 3, wherein said body part is arm portion, andsaid first part-length calculation unit calculates the part length of said arm portion by using a predetermined expression based on said draw-out extent and said angle. 5. The body composition measuring apparatus according to claim 3, further comprising: a second part-length calculation unit for calculating the part length of a leg portion of said subject based on said draw-out extent detected by said first detection unit;a second measurement unit for measuring the impedance of said leg portion by using said electrodes for foot and said electrode... | OMRON HEALTHCARE Co., Ltd. | A61B005/05
US07957795 | B2 | 2011-06-07 | Abdominal impedance measurement apparatus | 2008-03-26 | An abdominal impedance measurement apparatus includes a plurality of electrodes for measuring an abdominal impedance of a human subject, and includes an electrode supporting member for supporting the electrodes in such a manner that the electrodes protrude from the electrode supporting member. The electrode supporting member includes a plurality of segments aligned in a direction, the electrodes being respectively mounted on different segments, neighboring segments being connected rotatably with each other. The electrode supporting member further includes rotation-angle restricting parts for restricting relative rotation-angles between neighboring segments. | 1. An abdominal impedance measurement apparatus comprising: a plurality of electrodes for measuring an abdominal impedance of a human subject; andan electrode supporting member for supporting the electrodes in such a manner that the electrodes protrude from the electrode supporting member, the electrode supporting member comprising a plurality of segments aligned in a direction, each single electrode being mounted on a segment that is different from every other segment on which another electrode is mounted, such that no two electrodes are mounted on the same segment, neighboring segments being connected rotatably with each other, the electrode supporting member comprising rotation-angle restricting parts for restricting relative rotation-angles between neighboring segments,wherein the plurality of electrodes comprise two current-supplying electrodes adapted to contact an abdomen of the human subject for applying an electrical current to the abdomen of the human subject and two voltage-measurement electrodes adapted to be brought into contact with two parts of the abdomen of the human subject for measuring a voltage between the parts with which the voltage-measurement electrodes are in contact, the current-supplying electrodes being mounted on two of the segments of the electrode supporting member, the voltage-measurement electrodes being mounted on two of the segments of the electrode supporting member that are different from the segments on which the current-supplying electrodes are mounted,wherein at least one segment on which no electrode is mounted exists between the segment on which one of the current-supplying electrodes is mounted and the segment on which one of the voltage-measurement electrodes is mounted, andwherein the rotation-angle restricting parts of the electrode supporting member are provided such that the rotation-angle restricting parts located nearer a center of the electrode supporting member restrict a relative rotation-angle between neighboring segments within a larger range and such that the rotation-angle restricting parts located farther from the center of the electrode supporting member restrict a relative rotation-angle between neighboring segments within a smaller range. 2. The abdominal impedance measurement apparatus of claim 1, wherein segments each on which no electrode is mounted exist at all positions between neighboring segments on which the electrodes are mounted. 3. The abdominal impedance measurement apparatus of claim 1, wherein neighboring segments are rotatably jointed by pins, and the rotation-angle restricting parts are adapted to restrict relative rotation-angles about the pins between neighboring segments. 4. The abdominal impedance measurement apparatus of claim 1, wherein each of the electrodes comprises a cylindrical part protruding from the electrode supporting member and a convex curved surface adapted to be brought into contact with the abdomen of the human subject, the convex curved surface ... | Tanita Corporation | A61B005/05
US07953467 | B2 | 2011-05-31 | Method for non-invasive cancerous tissue diagnosis and tomography using terahertz imaging | 2009-12-02 | The primary objective of the present method and apparatus is to provide a portable and new diagnosis system for quickly and reliably examining tissue conditions. The method uses the most advance miniaturized micro-opto-electro-mechanical systems (MOEMS) for generating a rapid variable optical delay line capable of generating wideband terahertz pulses. The method detects and analyzes cancerous tissues by comparing a plurality of spectrum resolved images of suspected tissue without applying harmful agents into the tissue to facilitate interaction with illumination sources. The method employs non-evasive, real time terahertz imaging systems and techniques to diagnose tissue for detecting the presence of cancer. A map showing, which tissue is healthy and which is cancerous can aid in the accurate removal of cancerous tissue. | 1. A method for diagnosis of tissue, comprising: a) transmitting femtosecond pulsed laser light through a fiber carrier;b) splitting the laser light into a pump light and a reference light to create an optical path;c) inserting into the optical path a micro-opto-electro-mechanical system (MOEMS) rapid scanning delay line to generate long delays and to enhance sampling rate;d) inserting into the optical path a quasi optics chopper coupled to the MOEMS rapid scanning delay line to improve signal processing;e) inserting into the optical path at the output of the chopper a terahertz transmitter; andf) illuminating the tissue with pulses of terahertz waves coming from the terahertz transmitter for diagnosis of tissue. 2. The method of claim 1, further comprising using a sample stepper dynamically synchronized to the terahertz waves that are illuminating a tissue sample. 3. The method of claim 2, further comprising using a terahertz detector to detect a difference signal between a reflected signal from the sample and a reference signal from the femtosecond pulsed laser. 4. The method of claim 3, further comprising using matching amplifiers to improve the detected difference signal. 5. The method of claim 3, further comprising forming an image from the reflected signal at each tissue layer parallel to the tissue surface. 6. The method of claim 3, further comprising: a) generating images to represent the sample using the difference signal;b) comparing the images with a calibrated reference stored in memory;c) combining the images at different layers to obtain the tomography of the tissue;d) indicating regions of coincidence and regions of non-coincidence; ande) showing the result in a control and display system. 7. The method of claim 6, further providing a three dimensional image of the tissue in real time which includes compositional information about the tissue. 8. The method of claim 7, wherein tissue image is comprised of a plurality of horizontal frequency bands, each band being adjacent to another, with equal bandwidths. 9. The method of claim 8, wherein each horizontal frequency band consists of a plurality of pixels, each being adjacent to another. 10. The method of claim 3, wherein the detector is identical to the transmitter and wherein the terahertz antenna is designed to receive terahertz frequency radiation reflected from the tissue. 11. The method of claim 3 wherein the detector acts as a convolver to detect the reflected terahertz pulses from the tissue, which arrive at detector, synchronized to the split reference beam. 12. The method of claim 1, wherein the pulses of terahertz waves have duration from 10-1000 femtoseconds to cover terahertz bandwidth from 100-10000 GHz. 13. The method of claim 1, wherein the MOEMS rapid scanning delay line generates high speed delay line patterns using a blazed grating and a bounced mirror combined with a MOEMS scanner. 14. The method of claim 13, wherein the quasi optics chopper is a minia... | - | A61B005/00
US07957792 | B2 | 2011-06-07 | Spatial resolution determination for cardiac mapping | 2010-06-01 | A non-contact cardiac mapping method is disclosed that includes: (i) inserting a catheter into a heart cavity having an endocardium surface, the catheter including multiple, spatially distributed electrodes; (ii) measuring signals at the catheter electrodes in response to electrical activity in the heart cavity with the catheter spaced from the endocardium surface; and (iii) determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface. Related systems and computer programs are also disclosed. | 1. A method comprising: inserting a catheter into a heart cavity having an endocardium surface, the catheter comprising multiple, spatially distributed electrodes;measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the catheter electrodes with respect to the endocardium surface;computing values indicative of a degree of spatial resolution of the determined physiological information for at least some of the locations of the endocardium surface; anddisplaying information related to at least some of the determined physiological information and information related to at least some of the computed resolution values. 2. The method of claim 1, wherein the computed values are derived, at least in part, from a transformation function relating the physiological information at the multiple locations of the endocardium surface to the signals measured by the catheter electrodes. 3. The method of claim 1, further comprising: moving the catheter to each of multiple, different positions in the heart cavity; andfor each of the different catheter positions, determining the positions of the catheter electrodes with respect to the endocardium surface and measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;wherein the determination of the physiological information at the multiple locations of the endocardium surface is based further on the positions of the catheter electrodes and the measured signals at the different catheter positions. 4. The method of claim 3, wherein determining the position of the catheter electrodes comprises measuring information about at least one of a position and orientation of the catheter within the heart cavity and determining the position of the catheter electrodes is based further on information about the distribution of the electrodes on the catheter. 5. The method of claim 3, wherein determining the position of the catheter electrodes comprises directly measuring the position of each catheter electrode within the heart cavity. 6. The method of claim 1, wherein displaying information related to at least some of the determined physiological information further comprises overlaying on a display device at least some of the physiological information determined at the multiple surface locations. 7. The method of claim 1, wherein displaying information related to at least some of the determined physiological information further comprises displaying at least a portion of the endocardium surface to include a selected subset of the physiological information determined at the multiple surface locations, wherein the subset is selected based on at least some of the computed resolution values. 8. The method of claim 1, wherein the determined physiological information comprises isopotential lines or bands correspondi... | Rhythmia Medical, Inc. | A61B005/02
US07947001 | B1 | 2011-05-24 | Methods and devices for measuring structural and elastic properties of a hollow organ | 2010-06-07 | A spatial profile of structural and elastic parameters along a lumen of a hollow organ is determined based on obtaining the functional dataset of pressure P, balloon volume V, and balloon linear coordinate X while continuously moving the balloon along the lumen. A comprehensive characterization of elastic properties of the tissue and geometrical parameters of the lumen are then determined using a previously obtained calibration curve of balloon pressure vs. balloon volume when the balloon is expanded freely. As opposed to prior art devices for measuring tissues elasticity at one or several distinct locations by having deploying a balloon locally, the present invention describes a device in which the balloon is continuously moved through the lumen of the hollow organ whereby allowing elasticity determination over a significant portion or the entire length of the organ. The relationship between P, V, and X is then described by a family of constant pressure curves F(Pi=const,V,X), in the constant pressure method; or by family of constant volume curves F(P,Vi=const,X), in the constant volume method. The handheld device of the invention supports evaluation using either method. Repeating the test at different levels of balloon pressure or volume allows obtaining of a family of curves characterizing linear distribution of wall compliance along the lumen. Methods and devices of the invention can be used to evaluate cervix, urethra, esophagus, intestines, blood vessels, and other organs. | 1. A method for determination of structural and elastic properties of a hollow organ containing a lumen traversing therethrough, the method comprising the steps of: a) providing an inflatable balloon inside said lumen,b) inflating said balloon to a predetermined first pressure,c) moving said balloon along said lumen,d) continuously inflating or deflating said balloon to maintain said first pressure at a constant level during said movement of said balloon along said lumen, ande) measuring balloon pressure, balloon volume, and balloon coordinate along said lumen during said movement,whereby a first spatial profile of said structural and elastic properties of the hollow organ along said lumen is determined using said balloon pressure, said balloon volume, and said balloon coordinate along said lumen. 2. The method as in claim 1, wherein said steps (b) through (e) are repeated at a second predetermined pressure of said balloon so as to determine a second spatial profile of said structural and elastic properties of said hollow organ. 3. The method as in claim 1, wherein said steps (a) through (e) are preceded by obtaining a calibration curve for said balloon, said calibration curve comprising a function of balloon volume vs. balloon pressure while said balloon is not restricted to expand freely. 4. The method as in claim 1, wherein said hollow organ is selected from a group consisting of cervix, urethra, intestines, blood vessels, and esophagus. 5. The method as in claim 1, wherein said properties selected from a group consisting of tissue elasticity, Young's modulus, and lumen diameter. 6. A method for determination of structural and elastic properties of a hollow organ containing a lumen traversing therethrough, the method comprising the steps of: a) providing an inflatable balloon inside said lumen,b) inflating said balloon to a predetermined first volume,c) moving said balloon along said lumen,d) maintaining said first volume at a constant level during said movement of said balloon along said lumen, ande) measuring balloon pressure and balloon coordinate along said lumen during said movement,whereby a first spatial profile of said structural and elastic properties of the hollow organ along said lumen is determined using said balloon pressure, said first balloon volume, and said balloon coordinate along said lumen. 7. The method as in claim 6, wherein said steps (b) through (e) are repeated at a second predetermined balloon volume so as to determine a second spatial profile of said structural and elastic properties of said hollow organ. 8. The method as in claim 6, wherein said steps (a) through (e) are preceded by obtaining a calibration curve for said balloon, said calibration curve comprising a function of balloon volume vs. balloon pressure while said balloon is not restricted to expand freely. 9. The method as in claim 6, wherein said hollow organ is selected from a group consisting of cervix, urethra, intestines, blood vessels, and es... | - | A61B005/117 A61B005/103
US08587310 | B2 | 2013-11-19 | Magnetic resonance imaging device | 2009-07-22 | In the diffusion-weighted imaging, amounts of distortion and amounts of phase offset of k-space data due to a temporally changing magnetic field error induced by eddy currents and vibrations associated with application of a diffusion-weighted gradient magnetic field pulse are corrected with good precision to improve image quality. Characteristic data for correcting distortion of k-space data are calculated for every position in the slice direction as peak shifts of projections observed between the cases of applying and not applying an MPG pulse. As the characteristic data, amounts of distortion in the read-out direction and the phase encoding direction and phase offset amounts in a slice plane are calculated. | 1. A magnetic resonance imaging device comprising at least one of a hardware processor and circuitry effecting: an imaging portion configured to apply a radio frequency magnetic field and a gradient magnetic field to a subject placed in a static magnetic field and to detect magnetic resonance signals generated from the subject, a calculator configured to process the magnetic resonance signals detected by the imaging portion, and a controller configured to control the imaging portion and the calculator, wherein:the imaging portion including:a diffusion-weighted imaging execution portion configured to detect the magnetic resonance signals according to a pulse sequence including a diffusion-weighted gradient magnetic field pulse, anda reference data acquisition portion configured to obtain reference data for detecting distortion amounts of k-space data caused by the diffusion-weighted gradient magnetic field pulse at an arbitrary position in a slice direction, andthe calculator including:a characteristic data calculation portion configured to calculate distortion amounts for a read-out direction and a phase encoding direction and phase offset amounts at an arbitrary position in the slice direction from the reference data as characteristic data of the amounts of distortion of k-space data,a correction portion configured to correct the k-space data constituted by the magnetic resonance signals acquired by the diffusion-weighted imaging execution portion by using the characteristic data, andan image reconstruction portion configured to reconstruct an image from the data corrected by the correction portion. 2. The magnetic resonance imaging device according to claim 1, wherein: the reference data acquisition portion is configured to obtain the reference data by executing:a sequence corresponding to the pulse sequence executed by the diffusion-weighted imaging execution portion in which pulse strength of a phase encoding gradient magnetic field pulse is 0, and the diffusion-weighted gradient magnetic field pulse is applied,a sequence corresponding to the pulse sequence in which pulse strength of a phase encoding gradient magnetic field pulse is 0, and the diffusion-weighted gradient magnetic field pulse is not applied along any of axes,a sequence corresponding to the pulse sequence in which the read-out direction and the phase encoding direction of the pulse sequence are interchanged, pulse strength of a phase encoding gradient magnetic field pulse is 0, and the diffusion-weighted gradient magnetic field pulse is applied, anda sequence corresponding to the pulse sequence in which the read-out direction and the phase encoding direction of the pulse sequence are interchanged, pulse strength of a phase encoding gradient magnetic field pulse is 0, and the diffusion-weighted gradient magnetic field pulse is not applied along any of axes. 3. The magnetic resonance imaging device according to claim 2, wherein: the characteristic data calculation portion is co... | Hitachi Medical Corporation | G01V003/00
US08986207 | B2 | 2015-03-24 | Systems and methods for providing sensor arrays for detecting physiological characteristics | 2010-11-12 | Systems and methods for determining physiological parameters of a subject using a sensor array. In an embodiment, a sensor array may contain sensor elements for determining multiple physiological parameters. A combination of sensor elements and the physiological parameters determined may be selected based on signals obtained from the sensor elements of the sensor array. A sensor array may be connected to a monitoring device that may select an optimal sensor element or combination of sensor elements and one or more physiological parameters to be determined. The monitoring device may then determine physiological parameters using the selected combination of sensor elements and display information associated with the parameters on a monitor for use, for example, in monitoring a medical patient. | 1. A method for determining one or more physiological parameters of a subject, the method comprising: receiving physiological signals from sensor elements coupled to the subject;analyzing, using a processor, the physiological signals using a first metric associated with a first physiological parameter;analyzing, using a processor, the physiological signals using a second metric associated with a second physiological parameter;identifying, using a processor, whether the first physiological parameter can be determined based at least in part on the first metric;identifying, using a processor, whether the second physiological parameter can be determined based at least in part on the second metric; anddetermining, using a processor, at least one of the first and second physiological parameters of the subject based at least in part on the physiological signals received from the sensor elements and based at least in part on whether the first and second physiological parameters can be determined. 2. The method of claim 1, further comprising: receiving an indication of a set of predefined measurement types; andidentifying the first and second physiological parameters from the set. 3. The method of claim 1, wherein: analyzing the physiological signals using the first metric comprises: correlating the physiological signals with a first model for a first available measurement type; anddetermining whether the correlation between the physiological signals and the first model exceeds a threshold for the first available measurement type; andanalyzing the physiological signals using the second metric comprises: correlating the physiological signals with a second model for a second available measurement type; anddetermining whether the correlation between the physiological signals and the second model exceeds a threshold for the second available measurement type. 4. The method of claim 1, wherein further comprising: analyzing the physiological signals using the first metric comprises: comparing the physiological signals to the first metric, anddetermining whether the physiological signals meet a threshold level for the first metric; andanalyzing the physiological signals using the second metric comprises: comparing the physiological signals to the second metric, anddetermining whether the physiological signals meet a threshold level for the second metric. 5. The method of claim 4, further comprising identifying an optimal measurement type that can be determined from the physiological signals based on the comparisons of the physiological signals to the first and second metrics. 6. A system for determining one or more physiological parameters of a subject, the system comprising: a sensor interface capable of receiving physiological signals from sensor elements coupled to the subject; anda processor coupled to the sensor interface, wherein the processor is capable of: analyzing the physiological signals using a first metric associated with a first physiological ... | Covidien LP | A61B005/00 A61B005/1455 A61B005/0205 G06F019/00 A61B005/024 A61B005/08
US08489174 | B2 | 2013-07-16 | Method to detect a breathing movement of an examination subject corresponding to signal data by magnetic resonance | 2010-11-24 | In a method for the detection of signal data corresponding to a breathing movement of an examination subject by magnetic resonance (MR) first and second data sets are loaded, that each include complex k-space data acquired with a navigator sequence from a common excitation volume of the examination subject. The first and second data sets are processed to identify breathing movement at the acquisition time of at least one of the data sets, by comprising a transformation of the data sets in Cartesian space and calculating a phase difference between respective complex data pointes of the data sets having the same spatial position. The processing result is stored together with a time value that depends on a point in time of the acquisition of the first data set and/or the second data set. The data acquisition, processing and storage are repeated until a series of results has been stored that maps a breathing movement of interest, and in repetition at least one of the two data sets is acquired at a different point in time than the last two data sets. | 1. A method to detect signal data corresponding to a breathing movement of an examination subject by magnetic resonance (MR), comprising the steps of: a) loading a first data set and second data set into a computerized processor, said first and second data sets respectively being based on complex k-space data acquired with a navigator sequence from a common excitation volume of an examination subject in a magnetic resonance data acquisition unit;b) processing the first and second data sets in said processor to obtain a processing result related to the breathing movement at the acquisition time of at least one of said first and second data sets, said processing comprising a transformation of said first and said second data set in spatial space and calculating a phase difference between at least one complex data point of the first data set and a complex data point of the second data set which is associated with the same spatial position as said data point of the first data set;c) electronically storing the processing result from step b) together with a time value that depends on at least one of a point in time of acquisition of the k-space data on which the first data set is based and a point in time of acquisition of the k-space data on which the second data is based; andd) repeating steps a) through c) until a series of processing results has been stored that maps a breathing movement of interest and, in each repetition acquiring at least either the first or second data set loaded in step a) at a different point in time than the first and second data set loaded in previous repetition of step a). 2. A method as claimed in claim 1 comprising employing a navigator sequence with which the k-space data are acquired on which the first and second data set are based that generates at least one echo and scans every echo for the acquisition of the k-space data under a readout gradient. 3. A method as claimed in claim 2 comprising processing the first and second data sets by calculating respective Fourier transformation along the readout direction of the complex k-space data of the first and second data sets and thereby obtain a new first and a new second data set where different data points in the new first data set correspond to different spatial positions along the readout direction and different data points in the new second data set correspond to different spatial positions along the readout direction. 4. A method as claimed in claim 3 comprising processing the first and second data sets obtained after the Fourier Transformation by a pairwise complex conjugate multiplication of complex numbers in the first and the second data set which correspond to the same spatial position. 5. A method as claimed in claim 2 comprising acquiring k-space data on which the first and second data sets are based respective navigator sequences using at least two acquisition coils. 6. A method as claimed in claim 5 comprising processing the k-space data acquired by said... | Siemens Aktiengesellschaft | A61B005/05
US08951261 | B2 | 2015-02-10 | Subchondral treatment of joint pain | 2010-11-19 | The embodiments provide devices and methods that both strengthen the bone and stimulate the bone. Bone fractures or non-unions are stabilized, integrated or healed, which results in reduction of a bone defect, such as a bone marrow lesion or edema. In addition, the distribution of forces in a joint are restored or altered to relieve pain. In general, a joint is evaluated by taking an image of the joint and one or more subchondral defects are detected. At least one of the subchondral defects may be diagnosed as the source of pain and an extent of treatment for the subchondral defect is determined. The disclosed devices and techniques are particularly suited for treating chronic defects or injuries, where the patient's natural healing response has not resolved the defect. The present disclosure also provides several exemplary treatment modalities for the different extents of treatment needed. | 1. A system for treating joint pain comprising: an injectable material for biological stimulation of a healing response to a subchondral defect in a subchondral region of a bone adjacent to the joint;an injector tool for delivery of the material; andan instrument for guiding the tool to a target location in the bone, the instrument comprising: a first portion having a first guide section configured to guide the tool to the target location;a reference probe extending from the first portion having a tip configured for placement on the bone, without breaching an outer surface of the bone, and proximate a selected anatomical landmark; anda handle portion coupled to the first portion and having a second guide section configured to guide a tool to the target location;wherein the first guide section comprises portals configured to guide the tool towards the tip of the reference probe, and the second guide section comprises portals configured to guide the tool at an acute angle with respect to the reference probe. 2. The system of claim 1, wherein the reference probe is configured to rest against the anatomical landmark on the bone. 3. The system of claim 1, wherein the portals of the first and second guide sections are configured to guide the tool into the subchondral region of the bone. 4. The system of claim 1, wherein the instrument is a unitary body. 5. The system of claim 1, wherein the handle portion is detachable from the first portion. 6. The system of claim 1, wherein the handle portion is configured with a radius of curvature that provides a plurality of acute angles of approach to the target location. 7. The system of claim 1, further including a drilling tool for drilling into the subchondral region of the bone. 8. The system of claim 1, wherein the injectable material is a bone cement, a bone void filler, or a bone substitute material. 9. The system of claim 8, wherein the injectable material further includes a bone growth stimulating factor. 10. The system of claim 1, wherein the portals of the first guide section define trajectories that are substantially parallel to the reference probe. 11. The system of claim 1, wherein the portals of the second guide section comprise a set of differently angled portals. | Zimmer Knee Creations, Inc. | A61B017/56 A61B017/17 A61B017/68 A61B017/88
US09042954 | B2 | 2015-05-26 | Analyte sensors comprising hydrogel membranes | 2009-11-24 | Generally, embodiments of the present disclosure relate to analyte determining methods and devices (e.g., electrochemical analyte monitoring systems) that have improved signal response and stability by inclusion of a coating including a hydrogel, a crosslinker, and a swelling modulator, where the coating is disposed proximate to a working electrode of in vivo and/or in vitro analyte sensors, e.g., continuous and/or automatic in vivo monitoring using analyte sensors and/or test strips. Also provided are systems and methods of using the, for example electrochemical, analyte sensors in analyte monitoring. | 1. An electrode comprising: a sensing layer;a first membrane disposed over the sensing layer, wherein the first membrane is a flux limiting membrane; anda second membrane disposed over the flux limiting membrane, wherein the second membrane is a biocompatible membrane which comprises a hydrogel, a crosslinker, and a swelling modulator comprising a 3-arm epoxide sufficient to decrease or eliminate an early signal attenuation effect by the electrode. 2. The electrode of claim 1, wherein at least a portion of the electrode is adapted to be subcutaneously positioned in a subject. 3. The electrode of claim 1, wherein the second membrane is configured to reduce the absorption of water as compared to a membrane in the absence of the swelling modulator. 4. The electrode of claim 1, wherein the hydrogel comprises a heterocyclic nitrogen-containing polymer hydrogel. 5. The electrode of claim 4, wherein the crosslinker comprises a 4-arm polyethylene glycol epoxide. 6. The electrode of claim 5, wherein the 4-arm polyethylene glycol epoxide has a molecular weight of 1,000 to 20,000 Da. 7. The electrode of claim 1, wherein the 3-arm epoxide comprises triglycidyl glycerol. 8. The electrode of claim 1, wherein the hydrogel comprises a self-polymerizing hydrogel. 9. The electrode of claim 8, wherein the self-polymerizing hydrogel comprises a polyethylene glycol hydrogel or a polyethylene glycol derivative hydrogel. 10. The electrode of claim 9, wherein the self-polymerizing hydrogel comprises polyethylene glycol triacrylate. 11. The electrode of claim 8, wherein the crosslinker comprises a 2-arm polyethylene glycol acrylate or a 4-arm polyethylene glycol acrylate. 12. The electrode of claim 11, wherein the crosslinker has a molecular weight of 1,000 to 20,000 Da. 13. The electrode of claim 7, wherein the self-polymerizing hydrogel comprises an activating agent and an initiator. 14. The electrode of claim 13, wherein the activating agent comprises ferrous gluconate. 15. The electrode of claim 13, wherein the initiator comprises hydrogen peroxide. 16. The electrode of claim 8, wherein the self-polymerizing hydrogel is adapted to polymerize at room temperature. 17. The electrode of claim 1, wherein the electrode is a glucose sensor and the sensing layer further comprises a glucose-responsive enzyme. 18. The electrode of claim 17, wherein the sensing layer comprises a redox mediator. 19. The electrode of claim 18, wherein the redox mediator comprises a ruthenium-containing complex or an osmium-containing complex. 20. The electrode of claim 18, wherein the electrode is an in vivo or in vitro analyte sensor. 21. The electrode of claim 1, wherein the crosslinker is covalently bonded to the swelling modulator. 22. The electrode of claim 21, wherein the second membrane comprises a crosslinker having a molecular weight of 1,000 to 20,000 Da covalently bonded to a swelling modulator which does not comprise PEG. 23. The electrode of claim 1, wherein the... | Abbott Diabetes Care Inc. | A61B005/05 A61B005/1473 A61B005/145 A61B005/1486
US10041897 | B2 | 2018-08-07 | Hermetic implantable sensor | 2011-02-09 | At least one conductor is formed at a preselected location on a substrate made of a first insulating material having a high temperature resistance. The conductor is made from a solidified electrically conductive thick film material. A coating made of a second insulating material is formed over the substrate to hermetically seal at least a portion of the conductor. An exposed distal region of the conductor provides a detection electrode. The conductor has a reduced porosity that inhibits migration of fluid or constituents thereof through the conductor. | 1. A sensor, comprising: a substrate made of a first insulating material having a high temperature resistance;at least one conductor disposed over a preselected external surface location on the substrate, wherein the external surface location is planar and disposed on an upper or lower external surface of the substrate, and the at least one conductor is elongated along the external surface location and made from a solidified electrically conductive thick film material which contains particles of a metal selected from the group consisting of platinum, palladium, rhodium, ruthenium, iridium and osmium;a coating made of a second insulating material formed over at least a portion of the at least one conductor to hermetically seal at least a portion of the at least one conductor;an exposed distal region of the at least one conductor providing a detection electrode; andthe at least one conductor being sintered by heating to a temperature within a range that extends from below a melting point of the particles of metal to approximately 300° C. below the melting point of the particles of metal and having a reduced porosity that inhibits migration of fluid or constituents thereof through the at least one conductor, wherein the at least one conductor is hermetic along its elongated length. 2. The sensor of claim 1 wherein the first insulating material is selected from the group consisting of ceramic and glass/ceramic combinations. 3. The sensor of claim 2 wherein the ceramic is selected from the group consisting of oxides, carbides, borides, nitrides, and silicides of aluminum, zirconium, beryllium, silicon, titanium, yttrium, and zinc and mixtures thereof. 4. The sensor of claim 1, wherein the second insulating material is selected from the group consisting of glass, ceramic, glass/ceramic combinations, and polymers. 5. The sensor of claim 1 wherein the substrate has a generally planar configuration. 6. The sensor of claim 1 wherein there are at least two conductors that are elongated and parallel. 7. The sensor of claim 1 wherein at least one insulation-testing conductor is formed on the substrate in addition to the at least one conductor, and the coating of the second insulating material hermetically seals at least a portion of the insulation-testing conductor that may come into contact with a fluid. 8. The sensor of claim 1 further comprising a source of glucose oxidase (GO) positioned in operative contact with the detection electrode for catalyzing the reaction of glucose and oxygen. 9. The sensor of claim 1 wherein an exposed proximal region of the at least one conductor provides a contact pad. 10. The sensor of claim 1 wherein an electrical contact connects to the at least one conductor, the contact being selected from the group consisting of a metal wire, sintered metal particles, a sintered mixture of metal and inorganic oxide particles, and a metal plug. 11. A sensor suitable for implanting in living tissue for monitoring the level of a ... | Glysens, Incorporated, a California Corporation | A61B005/05 G01N027/12 A61B005/145 A61B005/1486
US09504435 | B2 | 2016-11-29 | Method and device for automated detection of the central line of at least one portion of a tubular tissue structure | 2010-07-01 | A method and a device are disclosed for automated detection of at least one part of the central line of at least one portion of a tubular tissue structure. In at least one embodiment of the method, characteristic landmarks of the tubular tissue structure are detected in a 3D data record of the tubular tissue structure; the detected characteristic landmarks of the tubular tissue structure and landmarks of a model of landmarks, which model belongs to the tubular tissue structure and takes into account the position of the landmarks relative to one another, are related to one another such that at least one portion of the tubular tissue structure is identified in the 3D data record of the tubular tissue structure. The central line of the identified portion of the tubular tissue structure is determined at least in part on the basis of at least one detected characteristic landmark and/or at least one landmark of the model of landmarks. | 1. A method for automated detection of at least one part of a central line of at least one portion of a tubular tissue structure of a current patient under examination, the method comprising: detecting, by a computer processor, characteristic landmarks of the tubular tissue structure in a 3D data record of the tubular tissue structure;identifying, by the computer processor, a set of anatomical landmarks of the at least one portion of the tubular tissue structure in the 3D data record based on a registration of the detected characteristic landmarks of the tubular tissue structure and landmarks of a model of landmarks, the identifying including, performing an affine transformation to register the characteristic landmarks and the model of the landmarks, andanatomically identifying each of the characteristic landmarks with a corresponding one of the landmarks of the model of landmarks such that each of the characteristic landmarks has a corresponding anatomical identification associated therewith, wherein the model of landmarks, is designed with respect to a build or physique of an average person,is scalable to fit a size of the current patient so that the landmarks of the model of landmarks indicate a usual spatial position or arrangement of parts of the tubular tissue structure, andis in the form of a knowledge base stored in a storage device;determining, by the computer processor, the central line of the at least one portion of the tubular tissue structure, at least in part, on a basis of a path-finding operation that uses the identified set of anatomical landmarks; anddisplaying, on a visualization instrument, a curved multiplanar reconstruction (CPR) of the at least one portion of the tubular tissue structure generated by the computer processor based on the central line thereof. 2. The method as claimed in claim 1, wherein the model of landmarks includes at least one further landmark assigned to at least one further portion of the tubular tissue structure, the at least one further portion being only partly contained in the 3D data record, wherein the at least one further landmark of the model of landmarks cannot be put into relation with a corresponding characteristic landmark of the tubular tissue structure because the corresponding characteristic landmarks of the at least one further portion are not imaged in the 3D data record. 3. The method as claimed in claim 2, wherein the at least one further portion of the tubular tissue structure only partly contained in the 3D data record is identified on a basis of the at least one further landmark of the model of landmarks, and the central line of the at least one portion of the tubular tissue structure is at least partly determined based on the identified at least one further portion. 4. The method as claimed in claim 1, wherein the at least one portion of the tubular tissue structure is provided with its anatomical designation. 5. The method as claimed in claim 1, wherein a CPR (curved multipla... | Siemens Aktiengesellschaft | G06K009/00 A61B005/05 A61B006/00 G06T007/00 G06T019/00 A61B005/055
US20110105877 | A1 | 2011-05-05 | APPARATUS AND METHOD FOR DETECTING DISCONNECTION OF AN INTRAVASCULAR ACCESS DEVICE | 2010-10-29 | An apparatus and method are disclosed for detecting the disconnection of a vascular access device such as a needle, cannula or catheter from a blood vessel or vascular graft segment. A pair of electrodes is placed in direct contact with fluid or blood in fluid communication with the vascular segment. In one embodiment, the electrodes are incorporated into a pair of connectors connecting arterial and venous catheters to arterial and venous tubes leading to and from an extracorporeal blood flow apparatus. Wires leading from the electrodes to a detecting circuit can be incorporated into a pair of double lumen arterial and venous tubes connecting the blood flow apparatus to the blood vessel or vascular graft. The detecting circuit is configured to provide a low-voltage alternating current signal to the electrodes to measure the electrical resistance between the electrodes, minimizing both the duration and amount of current being delivered. Detection of an increase in electrical resistance between the electrodes exceeding a pre-determined threshold value may be used to indicate a possible disconnection of the vascular access device. | 1. A system for detecting the disconnection of a vascular access device from a blood vessel or vascular graft, comprising: a fluid delivery device for providing fluid through a first conduit into a first site of the blood vessel or graft;a first electrode in contact with the lumen of the first conduit;a second electrode in fluid communication with a second site of the blood vessel or graft;an electronic circuit connected to the first and second electrodes, and configured to deliver a control signal to the first and second electrodes in order to measure the electrical resistance of a fluid between the first and second electrodes, whereinat least one of the electrodes is located closer to the blood vessel or graft than to the fluid delivery device. 2. The system of claim 1, wherein the fluid delivery device comprises a pump. 3. The system of claim 1, wherein the fluid delivery device comprises a hemodialysis blood flow circuit. 4. The system of claim 1, wherein the second electrode is in contact with the lumen of a second conduit accessing the blood vessel or graft at the second site. 5. The system of claim 4, wherein the second conduit comprises part of a fluid flow path from the blood vessel or graft to the fluid delivery device. 6. The system of claim 5, wherein the first conduit comprises a first connector connecting a first vascular access catheter to a first tube segment, the second conduit comprises a second connector connecting a second vascular access catheter to a second tube segment, and the first and second vascular access catheters are shorter than the first and second tube segments. 7. The system of claim 6, wherein the first connector includes the first electrode and the second connector includes the second electrode. 8. The system of claim 1, wherein the first conduit comprises a double lumen flexible tube having a first lumen for carrying the fluid, and a second lumen for carrying a wire connecting the first electrode to the electronic circuit. 9. The system of claim 1, further comprising a controller configured to receive a signal from the electronic circuit, the signal representing the electrical resistance between the electrodes, wherein the controller is programmed to trigger an alert signal when the electrical resistance value exceeds a pre-determined threshold. 10. The system of claim 9, wherein the alert signal comprises an electrical command to a tubing occluder apparatus, wherein the tubing occluder apparatus includes a mechanical occluder arranged and configured to occlude the conduit. 11. An apparatus for monitoring the continuity between a vascular access device and a blood vessel or vascular graft segment, comprising: a first and second vascular connector, the first connector being attached on a proximal end to a distal end of a fluid-carrying lumen of a first double-lumen tube, and the second connector being attached on a proximal end to a distal end of a fluid-carrying lumen of a second double-lumen tube,th... | DEKA Products Limited Partnership | A61B005/04
US20110118696 | A1 | 2011-05-19 | Device, system, and method for targeted delivery of anti-inflammatory medicaments to a mammalian subject | 2010-02-19 | A device, a system, or a method is described for treating a disease or condition of one or more joints of articulating bone in a mammalian subject. A device is described which includes one or more substrates configured to contact one or more body contours at one or more joints of articulating bone of a mammalian subject; and an applicator supported by the one or more substrates and configured to inject one or more medicaments to one or more joint tissues of the mammalian subject. | 1.-47. (canceled) 48. A method comprising: injecting one or more medicaments to one or more joint tissues of a mammalian subject, wherein the injecting of the one or more medicaments is via a device including one or more sheaths configured to contact one or more body contours in proximity to one or more joints of articulating bone of the mammalian subject; and one or more microjet applicators supported by the one or more sheaths and configured to inject one or more medicaments to one or more joint tissues of the mammalian subject. 49. The method of claim 48, wherein the one or more applicators include one or more high speed microjets. 50. The method of claim 49, wherein the high speed microjets are nanoliter-volume pulsed microjets. 51. The method of claim 48, wherein the one or more sheaths includes a sleeve-shaped article of clothing or a sleeve-shaped body covering. 52. The method of claim 51, wherein the one or more sheaths includes shirt sleeve, pant, leg covering, glove, stocking, bandage-like covering, brace, knee brace, elbow brace, ankle brace, foot brace, hand brace, or spinal brace. 53. The method of claim 48, wherein the one or more joint tissues includes one or more of a fibrous capsule, subsynovium, synovial membrane, synovium, joint cavity, synovial fluid, articular cartilage, subchondral bone, or ligament. 54. The method of claim 48, wherein the one or more joints includes a synovial joint, cartilaginous joint, or fibrous joint. 55. The method of claim 48, wherein the one or more applicators supported by the one or more sheaths are configured to respond to one or more sensors, wherein the one or more applicators are configured to inject the one or more medicaments to the one or more joint tissues of the mammalian subject. 56. The method of claim 48, further including providing one or more sensors configured to detect movement of the one or more body contours. 57. The method of claim 56, wherein the one or more sensors are configured to monitor joint movement, body movement, or gait of the mammalian subject. 58. The method of claim 56, wherein the one or more sensors are configured to monitor posture of the mammalian subject. 59. The method of claim 48, further including providing one or more sensors configured to detect one or more physiological conditions of the mammalian subject. 60. The method of claim 59, wherein the one or more sensors are configured to detect tissue swelling, tissue pressure, tissue color, tissue temperature, environmental temperature, electrical property of tissue, optical property of tissue, perspiration, or presence of an analyte. 61. The method of claim 60, wherein the analyte includes an inflammatory marker, antibody, or cytokine. 62. The method of claim 48, further including providing one or more sensors configured to determine quality of location or quality of contact placement of the one or more sheaths on the one or more body contours. 63. The method of claim 48, further including pr... | Searete LLC, a limited liability corporation of the State of Delaware | A61M031/00 A61B005/103
US07955260 | B2 | 2011-06-07 | Method for rating severity of psoriasis | 2003-10-20 | Methods for rating the severity of psoriasis in a patient are provided. Such methods comprise assessing the percent of the patient's body surface exhibiting signs of psoriasis, assessing the overall elevation, erythema, and scale of the patient's psoriasis plaques, and classifying the severity of the patient's psoriasis into at least four classes based upon the percent of the patient's body surface exhibiting signs of psoriasis adjusted by the overall elevation, erythema and scale of the patient's psoriasis plaques. In the present methods the weight given to one or more of the overall qualities of elevation, erythema or scale of the patient's psoriasis plaques in classifying the severity of the patient's psoriasis into a higher class is greater than the weight given to other of the overall qualities. A computer readable medium and system for practicing the present methods are also provided. | 1. A method for rating the severity of psoriasis in a patient, comprising: a) assessing a percent of the patient's body surface exhibiting signs of psoriasis,b) assessing overall qualities of the patient's psoriasis plaques, said qualities being elevation, erythema, and scale,c) classifying, in a computing device, the severity of the patient's psoriasis into at least four classes of increasing severity,wherein a classification into the lowest class indicates that the patient is free of psoriasis, and wherein a classification into the highest class indicates that the patient has a most severe form of psoriasis,wherein said classification is based upon the percent of the patient's body surface exhibiting signs of psoriasis adjusted by the overall elevation, erythema and scale of the patient's psoriasis plaques, andwherein the weight given to one or more of the overall qualities of elevation, erythema or scale of the patient's psoriasis plaques in classifying the severity of the patient's psoriasis into a higher class is greater than the weight given to other of the overall qualities. 2. The method of claim 1 wherein the overall elevation, erythema and scale of the patient's psoriasis plaques are categorized into a plurality of categories of increasing severity, with the lowest category encompassing plaques that display no elevation, no erythema, and no scale, respectively. 3. The method of claim 2 wherein the overall elevation, erythema and scale of the patient's psoriasis plaques are categorized into 2 to 8 categories of increasing severity. 4. The method of claim 3 wherein the overall elevation, erythema and scale of the patient's psoriasis plaques are categorized into 4 categories of increasing severity. 5. The method of claim 1 wherein the percent of the patient's body surface exhibiting signs of psoriasis is grouped into a plurality of percentage ranges of increasing involvement. 6. The method of claim 5 wherein the percent of the patient's body surface exhibiting signs of psoriasis is grouped into 4 to 10 ranges of increasing involvement. 7. The method of claim 5 wherein the lowest range is 0% body surface involvement and the highest range is approximately 51%-100% body surface involvement. 8. The method of claim 7 wherein the percent of the patient's body surface involvement is grouped into the following ranges: 0% of the body surface exhibiting signs of psoriasis,from about 1 to about 3% of the body surface exhibiting signs of psoriasis,from about 4 to about 9% of the body surface exhibiting signs of psoriasis,from about 10 to about 20% of the body surface exhibiting signs of psoriasis,from about 21% to about 29% of the body surface exhibiting signs of psoriasis,from about 30 to about 50% of the body surface exhibiting signs of psoriasis, andfrom about 51% to 100% of the body surface exhibiting signs of psoriasis. 9. The method of claim 5, wherein the percentage ranges are broader as the percent of body surface involvement increases... | DatAcquire, Inc. | A61B005/00
US07962199 | B2 | 2011-06-14 | Method and apparatus for determination of bone fracture risk using raman spectroscopy | 2006-06-29 | A method is disclosed for analyzing keratinized tissue, particularly fingernails, of a subject to diagnose osteoporosis and bone fracture risk. A Raman spectrum of a sample of keratinized tissue is generated. Broad spectral background features of the spectrum are removed, preferably by using Fourier transform analysis. Peak heights of Raman features of interest, particularly the S—S bond of cystine, are measured. These peak height measurements are normalized using reference peak heights of Raman features that are invariant between normal and osteoporotic subjects, such as the CH2 bending peak. | 1. A method of analyzing keratinized tissue of a subject to diagnose osteoporosis and bone fracture risk, comprising the steps of: (a) generating a Raman spectrum of a keratinized tissue sample using a Raman spectroscope, said spectrum including at least a peak of a measureable height representative of disulfide bonds and a peak of a measureable height representative of cystine;(b) removing broad spectral background features from the spectrum;(c) measuring the peak height of the peak representative of disulfide bonds; and(d) normalizing the cystine peak height using an invariant reference peak height. 2. A method as defined in claim 1, wherein the step of removing the broad spectral background features comprises: (a) performing a Fourier transform on the spectrum;(b) removing high frequency components to produce a modified transformed spectrum;(c) performing an inverse Fourier transform on the modified transformed spectrum to produce a modified spectrum, and(d) subtracting the modified spectrum from the original spectrum. 3. A method as defined in claim 1, wherein the step of measuring peak height comprises identifying peak maxima and baseline from a data array associated with the spectrum independent of a standard reference location. 4. A method as defined in claim 1, wherein the keratinized sample is selected from the group consisting of hair, skin and nails. 5. A method as defined in claim 1, wherein the sample resides in a live subject. 6. A stand-alone instrument for proving information about a risk of bone fracture of a subject, comprising a Raman spectroscope, a digital computer configured to receive and analyze spectra from the spectroscope and provides information about the bone fracture risk of the subject, and means for generating an output of the information from the digital computer in human readable form, and wherein the instrument is configured to carry out the method as defined in claim 1 upon the placement of a sample of keratinized tissue of the subject within scanning range of the instrument. | University of Wyoming | A61B005/05
US20110105912 | A1 | 2011-05-05 | CEREBRAL AUTOREGULATION INDICES | 2010-11-05 | Monitoring cerebral autoregulation may include determining one or more autoregulation indices incorporating cerebral blood flow and blood pressure measurements and/or indices. Measurement techniques may be invasive or non-invasive. Various combinations of data, e.g., oximeter data, electrocardiogram data, blood pressure data, hemoglobin data, and heart rate data, may be used to create various indices. Many of the indices may be based on correlations of data. A display may indicate several of the indices. | 1. A method comprising: receiving data relating to cerebral blood flow of a patient;receiving data relating to blood pressure of the patient;correlating the cerebral blood flow and blood pressure data; andutilizing the correlated data to monitor a cerebrovascular autoregulation state of the patient. 2. The method of claim 1, further comprising causing a change of blood pressure of the patient based on the cerebrovascular state of the patient determined based on the correlated data. 3. The method of claim 1, wherein the data relating to cerebral blood flow of a patient is at least one of oximeter data, electrocardiogram data, hemoglobin data, and heart rate data. 4. The method of claim 1, wherein the data relating to blood pressure of a patient is at least one of oximeter data, electrocardiogram data, blood pressure data, hemoglobin data, and heart rate data. 5. The method of claim 1, further comprising: determining a cerebral blood flow measurement based on the received data relating to cerebral blood flow; anddetermining a blood pressure measurement based on the received data relating to blood pressure. 6. The method of claim 1, further comprising: receiving the data relating to cerebral blood flow of the patient utilizing a cerebral oximetry measurement of regional oxygen saturation using a near infrared sensor;using the measurement as an index to cerebral blood flow; andreceiving the data relating to blood pressure as arterial distension changes as measured by arterial tonometry as an index in changes in blood pressure. 7. The method of claim 1, further comprising: receiving the data relating to cerebral blood flow of the patient utilizing a cerebral oximetry measurement of regional oxygen saturation using a near infrared sensor;using the measurement as an index to cerebral blood flow; andreceiving the data relating to blood pressure as signal changes from a non-invasive servo controlled cuff measurement of continuous blood pressure. 8. The method of claim 1, further comprising: receiving the data relating to cerebral blood flow of the patient utilizing a cerebral oximetry measurement of regional oxygen saturation using a near infrared sensor;using the measurement as an index to cerebral blood flow; andreceiving the data relating to blood pressure as changes in optical density measured by a non-invasive peripheral optical plethysmograph as an index of blood pressure changes. 9. The method of claim 1, further comprising: receiving the data relating to cerebral blood flow of the patient utilizing a cerebral oximetry measurement of regional oxygen saturation using a near infrared sensor;using the measurement as an index to cerebral blood flow; andreceiving the data relating to blood pressure as changes in optical density measured by a non-invasive peripheral pulse oximeter plethysmograph as an index of blood pressure changes. 10. The method of claim 1, further comprising: receiving the data relating to cerebral blood flow of the patient ... | - | A61B005/0205
US20110135571 | A1 | 2011-06-09 | HYBRID NANOPARTICLES AS ANTI-CANCER THERAPEUTIC AGENTS AND DUAL THERAPEUTIC/IMAGING CONTRAST AGENTS | 2009-02-23 | The presently disclosed subject matter provides nanoscale coordination polymers for use as anticancer agents and as dual anticancer/imaging agents. The nanoscale coordination polymers can comprise a plurality of platinum metal complexes; nonplatinum anticancer drug bridging ligands complexed to multiple metal centers; or combinations thereof. The nanoscale coordination polymers can be targeted for delivery to cancer cells. They can also comprise stabilizing agents to allow for controlled and/or sustained release of anticancer agents in vivo. | 1. A nanoparticle comprising a coordination polymer comprising a plurality of platinum metal complexes. 2. The nanoparticle of claim 1, wherein the plurality of platinum metal complexes consist of a plurality of platinum (II) metal complexes, a plurality of platinum (IV) metal complexes, or a combination thereof. 3. The nanoparticle of claim 1, wherein one or more of the platinum metal complexes comprises: a platinum metal atom;at least one nonbridging ligand, wherein the at least one nonbridging ligand is bonded to the platinum metal atom through at least one coordination bond; andat least one bridging ligand, wherein the at least one bridging ligand is bonded to the platinum metal atom through at least one coordination bond and comprises at least one linking moiety, wherein each of the at least one linking moiety is bonded to an additional metal atom via a coordination bond. 4. The nanoparticle of claim 3, wherein each of the at least one linking moiety is independently selected from the group consisting of a carboxylate, a carboxylic acid, an amine, a hydroxyl, a thiol, a carbamate, an ester, a phosphate, a phosphonate, a carbonate, and an amide. 5. The nanoparticle of claim 3, wherein each of the at least one bridging ligand is independently selected from the group consisting of a polymeric bridging ligand and a nonpolymeric bridging ligand. 6. The nanoparticle of claim 5, wherein each of the at least one bridging ligand is a nonpolymeric bridging ligand. 7. The nanoparticle of claim 6, wherein the bridging ligand comprises at least two carboxylate groups. 8. The nanoparticle of claim 6, wherein at least one platinum metal complex comprises two bridging ligands, wherein each of the two bridging ligands is bonded to the platinum metal atom through one coordination bond and comprises at least one linking moiety. 9. The nanoparticle of claim 8, wherein each of the two bridging ligands is independently selected from the group consisting of 1,4-benzene dicarboxylate; 1,3,5-benzene tricarboxylate; succinate; and ethylene diamine tetraacetate. 10. The nanoparticle of claim 6, wherein at least one platinum metal complex comprises one bridging ligand, wherein the one bridging ligand is bonded to the platinum metal atom through two coordination bonds and comprises at least two linking moieties. 11. The nanoparticle of claim 10, wherein the one bridging ligand is a bipyridine dicarboxylate. 12. The nanoparticle of claim 11, wherein the bipyridine dicarboxylate is selected from 2,2′-bipyridine-5,5′-dicarboxylate and 2,2′-bipyridine-4,4′-dicarboxylate. 13. The nanoparticle of claim 6, wherein one of the at least one bridging ligand is a nonplatinum anticancer drug. 14. The nanoparticle of claim 13, wherein the nonplatinum anticancer drug is selected from the group consisting of methotrexate, folic acid, leucovorin, vinblastine, vincristine, melphalan, pemetrexed, vindesine, anastrozole, doxorubicin, cytarabine, azathioprine, letrozole and car... | - | A61K009/14 C08F026/10 B32B005/00 A61K031/79 A61K051/00 A61K049/00 A61B005/055 A61K009/127 C12N005/00 A61P035/00 B82Y005/00
US07945312 | B2 | 2011-05-17 | Multisensor probe for tissue identification | 2008-08-22 | A multisensor probe for continuous real-time tissue identification. The multisensor probe includes a tissue penetrating needle, a plurality of sensors useful in characterizing tissue and a position sensor to measure the depth of the needle into the tissue being diagnosed. The sensors include but are not limited to an optical scattering and absorption spectroscopy sensor, an optical coherence domain reflectometry sensor, an electrical impedance sensor, a temperature sensor, a pO2 sensor, a chemical sensor and other sensors useful in identifying tissue. The sensors may take the form of a plurality of optical fibers extending through said needle. A retractable sheath may be disposed around the distal section of the needle to protect the needle when not in use. The sheath retracts when the probe is inserted into tissue and the position of the sheath relative to the needle may be measured to determine the needle's depth. Systems and methods for tissue identification are also provided. | 1. A multisensor probe for tissue identification, the probe connectable to a controller via a cable, the probe comprising: a handle;an elongate body having a distal tip and a lumen extending through the elongate body to the distal tip;a slideable sheath disposed over a distal section of the elongate body and being retractable from the elongate body toward the handle;a resilient member configured to bias the sheath toward the distal tip of the elongate body; anda plurality of optical fibers extending from the controller, through the cable, through the lumen, to the distal tip of the elongate body when the probe is connected to the controller, wherein at least two of the plurality of optical fibers are optical scattering and absorption spectroscopy fiber optics and wherein at least one of the plurality of optical fibers is an optical coherence domain reflectometry fiber optic, and wherein a position sensor is incorporated into the probe, the position sensor adapted to measure a depth to which the elongate body is inserted into tissue, the position sensor configured to read a position of the sheath relative to the elongate body. 2. The multisensor probe of claim 1 wherein the position sensor is a sensor selected from the group consisting of a resistive sensor and a linear optical encoder. 3. The multisensor probe of claim 1 having a conducting element extending through the lumen. 4. The multisensory probe of claim 1, further comprising a marking switch to identify a location in the tissue the distal section of the elongate body is inserted into the tissue. 5. The multisensory probe of claim 1, wherein the resilient member is a spring. 6. A tissue detection system comprising: a probe, the probe comprising a multisensor needle comprising a plurality of optical fibers, the probe further comprising: a handle;a slideable sheath disposed over a distal section of the needle and being retractable from the needle toward the handle;a resilient member configured to bias the sheath over the distal section of the needle;a position sensor adapted to sense depth of the needle inserted into tissue using the sheath wherein the sensor reads a position of the sheath relative to the needle; anda controller configured to deliver and collect light through the plurality of optical fibers wherein at least one of the plurality of optical fibers is utilized as an optical coherence domain reflectometry sensor configured for measuring backscattered and reflected light as a function of depth into the tissue and wherein at least one the optical fibers is utilized for optical scattering and absorption spectroscopy. 7. The system of claim 6 further comprising at least one electrode for sensing electrical information about the tissue. 8. The system of claim 6 wherein the sheath retracts when the needle is inserted into the tissue. 9. The system of claim 6, wherein at least two of the plurality of optical fibers are utilized for optical scattering and absorption spectroscopy... | Bioluminate, Inc. | A61B006/00 A61B005/00
US08206268 | B2 | 2012-06-26 | Display mode selection | 2011-01-11 | An apparatus, a method, and a computer program are disclosed. The apparatus comprises a processor. The processor is configured to obtain exercise data of a user from a measurement sensor, to identify a present exercise phase of an exercise from among a plurality of exercise phases on the basis of the exercise data, and to select a relevant display mode from among a plurality of display modes on the basis of the present exercise phase and a mapping between the display modes and the exercise phases, wherein the relevant display mode defines a set of display elements associated with the present exercise phase to be displayed to the user. | 1. An apparatus comprising: a processing device configured to obtain exercise data of a user from a plurality of measurement sensors associated with bicycling, the exercise data comprising training parameters relating to the user's actions during bicycling, at least one of the measurement sensors being a speed sensor, the exercise data comprising speed data, the processing device identifying a present exercise phase of an exercise from among a plurality of exercise phases on the basis of the speed data, the processing device selecting a relevant display mode from among a plurality of display modes on the basis of the present exercise phase and a mapping between the display modes and the exercise phases, the processing device automatically switching between display modes relevant to the present exercise phase based on the speed data, wherein the relevant display mode defines a set of display elements associated with the present exercise phase to be displayed to the user, wherein each display element comprises exercise data; anda display configured to display the relevant display mode to the user during bicycling, the plurality of display modes comprising a summary display mode comprising a present heart rate, a travelled distance, and an elapsed exercise time as display elements, and at least one of the following:an uphill display mode comprising increase in heart rate, starting from the bottom of the hill, steepness of the hill, and an elapsed time going uphill as display elements;a downhill display mode comprising decrease in heart rate, starting from the top of the hill, speed, and an elapsed time going downhill as display elements;a speed interval display mode comprising increase in heart rate, starting from the start of the speed interval, average speed during the speed interval, and an elapsed time since the start of the speed interval as display elements;a recovery period display mode comprising decrease in heart rate, starting from the start of the recovery period, a reaction diagram illustrating a previous speed interval, and an elapsed recovery time as display elements;a heart rate interval display mode comprising increase in heart rate, starting from the start of the heart rate interval, travelled distance during the heart rate interval, and an elapsed time since the start of the heart rate interval as display elements;a recovery period display mode comprising decrease in heart rate, starting from the start of the recovery period, a reaction diagram illustrating a previous heart rate interval, and an elapsed recovery time as display elements;a cadence interval display mode comprising cadence as rotations per minute, average speed during the cadence interval, and an elapsed time since the start of the cadence interval as display elements;a recovery period display mode comprising decrease in heart rate, starting from the start of the recovery period, a reaction diagram illustrating the previous cadence interval, and an elapsed recovery ti... | Polar Electro Oy | A63B071/00
US08606354 | B2 | 2013-12-10 | Visceral fat measuring device | 2011-01-19 | A visceral fat measuring device is capable of simply and noninvasively measuring a visceral fat amount, by calculating the visceral fat amount based on trunk measurement information, impedance information of the entire trunk, and impedance information of a surface layer of the trunk. The visceral fat measuring device includes: a belt to be wound around the trunk that has a pressing member with electrodes E to be pressed onto the dorsal side of a person subjected to measurement; a pair of belt portions respectively fixed to both sides of the pressing member; a buckle for connecting the pair of belt portions; and indicators on the pair of belt portions to show distances from a reference position of the belt. | 1. A visceral fat measuring device for calculating a visceral fat amount based on: trunk measurement information for use in obtaining an area of a cross-sectional plane of an abdominal part of a trunk which is orthogonal to a body axis of the trunk;impedance information of an entire trunk obtained by applying an electric current through the trunk from hands and legs and measuring a potential difference in part of a surface of the trunk; andimpedance information of a surface layer of the trunk obtained by winding a belt having a plurality of electrodes around the trunk so as to apply the electric current through a vicinity of the surface layer of the trunk and measure a potential difference in part of the surface of the trunk,said visceral fat measuring device comprising a belt assembly that comprises:a pressing member to be placed on a dorsal side of a person subjected to measurement and provided with the electrodes;a first belt portion extending from a first side of the pressing member to a free end, said first belt portion being provided with first indicators showing a distance from the free end of the first belt portion;a second belt portion extending from a second side, opposite to said first side, of the pressing member to a free end, said second belt portion being provided with second indicators showing a distance from the free end of the second belt portion; anda buckle for connecting the first and second belt portions to one another, wherebythe pressing member can be aligned by adjusting distances of the free ends of the first and second belt portions from the buckle based on the first and second indicators, said indicators being arranged along the first and second belt portions from the free ends towards the pressing member. 2. The visceral fat measuring device according to claim 1, comprising a potential difference detector for measuring the potential difference in part of a dorsal surface of the trunk. 3. The visceral fat measuring device according to claim 1, comprising a potential difference detector for measuring the potential difference in the body axis of the trunk. 4. The visceral fat measuring device according to claim 1, wherein the buckle is provided with a pair of rollers around which the first and second belt portions are respectively wound. 5. The visceral fat measuring device according to claim 1, wherein a transparent plate is provided in a center of the buckle, and an indicator showing a center in a width direction of the buckle is provided in the transparent plate. 6. The visceral fat measuring device according to claim 1, wherein an indicator is provided on a surface of the pressing member to show a center position between the first side and the second side of the pressing member. 7. The visceral fat measuring device according to claim 1, comprising a control unit configured to: calculate a lean body sectional area excluding fat from the impedance information of the entire trunk,calculate a subcutaneous fat secti... | Omron Healthcare Co., Ltd. | A61B005/05
US08449556 | B2 | 2013-05-28 | Method for detecting body parameters | 2011-01-28 | A method for detecting biometric parameters includes the steps of performing a bone graft procedure on a bone, providing a biometric sensor at the bone, and measuring a biometric parameter at the bone with the sensor. The sensor is capable of measuring parameters in an adjacent surrounding including pressure, tension, shear, relative position, and vascular flow. Data relating to the biometric parameter is transmitted to an external source and the data are analyzed to evaluate a biometric condition of the bone. A set of the sensors can be placed on the bone. | 1. A method for detecting biometric parameters, which comprises: performing a surgical procedure on at least one bone;implanting at least one biometric transceiver at the at least one bone;transmitting a first energy wave from the transceiver into a procedure area including at least one of the bone and an area adjacent the bone;quantitatively assessing the behavior of the energy wave with the transceiver;after transmitting the first energy wave: transmitting a second energy wave from the transceiver into the procedure area; andquantitatively assessing the behavior of the second energy wave;wherein at least one of the first and second energy waves are transmitted in a vibratory manner to stimulate the procedure area in accordance with at least one detected parameter of the procedure area; anddetermining a current status of the at least one parameter of the procedure area selected from the group consisting of pressure, tension, shear, relative position, bone density, fluid viscosity, temperature, strain, angular deformity, vibration, venous flow, lymphatic flow, load, torque, distance, tilt, shape, elasticity, motion, bearing wear, subsidence, bone integration, change in viscosity, particulate matter, kinematics, stability, and vascular flow with the transceiver based upon a comparison between the assessed behavior of the first and second energy waves. 2. The method according to claim 1, which further comprises: transmitting data relating to the at least one biometric parameter to an external source; andanalyzing the data to evaluate a biometric condition of the at least one bone. 3. The method according to claim 1, which further comprises providing a set of the transceivers on the at least one bone. 4. The method according to claim 2, which further comprises, based upon the evaluation of the biometric condition of the at least one bone, changing a currently ongoing interoperative procedure at the procedure area. 5. The method according to claim 2, which further comprises, based upon the evaluation of the biometric condition of the at least one bone, chronically monitoring the anatomic condition relating to the procedure area. 6. The method according to claim 1, which further comprises providing energy from outside the procedure area to the transceiver to power the transceiver and, thereby, create the energy wave and quantitatively assess the behavior of the energy wave. 7. The method according to claim 6, which further comprises providing the energy through at least one of an electromagnetic couple, a magnetic couple, a capacitive couple, an inductive couple, a sonic couple, an ultrasonic couple, a fiber optic couple, an optical couple, and an infrared couple. 8. The method according to claim 1, which further comprises: providing a set of the biometric transceivers at the procedure area;transmitting an energy wave from the transceivers into the procedure area;quantitatively assessing the behavior of the energy wave with at least one of the ... | - | A61B017/58 A61B017/60 A61F002/00
US08682608 | B2 | 2014-03-25 | Behavior recognition apparatus | 2010-05-27 | A behavior recognition apparatus includes a hand movement observation unit which outputs hand movement data indicating a hand movement made by a user; a velocity variation detection unit which detects a time at which the hand movement matches a predetermined movement pattern; and a feature selection unit which selects the hand movement data outputted at the time detected by the velocity variation detection unit. A standard behavior pattern storage unit stores a movement pattern data sequence represented by a hand movement data sequence, in association with a user behavior; and a behavior recognition unit calculates a degree of similarity between the movement pattern data sequence stored in the standard behavior pattern storage unit and the hand movement data selected by the feature selection unit, and recognizes, as the user behavior, a behavior represented by the movement pattern data sequence when the degree of similarity is the highest. | 1. A behavior recognition apparatus which recognizes a user behavior from a hand movement made by a user, the behavior recognition apparatus comprising: a hand movement observation unit configured to detect the hand movement made by the user and output hand movement data indicating the hand movement;a velocity variation detection unit configured to detect a time at which a variation in velocity of a hand of the user occurs, from the hand movement data outputted from the hand movement observation unit;a feature selection unit configured to select the hand movement data that is outputted from the hand movement observation unit at specific times including the time detected by the velocity variation detection unit and times before and after the detected time;a standard behavior pattern storage unit configured to store a movement pattern data sequence represented by a hand movement data sequence, in association with the user behavior; anda behavior recognition unit configured to calculate a degree of similarity between the movement pattern data sequence stored in the standard behavior pattern storage unit and the hand movement data selected by the feature selection unit, and to recognize, as the user behavior, a behavior represented by the movement pattern data sequence when the degree of similarity is highest. 2. The behavior recognition apparatus according to claim 1, further comprising an arm posture estimation unit configured to estimate an arm posture of the user relative to a body of the user, from the hand movement data outputted from the hand movement observation unit, wherein, when the arm posture matches a predetermined arm posture and the velocity variation detection unit detects the time at which the variation in the velocity of the hand of the user occurs, the feature selection unit is configured to select the hand movement data of the specific times including the detected time at which the variation in the velocity of the hand of the user occurs and the times before and after the detected time. 3. The behavior recognition apparatus according to claim 2, further comprising a body posture estimation unit configured to estimate a body posture of the user,wherein, when: the arm posture matches the predetermined arm posture; the body posture matches a predetermined body posture; and the velocity variation detection unit detects the time at which the variation in the velocity of the hand of the user occurs, the feature selection unit is configured to select the hand movement data of the specific times including the detected time at which the variation in the velocity of the hand of the user occurs and the times before and after the detected time. 4. The behavior recognition apparatus according to claim 2, wherein the standard behavior pattern storage unit is configured to store, in association with the user behavior, a pair of the movement pattern data sequence represented by the hand movement data sequence and the arm posture of the time a... | Panasonic Corporation | G01P015/00
US09259170 | B2 | 2016-02-16 | Imaging pallets for multi-modality imaging systems | 2009-12-21 | A multi-modality imaging system including first and second imaging modality units having respective field of views (FOVs) that are spaced apart from each other. The imaging system also having a positioning system that includes an imaging pallet. The pallet has an elongated support body that includes first and second portions that extend lengthwise along the support body. The first portion is shaped for imaging within the FOV of the first modality unit, and the second portion is shaped for imaging within the FOV of the second modality unit. The first and second portions are shaped differently than each other. The positioning system is configured to position the first portion of the pallet within the FOV of the first modality unit and configured to position the second portion of the pallet within the FOV of the second modality unit. | 1. A multi-modality imaging system comprising: a nuclear medicine (NM) modality unit and a computed-tomography (CT) modality unit, the NM and CT modality units having respective field of views (FOVs) spaced apart from each other; anda positioning system comprising a pallet support and an imaging pallet, the imaging pallet having an elongated support body that includes a mounting portion and first and second portions extending along a longitudinal axis of the support body, the mounting portion being mounted to the pallet support, the first portion extending between the mounting portion and the second portion and being shaped for imaging within the FOV of the NM modality unit, the second portion being shaped for imaging within the FOV of the CT modality unit, the first and second portions being shaped differently than each other and having first and second lengths, respectively, the second length being greater than the first length, wherein the positioning system is configured to position the first portion of the imaging pallet within the FOV of the NM modality unit while the mounting portion is mounted onto the pallet support and configured to position the second portion of the imaging pallet within the FOV of the CT modality unit while the mounting portion remains mounted onto the pallet support, wherein the FOVs of the NM and CT modality units are aligned with each other along an examination axis, the positioning system moving the imaging pallet along the examination axis to position the imaging pallet, wherein the first portion is shaped for the NM modality unit and is configured such that attenuation to 140 keV gamma rays does not exceed 10%, and wherein the second portion is shaped for the CT modality unit and is configured such that attenuation to 120 kVp X-ray does not exceed 27%. 2. The imaging system in accordance with claim 1 wherein the first portion has a plurality of cross-sections extending transverse to the longitudinal axis, the plurality of cross-sections including a medial cross-section that is proximate to the mounting portion and a distal cross-section that is proximate to the second portion, the medial and distal cross-sections having different dimensions, wherein a cross-sectional shape of the first portion reduces as the first portion extends from the medial cross-section to the distal cross-section, the first portion having a scan section located between the medial and distal cross-sections, the positioning system configured to position the scan section within the FOV of the NM modality unit. 3. The imaging system in accordance with claim 2 wherein the medial cross-section has a cross-sectional area that is greater than a cross-sectional area of the distal cross-section. 4. The imaging system in accordance with claim 1 wherein the imaging pallet has a bottom side, the bottom side having a width along the first portion that reduces as the first portion extends from the mounting portion toward the second portion, the width... | General Electric Company | A61B005/05 A61B005/055 A61B006/03 A61B006/04 A61B006/00 A61B005/00
US09282947 | B2 | 2016-03-15 | Imager focusing based on intraoperative data | 2010-11-18 | Various embodiments herein provide for imager focusing based on intraoperative data. Ideally, an imaging plane of an ultrasound transducer would be truly planar. That is not the case, though. Instead, ultrasound transducers image a volume that is closer to a rectangular volume, but that has focal depths or areas in the imaged volume that are “thinner” and provide a better resolution. In general, embodiments herein may include determining the pose of an imager, such as an ultrasound transducer, and the pose of a location of interest. Based on those poses, a focal adjustment may be determined in order to, for example, better focus the imager on the object of interest. Then data is generated and the focus of the imager is adjusted. Additionally, imaging data and/or the object of interest may be displayed. In other embodiments, estimated projections of medical devices are displayed to allow for better intraoperative planning. | 1. A method for imager focusing based on intraoperative data, implemented on one or more computing devices, comprising: determining, using the one or more computing devices, position and orientation of an imaging plane of a focusable imager used in a medical scene based at least in part on a detected position and orientation of a first tracking device coupled to the focusable imager;determining, using the one or more computing devices, position and orientation of a medical device used in the medical scene based at least in part on a detected position and orientation of a second tracking device coupled to the medical device;determining, using the one or more computing devices, an intersection of the imaging plane and a central axis of the medical device based at least in part on the determined position and orientation of the imaging plane and the determined position and orientation of the medical device;determining, using the one or more computing devices, position and orientation of at least one object of interest in the medical scene based at least in part on the determined intersection of the imaging plane and the central axis of the medical device;determining, using the one or more computing devices, a focal adjustment for the focusable imager based at least in part on the determined position and orientation of the at least one object of interest and a focal plane of the focusable imager; andcausing the focusable imager to adjust the focal plane based at least in part on the determined focal adjustment for the focusable imager. 2. The method of claim 1, further comprising displaying a 3D graphics representation of the object of interest. 3. The method of claim 1, wherein determining the position and orientation of the at least one object of interest comprises accessing stored position and orientation of an object whose position and orientation was previously indicated. 4. The method of claim 3, wherein accessing stored position and orientation of an already-known object of interest comprises accessing stored position and orientation of at least one of a highlight, marking, or annotation of the already-known object of interest. 5. The method of claim 1, wherein the medical device comprises a tracked ablation needle. 6. The method of claim 1, wherein determining the pose for the at least one object of interest comprises determining position and orientation of a body part of an operator. 7. The method of claim 1, wherein determining the focal adjustment based at least in part on the position and orientation of the at least one object of interest comprises determining a trajectory of the medical device onto an imaging plane associated with the imager. 8. The method of claim 1, wherein determining the focal adjustment based at least in part on the position and orientation of the at least one object of interest comprises determining a series of focal adjustments to be used over time based at least in part on the position and orientation of th... | InnerOptic Technology, Inc. | A61B005/05 A61B008/08 A61B008/00 G01S007/52 G01S015/89 A61B005/06
US09538922 | B2 | 2017-01-10 | Monitoring an interval within the cardiac cycle | 2009-10-30 | Aspects of this disclosure describe measuring intervals within a cardiac cycle to, for example, determine whether a patient is a candidate for cardiac therapy initiation or modification. The intervals may be measured in response to a trigger identifying a physiological event. The intervals and an identification of the physiological event may be stored. A physician or clinician may determine whether the patient is a candidate for cardiac therapy modification based on the measured intervals. | 1. A method comprising: identifying, with one or more monitors associated with an implantable medical device (IMD) implanted within a patient, at least one physiological event; andmeasuring, with an interval measurement unit, an interval within a cardiac cycle for one or more cardiac cycles in response to the identification of the at least one physiological event to generate one or more measured intervals. 2. The method of claim 1, further comprising: storing the one or more of the measured intervals and an indication the of the at least one physiological event. 3. The method of claim 1, wherein the physiological event comprises a cardiovascular event. 4. The method of claim 2, wherein the cardiovascular event comprises at least one of a heart rate, a change in heart rate, an impedance measurement, a heart pressure, bradycardia, tachycardia, and fibrillation. 5. The method of claim 1, wherein the physiological event indicates at least one of a worsening cardiovascular condition, a progression to a state of heart failure, and an improving cardiovascular condition. 6. The method of claim 1, wherein the physiological event comprises at least one of sleep apnea, activity level, and change in activity level. 7. The method of claim 1, further comprising: identifying the patient as a candidate for a cardiac therapy based on the measured intervals. 8. The method of claim 1, further comprising: transmitting the one or more measured intervals to an external computing device. 9. The method of claim 1, further comprising: presenting the one or more measured intervals to a clinician. 10. The method of claim 1, further comprising: comparing each one of the measured intervals to a threshold; andtransmitting an alert to an external computing device based on the comparison. 11. The method of claim 1, wherein the interval comprises a QRS width. 12. An implantable medical device (IMD) comprising: one or more monitors configured to identify at least one physiological event of a patient; andan interval measurement unit configured to measure an interval within a cardiac cycle for one or more cardiac cycles in response to the identification of the at least one physiological event to generate one or more measured intervals. 13. The IMD of claim 12, wherein the IMD further comprises: a memory configured to store the one or more measured intervals and an indication of the at least one physiological event. 14. The IMD of claim 12, wherein the physiological event comprises a cardiovascular event. 15. The IMD of claim 14, wherein the cardiovascular event comprises at least one of a heart rate, a change in heart rate, an impedance measurement, a heart pressure, bradycardia, tachycardia, and fibrillation. 16. The IMD of claim 12, wherein the physiological event indicates at least one of a worsening cardiovascular condition, a progression to a state of heart failure, and an improving cardiovascular condition. 17. The IMD of claim 12, wherein the physiological ev... | Medtronic, Inc. | A61B005/02 A61B005/024 A61B005/0245 A61B005/0452 A61N001/37 A61N001/372 A61B005/00 A61N001/362
US09307919 | B2 | 2016-04-12 | Electrode interface system | 2010-11-08 | An electrode interface system for providing a connection between at least one electrode and a maternal-fetal monitor, wherein the interface system converts electrical muscle activity captured by the electrode(s) into uterine activity data signals for use by the maternal-fetal monitor. The electrode interface system of the invention preferably includes a signal converter for processing the signals from the electrode(s) into signals similar to those produced by a tocodynometer. | 1. An interface system for interfacing an electrode with a maternal-fetal monitor, wherein the maternal-fetal monitor has a port configured to operably connect with a tocodynamometer to receive output from the tocodynamometer or to connect with an intrauterine pressure catheter to receive output from the intrauterine pressure catheter, and wherein the maternal fetal monitor is configured to display the tocodynarnometer output or the intrauterine pressure catheter output to a user, wherein the interface system comprises: a first component comprising: an electrode interface configured to operably and removably connect with the electrode and to receive an electrohysterogram signal from the electrode; anda wireless signal transmitter configured to receive the electrohysterogram signal from the electrode interface when the electrode interface is connected with the electrode and to wirelessly transmit the received electrohysterogram signal to a wireless signal receiver located remote from the first component; anda second component comprising: the wireless signal receiver configured to receive the electrohysterogram signal transmitted by the wireless signal transmitter;a signal converter configured to receive the electrohysterogram signal from the wireless signal receiver and to process the electrohysterogram signal into output data that mimics electrical output from a tocodynamometer or an intrauterine pressure catheter; anda maternal-fetal monitor port interface configured to operably and removably connect to the port of the maternal-fetal monitor, wherein the maternal-fetal monitor port interface is configured to receive from the signal converter the output data that mimics electrical output from a tocodynamometer or an intrauterine pressure catheter and to communicate to the maternal-fetal monitor the output data that mimics electrical output from a tocodynamometer or an intrauterine pressure catheter when the maternal-fetal monitor port interface is connected to the port of the maternal-fetal monitor. 2. The interface system of claim 1, wherein the electrode interface is configured to operably and removably connect with a plurality of electrodes and to receive an electrohysterograrn signal from each of the plurality of electrodes. 3. The interface system of claim 2, wherein the plurality of electrodes comprises between 2 and 6 electrodes. 4. The interface system of claim 2, wherein the plurality of electrodes are provided on a strip configured to attach to a body of a patient and wherein the electrode interface is configured to operably and removably connect with a connector of the strip. 5. The interface system of claim 1, wherein the first component is configured to be supported on a body of a patient when the electrode interface is operably and removably connected with the electrode. 6. The interface system of claim 5, wherein the first component is configured to be supported on an arm of the patient when the electrode interface is operably... |  | A61B005/103 A61B005/03 A61B005/0448 A61B005/0444 A61B005/0488 A61B005/04 A61B005/024
US09408538 | B2 | 2016-08-09 | Method and apparatus for pressure sore detection | 2011-01-21 | The present invention pertains to a method and apparatus for pressure sore detection. A modulated optical signal based on a digital code sequence is transmitted to human tissue. A temporal transfer characteristic is derived from the modulated optical signal. Tissue characteristics is determined based on the temporal transfer characteristic. | 1. A pressure sore detection system for detecting pressure sores in human tissue comprising: a memory device configured to store a plurality of values corresponding to a plurality of code sequences wherein said plurality of values is located by a plurality of cell addresses;a code selector configured to select a first code sequence and a second code sequence from said plurality of code sequences during operation of said pressure sore detection system and wherein said code selector changes from said first code sequence to said second code sequence during said operation of said pressure sore detection system;an address sequencer coupled to said memory device and said code selector configured to provide cell addresses from said plurality of cell addresses corresponding to said first selected code sequence and said second selected code sequence;a signal generator coupled to said code selector configured to generate a first digital modulation signal and second digital modulation signal representing said first selected code sequence and said second selected code sequence;an optical illumination source of a first wavelength coupled to said signal generator configured to receive said first digital modulation signal and said second digital modulation signal and configured to transmit a first modulated optical signal of said first wavelength along an optical transmission path to said human tissue in response to said first digital modulation signal and transmit a second modulated optical signal of said first wavelength along said optical transmission path to said human tissue in response to said second digital modulation signal;a detector configured to be optically coupled to said human tissue configured to receive said first and second modulated optical signal after transmitting to said human tissue;a second signal generator configured to generate a third digital modulation signal representing a third code sequence, wherein said third code sequence and said second code sequence have the same ordering of code elements and begin at different code elements;a second optical illumination source of a second wavelength coupled to said second signal generator configured to receive said third digital modulation signal and configured to transmit a third modulated optical signal of said second wavelength to said human tissue in response to said third digital modulation signal;a processor coupled to said detector configured to derive a plurality of temporal transfer characteristics for said first, second and third modulated optical signals and configured to detect said pressure sores based on said temporal transfer characteristics; anda blinking indicator coupled to said processor configured to increase alert level by adjusting frequency of blinking based upon detection of said pressure sores. 2. The pressure sore detection system of claim 1 further comprising: a delay component optically coupled to said optical illumination source configured to increase length of sa... | Triple Ring Technologies, Inc. Ocela Technologies, Inc. | A61B005/1455 A61B005/00 G01J003/02 G01J003/10 G01J003/28 G01J003/44
US20110152710 | A1 | 2011-06-23 | ADAPTIVE BRAIN-COMPUTER INTERFACE DEVICE | 2010-01-26 | Disclosed herein is an adaptive brain-computer interface device. The brain-computer interface device includes a stimulus generation unit, a signal collection unit, a preprocessing unit, an analysis unit, and a stimuli sequence determination unit. The stimulus generation unit generates stimuli and applies the stimuli to a user. The signal collection unit records the user's Electroencephalogram (EEG) signals generated by the stimuli. The preprocessing unit extracts the feature of P300 on the basis of one of the stimulus from the EEG signals. The analysis unit determines whether P300 is present in the signal extracted by the preprocessing unit. The stimuli sequence determination unit determines the sequence of the stimuli of the stimulus generation unit by inferring a current state from the observations of the analysis unit and selecting an optimum stimulus for the current state. | 1. An adaptive brain-computer interface device, comprising: a stimulus generation unit for generating stimuli and applying the stimuli to a user;a signal collection unit for recording the user's Electroencephalogram (EEG) signals generated by the stimuli;a preprocessing unit for extracting a feature of P300 based on one of the stimulus from the EEG signals;an analysis unit for determining whether P300 is present in the signal extracted by the preprocessing unit; anda stimuli sequence determination unit for determining a sequence of the stimuli of the stimulus generation unit by inferring a current state from observations of the analysis unit and selecting an optimum stimulus for the current state. 2. The adaptive brain-computer interface device as set forth in claim 1, wherein the stimuli sequence determination unit determines an optimum action policy using a Partially Observable Markov Decision Process (POMDP). 3. The adaptive brain-computer interface device as set forth in claim 2, wherein the stimuli sequence determination unit comprises: a belief state update unit for inferring a distribution of probabilities of a current state from previous stimuli given to the user and observations corresponding to the respective stimuli; andan optimum stimulus selection unit for selecting and executing an optimum stimulus corresponding to a belief state of the belief state update unit, and transmitting an instruction or a message to an external device. 4. The adaptive brain-computer interface device as set forth in claim 1, wherein the stimuli sequence determination unit determines an optimum action policy using a delayed observation POMDP. 5. The adaptive brain-computer interface device as set forth in claim 1, wherein the stimuli sequence determination unit determines an optimum action policy using only actions except for actions which had been taken for a period for which repetition blindness caused by repeated stimuli occurred. 6. The adaptive brain-computer interface device as set forth in claim 5, wherein the optimum action policy is calculated using a value function defined by the following equation:V*(b)=maxA-A′[R(b,a)+γ∑zP(zb,a)V*(τ(b,a,z))]where A is a set of actions and A′ is a set of actions performed within 500 ms from a reference time. 7. The adaptive brain-computer interface device as set forth in claim 1, wherein the preprocessing unit removes noise by averaging the EEG signals or uses a P300 extraction algorithm such as a spatial filter algorithm or Mexican hat wavelet. 8. The adaptive brain-computer interface device as set forth in claim 1, wherein the analysis unit comprises a P300 classifier using a classification algorithm such as Fisher's linear discriminant, a stepwise linear discriminant analysis, or a support vector machine. | KOREA ADVANCED INSTITUTE OF SCIENCE AND TECHNOLOGY | A61B005/0478
US20110144458 | A1 | 2011-06-16 | Method and apparatus providing timed visual images during childbirth | 2010-07-20 | Disclosed is an apparatus for providing a visual image timed to coincide with uterine contractions to assist a expectant mother during childbirth labor. The apparatus is particularly useful directed to help those expectant mothers using a regional anesthesia, such as an epidural anesthesia, in the second stage of labor to push effectively. The apparatus consists of a computer system that converts signals obtained from strategically positioned sensors such as a tocodynamometer or intrauterine pressure catheter into biofeedback visual vignettes that are understandable to the expectant mother. The signals are processed into the visual vignettes or a similar pictorial image based upon the pressure of the uterine contractions so that the expectant mother may understand her contraction pattern. This in turn will allow the expectant mother to visually understand when to start and stop pushing during the second stage of labor. | 1. A medical apparatus for assisting a expectant mother during childbirth comprising: a computer having a display monitor;at least one sensor positioned on a expectant mother to acquire a uterine activity signal from a uterine contraction, said uterine activity signal coupled to said computer;a software program for operating the computer to determine a change in the uterine activity signal and converting said signal into a vignette image viewable on said display monitor;wherein said vignette image provides timing of the contraction pattern and visual instruction of when to start and stop pushing during labor by altering said vignette image. 2. The medical apparatus according to claim 1 wherein said uterine contraction are measurements proportional to the amount of pressure that the expectant mother's uterine contraction creates. 3. The medical apparatus according to claim 1 wherein said uterine contraction are measurements proportional to the amount of pressure the expectant mother creates through valsalva. 4. The medical apparatus according to claim 1 wherein said uterine contraction are measurements proportional to the amount of pressure the expectant mother creates through abdominal muscle contraction. 5. The medical apparatus according to claim 1 wherein said sensor is a tocodynamometer placed on the external surface of uterus in an area that can be palpated easily. 6. The medical apparatus according to claim 1 wherein said sensor is an internal pressure transducer placed into the uterus for measuring the internal pressure of the uterine contractions. 7. The medical apparatus according to claim 2 wherein a baseline measurement is used to visually depict a vignette in a first position that continues to change until a peak measurement is obtained placing said vignette in a second position. 8. The medical apparatus according to claim 7 wherein said vignette begins to disintegrate after said second position is obtained. 9. The medical apparatus according to claim 1 wherein an audio sound is produced when a preset peak measurement is reached, said audio sound increasing in tone as said measurement exceeds said peak. 10. The medical apparatus according to claim 7 wherein said vignette is displayed in a dynamic format in response to measurement change proportional to the pressure received by a transducer. 11. A medical apparatus for assisting a expectant mother during childbirth comprising: a computer having a display monitor;at least one internal pressure transducer placed into the uterus for measuring the internal pressure of the uterine contractions on a expectant mother to acquire a uterine activity signal from uterine contraction measurements that are proportional to the amount of pressure that the expectant mother's uterine contraction creates, said uterine activity signal coupled to said computer;a software program for operating the computer to determine a change in the uterine activity signal and converting said signal into a vigne... | - | A61B005/00
US20110160615 | A1 | 2011-06-30 | Swallowing Frequency Measurement | 2009-12-29 | Implementations for measurements of swallowing frequency are generally disclosed. | 1. A swallowing frequency measurement method, comprising: receiving sound data captured by a microphone; anddetermining whether the sound data has characteristics of swallowing sound, based at least in part on a sound waveform of the sound data. 2. The method as recited in claim 1, further comprising: incrementing a number of swallowing occurrences based at least in part on the determination that the sound data has the characteristics of swallowing sound. 3. The method as recited in claim 1, wherein the characteristics of swallowing sound include a period of silence before a sound event, a period of silence after the sound event, pulse width of the sound event, and number of pulses. 4. A swallowing monitoring system, comprising: a sensor unit configured to capture sound data; andan analysis unit configured to determine whether the sound data has characteristics of swallowing sound, based at least in part on a sound waveform of the sound data from the sensor unit. 5. A swallowing monitoring system as recited in claim 4, wherein the sensor unit further comprises a communication device configured to send the sound data to the analysis unit; and wherein the analysis unit further comprises a second communication device configured to receive the sound data from the sensor unit. 6. A swallowing monitoring system as recited in claim 4, wherein the analysis unit is configured to increment a number of swallowing occurrences based at least in part on the determination that the sound data has the characteristics of swallowing sound. 7. A swallowing monitoring system as recited in claim 4, wherein the characteristics of swallowing sound include a period of silence before a sound event, a period of silence after the sound event, pulse width of the sound event, and number of pulses. 8. A sensor unit for detecting swallowing, comprising: an elastic collar configured to removably secure the sensor unit to a neck portion;a sensor coupled to the elastic collar and oriented to be adjacent the neck portion when the elastic collar is removably secured, the sensor unit configured to capture sound data; anda data processor operatively associated with the sensor unit, the data processor configured to determine whether the sound data has characteristics of swallowing sound, based at least in part on a sound waveform of the sound data from the sensor unit. 9. A sensor unit as recited in claim 8, wherein the data processor is configured to increment a number of swallowing occurrences based at least in part on the determination that the sound data has the characteristics of swallowing sound. 10. A sensor unit as recited in claim 8, wherein the characteristics of swallowing sound include a period of silence before a sound event, a period of silence after the sound event, pulse width of the sound event, and number of pulses. 11. An article comprising: a signal bearing medium comprising machine-readable instructions stored thereon, which, if executed by one or more pr... | EMPIRE TECHNOLOGY DEVELOPMENT LLC | A61B005/103
US20110152706 | A1 | 2011-06-23 | METHOD AND APPARATUS FOR SENSING RESPIRATORY PRESSURE IN AN IMPLANTABLE STIMULATION SYSTEM | 2009-05-15 | A method and system for sensing respiration in the treatment of sleep apnea is disclosed. | 1. A method of treating a respiratory dysfunction, the method comprising: positioning a sensor portion of a lead, via an intercostal access point, into an implanted location in which the sensor portion is coupled relative to respiratory pressures of a pleural space;sensing, via the sensor portion at the implanted location, at least one of a temperature and the respiratory pressure; andidentifying a respiratory event based on a substantial change in the amplitude of at least one of the sensed temperature and the sensed respiratory pressure. 2. The method of claim 1, wherein the respiratory event includes an obstructive sleep apnea event. 3. The method of claim 1, comprising: triggering the application of an electrical stimulation signal, from an implantable pulse generator coupled to the sensor portion, based on the substantial change in the amplitude of at least one of the sensed temperature and the sensed respiratory pressure. 4. The method of claim 3, comprising: synchronizing the application of the electrical stimulation signal according to cyclical respiration waveform parameters, including at least one of:a portion of inspiratory phase;a portion of an expiratory phase; ora transition region between the respective inspiratory and expiratory phases. 5. The method of claim 1, wherein identifying the respiratory event comprises: using the sensed temperature to indicate a polarity of the sensed respiratory pressure to differentiate inspiratory phases from expiratory phases of sensed respiratory pressure. 6. The method of claim 1, wherein sensing at least one of the pressure and the temperature includes: providing a first piezo-electric transducer within the sensor portion to sense the pressure and a second piezo-electric transducer within the sensor portion to sense the temperature, independently from sensing the pressure; andmaintaining electrical independence between the first and second piezo-electric transducers. 7. The method of claim 6, wherein identifying the respiratory event includes: identifying a substantial increase in the cyclical amplitudes of the sensed pressure relative to a pressure threshold and identifying, at generally the same time, a substantial decrease in the cyclical amplitudes of the sensed temperature. 8. The method of claim 6, comprising: summing a signal of the sensed pressure and a signal of the sensed temperature to produce a measured respiratory signal;wherein identifying the respiratory event includes:identifying a substantial change in the amplitude of the measured respiratory signal. 9. The method of claim 1, wherein positioning the sensor portion comprises: providing the sensor portion as a directional sensing element including a single side of the sensor portion configured to sense the respiratory pressure;facing the directional sensing element directly toward the lungs from the extrapleural position or the intrapleural position. 10. The method of claim 1, wherein positioning the sensor portion compri... | INSPIRE MEDICAL SYSTEMS, INC. | A61B005/08
US20110118620 | A1 | 2011-05-19 | Method and system for monitoring and displaying physiological conditions | 2010-10-18 | A method for providing an indication of a state of awareness for a patient, includes the steps of arranging data of an EEG and EMG power spectrogram to provide power versus frequency in a log-log arrangement; calculating a first best-fit line for a lower frequency region of the EEG power spectrogram; calculating at least a second best-fit line for a higher frequency region of the EEG power spectrogram. The display of these lines is augmented by displaying a template that identifies different regions on the display that help confirm the state of the patient. Secondly, the time domain EEG signals can be filtered and displayed such that different frequency bands can be simultaneously displayed or a single frequency band can be displayed according to different time scales. | 1. A method for providing an indication of a state of awareness for a patient, comprising the steps of: acquiring an EEG signal;filtering the EEG signal to generate a first frequency band of signals;filtering the EEG signal to generate a second frequency band of signals; andconcurrently displaying the first and second frequency bands of signals. 2. The method of claim 1, wherein concurrently displaying includes: superimposing the first and second frequency bands of signals within a same display window. 3. The method of claim 1, wherein concurrently displaying includes: displaying the first and second frequency bands of signals in separate display windows. 4. The method of claim 1, further comprising: filtering the EEG signal to generate a third frequency band of signals; andconcurrently displaying the first, second, and third frequency bands of signals. 5. The method of claim 4, further comprising: filtering the EEG signal to generate one or more additional frequency bands of signals; andconcurrently displaying at least one of the one or more additional frequency bands of signals. 6. The method of claim 1, further comprising: selecting a range of frequencies for the first frequency band based on one or more attributes of the EEG signal. 7. The method of claim 6, wherein the one or more attributes includes a peak frequency within one of the first frequency band of signals and the second frequency band of signals. 8. The method of claim 7, wherein the one or more attributes includes an alpha peak. 9. An apparatus for providing an indication of a state of awareness for a patient, comprising: a receiver configured to acquire an EEG signala first filter coupled with the receiver and configured to filter the EEG signal to generate a first frequency band of signals;a second filter coupled with the receiver and configured to filter the EEG signal to generate a second frequency band of signals; anda visual display screen configured to concurrently display the first and second frequency bands of signals. 10. The apparatus of claim 9, wherein the visual display screen includes both first and second frequency bands of signals superimposed on one another. 11. The apparatus of claim 9, wherein the visual display screen includes a first window displaying the first frequency band of signals and a second window displaying the second frequency band of signals. 12. A method for providing an indication of a state of awareness for a patient, comprising the steps of: acquiring an EEG signal;filtering the EEG signal to generate a first frequency band of signals;displaying the first frequency band of signals in a first time scale; andconcurrently displaying the first frequency band of signals in a second time scale, different than the first time scale. 13. The method of claim 12, wherein the first time scale is selected such that individual oscillations within spindle activity of the EEG signal is visible and the second time scale is selected such that packe... | - | A61B005/048
US20110105852 | A1 | 2011-05-05 | USING DATA IMPUTATION TO DETERMINE AND RANK OF RISKS OF HEALTH OUTCOMES | 2009-11-03 | Techniques for generating prediction of risks of medical outcomes and benefit scores for medical interventions, with imputation of missing patient data values, are disclosed. Apparatus or computer program products may be configured to receive a patient record for the patient from a database of a data storage unit, wherein one or more demographic data values or biometric data values in the patient record are missing or have null values; create and store a plurality of clone patient records in the database; impute a plurality of different substitute demographic data values or biometric data values and substitute a different one of the plurality of substitute values into each one of the clone patient records; determine, create and store a first metric, based at least in part on the clone patient records, wherein the first metric comprises a current health related metric for the patient; determine, create and store one or more medical intervention metrics, each based at least in part on an associated medical intervention and the clone patient records, representing a predicted health related metric for the patient when the associated medical intervention is performed; transform the database by updating the patient record to include the first metric and the one or more medical intervention metrics. | 1. A data processing apparatus, comprising: one or more processors;a data storage unit storing a database of patient information associated with a patient of a healthcare provider;query execution logic and an imputation engine coupled to the one or more processors and the data storage unit, and configured to: receive a patient record for the patient from the database of the data storage unit, wherein one or more demographic data values or biometric data values in the patient record are missing or have null values;create and store a plurality of clone patient records in memory;impute different substitute demographic data values or biometric data values and substitute a different one of the substitute values into each one of the clone patient records;determine, create and store a first metric, based at least in part on the clone patient records, wherein the first metric comprises a current health related metric for the patient;determine, create and store one or more medical intervention metrics, each based at least in part on an associated medical intervention and the clone patient records, representing a predicted health related metric for the patient when the associated medical intervention is performed;transform the database by updating the patient record to include the first metric and the one or more medical intervention metrics. 2. The apparatus of claim 1, wherein the query execution logic is configured to determine each of the one or more medical intervention metrics as a risk of a specified medical outcome. 3. The apparatus of claim 2, wherein the risk is any of myocardial infarction, stroke, onset of diabetes mellitus, or onset of complications of diabetes mellitus. 4. The apparatus of claim 1, wherein the query execution logic is further configured to determine one or more benefit scores comprising weighted sums of likelihoods of different medical outcomes within a specified time period. 5. The apparatus of claim 4, wherein each of the benefit scores measures a difference in a first predicted quality of life score in the specified time period if the patient receives an associated intervention compared to a second predicted quality of life score in the same specified time period if the patient does not. 6. The apparatus of claim 1, wherein the query execution logic is further configured to determine each of the one or more medical intervention metrics by determining average risks of a specified medical outcome and standard deviation of the risks across all the clone patient records. 7. The apparatus of claim 1, wherein the query execution logic is further configured to transform each of the one or more medical intervention metrics according to one or more medical rules when an action specified by any of the medical intervention metrics is inconsistent with the medical rules. 8. The apparatus of claim 1, wherein one of the intervention metrics is a combination metric based on a combination of two or more of the intervention metrics ... | - | A61B005/00 G06Q050/00
US08321002 | B2 | 2012-11-27 | Arrhythmia discrimination using ECG's sensed from multiple implanted electrodes | 2011-02-15 | Cardiac monitoring and/or stimulation methods and systems provide for monitoring, diagnosing, defibrillation and pacing therapies, or a combination of these capabilities, including cardiac systems incorporating or cooperating with neuro-stimulating devices, drug pumps, or other therapies. Embodiments of the present invention relate generally to implantable medical devices employing automated cardiac activation sequence monitoring and/or tracking for arrhythmia discrimination. Embodiments of the invention are directed to devices and methods involving sensing a plurality of composite cardiac signals using a plurality of implantable electrodes. A source separation is performed using the sensed plurality of composite cardiac signals and the separation produces one or more cardiac signal vectors associated with one or more cardiac activation sequences that is indicative of ischemia. A change of the one or more cardiac signal vectors is detected using the one or more cardiac signal vectors. Cardiac arrhythmias are discriminated using the one or more cardiac signal vectors. | 1. A processor-implemented method, comprising: sensing a plurality of composite cardiac signals at a plurality of patient-internal locations for a multiplicity of cardiac cycles;performing a plurality of source separations using the sensed composite cardiac signals;producing a cardiac signal vector indicative of a cardiac activation sequence of an area of the patient's heart based on each of the plurality of source separations;measuring an orientation of the cardiac signal vectors relative to a baseline established for the patient;detecting ventricular tachycardia (VT) based on the cardiac signal vector orientation measurements; andin response to detecting VT: determining variation of the cardiac signal vector orientation measurements, wherein determining the variation of the cardiac signal vector orientation measurements comprises: performing a subsequent source separation using sensed composite cardiac signals for each of a plurality of consecutive cardiac cycles after VT is detected;producing a subsequent cardiac signal vector based on each of the subsequent source separations;measuring an orientation of each subsequent cardiac signal vector relative to the baseline; anddetermining the variation of the subsequent cardiac signal vector orientation measurements; anddiscriminating between polymorphic ventricular tachycardia (PVT) and monomorphic ventricular tachycardia (MVT) based on the variation of the cardiac signal vector orientation measurements, wherein at least some of the method is implemented by a processor. 2. The method according to claim 1, wherein determining the variation comprises determining the variation of cardiac signal vector orientation measurements for a plurality of consecutive cardiac cycles subsequent to VT detection. 3. The method according to claim 1, wherein discriminating between PVT and MVT comprises: comparing the variation to a pre-determined value; anddiscriminating between PVT and MVT based on the comparison. 4. The method according to claim 1, wherein discriminating between PVT and MVT comprises: comparing the variation to a pre-determined value;determining presence of PVT if the variation is greater than the pre-determined value; anddetermining presence of MVT if the variation is less than the pre-determined value. 5. The method according to claim 1, comprising establishing the baseline by performing an initial source separation using sensed composite cardiac signals. 6. The method according to claim 1, comprising establishing the baseline using clinical data. 7. The method according to claim 1, further comprising differentiating between left ventricular tachycardia (LVT) and right ventricular tachycardia (RVT) in response to detecting MVT. 8. The method according to claim 7, comprising: determining a shift in orientation of axes of the cardiac signal vectors relative to the baseline; anddifferentiating between LVT and RVT based on a direction of the shift of the cardiac signal vector axes. 9. The metho... | Cardiac Pacemakers, Inc. | A61B005/0464
US07981061 | B2 | 2011-07-19 | Power apparatus, power system, and power control method | 2010-05-12 | A power apparatus comprising includes: a moving section that moves in accordance with an action of a body part; a motor that outputs, during action of the body part in a first action direction, power to the moving section so as to cause the moving section to move in the first action direction and that recovers, during action of the body part in a second action direction, power generated by the movement of the moving section in the second action direction; and a capacitor that stores regenerative electric power generated by the recovery of the power and that supplies the stored regenerative electric power to the motor when the motor outputs the power to the moving section. | 1. A power apparatus comprising: a moving device configured to move in accordance with an action of a body part;a motor that outputs, during an action of the body part in a first action direction, power to the moving device so as to cause the moving device to move in the first action direction and that recovers, during an action of the body part in a second action direction, power generated by the movement of the moving device in the second action direction;a capacitor that stores regenerative electric power generated by the recovery of the power and that supplies the stored regenerative electric power to the motor when the motor outputs the power to the moving device;a battery that supplies battery electric power to the motor when the motor outputs the power to the moving device; anda controlling device configured to control switching between a first mode in which the motor outputs the power to the moving device and a second mode in which the power of the motor is recovered;wherein, in the first mode, when the regenerative electric power stored in the capacitor is greater than or equal to a predetermined value, the controlling device performs control so as to block supply of the battery electric power from the battery to the motor and so as to supply the regenerative electric power from the capacitor to the motor, and when the regenerative electric power stored in the capacitor is smaller than the predetermined value, the controlling device performs control so as to supply the battery electric power from the battery to the motor, andwherein, when the moving device is not moving and the regenerative electric power is stored in the capacitor, the controlling device performs control so as to supply the regenerative electric power from the capacitor to the battery. 2. The power apparatus according to claim 1, wherein, in the second mode, the controlling device performs control so as to supply the regenerative electric power to the capacitor when a voltage value of the capacitor is not a maximum value and so as to supply the regenerative electric power to a load when the voltage value of the capacitor is the maximum value. 3. The power apparatus according to claim 1, further comprising an acceleration sensor that measures acceleration of the action of the body part, wherein, when the measured acceleration contains a component in a direction opposite to a direction of gravity, the controlling device switches a mode to the first mode to perform control, and when the measured acceleration contains a component in the direction of gravity, the controlling device switches a mode to the second mode to perform control. 4. The power apparatus according to claim 1, further comprising an electromyograph that measures a myoelectric potential of the body part, wherein, in the second mode, the controlling device controls a recovery load of the motor in accordance with the measured myoelectric potential so that no myoelectric potential is generated. 5. The powe... | Empire Technology Development LLC | A61F005/00 A61H001/00
US07935063 | B2 | 2011-05-03 | System and method for breaking a sterility seal to engage a lancet | 2006-10-20 | An integrated lancing test strip includes a test strip and a lancet packet coupled to the test strip. The lancet packet includes a sterility sheet enclosing a lancet to maintain the sterility of the lancet and prevent cross-contamination between the test strip and the lancet. The sterility sheet allows the lancet to be sterilized separately from the test strip. In one form, a firing mechanism has a blade that punctures and slits the sterility sheet during the firing cycle to reduce the chance of cross contamination between the lancet and the test strip prior to use. This allows an operator to load and subsequently remove an unused integrated lancing test strip without compromising the sterility of the lancet. The lancet includes an engagement opening in which the blade of the firing mechanism engages the lancet. The engagement opening is configured to reduce the chance of the blade becoming dull. | 1. An apparatus, comprising: a lancet defining an engagement opening, wherein the engagement opening is in the lancet;a firing mechanism to fire the lancet, the firing mechanism including an engagement blade that is configured to engage with the engagement opening of the lancet, the engagement blade having one or more cutting edges;the engagement opening having one or more cut out portions at the one or more cutting edges when the blade engages the engagement opening to minimize dulling of the one or more cutting edges of the engagement blade; andwherein the lancet is sealed inside a sterility sheet to form a lancet packet. 2. The apparatus of claim 1, wherein the engagement opening of the lancet has one or more engagement protrusions configured to contact the blade at a location other than at the one or more cutting edges. 3. The apparatus of claim 2, wherein the one or more engagement protrusions are rounded to prevent damage to the blade. 4. The apparatus of claim 1, wherein the one or more cut out portions have a circular shape. 5. The apparatus of claim 4, wherein the engagement opening has a recessed section that is diamond shaped. 6. The apparatus of claim 5, wherein: the engagement blade has two of the cutting edges; andthe engagement opening has two of the circular shaped cut out portions that are disposed on opposite ends of the diamond shaped recessed section. 7. The apparatus of claim 1, wherein the one or more cutting edges of the engagement blade are configured to cut one or more slits in the sterility sheet of the lancet packet to promote smooth actuation of the lancet. 8. The apparatus of claim 1, wherein the engagement blade of the firing mechanism has a tip configured to puncture the sterility sheet during engagement with the engagement opening of the lancet. 9. The apparatus of claim 8, wherein the firing mechanism is configured to puncture the sterility sheet with the tip of the blade during firing of the lancet. 10. The apparatus of claim 9, wherein the firing mechanism includes: a firing arm secured to the engagement blade, the firing arm including one or more cam pins;a plunger defining one or more cam slots in which the one or more cam pins are slidably received; andwherein the one or more cam slots are shaped to move the engagement blade to puncture the sterility sheet with the tip of the blade when the plunger is fired. 11. The apparatus of claim 1, further comprising a test strip attached to the lancet packet. 12. The apparatus of claim 1, wherein the lancet defines a capillary groove configured to draw body fluid via capillary action. 13. An apparatus, comprising: a lancet defining an engagement opening where the lancet is coupled to a firing mechanism, wherein the engagement opening is in the lancet; andwherein the engagement opening is shaped like a rounded diamond with circular cut outs at opposite ends of the diamond. 14. The apparatus of claim 13, further comprising a test strip coupled to the lancet ... | Roche Diagnostics Operations, Inc. | A61B005/00
US20110144472 | A1 | 2011-06-16 | METHODS OF GRAVITY AND/OR MAGNETIC HOLOGRAPHIC IMAGING USING VECTOR AND/OR TENSOR DATA | 2010-09-10 | A method for holographic imaging an object having density and/or magnetization is described, the object being located in an examined medium using potential field data including but not limited to gravity and/or magnetic total field and/or vector and/or tensor data. The potential field sensors may measure the gravity and/or magnetic total field and/or vector and/or tensor data at least one receiving position with respect to the examined medium. At least one component of the measured potential field in at least one receiver location (potential field data) may be used as at least one artificial source of the potential field data. Artificial sources may produce a back-propagating (migration) field. An integrated sensitivity of the potential field data to density and/or magnetization perturbation may be calculated. A spatial weighting of at least one of the back-scattering (migration) fields may form a potential field holographic image. At least one desired property of the medium, such as density and/or magnetization, may be derived from this holographic image. | 1. A method for imaging an object having density and/or magnetization, the object being located in an examined medium, the method comprising: a. placing at least one actual gravity vector and/or tensor (GVT) and/or magnetic vector and/or tensor (MVT) sensor at least one receiving position with respect to the examined medium;b. measuring at least one GVT and/or MVT component of the GVT and/or MVT data fields with at least one actual GVT and/or MVT sensor;c. conceptually replacing the at least one actual GVT and/or MVT sensor with at least one conceptual source of GVT and/or MVT data, the at least one conceptual source having a scalar density and/or vector magnetization which directly corresponds to the at least one measured GVT and/or MVT component;d. obtaining a back-propagating (migration) tensor field equivalent to that produced by the at least one conceptual source that replaced the at least one actual GVT and/or MVT sensor;e. obtaining an integrated sensitivity of the GVT and/or MVT data acquisition system by estimating a least square norm of values of perturbation of the at least one GVT and/or MVT component at the at least one receiving position due to a density and/or magnetization perturbation at a specific local area of the examined medium; andf. producing a holographic image of the object by spatially weighting the back-propagating (migration) field. 2. The method of claim 1, wherein the at least one actual GVT and/or MVT sensor comprises a plurality of GVT and/or MVT sensors arranged in an array above and/or on the surface and/or within the volume of the examined medium. 3. The method of claim 2, wherein the plurality of sensors include both GVT and MVT sensors. 4. The method of claim 1, wherein the measured at least one GVT and/or MVT component of GVT and/or MVT data is input to a processor, and the processor includes executable instructions to: analyze said GVT and/or MVT fields;compute the back-propagating (migration) tensor field by simulating the replacement of the actual GVT and/or MVT sensors with an array of conceptual sources of the GVT and/or MVT data, each conceptual source with a scalar density and/or vector magnetization which is determined by the actually measured GVT and/or MVT components measured in the locations of said actual GVT and/or MVT sensors;compute the integrated sensitivity of the GVT and/or MVT data acquisition system; andconstruct a volume image of density and/or magnetization by calculating a spatial distribution of said back-propagating (migration) fields weighted with said integrated sensitivity. 5. The method of claim 1, wherein the GVT and/or MVT data is gravity total field and/or vector and/or tensor data and/or magnetic total field and/or vector and/or tensor data. 6. The method of claim 1, further comprising deriving one or more properties of the examined medium from the holographic image of the object. 7. The method of claim 6, wherein the one or more properties include density and/or magneti... | TECHNOIMAGING, LLC | A61B005/05 G01R033/02 G01C009/10
US20110124976 | A1 | 2011-05-26 | MODEL ENHANCED IMAGING | 2009-07-22 | A therapy treatment response simulator includes a modeler (202) that generates a model of a structure of an object or subject based on information about the object or subject and a predictor (204) that generates a prediction indicative of how the structure is likely to respond to treatment based on the model and a therapy treatment plan. In another aspect, a system includes performing a patient state determining in silico simulation for a patient using a candidate set of parameters corresponding to another patient and producing a first signal indicative of a predicted state of the patient, and generating a second signal indicative of whether the candidate set of parameters are suitable for the patient based on a known state of the patient. | 1. A therapy treatment response simulator, comprising: a modeler that generates a model of a structure of an object or subject to be treated based on information about the object or subject; anda predictor that generates a predicted response indicative of how the structure is likely to respond to the treatment based on the model and a therapy treatment plan. 2. The simulator of claim 1, further including: a parameter map generator that generates a parameter map that includes quantitative information indicative of the predicted response. 3. The simulator of claim 2, wherein the quantitative information includes quantitative information indicative of tracer uptake related to the structure. 4. The simulator of claim 3, wherein the structure includes macrophages processing cells killed by the treatment. 5. (canceled) 6. The simulator of claim 4, wherein the tracer uptake by the macrophages is similar to tracer uptake by a tumor being treated, wherein the macrophages process normal cells around tumor that have been killed by the treatment. 7. The simulator of claim 1, wherein the treatment includes radiation therapy, chemotherapy, particle therapy, high intensity focused ultrasound, ablation or combinations thereof. 8. The simulator of claim 1, wherein the information about the object or subject includes image data generated from data acquired prior to the treatment. 9. The simulator of claim 1, wherein the information about the object or subject includes one or more of histological data, patient health information, medical history, genetics, laboratory test results, or pathological information. 10. The simulator of claim 1, wherein at least one of the model, the prediction or the parameter map is generated in silico. 11. A therapy system, comprising: a treatment response simulator that generates a parameter map that includes quantitative information indicative of how a first structure of an object or subject is likely to respond to treatment based on a model of the object or subject and a therapy treatment plan for the object or subject; anda treatment monitoring system that enhances image data generated from data acquired after the treatment based on the parameter map. 12. (canceled) 13. The therapy system of claim 11, wherein the image data includes information indicative of tracer uptake by the first structure and at least one different structure of the object or subject. 14. (canceled) 15. The therapy system of claim 11, wherein the treatment monitoring system includes: an image data processor that processes the image data to generate quantitative information about two or more structures receiving treatment, wherein one of the two or more treated structures includes the first structure; andan image data enhancer that enhances a second structure of the two or more structures in the image data by subtracting the quantitative information about the first structure from the image data. 16. The therapy system of claim 11, wherein the tre... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/00 G06G007/60
US20110130635 | A1 | 2011-06-02 | Medical self-service | 2009-11-30 | Techniques for medical self-service are provided. A kiosk interacts with a patient to establish a patient identity, acquire medical records for the patient, and acquire one or more medical vital readings for the patient. The kiosk then packages the information in a summary that is electronically provided to medical personnel before a patient is addressed by the medical personnel at a medical facility. | 1. A processor-implemented method to execute on a processor configured to execute the method, comprising: identifying, via a biometric sensor, an identity of a medical patient;accessing medical data, over a network, to acquire medical records for the medical patient using the identity;interacting with the medical patient via one or more automated sensors to acquire one or more readings for medical vitals experienced by the medical patient; andautomatically assembling the medical records and the readings into an information packet for medical personnel to address the medical patient. 2. The method of claim 1, wherein identifying further includes acquiring one or more of the following from the medical patient via the biometric sensor to establish the identity: a fingerprint reading, a retinal scan, and a facial scan. 3. The method of claim 1, wherein interacting further includes using the automated sensors to acquire the one or more readings as one or more of the following: a height, a weight, a blood pressure, a pulse rate, an oxygenation level, a respiration rate, and a blood-sugar level. 4. The method of claim 3, wherein using further includes automatically calibrating at least one of the automated sensors based on the medical records to acquire the blood-sugar level of the medical patient. 5. The method of claim 1, wherein interacting further includes automatically uploading some of the readings or a history of some readings from a portable medical device in the possession of the medical patient. 6. The method of claim 5, wherein automatically uploading further includes automatically acquiring a medical report, an x-ray, a doctor's written order, a doctor's prescription, or transferred medical records no available in the medical data from a portable medical device in the possession of the medical patient. 7. The method of claim 1 further comprising, processing the method as a standalone medical kiosk at a medical facility. 8. The method of claim 1 further comprising, processing the method as a portable device operated by a medic on behalf of the medical patient. 9. The method of claim 1 further comprising, automatically prioritizing a schedule to address the medical patient by the medical personnel in response to the readings and the medical records. 10. A processor-implemented method to execute on a processor configured to execute the method, the method, comprising: automatically, interacting with a patient, via a kiosk having the processor, the interaction including:establishing a patient identity for the patent via biometric sensors integrated into the kiosk;acquiring medical records for the patent from one or more data sources via one or more network connections to and from the kiosk; andassembling an intake packet for medical personal to address the patient using the patient identity and the medical records. 11. The method of claim 10, wherein establishing further includes taking biometric readings and searching a biometric datab... | NCR Corporation | G06Q050/00 G06Q010/00 A61B005/00 G06F017/30 G05B019/00
US07962203 | B2 | 2011-06-14 | Arrhythmia display | 2003-12-29 | Systems, devices, structures, and methods are provided to present a visual display based on data from an implantable medical device. The display includes a chart showing the frequency of a detected type of arrhythmia over a predetermined period of time. | 1. A programmer for an implantable medical device, the programmer comprising: a receiver to receive from the implantable medical device information indicative of episodes of a plurality of types of arrhythmia detected by the implantable medical device, the plurality of types of arrhythmia including at least ventricular tachycardia and ventricular fibrillation; anda display including at least one column graph having a plurality of distinguishing types of column symbol, the distinguishing types each including a plurality of column symbols associated with one type of arrhythmia of the plurality of types of arrhythmia and indicative of a frequency of episodes of the one type of arrhythmia detected over a second predetermined period including a plurality of first predetermined periods, wherein each column symbol of the plurality of column symbols indicates the frequency of episodes of the one type of arrhythmia detected over a first predetermined period of the plurality of first predetermined periods. 2. The programmer of claim 1, further comprising a transmitter to transmit detection zones of interest each classifying one type of arrhythmia of the plurality of types of arrhythmia to the implantable medical device. 3. The programmer of claim 1, wherein the display comprises a legend associating a text label with a sample of each type of the plurality of distinguishing types of column symbol. 4. The programmer of claim 3, wherein the display comprises numerical data labels each indicative of a number of episodes of the plurality of types of arrhythmia detected over each of the one or more first predetermined periods and the second predetermined period. 5. The programmer of claim 3, wherein the column graph further comprises grid lines to visually communicate a level of frequency with respect to the column symbols. 6. The programmer of claim 1, wherein the plurality of distinguishing types of column symbol comprises column symbols diverse in color. 7. The programmer of claim 6, wherein the plurality of distinguishing types of column symbol comprises column symbols diverse in gray scale. 8. The programmer of claim 1, wherein the plurality of distinguishing types of column symbol comprises column symbols diverse in pattern. 9. The programmer of claim 8, wherein the plurality of distinguishing types of column symbol comprises column symbols diverse in pattern and color. 10. The programmer of claim 1, wherein the display comprises an electrogram display to present at least one electrogram. 11. The programmer of claim 10, wherein the electrogram display comprises presentation of markers including labels each indicating a type of arrhythmia on a region of the electrogram associated with the type of arrhythmia. 12. The programmer of claim 1, wherein the at least one column graph comprises a first column graph having the plurality of distinguishing types of column symbol, the distinguishing types each associated with one type of ventricular arrhythmi... | Cardiac Pacemakers, Inc. | A61B005/044
US07959539 | B2 | 2011-06-14 | Body motion detection device capable of properly managing information on number of steps in walking exercise | 2007-10-02 | A Nordic walking pole has a transmission device attached thereto and a switch thereof is pushed to transmit a signal indicating the start or end of Nordic walking. A pedometer receives the signal from the transmission device to thereby detect that the start or end of Nordic walking is instructed (S107). Assuming that the Nordic mode is specified when the current walking mode is a normal walking mode, the parameter for Nordic walking is set as a parameter for the arithmetic operation (S111). Assuming that the normal walking mode is specified when the current walking mode is a Nordic mode, the parameter for the normal walking mode is set as a parameter for the arithmetic operation (S113). | 1. A body motion detection device comprising: a count unit counting a number of steps in a walking mode by detecting body motion, said walking mode including a first walking mode and a second walking mode;a calculation unit calculating an amount of exercise based on said counted number of steps;a receiving unit receiving from another device a signal indicating start or end of said second walking mode, said signal containing information identifying said another device; anda setting unit setting, as a parameter used in said calculation unit, a parameter for said first walking mode when said signal indicates said end of said second walking mode and a parameter for said second walking mode when said signal indicates said start of said second walking mode. 2. The body motion detection device according to claim 1, further comprising a first storage unit storing the parameter for said first walking mode and the parameter for said second walking mode, wherein said setting unit reads the parameter for said first walking mode or the parameter for said second walking mode from said first storage unit for setting as a parameter used in said calculation unit. 3. The body motion detection device according to claim 1, further comprising a second storage unit storing the parameter for said first walking mode, anda third storage unit storing a coefficient for calculating the parameter for said second walking mode based on the parameter for said first walking mode, whereinsaid setting unit reads said coefficient from said third storage unit and calculates the parameter for said second walking mode from the parameter for said first walking mode based on said signal for setting as a parameter used in said calculation unit. 4. The body motion detection device according to claim 1, further comprising a display unit displaying said calculated amount of exercise, anda display control unit controlling to display a mark on said display unit identifying whether said walking mode is said first walking mode or said second walking mode. 5. The body motion detection device according to claim 1, wherein said setting unit sets an equation for said first walking mode or an equation for said second walking mode as an equation used in said calculation unit. 6. The body motion detection device according to claim 1, further comprising a storing unit storing said calculated amount of exercise in a storage area, wherein said storing unit stores the amount of exercise calculated using the parameter for said first walking mode in a first area of said storage area and stores the amount of exercise calculated using the parameter for said second walking mode in a second area of said storage area. 7. The body motion detection device according to claim 1, wherein said setting unit sets a parameter corresponding to said another device as a parameter for said second walking mode based on said information contained in said signal. 8. The body motion detection device according to claim 1, ... | Omron Healthcare Co., Ltd. | A63B071/00 A61B005/103 G01C021/00
US08632485 | B2 | 2014-01-21 | Patient treatment and monitoring systems and methods | 2009-11-05 | The invention provides apparatus and methods for home or other remote delivery of health care that gather subjective and objective measures of patient health and treatment, analyzing them (e.g., correlating them with one another and/or with norms) and reporting them to aid in on-going patient diagnosis and treatment (both on acute and chronic bases), as well as to aid physicians, nurses and other caregivers in decision support, monitoring treatment compliance, facilitating regulatory compliance, billing, and so forth. Thus, for example, in some aspects a health care delivery device comprising a medical treatment apparatus, such as a home hemodialysis or home peritoneal dialysis unit, that is coupled to a processor. The processor generates patient queries in connection with treatments rendered by the dialysis equipment (or other treatment apparatus). The queries are directed, at least in part, to subjective topics, such as the state of the patient's mental health and well being, quality of life, degrees of pain, views on success of therapy, and so forth. They are presented on an LCD screen or other output device coupled to the processor to elicit responses on a keyboard or other input device, also coupled to the processor. | 1. A health care delivery device comprising: A. medical treatment apparatus having one or more sensors that take biometric readings of a patient receiving medical treatment,B. an output device that presents information or other output for the patient,C. an input device that accepts information or other input in response to a query, andD. a processor that is coupled to the medical treatment apparatus and that generates subjective queries in connection with treatment of the patient thereby, the processor being coupled to the output device for presentation of those subjective queries for the patient and being coupled to the input device for accepting responses to those subjective queries, the processor comprising an analysis module that correlates the patient's responses and the biometric readings so as to monitor at least one of the treatment of the patient and operation of the medical treatment apparatus, andE. wherein the subjective queries include one or more of the patient's mental health and well being, quality of life, degrees of pain, and views on success of therapy. 2. The health device of claim 1, wherein the medical treatment apparatus comprises dialysis equipment. 3. The health device of claim 2, wherein the dialysis equipment comprises equipment for any of hemodialysis and peritoneal dialysis. 4. The health device of claim 1, adapted for treatment of the patient in a patient's home or other location remote from a dialysis treatment center. 5. The health device of claim 1, comprising communication functionality that transmits messages, input by the patient via the input device, to a remote care provider. 6. The health device of claim 5, the communication functionality accepts messages transmitted from a remote care provider and that presents those messages on the output device. 7. The health device of claim 6, wherein the messages are text messages. 8. A health care delivery system comprising: A. a digital data processing system,B. a health care delivery device disposed remotely from the digital data processing system, the health care delivery device being coupled for at least occasional communication with the digital data processing system and including i. medical treatment apparatus having one or more sensors that sense operating conditions of the medical treatment apparatus as the patient receives medical treatment therefrom,ii. an output device that presents information or other output for a patient,iii. an input device that accepts information or other input in response to a query,iv. a processor that is coupled to the medical treatment apparatus and that generates subjective queries in connection with treatment of the patient thereby, the processor being coupled to the output device for presentation of those subjective queries for the patient and being coupled to the input device for accepting responses to those subjective queries, the processor comprising an analysis module that correlates the patient's responses and the op... | Fresenius Medical Care Holdings, Inc. | A61M037/00 C02F001/44 G06Q050/00
US08812083 | B2 | 2014-08-19 | Methods of optically monitoring wound healing | 2009-04-21 | Optical changes of tissue during wound healing measured by Near Infrared and Diffuse Reflectance Spectroscopy are shown to correlate with histologic changes. Near Infrared absorption coefficient correlated with blood vessel in-growth over time, while Diffuse Reflectance Spectroscopy (DRS) data correlated with collagen concentration. Changes of optical properties of wound tissue at greater depths are also quantified by Diffuse Photon Density Wave (DPDW) methodology at near infrared wavelengths. The diffusion equation for semi-infinite media is used to calculate the absorption and scattering coefficients based on measurements of phase and amplitude with a frequency domain or time domain device. An increase in the absorption and scattering coefficients and a decrease in blood saturation of the wounds compared to the non wounded sites was observed. The changes correlated with the healing stage of the wound. The methodologies used to collect information regarding the healing state of a wound may be used to clinically assess the efficacy of wound healing agents in a patient (e.g., a diabetic) and as a non-invasive method to detect the progress of wound healing, particularly chronic wounds due to diabetes. The methodology applies to ischemic environments, impaired healing states, and emerging subsurface tissue deterioration, such as in pressure ulcers, venous ulcers, and ubiquitous ulcers. | 1. A method of determining if a wound in a human patient is healing, comprising the steps of: illuminating the wound tissue in the patient with light from a light source;measuring over time less than ten weeks one or both of the amplitude and phase shift of the light as it propagates through the wound tissue; wherein the illuminating and measuring steps are performed using a diffuse photon density wave (DPDW) device that is a continuous wave, a frequency domain, or a time domain measurement device,calculating optical absorption coefficients using the measured amplitude or phase shift of the light;quantifying the concentration of the oxygenated hemoglobin or the total hemoglobin in the wound tissue over a period of time of less than ten weeks from the optical absorption coefficients; anddetermining the rate of change of the oxygenated hemoglobin or total hemoglobin concentration over the period of time, wherein a negative rate of change of either the oxygenated hemoglobin or total hemoglobin indicates the wound is healing and a negligible or positive rate of change indicates the wound is not healing. 2. The method of claim 1, wherein the light has a wavelength between 650 and 870 nm. 3. The method of claim 1, wherein the light has a near infrared wavelength selected from the group consisting of 685 nm light, 780 nm light, 830 nm light, and 950 nm light. 4. The method of claim 1, wherein the wound is selected from the group consisting of pressure ulcers, diabetic foot ulcers, ubiquitous ulcers, and venous ulcers. 5. The method of claim 1, further comprising differentiating a healing wound from a non-healing wound using the rate of change of total hemoglobin concentration over the period of time. 6. The method of claim 1, wherein the measurements are taken from at least 2 mm below the surface of the wound. 7. The method of claim 1, comprising determining the statistical variability of the measurements, wherein a higher degree of statistical variability between consecutive measurements relative to the variability in a healing wound indicates a non-healing wound. 8. The method of claim 1, wherein the period of time is less than five weeks. 9. The method of claim 1, wherein the concentration is measured up to five times over the period of time. 10. The method of claim 1, wherein the threshold between the negative rate of change and the negligible rate of change is a change of −0.003 cm−1/wk in the optical absorption coefficient. | Philadelphia Health & Education Corporation Drexel University | A61B006/00
US08790271 | B2 | 2014-07-29 | Portable device for calculating consumed calories | 2010-12-01 | A portable device for measuring a consumed calories includes: a respiration measurement device measuring an amount of air inhaled into a user's nose to acquire and output a respiration signal; a plurality of movement detection devices acquiring and outputting acceleration signals reflecting (or indicating) the magnitude and direction of a movement of each part of the user's body; and a controller recognizing the amount of a user's movements and movement patterns by analyzing the acceleration signal, recognizing an intensity of an exercise by analyzing the respiration signal, and calculating a consumed calories in consideration of the amount of the user's movements, the movement patterns, and the intensity of exercise. A user's consumed calories can be precisely calculated by recognizing the user's movement and even the user's respiration rate. | 1. A portable device for measuring a consumed calorie, the device comprising: a respiration measurement device measuring an amount of air inhaled into a user's nose to acquire and output a respiration signal, and including a signal generation unit having a cylindrical body;a plurality of movement detection devices acquiring and outputting acceleration signals at least one of reflecting or indicating a magnitude and direction of a movement of corresponding parts of the user's body; anda controller recognizing an amount of a user's movements and movement patterns by analyzing the acceleration signals, recognizing an intensity of an exercise by analyzing the respiration signal, and calculating consumed calories in consideration of the amount of a user's movements, the movement patterns, and the intensity of the exercise;wherein the cylindrical body of the signal generation unit is coaxial with an air tube insertable at least partly into a user's nose. 2. The device of claim 1, further comprising one or more temperature sensors measuring one or more of a user's mouth temperature, a user's skin temperature, and an external body temperature. 3. The device of claim 2, wherein the controller calculates the consumed calories further in consideration of one or more of the user's mouth temperature, the user's skin temperature, and the external body temperature, in addition to the amount of movements, the movement patterns, and the respiration rate. 4. The device of claim 1, wherein the respiration measurement device comprises: a respiration measurement sensor measuring the amount of air inhaled into the user's nose; anda signal processing unit calculating the respiration rate from the amount of air which has been measured by the respiration measurement sensor, converting the respiration amount into a signal recognizable to the controller, and outputting the converted signal. 5. The device of claim 4, wherein the respiration measurement sensor comprises: a glasses frame having a nose cover formed to cover the user's nose; andwherein the signal generation unit is placed on the nose cover and generates a signal corresponding to the amount of air inhaled into the user's nose or exhaled from the user's nose. 6. The device of claim 5, wherein the cylindrical body comprises: an inhalation passage;a turbine installed in the inhalation passage so as to rotate according to inhalation through the inhalation passage; anda plurality of coils installed at an inner side of the inhalation passage in order to generate an AC signal according to alternate magnetic fields resulting from the rotation of the turbine. 7. The device of claim 6, wherein the cylindrical body further comprises: an exhalation passage accommodating the inhalation passage therein;a suction valve plate installed in the inhalation passage and opened only for inhalation;a discharge valve plate installed in the exhalation passage and opened only for exhalation; anda frame fixing the inhalation passage... | Electronics and Telecommunications Research Institute | A61B005/08
US08781588 | B2 | 2014-07-15 | MRT optocoupler | 2010-12-16 | Device and method for detecting electromagnetic fields occurring in imaging magnetic resonance tomography MRT/MRI tests. Relates to an implantable medical device (IMD) comprising a hermetically sealed housing, control unit(s), detection unit(s) for MRT interference fields connected/connectable to control unit(s) and to electrode(s) and/or to antenna(s) and/or coil(s), wherein the MRT interference detection unit contains at least one electro-optical converter which converts induced voltages from the electrode(s) and/or the antenna(s) and/or the coil(s) to optical signals, which are optically transmitted in a potential-free manner within the detection unit for MRT interference fields to an evaluation unit for the detection unit for MRT interference fields, and when a threshold for the optical signal and/or a predetermined periodic occurrence of the optical signals is exceeded, the evaluation unit triggers switching to an MRI-safe state or transmits a corresponding signal to the control unit(s). | 1. An implantable medical device comprising: a hermetically sealed housing;at least one control unit situated within said hermetically sealed housing;at least one conducting element coupled with said implantable medical device wherein said at least one conducting element comprises any combination of one or more of at least one electrode,at least one antenna,at least one coil;at least one detection unit configured to detect magnetic resonance tomography interference fields which is connected to the at least one control unit and to said at least one conducting element;an evaluation unit coupled with the at least one detection unit;wherein the at least one detection unit contains at least one electro-optical converter configured to convert induced voltages for said magnetic resonance tomography interference fields from said at least one conducting element to optical signals, which are optically transmitted in a potential-free manner within the detection unit to the evaluation unit; and,wherein when one or more of a threshold for the optical signals and a predetermined periodic occurrence of the optical signals is exceeded, the evaluation unit is configured to trigger the implantable medical device to switch to a magnetic resonance imaging-safe state. 2. The implantable medical device according to claim 1, wherein said at least one electro-optical converter comprises multiple electro-optical converters, wherein the at least one detection unit for magnetic resonance tomography interference fields has at least one sensor configured to convert the optical signals from said at least one electro-optical converter to electrical signals which are used for the magnetic resonance tomography detection, and wherein the optical signals from said multiple electro-optical converters are configurably sent to a second sensor. 3. The implantable medical device according to claim 1, wherein the at least one electrode is configured for therapeutic and/or or diagnostic purposes, wherein the electrode extends outside the hermetically sealed housing, or is situated on an exterior of the hermetically sealed housing or is part of the hermetically sealed housing. 4. The implantable medical device according to claim 3, wherein the at least one electrode is a connection to one or more of at least one sensor and at least one actuator distally situated with respect to the implantable medical device. 5. The implantable medical device according to claim 1, wherein the electro-optical converter is an optocoupler. 6. The implantable medical device according to claim 1, wherein the electro-optical converter is a sensor for electromagnetic fields which is based on optical properties of a crystal which depend on external electromagnetic fields. 7. The implantable medical device according to claim 1, wherein the electro-optical converter transfers only information concerning one or more of: amplitudes;amplitude envelopes;frequencies and;phases, in configurations having multiple el... | BIOTRONIK CRM Patent AG | A61B005/05 A61N001/37
US09042972 | B2 | 2015-05-26 | Pain judging device to judge pain based on a frequency component of a peak-relevant value | 2009-06-23 | A peak-relevant value device acquires a peak-relevant value (for example, the peak value of an R wave (R peak value)) every cycle from an electrocardiogram acquired. The frequencies of the peak-relevant value acquired as time-series data and the magnitudes for the respective frequencies are analyzed. A peak-relevant value LF calculating device calculates an LF component (peak-relevant value LF component) from the frequency component of the peak-relevant value. An interval acquiring device acquires the interval between characteristic points of the electrocardiographic complex from the electrocardiogram acquired and the frequencies of the feature point interval acquired as time-series data to acquire the magnitudes of the respective frequency component are analyzed. An interval HF calculating device calculates the HF component from the frequency components of the feature point interval acquired and pain is judged on the basis of the variations of the peak-relevant value LF components and the interval HF components. | 1. A pain judging device comprising: an electrocardiographic information acquiring unit for acquiring electrocardiographic information measured in a measuring unit;pain judging means for judging a degree of pain based on analysis of a frequency of a peak value of a P wave, a Q wave, an R wave, an S wave, a T wave, or an ST of the electrocardiographic information; andoutputting means for outputting a judgment result by the pain judging means for treatment of a patient,wherein the pain judging means comprises:peak value frequency analyzing means for analyzing a frequency of the peak value acquired as time-series data;peak value Low Frequency (LF) component calculating means for calculating as a peak value LF component an LF component based on a frequency component of a peak value acquired by the peak value frequency analyzing means, andwherein the peak value LF component is acquired as characteristics of fluctuation of a peak value. 2. The pain judging device according to claim 1, wherein the pain judging means comprises noise eliminating means for improving accuracy of pain judgment based on the fluctuation related to the peak value. 3. The pain judging device according to claim 2, wherein the noise eliminating means for eliminating influence of the noise using, as an index, fluctuation of an interval between characteristic points of the electrocardiographic complex, myoelectric information, or a skin resistance or a skin potential measured by a second measuring unit. 4. The pain judging device according to claim 1, wherein the pain judging means judges that there is pain when the peak value Low Frequency (LF) component increases. 5. The pain judging device according to claim 1, wherein the pain judging means comprises: interval frequency analyzing means for analyzing a frequency of an interval between waveform characteristic points of the electrocardiographic information acquired as time-series data;interval High Frequency (HF) component calculating means for calculating as an interval HF component an HF component based on a frequency component of the interval between the waveform characteristic points that are acquired by the interval frequency analyzing means, andwherein the interval HF component is acquired as characteristics of fluctuation of an interval. 6. The pain judging device according to claim 5, wherein the interval HF component is an index indicating a respiratory fluctuation component acquired from a frequency component of an interval between waveform characteristic points that are acquired by the interval frequency analyzing means. 7. The pain judging device according to claim 5, further comprising: peak value LF component calculating means for calculating as a peak value LF component an LF component based on a frequency component of a peak value acquired by the peak value frequency analyzing means, andwherein the pain judging means judges that there is pain when an interval HF component does not decrease and a peak value LF ... | NIHON KOHDEN CORPORATION | A61B005/0456 A61B005/04 A61B005/0452 A61B005/00
US09014780 | B2 | 2015-04-21 | Image-based biopsy guidance method | 2010-11-09 | When performing image-guided biopsy of an anatomical structure in a patient, a target anatomical patient region containing biopsy target is imaged using both SPECT and XCT concurrently. 3D SPECT and XCT image data is fused to generate a fused 3D reference image that is overlaid on 2D patient image(s) generated during the biopsy procedure to generate an overlay image. The overlay image also includes a planned path or trajectory for a biopsy instrument. The 2D patient images are generated using SPECT and/or XCT, and are updated periodically to show biopsy instrument position and progress. | 1. A system that facilitates image-guided biopsy, including: a multi-modal imaging device comprising at least two imaging modalities, which generates images of an anatomical region of a patient including a target region;a processor that executes computer-executable instructions stored in a memory, the instructions comprising: mapping a biopsy target in the target region by generating images of the biopsy target using at least one of the imaging modalities;monitoring a position of a biopsy instrument in the patient during a biopsy procedure by imaging the target region using one of the at least two imaging modalities provided by the multi-modal imaging device and generating at least two real-time planar 2D images that are orthogonal to each other;generating 2D overlay images by projecting a planned biopsy instrument trajectory and biopsy instrument position into the at least two orthogonal planar real-time 2D images of the target region;generating, from the at least two 2D overlay images, a fused 3D reference image of the target region and the planned biopsy instrument trajectory;monitoring and compensating for patient movement by updating patient position in the 2D overlay images and the fused 3D reference image during the biopsy procedure, wherein updating patient position is performed by rotating an x-ray source and a flat panel x-ray detector on a gantry and generating image data that is compared with the fused 3D reference image to determine patient movement;wherein monitoring the position of the biopsy instrument further comprises periodically generating and displaying updated 2D overlay images of the target region that show biopsy instrument position and progress relative to the target region and the biopsy target; andwherein the x-ray source and x-ray detector are positioned in a 90° orientation relative to the biopsy path thereby mitigating foreshortening effects in x-ray image data. 2. The system according to claim 1, further comprising: mapping the biopsy target using at least two of the imaging modalities; andgenerating the fused 3D reference image of the target region and planned trajectory by fusing 3D image data acquired, while mapping the biopsy target, in a common imaging plane by the at least two imaging modalities;wherein the fused 3D reference image includes 3D nuclear image data combined with 3D x-ray computed tomography (XCT) image data. 3. The system according to claim 1, wherein the multi-modal imaging device includes at least two nuclear cameras, an x-ray source, and a flat-panel x-ray detector. 4. The system according to claim 3, wherein the nuclear cameras include single photon emission computed tomography (SPECT) cameras. 5. The system according to claim 3, wherein the instructions further include: prior to the biopsy procedure, retracting the nuclear cameras into a stowed position;monitoring the position of the biopsy instrument in the patient during the biopsy procedure by imaging the target region using the x-ray... | Koninklijke Philips N.V. | A61B005/05 A61B006/03 A61B006/00 A61B019/00 A61B010/02 G01R033/48
US09706933 | B2 | 2017-07-18 | Blood pressure measurement device | 2009-05-21 | In a sphygmomanometer, a value representing a perimeter of a measuring portion is obtained. A parameter for controlling a drive voltage of a pump is determined based on the value. The drive voltage is determined based on the parameter and an internal pressure of a fluid bladder to pressurize the fluid bladder. In pressurizing process, a diastolic blood pressure value is calculated. A systolic blood pressure value is estimated, and when the internal pressure of the fluid bladder reaches the pressure, the pressurizing is stopped. A gap of a valve for discharging the fluid from the fluid bladder is determined based on a perimeter of a measuring portion, and the fluid bladder is depressurized with the gap constant. In the depressurizing process, the systolic blood pressure value is calculated. When the systolic blood pressure value is calculated, the fluid is discharged from the fluid bladder, and the measurement is terminated. | 1. A blood pressure measurement device comprising: a fluid bladder;a pressurizing unit for pressurizing the fluid bladder by injecting a fluid into the fluid bladder;a depressurizing unit for depressurizing the fluid bladder by discharging the fluid from the fluid bladder;a sensor for measuring an internal pressure change of the fluid bladder;a blood pressure measurement unit for calculating a blood pressure value based on the internal pressure change of the fluid bladder obtained by the sensor in the depressurizing process in which the depressurizing unit discharges the fluid from the fluid bladder; anda control unit for controlling the pressurizing unit, the depressurizing unit, and the blood pressure measurement unit,wherein the depressurizing unit includes only one valve arranged on the fluid bladder, so that a gap of an opening on the fluid bladder is controlled by the one valve,wherein the gap of the one valve is controlled by a drive voltage,wherein the control unit includes an acquisition unit for obtaining information about a perimeter of a measuring portion,wherein the control unit determines the drive voltage of the one valve according to the perimeter so that the gap is set to a constant determined gap with respect to the obtained perimeter,wherein the control unit controls the gap of the one valve to be maintained constant at the constant determined gap during the depressurizing process, resulting in rapid deflation at a beginning of the depressurizing process and slow deflation at an end of the depressurizing process, andwherein the constant determined gap changes relatively to the perimeter such that the constant determined gap becomes great as the perimeter becomes large. 2. The blood pressure measurement device according to claim 1, wherein the control unit determines the gap of the one valve such that the internal pressure of the fluid bladder attains a depressurizing rate in which a predetermined number of more pulses are included in a time varying from a systolic blood pressure to a diastolic blood pressure. 3. The blood pressure measurement device according to claim 1, wherein the blood pressure measurement unit further calculates a blood pressure value based on the internal pressure change of the fluid bladder obtained by the sensor in the pressurizing process in which the pressurizing unit injects the fluid into the fluid bladder, andwherein the control unit determines the gap of the one valve according to the blood pressure value calculated based on the internal pressure change of the fluid bladder in the pressurizing process. 4. The blood pressure measurement device according to claim 1, wherein the blood pressure measurement unit further calculates a cycle of a pulse wave based on the internal pressure change of the fluid bladder obtained by the sensor in the pressurizing process in which the pressurizing unit injects the fluid into the fluid bladder, andwherein the control unit determines the gap of the one valve acc... | OMRON HEALTHCARE Co., Ltd. | A61B005/02 A61B005/0225 A61B005/022
US10398345 | B2 | 2019-09-03 | Method and system to reposition an imager based on the orientation of a medical intervention device | 2011-01-31 | System for registering a first image with a second image, the system including a first medical positioning system for detecting a first position and orientation of the body of a patient, a second medical positioning system for detecting a second position and orientation of the body, and a registering module coupled with a second imager and with the second medical positioning system, the first medical positioning system being associated with and coupled with a first imager, the first imager acquiring the first image from the body, the first imager producing the first image by associating the first image with the first position and orientation, the second medical positioning system being associated with and coupled with the second imager, the second imager acquiring the second image and associating the second image with the second position and orientation, the registering module registering the first image with the second image, according to the first position and orientation and the second position and orientation. | 1. A system for adjusting an imager, by means of a moving mechanism, to a desired orientation with respect to a section of the body of a patient to acquire a visual representation of said section, the visual representation including an optimal representation of a portion of interest of a medical intervention device when inserted into said section, the system comprising: a medical positioning system associated with a coordinate system, said medical positioning system including a transmitter configured to transmit an electromagnetic field;a processor coupled with said medical positioning system and with said moving mechanism; anda device position and orientation detector coupled at a distal end of said medical intervention device at said portion of interest that is configured for insertion into said section of the body of the patient, said device detector comprising a first electromagnetic field sensor configured to detect said electromagnetic field and produce a first signal respective of the device position and orientation, said device detector being further coupled with said medical positioning system,an imager position and orientation detector coupled with said imager and with said medical positioning system, said imager detector comprising a second electromagnetic field sensor configured to detect said electromagnetic field and produce a second signal respective of the imager position and orientation,wherein said medical positioning system is configured to detect a device position and orientation of said device position and orientation detector in said coordinate system based on said first signal, said medical positioning system is configured to provide said device position and orientation to said processor, wherein said medical positioning system is configured to detect an imager position and orientation of said imager position and orientation detector in said coordinate system based on said second signal, said medical positioning system is configured to provide said imager position and orientation to said processor,wherein said processor is configured to determine a vector indicative of an orientation of said medical intervention device in said coordinate system according to said device position and orientation;wherein said processor is configured to determine said desired orientation of said imager according to said device position and orientation and said vector in said coordinate system, andwherein said processor according to the detected position and orientation of said imager position and orientation detector is configured to direct said moving mechanism to move said imager to said desired orientation in which an imaging axis associated with said imager is in a desired relationship with said determined vector wherein said desired relationship comprises said imaging axis being normal to said vector and wherein when alignment of said imaging axis of said imager normal to said vector is not possible, said imager is directed to an orientati... | St. Jude Medical Internation Holding S.à.r.l. | A61B005/055 A61B005/06 A61B005/00 A61N005/10 A61B034/20 G06T007/33 A61B034/00 A61B090/00
US09545222 | B2 | 2017-01-17 | Garment with noninvasive method and system for monitoring physiological characteristics and athletic performance | 2010-08-26 | A physiological monitoring system for noninvasively monitoring physiological parameters of a subject, comprising a fitness monitoring system for monitoring physiological parameters of a subject engaged in a physical activity, in accordance with one embodiment of the invention, includes (i) a monitoring garment adapted to cover at least a portion of a subject's torso, (ii) a magnetometer subsystem including a first magnetometer and a second magnetometer, wherein the first and second magnetometers are responsive to changes in distance therebetween, wherein the magnetometer subsystem is configured to generate and transmit a signal representative of a change in the distance between the first and second magnetometers, wherein the first and second magnetometers are incorporated into the monitoring garment, and wherein the first and second magnetometers are proximate to the subject's chest region when the monitoring garment is worn by the subject. | 1. A fitness monitoring system for monitoring parameters of a subject engaged in a physical activity, comprising: a monitoring garment configured to cover at least a portion of the subject's torso;a sensor subsystem including a first sensor, a second sensor, a third sensor, and a fourth sensor, wherein:the first, third, and fourth sensors are on the subject's front chest region when the garment is worn by the subject,the second sensor is on the subject's back chest region when the garment is worn by the subject,the first and second sensors are responsive to changes in distance therebetween,the third and second sensors are responsive to changes in distance therebetween,the fourth and second sensors are responsive to changes in distance therebetween,the sensor subsystem is configured to generate and transmit a distance signal representative of at least the distance between the first and second sensors without initial calibration, andthe first and second sensors are incorporated into the garment; anda processor subsystem in communication with the sensor subsystem and configured to receive the distance signal, wherein the processor subsystem is configured to control the sensor subsystem and to process the distance signal, wherein the processor subsystem includes a programmed algorithm that includes an empirical relationship for determining a respiratory parameter from the distance signal, is adapted to store benchmark respiratory parameter signals, is configured to compare the distance signal with the benchmark respiratory parameter signals, is configured to generate and transmit a respiratory parameter signal representative of the respiratory parameter as a function of the comparison of the distance signal with the benchmark respiratory parameter signals, the processor subsystem includes stored adverse physiological parameters and the processor subsystem is further configured to compare a detected physiological parameter to the stored adverse physiological parameters, and to generate and transmit an adverse physiological parameter signal representative of the adverse physiological parameter associated with the warning signal if the detected physiological parameter corresponds to one of the stored adverse physiological parameters; anda data monitoring subsystem programmed and configured to receive the respiratory parameter signal, the data monitoring subsystem being programmed to recognize and display the respiratory parameter associated with the respiratory parameter signal, wherein the processor subsystem is further programmed and configured to generate a three-dimensional model of the subject's chest wall from the magnetometer signal, and wherein the data monitoring subsystem is further programmed and configured to display the generated three-dimensional model of the subject's chest wall. 2. The fitness monitoring system of claim 1, wherein the data monitoring subsystem is configured to receive the respiratory parameter signal, wherein the data m... | adidas AG | A61B005/0205 A61B005/113 A61B005/00 G06F019/00 A61B005/024 A61B005/0402 A61B005/0476 A61B005/11 A61B005/1455 A61B007/00
US07970628 | B2 | 2011-06-28 | Method and system for providing dynamic orthodontic assessment and treatment profiles | 2006-10-13 | Method and system for receiving one or more dentition conditions of a patient, associating a first set of one or more treatment goals based on the one or more dentition conditions, generating a case difficulty assessment based on the one or more treatment goals, detecting a modification to the case difficulty assessment, retrieving a second set of one or more treatment goals associated with the modified case difficulty assessment, and displaying the second set of one or more treatment goals are provided. | 1. A computer-implemented method, comprising: receiving a first set of one or more dentition conditions of a patient;retrieving a second set of one or more dentition conditions from a database wherein each of the second set of conditions has a predetermined level of difficulty based on at least one of expert opinion, computational algorithm, and historical case content and is associated with a corresponding treatment goal;comparing the first set of conditions with the second set of conditions;estimating a case difficulty assessment level for each of the first set of conditions based on the comparison and associating a respective minimum recommended certification or experience level of a treatment provider with each assessment level;generating with the computer a first treatment goal of the patient based on the first set of conditions;generating with the computer, at a particular stage in a treatment of the patient, a first treatment plan based on the first set of conditions and the first treatment goal of the patient;calculating and displaying a case difficulty assessment level and associated minimum recommended certification or experience level for the first treatment plan;receiving a request for at least one of a modified case difficulty assessment level and an associated minimum recommended certification or experience level; andgenerating with the computer, at the particular stage in treatment, a second treatment plan based on at least one of the first set of conditions and the received request, the second treatment plan having the modified case difficulty assessment level, and the associated minimum recommended certification or experience. 2. The computer-implemented method of claim 1 further comprising retrieving a second treatment goal of the patient based on the modified case difficulty assessment level. 3. The computer-implemented method of claim 1 wherein the first treatment goal of the patient includes a first set of one or more treatment goals of the patient, and the second treatment goal of the patient includes a second set of one or more treatment goals of the patient. 4. The computer-implemented method of claim 3 wherein the first treatment goal set and the second treatment goal set include at least one common treatment goal. 5. The computer-implemented method of claim 4 wherein the at least one common treatment goal is associated with a lowest case difficulty assessment level. 6. The computer-implemented method of claim 1, wherein receiving the request for the modified case difficulty assessment level includes detecting one of a higher or lower case difficulty assessment level associated with the first set of one or more dentition conditions of the patient. 7. The computer-implemented method of claim 1 further comprising selecting a first dentition condition from the second set of the one or more dentition conditions stored in the database that is similar to a second dentition condition in the first set of the one or more of ... | Align Technology, Inc. | A61C003/00 G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US20110144457 | A1 | 2011-06-16 | INSTRUMENTED, COMMUNICATING PORTABLE CLOTHING AND ACCESSORY | 2009-08-04 | A garment or accessory includes: at least one sensor for measuring at least one physiological parameter of the wearer of the garment or accessory; at least one element for determining the stress as a function of at least one measurement of the physiological parameter; and a remote transmission link, by wireless element, of data representative of measurements from each sensor. In some embodiments, a sensor is adapted to measure heart rate, the blood's oxygen saturation rate and the body temperature of the wearer, preferably by pinching the lobe of an ear of the wearer. For example, this sensor is an optical sensor that uses two separate wavelengths. In some embodiments, the element for determination of the composite stress determines the composite stress as function of the variability of heart rate, the body temperature, the heart rate of the user, and the voice and/or orientation. | 1-27. (canceled) 28. Garment or accessory module, that comprises: at least one sensor for measuring at least one physiological data of the wearer,at least one means for determining the stress as a function of at least one measurement of a said physiological data anda remote transmission link, by wireless, of data representative of measurements from each said sensor. 29. The garment or accessory module according to claim 28, wherein at least one sensor is adapted to measure the wearer's heart rate. 30. The garment or accessory module according to claim 28, wherein at least one sensor is adapted to measure the blood's oxygen saturation of the wearer 31. The garment or accessory module according to claim 28, wherein at least one sensor is adapted to measure the body temperature of the wearer. 32. The garment or accessory module according to claim 28, wherein at least one sensor clamps the lobe of an ear of the wearer. 33. The garment or accessory module according to claim 28, wherein at least one sensor is an optical sensor that uses at least two separate wavelengths. 34. The garment or accessory module according to claim 28, wherein at least one sensor is adapted to measure the respiration rate of the wearer. 35. The garment portable accessory according to claim 28, wherein at least one sensor, adapted to measure the respiration rate of the wearer determines the respiration rate from oximetry data. 36. The garment or accessory module according to claim 28, wherein at least one sensor is adapted to measure heart rate variability. 37. The garment or accessory module according to claim 36, that comprises means for determining a composite stress, according to said variability and data representative of the body temperature as well as the heart rate of the wearer. 38. The garment or accessory module according to claim 28, that comprises means of determining a composite stress, according to a voice signal of the wearer. 39. The garment or accessory module according to claim 28, that comprises means of determining a composite stress, depending on the evolution of the position and/or orientation of the wearer. 40. The garment or accessory cell according to claim 28, that further comprises at least one sensor for measuring an environment parameter value. 41. The garment or accessory module according to claim 40, wherein at least one sensor for measuring an environmental parameter value includes an accelerometer and/or is adapted to measure movements of the wearer. 42. The garment or accessory module according to claim 40, wherein at least one sensor for measuring an environmental parameter value is adapted to determine the geographical position and/or orientation of the wearer. 43. The garment or portable accessory according to claim 28, that comprises a wireless point-to-point communication means with similar garments or portable accessories. 44. The garment or accessory cell according to claim 43, that comprises means for detecting danger... | BODYSENS S.A.S | A61B005/00
US20110130655 | A1 | 2011-06-02 | IMPLANTABLE DEVICES AND METHODS FOR EXTERNAL BEAM RADIATION TREATMENTS | 2010-11-30 | Implantable devices and methods for external beam radiation treatments are disclosed. One embodiment of an implantable device for guided radiation therapy comprises an active marker configured to be positioned within the patient and to transmit a non-ionizing wireless signal in response to a non-ionizing wirelessly transmitted source energy. The device further includes a fastening unit coupled to the active marker and configured to (a) hold the marker within a desired distance of a target in the tissue and (b) inhibit deformation of tissue from moving the active marker relative to the target. | 1. An implantable device for guided radiation therapy, comprising: an active marker having a circuit configured to be implanted in a patient and to transmit a wirelessly transmitted location signal from within the patient; anda fastening unit having a first portion coupled to the active marker and a second portion configured to be at least proximate to the dermis of the patient. 2. The device of claim 1 further comprising a plurality of active markers. 3. The device of claim 1 wherein the active marker comprises a wired marker having a casing and a core within the casing, and wherein the circuit comprises a coil within the casing around the core. 4. The device of claim 1 wherein the active marker comprises a leadless marker having a casing and a core within the casing, and wherein the circuit comprises a coil within the casing around the core. 5. The device of claim 1 wherein the fastening unit comprises a stability element configured to fix the marker to the tissue and a decoupling element configured to inhibit deformation of tissue from moving the marker relative to the target. 6. The device of claim 1 wherein the fastening unit comprises a stability element configured to fix the marker to the tissue without a decoupling element configured to inhibit deformation of tissue from moving the marker relative to the target. 7. The device of claim 5 wherein the stability element comprises a first coil having a first loop diameter and the decoupling element comprises a second coil having a second loop diameter less than the first loop diameter. 8. The device of claim 7 wherein the first coil has a first flexibility and the second coil has a second flexibility greater than the first flexibility. 9. The device of claim 7 wherein the first coil is made from a first material and the second coil is made from a second material and construction different than the first material. 10. The device of claim 7 wherein the first coil is made from a first material and the second coil is made from the same material wherein the material of the second coil is processed differently than the material of the first coil. 11. The device of claim 7 wherein the first coil has a first filament diameter and the second coil has a second filament diameter less than the first filament diameter. 12. The device of claim 7 wherein the first coil has not more than two loops and the second coil has at least two loops. 13. The device of claim 12 wherein the first coil has a single-loop. 14. The device of claim 7 wherein the first coil has not more than 10 loops. 15. The device of claim 5 wherein the decoupling element comprises a coil. 16. The device of claim 5 wherein the decoupling element comprises a flexible casing and a serpentine section of filament. 17. The device of claim 5 wherein the decoupling element comprises a loop of filament having a zigzag configuration. 18. The device of claim 5 wherein the decoupling element comprises an accordion-like member. 19. The... | - | A61B005/05
US20110137367 | A1 | 2011-06-09 | EVALUATING A PATIENT CONDITION USING AUTONOMIC BALANCE INFORMATION IN IMPLANTABLE CARDIAC DEVICES | 2011-02-15 | Systems and methods for evaluating a patient condition using autonomic balance information involve providing an implantable cardiac device that acquires a cardiac waveform from a patient. One or more characteristics associated with autonomic balance of the patient are detected and used to evaluate a patient condition, such as sleep onset, sleep stage, cardiac vulnerability over a predetermined duration, and sleep disordered breathing. Patient activity levels may be sensed and used to evaluate the patient's condition, such as for determining a level of systemic stress. Characteristics associated with the autonomic balance include calculating an LF/HF ratio waveform and/or determining one or more morphological features of the LF/HF ratio waveform. Coordination with a patient-external device may facilitate transmission of information about one or more of the cardiac waveform, the one or more characteristics associated with the autonomic balance, and a marked cardiac waveform. | 1. A method, comprising: acquiring a cardiac waveform from a patient using an implantable cardiac device;detecting one or more characteristics associated with autonomic balance of the patient; andgenerating a marked waveform including the acquired cardiac waveform and one or more indicators representative of the one or more characteristics. 2. The method of claim 1, comprising displaying the marked waveform on a screen. 3. The method of claim 1, comprising printing the marked waveform. 4. The method of claim 1, wherein generating the marked waveform comprises generating symbols indicative of discrete portions of the cardiac waveform corresponding to the occurrence of the one or more characteristics. 5. The method of claim 1, wherein generating the marked waveform comprises generating symbols indicative of general conditions pertaining to a continuous portion of the cardiac waveform. 6. The method of claim 1, wherein detecting the one or more characteristics comprises determining an LF/HF ratio waveform from the cardiac waveform, the LF/HF ratio waveform being a measure of the hemodynamic balance determined by dividing a power of low frequency (LF) signals of the cardiac waveform by a power of high-frequency (HF) signals of the cardiac waveform. 7. The method of claim 6, wherein detecting the one or more characteristics comprises determining one or more morphological features of the LF/HF ratio waveform. 8. The method of claim 1, comprising: acquiring a respiration waveform;time-aligning the marked waveform with the respiration waveform; anddisplaying the marked waveform and the respiration waveform. 9. The method of claim 8, comprising deriving respiration characteristics from the respiration waveform. 10. The method of claim 9, wherein deriving the respiration characteristics comprises deriving parameters of the respiration waveform morphology. 11. The method of claim 9, wherein deriving the respiration characteristics comprises deriving parameters characterizing respiration abnormalities. 12. The method of claim 9, wherein detecting the one or more characteristics comprises determining an LF/HF ratio waveform from the cardiac waveform, the LF/HF ratio waveform being a measure of the hemodynamic balance determined by dividing a power of low frequency (LF) signals of the cardiac waveform by a power of high-frequency (HF) signals of the cardiac waveform, and comprising detecting symptoms or physiological conditions from the LF/HF ratio waveform and the respiration waveform. 13. The method of claim 9, wherein generating the marked waveform including generating symbols indicating respiration abnormalities. 14. The method of claim 9, wherein generating the marked waveform including generating symbols indicating conditions affecting respiration. 15. A system configured to communicate with a patient-internal medical device implanted in a patient, the system comprising: a patient-external medical device configured to be communicatively cou... | - | A61N001/365 A61B005/0205
US20110160608 | A1 | 2011-06-30 | Brain-related chronic pain disorder diagnosis and assessment method | 2008-12-18 | A method for diagnosing and assessing a brain-related chronic pain disorder. The method includes assessing a subject's brain function, determining the probability that a subject is suffering from chronic pain as a result of an abnormal brain function condition by obtaining a quantitative assessment of the subject's brain function, and making a statistical comparison between the subject's quantitative brain function assessment and either a database of quantitative assessments of the brain functions of normal, healthy individuals, or a database of quantitative assessments of the brain functions of individuals known to have been suffering from chronic pain as a result of the abnormal brain function condition. | 1. A method for diagnosing and assessing a brain-related chronic pain disorder, the method including the steps of: assessing a subject's brain function;determining the probability that a subject is suffering from chronic pain as a result of an abnormal brain function condition by obtaining a quantitative assessment of the subject's brain function; andmaking a statistical comparison between the subject's quantitative brain function assessment and either: a database of quantitative assessments of the brain functions of normal, healthy individuals, ora database of quantitative assessments of the brain functions of individuals known to have been suffering from chronic pain as a result of the abnormal brain function condition. 2. The method of claim 1 in which: the step of assessing a subject's brain function includes obtaining an electroencephalogram (EEG) of the subject's electrical brain activity; andthe step of determining the probability includes determining the probability that the subject is suffering from the abnormal brain function condition by: obtaining a quantitative assessment of the subject's EEG (qEEG), andmaking a statistical comparison between the subject's qEEG and either: a database of qEEGs of normal, healthy individuals, ora database of qEEGs of individuals known to have been suffering from the abnormal brain function condition. 3. The method of claim 2 in which: the step of assessing a subject's brain function includes obtaining an electroencephalogram (EEG) of the subject's electrical brain activity;the step of determining the probability that a subject is suffering from chronic pain as a result of abnormal brain function includes: obtaining a quantitative assessment of the subject's EEG (qEEG), andmaking a statistical comparison between the subject's qEEG and a database of qEEGs of individuals known to have been suffering from chronic pain as a result of abnormal brain function; andthe step of determining the probability that the subject is suffering from chronic pain as a result of the abnormal brain function condition includes making a statistical comparison between the subject's qEEG and a database of qEEGs of the brain functions of individuals known to have been suffering from the abnormal brain function condition. 4. (canceled) 5. (canceled) 6. The method of claim 2 in which the step of assessing a subject's brain function includes obtaining a “TPP EEG” record of EEG data obtained over a period of time during which at least one tender point on the subject's body is palpated. 7. The method of claim 6 in which an algometer is used to palpate at least one tender point on the subject's body while the “TPP EEG” record is being obtained. 8. The method of claim 6 in which: the step of making a statistical comparison includes making such a comparison between the subject's qEEG and a database of qEEGs of individuals known to have been suffering from fibromyalgia; andthe step of palpating at least one tender point includes se... | CEREPHEX CORPORATION | A61B005/0476
US20110144520 | A1 | 2011-06-16 | METHOD AND DEVICE FOR POINT-OF-CARE NEURO-ASSESSMENT AND TREATMENT GUIDANCE | 2009-12-16 | A method and apparatus for providing an objective assessment of the neurological state of a patient using a field-portable neuro-assessment device is described. The method includes placing an electrode set on the patient's head, acquiring spontaneous brain electrical signals and evoked potential signals from the patient through the electrode set, processing the signals using a handheld base unit, and displaying a result indicating the probability of the patient's neurological signal being normal or abnormal. The neuro-assessment device allows for a rapid, on-site neurological evaluation by an emergency medical technician, triage nurse, or any other medical personnel to identify patients with neurological disorders who may require immediate medical attention. | 1. A method of performing neurological triage on a patient, comprising the steps of: providing a patient sensor comprising at most eight disposable neurological electrodes and at least one ear phone;acquiring spontaneous brain electrical signals using the neurological electrodes;providing a handheld base unit comprising a signal processor, the base unit being operatively coupled to the patient sensor for processing the acquired spontaneous brain electrical signals;calculating an index indicating a statistical probability of the patient's brain electrical signals being normal or abnormal using discriminant classification analysis; anddisplaying the index on the handheld base unit. 2. The method of claim 1, further comprising the step of displaying tables of probabilistic values of a large set of quantitative features derived from the acquired spontaneous brain electrical signals. 3. The method of claim 1, further comprising the step of providing auditory stimuli to the patient using the at least one ear phone. 4. The method of claim 3, further comprising the steps of acquiring auditory evoked potential signals using the neurological electrodes; and processing the acquired auditory evoked potential signals using the handheld base unit. 5. The method of claim 4, wherein the acquired auditory evoked potential signals comprised auditory brainstem response (ABR) signals. 6. The method of claim 4, further comprising the step of displaying the auditory evoked potential signal waveforms and a measure of the signal-to-noise ratio associated with the evoked potential signals. 7. The method of claim 1, further comprising the step of displaying the acquired spontaneous electrical signals in real-time. 8. The method of claim 1, wherein the step of calculating an index further comprises the steps of: performing analog-to-digital conversion of the signals;automatically identifying and removing artifacts from the signals; andextracting quantitative features and computing Z-transform scores. 9. The method of claim 8, wherein the computation of Z-transform scores and the discriminant classification is performed using a stored population database comprising brain electrical activity data from a plurality of individuals. 10. The method of claim 9, wherein the population database comprises neurological reference data from a plurality of individuals in the presence or absence of different types of acute neurological conditions. 11. The method of claim 8, wherein the quantitative features are extracted using Fast Fourier Transform. 12. The method of claim 8, wherein the quantitative features are extracted using wavelet transform. 13. The method of claim 1, further comprising the step of measuring periodically the impedance of each electrode. 14. An apparatus for performing neurological triage on a patient, comprising: a patient sensor comprising at most eight disposable neurological electrodes and at least one ear phone;a handheld base unit operatively conn... | - | A61B005/0484
US07962191 | B2 | 2011-06-14 | Nerve surveillance cannulae systems | 2007-10-31 | An expandable tip cannula system, comprising: a hollow cannula shaft having a proximal end and a distal end; and an expandable tip mounted at the distal end of the hollow cannula shaft, the expandable tip comprising a plurality of generally-triangular shaped petals held together in a radially-inwardly tapered arrangement between adjacent petals, each petal comprising a nerve sensing electrode disposed therein. | 1. A method of performing spine surgery, comprising: minimally invasively advancing a nerve surveillance instrument and a surgical access instrument towards a spinal target site, wherein the nerve surveillance instrument is dimensioned to be slidably received within the surgical access instrument and the surgical access instrument has an inner lumen dimensioned to pass an implant therethrough for introduction into said spinal target site, the nerve surveillance instrument having at least one nerve stimulation electrode disposed proximate a distal end thereof;energizing at least one nerve stimulation electrode disposed proximate a distal end of the nerve surveillance instrument while advancing the nerve surveillance instrument towards the surgical target site in an amount sufficient to depolarize nerves adjacent said nerve surveillance instrument such that said nerves may be detected and avoided without effecting tissue ablation or cauterization;removing said nerve surveillance instrument from said inner lumen after said surgical access instrument has been advanced to said spinal target site to thereby enable a surgeon to introduce said implant into said spinal target site; andintroducing said implant into said spinal target site. 2. The method of claim 1, wherein said at least one nerve stimulation electrode comprises a plurality of nerve stimulation electrodes and wherein the step of advancing comprises the further sub-step of sensing the presence of said nerves in a radial direction relative to a central axis of said surgical access instrument through the use of the plurality of nerve stimulation electrodes which are radially disposed around said distal end of said nerve surveillance instrument. 3. The method of claim 1, further comprising: slideably engaging an expandable mesh with said surgical access instrument when said mesh is in a first position such that said mesh is received on the exterior of said surgical access instrument, said expandable mesh being expandable from said first position to a second position. 4. The method of claim 3, further comprising: advancing said mesh in a distal direction over said surgical access instrument such that said mesh extends substantially the length of said surgical access instrument; andexpanding said mesh from said first position to said second position. 5. The method of claim 4, further comprising: minimally invasively advancing a second surgical access instrument to said spinal target site over said mesh when said mesh is in said second position, wherein said second surgical access instrument has an inner lumen configured to pass surgical instruments therethrough. 6. The method of claim 1, wherein said step of introducing said implant includes using at least one of a bone decorticator, a camera, an articulating forcep and an intervertebral positioning system. 7. The method of claim 1, wherein said distal end of said surgical access instrument comprises a plurality of expandable elements, each... | NuVasive, Inc. | A61B005/04 A61B017/08 A61M029/00
US20110152697 | A1 | 2011-06-23 | Circulatory Pressure Monitoring Using Infusion Pump Systems | 2010-12-17 | A low cost, transportable system for monitoring the central venous pressure of a patient receiving an infusion is provided. The pressure monitoring system of the present invention employs a pump and a flow meter in order to supply infusion fluids to a patient. Based upon the control factors and changes thereof communicated to the pump by a controller in order to achieve and maintain a desired infusion fluid flow rate, relative changes in patient's venous pressure and/or quantitative pressure data is obtained. | 1. A method for monitoring patient circulatory pressure comprising: generating a first infusion fluid pressure for delivery of an infusion fluid to a patient;generating a second infusion fluid pressure responsive to a patient circulatory pressure;determining patient circulatory pressure data based on a difference in control factors employed to generate the first infusion fluid pressure and the second infusion fluid pressure. 2. The method of claim 1 wherein the step of generating a first infusion fluid pressure for delivery of an infusion fluid to a patient comprises pumping infusion fluid through a piezoelectric driven pump. 3. The method of claim 1 wherein the step of generating a first infusion fluid pressure for delivery of an infusion fluid to a patient comprises determining an infusion fluid flow rate with a flow meter. 4. The method of claim 1 wherein the step of generating a first infusion fluid pressure for delivery of an infusion fluid to a patient comprises generating a first infusion fluid pressure greater than the patient circulatory pressure. 5. The method of claim 1 wherein the step of generating a second infusion fluid pressure responsive to a patient circulatory pressure comprises determining an infusion fluid flow rate with a flow meter. 6. The method of claim 1 wherein the step of generating a second infusion fluid pressure responsive to a patient circulatory pressure comprises generating a second infusion fluid pressure greater than the patient circulatory pressure. 7. The method of claim 1 wherein the step of generating a second infusion fluid pressure responsive to a patient circulatory pressure comprises providing a control factor to an infusion pump. 8. The method of claim 7 wherein the step of providing a control factor to an infusion pump comprises determining the control factor based upon infusion fluid flow data received from a flow meter. 10. The method of claim 1 wherein the step of determining patient circulatory pressure data based on a difference in control factors employed to generate the first infusion fluid pressure and the second infusion fluid pressure comprises analyzing control factors employed to direct operation of an infusion pump. 11. The method of claim 1 further comprising the step of correlating specific control factors to a range of different pressures. 12. The method of claim 1 further comprising the step of designating control factors for an infusion pressure equal to zero. 13. A patient circulatory pressure monitoring system comprising: an infusion pump;a flow meter in fluid communication with the infusion pump; anda controller in electrical communication with the infusion pump and the flow meter, the controller further comprising a data correlation of infusion pump control factors and circulatory pressures. 14. The system of claim 13 wherein the infusion pump is a piezoelectric driven infusion pump. 15. The system of claim 13 wherein a control factor comprises a voltage. 16. The sy... | K&Y Corporation | A61B005/021
US07970627 | B2 | 2011-06-28 | Method and system for providing dynamic orthodontic assessment and treatment profiles | 2006-10-13 | Method and system for providing dynamically generated orthodontic profile and associated treatment information including receiving an orthodontic condition and one or more related treatment goal options, retrieving a predetermined set of parameters associated with the orthodontic condition and the one or more related treatment goal options, determining a weighted parameter associated with each of the one or more treatment goal options for the orthodontic condition, and generating treatment plan information for the orthodontic condition based on the determined weighted parameter are provided. | 1. A computer-implemented method, comprising: receiving an orthodontic condition of a patient;providing one or more treatment goal options from a database query of previous treatment goals for similar orthodontic conditions;receiving a selection of one of the options;providing a predetermined set of one or more orthodontic treatment goal parameters;receiving an importance rating from a user for at least one of the treatment goal parameters based on a predetermined scale, wherein the at least one importance rating represents an importance of a particular treatment goal parameter to the selected option;parsing the orthodontic condition into one or more predetermined diagnostic parameters;weighting each of the diagnostic parameters according to a corresponding weighted function based on similarity of the diagnostic parameters to past treatment information; andgenerating a treatment plan based on the weighted diagnostic parameters, the selected treatment goal option, and the at least one importance rated treatment goal parameter. 2. The computer-implemented method of claim 1 wherein the one or more orthodontic treatment goal parameters include at least one of a treatment difficulty assessment, a type of appliance available for treatment, or a treatment duration. 3. The computer-implemented method of claim 1 wherein the importance rating is associated with the orthodontic condition. 4. The computer-implemented method of claim 1 wherein determining the weighted parameter further includes: associating a difficulty rating with each of the one or more treatment goal options for the orthodontic condition; andincreasing the weighting for each one of the diagnostic parameters when the difficulty rating transcends a predetermined value in a first direction, and reducing the weighting when the difficulty rating transcends a predetermined value in a second direction. 5. The computer-implemented method of claim 4 further including correlating each difficulty rating with one or more visual indicators. 6. The computer-implemented method of claim 5 wherein the one or more visual indicators are associated with one or more of a color indicator, a alphanumeric indicator, a symbol indicator, or combinations thereof. 7. The computer-implemented method of claim 5 further including displaying the one or more of the visual indicators correlated with each difficulty rating for the one or more treatment goal options. 8. The computer-implemented method of claim 1 wherein weighting each of the diagnostic parameters includes determining a distance function associated with each of the predetermined set of one or more orthodontic treatment goal parameters for the selected orthodontic condition and the one or more related treatment goal options. 9. The computer-implemented method of claim 8 wherein determining the distance function includes: retrieving historical treatment information associated with the orthodontic condition;comparing the historical treatment information ... | Align Technology, Inc. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00 A61C003/00
US08264969 | B2 | 2012-09-11 | Communications link health monitoring | 2009-10-30 | A reference signal of reference pulses occurring at a uniform reference pulse rate is received. An error event signal of error event pulses occurring at a variable error event pulse rate corresponding to a rate of detected error event occurrences is received. An integrator output signal is produced from the reference signal and the error event signal. The integrator output signal has a value that approaches a first boundary threshold when the reference pulse rate exceeds the error event pulse rate and approaches a second boundary threshold when the error event pulse rate exceeds the reference pulse rate. A decision signal is set to a first value in response to a determination that the value of the received integrator output signal meets a first decision threshold and set to a second value in response to a determination that the value of the received integrator output signal meets a second decision threshold. | 1. Apparatus, comprising: an integrator circuit comprising a first input for receiving a reference signal of reference pulses occurring at a uniform reference pulse rate and a second input for receiving an error event signal of error event pulses occurring at a variable error event pulse rate corresponding to a rate of detected error event occurrences, wherein the integrator circuit is operable to produce from the reference signal and the error event signal an integrator output signal with a value that approaches a first boundary threshold when the reference pulse rate exceeds the error event pulse rate and approaches a second boundary threshold when the error event pulse rate exceeds the reference pulse rate; anda decision circuit that has an input coupled to receive the integrator output signal from the integrator circuit, wherein the decision circuit is operable to set a decision signal with a first value in response to a determination that the value of the received integrator output signal meets a first decision threshold, wherein the decision circuit is operable to set the decision signal to a second value in response to a determination that the value of the received integrator output signal meets a second decision threshold. 2. The apparatus of claim 1, wherein the integrator circuit comprises a counter that tallies a count from a combination of the reference pulses counted in a first count direction and the error event pulses counted in a second count direction. 3. The apparatus of claim 2, wherein the counter saturates at the first boundary threshold in the first count direction and saturates at the second boundary threshold in the second count direction, and the first and second decision thresholds are between the first boundary threshold and the second boundary threshold. 4. The apparatus of claim 3, wherein the counter is incremented in response to each reference pulse and is decremented in response to each error event pulse, the first boundary threshold is higher than the second decision threshold, the first decision threshold is higher than or equal to the second decision threshold, and the second decision threshold is higher than the second boundary threshold. 5. The apparatus of claim 1, wherein the decision circuit comprises: a first comparator operable to compare the value of the received integrator output signal with the first decision threshold;a second comparator operable to compare the value of the received integrator output signal with the second decision threshold; andan output circuit coupled to the first comparator and the second comparator and operable to produce the decision signal. 6. The apparatus of claim 1, further comprising multiple monitor circuits each comprising a respective instance of the integrator circuit and a respective instance of the decision circuit that produces a respective decision signal, and a combination circuit having a respective input for each decision signal and being operable to produce... | Hewlett-Packard Development Company, L.P. | H04L001/00
US07953475 | B2 | 2011-05-31 | Preprocessing for cardiac mapping | 2009-03-16 | A non-contact cardiac mapping method is disclosed that includes: (i) inserting a catheter into a heart cavity having an endocardium surface, the catheter including multiple, spatially distributed electrodes; (ii) measuring signals at the catheter electrodes in response to electrical activity in the heart cavity with the catheter spaced from the endocardium surface; and (iii) determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface. Related systems and computer programs are also disclosed. | 1. A method comprising: inserting a catheter into a heart cavity having an endocardium surface, the catheter comprising multiple, spatially distributed electrodes;prior to determining physiological information, generating a representation of the endocardium surface of a patient's heart cavity;registering the representation of the endocardium surface with a first coordinate system used to determine positions of the catheter electrodes relative to the endocardium surface after the catheter is inserted into the heart cavity;measuring signals at the catheter electrodes in response to electrical activity in the heart cavity; anddetermining the physiological information at multiple locations of the endocardium surface based on the measured signals, the positions of the catheter electrodes relative to the endocardium surface, and the registered representation of the endocardium surface;wherein determining the physiological information comprises determining the physiological information based at least in part on a mathematical operator applied to the measured signals; andwherein generating the representation of the endocardium surface comprises segmenting a representation of the heart cavity. 2. The method of claim 1, further comprising: moving the catheter to each of multiple, different positions in the heart cavity; andfor each of the different catheter positions, determining the positions of the catheter electrodes with respect to the endocardium surface and measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;wherein the determination of the physiological information at the multiple locations of the endocardium surface is based further on the positions of the catheter electrodes, the measured signals at the different catheter positions and the registered representation of the endocardium surface. 3. The method of claim 2, wherein determining the position of the catheter electrodes comprises measuring information about at least one of a position and orientation of the catheter within the heart cavity. 4. The method claim 3, wherein determining the position of the catheter electrodes is based further on information about the distribution of the electrodes on the catheter. 5. The method of claim 4, wherein measuring information about at least one of a position and orientation of the catheter within the heart cavity comprises measuring the position of one or more catheter electrodes within the heart cavity. 6. The method of claim 2, wherein determining the position of the catheter electrodes comprises directly measuring the position of each catheter electrode within the heart cavity. 7. The method of claim 1, wherein the representation is obtained from at least one of: a coherence tomography (CT) image; a magnetic resonance imaging (MRI) image; and an ultrasound image. 8. The method of claim 1, wherein the representation is segmented into a substantially closed surface representation of the endocardiu... | Rhythmia Medical, Inc. | A61B005/02
US07945082 | B2 | 2011-05-17 | Analysis method for data of an examination object, featuring correlation of a volume data record with a number of projection images | 2007-10-05 | Data of an examination object comprises a volume-data record and a plurality of two-dimensional projection images. The volume-data record includes voxels where each voxel is assigned to a location in three-dimensional space. Each projection image includes pixels where each pixel is assigned to a location in a two-dimensional-projection plane and has a value. Each pixel is assigned a projection volume, this being specified in that it is mapped by the radioscopy onto the pixel to which it is assigned. A sub-volume of the volume-data record is selected. The projection images are registered in relation to the volume-data record. A functional parameter of the examination object is specified for the pixels of the projection images, depending on their values. For each pixel, when specifying the functional parameter, consideration is given to the locations and/or the number of those voxels which are positioned both within the sub-volume and within the projection volume. | 1. An analysis method for data of an examination object, selecting a sub-volume of a volume data record, the data comprising at least one volume data record of the examination object and a plurality of two-dimensional projection images of the examination object,the volume data record includes a plurality of voxels and each voxel is assigned to a location in the three-dimensional space;registering a plurality of projection images in relation to the volume data record,each projection image includes a multiplicity of pixels, each pixel is assigned to a location in a two-dimensional projection plane and has a pixel data value, and a projection volume is assigned to each pixel,each of the projection volumes are specified in that they are mapped by the radioscopy onto the pixels to which they are assigned; andspecifying a functional parameter of the examination object for the pixels of the projection images depending on the pixel data values,wherein for each pixel, when specifying the functional parameter, consideration is given to the locations and the number of those which lie both within a sub-volume and within the projection volume that is assigned to the pixel concerned, andwherein the sub-volume is specified on the basis of the selected segmented vessels. 2. The analysis method as claimed in claim 1, wherein the examination object includes a vascular system with vessels and its surroundings, that the vascular system is segmented in the volume data record, that some of the segmented vessels are selected. 3. The analysis method as claimed in claim 2, wherein the projection images are captured following the selection of the segmented vessels, and that a positioning of a recording arrangement to capture the projection images is specified depending on which of the segmented vessels were selected. 4. The analysis method as claimed in claim 1, wherein the examination object includes a vascular system with vessels and its surroundings, that at least locations of segmented vessels are specified in the projection images, and that, based on the pixel data values of pixels which correspond to the segmented vessels, a standard influence factor is determined for all pixels of the projection images, which factor is included in the specification of the at least one functional parameter. 5. The analysis method as claimed in claim 1, wherein the examination object includes a vascular system with vessels and its surroundings, that at least locations of segmented vessels are specified in the projection images, and that the pixels which correspond to the segmented vessels are not taken into consideration when specifying the at least one functional parameter. 6. The analysis method as claimed in claim 5, Wherein, based on the pixel data values of pixels which correspond to the segmented vessels, a standard influence factor is determined for all pixels of the projection images, which factor is included in the specification of the at least one functional parameter.... | Siemens Aktiengesellschaft | G06K009/00 A61B005/05
US08911505 | B2 | 2014-12-16 | Prosthetic socket stabilization apparatus and technique | 2010-11-18 | A portable device, with method, advantageously applies dynamic stimulation of enclosed muscle tissue to stabilize a prosthetic socket on a residual limb. Dynamic stimulation is in response to physical conditions such as prosthesis motion, position and/or internal pressures. Tissue volume contained within the socket may be stabilized by varying average stimulation levels in response to internal socket pressure. | 1. A system to stabilize a prosthesis affixed to an appendage comprising: a shell to contain the appendage and form an interface to the prosthesis;means to measure at least one of (a) a spatial attribute of the shell and (b) a force on the shell;means to measure average pressure within the shell;means to (c) dynamically stimulate at least one area of muscle tissue contained within said shell and (d) stabilize the interface in response to the dynamic stimulation; andmeans to control the dynamic stimulation in response to the average pressure and the at least one of the spatial attribute and the force. 2. The apparatus of claim 1, wherein said means to control the dynamic stimulation comprises at least one of digital circuitry and software executed by a processor. 3. The apparatus of claim 1, wherein said means to measure comprises an accelerometer. 4. The apparatus of claim 1, wherein said means to dynamically stimulate applies electrical stimulation. 5. The apparatus of claim 1, comprising an elastomeric liner to be worn within said shell. 6. The apparatus of claim 1, wherein said means to dynamically stimulate includes switching amplification. 7. A prosthesis comprising: first and second sensor nodes; anda controller, when coupled to the first and second sensor nodes, configured to perform operations comprising: determining a first force, based on the first sensor node, over a first period of time;determining an environmental condition based on the second sensor node, wherein the environmental condition (a) corresponds to a shell coupled to an appendage, and (b) comprises force;determining a first component of first stimulus energy based on the environmental condition and the first force; andstimulating muscle tissue coupled to the prosthetic with the first stimulus energy to stabilize an interface between the appendage and the prosthesis;wherein the first component includes pulse amplitude that is inversely proportional to the first force, the first force includes a composite force based on inputs from the first sensor node and at least one of the second sensor node and a third sensor node, and the first force includes an average force based on the inputs. 8. The prosthesis of claim 7, wherein the first period of time extends at least 1 day in duration and the first force includes an average of forces sensed during the first period of time. 9. The prosthesis of claim 7 including a plurality of stimulation nodes and each of the plurality of stimulation nodes receives composite force as an input. 10. A system comprising: a shell to form an interface between an appendage and a prosthesis;a plurality of sensors;a plurality of stimulation nodes; anda controller configured to couple to the sensors and the shell and (a) determine a spatial attribute and a force, both of which correspond to the shell; (b) determine average pressure corresponding to the shell based on determining the force, (c) stimulate the appendage, via the plurality of stim... | Articulate Labs, Inc. | A61F002/80 A61F002/76 A61B005/103 A61B005/00 A61F002/78 A61N001/04 A61N001/08 A61N001/36 A61B005/11 A61F002/50
US08498697 | B2 | 2013-07-30 | Classification of somatosensory evoked potential waveforms | 2010-10-29 | Embodiments are disclosed relating to classification of somatosensory evoked potential waveforms. | 1. A method, comprising: using a computing platform,classifying at least one evoked potential recording based at least in part on a median frequency of said at least one evoked potential recording and a measured zero-crossing rate of said at least one evoked potential recording; andrejecting said at least one evoked potential recording based at least in part on said median frequency and said zero-crossing rate being approximately within a particular range, wherein a threshold median frequency is approximately within a range between 10 Hz and 150 Hz. 2. The method of claim 1, further comprising recording said at least one evoked potential recording in response to one or more somatosensory stimulus. 3. The method of claim 1, wherein said at least one evoked potential recording is obtained from a human or animal body. 4. The method of claim 1, further comprising averaging at least two acceptable evoked potential recordings to generate a composite evoked potential recording. 5. The method of claim 1, wherein said median frequency is computed as∑k=0fmedian⁢P⁡(fk)=∑k=fmedianfc2⁢P⁡(fk)where, P(fk) comprises a kth sample of a power spectrum of said at least one evoked potential recording, fc comprises a sampling frequency, and fmedian comprises said median frequency. 6. The method of claim 1, further comprising presenting said median frequency on a display device. 7. The method of claim 1, wherein said zero-crossing rate is computed asz⁢⁢c⁢⁢r=1T⁢∑t=0T-1⁢Π⁢{st⁢st-1&lt;0}where s comprises said at least one evoked potential recording of time length T and an indicator function Π{A} is 1 if argument A is true and 0 if otherwise. 8. The method of claim 7, wherein a threshold zero-crossing rate is approximately within a range of 1 to 100 during said time length T. 9. The method of claim 8, wherein said threshold zero-crossing rate is user-selectable. 10. The method of claim 8, wherein at least one evoked potential recording is classified as rejected if the measured zero-crossing rate is substantially equal to or higher than said threshold zero-crossing rate. 11. The method of claim 1, further comprising presenting said measured zero-crossing rate on a display device. 12. The method of claim 1, wherein said evoked potential recording comprises a somatosensory evoked potential (SEP) recording. 13. The method of claim 1, wherein said evoked potential recording comprises at least one of the following: an electroencephalogram, nerve conduction recording, visual evoked potentials, audio evoked potentials, or any combination thereof. 14. The method of claim 1, further comprising receiving one or more binary digital signals representative of said at least one evoked potential recording. 15. A method, comprising: using a computing platform,classifying at least one evoked potential recording based at least in part on a median frequency of said at least one evoked potential recording and a measured zero-crossing rate of said at least one evoked potential rec... | The University of Hong Kong | A61B005/04
US08611990 | B2 | 2013-12-17 | Systems, methods and computer program products for heart monitoring | 2011-02-25 | The present invention provides systems, methods and computer program products for monitoring a heart. According to one embodiment, the system includes an implantable registering unit. The registering unit comprises a first controller structured to register an electrical signal from the heart. The system includes a second controller in operable communication with the first controller. The second controller comprises a data repository structured to receive data corresponding to the registered electrical signal and being structured to store the data. The data repository stores data corresponding to a baseline electrical signal of the heart. The second controller is structured to receive the data from the first controller corresponding to the registered electrical signal and to compare the registered electrical signal to the baseline electrical signal to determine whether the heart is functioning properly. | 1. A system for monitoring a patient's heart, comprising: a registering unit comprising a first controller in electrical communication with the patient's heart, said first controller being configured to register an electrical signal emitted by the patient's heart; anda second controller in operable communication with said first controller, said second controller being configured to receive data from said first controller corresponding to the registered electrical signal; to generate a registered template comprising a digital graphical representation of an intracardiac electrogram for the registered electrical signal; and to compare the registered template to a baseline template comprising a digital graphical representation of an intracardiac electrogram for a baseline electrical signal emitted by the patient's heart to determine whether the patient's heart is functioning properly, wherein the second controller is adapted to compare the baseline template and the registered template through a wavelet transformation in which the templates are transformed to create one or more wavelets in the form of corresponding fragments of the templates, the second controller being adapted to compare one or more registered wavelets to one or more corresponding baseline wavelets to determine whether the patient's heart is functioning properly. 2. The system of claim 1, wherein said registering unit further comprises a transmitter in operable communication with said second controller. 3. The system of claim 1, further comprising a relay unit independent of said first controller and said second controller and configured to communicate with said first controller and said second controller, said relay unit being configured to receive data from said first controller corresponding to the registered electrical signal and to transmit the received data to said second controller. 4. The system of claim 3, wherein the relay unit is configured to wirelessly communicate with said registering unit. 5. The system of claim 1, wherein said second controller is located remotely from said registering unit. 6. The system of claim 1, wherein said second controller includes a data repository for storing data corresponding to the registered template and data corresponding to the baseline template of the patient's heart. 7. The system of claim 1, further comprising a plurality of electrodes in electrical communication with said registering unit and configured to be coupled directly to the heart. 8. The system of claim 1, wherein said second controller is configured to determine whether a predetermined critical match-percent threshold has been exceeded based on a comparison of the registered template and the baseline template. 9. The system of claim 8, wherein the critical match-percent threshold comprises at least 70% correlation between the registered template and the baseline template, wherein a critical-match percent threshold less than 70% indicates that the heart is not func... | Lifescience Solutions LLC | A61B005/0402
US08439896 | B2 | 2013-05-14 | Device, system, and method for targeted delivery of anti-inflammatory medicaments to a mammalian subject | 2009-11-13 | A device, a system, or a method is described for treating a disease or a condition of one or more joints of articulating bone in a mammalian subject. The device provides one or more medicaments to one or more joints of the mammalian subject. A device is described that includes an enclosure including one or more sensors, a controller, and one or more applicators configured to surround one or more joints of articulating bone of a mammalian subject, wherein the one or more sensors are configured to detect one or more physiological conditions of the one or more joints of the mammalian subject, and the controller, configured to communicate with the one or more sensors, is configured to activate the one or more applicators, and wherein the one or more applicators are configured to inject one or more medicaments to one or more joint tissues of the mammalian subject. | 1. A device comprising: an enclosure including one or more sensors, a controller, and one or more applicators configured to surround one or more joints of articulating bone of a mammalian subject; wherein, the one or more sensors are configured to detect one or more physiological conditions of the one or more joints of the mammalian subject,the controller is configured to communicate with the one or more sensors, and further configured to position the one or more applicators near the one or more joints detected by the one or more sensors as exhibiting the one or more physiological conditions and configured to activate the one or more applicators, andthe one or more applicators are configured to inject one or more medicaments to one or more joint tissues of the mammalian subject. 2. The device of claim 1, wherein the one or more sensors are configured to detect the one or more physiological conditions of the one or more joints by ultrasonic imaging of the one or more joints. 3. The device of claim 2, further including a liquid configured between an ultrasonic transducer and the one or more joints for the ultrasonic imaging of the one or more joints of the mammalian subject. 4. The device of claim 1, wherein the one or more sensors are configured to detect the one or more physiological conditions of the one or more joints by X-ray imaging of the one or more joints. 5. The device of claim 1, wherein the one or more applicators includes one or more high speed microjets, one or more microneedles, one or more microfine lances, one or more microfine cannulas, one or more microprojections, or one or more electrodes. 6. The device of claim 5, wherein the one or more applicators includes one or more high speed microjets. 7. The device of claim 6, wherein the one or more applicators further includes one or more microneedles, one or more microfine lances, or one or more microfine cannulas. 8. The device of claim 5, wherein the one or more applicators includes electrodes on microprojections configured to apply electrical energy to skin of the mammalian subject. 9. The device of claim 8, wherein the electrodes on microprojections provide ablation of stratum corneum in an area beneath the electrodes thereby generating a plurality of hydrophilic microchannels in the skin of the mammalian subject. 10. The device of claim 1, wherein the one or more applicators are configured to respond to the one or more sensors, wherein the one or more applicators are configured to inject the one or more medicaments to the one or more joint tissues of the mammalian subject. 11. The device of claim 1, wherein the one or more sensors are configured to detect movement of the one or more body contours. 12. The device of claim 11, wherein the one or more sensors are configured to monitor joint movement, body movement, or gait of the mammalian subject. 13. The device of claim 11, wherein the one or more sensors are configured to monitor posture of the mammalian subject. 14... | The Invention Science Fund I, LLC | A61K009/22
US08506482 | B2 | 2013-08-13 | Method and system for providing continuous calibration of implantable analyte sensors | 2011-02-07 | Method and system for providing continuous calibration of analyte sensors includes calibrating a first sensor, receiving data associated with detected analyte levels from the first sensor, and calibrating a second sensor based on a predetermined scaling factor and data associated with detected analyte levels from the first sensor, is disclosed. | 1. A system for providing analyte sensor calibration, comprising: a first analyte sensor configured for fluid contact with an analyte under a skin layer to monitor an analyte level;a second analyte sensor configured for fluid contact with the analyte under the skin layer to monitor the analyte level; andone or more processors in signal communication with the first analyte sensor and the second analyte sensor, the one or more processors configured to receive data associated with the monitored analyte level from the first analyte sensor, and to calibrate the second analyte sensor with reference to the received data associated with the monitored analyte level received from the first analyte sensor by determining a scaling factor that includes comparing data from the first analyte sensor with the data from the second analyte sensor and based on an initial scaling factor with a predetermined level of average error between the data of the first analyte sensor and the data of the second analyte sensor, applying the determined scaling factor to the data from the second analyte sensor, and determining a correlation level of signals from the first analyte sensor and the signals from the second analyte sensor. 2. The system of claim 1 further including a third analyte sensor in signal communication with the one or more processors, wherein the one or more processors are configured to calibrate the third analyte sensor based on a second scaling factor and data associated with detected analyte levels from the second analyte sensor. 3. The system of claim 1 wherein the one or more processors are configured to start the calibration of the second analyte sensor after a predetermined time period has elapsed since the first analyte sensor is maintained in fluid contact with the analyte under the skin layer. 4. The system of claim 3 wherein the predetermined time period includes at least approximately 90% of the life of the first analyte sensor. 5. The system of claim 3 wherein the predetermined time period includes at least approximately 50% of the life of the first analyte sensor. 6. The system of claim 1 wherein the one or more processors are configured to determine a sensitivity associated with the first analyte sensor. 7. The system of claim 1 wherein the one or more processors are configured to receive data associated with detected analyte levels from the second analyte sensor. 8. The system of claim 1 wherein the one or more processors are configured to calibrate the second analyte sensor by determining an analyte level and comparing the determined analyte level with the data associated with the detected analyte level from the first analyte sensor. 9. The system of claim 1 wherein the one or more processors are configured to compare the correlation level with a predetermined correlation threshold. 10. The system of claim 9 wherein the predetermined correlation threshold defines an acceptable stability level of the second analyte sensor, and further, ... | Abbott Diabetes Care Inc. | A61B005/05
US08433386 | B2 | 2013-04-30 | Subdermal needle electrode cable assembly having movable needle safety cover integral therewith | 2010-11-01 | An apparatus for monitoring neurological and/or neurophysiological signals of a patient has a signal conductor, an electrode, and a safety cover. The signal conductor has a first end portion and a second end portion. The electrode has a first end portion and a second end portion. The first end portion is a skin piercing portion of the electrode and the second end portion is electrically connected to the first end portion of the signal conductor. The safety cover has an electrode shrouding space therein. The electrode shrouding space is configured in a manner allowing the first end portion of the electrode to be positioned within the electrode shrouding space as a result of the safety cover being moved along a length of the signal conductor toward the first end portion of the electrode. | 1. An apparatus for monitoring at least one of neurological signals and neurophysiological signals of a patient, comprising: a signal conductor having a first end portion and a second end portion;an electrode having a first end portion and a second end portion, wherein the first end portion is a skin piercing portion of the electrode and wherein the second end portion is electrically connected to the first end portion of the signal conductor; anda safety cover having an electrode shrouding space therein, wherein the signal conductor extends through a central passage of the safety cover, wherein the electrode shrouding space is configured in a manner allowing the first end portion of the electrode to be positioned within the electrode shrouding space as a result of the safety cover being moved along a length of the signal conductor toward the first end portion of the electrode. 2. The apparatus of claim 1 wherein: the safety cover has a first end portion and a second end portion;the central passage extends through the safety cover between the first end portion and the second end portion thereof; andthe electrode shrouding space is defined by the central passage and extends through the first end portion of the safety cover. 3. The apparatus of claim 2, further comprising: an electrical connector connected between the electrode and the signal conductor for providing electrical conductivity between the electrode and the signal conductor, wherein the central passage includes a reduced cross sectional portion at the second end portion of the safety cover such that the safety cover engages the electrical connector for defining a maximum amount of displacement of the safety cover with respect to the signal conductor when the safety cover is being moved toward the first end portion of the electrode. 4. The apparatus of claim 3, further comprising: a protective cover removably secured over the first end portion of the electrode for concealing the first end portion of the electrode prior to its use, wherein the safety cover is inhibited from being moved along the length of the signal conductor toward the first end portion of the electrode for causing the first end portion of the electrode to become positioned within the electrode shrouding space until the protective cover is removed from its secured position over the first end portion of the electrode. 5. The apparatus of claim 3 wherein: said apparatus is one of a set of apparatuses for monitoring at least one of neurological signals and neurophysiological signals;the electrical connector includes identifying information uniquely identifying said one apparatus of the set with respect to each other one of said apparatuses of the set. 6. The apparatus of claim 5 wherein said identifying information includes at least one of color-coded identifying information and alphanumeric identifying information. 7. The apparatus of claim 6 wherein said identifying information includes both said color-coded identify... | - | A61B005/0492 A61N001/05
US08874186 | B2 | 2014-10-28 | Apparatus and method for monitoring physiological parameters using electrical measurements | 2009-12-30 | A system for monitoring a physiological parameter comprises a substrate, a pair of drive electrodes, a pair of detection electrodes, and an RFID apparatus. The substrate is arranged to be removably securable to a biological organism. At least the pair of drive electrodes or the pair of detection electrodes is secured to the substrate. The RFID apparatus is arranged to be in electrical communication with at least the pair of drive electrodes or the pair of detection electrodes. Methods of using the device are also provided. | 1. A system for monitoring a physiological parameter comprising: a substrate arranged to be removably securable to a biological organism;a first drive electrode, a second drive electrode, a first detection electrode, a second detection electrode, and an RFID device having a chip with a unique identifier, each disposed on the substrate;an RFID transceiver; andwherein the RFID device is in electrical communication with each of the first drive electrode, the second drive electrode, the first detection electrode, and the second detection electrode,wherein the RFID device is configured to provide an electrical signal to each of the first and second drive electrodes and the RFID device is configured to detect a voltage at each of the first and second detection electrodes,wherein the RFID device is configured to determine a voltage profile based at least in part on the voltages detected at the first and second detection electrodes, andwherein the RFID device is configured to determine the physiological parameter, which is a wound depth, by applying AC currents of varying frequency, and wherein the RFID transceiver is arranged to transmit radio signals to, and receive radio signals from, the RFID device. 2. A system as recited in claim 1, wherein the RFID device further comprises: logic configured to determine an additional physiological parameter from the voltage profile; andthe additional physiological parameter includes at least one of: heart rate, respiration rate, body temperature, perspiration rate, the presence or level of bacteria, body hydration, blood oxygen levels or combinations thereof. 3. A system as recited in claim 1, wherein the RFID device further includes a memory device. 4. A system as recited in claim 1, wherein the RFID device includes an on-board power supply. 5. A system as recited in claim 4, wherein the on-board power supply is a battery. 6. A system as recited in claim 1, wherein the RFID device is configured to determine the voltage profile at various wound depths. 7. A system as recited in claim 1, wherein the electrical signal provided to the first drive electrode comprises a first voltage and the electrical signal provided to the second drive electrode comprises a second voltage. 8. A system for monitoring a physiological parameter, comprising a patch, a first drive electrode, a second drive electrode, a first detection electrode, a second detection electrode, a passive RFID apparatus, and an RFID transceiver; wherein: the patch is arranged to be removably securable to a biological organism;the first drive electrode and the first detection electrode are each coupled to the patch;the second drive electrode is coupled to the patch and located proximate to the first drive electrode;the second detection electrode is coupled to the patch and located proximate to the first detection electrode;the first drive electrode, the second drive electrode, the first detection electrode, and the second detection electrode are coupled... | Avery Dennison Corporation | A61B005/053 A61B005/00
US08560073 | B2 | 2013-10-15 | System and apparatus for automated quantitative assessment, optimization and logging of the effects of a therapy | 2010-03-22 | A method for assessment, optimization and logging of the effects of a therapy for a medical condition, including (a) receiving into a signal processor input signals indicative of the subject's brain activity; (b) characterizing the spatio-temporal behavior of the brain activity using the signals; (c) delivering a therapy to a target tissue of the subject; (d) characterizing the spatio-temporal effect of the therapy on the brain activity; (e) in response to the characterizing, optimizing at least one parameter of the therapy if the brain activity has not been satisfactorily modified and/or has been adversely modified by the therapy; (f) characterizing the spatio-temporal effect of the at least one optimized parameter; and (g) logging to memory the at least one optimized parameter. A computer readable program storage unit encoded with instructions that, when executed by a computer, performs the method. | 1. A computer readable program storage unit encoded with instructions that, when executed by a computer, performs a method for assessment, optimization and logging of the effects of a therapy for a medical condition, said method comprising the steps of: (a) receiving into a signal processor input signals indicative of the subject's electrochemical brain activity;(b) characterizing the spatio-temporal behavior of said electrochemical brain activity using said signals;(c) delivering a therapy to a target tissue of the subject, wherein the therapy is delivered by a single therapy source;(d) characterizing the spatio-temporal effect of said therapy on said electrochemical brain activity;(e) in response to said characterizing said spatio-temporal effect, optimizing at least one parameter of the therapy if said electrochemical brain activity has not been satisfactorily modified and/or has been adversely modified by said therapy;(f) characterizing the spatio-temporal effect on said electrochemical brain activity of said at least one optimized parameter; and(g) logging to memory said at least one optimized parameter. 2. The computer readable program storage unit of claim 1, wherein said method further comprises logging to memory at least one effect of the optimized therapy on the subject's electrochemical brain activity. 3. The computer readable program storage unit of claim 1, wherein step (d) of said method comprises quantitatively characterizing said spatio-temporal effect of said therapy on said electrochemical brain activity. 4. The computer readable program storage unit of claim 1, wherein step (d) of said method comprises semi-quantitatively or qualitatively characterizing said spatio-temporal effect of said therapy on said electrochemical brain activity. 5. The computer readable program storage unit of claim 1, wherein said medical condition is epilepsy. 6. The computer readable program storage unit of claim 5, wherein steps (a) through (e) of said method are performed at or before one or more of the electrographic onset of a seizure, the clinical onset of a seizure, a loss of responsiveness, a loss of consciousness, the electrographic termination of a seizure, a recovery of consciousness, or a recovery of responsiveness. 7. The computer readable program storage unit of claim 1, wherein said input signals are one or more of electrical signals, magnetic signals, thermal signals, optical signals, chemical signals, or cognitive signals. 8. The computer readable program storage unit of claim 7, wherein in said method: when said input signals are electrical signals, characterizing said spatio-temporal behavior is performed in real-time using time-frequency-energy information in one, or more dimensions, at one or more time scales; orwhen said input signals are magnetic signals, characterizing said spatio-temporal behavior is performed in real-time using time-frequency-energy information in one or more dimensions and at one or more time scales;wh... | Flint Hills Scientific, LLC | A61N001/08 A61B005/0482
US10898153 | B2 | 2021-01-26 | Multiple cannula image guided tool for image guided procedures | 2011-01-24 | Apparatus and methods are disclosed for use within an image-guided surgical navigation system for facilitating the combined positioning and orientation of multiple surgical implements. A tool guide having multiple cannulas is tracked by a surgical navigation system in real time. Position data of the tool guide is registered and combined with pre-acquired images by the navigation computer. Concurrent graphical representations of the plurality of cannulas are superimposed over the images and displayed. The display allows the surgeon to place the tool guide into the patient's body and position and orient the plurality of cannulas which are then used to place each of the implements. | 1. An apparatus for the placement of surgical implements, comprising: a plurality of cannulas;a fixture coupled to the plurality of cannulas, (a) wherein the fixture includes a body member having a plurality of attachment locations such that each cannula of the plurality of cannulas is attached at one attachment location of the plurality of attachment locations, wherein each one attachment location is spaced apart and fixed relative to each other one attachment location, (b) a first cannula of the plurality of cannulas is further individually adjustable via the body member at a first attachment location of the plurality of attachment locations to (i) vary and select at least an angular position of the first cannula relative to the body member of the fixture and (ii) is adjustable to vary a length and select the length of the first cannula relative to the body member of the fixture, and (c) a post extending from the fixture; anda frame having at least one point, wherein at least one trackable marker is connected to the frame at the at least one point, wherein the frame is selectively and removably interchangeably connected to the fixture via the post at a location spaced apart from the plurality of attachment locations and the plurality of cannulas;wherein the at least one trackable marker is configured to be connected and fixed to the fixture by the frame and tracked with a detector in a detector space to determine a three-dimensional location of the plurality of cannulas in the detector space when the at least one trackable marker is selectively connected and fixed to the fixture;wherein the determined location of the plurality of cannulas is determined relative to a patient based at least on the tracked at least one trackable marker; andwherein a graphical representation is operable to be displayed with a monitor, wherein the graphical representation simultaneously displays all of the plurality of cannulas at the determined three-dimensional location of the plurality of cannulas, wherein the graphical representation is operable to be superimposed on a pre-acquired image at the determined three-dimensional location of the plurality of cannulas including at least a tip of each cannula of the plurality of cannulas. 2. The apparatus of claim 1, further comprising: a locking mechanism having a moveable portion to lock at least one of the selected length or the selected angular position of the first cannula relative to the fixture at the t attachment location of the plurality of attachment locations. 3. The apparatus of claim 2, wherein each cannula of the plurality of cannulas is configured to vary at least one of its length relative to the fixture or vary its angular position relative to the fixture. 4. The apparatus of claim 3, wherein the locking mechanism is configured to lock each cannula of the plurality of cannulas in respective at least one selected length or selected angular position relative to the fixture. 5. The apparatus of claim 4,... | Medtronic Navigation, Inc. | A61B090/10 A61B034/20 A61B006/00 A61B005/06 A61B017/17 A61B090/11 A61B005/00 A61B090/00 A61B034/00 A61B034/10
US09326728 | B2 | 2016-05-03 | BioMEMS sensor and apparatuses and methods therefor | 2011-02-25 | Electronic devices, apparatus, systems, and methods of operating and constructing the devices, apparatus, and/or systems include a wireless sensor configured to measure strain of hardware implanted in a subject. In various embodiments, temporal measurement of the hardware strain includes monitoring changes of the resonant frequency of the sensor. The sensor can be realized as an inductively powered device that operates as an all-on-chip resonator, where the components of the sensor are biocompatible. Additional apparatus, systems, and methods are disclosed. | 1. A method of monitoring changes in a hardware implant device using a sensor device mounted to or integrated with the hardware device, which is implanted in a subject, wherein the sensor device comprises:a substrate; anda resonator configured as a strain gauge composed essentially of: a solid dielectric material disposed over the substrate as a dielectric capacitor;a conductive layer between the substrate and the solid dielectric material and in contact with the substrate and the solid dielectric material; andat least one conductive coil disposed on the solid dielectric material,wherein the substrate, the conductive layer, the solid dielectric material, and the conductive coil are configured to undergo strain as the hardware implant device mounted to or integral with the sensor device undergoes strain,wherein the substrate, the conductive layer, the solid dielectric material, and the conductive coil are configured as a biocompatible sensor device for implantation in a biological subject,wherein the resonator is configured to be inductively powered without an incorporated power supply,wherein the resonator is configured to reflect electromagnetic waves corresponding to a resonant frequency of the resonator in response to an alternating magnetic or electromagnetic field directed toward the sensor device, andwherein the resonant frequency of the resonator shifts in correspondence with strain applied to the resonator by the hardware implant device,wherein the method comprises the steps of:inductively powering the resonator by directing an alternating magnetic or electromagnetic field toward the sensor device at different times over a period;receiving electromagnetic waves reflected from the resonator, which undergoes strain as the hardware implant device undergoes strain, at different times over the period;determining a shift in a resonant frequency of the resonator based on the received reflected electromagnetic waves at different times over the period; anddetermining a temporal change in strain of the hardware implant device based on the determined shift in the resonant frequency of the resonator over the period. 2. The method of claim 1, wherein the hardware implant device is implanted in a biological subject. 3. The method of claim 2, wherein: the resonant frequency determining step determines the resonant frequency of the resonator by inputting the received signals to a spectrum analyzer, andthe resonator is configured to have a resonant frequency ranging from 50 MHz to 7 GHz at no load. 4. The method of claim 2, further including the step of determining a change in the biological subject based on the temporal change in strain of the hardware implant device. 5. The method of claim 2, wherein the temporal change in strain of the hardware implant device is determined using a strain-frequency calibration of the hardware implant device. 6. The method of claim 2, further including the step of monitoring surface bending strains of the hardware i... | INNOVATIVE IN VIVO SENSING, LLC | A61B005/103 A61B005/117 A61B005/00 H01F041/04 H01F017/00
US09345419 | B2 | 2016-05-24 | Method and apparatus for the objective detection of auditive disorders | 2009-07-24 | This invention relates to the apparatus for the objective detection of auditive disorders, by recording arising potentials of a steady state with simultaneous acoustic stimulation via the bones or the airway to be used in newborns. The apparatus comprises the necessary means for the independent generation of two different sound stimuli presented to the subject simultaneously via osseous, and airway electroacoustic transducers, the synchronous recording of the cerebral electrical activity of the subject, the continuous evaluation of the contact of the electrodes, the interactive measurement of the levels of ambient noise, and the wireless digital transmission of the bioelectric activity towards a computer. | 1. An apparatus for the objective evaluation of hearing through brain responses or steady state auditive potentials, comprising: means for simultaneously providing a patient with different stimuli through electro-acoustic transducers for both air and bone conduction modes;means for synchronous recording of the neuroelectric activity of the patient, further comprising first and second audio channels, said stimuli comprise a first stimulus comprising one or several mixed low frequencies, amplitude modulated, carrier tones with a M1 modulation frequency through the bone transducer and first audio channel, and the second stimulus comprises one or several mixed high frequencies, amplitude modulated, carrier tones with a M2 modulation frequency through the airway transducer and second audio channel, with synchronous recording of the neuroelectric activity caused by the simultaneous bone and airway stimulation;andmeans for processing and analyzing the neuroelectric activity to obtain a result of the primary hearing test for classifying, the hearing disorder as a conductive or sensorineural disorder. 2. The apparatus according to claim 1, further comprising means for collecting and measuring the ambient noise level. 3. The apparatus of claim 2, further comprising means for evaluating the ambient noise value and the impedance values for determining whether a pre-established rejection criterion is met for interruption of the hearing evaluation. 4. An apparatus for objectively evaluating the hearing of a subject comprising: (a) a first electro-acoustical transducer comprising means for generating a first stimuli comprising first frequencies specific for bone stimulation and providing a resultant first impedance value;(b) a second electro-acoustical transducer comprising means for generating a second stimuli comprising second frequencies specific for airway stimulation providing a resultant second impedance value, said first frequencies being lower than said second frequencies;(c) said (a) and (b) further comprise means for continuously and simultaneously stimulating of the respective bone and airway; and(d) means for calculating the respective impedance values of (a) and (b) with the simultaneous stimulation to determine whether there is a conductive or sensorineural disorder in the subject. 5. The apparatus of claim 4, further comprising means for measuring the ambient noise level, wherein the ambient noise level and impedance value are calculated according to a pre-established rejection criterion as to whether or not to interrupt the evaluation of the hearing of the subject. 6. The apparatus of claim 4, said first transducer comprises a bone conductor and said second transducer comprises an airway conductor, and further comprising means for synchronous recording of the neuroelectric activity from the respective bone and airway conductors. 7. The apparatus of claim 5, said first transducer comprises a bone conductor and said second transducer compris... | CENTRO DE NEUROSCIENCIAS DE CUBA | A61B005/0484 A61B005/12
US09282903 | B2 | 2016-03-15 | Device for measuring biological information | 2010-11-19 | A device for measuring biological information by being worn on a person includes a sensor unit subjected to measure biological information, a device main unit provided with a concavity and subjected to accommodate the sensor unit, and a holding unit supported on one side by the sensor unit and on the other side by the device main unit; wherein the holding unit is elastic and has a space on one side of the concavity to hold the sensor unit so that at least part of the sensor unit extends outside from the device main unit, and wherein the sensor unit is displaced relative to the device main unit by elastic deformation in accordance with an external force applied to the sensor unit. | 1. A device for measuring biological information by being worn on a measured object, the device comprising: a sensor unit configured to measure biological information;a device main unit provided with a concavity accommodating the sensor unit in the concavity; anda holding unit supporting the sensor unit and the device main unit with a space therebetween, the holding unit being elastic,the holding unit including an inner periphery part, an outer periphery part with an outer periphery surface, and an inclining part,the inner periphery part being coupled to the sensor unit, the inner periphery part having an inner periphery surface, a first surface, and a second surface, the inner periphery surface defining an opening of the holding unit, the first surface and the second surface being opposite to each other such that the inner periphery surface is disposed between the first surface and the second surface, the opening penetrating through the holding unit from the first surface to the second surface, both of the first surface and the second surface contacting the sensor unitthe outer periphery surface being coupled to the device main unit,the inclining part being disposed between the inner periphery part and the outer periphery part,a thickness of the inner periphery part and a thickness of the outer periphery part being greater than a thickness of the inclining part,the sensor unit being displaced relative to the device main unit by elastic deformation of the holding unit in accordance with an external force applied to the sensor unit. 2. The device of claim 1, wherein the sensor unit is provided with a light sensor. 3. The device of claim 1, wherein the device is configured so as to be worn on a person's wrist. 4. The device of claim 1, wherein the sensor unit and the holding unit are made waterproof. 5. The device of claim 4, further comprising an airflow unit configured to provide airflow between an interior of the device main unit and the outside air. 6. The device of claim 5, wherein the airflow unit includes a threaded section, and the airflow or sealing is performed by advancing or retracting the threaded section in accompaniment with rotation. 7. The device of claim 1, wherein the sensor unit is displaced relative to the device main unit in a displacement direction, and the thickness extending in the displacement direction of the internal periphery part and the thickness extending in the displacement direction of the outer periphery part are greater than the thickness extending in the displacement direction of the inclining part. 8. The device of claim 1, wherein the holding unit is integrated as a one piece, unitary member. 9. The device of claim 1, wherein the sensor unit includes a sensor retainer holding the holding unit,the device main unit includes a holding unit retainer protruding, toward the concavity, inward relative to the outer periphery surface of the holding unit, andthe holding unit retainer is configured to contact the... | Seiko Epson Corporation | A61B005/02 A61B005/024 A61B005/00
US20110160613 | A1 | 2011-06-30 | FLUID COLLECTION/INJECTION DEVICE HAVING SYRINGE SAMPLE CONTAINER AND METHOD OF MAKING AND USING THE SAME | 2010-12-21 | Fluid collection/injection device/system includes a fluid sampling syringe having a double-ended needle arranged on a forward end and a rear end. A syringe sampling device has a forward end insertable into the rear end of the fluid sampling syringe. The forward end of the syringe sampling device being puncturable by one needle of the double-ended needle. | 1. A fluid collection/injection device comprising: a body having a front end, a back end, and a main hollow section arranged between the front and back ends;a needle hub securing section arranged on the front end and being structured and arranged to receive therein a needle member;a double-ended needle removably and/or movably coupled to the needle hub securing section and/or removably coupled to a retractable needle hub; anda syringe sampling device having an insertable end which is puncturable by one needle of the double-ended needle. 2. The device of claim 1, wherein the syringe sampling device comprises a body and a plunger having a forward end in sealing engagement with an inner surface of the body. 3. The device of claim 2, wherein the plunger has a rear end configured to allow a user to move the plunger in a rear-ward direction so as to limit an amount of fluid that is withdrawn with the device. 4. The device of claim 2, wherein the plunger has a piston and a breakable and/or frangible and/or separable section. 5. The device of claim 1, wherein the insertable end of the syringe sampling device comprises a re-sealable member that prevents fluid from leaking out of the syringe sampling device when the one needle is withdrawn. 6. The device of claim 1, wherein the syringe sampling device comprises a plunger having a piston and a plurality of breakable and/or frangible sections. 7. The device of claim 1, wherein the syringe sampling device comprises a mechanism that prevents forward movement of a plunger and allows rearward movement of the plunger. 8. The device of claim 1, wherein the syringe sampling device comprises a ratchet mechanism that prevents forward movement of a plunger and allows rearward movement of the plunger. 9. The device of claim 1, wherein the insertable end of the syringe sampling device comprises a removable re-sealable member that prevents fluid from leaking out of the syringe sampling device when the one needle is withdrawn. 10. The device of claim 1, wherein the body, the needle hub securing section, and the double ended needle is a conventional device. 11. The device of claim 1, wherein the body and the needle hub securing section comprise a one-piece member. 12. The device of claim 1, wherein the body and the needle hub securing section comprise separate members. 13. A method of taking a fluid sample using the device of claim 1, the method comprising: inserting the syringe sampling device into the body; andremoving the syringe sampling device from the body. 14. A method of taking a fluid sample using the device of claim 1, the method comprising: inserting the syringe sampling device into the body;withdrawing a plunger of the syringe sampling device by a desired amount; andremoving the syringe sampling device from the body. 15. A fluid collection/injection system comprising: a fluid sampling syringe having a double-ended needle arranged on a forward end and a rear end;a syringe sampling device having a f... | STAT MEDICAL DEVICES, INC. | A61B005/15
US20110152722 | A1 | 2011-06-23 | UTERINE RUPTURE WARNING METHOD | 2010-12-15 | An abnormal uterine wall perforation is inferred by the failure of the distal balloon of a conventional ablation apparatus, to reach a prescribed, “normal” or “expected” pressure within a prescribed time. | 1. A method for detecting a perforation in the uterine wall of a patient, comprising the steps of: a) filling a balloon with a liquid of a predetermined temperature suitable for ablation treatment of a patient;b) inserting said balloon into the uterus of the patient and inflating said balloon into the uterus and inflating it to contact substantially the entire uterine cavity;c) determining the pressure of said inserted and inflated balloon;d) detecting if the pressure in said balloon has reached a prescribed pressure within a prescribed time that signifies normal pressure conditions; ande) taking remedial measures if said prescribed pressure has not been reached within said prescribed time. 2. An apparatus for detecting a perforation in the uterine wall of a patient, comprising: a) a balloon with a liquid of a predetermined temperature suitable for ablation treatment of a patient;b) pressurizing mechanism that inflates and deflates said balloon;c) pressure sensor that senses pressure of inflated balloon;d) timer that measures if said balloon has reached a prescribed pressure within a prescribed time that signifies normal pressure conditions; ande) an micro-controller that is programmed to operate said timer and said pressure sensor. | - | A61B005/103
US20110105855 | A1 | 2011-05-05 | EARLY DETECTION OF COGNITION IMPAIRMENT IN ELDERLY POPULATION | 2010-10-19 | We have discovered novel methods and systems for detecting cognitive impairment in a subject by analyzing the level of xantophylls in macula lutea or red blood cell sample from the subject. We have discovered that decrease of xanthopyll level in the subject is indicative of cognitive impairment. | 1. A method of diagnosing of mild cognitive impairment (MCI) in a subject comprising measuring a xanthophyll level in macula lutea or in a red blood cell (RBC) sample from the subject and comparing the xanthophyll level with a reference xanthophyll level, wherein if the measured xanthophyll level from the subject is lower than the reference xanthophyll level the subject is affected with MCI. 2. The method of claim 1, wherein the xanthophyll is lutein. 3. The method of claim 1, wherein the xanthophyll is a combination of lutein and zeaxanthin. 4. The method of claim 1, wherein the xanthophyll is zeaxanthin. 5. The method of claim 1, wherein the measuring is performed using light absorption of xanthophylls in a densitometer, and the wavelength of the absorbed light is between 450-495 nm. 6. The method of claim 1, wherein the measuring is performed using heterchromatic flicker photometry (HFP). 7. The method of claim 1, wherein the xanthophyll level is measured from RBCs and the xanthophyll is extracted from the RBCs with an organic solvent. 8. The method of claim 7, wherein the organic solvent is selected from a group consisting of chloroform, methanol and hexane. 9. The method of claim 7, wherein the level of the extracted xanthophylls is measured by high performance liquid chromatography (HPLC). 10. The method of a claim 1, wherein the subject is 65 years old or older. 11. The method of claim 1, wherein the subject is not currently affected with cancer, Alzheimer's disease, dementia, Creutzfeldt-Jakob Disease, abnormal thyroid stimulating hormone, vitamin B 12 deficiency, syphilis, sleep apnea, or Parkinson disease. 12. The method of claim 1, wherein the reference xanthophyll level is an average level obtained from a population of healthy subjects without symptoms associated with MCI and is measured by the same method as that used for measuring the xanthophyll level in the subject. 13. A method of determining a need for a dietary supplementation of xanthophylls in a subject, the method comprising: (a). measuring a xanthophyll level in a macula lutea or in a RBC sample from a subject at a first time point;(b). measuring a xanthophyll level in the subject at a second time point using the same method as in step (a), wherein the xanthophyll measured in step (a) and (b) are the same;(c). comparing the xanthophyll level obtained at the first time point with that obtained at the second time point, wherein if there is a decrease in the level of xanthophyll at the second time point compared to the first time point, the subject is in need of dietary supplementation of xanthophylls. 14. The method of claim 13, wherein the first and second time points are separated by 3-6 months. 15. The method of claim 13 further comprising measuring a level of xanthophyll at a third time point and comparing the level of xanthophyll obtained at the third time point with that obtained at the first and/or second time point, wherein the second and third time poin... | TRUSTEES OF TUFTS COLLEGE | A61B005/00 C12Q001/02
US20110160601 | A1 | 2011-06-30 | Wire Free Self-Contained Single or Multi-Lead Ambulatory ECG Recording and Analyzing Device, System and Method Thereof | 2009-12-30 | A device, method and system are provided to continuously store and analyze single or multi-lead ambulatory ECG signals. Each ECG device is self-contained, wire-free, reusable and capable of non-invasive continuous recording, analyzing and indexing of ECG data independently. A pair of snap connectors on the device body attaches the device to standard ECG electrodes to minimize noise caused by the wires. Each device contains a real-time clock to link the ECG data and a detected cardiac event with actual date and time. The device contains a USB port to upload the data to a host system. Multiple such devices are wearable by a patient to form a multi-lead ECG system. There is provided a method to continuously record and analyze ECG data using a self-contained, wire-free ECG device or multiple such devices; and further a method to synchronize the ECG data in multiple devices to realize multi-lead ECG recording. | 1. A self contained, wire-free, ambulatory ECG device for non-invasive continuous recording and analysis of ECG signals, the device comprising: a device body having a flexible substrate, the substrate having embedded electronic components including a microcontroller, a memory embedded with a diagnosis software, a second memory for storing indexed ECG data, a real time clock, a functional switch, electronic circuitry, a plurality of conductors, a battery, and an interface for data communication;two or more snap connectors integrated onto the substrate for engaging with two or more ECG electrodes separate from the device, the electrodes having a means to adhere to the skin of the body;wherein the device body is detachable from the electrodes via the snap connectors and reusable with two or more new electrodes for a subsequent recording and analysis. 2. The device according to claim 1, wherein the flexible substrate is a printed circuit board on which the electronic components reside. 3. The device according to claim 1, wherein the flexible substrate is a soft polymeric material encapsulating all the embedded electronic components. 4. The device according to claim 1, wherein the interface for data communication is an embedded USB interface connector. 5. The device according to claim 1, further comprising a feedback indicator located on the device body for identifying a status of the device. 6. The device according to claim 1, wherein the embedded electronic components are positioned in a middle section of the substrate, one snap connector is positioned on one side of the embedded electronic components, and one or more snap connectors are positioned on an opposing side of the embedded electronic components. 7. The device according to claim 1, wherein the flexible substrate of the device body has two defined sections with a snap connector in each section, and wherein the embedded electronic components are distributed throughout the two sections. 8. The device according to claim 1, wherein the memory comprises an internal flash memory embedded with a cardiac diagnosis algorithm which detects, analyzes, and indexes the recorded ECG data in real time using the microcontroller. 9. The device according to claim 8, wherein the real time clock registers a first continuous ECG data at each time the device is activated for recording from a non recording mode, and indexes recorded ECG data with a time stamp. 10. The device according to claim 1, wherein the second memory comprises an external flash memory and the indexed ECG data comprises recorded raw ECG data, analyzed ECG data and real time clock information. 11. An ECG monitoring and recording system comprising: a self-contained, wire free, ambulatory single lead ECG device which continuously records and analyzes ECG data, the device comprising: electronic components operatively coupled and embedded within a flexible substrate of the device, the components including an embedded microcontroller, cir... | - | A61B005/0402
US20110152720 | A1 | 2011-06-23 | SAMPLE PREPARATION METHODS AND SYSTEMS | 2010-10-20 | Methods and systems to prepare a sample, including a relatively small amount of a biological sample, and to mix the sample with preparation fluid stored in a device. The device may include multiple fluid chambers and a stepped plunger to force a fluid in at least one of the chambers into a sample receiving chamber and then to a storage chamber. Methods and systems disclosed herein may be implemented to collect and controllable mix a sample for analysis and/or storage. | 1. A system, comprising: a portable housing having a sample region, the sample region having a fluid inlet and a fluid outlet, the portable housing further having a fluid chamber and a fluid passage from the fluid chamber to the sample region fluid inlet;a movable cover to enclose the sample region; anda mechanically actuated fluid controller to force fluid from the fluid chamber, through the fluid passage, the sample region fluid inlet, and the sample region fluid outlet, to move at least a portion of a sample from the sample region through the fluid outlet. 2. The system of claim 1, wherein the movable cover is configured to prevent the mechanically actuated fluid controller from forcing fluid flow when the movable cover is in an open position. 3. The system of claim 1, wherein the sample region is located on an enclosable portion of a surface of the portable housing. 4. The system of claim 1, wherein the moveable cover includes a mechanically slideable cover. 5. The system of claim 4, wherein an outer surface of the portable housing has a substantially cylindrical shape, and wherein the moveable cover includes a slideable sleeve surrounding a portion of the portable housing. 6. The system of claim 1, wherein the mechanically actuated fluid controller includes a mechanical actuator to close the cover and to actuate fluid flow when the movable cover is closed. 7. The system of claim 6, wherein: the movable cover and the mechanical actuator are releasably coupled to one another; andthe mechanical actuator is configured to close the cover and to thereafter disengage from the cover and actuate fluid flow when the cover is closed. 8. The system of claim 7, further including: a breakable tab to releasably couple the mechanical actuator to the moveable cover. 9. The system of claim 8, wherein the moveable cover includes a slideable sleeve surrounding a portion of the portable housing. 10. The system of claim 1, wherein the moveable cover includes a hinged cover. 11. The system of claim 1, wherein the moveable cover includes a rotating cover. 12. The system of claim 1, wherein the moveable cover includes a movable window. 13. The system of claim 1, wherein the moveable cover includes a removable window. 14. The system of claim 1, further including: a sample collection pad disposed within the sample region to receive a biological sample. 15. The system of claim 1, further including: a sample collection pad disposed within the sample region to receive an environmental sample. 16. The system of claim 1, further including: a sample collection pad disposed within the sample region to absorb a sample. 17. The system of claim 16, wherein the sample region and the sample collection pad are dimensioned to receive less than approximately 250 micro-liters of the biological sample. 18. The system of claim 1, further including: a lance device to pierce a membrane to release a biological sample therefrom. 19. The system of claim 1, further includin... | BOSTON MICROFLUIDICS | A61B005/151 G01N033/48 C12M001/00
US07962205 | B2 | 2011-06-14 | Human subject index estimation apparatus and method | 2007-10-29 | A CPU (170) of a human subject index estimation apparatus (1) computes a waist circumference based on body weight measured by a weight scale (110) and bioelectric impedance measured by a bioelectric impedance measurement unit (200A), and information such as age and height which was input through input unit (150). The computed waist circumference is stored, along with the information such as age and height, in a third storage unit (140) that is a rewritable non-volatile memory. The CPU (170) displays the waist circumference on display unit (160). | 1. A human subject index estimation apparatus that estimates a circumference of a human subject at a predetermined position between levels of the waist and the hips inclusive, comprising: a weight scale that measures a weight of the human subject; anda computer that executes a computation in accordance with an equation (A) thereby estimating the circumference, wherein the equation (A) is: Y=a1*W+f(k) (A)wherein:Y is the circumference,W is the weight,a1 is a constant,k is a variable, andf(k) is a function of at least one variable, the at least one variable including at least one of sex, race, age, height, body mass index, bioelectric impedance, body fat percentage, and body fat mass,wherein the predetermined position is at the level of the waist,wherein the computer executes the computation of the equation (A) using, as the function f(k), a function f(Z,H,E) as shown in an equation (B) which is: f(Z,H,E)=a2*Z/H+a3*E (B)wherein:Z is bioelectric impedance,H is height,E is age, anda2 and a3 are constants. 2. A human subject index estimation apparatus according to claim 1, wherein the computer changes the constant al and a constant contained in the function f(k) depending on at least one of sex and race of the human subject. 3. A human subject index estimation apparatus according to claim 1, wherein the predetermined position is at the level of the hips,wherein the computer executes the computation of the equation (A) using, as the function f(k), a function f(% Fat,BMI) as shown in an equation (C) which is: f(% Fat,BMI)=a4*% Fat*BMI (C)wherein: % Fat is body fat percentage,BMI is body mass index, anda4 is a constant. 4. A human subject index estimation apparatus according to claim 1, further comprising a bioelectric impedance measurement device that measures bioelectric impedance between the left foot and the right foot, wherein the computer executes the computation of the equation (A) using the measured bioelectric impedance. 5. A human subject index estimation apparatus that estimates a circumference of a human subject at a predetermined position between levels of the waist and the hips inclusive, comprising: a weight scale that measures a weight of the human subject; anda computer that executes a computation in accordance with an equation (A) thereby estimating the circumference, wherein the equation (A) is: Y=a1*W+f(k) (A)wherein:Y is the circumference,W is the weight,a1 is a constant,k is a variable, andf(k) is a function of plural variables, the plural variables including bioelectric impedance and at least one of sex, race, age, height, body mass index, body fat percentage, and body fat mass,the human subject index estimation apparatus further comprising a bioelectric impedance measurement device that comprises: first, second, third, and fourth electrodes contacting a left foot, a right foot, a left hand, and a right hand, respectively;fifth, sixth, seventh, and eighth electrodes contacting a left foot, a right foot, a left hand, and a r... | Tanita Corporation | A61B005/00 A61B005/05 A61B005/103 A61B005/117 C30B033/06 C30B015/00 G01B005/08
US07955274 | B2 | 2011-06-07 | Device for sizing a human female breast | 2006-08-30 | A device for intraoperative sizing of a human female breast during resizing and/or shaping of the breast composes: a (a) substantially hemispherical, rigid, cup-shaped foraminous body, i.e., a foraminous cup, defining a desired size and shape of the breast; and (b) a rim member to which the foraminous cup is affixed. | 1. A method of resizing a human female breast, comprising: obtaining a substantially hemispherical, rigid, cup-shaped foraminous body defining a desired size and/or shape of the breast, said body comprising an annular rim member wherein the foraminous body is affixed to the annular rim member, said foraminous body being configured such that the breast is viewable in relation to the foraminous body by visual observation through the foraminous body from an exterior thereof;exposing a pedicle of the tissue of the breast to be resized and/or reshaped during an operation;resizing and/or reshaping the breast; andpositioning said cup-shaped foraminous body over the breast to determine if the size and/or shape of the breast matches the size and/or shape within the cup-shaped foraminous body by the visual observation from the exterior of the foraminous body;if the size and/or shape of the breast differs from the size and/or shape within the cup-shaped foraminous body, removing said cup-shaped foraminous body from the breast, and continuing said resizing and/or reshaping the breast to revise a size and/or a reshaped outcome during said operation; andrepositioning said cup-shaped foraminous body over the breast to determine if the size and/or shape of the breast matches the size and/or shape within the cup-shaped foraminous body by repeating the visual observation from the exterior of the foraminous body. 2. The method of claim 1 further comprising positioning said cup-shaped foraminous body over the other breast of the female for comparing the two breast sizes, and iteratively adjusting the size of the breast being operated upon so that the breast being operated upon is provided with the same size as the other breast. 3. A method of gauging a desired resizing and/or reshaping a human female breast of a patient and compensating for deviations therefrom intraoperatively, comprising: selecting a device including a substantially hemispherical grating defining a concave receiving space within which the breast is receivable having a cup size and/or shape representative of a desired breast size and/or shape to be achieved, said hemispherical grating being configured such that the breast is viewable in relation to the hemispherical grating which represents the desired breast size and/or shape by visual observation through the hemispherical grating from an exterior of the device;positioning said device with the breast received within the concave receiving space to determine if a size and/or a shape of the breast matches said cup size and/or shape by the visual observation from the exterior of the device;removing said device from the breast if the size and/or shape of the breast differs from the cup size and/or shape;resizing and/or re-shaping the breast; andrepositioning said device with the breast once again received within the concave receiving space to determine, by repeating the visual observation from the exterior of the device, if a resized size and/or a re... | Accurate Surgical & Scientific Instruments Corporation | A61B005/103
US08840561 | B2 | 2014-09-23 | Suprasystolic measurement in a fast blood-pressure cycle | 2009-12-31 | Disclosed herein is a system for monitoring a patient that includes a cuff configured to inflate to at least partially occlude an artery of the patient and a cuff controller configured to inflate the cuff during a dynamic phase and generally maintain inflation of the cuff at about a target pressure during a static phase. The system also includes a sensor configured to receive a signal associated with the at least partially occluded artery and generate an output signal based on the received signal, and a cuff control module configured to determine the target pressure during the dynamic phase and based on the output signal, and control the cuff controller during the dynamic phase and the static phase. | 1. A system for monitoring a patient, comprising: a cuff configured to inflate to at least partially occlude an artery of the patient;a cuff controller configured to inflate the cuff during a dynamic phase and generally maintain inflation of the cuff at about a target pressure during a static phase, wherein the target pressure is a suprasystolic pressure and the static phase is maintained for at least six cardiac cycles;a sensor configured to receive a signal associated with the at least partially occluded artery and generate an output signal based on the received signal; anda cuff control module configured to determine both a systolic pressure and the target pressure during the dynamic phase and based on the output signal, and control the cuff controller during the dynamic phase and the static phase. 2. The system of claim 1, wherein the systolic pressure is determined in less than about 15 seconds. 3. The system of claim 1, wherein the mean error has a standard deviation of less than about ±5 mmHg. 4. The system of claim 1, wherein the sensor includes a pressure sensor. 5. The system of claim 1, wherein the sensor is configured to operate with an oscillometric method. 6. The system of claim 1, further including a signal analysis module configured to determine a hemodynamic parameter based on the output signal of the sensor, wherein the hemodynamic parameter includes at least one of a heart rate, a blood pressure, a vessel compliance, an aortic index, an augmentation index, a reflected wave ratio, and an indication of treatment. 7. The system of claim 6, wherein the hemodynamic parameter is determined based on a first set of data obtained during the dynamic phase and a second set of data obtained during the static phase. 8. The system of claim 7, wherein the first set of data are obtained over more than two cardiac cycles and the second set of data obtained over more than two cardiac cycles. 9. The system of claim 1, wherein the target pressure is determined based on the systolic pressure and the systolic pressure is determined with a mean error of less than about 5 mmHg. 10. A system for monitoring a patient, comprising: a cuff configured to inflate to at least partially occlude an artery of the patient;a cuff controller configured to inflate the cuff during a single dynamic phase and generally maintain inflation of the cuff at about a target pressure during a static phase;a sensor configured to receive a signal associated with the at least partially occluded artery and generate an output signal based on the received signal;a cuff control module configured to determine both a systolic pressure and the target pressure during the single dynamic phase and based on the output signal, and control the cuff controller during the single dynamic phase and the static phase; anda signal analysis module configured to determine a parameter based on the output signal of the sensor, wherein the parameter is determined based on a first set of data obt... | Welch Allyn, Inc. | A61B005/02 A61B005/0225 A61B005/022
US20110125046 | A1 | 2011-05-26 | METHODS AND APPARATUS FOR MONITORING CONSCIOUSNESS | 2010-08-09 | The systems of the present invention provide improved accuracy in monitoring, analysing, detecting, predicting and/or providing alerts and alarms associated with depth of anaesthesia, depth of consciousness, hypnotic state, sedation depth, fatigue or vigilance of a subject, with as few as 3 surface electrodes. The systems incorporate real-time phase, amplitude and frequency analysis of a subject's electro-encephalogram. The systems weight outputs of various types of analyses to produce an integrated analysis or display for precise indication or alert to users of the systems including anaesthetists, nurses and other medical personnel, transport drivers and machine workers. The systems weight the outputs of one or more analysis algorithms including combinations of simultaneous, real-time R&K analysis, AEP spectral analysis-SEF-MF, Bi-coherence analysis, initial wave analysis, auditory response, arousal analysis, body movement analysis, 95% spectral edge analysis and anaesthetic phase and spectral energy variance measurement in association with a subject's state of consciousness. | 1. An apparatus for acquiring physiological data from a living being for determining the state of consciousness of said living being, comprising: means for acquiring at least one continuous biosignal;means for stimulating at least one evoked potential signal in said being;means for acquiring said at least one evoked potential biosignal;means for calculating at least two indices from each acquired biosignal and selecting an index from the calculated indices to represent the state of consciousness from the living being wherein a first of said indices is derived from a transformation of raw signal data according to a first transformation method and the second and subsequent of said indices is derived from a transformation of raw signal data according to a different transformation method or methods from said first transformation method;wherein said first means for acquiring said first biosignal and said means for acquiring at least one evoked signal includes a single means. 2. The apparatus according to claim 1, wherein the at least one biosignal is any one of, or a combination of, at least one EEG signal or at least one muscular activation signal. 3. The apparatus according to claim 2, wherein the at least one muscular activation signal is a measure of eyelid movement. 4. The apparatus according to claim 2, wherein the EEG signal is a continuous signal. 5. The apparatus according to claim 2, including means for deriving said evoked potential signal from the EEG signal. 6. The apparatus according to claim 1, wherein the means for a acquiring the at least one continuous biosignal or means for acquiring the at least one evoked potential signal includes at least one electrode sensor. 7. The apparatus according to claim 1, including means for monitoring the at least one sensor for signal integrity. 8. The apparatus according to claim 1, including means for monitoring the at least one sensor for signal quality. 9. The apparatus according to claim 1, wherein the means for acquiring the at least one continuous biosignal or means for acquiring the at least one evoked potential signal includes at least one disposable or semi-disposable sensor. 10. The apparatus according to claim 9, wherein said at least one disposable or semi-disposable sensor includes means for activating an electrical energy source. 11. The apparatus according to claim 10, wherein the means for activating said energy source includes the packaging of said electrical energy source. 12. The apparatus according to claim 1, wherein the means to acquire the at least one biosignal is an electrode sensor activatable in response to pressure from an operator or user of said apparatus. 13. The apparatus according to claim 1, wherein the means for stimulating the evoked potential signal stimulates any one or a combination of somatosensory, auditory, or visual evoked response. 14. The apparatus according to claim 13, wherein the means for stimulating said auditory evoked response signal is... | - | A61B005/0484 A61B005/11
US20110160641 | A1 | 2011-06-30 | BATTERY PACK AND ELECTRONIC DEVICE ASSEMBLY INCLUDING BATTERY PACK | 2009-08-28 | This invention relates to an improvement in battery packs. The invention intends to provide a highly flexible battery pack and an electronic device assembly using the same. The battery pack of the invention includes an elastic sheet and at least one coin battery embedded in the elastic sheet. In a preferable embodiment of the invention, the electronic device assembly includes a laminate of the battery pack and a flexible electronic device. The electronic device is driven by power supplied through positive and negative terminals of the battery pack. In another preferable embodiment of the invention, the electronic device assembly includes the battery pack and a device embedded in the elastic sheet of the battery pack, and the device is driven by power supplied by the at least one coin battery. | 1. A battery pack comprising an elastic sheet and at least one coin battery embedded in the elastic sheet. 2. The battery pack of claim 1, wherein a positive electrode of the at least one coin battery or a positive terminal connected to the positive electrode and a negative electrode of the at least one coin battery or a negative terminal connected to the negative electrode are exposed at a surface of the elastic sheet. 3. The battery pack of claim 1, wherein the elastic sheet comprises silicone rubber. 4. The battery pack of claim 1, wherein the at least one coin battery comprises two or more coin batteries disposed in a horizontal direction of the elastic sheet,the two or more coin batteries are electrically connected by flexible wiring, andthe wiring is disposed mainly in a central part of the thickness of the elastic sheet. 5. The battery pack of claim 4, wherein the wiring has slack so that it is longer than the shortest distance between electrodes of the two or more coin batteries, andwhen the elastic sheet bends, the wiring can deform together with deformation of the elastic sheet due to disappearance of the slack. 6. The battery pack of claim 5, wherein the slack is contained in a void that is formed in the elastic sheet at a given position along the wiring. 7. The battery pack of claim 1, wherein the at least one coin battery comprises two or more coin batteries connected in series. 8. The battery pack of claim 1, wherein the at least one coin battery comprises two or more coin batteries connected in parallel. 9. An electronic device assembly comprising a laminate of the battery pack of claim 1 and a flexible electronic device, the electronic device being driven by power supplied through positive and negative terminals of the battery pack. 10. An electronic device assembly comprising: the battery pack of claim 1; and a device embedded in the elastic sheet of the battery pack, the device being driven by power supplied by the at least one coin battery of the battery pack. 11. The electronic device assembly of claim 9, which is a portable device that operates in contact with the skin of a living body. 12. The electronic device assembly of claim 11, which is a biological information measuring device or a transdermal iontophoretic device. 13. The electronic device assembly of claim 10, which is a portable device that operates in contact with the skin of a living body. 14. The electronic device assembly of claim 13, which is a biological information measuring device or a transdermal iontophoretic device. | - | A61N001/30 H01M010/02 H01M002/24 A61B005/00
US08340753 | B2 | 2012-12-25 | Binaural beat augmented biofeedback system | 2008-11-14 | A biofeedback system that incorporates binaural beat technology to augment the EEG feedback provided to the trainee. Specifically, the system will provide binaural beats at the frequency of the brain's naturally occurring peak energy in the EEG band or bands subject to the biofeedback. | 1. A method for training a person to develop useful degrees of voluntary control of personal electroencephalographic (EEG) activity, said method comprising the steps of: (a) securing EEG leads to the person;(b) receiving an EEG signal from the person;(c) transmitting the resultant EEG signal to a filter;(d) filtering the EEG signal; andusing a computing device:(e) monitoring a predetermined EEG frequency band or subband;(f) determining an EEG frequency value of a preselected aspect of said EEG signal; and(g) generating a first audio signal at a first frequency that is delivered to said person's first ear with minimal delay and generating a second audio signal that is delivered to said person's second ear concurrently with said first audio signal, said second audio signal having a second frequency that differs from said first frequency by the value of said EEG frequency value;wherein said preselected aspect is the frequency at which the peak energy of the EEG signal occurs. 2. The method of claim 1 wherein the frequency of the first audio signal is constant and the volume of first audio signal varies in proportion with the amplitude of said EEG signal. 3. The method according to claim 1 wherein said first audio signal tone is between about 400 Hz and about 800 Hz as a representative signal for alpha training. 4. The method according to claim 1 wherein said first audio signal has a tone above about 800 Hz as a representative signal for theta training. 5. The method of claim 1 wherein said minimal delay is less than 350 MS. 6. A method for training a person to develop useful degrees of voluntary control of personal electroencephalographic (EEG) activity, said method comprising the steps of: (a) securing EEG leads to the person;(b) receiving an EEG signal from the person;(c) transmitting the resultant EEG signal to a filter;(d) filtering the EEG signal; andusing a computing device:(e) monitoring a predetermined EEG frequency band or subband;(f) determining an EEG frequency value of a preselected aspect of said EEG signal; andg) generating a first audio signal at a first frequency that is delivered to said person's first ear with minimal delay and generating a second audio signal that is delivered to said person's second ear concurrently with said first audio signal, said second audio signal having a second frequency that differs from said first frequency by the value of said EEG frequency value;wherein said preslected aspect is the frequency at which the greatest average energy of the EEG signal occurs over a pre-selected interval of time. 7. The method of claim 6 wherein the frequency of the first audio signal is constant and the volume of first audio signal varies in proportion with the amplitude of said EEG signal. 8. A method for training a person to develop useful degrees of voluntary control of personal electroencephalographic (EEG) activity, said method comprising the steps of: (a) securing EEG leads to the person;(b) receiving an EEG s... |  | A61B005/04
US08255041 | B2 | 2012-08-28 | Wireless ECG system | 2011-02-03 | A wireless monitoring system and, more particularly, a wireless monitoring system for detecting and transmitting physiological data. The present invention detects physiological data relating to a patient's cardiac activity and respiration rate and transmits the data to a remote base station via telemetry. The base station processes the data so that the data can be display by an ECG monitor. | 1. A system for transmitting data comprising: a chest assembly comprising at least a first electrode connector and a second electrode connector configured to couple electrodes to traces that run along the chest assembly;a discrete body electronics unit separable from the chest assembly, the chest assembly being removably connected proximate to and in electrical contact with the body electronics unit, wherein the body electronics unit acquires physiological data from the chest assembly and wirelessly transmits the physiological data;a plurality of base stations for receiving the physiological data transmitted from the body electronics unit, each base station capable of being simultaneously paired to the body electronics unit by activation of a switch on the body electronics unit and then pairing the body electronics unit with the plurality of base stations such that each of the base stations is enabled to receive physiological data transmitted by the body electronics unit, each base station comprising a plurality of terminals for transmitting the physiological data to a monitor. 2. The system of claim 1, wherein the physiological data is transmitted from the body electronics unit to the plurality of base stations using two-way communication. 3. The system of claim 1, wherein the body electronics unit is configured to transmit non-physiological data to each base station. 4. The system of claim 3, wherein the physiological data corresponds to a patient's respiration rate and cardiac activity and the non-physiological data corresponds to pacemaker pulse data. 5. The system of claim 1, wherein each base station is configured to transmit non-physiological signals to the body electronics unit. 6. The system of claim 1, wherein the body electronics unit is configured to continuously transmit physiological data to a first base station and to simultaneously transmit physiological data to a second base station for a temporary amount of time. 7. The system of claim 1, wherein the chest assembly further comprises: a base layer having a first side and a second side, wherein the first side contains a plurality of electrically conductive elements, the electrically conductive elements configured to be removably coupled to electrodes, the base layer positioned between a first insulating layer and a second insulating layer. 8. The system of claim 7, wherein the chest assembly further comprises a plurality of electrode housings, each electrode housing positioned over an aperture formed in the chest assembly and containing an elastomeric portion defining a female void for receiving a portion of an electrode. 9. The system of claim 8, wherein the aperture forms at least one flap for engaging an electrode positioned through the aperture. 10. The system of claim 7, wherein the chest assembly further comprises a plurality of electrode housings, each electrode housing positioned over an aperture formed in the chest assembly and containing an elastomeric portion d... | LifeSync Corporation | A61B005/04
US08287468 | B2 | 2012-10-16 | Blood component measurement device and tip for blood measurement | 2011-02-16 | A blood component measurement device has a puncturing unit provided inside a housing and having a puncture needle; a button unit for advancing and retreating the puncture needle; a contact mechanism coming into contact with the skin of a person to be measured; and a holder for displaceably holding the puncturing unit. The blood component measurement device also has an introduction section at the substantially center of a contact member. After the skin is punctured by the puncture needle provided at the substantially center of the mechanism, the introduction section introduces blood to testing paper when the contact member of the contact mechanism is displaced in a sliding manner along a guide member. | 1. A tip for blood measurement, comprising: a mounted section to be mounted to a blood component measurement device;a puncture needle possessing an axis and movable for being advanced in an axial direction of the needle to puncture skin and produce a blood sample, and for being retracted in the axial direction of the puncture needle;a contact section which has a cavity permitting said puncture needle to pass therethrough and which is contactable with the skin;a guide section adjacent the contact section;a measurement unit for measuring a component of the blood sample; anda blood introduction section in communication with the measurement unit,wherein said contact section is configured to obliquely slide relative to the blood introduction section so that the contact section is obliquely displaceable relative to the axial direction of said puncture needle, andsaid contact section is provided with a first slant surface opposed to the guide section and inclined at a predetermined angle relative to the axial direction of said puncture needle, said guide section is provided with a second slant surface in contact with said first slant surface and being at substantially the same angle as said first slant surface, and said contact section is displaceable along said second slant surface via said first slant surface. 2. The tip for blood component measurement as set forth in claim 1, wherein the puncture needle is held in a hub that is slidably positioned inside the guide section. 3. The tip for blood component measurement as set forth in claim 1, wherein the first slant surface is angled at an oblique angle relative to the axial direction of said puncture needle. 4. The tip for blood component measurement as set forth in claim 1, wherein the puncture needle is movably mounted in a holding hole in the guide section for movement in the axial direction,the blood introduction section comprises an introduction port through which the blood sample is passable from the skin to the measurement unit,the holding hole opens to one side surface of the guide section, andthe introduction port extends outwardly from the one side surface of the guide section. 5. The tip for blood component measurement as set forth in claim 1, wherein the contact section is engageable with a stopper in the blood component measurement device, the stopper for restricting displacement of the contact section when engaged with the contact section, and the contact section is movable out of engagement with the stopper after puncture of the skin by the puncture needle to permit the displacement of the contact section. 6. A tip for blood component measurement comprising: a puncture needle possessing an axis, the needle being movable in an axial direction of the needle to perform advancing movement during which the needle is configured to puncture skin to produce a blood sample, and being movable in the axial direction of the needle to perform retracting movement;a contact section adapted to cont... | Terumo Kabushiki Kaisha | A61B005/00 A61B017/32 A61B017/14 B65D081/00
US08633695 | B2 | 2014-01-21 | Adiabatic multi-band RF pulses for selective signal suppression in a magnetic resonance imaging | 2009-08-13 | A magnetic resonance imaging (MRI) system, comprising: a magnetic resonance imaging scanner comprising: a main magnet providing a substantially uniform main magnetic field B0 for a subject under observation; and a radio frequency (RF) coil configured to irradiate a radio frequency (RF) pulse into a region of interest of the subject under observation, wherein the RF pulse comprises a base pulse comprising an adiabatic pulse having a first bandwidth time product (BWTP), wherein the RF pulse selectively suppresses magnetic resonance signals from more than one chemical component or more than one spatial region within the region of interest of the subject under observation, and wherein the adiabatic pulse is characterized by an amplitude modulation function and a frequency modulation function. | 1. A magnetic resonance imaging (MRI) system, comprising: a magnetic resonance imaging scanner comprising:a main magnet providing a substantially uniform main magnetic field B0 for a subject under observation; anda radio frequency (RF) coil configured to irradiate a radio frequency (RF) pulse into a region of interest of said subject under observation,wherein said RF pulse is a resultant pulse including a summation of a first base pulse comprising a first adiabatic pulse and a second base pulse comprising a second adiabatic pulse,wherein said RF pulse selectively suppresses magnetic resonance signals from more than one chemical component or more than one spatial region within said region of interest of said subject under observation, andwherein said first adiabatic pulse is characterized by a first amplitude modulation function and a first frequency modulation function, and said second adiabatic pulse is characterized by a second amplitude modulation function and a second frequency modulation function. 2. The MRI system according to claim 1, wherein the at least one of the first and second adiabatic pulses is a form of a hyperbolic-secant (HS) pulse. 3. The MRI system according to claim 1, wherein at least one of said first and second base pulses further comprises a blunt pulse to form a hyper-pulse. 4. The MRI system according to claim 3, wherein the blunt pulse is a Numerically-Optimized-Modulation (NOM) pulse or a HS pulse with an order above 2. 5. The MRI system according to claim 3, wherein at least one of said first and second bases pulses has a first bandwidth and time product (BWTP), and said blunt pulse has a second bandwidth and time product (BWTP) that is substantially smaller than said first BWTP. 6. The MRI system according to claim 1, wherein said RF pulse further comprises at least a new instance of at least one of said base pulses with a new scaling factor for said amplitude modulation or said frequency modulation. 7. The MRI system according to claim 1, wherein at least one of said base pulses is a hyper-pulse. 8. The MRI system according to claim 1, wherein said more than one chemical components comprises at least one of water, triacylglycerol, N-acetyleaspartase (NAA), or combinations thereof. 9. The MRI system according to claim 1, wherein said subject is a human, an animal, a phantom, a sample, or combinations thereof. 10. The MRI system according to claim 1, further comprising: a signal processing unit in communication with said magnetic resonance imaging scanner to receive a plurality of magnetic signals from said region of interest in response to said RF pulse. 11. The MRI system according to claim 10, wherein said signal processing unit performs spectral editing for said received magnetic resonance signals. 12. The MRI system according to claim 1, wherein said magnetic resonance imaging scanner further comprises a gradient system to generate spoiler gradients to dephase said MR signals from said more than one c... | The Johns Hopkins University | G01V003/00
US08989841 | B2 | 2015-03-24 | Interface devices for use with intracavity probes for high field strength magnetic resonance systems | 2011-02-01 | An interface device enables an intracavity probe to be used with a magnetic resonance (MR) system for the purpose of obtaining images or spectra of a region of interest within a cavity of a patient. The probe includes a shaft, a balloon at one end thereof, and a coil loop within the balloon. The coil loop preferably includes two drive capacitors and a tuning capacitor, all of which in series. A junction node between the drive capacitors serves as a ground for electrically balancing the coil loop. Diametrically opposite the junction node, the tuning capacitor enables the coil loop to resonate at the operating frequency of the MR system. Across each drive capacitor is connected an output cable having an electrical length of SL+n(λ/4). The output cables terminate in a plug that is used to connect the coil loop to the interface device. | 1. An interface device for interfacing an intracavity probe with a magnetic resonance system having a receive cycle and a transmit cycle of operation, said intracavity probe having a coil loop and a pair of output cables terminating in a cable connector therefor for connecting said coil loop to said interface device, said interface device comprising: (a) an interface device connector to which said intracavity probe is connectable via said cable connector thereof;(b) a phase shifting network linked to said interface device connector, said phase shifting network having a first and a second sub-network to which a first and a second of said pair of output cables are joined, respectively, when said cable and said interface device connectors of said intracavity probe and said interface device are connected, each of said output cables connected at the other end thereof across a drive capacitor in said coil loop and having an electrical length of SL+n(λ/4) wherein SL is a supplemental length whose reactance is of a same magnitude as a reactance of said drive capacitor corresponding thereto, n is an odd integer, and λ is a wavelength of the operating frequency of said magnetic resonance system; and(c) a pair of PIN diodes of which a first PIN diode is connected across an output of said first sub-network and a second PIN diode is connected across an output of said second sub-network such that an electrical length of each of said first and said second sub-networks from said interface device connector to said PIN diode corresponding thereto is λ/4;wherein when said cable and said interface device connectors of said intracavity probe and said interface device are connected, said phase shifting network enables said coil loop (i) during said receive cycle when said PIN diodes are reverse biased to be coupled through said output cables to a probe input port of said magnetic resonance system thus allowing magnetic resonance signals received from each of said output cables to be constructively combined and routed to said probe input port, and (ii) during said transmit cycle when said PIN diodes are forward biased to be decoupled from a transmit field of said magnetic resonance system via a combined electrical length of SL +n(λ/2) of each of said output cables and said sub-network corresponding thereto. 2. The interface device of claim 1, wherein said phase shifting network includes: a preamplifier for providing gain and impedance matching between said phase shifting network and said probe input port so that said magnetic resonance signals received from said phase shifting network are passed with enhancement of signal-to-noise ratio to said probe input port during said receive cycle. 3. The interface device of claim 2, wherein said preamplifier includes: (a) a GASFET having a gate, a source, and a drain; and(b) a series resonant circuit connected between said phase shifting network and said gate, said series resonant circuit for coupling said intracavity probe to... | Bayer Medical Care Inc. | A61B005/055 G01R033/28 G01R033/36 G01R033/34
US09261407 | B2 | 2016-02-16 | Thermometer for determining the temperature of an animal's ear drum and method of using the same | 2010-08-19 | The present invention relates to a thermometer for determining the temperature of an animal's ear drum. The thermometer includes a probe, an infrared-radiation detector adapted to receive infrared radiation emitted by the ear drum, and devices that help determine the probe's position in the ear canal so as to optimize the infrared radiation received from the ear drum, and to minimize the infrared radiation received from other ear parts. A method of using the thermometer is also disclosed. | 1. A thermometer for determining a temperature of a vertebrate animal's ear drum disposed at the inner end of the animal's ear canal defined by an ear canal wall, said thermometer comprising: a probe adapted to be inserted into the ear canal, said probe possessing a substantially frusto-conical configuration about a substantially longitudinal axis and possessing an outer peripheral surface;a detector adapted to sense infrared radiation emitted by the ear canal and the ear drum, said detector being operatively coupled to said probe;an array of at least three electrically conductive zones disposed adjacent to and arranged about said outer peripheral surface;at least one electrode overlapping at least one of said zones;at least one sensor electrically connected to each of said zones and also electrically connected to said at least one electrode, said at least one sensor adapted to sense the conductivity of each of said zones as said probe is inserted into the ear canal and adapted to sense the conductivity of said at least one electrode as said probe is inserted into the ear canal; anda microprocessor configured to assess the orientation of the longitudinal axis of said probe relative to the wall of said ear canal in the vicinity of said probe at least partially based on the conductivity sensed by said at least one sensor. 2. The thermometer according to claim 1 wherein each of said at least three zones comprises an electrode pad, wherein the microprocessor is configured to assess the conductivity relative to a threshold value associated with each of said at least three zones, wherein the microprocessor is configured to assign a weighting value to each of said assessments, and wherein the microprocessor is configured to inhibit the temperature from being determined at least partially based on said assessments and said assignments. 3. The thermometer according to claim 1 wherein each of said at least three zones comprises an electrode pad, wherein the microprocessor is configured to assess the conductivity relative to a threshold value associated with each of said at least three zones, and wherein the microprocessor is configured to inhibit the temperature from being determined unless at least one of said assessments attains at least one predetermined condition. 4. The thermometer according to claim 1 wherein the microprocessor is in operative electrical communication with said at least one sensor and configured to assess the conductivity of each of said zones. 5. The thermometer according to claim 1 wherein the microprocessor is configured to assess whether the aggregate conductivity of each of said zones exceeds a predetermined value. 6. The thermometer according to claim 1 wherein the microprocessor is configured to assess whether the conductivity of a preselected percentage greater than zero percent of the number of said zones exceeds a predetermined value. 7. The thermometer according to claim 1 wherein said outer peripheral surface is for... |  | A61B006/00 G01J005/00 G01J005/02 G01J005/04 G01J005/08
US07965907 | B2 | 2011-06-21 | Medical image handling system and method | 2003-11-18 | A medical image handling system includes a monitor for displaying a medical image, an input device for inputting an image reading report corresponding to the medical image displayed on the monitor, and a processor. The processor judges the presence or absence of an image reading report corresponding to the medical image displayed on the monitor, and restricts a change of the medical image in a case where the image reading report is judged as being absent. | 1. A medical image handling system comprising: a) a monitor for displaying a medical image;b) a processor configured to process a control of judging presence or absence of an inputting of a diagnosis report corresponding to the medical image displayed on said monitor, wherein the diagnosis report corresponding to the medical image displayed on said monitor includes information input by the user while viewing the medical image, displaying an image for inputting the diagnosis report corresponding to the medical image displayed on said monitor in a case where the input of the diagnosis report is judged to be absent, and restricting a change of displaying the medical image in a case where the inputting of the diagnosis report is judged to be absent; andc) an input device for inputting a diagnosis report corresponding to the medical image displayed on said monitor on the basis of a user instruction. 2. A medical image handling system according to claim 1, wherein said processor judges presence or absence of the inputting of the diagnosis report corresponding to the medical image displayed on said monitor when the medical image displayed on said monitor is changed. 3. A medical image handling system according to claim 1, wherein the processor requests an inputting of the diagnosis report, in a case where the inputting of the image reading report is judged to be absent. 4. A medical image handling system according to claim 3, wherein said processor measures a time elapsing from the display of the medical image on said monitor and judges presence or absence of the inputting of the diagnosis report corresponding to the displayed medical image when the measured time exceeds a predetermined time. 5. A medical image handling system comprising: a) a monitor for displaying a medical image;b) a processor configured to process a control of judging presence or absence of the inputting of a diagnosis report corresponding to the displayed medical image, wherein the diagnosis report corresponding to the medical image displayed on said monitor includes information input by the user while viewing the medical image, to display an image for inputting a diagnosis report corresponding to the medical image displayed on said monitor in a case where the input of the diagnosis report is judged to be absent, and to input a diagnosis report that indicates absence of observation instead of a diagnosis report input by the input device, in a case where the inputting of the diagnosis report is judged to be absent and the medical image displayed on the monitor is changed, or in a case where a predetermined time has elapsed; andc) an input device for inputting a diagnosis report corresponding to the medical image displayed on said monitor on the basis of a user instruction. 6. A medical image handling method comprising: a) a step of displaying a medical image on a monitor;b) a step of judging presence or absence of an inputting of a diagnosis report corresponding to the displaye... | Canon Kabushiki Kaisha | G06K009/54 G06K009/60 H04N001/00 G06Q010/00 A61B005/00 G06F003/048 G06F003/00
US20110160563 | A1 | 2011-06-30 | DIAGNOSTIC SKIN MAPPING BY MRS, MRI AND OTHER METHODS | 2009-02-26 | This invention pertains to improved methods of classifying skin types as well as improved methods for determining the appropriateness of products and evaluating methods for treating particular skin. The methods typically utilize a “skin type” database containing one or more quantitative measures (e.g., NMR data) of skin properties. The database can optionally include various qualitative measures of skin as well (e.g., Glogau scale and/or Fitzpatrick scale values). | 1. A method of characterizing skin type in a human subject, said method comprising: performing nuclear magnetic resonance (NMR) measurements at a plurality of skin depths at one or more locations on said subject; andcalculating from data provided by said NMR measurements one or more skin-type values indicative of a skin characteristic. 2. The method of claim 1, wherein said skin-type values are indicative of a Glogau and/or a Fitzpatrick scale value. 3. The method of claim 1, wherein said NMR measurements are made using a portable NMR device. 4. The method of claim 1, wherein said calculating comprises determining a Fourier transform of an NMR signal depth profile. 5-7. (canceled) 8. The method of claim 1, wherein said calculating comprises determining an NMR signal amplitude as a function of skin depth at a skin location. 9. The method of claim 8, wherein said calculating comprises determining an NMR signal amplitude as a function of skin depth wherein said signal amplitude as a function of skin depth defines a step. 10. The method of claim 8, wherein said calculating comprises identifying a step depth (d0), optionally identifying a step height (Δf), and optionally identifying a step width (σ). 11. The method of claim 8, wherein said skin-type values are a function of a step depth (d0) and/or a step height (Δf), and/or a step width (σ). 12. The method of claim 8, wherein said calculating comprises determining a location of a transition between cutis and subcutis and/or comprises determining a thickness of cutis and/or subcutis. 13. (canceled) 14. The method of claim 1, wherein said NMR measurements comprise measurements of one or more parameters selected from the group consisting of: relaxation time T2 (true), relaxation time T2 (effective), relaxation time T1 (true), relaxation time T1 (effective), self-diffusion coefficient D, signal amplitude, spin modes, pulse sequence CPMG, dipolar encoded longitudinal magnetization, multiquantum build-up, multiquantum decay, diffusion coefficients, chemical shift resolved spectra, component amplitudes in chemical shift resolved spectra, and/or a mathematical function thereof. 15. The method of claim 1, further comprising outputting to a patient medical record said one or more NMR measurements and/or said one or more skin-type values. 16. The method of claim 1, further comprising outputting to a display or printer and/or storing to a computer readable medium said one or more NMR measurements and/or said one or more skin-type values. 17. The method of claim 15, wherein said patient medical records comprise one or more of the following: Glogau value for the same skin, Fitzpatrick value for the same skin, skin thickness, skin hardness, skin water content, freckles, scaling, subject identifier, subject age, subject ethnicity, subject gender, and location of skin measurement. 18. A method of identifying a skin type for a region of skin of a treatment subject, said method comprising: providing a sk... | - | A61B005/055
US20110160545 | A1 | 2011-06-30 | ELECTRONIC INVESTIGATIVE DEVICE AND METHOD FOR IDENTIFYING THE TRUTH | 2008-05-30 | The present invention provides a novel electronic investigative device for identification of truth from individuals who have committed an act of offence. The invention also provides a method for identification of truth based on bio-electric responses that are elicited by presentation of a unique design of nestled probes. The method advantageously utilizes the experiential knowledge present in a subject's brain that elicits a bio electric response to the presentation of probe. The invention is completely non-invasive and does not require an active participation of the subject. Further, the method is fully automated to enable tamper proof results. | 1. An electronic investigative device for finding the truth comprising of At least one bio signal conditioning unitSignal acquisition, Processing and Analysis UnitVisual and Auditory Presentation Control UnitAnalysis UnitPower Conditioning UnitOptionally Temperature and Humidity Control Units andMeans for Printing the output from the analysis unit 2. A electronic investigative device of claim 1, wherein said system comprises of at least one Bio signal conditioning unit that receives the bio electric information from the subject and reconditions the signals and wherein said reconditioning of the signal is achieved by means of an array of Filters and said signal is digitized for storage and analysis. 3. A electronic investigative device of claim 1, wherein said system comprises of signal acquisition and processing unit provided with means to receive information about the start and stop of the stimuli;store information about the start and stop of the stimuli denoting the start of an Event, or stop of an Event andhaving means for providing tamper proof neuro signatures. 4. A electronic investigative device of claim 1, wherein said signal processing unit comprises of video monitoring system that locks into the subject and captures video related information such as facial expression, eye ball movement and skin condition. 5. A electronic investigative device of claim 1, wherein the said system comprises of a Visual and Auditory Stimulus Presentation Unit (VASP) wherein the said unit is utilized for presenting visual and/or auditory stimulus to the subject. 6. A novel method of administering truth finding tests employing the device claimed in claim 1, wherein the said method comprises of Collation of informationValidation of the probesCreation of nestled probesPresentation of the probes andDetection of bio electric responses from brain. 7. A method of claim 6, wherein the said collation of information comprises of identification and categorization of at least one auditory and/or one visual event experienced by the subject and wherein the said collation of said at least one auditory and/or visual event constitutes a probe. 8. A probe of claim 7, wherein the said probe can be categorized into categories such as Neutral probes that comprises of at least one auditory stimulus of generic and/or semantic nature unconnected to the test;Control probes that comprises of auditory and/or visual stimulus related to the personal life events of the subject which are utilized to validate the findings of the target probes andTarget probes that comprises of auditory and/or visual stimulus that are responsible for detection of truthAnd collation of Neutral, Control and Target probe constitute a nestled probe. 9. A method of claim 6, wherein a plurality of the said nestled probe maybe presented to the subject for detection of neuro signatures. 10. A method of claim 6, wherein the said presentation of the said probes is achieved by the visual and auditory stimulus ... | - | A61B005/00
US07966053 | B2 | 2011-06-21 | Method for magnetic resonance spectroscopic imaging | 2007-11-26 | A method for performing spectroscopy using an interleaved readout for at least two species. A B0 field is applied. A first spatial-spectral (SPSP) position resolved spectroscopy sequence (PRESS) excitation with a sufficiently narrow band to excite a first species without exciting a second species is applied. A first readout that measures the first species is performed. A second SPSP PRESS excitation with a sufficiently narrow band to excite the second species without exciting the first species is applied. A second readout that measures the second species is performed. | 1. A method for performing spectroscopy using an interleaved readout for at least two species, comprising of: a) applying a B0 field;b) applying a first spatial-spectral (SPSP) position resolved spectroscopy sequence (PRESS) excitation with a sufficiently narrow band to excite a first species without exciting a second species;c) performing a first readout that measures the first species;d) applying a second SPSP PRESS excitation with a sufficiently narrow band to excite the second species without exciting the first species;e) performing a second readout that measures the second species; andf) generating a spectroscopic output for a portion of k-space from the first readout and the second readout, wherein steps b-e are performed a plurality of times so that the first SPSP PRESS excitation is interleaved with the second SPSP PRESS excitation. 2. The method of claim 1, wherein the spectroscopic output is a spectrum for the first species and the second species. 3. The method of claim 1, wherein the spectroscopic output is a spectroscopic image for the first species and the second species from the first and second readouts. 4. The method of claim 3, wherein the applying a first spatial-spectral (SPSP) position resolved spectroscopy sequence (PRESS) excitation excites a third species, wherein the first readout measures the third species and the spectroscopic image also shows the third species. 5. The method of claim 4, wherein the first SPSP PRESS excitation comprises: a SPSP 90° pulse;a first adiabatic SPSP 180° pulse; anda second adiabatic SPSP 180° pulse, wherein the SPSP 90° pulse and the first and second adiabatic SPSP 180° pulses each have a spectral profile centered on the resonant frequency for exciting the first and third species and a bandwidth sufficiently narrow to avoid exciting the second species; andwherein the second SPSP PRESS excitation comprises: a SPSP 90° pulse;a first adiabatic SPSP 180° pulse; anda second adiabatic SPSP 180° pulse, wherein the SPSP 90° pulse and the first and second adiabatic SPSP 180° pulses each have a spectral profile centered on the resonant frequency for exciting the second species and a bandwidth sufficiently narrow to avoid exciting the first and third species. 6. The method of claim 5, wherein each SPSP 90° pulse has a same spectral profile as the associated first and second adiabatic SPSP 180° pulses. 7. The method of claim 6, wherein the first species is choline containing compounds, the second species is N-acetyl aspartate, and the third species is creatine/phosphocreatine. 8. The method of claim 7, wherein the first and second adiabatic SPSP 180° pulses for the first excitation and the second excitation were designed by first creating an adiabatic sech/tanh pulse wherein the amplitude A(t) of the adiabatic sech/tanh pulse is A(t) =A0sech(βt)−T/2≦t ≦T/2and wherein the frequency modulation function Δω(t) of the adiabatic sech/tanh pulse is Δω(t)=−μβtanh(βt)−T/2≦t ≦T/2where β is a modulation angul... | The Board of Trustees of the Leland Stanford Junior University | A61B005/05 G01N024/00 G01V003/00
US07955258 | B2 | 2011-06-07 | External data processing device to interface with an ambulatory repeater and method thereof | 2007-09-17 | An external data processing device to interface with an ambulatory repeater and method thereof is presented. An external data processing device is interfaced in far field telemetric communication with an ambulatory repeater. The external data processing device receives sensitive information preencrypted prior to implant under a cryptographic key uniquely assigned to an implantable medical device over a secure connection from the ambulatory repeater. Physiological measures retrieved from the implantable medical device by the ambulatory repeater are received by the external data processing device over a non-secure connection. | 1. A method for retrieving information from an implantable medical device, comprising: activating an ambulatory repeater;securely obtaining a cryptographic key uniquely assigned to an implantable device from the implantable device to the ambulatory repeater;interrogating the implantable medical device with the ambulatory repeater to access patient health information, wherein the patient health information comprise sensitive information and physiological measures;securely receiving sensitive information from the implantable medical device with the ambulatory repeater, wherein the sensitive information is preencrypted under the cryptographic key;non-securely receiving physiological measures retrieved from the implantable medical device via the ambulatory repeater;interfacing a base repeater with the ambulatory repeater through far field telemetric communication;opening a secure connection between the ambulatory repeater and the base repeater;securely exchanging the sensitive information from the ambulatory repeater with the base repeater;closing the secure connection between the ambulatory repeater and the base repeater; andtransmitting the patient health information from the base repeater to a server. 2. The method according to claim 1, further comprising one of: providing a dedicated monitoring module matched to the implantable medical device;storing the physiological measures, the processor further comprising a processing module operative on the physiological measures; andproviding a clinical interrogation interface to a plurality of implantable medical devices. 3. The method according to claim 1, further comprising: issuing a data transfer trigger over the ambulatory repeater interface to the ambulatory repeater. 4. The method according to claim 3, wherein the data transfer trigger comprises a system-based data transfer trigger occurring in response to a device-specific circumstance. 5. The method according to claim 3, wherein the data transfer trigger comprises a health-based data transfer trigger occurring in response to at least one of a detected prescribed health status and a detected alert condition. 6. The method according to claim 1, further comprising: scheduling interrogations of the implantable medical device by the ambulatory repeater; anduploading the schedule to the ambulatory repeater. 7. The method according to claim 1, wherein the far field telemetry is selected from at least one of radio frequency telemetry and cellular telemetry. 8. The method according to claim 1, wherein the physiological measures comprises at least one of physiological and behavioral data. 9. The method according to claim 1, further comprising: opening a non-secure connection between the ambulatory repeater and the base repeater;exchanging the sensitive information from the ambulatory repeater with the base repeater; andclosing the non-secure connection between the ambulatory repeater and the base repeater. 10. The method according to claim 9, wher... | Cardiac Pacemakers, Inc. | A61B005/00
US20110111394 | A1 | 2011-05-12 | CTGF AS A BIOMARKER, THERAPEUTIC AND DIAGNOSTIC TARGET | 2007-10-04 | The invention provides CTGF which is associated with the cardiovascular diseases and hematological diseases. The invention also provides assays for the identification of compounds useful in the treatment or prevention of cardiovascular diseases and hematological diseases. The invention also features compounds which bind to and/or activate or inhibit the activity of CTGF as well as pharmaceutical compositions comprising such compounds. The invention also provides CTGF as a biomarker for diseases as cardiovascular diseases and hematological diseases | 1. A method for diagnosing a disease using CTGF as a biomarker, comprising: i) obtaining a biological sample from a mammal;ii) measuring the level of CTGF in the biological sample; andiii) comparing the level of CTGF in the biological sample with the level of CTGF in a control sample,wherein a disease is diagnosed based upon the deviation between the CTGF level of the biological sample in comparison to the control sample. 2. A method for guiding a therapy of a disease using CTGF as a biomarker, comprising: i) obtaining a baseline level of CTGF in biological sample from a diseased mammal;ii) administering to the diseased mammal a treatment for the disease;iii) obtaining one or more subsequent biological samples from the diseased mammal;iv) measuring the level of CTGF in the one or more subsequent biological samples;v) comparing the level of CTGF in the one or more subsequent biological samples with the baseline sample; andvi) determining whether increased dosages, additional or alternative treatments are necessary based on CTGF levels obtained from one or more subsequent biological samples compared to the baseline CTGF level, thereby guiding a therapy of a disease. 3. The method of claim 1, wherein the biological sample is selected from a group consisting of a blood sample, a plasma sample, a serum sample, a tissue sample, a oral mucosa sample, a saliva sample, an interstitial fluid sample and a urine sample. 4. The method of claim 1, wherein the level of CTGF is determined by determining the level of CTGF polynucleotide. 5. The method of claim 1, wherein the level of CTGF is determined by determining the level of CTGF polypeptide. 6. The method of claim 1, wherein the level of CTGF is determined by determining the level of CTGF activity. 7. The method of claim 1, wherein the disease is selected from the group consisting of cardiovascular disease, hematological disease, neurological disease, respiratory disease, gastroenterological disease, and urological disease. 8. (canceled) 9. The method of claim 1, wherein the level of CTGF of the biological sample is elevated compared to the control sample. 10. (canceled) 11. The method of claim 1, further comprising the use of one or more biomarkers selected from the group consisting of CRTAC, PRSS23, FN1, LTBP2, TGFB2, and NPR3, BNP, ANP, Troponin, CRP, Myoglobin, CK-MB, and metabolites. 12. (canceled) 13. The method of claim 1, wherein measuring the level of CTGF is combined with the use of one or more clinical biomarkers selected from the group consisting of blood pressure, heart rate, pulmonary artery pressure, or systemic vascular resistance. 14. The method of claim 1, wherein measuring the level of CTGF involves the use of one or more diagnostic imaging methods selected from the group consisting of PET (Positron Emission Tomography), CT (Computed Tomography), ultrasonic, SPECT (Single Photon Emission Computed Tomography), Echocardiography, or Impedance Cardiography. 15-16. (canceled) 17... | Bayer Schering Pharma Aktiengesellschaft | C12Q001/68 C12Q001/00 G01T001/164 A61B006/03 G01N029/04 A61B005/05
US20110144460 | A1 | 2011-06-16 | PULSE FREQUENCY MEASURING METHOD AND APPARATUS | 2010-12-14 | A method and an apparatus are provided for measuring a pulse frequency in a bio-signal measurement device. A bio-signal collected by a sensor is applied as an input signal of a notch filter. A filter coefficient of the notch filter is adaptively changed according to a result of tracking the bio-signal in the notch filter and calculating a pulse frequency corresponding to the filter coefficient of the notch filter. | 1. A method for measuring a pulse frequency in a bio-signal measurement device, the method comprising the steps of: applying a bio-signal collected by a sensor as an input signal of a notch filter; andadaptively changing a filter coefficient of the notch filter according to a result of tracking the bio-signal in the notch filter and calculating a pulse frequency corresponding to the filter coefficient of the notch filter. 2. The method of claim 1, wherein calculating a pulse frequency comprises: determining a degeneration period of the bio-signal by using the input signal and an output signal of the notch filter;determining an impulse noise period in which an impulse noise signal is introduced, by measuring a power of the bio-signal tracked by the notch filter;delaying a tracking speed of the notch filter in the degeneration period and the impulse noise period; andcalculating the pulse frequency corresponding to the filter coefficient of the notch filter. 3. The method of claim 2, wherein determining a degeneration period of the bio-signal comprises determining a corresponding period as a degeneration period when a difference between a power of the input signal and a power of the output signal is less than a predetermined power reference value and continuously maintaining a filter coefficient from before the degeneration period during the degeneration period. 4. The method of claim 3, wherein the filter coefficient of the notch filter detected in a normal period of the bio-signal tracked by the notch filter is stored and updated during a predetermined period. 5. The method of claim 2, wherein the impulse noise period is determined by comparing two signals that are envelope-estimated by applying different attack time constants to the power of the bio-signal tracked by the notch filter. 6. The method of claim 5, wherein determining an impulse noise period comprises: detecting a first envelope envelope-estimated by applying a first attack time constant to the power of the bio-signal tracked by the notch filter;detecting a second envelope envelope-estimated by applying a second attack time constant, which is slower than the first attack time constant, to the power of the bio-signal tracked by the notch filter; anddetermining the impulse noise period, when a ratio of the second envelope to the first envelope is less than an envelope reference value. 7. The method of claim 1, wherein, before the bio-signal collected by the sensor is applied to the notch filter, motion artifacts are removed from the bio-signal by a High Pass Filter (HPF). 8. The method of claims 1, wherein the bio-signal is one of an ElectroCardioGram (ECG) and a PhotoPlethysmoGraphy (PPG). 9. An apparatus for measuring a pulse frequency in a bio-signal measurement system, the apparatus comprising: a bio-signal processor for adaptively changing a filter coefficient of a notch filter according to a result of tracking a bio-signal in the notch filter when the bio-signal collected ... | Samsung Electronics Co., Ltd. Industry-Academic Cooperation Foundation Yonsei University | A61B005/1455 A61B005/0452 G06F019/00
US20110105177 | A1 | 2011-05-05 | Cell/blood-monitor combination cellphone/blood glucose monitor | 2010-12-09 | The external body of the invention is six and ¾ in length when flipped open and 3 inches when flipped close and 1¾ wide, all components parts including the external and internal will snap into place cut and design from harden plastic. | 1. The system claims when it is flipped open it reveals on the center top an open port for the placement of the strip tester where blood drops can be apply to begging the test of sugar in the blood. on FIG. 2 which shows the device flipped close from item 6 to 9, revels the name in short CBGM Cell phone Blood Glucose Monitor the dual dialog box and the up and down dialers; the number (10) item is electric port charger. 2. The system claims on the center top portion, an screen dialog dual display box is visible where the result of glucose test will be reveal in legible numbers, the screen also serves as visual reader for names and phones numbers of incoming and outgoing calls, and the time and date. 3. The system claims just below the hinge which serves to flip open or closing the device shows a four corner tactile to track names and phone numbers; just below that on figure (4) are number for dialing phone numbers follow by FIGURE (5) Which shows the up and down dialers that will serve to read the previous results of glucose reading saved in the memory of the system, the “Ok” bottom will serve as the stop dialer once comparison to current and previous reading is found. | - | A61B005/157 H04M001/00
US07953496 | B2 | 2011-05-31 | Implantable lead with isolated contact coupling | 2003-04-11 | An implantable lead for a medical device with an isolated contact connection for connecting a conductor to a contact reduces the opportunity for conductor material to migrate to a contact or into a patient. The implantable lead comprises a lead body having a proximal end and a distal end, at least one conductor, at least one contact carried on the proximal end, at least one contact carried on the distal end, at least one coupling. The lead has an exterior surface. The conductor is contained in the lead body and extends from the lead proximal end to the distal end. The conductor is electrically insulated. The contact carried on the proximal end is electrically connected to the conductor. The coupling has a conductor coupling and a contact coupling. The conductor coupling is placed over the conductor and attached to the conductor. The contact coupling exits the lead body and has a weld to connect the contact coupling to the contact. There is an isolation space created between the conductor and the contact to prevent the weld from containing conductor material. | 1. An implantable lead with an isolated contact connection, comprising: a lead body having a proximal and distal end;at least one conductor within the lead body traversing from the proximal end to the distal end;at least one contact carried on the proximal end and electrically connected to the conductor;at least one contact carried on the distal end and electrically connected to the conductor and having an inside surface and an outside surface and having a contact slot extending from the inside surface to the outside surface, the contact slot having a first and a second edge parallel to the conductor and a proximal or distal end which is open with no edge; and,at least one conductor coupling being partially placed over the conductor and being attached to the conductor; the conductor coupling being placed in the contact slot on the at least one contact, and the conductor coupling having a weld to connect the conductor coupling to the distal end contact,wherein the conductor coupling creates an isolation space between the conductor and the contact to prevent directly welding the conductor to the contact. 2. The implantable lead as in claim 1, wherein the isolation space separates the conductor from the weld to substantially prevent the conductor from contacting the weld. 3. The implantable lead as in claim 1, wherein the isolation space is configured to reduce communication between patient tissue and the conductor. 4. The implantable lead as in claim 1, wherein the isolation space reduces conductor ion migration outside of the lead body. 5. The implantable lead as in claim 4, wherein the conductor is braided strand wire with a silver core. 6. The implantable lead as in claim 1, wherein the conductor coupling is attached to the conductor by crimping. 7. The implantable lead as in claim 1, wherein the weld is performed substantially on at least one contact exterior surface for ease of manufacturing. 8. The implantable lead as in claim 1, wherein the at least one proximal contact comprises MP35N. 9. The implantable lead as in claim 8, wherein the at least one distal contact comprises a platinum-iridium alloy. 10. The implantable lead as in claim 1, wherein the contact slot extends from a proximal end to a distal end of the contact. 11. The implantable lead as in claim 6, wherein the conductor coupling comprises two side arms crimped inward to hold the conductor. 12. An implantable lead with an isolated connection, comprising: a lead body having a proximal end and a distal end;at least one conductor within the lead body traversing from the proximal end to the distal end;at least one electrode positioned along the length of the lead body and electrically connected to the conductor, the electrode comprising a hollow cylindrical tube with a first opening and a second opening and a slot extending from the first or second opening along the electrode substantially parallel with the lead body; and,a conductor coupling means for attaching the electr... | Medtronic, Inc. | A61N001/00 A61B005/04
US08317721 | B2 | 2012-11-27 | Reuseable skin testing device | 2010-10-14 | The present invention provides a device and method for skin testing. The device includes a handle assembly connected to a cylinder assembly having a first and second cylinder housing. Each cylinder housing is connected to at least one testing head assembly that includes a plurality of pins having a plurality of tips positionable against the skin and a socket assembly that is structured and operable to pivot the testing head to apply equal and even pressure to all the pins. Further, the method provides positioning the device tangent to the surface of the skin, and pivoting the testing head assembly horizontally to apply equal and even pressure to the pins. | 1. A device for performing skin testing, comprising: a handle assembly connected to a cylinder assembly having a first cylinder housing and a second cylinder housing, each cylinder housing connected to at least one testing head assembly, wherein the testing head assembly includes: a) a plurality of pins having a plurality of tips positionable against the skin;b) a socket assembly, wherein the socket assembly is structured and operable to pivot the testing head assembly to apply equal and even pressure to all the pins; and,wherein the handle assembly further includes a trigger mechanism, the trigger mechanism includes a biasing device connected to a cylinder bearing shaft of the cylinder assembly that is removably connectable to the testing head assembly. 2. The device according to claim 1, wherein the cylinder bearing shaft includes a pocket structured to receive a bearing of the cylinder assembly. 3. The device according to claim 2, wherein the trigger mechanism is structured and operable to compress the biasing device such that the cylinder bearing shaft receives the bearing in the pocket thereby disconnecting the testing head assembly from the cylinder housing. 4. A device for performing skin testing, comprising: a handle assembly connected to a cylinder assembly having a first cylinder housing and a second cylinder housing, each cylinder housing connected to at least one testing head assembly, wherein the testing head assembly includes: a) a plurality of pins having a plurality of tips positionable against the skin;b) a socket assembly, wherein the socket assembly is structured and operable to pivot the testing head assembly to apply equal and even pressure to all the pins; and,wherein the socket assembly further includes a release mechanism having a ball bearing connected to a spring, wherein the ball bearing is structured and operable to rest in a slot located in the testing head assembly. 5. The device according to claim 4, wherein the testing head assembly is structured and operable to pivot when the ball bearing depresses the spring. 6. The device according to claim 4, wherein the slot is coaxial to a longitudinal axis of the cylinder housing. 7. The device according to claim 4, wherein the release mechanism is structured and operable in a locked position when the ball bearing is positioned within the slot, and the release mechanism is structured and operable in an unlocked position when the ball bearing is positioned away from the slot. 8. The device according to claim 4, wherein the release mechanism is structured and operable to be unlocked when the cylinder housing applies pressure to the testing head assembly resulting in the ball bearing being positioned away from the slot. 9. The device according to claim 4, wherein the testing head assembly is structured and operable to pivot based on a pressure applied by the cylinder housing to the testing head assembly, resulting in the ball bearing compressing the spring, thereby allo... |  | A61B005/00
US08140149 | B2 | 2012-03-20 | Drowsiness detector | 2009-07-02 | A drowsiness detector capable of detecting light drowsiness of a subject with high precision is provided. The drowsiness detector acquires the measurement data of a measurement device first, and acquires a value of the pulse rate by performing preprocessing on the measurement data and extracts a pulse fluctuation from the value of the pulse rate. Then, the drowsiness detector sets the reference section width of the pulse feature value referred to in order to acquire the standard deviation of the pulse feature value (value of the pulse rate and pulse fluctuation) and then calculates the standard deviation value of the pulse feature value in the reference section width and acquires the corrected standard deviation value of the pulse feature value by dividing the standard deviation value of the pulse feature value by the pulse feature value. Then, the drowsiness detector determines whether or not the driver is lightly drowsy using the corrected standard deviation value of the pulse feature value. | 1. A drowsiness detector comprising: a measurement means for measuring a heartbeat or a pulse of a subject;a pulse feature value extracting means for extracting a pulse feature value from the heartbeat or the pulse measured by the measurement means;a variation distribution calculating means for calculating variation distribution of the pulse feature value extracted by the pulse feature value extracting means; anda drowsiness level determining means for determining a drowsiness level of the subject using the variation distribution of the pulse feature value calculated by the variation distribution calculating means,wherein the variation distribution calculating means has a means for calculating a standard deviation of the pulse feature value as the variation distribution of the pulse feature value and a means for correcting the standard deviation of the pulse feature value by dividing the standard deviation of the pulse feature value by the pulse feature value acquired in a state where there is no drowsiness, andthe drowsiness level determining means determines the drowsiness level of the subject on the basis of the standard deviation of the pulse feature value, andthe pulse feature value includes at least one of a pulse fluctuation low-frequency component relevant to the activity of the sympathetic nerve, a pulse fluctuation high-frequency component relevant to the activity of the parasympathetic nerve, and a ratio of the pulse fluctuation low-frequency component and the pulse fluctuation high-frequency component. 2. The drowsiness detector according to claim 1, further comprising: an average value calculating means for calculating an average value of the pulse feature value extracted by the pulse feature value extracting means,wherein the variation distribution calculating means calculates a standard deviation of the pulse feature value as the variation distribution of the pulse feature value, andthe drowsiness level determining means determines the drowsiness level of the subject on the basis of the standard deviation of the pulse feature value and the average value of the pulse feature value. 3. The drowsiness detector according to claim 2, further comprising: a reference time width setting means for setting a reference time width of the pulse feature value referred to in order to obtain the standard deviation of the pulse feature value,wherein the variation distribution calculating means calculates the standard deviation of the pulse feature value within the reference time width of the pulse feature value. 4. The drowsiness detector according to claim 3, wherein the reference time width setting means extracts a peak frequency by frequency analysis of the pulse feature value and sets a period corresponding to the peak frequency as the reference time width. 5. The drowsiness detector according to claim 1 which detects drowsiness of a driver of a vehicle, wherein a standard deviation calculating means for calculating a standard deviation of ... | Toyota Jidosha Kabushiki Kaisha | A61B005/0402
US09066693 | B2 | 2015-06-30 | Device for measuring the blood flow of a body tissue | 2009-07-07 | A device for measuring the blood flow of a body tissue comprises a catheter having a catheter head for the insertion into the inside of a body tissue and a center piece having a light emission surface, out of which an optical conductor leads, and having a reflection surface, which is disposed opposite of the light emission surface and oriented obliquely to the longitudinal axis of the optical conductor. The optical conductor is disposed such that an emitted light beam is directed at the reflection surface, the emitted light beam can be deflected at the reflection surface and reflected into the body tissue, and a reflected light beam can be reflected out of the body tissue at the reflection surface and fed into the optical conductor. The catheter head is divided into an insertion region and a connecting region, wherein the insertion region comprises a plurality of recesses on the surface thereof. In the direction of the connecting region, the insertion region has an increasing diameter. The recesses are provided in the insertion region such that webs are formed in the direction of the connecting region between the recesses along the surface of the catheter head. | 1. A device for measuring the flow of blood through a body tissue, comprising: a catheter with a catheter head for insertion into the interior of a body tissue;an optical conductor inside the catheter;a light source for emitting a beam of light into the body tissue by means of the optical conductor; anda processing unit for determining the rate of blood flow by means of a beam of light reflected from the body tissue; wherein the catheter has a middle piece with a recess comprising a light exit surface area, from which the optical conductor emerges, and a reflection surface area which lies opposite the light exit surface area and which is oriented at least partially obliquely with respect to the longitudinal axis of the optical conductor;the optical conductor being arranged in such a way that a beam of light emitted from the light source is directed to the reflection surface area, the emitted beam of light can be deflected on the reflection surface area and can be reflected into the surrounding body tissue; andwherein the distance between the catheter head and the middle piece is varied using connection elements of different lengths. 2. The device of claim 1 wherein the reflection surface area and the optical conductor are arranged in such a way that a beam of light reflected on the body tissue can be reflected on the reflection surface area and fed into the optical conductor. 3. The device of claim 1 wherein at least one photodetector is arranged in the middle piece and light reflected on the body tissue can be detected by the at least one photodetector. 4. The device of claim 3 wherein a circuit board is present in the recess and wherein the at least one photodetector is arranged on the circuit board. 5. The device of claim 1 wherein the reflection surface area is a mirror. 6. The device of claim 5 wherein the middle piece has an inclined surface on which a mirror is secured, wherein the mirror forms the light reflection surface area. 7. The device of claim 1 wherein the middle piece is rigid. 8. The device of claim 1 wherein the connection elements comprise a first connection tube and wherein the device further comprises a second connection tube, wherein the middle piece has a distal fastening piece for fastening to the first connection tube and a proximal fastening piece for fastening to the second connection tube. 9. The device of claim 8 wherein the catheter head is rigid, and wherein the catheter head has a third fastening piece for fastening to the first connection tube, such that the head part is arranged at a distance from the middle piece. 10. The device of claim 8 wherein at least the first connection tube is flexible. 11. The device of claim 1 wherein the middle piece is made of metal or plastic. 12. The device of claim 1 wherein a pressure sensor and/or a temperature sensor are arranged in the middle piece. 13. The device of claim 1 wherein the middle piece comprises a lower part and a lid, the lower part and the lid bein... | CARAG AG | A61B005/02 A61B005/1459 A61B005/026 A61B005/1455 A61M025/00
US20110112423 | A1 | 2011-05-12 | INCREASING CPR EFFECTIVENESS USING PHONOCARDIOGRAM ANALYSIS | 2010-07-09 | Embodiments of the invention include a system and methods for providing status information about resuscitation efforts of a person receiving chest compressions as part of Cardiopulmonary Resuscitation (CPR). A microphone generates a soundtrack by sampling sounds within or around the body of the person receiving CPR. The soundtrack is gated in one or more various ways to eliminate portions of the soundtrack, and analysis performed on the remaining portions. By evaluating the remaining portions of the soundtrack, the system can determine a cardiovascular effect of the compressions and provide status information to the rescuer about the determined cardiovascular effect. | 1. A system for providing status information about resuscitation efforts of a person receiving chest compressions as part of Cardiopulmonary Resuscitation (CPR), comprising: a microphone for sampling sounds within a body of the person to generate a soundtrack;a processor structured to determine a cardiovascular effect of the compressions from gated portions of the soundtrack but not from other portions, the gated portions and the other portions of the soundtrack determined by a gating signal that is presumed to be correlated with the chest compressions, the processor further structured to determine the status information from the cardiovascular effect; anda status unit coupled to the processor and structured to output the status information. 2. The system of claim 1, in which the cardiovascular effect is a heart valve closure of the person receiving chest compressions. 3. The system of claim 2, in which the processor is further structured to determine a time interval from a compression action of the compressions until the heart valve closure. 4. The system of claim 1, in which the cardiovascular effect is a Return of Spontaneous Circulation (ROSC) of the person receiving chest compressions. 5. The system of claim 1, in which the status information is a determination that the chest compressions are not effective for the person receiving chest compressions for a physiological reason. 6. The system of claim 1, in which the gating signal correlates with a compressing action of the compressions. 7. The system of claim 6, in which the gating signal correlates with a releasing action between two compressing actions. 8. The system of claim 1, in which the gating signal is one of an electrical signal and a software value. 9. The system of claim 1, in which the gating signal is generated based on the soundtrack. 10. The system of claim 1, in which the gating signal is received as issued by an automated chest compression machine. 11. The system of claim 1, further comprising: a signal generator for producing an artificial sound; anda gating signal generator for using the artificial sound as an input and structured to generate the gating signal in response thereto. 12. The system of claim 11, in which producing the artificial sound comprises injecting an artificial sound to the person's body. 13. The system of claim 11, in which the signal generator is coupled to a hand of a rescuer providing the chest compressions. 14. The system of claim 11, in which the artificial sound is substantially periodic. 15. The system of claim 11, in which the artificial sound is correlated with an action of a mechanical chest compression device that provides the chest compressions. 16. The system of claim 1, further comprising: a sensor for generating a compression input related to the chest compressions; anda gating signal generator coupled to the sensor and structured to generate the gating signal from the compression input. 17. The system of claim 16, in whic... | PHYSIO-CONTROL, INC. | A61B005/02
US07939664 | B2 | 2011-05-10 | Pyridinium boronic acid quenchers for use in analyte sensors | 2008-05-01 | Novel pyridinium salts functionalized with boronic acid and methods of making them are disclosed. When combined with a fluorescent dye, the compounds are useful in the detection of polyhydroxyl-substituted organic molecules. | 1. The reactive compound:wherein: X— is a counterion;X1 is —O— or —NH—;X2 is —O— or —NH—;L is a divalent linking group comprising a direct bond or a lower alkylene having 1 to 8 carbon atoms, optionally terminated with or interrupted by one or more divalent connecting groups selected from the group consisting of sulfonamide (—SO2NH—), amide —(C═O)N—, ester —(C═O)—O—, ether —O—, sulfide —S—, sulfone (—SO2—), phenylene —C6H4—, urethane —NH(C═O)—O—, urea —NH(C═O)NH—, thiourea —NH(C═S)—NH—, amide —(C═O)NH—, and amine —NR—, where R is defined as alkyl having 1 to 6 carbon atoms; andZ is either a reactive, ethylenically unsaturated group or a reactive functional group, capable of forming a covalent bond with a polymer or matrix. 2. The compound of claim 1, wherein Z is a reactive, ethylenically unsaturated group selected from the group consisting of methacrylamido-, acrylamido-, methacryloyl-, acryloyl- and styryl-. 3. The compound of claim 1, wherein Z is a reactive functional group capable of forming a covalent bond with a polymer or matrix selected from the group consisting of —Br, —OH, —SH, —CO2H, and —NH2. 4. The compound of claim 1, wherein Z iswherein R is H or CH3. 5. The reactive compound of claim 1, wherein said compound is: 6. A method of making the compound of claim 5, comprising the steps of: 7. The compound:wherein: Z is a reactive, ethylenically unsaturated group selected from the group consisting of methacrylamido-, acrylamido-, methacryloyl-, acryloyl- and styryl-, or wherein Z is a reactive functional group capable of forming a covalent bond with a polymer or matrix;Y is a trivalent connecting group selected from where R is H or a lower alkyl, andX1 and X2 are —O— or —NH—; andL1, L2, and L3 are selected from a direct bond or, a lower alkylene having 1 to 8 carbon atoms, optionally terminated with or interrupted by one or more divalent connecting groups selected from sulfonamide (—SO2NH—), amide —(C═O)N—, ester —(C═O)—O—, ether —O—, sulfide —S—, sulfone (—SO2—), phenylene —C6H4—, urethane —NH(C═O)—O—, urea —NH(C═O)NH—, thiourea —NH(C═S)—NH—, amide —(C═O)NH—, amine —NR—(where R is defined as alkyl having 1 to 6 carbon atoms) or combinations thereof. 8. The compound of claim 7, wherein Z is a polymerizable, ethylenically unsaturated group selected from the group consisting of methacrylamido-, acrylamido-, methacryloyl-, acryloyl- and styryl-. 9. The compound of claim 7, wherein Z is a reactive functional group capable of forming a covalent bond with a polymer or matrix selected from the group consisting of —Br, —OH, —SH, —CO2H, and —NH2. 10. The compound:wherein n is an integer from 1-10. 11. A method of making the compound of claim 10, comprising the steps of: 12. The compound:wherein:X− is a counterion;X1 is —O— or —NH—;X2 is —O— or —NH—;L is a divalent linking selected from a direct bond or, a lower alkylene having 1 to 8 carbon atoms, optionally terminated with or interrupted by one or more divalent connecting groups selec... | Glumetrics Inc. | C07F005/02 A61B005/145
US07959575 | B2 | 2011-06-14 | Blood rheology measuring apparatus | 2006-05-02 | A blood rheology measuring apparatus has a measuring portion for measuring a flow velocity of the blood flowing in a blood vessel of a person in a mode of a Doppler shift signal by transmitting and receiving a wave to and from a surface of the person's skin. An information processing portion calculates an intensity at each of frequency components of the Doppler shift signal, extracts a maximum frequency in a signal at an intensity level equal to or larger than a threshold in the histogram or a maximum frequency when an integrated value from a low frequency component reaches a predetermined rate of a total thereof in the histogram, and provides a temporal change waveform of the extracted frequency. The blood rheology is analyzed based on the maximum frequency in the frequency waveform. | 1. A blood rheology measuring apparatus comprising: a measuring portion for measuring a flow velocity of the blood flowing in a blood vessel of a person in a mode of a Doppler shift signal by transmitting and receiving a wave to and from a surface of the person's skin; andan information processing portion having means for calculating an intensity at each of frequency components of the Doppler shift signal, means for extracting a maximum frequency in a signal at an intensity level equal to or larger than a threshold in the histogram or a maximum frequency when an integrated value from a low frequency component reaches a predetermined rate of a total thereof in the histogram, and means for providing a temporal change waveform of the extracted frequency;wherein a blood rheology is analyzed based on the maximum frequency in the frequency waveform. 2. A blood rheology measuring apparatus according to claim 1; wherein the measuring portion comprises an ultrasonic wave transmitter and receiver for transmitting and receiving an ultrasonic wave to and from an artery at a tip portion of the person's finger. 3. A blood rheology measuring apparatus according to claim 1; wherein the measuring portion comprises an ultrasonic wave transmitter and receiver for transmitting and receiving an ultrasonic wave to and from an artery at the person's finger. 4. A blood rheology measuring apparatus according to claim 1; further comprising a temperature measuring portion for measuring a temperature of the blood vessel portion for detecting the blood flow; and wherein the information processing portion has means for operating and compensating for an amount of changing the flow velocity of the blood based on expansion and contraction of the blood vessel by a value of the temperature measured by the temperature measuring portion. 5. A blood rheology measuring apparatus according to claim 1; further comprising blood pressure measuring means for measuring a blood pressure of the person; wherein the information processing portion has means for operating and compensating for a change amount based on a blood pressure by dividing by a value of the blood pressure measured by the blood pressure measuring means. 6. A blood rheology measuring apparatus according to claim 1; further comprising a blood pressure measuring portion for measuring the person's maximum blood pressure; and wherein the information processing portion has means for operating and compensating for a change amount based on a blood pressure by dividing the maximum frequency in the frequency waveform by a value of the maximum blood pressure measured by the blood pressure measuring portion. 7. A blood rheology measuring apparatus according to claim 1; wherein the measuring portion and the information processing portion are integrally or individually made portable to thereby enable the information concerning the blood rheology to be obtained continuously or daily. 8. A blood rheology measuring apparatus according... | Seiko Instruments Inc. | A61B005/02 A61B008/00
US07951088 | B2 | 2011-05-31 | Blood-flow-occluding, temperature-sensing catheters with elastic covers | 2007-01-04 | A catheter assembly for sensing the temperature of an arterial wall or other body lumen includes a blood-flow-occluding feature to increase the accuracy the temperature measurements. In addition to the flow-occluding feature, the catheter includes a distal end with a temperature sensing structure and a proximal end including a manually operated expansion control. The temperature sensing structure includes one or more presentation elements in the form of a basket or braided structure having at least one temperature sensor supported thereon, each sensor being operative to generate an electrical signal indicative of temperature. In the preferred embodiment, the feature to occlude blood flow is an inflatable balloon disposed near the expandable basket or braid structure. The temperature sensors are preferably thermistors, and an elastic sleeve covering the expandable basket or braid structure is also preferably provided to further insulate the temperature sensors from the cooling effects of blood flow. | 1. A method of determining whether a region of the inner wall of a blood vessel exhibits an elevated temperature, comprising the steps of: providing a temperature-sensing catheter having a distal end with a gas-filled expansion device in the form of a basket or braided structure covered with a flexible, elastic material and at least one sensor operative to generate an electrical signal indicative of temperature, the catheter having a proximal end including a control operative to mechanically expand and contract the basket or braided structure without the use of inflation;inserting the catheter into the vessel such that the temperature sensor is in opposition to a point on the vessel wall to be measured;expanding the expansion device by operating the control, thereby reducing the pressure of the gas contained within the expansion device; andobtain a temperature measurement of the point on the vessel wall as a function of the electrical signal generated by the sensor. 2. The method of claim 1, wherein a vacuum is established in the expansion device covered with the flexible, elastic material when the device is expanded using the expansion control. 3. The method of claim 1, further including the steps of: obtaining a base temperature reading, Tflowing, with at least some blood flowing and a temperature, Toccluded with blood flow substantially occluded; anddetermining if the point on the vessel wall exhibits an elevated temperature using ΔT=Toccluded−Tflowing. 4. The method of claim 3, further including the step of assessing blood flow to determine the extent of occlusion. 5. The method of claim 4, wherein the step of assessing blood uses a contrast dye. 6. The method of claim 4, wherein the step of obtaining a temperature measurement of the point on the vessel wall, Toccluded, is not performed until the assessment indicated that the flow of blood has been substantially occluded. 7. The method of claim 1, wherein blood flow is substantially occluded using the expansion device. 8. The method of claim 1, further including the step of providing a separate blood-flow-occluding feature. 9. The method of claim 1, further including the step of waiting for a predetermined period of time before obtaining the temperature measurements. 10. The method of claim 1, comprising the steps of: providing a control wire disposed lengthwise within the catheter, the wire being attached to a distal portion of the basket or braided structure, such that pulling on the proximal end of the wire expands the basket or braided structure, and pushing on the proximal end of the wire contracts the basket or braided structure. 11. A method of determining whether a region of the inner wall of a blood vessel exhibits an elevated temperature, comprising the steps of: providing a temperature-sensing catheter having a distal end with a gas-filled expansion device covered with a flexible, elastic material and at least one sensor operative to generate an electrical signal indicati... | Accumed Systems, Inc. | A61B005/00
US08818484 | B2 | 2014-08-26 | Methods of predicting musculoskeletal disease | 2010-11-17 | This invention is directed to methods of predicting bone or joint disease in a subject. The invention is also directed to methods of determining the effect of a candidate agent on any subject's risk of developing bone or joint disease. | 1. A method for assessing bone condition in a subject by obtaining information regarding one or more bone parameters from an image of the subject in a computer system, the method comprising: (a) obtaining the image comprising image of a bone of the subject;(b) defining two or more regions of interest (ROIs) in the image;(c) analyzing a plurality of positions in the ROIs to determine one or more bone microarchitecture parameters and one or more bone macro-anatomy parameters;(d) combining the parameters into a numerical index; and(e) comparing the numerical index against a reference database. 2. The method of claim 1, wherein the ROIs are overlapping. 3. The method of claim 1, wherein the positions analyzed in the ROIs are at regular intervals in the image. 4. The method of claim 1, wherein the positions analyzed in the ROIs are at irregular intervals in the image. 5. The method of claim 1, wherein the positions analyzed are at regular intervals. 6. The method of claim 1, further including determining one or more biomechanical parameters of the bone of the subject. 7. The method of claim 1, wherein the image is two-dimensional. 8. The method of claim 7, wherein image is an x-ray image. 9. The method of claim 1, wherein the image is three-dimensional. 10. The method of claim 1, wherein the image is an electronic image. 11. The method of claim 1, wherein the subject is an osteoporosis subject. 12. A method of generating a map of one or more bone parameters, comprising: (a) obtaining information on bone parameters according to the method of claim 1; and(b) identifying regions of the image that exhibit similar parameter characteristics, thereby creating a parameter map of the image. 13. A method of predicting a fracture path in a subject, comprising: (a) generating two or more parameter maps according to the method of claim 12;(b) generating a composite parameter map from the two or more parameter maps of step (a); and(c) analyzing the composite parameter map to identify possible fracture paths. 14. A method of determining the risk of fracture in a subject comprising: (a) predicting a fracture path according to the method of claim 13; and(b) evaluating one or more selected bone parameters along the predicted fracture path, thereby estimating the risk of fracture. 15. A method of predicting the risk of fracture in a subject comprising: (a) generating a finite element model from one or more parameter maps obtained according to the method of claim 12;(b) applying simulated force vectors that would occur during a fracture incident to the finite element model generated in step (a); and(c) determining the minimum forces required for fracture to occur, thereby estimating the risk of fracture. | ImaTx, Inc. | A61B005/05 A61B008/08 A61B006/00 G06T005/00 G06T007/00
US08591427 | B2 | 2013-11-26 | Heart monitoring system or other system for measuring magnetic fields | 2010-11-09 | A system includes at least one first magnetic field sensor configured to measure first and second magnetic fields. The system also includes at least one second magnetic field sensor configured to measure the second magnetic field substantially without measuring the first magnetic field. The system further includes processing circuitry configured to perform signal cancellation to generate measurements of the first magnetic field and to generate an output based on the measurements of the first magnetic field. The sensors could represent magneto-electric sensors. The magneto-electric sensors could be configured to up-convert electrical signals associated with the first and/or second magnetic fields to a higher frequency. The processing circuitry could be configured to identify one or more problems associated with a patient's heart. | 1. A system comprising: at least one first magnetic field sensor configured to measure first and second magnetic fields;at least one second magnetic field sensor configured to measure the second magnetic field substantially without measuring the first magnetic field;wherein the first and second magnetic field sensors comprise magneto-electric sensors, each magneto-electric sensor comprising alternating layers of magnetostrictive material and piezo-electric material;wherein each magneto-electric sensor further comprises: a magnet configured to sensitize the sensor;an electromagnet configured to at least partially desensitize the sensor;wherein the electromagnet in each magneto-electric configured sensor is configured to receive a carrier signal;the at least one first magnetic field sensor is configured such that the carrier signal modulates the biasing magnetic field thereby emulating an up-converting heterodyne mixer, wherein the low frequency magnetic field signal is outputted from the sensor as coherent sidebands near the carrier frequency;the at least one second magnetic field sensor is configured such that the carrier signal modulates the biasing magnetic field thereby emulating an up-converting heterodyne mixer, wherein the low frequency magnetic field signal is outputted from the sensor as coherent sidebands near the carrier frequency;thereby repeatedly sensitizing and at least substantially desensitizing the at least one first and second magnetic field sensors; andprocessing circuitry configured to perform signal cancellation to generate measurements of the first magnetic field and to generate an output based on the measurements of the first magnetic field. 2. The system of claim 1, wherein the first and second magnetic field sensors comprise one of: magneto-resistive sensors, magneto-electric sensors, giant magnetic tunneling junctions, spin-valve sensors, cesium vapor sensors, flux gate sensors, optically-pumped sensors, and micro-electro-mechanical system (MEMS) sensors. 3. The system of claim 1, wherein: the system comprises multiple channels, each channel comprising one or more first magnetic field sensors; andthe system comprises a single second magnetic field sensor. 4. The system of claim 1, wherein the processing circuitry is configured to: generate a waveform associated with a patient's heart;identify T-waves in the waveform; anddetermine whether a difference between average amplitudes of alternating beat T-waves exceeds a threshold. 5. The system of claim 1, wherein: the processing circuitry is configured to generate and process a waveform associated with a patient's heart; andthe system further comprises a triggering unit configured to synchronize at least one of the processing circuitry and a display device with heartbeats of the patient. 6. An apparatus comprising: a signal cancellation unit configured to:receive signals from at least one first magnetic field sensor configured to measure first and second magnetic fields;... | National Semiconductor Corporation | A61B005/02
US08372105 | B2 | 2013-02-12 | Puncture instrument | 2008-08-22 | A puncture instrument includes a needle assembly, a hub passage for axially guiding a hub, a needle passage communicating with the hub passage and having a distal end opening for allowing the tip end of a needle to project instantaneously therethrough, two arms having first junctions connected to the hub and inclined rearward in a direction away from the hub in an initial state, two push rods connected to second junctions of the arms and extending rearward, a button for pushing the push rods forwardly, and fulcrum members disposed more closely to the distal end than the arms for abutment against portions of the arms when the arms are moved. When the push rods are pushed forwardly by the button, the arms push the needle assembly forwardly, thereafter have portions thereof brought into abutment against the fulcrum members, and turn to pull the needle assembly rearward. | 1. A puncture instrument comprising: a case;a needle assembly movably disposed in the case and including a needle and a hub holding the needle;a hub passage for axially guiding the hub, the hub passage being disposed in the case;a needle passage disposed in the case and communicating with the hub passage, the needle passage having a distal end opening for allowing a tip end of the needle to project therethrough;at least one arm having a first junction on an end thereof connected to the hub and a second junction on another end thereof, the arm being inclined from the first junction toward a proximal end of the case in an initial state;a push rod joined to the second junction and extending toward the proximal end;a pusher for pushing the push rod toward a distal end of the case; anda fulcrum member separate from the arm in the initial state and abutting against a portion of the arm when the arm moves;wherein when the push rod is pushed toward the distal end by the pusher, the arm pushes the needle assembly toward the distal end, brings the portion thereof into abutment with the fulcrum member, and turns about the fulcrum member, so that the arm is inclined from the first junction toward the distal end thereby to pull back the needle assembly toward the proximal end. 2. A puncture instrument according to claim 1, wherein the push rod comprises a leaf spring which is elastically deformable under force applied from the arm when the push rod is pushed toward the distal end by the pusher. 3. A puncture instrument according to claim 1, wherein the arm is inclined toward the proximal end in a direction away from the hub in the initial state, and when the push rod is pushed toward the distal end as the pusher is pushed, the arm turns about the fulcrum member and inclined toward the distal end in a direction away from the hub. 4. A puncture instrument according to claim 1, wherein at least one of the first junction and the second junction is narrower than the arm and is plastically deformable when the arm turns and changes its inclination direction. 5. A puncture instrument according to claim 1, wherein the hub includes a base on a proximal end portion thereof and having a flat surface parallel to a plane in which the arm turns, the first junction being rotatably connected to the base at a position which is spaced a predetermined distance from a longitudinal axis of the hub. 6. A puncture instrument according to claim 1, wherein the pusher includes a convex or concave engaging portion, and the case includes a first engageable portion for engaging the engaging portion when the pusher is in an initial position and a second engageable portion for engaging the engaging portion when the pusher is in a stroke end position. 7. A puncture instrument according to claim 6, wherein the pusher is movably disposed in an inner cavity formed in the case and has a proximal end projecting from the case by a distance ranging from 5 to 30 mm when the pusher is in an ini... | Terumo Kabushiki Kaisha | A61B017/14 A61B017/32
US08849366 | B2 | 2014-09-30 | Catheter system having an optical probe and method for the application of an optical probe in a catheter system | 2008-07-25 | A catheter system includes a flexible, elongated base body adapted to be applied to a vein central-venously and having a base distal end; a fiber-optic probe having a probe distal part; a fiber-optic lumen receiving the fiber-optic probe; and an attachment element configured to avoid a longitudinal displacement of the fiber-optic probe relative to the fiber-optic lumen and configured to detach so as to allow removal of the fiber-optic probe through the longitudinal displacement of the fiber-optic probe relative to the fiber-optic lumen. The attachment element has a connector piece firmly connected to the fiber-optic probe and a counter-piece firmly connected to the base body, wherein the connector piece is connectable to the counter-piece, the connector piece having a shaft piece adjustable lengthwise and disposed on the connector piece so as to vary a length of the connector piece. | 1. A catheter system comprising: a fibre-optic probe having a probe distal part;a flexible, elongated base body adapted to be applied to a vein central-venously and having a base distal end, the base body including a lumen that receives the fibre-optic probe;andan attachment element including a connector piece firmly connected to the fibre-optic probe and a counter-piece firmly connected to the base body, the connector piece being detachably connectable to the counter-piece so as to attach the fibre-optic probe with respect to the base body and provide removability of the fibre-optic probe from the base body by longitudinal displacement of the fibre-optic probe relative to the lumen when the connector piece and counter-piece are detached, the connector piece having a shaft piece that is adjustable in length so as to vary a length of the connector piece,wherein a distance between the distal end of the base body and the distal end of the fibre-optic probe is adjustable to a predetermined value in a connected state of the attachment element by the connection between the fibre-optic probe and the connector piece and the connection between the base body and the counter-piece. 2. The system as recited in claim 1, wherein the shaft piece includes a tube section displaceable in a direction of the fibre-optic lumen in an inner space of the connector piece. 3. The system as recited in claim 2, wherein the tube section includes a sealing lip disposed at a distal end of the tube section, and wherein the tube section is sealed off via the sealing lip at the inner space. 4. The system as recited in claim 1, wherein the shaft piece includes at least one length mark. 5. The system as recited in claim 1, wherein the connector piece includes a locking device configured to fix the shaft piece. 6. The system as recited in claim 1, wherein the fibre-optic probe includes a balloon disposed at the distal end. 7. The system as recited in claim 1, wherein the fibre-optic probe includes more than one optical fibre. 8. A method for an application of a fibre-optic probe in a catheter system comprising: introducing a fibre-optic probe into a connector piece that has a shaft piece with an adjustable length;setting the shaft piece to a predetermined length so as to accurately position a distal end of the fibre-optic probe with respect to the connector piece; andfixing the fibre-optic probe with reference to the connector piece. 9. The method as recited in claim 8, further comprising locking the shaft piece at the predetermined length. 10. The method as recited in claim 8, further comprising measuring an intensity of reflected optical radiation in the fibre-optic probe during a measuring of oxygen saturation using the optical-fibre probe and a connected evaluation apparatus, and checking a plausibility of measured intensity values. 11. The method as recited in claim 10, further comprising emitting a signal if the fibre-optic probe is correctly placed. | Up-Med GmbH | A61B005/1455 A61B005/145 A61B005/00 A61M025/00 A61B005/1459 A61M025/01 A61B005/027
US08532740 | B2 | 2013-09-10 | Measurement device and measurement method | 2009-07-07 | A measurement device and measurement method make it possible to eliminate the load generated by stopping on a measurement subject that is moved among multiple magnetic field generating devices. The measurement device includes a first external magnetic field generating device that generates a magnetic field of a set size, a second external magnetic field generating device that generates a magnetic field of a size that differs from that of the magnetic field of the first external magnetic field generating device, a rotating table that causes the subject of measurement to pass in sequence through the magnetic fields of the first and second external magnetic field generating devices by causing the subject of measurement to move rotationally, and an OMRI measurement processing part and MRI measurement processing part that measures images such as functional images or structural images of the subject of measurement while it is being moved rotationally by the rotating table. | 1. A measurement device, comprising: a first magnetic field generating device for generating a magnetic field of a set size;a second magnetic field generating device for generating a magnetic field of a size that differs from that of the magnetic field of the first magnetic field generating device;a rotating movement device for causing a measurement subject to pass in sequence through the magnetic fields of the first and second magnetic field generating devices by causing the measurement subject or the first and second magnetic field generating devices to move rotationally without stopping; anda measurement processing part that measures images of the measurement subject in different magnetic fields without stopping the measurement subject or the magnetic field generating devices while the measurement subject is or the magnetic field generating devices are being moved rotationally by the rotating movement device. 2. The measurement device according to claim 1, wherein one of the first and second magnetic field generating devices is for exciting and measuring nuclear magnetic resonance, and the other is for exciting and measuring electron spin resonance. 3. The measurement device according to claim 1, wherein the second magnetic field generating device generates a magnetic field larger than the magnetic field of the first magnetic field generating device. 4. The measurement device according to claim 1, wherein the first magnetic field generating device generates a magnetic field larger than the magnetic field of the second magnetic field generating device. 5. The measurement device according to claim 1, wherein the first or second magnetic field generating device is for exciting and measuring magnetic resonance. 6. A measurement method, comprising: causing a measurement subject to pass in sequence through the magnetic fields of a first magnetic field generating device that generates a magnetic field of a set size and a second magnetic field generating device that generates a magnetic field of a size that differs from that of the magnetic field of the first magnetic field generating device by causing the measurement subject or the first and second magnetic field generating devices to move rotationally without stopping; andmeasuring images of the measurement subject in different magnetic fields without stopping the measurement subject or the first and second magnetic field generating devices while the measurement subject is or the first and second magnetic field generating devices being moved rotationally by the rotating movement device. 7. The measurement method according to claim 6, further comprising: exciting and measuring nuclear magnetic resonance by either the first or the second magnetic field generating device; andexciting and measuring electron spin resonance by the other of the first or the second magnetic field generating device. 8. The measurement method according to claim 6, further comprising generating a magnetic field by the s... | Kyushi University | A61B005/05
US08669113 | B2 | 2014-03-11 | Advanced detection of sepsis | 2009-04-02 | The present invention relates to methods, monitors and systems measuring lysophosphatidylcholine, its derivatives and/or procalcitonin as well as at least one clinical marker (e.g. temperature or respiration rate) and/or at least one biomarker for the early detection of sepsis in a subject. | 1. A method for the advanced detection of sepsis in a SIRS-positive subject, comprising the steps of: (a) measuring at a plurality of time points, prior to laboratory confirmation of a clinically significant infection causative of sepsis, an amount of lysophosphatidylcholine in fluid or tissue of the SIRS-positive subject;(b) measuring at a plurality of time points, prior to laboratory confirmation of a clinically significant infection causative of sepsis, two or more clinical markers of the SIRS-positive subject selected from the group consisting of respiratory rate, temperature, heart rate, systolic blood pressure, diastolic blood pressure, mean artery pressure, white blood cell count, monocyte count, lymphocyte count, granulocyte count, neutrophil count, immature neutrophil to total neutrophil ratio, platelet count, serum creatinine concentration, urea concentration, lactate concentration, glucose concentration, base excess, pO2 and HCO3− concentration; provided respiratory rate and temperature are measured; and(c) measuring at a plurality of time points, prior to laboratory confirmation of a clinically significant infection causative of sepsis, an amount of one or more biomarkers in fluid or tissue of the SIRS-positive subject selected from the group consisting of ACTH, activated partial thromboplastin, albumin, antithrombin III, bacterial DNA, carcinoembryonic antigen-related cell adhesion molecule 1 (CEACAM1), cell surface proteins CD-14 and CD-64, copeptin, C-reactive protein, cortisol, D-dimer, E-selectin, extra-hepatic arginase (ARG2), endotoxin, fibrin degrading products, high mobility group box 1 (HMGB1), HLA-DRA, interleukins IL-6, IL-8 and IL-10, LBP LPS-binding protein, plasminogen activator inhibitor-1, procalcitonin, protein C, protein S, prothrombin, soluble thrombomodulin, surface-bound tumor necrosis factor receptor I (sTNF-RI), surface-bound tumor necrosis factor receptor II (sTNF-RII), thrombin activatable fibrinolysis inhibitor, TRAF-interacting protein with a forkhead-associated domain (TIFA), triggering receptor expressed on myeloid cells 1, and tumor necrosis factor alpha (TNFα);wherein (i) a decrease in the amount lysophosphatidylcholine measured over time, such that a second amount is less than 75% of a previous amount over a 24 hour interval; (ii) a body temperature greater than 38° C. or less than 36° C.; (iii) a respiratory rate greater than 20 breaths per minute; and (iv) a difference in the amount of one or more biomarkers from a previous amount over a 24 hour interval,detect sepsis in the subject. 2. The method of claim 1, wherein said lysophosphatidylcholine comprises 1-O-palmitoyl-2-lyso-sn-glycero-3-phosphocholine and 1-O-stearoyl-2-lyso-sn-glycero-3-phosphocholine. 3. The method of claim 1, wherein said advanced detection is at least 12 hours prior to the laboratory confirmation of a clinically significant infection causative of sepsis. 4. The method of claim 1, wherein free lysophosphatidylcholine is measu... | Becton, Dickinson and Company | G01N033/00
US08814800 | B2 | 2014-08-26 | Apparatus and method for enhancing and analyzing signals from a continuous non-invasive blood pressure device | 2010-10-29 | A system and method of enhancing a blood pressure signal is disclosed. The volume of an artery in a finger is measured by a photo-plesthysmographic (PPG) system, which produces a PPG signal. This PPG system is placed inside a cuff, and the cuff pressure is controlled by the PPG signal. The portion or component of the PPG signal having a frequency higher than a predefined threshold frequency is then modified or enhanced, such as by multiplying the high frequency component by a calibration factor. A blood pressure signal is then calculated using the cuff pressure and the modified PPG signal. A blood pressure contour curve may then be generated, and a variety of parameters may be calculated using the curve. | 1. A method for determining a blood pressure contour curve comprising: placing a cuff having a photo-plethysmographic (PPG) system over an artery in a human finger, the PPG system producing a PPG signal based on volume of the artery, the PPG system including at least one light source and at least one light detector;modifying, by a computing device, a component of the PPG signal having a frequency higher than a predefined threshold frequency by eliminating from the PPG signal an undesired portion of the PPG signal and reconstructing the PPG signal from the remaining portion of the PPG signal;the cuff pressure being controlled by one or more components of the modified PPG signal; andcalculating, by the computing device, a blood pressure signal using at least the modified component of the PPG signal. 2. The method of claim 1 wherein a new blood pressure is calculated using the cuff pressure and the modified PPG signal. 3. The method of claim 2 wherein modifying a component of the PPG signal having a frequency higher than a predefined threshold frequency further comprises: separating the PPG signal into a first component having a frequency higher than the predefined threshold frequency and a second component having a frequency lower than the predefined threshold frequency;modifying the first component; andadding the modified first component to the second component to create a modified PPG signal;using the modified PPG signal and the cuff pressure to calculate the blood pressure signal. 4. The method of claim 2 wherein the modification further includes calibrating a component of the blood pressure signal having a frequency higher than a predefined threshold frequency using a value obtained for blood pressure by a sphygmomanometer placed on an artery in a human upper arm. 5. The method of claim 2 wherein the modification further includes multiplying a component of the blood pressure signal having a frequency higher than a predefined threshold frequency by a calibration factor, the calibration factor being calculated from a blood pressure measurement from a sphygmomanometer placed on an artery in a human upper arm. 6. The method according to claim 2, wherein the threshold frequency is about 0.3 Hz. 7. The method according to claim 2, wherein the calculation uses anthropometric parameters. 8. The method according to claim 2, further comprising calculating physiological parameters from the blood pressure contour curve. 9. The method according to claim 8, wherein the parameters are calculated by using multiport algorithms. 10. The method according to claim 8, wherein the parameters are calculated by using one or more markers of input signals. 11. The method according to claim 1, wherein the reconstruction is calculated from the pulsatile part of the remaining portion of the PPG signal. 12. The method according to claim 11, wherein the reconstructed PPG signal ispTn=pT⁢⁢0+gI·∑0n⁢Pn+gP·Pn⁢⁢and⁢⁢Pn=∫tn-1t⁢vAC⁡(t)⁢ⅆt. 13. The method according to cl... | CNSystems Medizintechnik AG | A61B005/02 A61B005/021 A61B005/024 A61B005/0225 A61B005/022 A61B005/025 A61B005/00
US08847769 | B2 | 2014-09-30 | Signal detecting device for detecting a difference signal for an electrical measurement of a vital parameter of a living being | 2010-10-27 | What is described are a signal detecting device for detecting a vital parameter of a living being, having a plurality of electrodes and an electrode selecting device for selecting a pair of electrodes from the plurality of electrodes such that a difference signal suitable for measuring the vital parameter is derivable, an electrode arrangement and a method. | 1. A signal detecting device for determining a difference signal for an electrical measurement of a vital parameter of a living being, comprising: a plurality of electrodes configured to contact the living being and provide electrode signals; andan electrode selecting device configured to select a pair of electrodes for determining the difference signal from the plurality of electrodes based on the electrode signals so that a difference signal suitable for measuring the vital parameter is derivable from the electrode signals of the selected electrodes;wherein the electrode selecting device is configured to: apply a reference signal to a first electrode from the plurality of electrodes,to detect an electrode signal at a second electrode from the plurality of electrodes,to decide, based on a signal strength of the electrode signal at the second electrode, whether the first electrode and the second electrode contact the living being sufficiently, andto select, responsive to a decision that the first electrode and the second electrode contact the living being sufficiently well, the first electrode and the second electrode as the pair of electrodes for determining the difference signal, or to mark the first electrode and the second electrode as candidate electrodes for the selection of the pair of electrodes for determining the difference signal. 2. The signal detecting device in accordance with claim 1, wherein the electrode selecting device is configured to detect signals provided by individual electrodes from the plurality of electrodes and compare these signals to a reference signal or several reference signals so as to establish electrodes which are in electrically or capacitively coupled contact to the living being and enable same for a measurement of the difference signal. 3. The signal detecting device in accordance with claim 1, wherein the electrode selecting device is configured to compare, when selecting a pair of electrodes for measuring the vital parameter, a first difference signal of a first pair of electrodes from the plurality of electrodes to a second difference signal of a second pair of electrodes from the plurality of electrodes, andto select that pair of electrodes for determining the difference signal for the electrical measurement of the vital parameter, comprising, when selecting, the greatest difference signal. 4. The signal detecting device in accordance with claim 1, wherein the electrode selecting device is configured to apply a reference signal to a first electrode from the plurality of electrodes so as to detect an electrode signal at a second electrode from the plurality of electrodes, andto decide, based on a signal strength of the electrode signal at the second electrode, whether the first electrode and the second electrode contact the living being sufficiently, andto, responsive to a decision that the first electrode and the second electrode contact the living being sufficiently well, mark the first electrode and... | Fraunhofer-Gesellschaft zur Foerderung der Angewandten Forschung E.V. | G08B023/00 A61B005/18 A61B005/0428 A61B005/04 A61B005/0245
US08548558 | B2 | 2013-10-01 | Electrode capable of attachment to a garment, system, and methods of manufacturing | 2011-01-13 | An electrode for selective attachment to at least one of a garment and a subject, is provided and includes a conductive member defining a first side and a second side; a conductive composition disposed on the first side of the conductive member, wherein the conductive composition has a first adhesive strength; a contact layer disposed on the second side of the conductive member, wherein the contact layer includes a pressure sensitive adhesive portion and a conductive hydrogel portion, wherein the pressure sensitive adhesive portion has a second adhesive strength that is greater than the adhesive strength of the conductive composition; whereby the electrode is adherable to the garment and to the subject such that the conductive composition is adhered to the subject and the pressure sensitive adhesive portion is adhered to the garment, wherein removal of the garment from the subject results in removal of the electrode from the subject. | 1. An electrode for selective attachment to at least one of a garment and a subject, the electrode comprising: a conductive member defining a first side and a second side;a conductive composition disposed on the first side of the conductive member, wherein the conductive composition has a first adhesive strength; anda contact layer disposed on the second side of the conductive member, wherein the contact layer includes a first and a second pressure sensitive adhesive portions and a conductive hydrogel portion, wherein each of the first and second pressure sensitive adhesive portions has a second adhesive strength that is greater than the first adhesive strength of the conductive composition;wherein the electrode is configured to be disposed between the garment and the subject in a wireless and stud-free manner such that the conductive composition is adhered to the subject and the first and second pressure sensitive adhesive portions are adhered to the garment, wherein removal of the garment from the subject results in removal of the electrode from the subject; andwherein the hydrogel portion separates the first pressure sensitive adhesive portion from the second pressure sensitive adhesive portion such that the first pressure sensitive adhesive portion abuts one end of the hydrogel portion and the second pressure sensitive adhesive portion abuts another end of the hydrogel portion. 2. The electrode according to claim 1, wherein the conductive composition is a hydrogel. 3. The electrode according to claim 1, wherein the hydrogel portion of the contact layer has an adhesive strength that is less than the second adhesive strength of the first and second pressure sensitive adhesive portions. 4. The electrode according to claim 1, wherein the hydrogel portion of the contact layer extends at least partially across a length of the electrode. 5. The electrode according to claim 1, wherein the hydrogel portion of the contact layer extends completely across a length of the electrode. 6. The electrode according to claim 1, further comprising a layer of a silver composition interposed between the conductive member and the conductive composition. 7. The electrode according to claim 6, wherein the layer of silver composition is a layer of Silver/Silver-Chloride (Ag/AgCl). 8. The electrode according to claim 7, wherein the layer of Silver/Silver-Chloride (Ag/AgCl) is in the form of a coating of Silver/Silver-Chloride (Ag/AgCl) ink. 9. An electrode, comprising: a first layer adapted for selective disposition on a garment; a second layer adapted for selective disposition on a subject; and a third layer adapted for disposition and electrical communication between the first layer and the second layer; wherein the first layer includes a hydrogel portion, a first pressure sensitive adhesive portion, and a second pressure sensitive adhesive portion;wherein the electrode is configured to be disposed between the garment and the subject in a wire-less and stud-fr... | Covidien LP | A61B005/04 A61N001/00
US08594774 | B2 | 2013-11-26 | Sleep determination device and sleep determination method | 2009-08-27 | An object of the present invention is to provide a sleep determination device and a sleep determination method capable of precisely determining a sleep stage based on heartbeat information of a subject. The sleep determination device and the sleep determination method according to the present invention obtain the heartbeat of the subject, calculate a heartbeat cycle for each heartbeat based on the obtained heartbeat, calculate difference values until obtaining Nth-degree difference values set in advance with difference values between the successive heartbeat cycles as first-degree difference values and difference values between the successive first-degree difference values as second-degree difference values, generate a return map and a histogram indicative of a distribution of the respective values in a predetermined period of time for each of the heartbeat cycles and the difference values from the first-degree difference values to the Nth-degree difference values, and determine a sleep stage of the subject based on the generated return map and the histogram by referring to a return map and a histogram for determination set in advance for each sleep stage. | 1. A sleep determination device comprising: a heartbeat obtaining unit that obtains a heartbeat of a subject;a difference calculating unit that calculates a heartbeat cycle for each heartbeat based on the heartbeat obtained by the heartbeat obtaining unit, and calculates difference values until obtaining Nth-degree difference values set in advance with difference values between the successive heartbeat cycles as first-degree difference values and difference values between the successive first-degree difference values as second-degree difference values;a distribution data generating unit that generates distribution data indicative of a distribution of the respective values in a predetermined period of time for each of the heartbeat cycles and the difference values from the first-degree difference values to the Nth-degree difference values calculated by the difference calculating unit; anda sleep determining unit that determines a sleep stage of the subject based on the distribution data generated by the distribution data generating unit by referring to distribution data for determination set in advance for each sleep stage. 2. The sleep determination device according to claim 1, wherein the distribution data generating unit generates a return map that uses as parameters successive values of the heartbeat cycles and the difference values in a predetermined period of time for each of the heartbeat cycles and the difference values from the first-degree difference values to the Nth-degree difference values calculated by the difference calculating unit; andthe sleep determining unit compares the return map generated by the distribution data generating unit with a return map for determination set in advance for each sleep stage to determine a sleep stage of the subject. 3. The sleep determination device according to claim 2, wherein the sleep determining unit sets a return map for discrimination in advance in which the degree of separation is high in neighboring sleep stages and compares the return map generated by the distribution data generating unit with the return map for discrimination to determine a sleep stage of the subject. 4. The sleep determination device according to claim 1, wherein the distribution data generating unit generates a histogram for each of the heartbeat cycles and the first-degree difference values to the Nth-degree difference values calculated by the difference calculating unit, andthe sleep determining unit compares the histogram generated by the distribution data generating unit with a histogram for determination set in advance for each sleep stage to determine a sleep stage of the subject. 5. The sleep determination device according to claim 4, wherein the sleep determining unit sets a histogram for discrimination in advance in which the degree of separation is high in neighboring sleep stages and compares the histogram generated by the distribution data generating unit with the histogram for discrimination to determin... | Toyota Jidosha Kabushiki Kaisha | A61B005/04
US09119573 | B2 | 2015-09-01 | Stent marker detection using a learning based classifier in medical imaging | 2010-12-06 | Stent marker detection is automatically performed. Stent markers in fluoroscopic images or other markers in other types of imaging are detected using a machine-learnt classifier. Hierarchal classification may be used, such as detecting individual markers with one classifier and then detecting groups of markers (e.g., a pair) with a joint classifier. The detection may be performed in a single image and without user indication of a location. | 1. A method for marker detection in fluoroscopic imaging of a stent, the method comprising: receiving, from a fluoroscopic scanner, a fluoroscopic scan of a patient;masking the fluoroscopic scan of the patient to limit the number of locations to which marker detection is applied;individually detecting a plurality of possible markers for a stent in one frame of fluoroscopic data from the fluoroscopic scan without information from other frames of fluoroscopic data from the fluoroscopic scan, the individual detecting being with a first machine-trained classifier, wherein the first machine-trained classifier does not detect masked out locations;jointly detecting a pair of markers from the plurality of possible markers for the stent output from the individually detecting, the joint detecting being with a second machine-trained classifier, wherein the second machine-trained classifier does not detect masked out locations; andoutputting an image representing information about the pair of markers. 2. The method of claim 1 wherein individually detecting comprises detecting with the first machine-trained classifier comprising a probabilistic boosting tree classifier. 3. The method of claim 1 wherein individually detecting comprises detecting with first input features comprising Haar features and wherein jointly detecting comprises detecting with second input features comprising gradient features. 4. The method of claim 1 wherein individually detecting further comprises removing some of the possible markers with a third machine-trained classifier, the third machine-trained classifier operating on an output of the first machine-trained classifier and the second machine-trained classifier operating on an output of the third machine-trained classifier. 5. The method of claim 1 wherein jointly detecting comprises fitting a rectangular box defining a region over the pair of markers and detecting as a function of an input feature set comprising a gradient magnitude, a gradient angle, or both relative to the rectangular box. 6. The method of claim 1 wherein jointly detecting comprises detecting with the second machine-trained classifier comprising a probabilistic boosting tree classifier. 7. The method of claim 1 wherein individually detecting comprises determining a first score for each of the possible markers, and wherein jointly detecting comprises determining a second score for each of paired combinations of the possible markers; further comprising calculating a third score for each of the paired combinations, the third score being a function of the first and second scores associated with the possible markers in the paired combination, the pair of markers being the paired combination with a highest score. 8. The method of claim 1 further comprising: steerably filtering the frame of fluoroscopic data such that guide wire regions are distinguished from background; andwherein individually detecting comprises sampling locations along the guide wire regions.... | Siemens Aktiengesellschaft | A61B006/12 A61B006/00
US08708901 | B2 | 2014-04-29 | Health monitoring system with a waveguide to guide a wave from a power source | 2009-12-30 | A health monitoring system includes a waveguide that receives a wave transmitted from an external power source and that guides the wave to reach within a width of a rectenna. The waveguide may include a negative refractive index medium and/or a surface plasmon medium. | 1. A health monitoring system comprising: a detection device configured to detect a health status of a creature;a rectenna;a waveguide configured to receive a wave transmitted from an external power source and to guide the wave to reach within a width of the rectenna, the waveguide having a width greater than the width of the rectenna, the external power source positioned on a first side of the waveguide;wherein portions of the wave in the waveguide that are outside of the width of the rectenna are guided to reach the rectenna,wherein the rectenna is positioned on a second side of the waveguide and configured to receive and to rectify the guided wave to generate an electric power to drive the detection device; anda transmitter configured to transmit the detected health status to an external device. 2. The health monitoring system of claim 1, wherein the detection device comprises at least one of a chemosensor and a biosensor to sense the health status. 3. The health monitoring system of claim 1, wherein the health monitoring system is configured to be inserted into a body of the creature or configured to be attached on an outer surface of the creature. 4. The health monitoring system of claim 1, wherein the external device comprises a monitor to display the detected health status. 5. The health monitoring system of claim 1, wherein the waveguide includes a negative refractive index medium to refract the wave to reach the width of the rectenna. 6. The health monitoring system of claim 5, wherein the negative refractive index medium comprises a metamaterial. 7. The health monitoring system of claim 6, wherein the metamaterial comprises a two-dimensional array of repeated unit cells of copper strips and split ring resonators on interlocking strips of standard circuit board material. 8. The health monitoring system of claim 6, wherein the metamaterial comprises a quantum metamaterial having a manipulated chemical composition and/or a structure of a material. 9. The health monitoring system of claim 1, wherein the waveguide comprises a surface plasmon medium to generate a surface plasmon resonance. 10. The health monitoring system of claim 9, wherein the surface plasmon medium comprises two or more components and one or more dielectric components interposed between the two or more components, wherein each of the two or more components is composed of a surface plasmon material and each of the one or more dielectric components is composed of a dielectric material, and wherein a permittivity of the two or more components is greater than a permittivity of the one or more dielectric components. 11. The health monitoring system of claim 9, wherein the surface plasmon medium comprises a loop-wire structure, the loop-wire structure comprises a plurality of wire elements, and each wire element comprises at least one vertical wire with at least one planar loop wound at intervals. 12. The health monitoring system of claim 9, wherein the surface plasmo... | University of Seoul Industry Cooperation Foundation | A61B005/00 H01P003/12 H02J007/02 H02J005/00 H01Q001/24 H01Q009/06 H01Q001/27 H01Q001/32 H01Q023/00 H01P003/00 H01P003/16 H02J007/00 H01Q001/00 H01Q015/00
US09427186 | B2 | 2016-08-30 | Magnetic probe apparatus | 2010-12-06 | A system and method for locating magnetic material. In one embodiment the system includes a magnetic probe; a power module in electrical communication with the magnetic probe to supply current to the magnetic probe; a sense module in electrical communication with the magnetic probe to receive signals from the magnetic probe; and a processing module in electrical communication with the power module and the sense module. The processing module generates a waveform that controls the supply of current from the power module and receives a signal from the sense module that indicates the presence of magnetic material. The magnetic probe is constructed from a material having a coefficient of thermal expansion of substantially 10−6/° C. or less and a Young's modulus of substantially 50 GPa or greater. In one embodiment magnetic nanoparticles collect in the lymph nodes. In one embodiment the particles have a mean hydrodynamic diameter of between 5-200 nm. | 1. A system for locating a magnetic marker in a patient comprising: a magnetic probe comprising; a drive coil for producing an alternating magnetic field; anda sense coil for detecting a magnetic field as generated by the magnetic marker in response to the alternating magnetic field produced by the drive coil,wherein the drive coil and sense coil having different radii, and being spaced axially from each other; andwherein the axial spacing between the drive coil and the sense coils is such that as the radius of the larger coil varies from an initial radius the change in the mutual inductance between the sense coil and the drive coil caused by the change in radius of the larger coil is zero,a power module in electrical communication with the magnetic probe to supply current to the magnetic probe;an amplifier in electrical communication with the magnetic probe to receive signals from the magnetic probe;a processing module in electrical communication with the power module and the amplifier; anda plurality of magnetic nanoparticles,wherein the processing module generates a waveform that controls the supply of current from the power module to the magnetic probe;wherein the processing module receives a signal from the amplifier that indicates the proximity to the injectable magnetic nanoparticles; andwherein the magnetic nanoparticles have a mean diameter size of greater than 200 nm thereby not to migrate but remain positioned at the site of introduction. 2. The system of claim 1 wherein the injectable magnetic nanoparticles have a mean diameter size of greater than 1000 nm. 3. The system of claim 1 further comprising a second drive coil having a different radius that the sense coil. 4. The system of claim 3 wherein said drive coils form a first order gradiometer. 5. The system of claim 1 further comprising a second sense coil having a different radius that the drive coil. 6. The system of claim 5 wherein said sense coils form a first order gradiometer. | ENDOMAGNETICS LTD. | A61B005/00 A61B005/05 A61B005/06 A61B010/00 A61K049/18 B82Y005/00
US09622669 | B2 | 2017-04-18 | Diagnosis of neurodegenerative disorders | 2009-06-15 | Methods are provided for non-invasive early screening of a subject for a neurodegenerative disorder, by analyzing one or more parameters associated with the neurodegenerative disorder. | 1. A method of detecting an increased risk of developing Parkinson's disease or Parkinson's-like disease indicated by a loss of standard deviation of heart rate (SDHR), said method comprising: a) selecting a subject that does not exhibit motor symptoms indicative of Parkinson's disease or Parkinson's-like disease; b) obtaining 5 or fewer minutes of heart rate data;c) determining a SDHR of said subject from the heart rate data;d) detecting the increased risk for developing Parkinson's disease or Parkinson's-like disease if said SDHR of said subject is less than a control SDHR; ande) taking a sample for genetic testing, if the increased risk for developing Parkinson's disease or Parkinson's-like disease has been detected in said subject in step d). 2. The method of claim 1, wherein the subject is in a wakeful state or awake while measuring heart rate. 3. The method of claim 1, further comprising treating said subject. 4. The method of claim 1, wherein the subject scores a 0 on a scale selected from a Unified Parkinson's Disease Rating Scale (UPDRS) and a Hoehn and Yahr scale. 5. The method of claim 1, wherein the control SDHR is obtained from one or more healthy subjects. 6. The method of claim 1, wherein the control SDHR is obtained from one or more subjects selected from a subject that is not diagnosed with Parkinson's disease or Parkinson's-like disease. 7. The method of claim 1, wherein the subject does not exhibit motor symptoms indicative of Parkinson's Disease or Parkinson's like disease as assessed by a scale selected from the Hoehn and Yahr Scale and the Unified Parkinson's Disease Rating Scale (UPDRS). 8. The method of claim 1, wherein the SDHR of said subject is below the average for the control SDHR. 9. A method of detecting an increased risk of developing Parkinson's disease or Parkinson's-like indicated by a loss of standard deviation of heart rate (SDHR), said method comprising: a) selecting a subject that does not exhibit motor symptoms indicative of Parkinson's disease or Parkinson's-like disease as assessed by the Hoehn and Yahr Scale or the Unified Parkinson's Disease Rating Scale (UPDRS);b) obtaining 5 or fewer minutes of heart rate data from said subject;c) determining a SDHR of said subject from the heart rate data;d) detecting the increased risk for developing Parkinson's disease or Parkinson's-like disease if said SDHR of said subject is less than a control SDHR, ande) taking a sample for genetic testing, if the increased risk for developing Parkinson's disease or Parkinson's-like disease has been detected in said subject in step d). | PARKINSON'S INSTITUTE | A61B005/024 A61B005/0245 A61B005/00
US09452022 | B2 | 2016-09-27 | Disposable wand and sensor for orthopedic alignment | 2010-12-31 | A disposable tool suitable for use in orthopedic alignment that registers points of interest on a first and second bone and transmits location data related to the points of interest to the sensor to assess orthopedic alignment with the points of interest. A display can report and visually display alignment information in real-time. | 1. A sensor for a surgical orthopedic alignment system comprising: at least two ultrasonic transducers located at predetermined positions relative to each other on the sensor;an accelerometer configured to identify axial tilt; anda controller operatively coupled to the at least two ultrasonic transducers and the accelerometer, where at least one of the at least two ultrasonic transducers are configured to measure ultrasonic signals, where the measured ultrasonic signals are used to calculate a time of flight, where the time of flight and a differential time of flight is used to calculate a location of an object in three-dimensional space within a surgical field, where the sensor is configured to be attached to a portion of the object, where the object is a portion of a muscular skeletal system, where the location of the object can be used by a processor to determine alignment of the muscular skeletal system and where a trigger system is configured to initiate a measurement process where a trigger signal is faster than an ultrasonic signal. 2. The sensor of claim 1 where the sensor further includes a mounting structure to attach the sensor in a reference position and where the mounting structure is configured to attach to bone, cutting block, or jig. 3. The sensor of claim 1 further including an infrared detector for receiving a trigger signal. 4. The sensor of claim 1 where the sensor further includes: a printed circuit board and where the at least two ultrasonic transducers are coupled to a first side of the printed circuit board where each of the at least two ultrasonic transducers have a predetermined location on the printed circuit board for receiving signals directed thereto. 5. The sensor of claim 4 where a second side of the printed circuit board includes at least three ultrasonic transducers each at a predetermined location on the second side of the printed circuit board for receiving signals directed thereto. 6. The sensor of claim 1 further including an optical detector for receiving an optical trigger signal. 7. The sensor of claim 1 further including: a power source coupled for powering the sensor; andwireless communication circuitry coupled to the controller for transmitting alignment information. 8. The sensor of claim 1 wherein the sensor is line of sight with at least one wand. 9. The sensor of claim 1 where the sensor includes a user interface coupled to the controller that can visually or audibly report a status. 10. A wand for a surgical orthopedic alignment system comprising: at least two ultrasonic transducers located at predetermined positions relative to each other on the wand;a trigger system configured to initiate a measurement process and where a trigger signal is faster than an ultrasonic signal;an accelerometer configured to identify axial tilt; anda controller operatively coupled to the at least two ultrasonic transducers and the accelerometer, where at least one of the at least two ultrasonic transducers emi... | ORTHOSENSOR INC | A61B017/58 A61B017/60 A61F002/00 A61B019/00 A61B008/00
US09398861 | B2 | 2016-07-26 | Methods and devices for assessing intracranial pressure | 2010-12-03 | Provided are methods for noninvasively assessing intracranial pressure (“ICP”) based on optic nerve sheath diameter (“ONSD”) and a blood velocity metric, such as pulsatility or resistivity index. Also provided are related devices and systems for performing the claimed methods. | 1. A method, comprising: providing an imaging device operatively connected to an intracranial pressure estimation component that comprises a processor;operating the imaging device so as to obtain optic nerve sheath diameter information and blood velocity information from a periocular blood vessel of a subject;outputting the optic nerve sheath diameter information and blood velocity information to the intracranial pressure estimation component; andexecuting instructions on the processor of the intracranial pressure estimation component,the instructions including receiving the optic nerve sheath diameter information and blood velocity information, estimating an intracranial pressure of the subject based on a mathematical combination of the optic nerve sheath diameter information and blood velocity information, and outputting the estimate of the intracranial pressure of the subject. 2. The method of claim 1, wherein the blood velocity information comprises pulsatility index, resistivity index, systolic velocity, diastolic velocity, or any combination thereof. 3. The method of claim 2, wherein the blood velocity information comprises pulsatility index, resistivity index, or both. 4. The method of claim 1, wherein the mathematical combination comprises multiplying at least one of the optic nerve sheath diameter or the blood velocity information by a factor, adding together the optic nerve sheath diameter or the blood velocity information, comparing at least one of the optic nerve sheath diameter or the blood velocity information to a control value, or any combination thereof. 5. The method of claim 1, wherein the estimating further comprises comparing the blood velocity information to a control value. 6. The method of claim 4, further comprising assigning a score value to the subject based on the mathematical combination of the optic nerve sheath diameter and blood velocity information. 7. The method of claim 6, wherein the score value correlates to the subject's estimated intracranial pressure. 8. The method of claim 1, wherein at least one of optic nerve sheath diameter and blood velocity information are obtained with ultrasound. 9. The method of claim 8, wherein the blood velocity information is obtained with Doppler ultrasound. 10. The method of claim 8, wherein the optic nerve sheath diameter is obtained with ultrasound. 11. The method of claim 1, further comprising selection of a triage plan based on the estimate of intracranial pressure. 12. The method of claim 1, wherein the obtaining optic nerve sheath diameter and blood velocity information from a periocular blood vessel is performed noninvasively. 13. The method of claim 1, further comprising placing the estimate of intracranial pressure on a computer-readable medium. 14. The method of claim 1, wherein the optic nerve sheath diameter and blood velocity information from a periocular blood vessel of a subject are collected in an automated fashion. 15. The method of claim 1, wher... | Third Eye Diagnostics, Inc. | A61B005/03 A61B005/00 A61B008/04 A61B008/06 A61B008/10 A61B006/03 A61B006/00 A61B008/00 A61B008/08
US07953615 | B2 | 2011-05-31 | System and method of administering, tracking and managing of claims processing | 2001-04-03 | A system and method of administering, tracking and managing of claims processing. | 1. A computer comprising: a processor; anda memory, coupled to the processor and having computer program code embodied therein for enabling the processor to: receiving data related to the insurance claim, the data comprising a plurality of data elements, a data element serving as an assignment criterion;assign a score to first and second data elements of the plurality of data elements based on scoring rules, wherein each data element includes one or more data from the group consisting of policy information, vehicle information, number of vehicles involved, repair cost, number of parties involved, time of incident, and location of incident;determine an overall score of the insurance claim based on the assigned scores;determine a class of the insurance claim according to classing rules;anddetermine a type of assignee to whom to assign the insurance claim according to the application of business rules to the overall score of the insurance claim and the class of the insurance claim, wherein the business rules weight the class more highly than the score. 2. A computer comprising: a processor, anda memory, coupled to the processor and having computer program code embodied therein for enabling the processor to:determine, based on data elements of the insurance claim, a claim score for first and second data elements, wherein each data element includes one or more data from the group consisting of policy information, vehicle information, number of vehicle involved, repair cost, number of party involved, time of incident, and location of incident;determine an overall score for the insurance claim using each of the claim scores;determine, based on the insurance claim, a class of the insurance claim;determine, based on the overall claim score and the class, the type of assignee; andassign the insurance claim to the determined assignee. 3. The computer of claim 2, wherein an effect of the class on the type of assignee is weighted greater than an effect of the claim score on the type of assignee. 4. The computer of claim 2, wherein the claim score reflects a severity of the insurance claim. 5. The computer of claim 2, wherein the claim score is determined based on an element of a group comprising information regarding an insurance policy, information regarding a party involved in a loss, and information regarding how a loss was reported. 6. The computer of claim 2, further comprising determining, based on the claim score, a priority of the insurance claim. 7. The computer of claim 2, wherein an assignee comprises an element of a group comprising an insurance adjuster, a repair facility, an appraiser, and a rental provider. 8. The computer of claim 2, wherein an effect of the first data element on the claim score is greater than an effect of the second data element on the claim score. 9. The computer of claim 2, wherein each data element includes an element score, and wherein the claim score is based on the element scores of the data elements. 10. The ... | Mitchell International, Inc. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00 G06Q040/00
US07941200 | B2 | 2011-05-10 | System and method for determining drug administration information | 2005-12-08 | A system for determining drug administration information may comprise an input device providing for user input of feed forward information having a first parameter component and a second parameter component, a data storage device and a processor. The data storage device may have stored therein a map correlating values of the first and second parameters to drug administration information. The processor may be responsive to user input of the feed forward information to determine corresponding drug administration information according to the map. | 1. A system for determining drug administration information for a user based at least on the user ingesting a meal, the system comprising: an input device providing for user input of feed forward information in the form of meal information having a first parameter component corresponding to carbohydrate content of the meal and a second parameter component corresponding to an expected speed of overall glucose absorption from the meal by the user, wherein the meal information relates to a meal that is to be ingested in the future,a data storage device having stored therein a map correlating values of the first and second parameter components to drug administration information, anda processor responsive to the user input of the feed forward information to determine corresponding drug administration information according to the map,wherein the input device includes a display unit, and wherein the processor controls the display unit to display a graphical user interface,wherein the graphical user interface has a first axis defined by values of the first parameter component and a second axis defined by values of the second parameter component, the graphical user interface providing for the user input of feed forward information in the form of a user selection of a corresponding pair of values of the first and second parameter components; andwherein the input device and the processor provide the user with the ability to modify one or more of time and date stamp information that is associated with previously entered meal information subsequent to actual meal ingestion. 2. The system according to claim 1 wherein the graphical user interface includes a touch responsive interface and defines a grid-type user interface providing for the user selection of a corresponding pair of values of the first and second parameter components. 3. The system according to claim 1 wherein the graphical user interface defines a continuous function of the first and second parameter components. 4. The system according to claim 1 wherein the drug administration information relates to a blood glucose lowering drug or insulin. 5. The system according to claim 1 wherein the processor controls the display unit to display via the graphical user interface values of the carbohydrate content in a format selected from direct estimates of carbohydrate weight, meal size values, and meal size values relative to a reference meal size. 6. The system according to claim 1 wherein the processor controls the display unit to display via the graphical user interface values of the expected speed in a format selected from meal duration values, meal duration values relative to a reference meal duration, and total glycemic index values. 7. The system according to claim 1 wherein the processor controls the display unit to display via the graphical user interface values of the expected speed in a format comprising a fat amount, a protein amount, and a carbohydrate amount, and displays the values of... | Roche Diagnostics Operations, Inc. | A61B005/00
US07937138 | B2 | 2011-05-03 | Method of monitoring brain function | 2004-01-14 | A method for assessing brain state by analysing mammalian brain electroencephalogram (“EEG”) recordings using an eighth order autoregressive and fifth order moving average discrete time equation. | 1. A method for assessing brain state comprising: analysing mammalian brain electroencephalogram (EEG) recordings using an eighth order autoregressive and fifth order moving average discrete time equation;obtaining a z-domain equation by taking a z-transform for said eighth order autoregressive and fifth order moving average discrete time equation;determining, using a processing unit, poles and zeroes in the solution of the z-domain equation; andplotting the poles onto the complex plane; anddisplaying positions of the poles. 2. The method as claimed in claim 1 further comprising a system having means for analysing mammalian brain electroencephalogram (EEG) recordings using an eighth order autoregressive and fifth order moving average discrete time equation. 3. A method of assessing the state of a mammalian brain including the steps of: obtaining an electroencephalogram (EEG) from the brain;(ii) digitising the EEG to define a digitised EEG data signal;(iii) segmenting the EEG data signal into time frames of fixed length, y[n];(iv) approximating, using a processing unit, each digitised time frame by a first equation:y⁡[n]=-∑k=18⁢ak⁢y⁡[n-k]+∑k=05⁢bk⁢u⁡[n-k]where ak and bk are coefficients to be determined for the EEG data signal, y[n] represents the digitized EEG data signal, and u[n] represents a Gaussian white noise process;(v) solving the first equation to determine coefficients a1 to a8 and b0 to b5;(vi) performing a z-transform, using a processing unit, on the first equation to obtain a second, z-domain equation:Y⁡(z)=∑k=05⁢bk⁢z-k1+∑k=18⁢ak⁢z-k⁢U⁡(z)=B⁡(z)A⁡(z)⁢U⁡(z)whereY[z] represents y[u] in the z-domain and U(z) represents u[n] in the z-domain, and:A⁡(z)=1+∑k=18⁢⁢ak⁢z-kB⁡(z)=∑k=05⁢⁢bk⁢z-k(vii) substituting each of the values of the coefficients into the z-domain equation;(viii) solving A(z)=0 for z in the second equation to determine the poles;(ix) determining, using a processing unit, the positions of the poles in the complex plane;(x) repeating steps (iv) to (ix) for each frame in the sample to determine clusters of poles in the complex plane; and(xi) displaying the position and distribution of at least some of said clusters of poles in the complex plane. 4. A method as claimed in claim 3 including the step of filtering the EEG to remove noise signals therefrom prior to carrying out step (iii). 5. A method as claimed in claim 3, wherein said EEG is obtained and recorded before it is processed. 6. A method as claimed in claim 5 including the step of filtering the EEG to remove noise signals therefrom prior to carrying out step (iii). 7. A method as claimed in claim 3, wherein the step of step (x) is repeated continuously to track the motion of the poles from each segment. 8. A method as claimed in claim 3, wherein the step of step (xi) includes the steps: (xi)(a) taking the centroid of the poles for each cluster of poles; and(xi)(b) monitoring and comparing the movement of said centroids. 9. A method as claimed in claim 8 including... | Cortical Dynamics Pty Ltd | A61B005/04
US07951090 | B2 | 2011-05-31 | Cannula equipped with a protective housing | 2007-03-16 | The invention relates to a cannula, preferably equipped with a cannula holder for, in particular a blood sampling device, which comprises a pivotable protective housing (102) for a free end of the cannula. Said protective housing (102) is closed on the periphery except for a cannula insertion opening (8) provided between opposing side walls. Said insertion opening is produced by at least one pair of lamella (8a, 8b), formed on the side walls, the lamella tapering towards one another in the manner of a funnel and forming an expandable funnel slot. The protective housing (102) comprises two housing parts which can be inserted into each other. The inner housing (2a) comprises the pair of lamella (8a, 8b) and an outer housing (11), which surrounds the inner housing (2a), is embodied as a sleeve that is less elastic than the inner housing (2a). | 1. A cannula equipped with a cannula holder for a blood-sampling device and comprising a pivotal protective casing for a free end of the cannula, the protective casing being laterally closed except for a cannula insertion slot provided between opposing side walls, the insertion slot being formed by at least one pair of lips formed on the side walls and angled toward each other like an outwardly flaring funnel and forming a widenable funnel slot, wherein the protective casing comprises inner and outer casing parts concentrically inserted into each other,the inner casing part comprises the pair of lips, andthe outer casing part surrounds the inner casing part, is formed as a sleeve that is less elastic than the inner casing part, and presses inward on the inner casing part to bias the lips into engagement with each other and thereby bias the slot closed. 2. A cannula equipped with a cannula holder for a blood-sampling device and comprising a pivotal protective casing for a free end of the cannula, the protective casing being laterally closed except for a cannula insertion slot provided between opposing side walls, the insertion slot being formed by at least one pair of lips formed on the side walls and angled toward each other like a funnel and forming a widenable funnel slot, wherein the protective casing comprises inner and outer casing parts concentrically inserted into each other, the inner casing part comprising the pair of lips and the outer casing part, which surrounds the inner casing part, is formed as a sleeve that is less elastic than the inner casing part; andthe outer casing part is equipped with contact surfaces that exert an inward locking force toward the funnel slot on opposing, complementarily provided contact surfaces of the inner casing part so as to urge the lips together and thereby bias the slot closed. 3. In combination with a cannula having a free end projecting along an axis from a holder base, a protective casing comprising: a generally tubular inner part formed with a pair of axially extending lips having edges releasably engaging each other at an axially extending normally closed slot, at least one of the lips being elastically flexible and bearing angularly on the lip, the lips converging radially inward at the slot;a generally tubular outer part that is substantially more rigid than the inner part, that generally coaxially surrounds the inner part, and that is formed with an axially extending outer slot aligned radially with the inner slot; andhinge means for supporting the inner and outer parts on the base for movement between an open position with the cannular free end outside of the inner and outer parts and available for use and a closed position generally coaxially surrounding the cannula free end and wholly enclosed inside the inner part, the one lip being sufficiently flexible that it deflects and lets the cannula free end pass through the inner slot on movement of the parts from the open to the closed positi... | Sarstedt AG & Co | A61B005/00 A61M005/00 A61M005/32 B65D081/00
US20110105891 | A1 | 2011-05-05 | Method and device for generating a passive movement in a diagnostic device | 2010-10-23 | A method and an apparatus for generating a passive movement of a patient in a magnetic resonance tomograph. The apparatus includes a support for supporting at least one body part of the patient. The support, together with a patient bed can be moved into a passage of the magnetic resonance tomograph. Driven by a motor, the support can be swiveled about at least one axis inside the passage of the magnetic resonance tomograph. The components of the support and its drive, which are disposed inside the passage of the magnetic resonance tomograph, are made of non-ferromagnetic materials. | 1-10. (canceled) 11. A method of acquiring images of a movement of a body or at least one part of a body of an object to be imaged in an imaging passage of an imaging modality comprising: fixing the body or the at least one part of a body on a support to be disposed in the imaging passage;generating a defined reproducible passive movement about at least one axis of the body or the at least one part of a body or the support by a non-ferromagnetic motorized drive while disposing the support and the drive and the body or the at least one part of a body within the imaging passage; andacquiring a series of images continuously in real time while and during imaging and while or during the passive movement of the body or the at least one part of a body or the support. 12. The method according to claim 11 comprising providing shielding against magnetic flux artifacts in the acquired images or the images as acquired. 13. The method according to claim 11 comprising providing electrical grounding against artifacts in the acquired images or the images as acquired. 14. The method according to claim 11 wherein the passive movement comprises applying pressure and/or movement and/or swiveling and/or rotating and/or pivoting to the body or the at least one part of a body or the support. 15. The method according to claim 14 wherein the applied pressure and/or movement and/or swiveling and/or rotating and/or pivoting is selected from the group consisting of continuous, constant, adjustable, step-by-step, variable and a combination thereof. 16. The method according to claim 11 wherein the passive movement comprises applying a load on the body or the at least one part of a body or the support. 17. The method according to claim 11 wherein the passive movement imitates the physiological movement of the body or the at least one part of a body. 18. The method according to claim 11 wherein the images are kinematic or static images of a position of the body or the at least one part of a body disposed in the imaging passage. 19. The method according to claim 11 wherein a predefined continuous movement of the body or the at least one part of a body or the support is generated or reproduced by the drive. 20. The method according to claim 11 wherein the body or the at least one part of a body includes a joint of a living object. 21. The method according to claim 11 wherein the body or the at least one part of a body is an ankle joint. 22. The method according to claim 11 wherein the imaging modality is a magnetic resonance tomograph. 23. The method according to claim 11 wherein object is an anatomical body part having a joint. 24. The method according to claim 11 wherein the passive movement is applied pressure and/or rotation and/or swiveling and/or movement and/or pivoting to the support while the body or the at least one part of the body is at rest. 25. The method according to claim 11 wherein the passive movement imitates a physiological movement of the body ... | - | A61B005/055
US20110152700 | A1 | 2011-06-23 | ELECTRONIC SPHYGMOMANOMETER FOR ENHANCING RELIABILITY OF MEASUREMENT VALUE | 2011-03-04 | An electronic sphygmomanometer has a cuff to be attached to a measurement site, a pressurization/depressurization unit for adjusting a pressure to be applied to the cuff, a pressure detection unit including a plurality of pressure sensors, for detecting cuff pressures in the cuff based on pressure information output from the plurality of pressure sensors, and a blood pressure calculating unit for calculating a blood pressure based on a change in the cuff pressures detected by the pressure detection unit. Blood pressure measurement and detection of abnormality on the plurality of pressure sensors are carried out based on the cuff pressures respectively corresponding to the plurality of pressure sensors detected according to the pieces of pressure information output from the plurality of pressure sensors. | 1. An electronic sphygmomanometer comprising: a cuff to be attached to a measurement site;a pressurization/depressurization unit for adjusting a pressure to be applied to the cuff;a pressure detection unit including a plurality of pressure sensors, for detecting cuff pressures in the cuff based on pressure information output from the plurality of pressure sensors; anda blood pressure calculating unit for calculating a blood pressure based on a change in the cuff pressures detected by the pressure detection unit,wherein blood pressure measurement and detection of abnormality on the plurality of pressure sensors are carried out based on the cuff pressures respectively corresponding to the plurality of pressure sensors detected according to the pieces of pressure information output from the plurality of pressure sensors. 2. The electronic sphygmomanometer according to claim 1, wherein the blood pressure calculating unit calculates the blood pressure based on the cuff pressures respectively corresponding to the plurality of pressure sensors. 3. The electronic sphygmomanometer according to claim 1, further comprising: an abnormality detecting portion for detecting abnormality of the plurality of pressure sensors,wherein the abnormality detecting portion mutually compares the pieces of pressure information respectively corresponding to the plurality of pressure sensors and detects that at least one of the plurality of pressure sensors is abnormal based on a comparison result. 4. The electronic sphygmomanometer according to claim 3, wherein the blood pressure calculating unit excludes the cuff pressure corresponding to the pressure sensor that is detected as being abnormal, out of the plurality of pressure sensors from data used to calculate the blood pressure. 5. The electronic sphygmomanometer according to claim 3, further comprising: a storage unit for storing data of the blood pressure calculated by the blood pressure calculating unit,wherein a detection result by the abnormality detecting portion is stored in the storage unit in association with the data of the blood pressure. 6. The electronic sphygmomanometer according to claim 5, wherein the pressure information is detected in a process of increasing or in a process of decreasing the pressure to be applied to the cuff by the pressurization/depressurization unit. 7. The electronic sphygmomanometer according to claim 3, wherein the pressure information is detected in a process of increasing or in a process of decreasing the pressure to be applied to the cuff by the pressurization/depressurization unit. 8. The electronic sphygmomanometer according to claim 3, wherein the abnormality detecting portion mutually compares the pieces of pressure information respectively corresponding to the plurality of pressure sensors at a predetermined cuff pressure, and detects that at least one of the plurality of pressure sensors is abnormal based on a comparison result, andthe predetermined cuff pressure c... | OMRON HEALTHCARE CO., LTD. | A61B005/0225
US20110152634 | A1 | 2011-06-23 | Measuring Human Biological Fluid Levels | 2009-12-17 | This disclosure describes measuring and analyzing a biological fluid level using a telecommunications device. Measuring the biological fluid level includes initiating a measurement and analysis application to activate measuring and analyzing the biological fluid level. The telecommunications device transmits a first signal to activate a measurement device placed on a human body. Once the biological fluid level is measured, the telecommunications device detects a second signal. Upon detecting the second signal, the biological fluid level is received and compared against a baseline range as predetermined by a user. Next, the measured biological fluid level and a symbol illustrating conditions of the level are presented on a user interface of the display of the telecommunications device. | 1. A method comprising: under control of at least partially by a telecommunications device configured with executable instructions: presenting a user interface on a display of the telecommunications device, the user interface including a measurement and analysis application stored in memory of the telecommunications device;initiating the measurement and analysis application to activate measuring a glucose level;transmitting a first signal to activate a measurement device placed on a human body through a wireless transmitter, the measurement device to measure the glucose level;measuring the glucose level;detecting a second signal received from the wireless transmitter connected to a measurement device, upon completion of measuring the glucose level;receiving the glucose level from the measurement device to a wireless receiver;comparing the received glucose level to a baseline glucose level predetermined by a user, the baseline glucose level identified as before a meal; andvisually presenting the received glucose level and a symbol on the user interface of the telecommunications device. 2. The method of claim 1, wherein the telecommunications device is a cellular device, a smart phone, a portable media player, a computing device, or a handheld computing device that operates in wireless networks such as a Personal Area Network (PAN), a Local Area Network (LAN), or a Wide Area Network (WAN). 3. The method of claim 1, wherein the initiating the measurement and analysis application to measure a glucose level comprises at least one of manual activation by receiving input on the user interface or automatic activation by identifying predetermined time periods to measure the glucose level. 4. The method of claim 1, further comprising: presenting the received glucose level as a numerical number or on a graph within a predetermined range;determining the symbol based on comparing the received glucose level to the baseline glucose level, when the received glucose is approximately about the same number as the baseline glucose level, the symbol is a smiley face, a thumbs up, or a color indicator; anddetermining the symbol based on comparing the received glucose level to the baseline glucose level, when the received glucose is approximately 15 mg/dl less than or greater than the baseline glucose level, the symbol is a frown face, a thumbs down, or the color indicator. 5. The method of claim 1, further comprising: determining whether the received glucose level is outside or within a glucose range predetermined by the user; in response to the received glucose level being within the glucose range, storing the received glucose level in memory of the telecommunications device; andin response to the received glucose level being outside the glucose range, generating an alert on the telecommunications device to vibrate, to play an audible sound, to vibrate and to play the audible sound, to send an email to an identified address, to send a text message to an identifie... | - | A61B005/145 A61B005/00
US20110105872 | A1 | 2011-05-05 | SYSTEMS AND METHODS FOR APPLICATION TO SKIN AND CONTROL OF ACTUATION, DELIVERY, AND/OR PERCEPTION THEREOF | 2010-10-29 | The present invention generally relates, in certain aspects, to systems and methods for application to the skin, and control of actuation, delivery, and/or perception thereof. For example, certain aspects of the invention are generally directed to devices for delivering to and/or withdrawing fluid or from subjects, e.g., to or from the skin and/or beneath the skin. In one aspect, the delivery and/or withdrawal of fluid is at least partially obscured. For example, the obscuration may be by time, and/or by sensory obscuration (e.g., by providing tactile, olfactory, auditory, and/or visual sensations). Certain aspects of the invention are generally directed to devices able to automatically deliver to and/or withdraw fluid from the skin and/or beneath the skin after activation, e.g., where the fluid is delivered and/or withdrawn without the need for any additional intervention. The devices may be activated by any suitable technique. In other aspects, a device may be contained in a package, and when removed from the package, the device may be constructed and arranged to be able to deliver and/or withdraw fluid when the device is applied to the skin, and in some cases, automatically and/or without any further intervention by the subject or another person. | 1-35. (canceled) 36. A device for extracting a fluid from the skin and/or from beneath the skin of a subject, the device comprising: a support structure for application to the skin of the subject;a fluid transporter associated with the support structure;an extraction activator associated with the support structure able to activate the fluid transporter to extract the fluid from the skin and/or from beneath the skin of the subject; anda sensory obscurer for producing a tactile, olfactory, auditory, and/or visual sensation able to at least partially obscure sensation produced by the fluid transporter. 37. The device of claim 36, wherein the support structure further comprises an adhesive for adhesion of the support structure to the skin of the subject. 38-42. (canceled) 43. The device of claim 36, wherein the extraction activator is able to be activated remotely. 44. The device of claim 36, wherein the extraction activator comprises a timing mechanism for timed activation of the extraction activator after removal from packaging containing the device. 45. A device for extracting a fluid from the skin and/or from beneath the skin of a subject, the device comprising: a support structure for application to the skin of the subject;a fluid transporter associated with the support structure;an extraction activator associated with the support structure able to activate the fluid transporter to extract the fluid from the skin and/or from beneath the skin of the subject; andan applicator positioned to apply an analgesic to the skin of the subject. 46-146. (canceled) 147. A device for multiple analyses of a fluid withdrawn from the skin of a subject, comprising: a plurality of means for extracting a fluid from the skin and/or from beneath the skin of the subject,wherein the plurality of means for extracting are initially fastened to the device and/or each other for use with a first extraction and, after first extraction and removal of the device from the skin of the subject, one means for extracting is removed from the device. 148. A device for extracting a fluid from the skin and/or from beneath the skin of a subject, the device comprising: a support structure for application to the skin of the subject;a fluid transporter associated with the support structure; andan extraction activator associated with the support structure able to activate the fluid transporter to extract the fluid from the skin and/or from beneath the skin of the subject,wherein the extraction activator is constructed and arranged to activate the fluid transporter by application of the support structure to the skin of the subject, and/or by one-time manipulation of an actuator, whereby extraction of the fluid is completed without additional manipulation of the extraction activator. 149. The device of claim 148, wherein the support structure further comprises an adhesive for adhesion of the support structure to the skin of the subject. 150-158. (canceled) 159. The device of claim... | Seventh Sense Biosystems, Inc. | A61B005/157 A61B005/15
US20110160549 | A1 | 2011-06-30 | METHOD, SYSTEM AND APPARATUS FOR USING ELECTROMAGNETIC RADIATION FOR MONITORING A TISSUE OF A USER | 2008-09-04 | A wearable monitoring apparatus for monitoring at least one biological parameter of an internal tissue of an ambulatory user. Said wearable monitoring apparatus comprises at least one transducer configured for delivering electromagnetic (EM) radiation to said internal tissue and intercepting a reflection of said EM radiation said reform in a plurality of transmission sessions during at least 24 hours, a processing unit configured for analyzing respective said reflection and identifying a change in said at least one biological parameter accordingly, a reporting unit configured for generating a report according to said change, and a housing for containing said at least one transducer, said reporting unit, and said processing unit, said housing being configured for being disposed on said body of said ambulatory user. | 1. A wearable monitoring apparatus for monitoring at least one biological parameter of an internal tissue of an ambulatory user, comprising: at least one transducer having at least one antenna configured for at least one of delivering electromagnetic (EM) radiation to the internal tissue and intercepting said EM radiation from said internal tissue in a plurality of transmission sessions during a period of at least 24 hours;a processing unit configured for analyzing said intercepted EM radiation and identifying a change in the at least one biological parameter accordingly; anda reporting unit configured for generating a report according to said change;wherein said at least one antenna is configured for being disposed on the body of the ambulatory user. 2. The wearable monitoring apparatus of claim 1, wherein said processing unit comprises a communication module for communicating with a remote processing unit thereby allowing performing of at least one of said analyzing and said identifying by said remote processing unit. 3. The wearable monitoring apparatus of claim 1, wherein said processing unit is configured for identifying said change by detecting at least one of a trend, a biological process, and a pattern according to said intercepted EM radiation of said plurality of transmission sessions. 4. The wearable monitoring apparatus of claim 1, wherein said processing unit is configured for evaluating a property change in a dielectric related property of the internal tissue in at least one of said plurality of transmission sessions and performing said identification according to said property change. 5. The wearable monitoring apparatus of claim 1, wherein said plurality of transmission sessions are performed in an adaptive rate. 6. The wearable monitoring apparatus of claim 5, wherein said adaptive rate is determined according to a clinical state of the user, said processing unit being configured for calculating said clinical state according to at least one output of a non EM radiation sensor and said intercepted EM radiation. 7. The wearable monitoring apparatus of claim 5, further comprising a posture detection unit configured for detecting a posture of the user, said adaptive rate being determined according to said posture. 8. The wearable monitoring apparatus of claim 1, wherein said reporting unit is configured for generating said report in real time. 9. (canceled) 10. The wearable monitoring apparatus of claim 1, wherein said change is indicative of a fluid content change in the internal tissue during said period. 11. The wearable monitoring apparatus of claim 1, wherein said change is indicative of a member of a group consisting of: a trauma a degenerative process, atelectasis, a post-operative atelectasis, an acute respiratory deficiency syndrome (ARDS), an infectious cause, an inhaled toxins, a circulating exogenous toxins, a vasoactive substance, a disseminated intravascular coagulopathy (DIC), a burn, an emphysema, a immunolo... | - | A61B005/00 A61B005/103 H01Q001/00
US07966057 | B2 | 2011-06-21 | Methods and apparatus for guided transluminal interventions using vessel wall penetrating catheters and other apparatus | 2002-02-06 | Electro-anatomically navigated catheters (e.g., mapping catheters, tissue penetrating catheters, delivery catheters and/or sheaths) and associated methods whereby devices or substances may be delivered to specific locations within a patient's body and/or penetration tracts or passageways are formed at specific locations between anatomical structures. The catheters are equipped with sensors and a sensing field is created around the body of the patient. The sensor-equipped catheters are then inserted into the patient's body and the position of the catheter-mounted sensor(s) is/are observed on a display representing the sensing field. Apparatus may be included for propelling the catheter(s) to specific locations and/or specific rotation orientations, within the patient's body. | 1. A method for using a navigable tissue penetrating catheter system in a human or veterinary patient to penetrate from a location within the lumen of an artery or vein, outwardly through the wall of that artery or vein and to a target location outside of that artery or vein, said method comprising the steps of: (A) providing a navigable tissue penetrating catheter system comprising; i) a field generator for generating a sensing field about at least a portion of the body of a human or veterinary patient;ii) a tissue penetrating catheter comprising a catheter body, a penetrator that is advanceable laterally from the catheter body so as to penetrate outwardly through the wall of an artery or vein in which the catheter body is positioned to a target location and a sensor mounted on or in the catheter body for sensing components of the sensing field and emitting signals indicative of the rotational orientation of the sensor within the sensing field; andiii) a sensor-monitor adapted to receive signals from the sensor and to provide, on the basis of those signals, indicia of the rotational orientation of the sensor within the sending field;(B) providing a sensor-equipped mapping catheter;(C) using the sensor equipped mapping catheter to map or mark the target location;(D) using the field generator to create a sensing field;(E) positioning at least the portion of the patient's body in which the target location is located within the sensing field;(F) inserting the tissue penetrating catheter into the patient's body and advancing the catheter into the lumen of an artery or vein located within the sensing field;(G) actuating the sensor of the tissue penetrating catheter such that signals from the sensor are received by the sensor monitor;(H) using the indicia provided by the sensor monitor to rotationally orient that catheter body within the anatomical conduit such that, when the penetrator is subsequently advanced laterally from the catheter body, it will penetrate outwardly through the wall of the anatomical conduit in which the catheter body is positioned and to the target location outside of the artery or vein in which the tissue penetrating catheter is positioned; and(I) advancing the penetrator from the catheter body, through the wall of the artery or vein in which the catheter body is positioned and to the target location. 2. A method according to claim 1 wherein the target location is selected from the group consisting of: an artery other than one in which the tissue penetrating catheter is positioned;a vein other than one in which the tissue penetrating catheter is positioned;a location within the parenchyma of an organ;a cavity within an organ;a body cavity;an aberrant anatomical structurean anatomical malformation;a location to which a drug, substance, flow of energy or apparatus is to be delivered; anda location from which a sample, reading image or diagnostic information is to be obtained. 3. A method according to claim 1 further comprising ... | Medtronic Vascular, Inc. | A61B005/05
US07941203 | B2 | 2011-05-10 | Multimodality imaging system | 2007-09-21 | A multimodality imaging system, comprising: a first imaging system for forming a first image; a second imaging system for forming a second image; and a rotating device on which the first imaging system and the second imaging system are fixed so that the first imaging system and the second imaging system are selectively rotated to a scanning position. | 1. A multimodality imaging system, comprising: a first imaging system for forming a first image;a second imaging system for forming a second image; anda moving device on which the first imaging system and the second imaging system are fixed so that the first imaging system and the second imaging system are operable to be selectively moved to a scanning position,wherein the moving device comprises a rotating device on which the first imaging system and the second imaging system are fixed so that the first imaging system and the second imaging system are operable to be selectively rotated to the scanning position; andwherein the rotating device comprises: a base disk on which the first imaging system and the second imaging system are fixed;a base disk stand on which the base disk is rotatably supported; anda driving device adapted to rotate the base disk with respect to the base disk stand so as to rotate the first or second imaging system to the scanning position. 2. The multimodality imaging system according to claim 1, wherein the rotating device further comprises: a detecting device for detecting a rotating angle of the base disk;a braking device adapted to brake the rotating device according to the rotating angle of the base disk detected by the detecting device, so that the first or second imaging system is stopped at the scanning position. 3. The multimodality imaging system according to claim 1, wherein the rotating device further comprises: an external gear type rotary bearing having an outer ring connected with the base disk, and an inner ring, the inner ring of the external gear type rotary bearing being fixed on the base disk stand;a pinion meshing with an external gear on the outer ring of the external gear type rotary bearing; anda motor with a brake, the motor being connected with the pinion through a speed reducer, so that the motor is operable to rotate the base disk through the speed reducer, the pinion and the external gear on the outer ring of the external gear type rotary bearing. 4. The multimodality imaging system according to claim 3, further comprising: a capacitive proximity switch support fixed at a position on an upper surface of the outer ring of the external gear type rotary bearing,a capacitive proximity switch fixed at a position on the capacitive proximity switch support, andan inclined metal block fixed at a position on the inner ring, wherein when the base disk rotates to a first angle, the capacitive proximity switch approaches the inclined metal block so that the capacitive proximity switch sends a control signal to control the motor to be decelerated; and when the base disk rotates to a second angle, the capacitive proximity switch sends control signal to control the motor to be braked and stopped. 5. The multimodality imaging system according to claim 3, further comprising: a positioning device for selectively positioning the first imaging system and the second imaging system at the scanning position. 6. ... | Neusolft Positron Medical Systems Co., Ltd. | A61B005/05
US07955273 | B2 | 2011-06-07 | Device for verifying and monitoring vital parameters of the body | 2002-07-24 | The invention lies in the domain of medical technology and relates to a diagnosis and monitoring device for the rapid diagnosis and monitoring of vital parameters of the human or animal body, in particular of the heart and/or lungs, said device being compact, without cables and/or tubes and easy to use for the user. Devices such as a bell (1) comprising a membrane (2) and/or measuring electrodes (4, 5) for receiving and forwarding acoustic and/or electric signals of the body are arranged in a housing with a cross-section that is approximately the size of the palm of a hand and a height of approximately half a hand-width, on the side of a housing that is to be placed on the body. Said devices are connected to a device in the housing, which analyses, filters and stores the signals of the receiving device and to additional devices for visually reproducing the evaluated signals in digital or analogue form using display fields (14, 15, 16, 17) and/or for acoustically reproducing said signals using a loudspeaker (9) located in the housing. The diagnosis and monitoring device also comprises interfaces (10, 11) for connecting external devices and equipment (e.g. computer, earphones, printer). The principal characteristics of the invention are illustrated in FIGS. 1 and 3. | 1. A device for examination and monitoring of vital parameters, from for example lung and respiration, heart and circulation, of a human or animal body, the device configured as a compact diagnostic device functioning independently of other devices and independently of communication with other devices, adapted to be used single-handedly and within reach of a user without cables, tubes, poking out parts, invasive procedures or implantable parts and comprising a common casing of a size fittable into pockets of a doctor's coat of substantially a size of a palm of a hand with a width and a length of substantially half a hand; devices interacting with the body and devices recording and transmitting one or more signals spontaneously emitted by the body or as a reflex of the interactions with the body all together integrated in said common casing, said recording and transmitting devices including a device performing electronic analysis, filtering and data storage, with intermediate storage of signals supplied by the recording devices; devices providing a parallel playback of a freely definable selection of analyzed signals consisting of devices providing visual playback in digital and analog form, devices providing acoustic playback, and both devices providing visual playback in digital and analog form and devices providing acoustic playback, wherein said body signals recording device includes a multifunctional sensor unit recording electrical voltage fluctuations, impedance measurements and temperature measurements with the same sensor, said body signals recording device including at least two measuring electrodes which are positioned on one level on an underside of said casing for placing over or onto body areas being examined; wherein the multifunctional sensor unit of said body signals recording device also producing, recording and transmitting acoustic signals of different frequency patterns and mechanical strain and elasticity of a body surface and includes a knocking mechanism with an electro-magnetic drive and a striking area so that standardized sound waves are produced and a sound funnel, a membrane and a microphone, said membrane being positioned on an underside of said housing for placing over or onto body areas being examined and said sound funnel is slidable vertically in said casing and thus transfers the strain of the body surface to a pressure recorder, and wherein the multifunctional sensor unit of said recording and transmitting body signals device also recording and transmitting the pulse oxymetry and the arterial capillary and venous blood pressure levels and the recapillarisation time and includes an opening for the insertion of a body part with one side on the exterior wall of the module and with the opposite side of the body part on a dish, and includes a pressure device which is horizontally movable on a guide track, whereby the dish is repulatable—pushable or releasable—by a pressure shaft and a pressure shaft motor against the... | - | A61B005/103
US07946996 | B2 | 2011-05-24 | Vascular introducer with temperature monitoring probe and systems for endovascular temperature control | 2007-02-20 | An introducer sheath/temperature probe assembly that is insertable into a blood vesel of a human or veterinary patent to measure the temperature of blood flowing through that blood vessel. The introducer sheath/temperature probe assembly may be used in conjunction with an indwelling heat exchange catheter system to warm or cool all or a portion of the patient's body to a desired target temperature and to maintain such target temperature for a desired period of time. | 1. A system for controlling the body temperature of a human or veterinary patient, said system comprising: an introducer sheath/temperature probe assembly comprising a tubular introducer sheath having a proximal end, a distal end and a lumen which extends longitudinally therethrough, said introducer sheath being insertable into a blood vessel and an elongate temperature probe having a distal end and a proximal end, said temperature probe being insertable, distal end first, into the lumen of the introducer sheath and useable, when inserted into the lumen of the introducer sheath, to sense the temperature of blood flowing through the blood vessel into which the introducer sheath is inserted, an atraumatic tip member being disposed on the distal end of the temperature probe such that the tip member protrudes beyond the distal end of the introducer sheath when the temperature probe is inserted into the lumen of the introducer sheath, wherein said temperature probe has a temperature sensor, said temperature sensor being located within the elongate temperature probe proximal of said tip member so that said tip member may extend substantially out ofthe distal end of said introducer sheath while said temperature sensor resides within said lumen of the introducer sheath;a heat exchange catheter having a heat exchanger that is positionable within the vasculature of the patient and useable to alternately add heat to and remove heat from the patient's blood;a heater/cooler apparatus for alternately heating and cooling the heat exchanger of the heat exchange catheter, thereby adjusting the amount and direction of heat exchange occurring between the heat exchanger and the patient's blood; anda controller which receives i) a target temperature input and ii) indications of the body temperature sensed by the temperature probe when the temperature probe is inserted into the introducer sheath, said controller being operative to control the heater/cooler apparatus in response to the indications of the body temperature sensed by the temperature probe such that the patient's body temperature as sensed by the temperature probe will be heated or cooled approximately to the target temperature and thereafter maintained at approximately said target temperature;wherein said tubular introducer sheath has a proximal opening diameter on said proximal end, said temperature probe has a distal length having a first diameter and a proximal length having a second diameter, said first diameter smaller than said opening diameter and said second diameter larger than said opening diameter, and the length of the probe distal of said second diameter up to said sensor is less than the length of said introducer lumen, and the length of said distal portion of said probe between said second diameter and the distal tip of said distal tip member is longer than said introducer sheath, such that when said temperature probe is inserted through said proximal opening the portion of the temperature ... | ZOLL Circulation, Inc. | A61B005/00
US07941204 | B1 | 2011-05-10 | Magnetic resonance imaging concepts | 2005-11-16 | Methods of acquiring magnetic resonance imaging (MRI) data for angiography. The present invention includes novel magnetization preparation schemes where the navigator and fat saturation pulses are executed in steady state after the preparatory pulses in order to minimize the delay between the magnetization preparation and the image echoes. The present invention also provides for improved methods of contrast-enhanced MRI where data are collected along non-linear trajectories through k-space and may also involve novel view ordering. In addition, the present methods employ novel motion corrections that minimize motion artifacts. The present invention further provides novel methods of self-calibrated sensitivity-encoded parallel imaging that allow for accurate and rapid scanning of subjects. | 1. A method of acquiring magnetization prepared steady-state magnetic resonance imaging data from a tissue, comprising the steps of: defining a plane within k-space;dividing said plane within k-space into a plurality of pairs of radial wedges, wherein each pair of radial wedges has a common vertex at the center of k-space, further wherein each pair of radial wedges reside on opposite sides of the center of k-space, and further wherein said plurality of pairs of radial wedges represent all of said plane within k-space;collecting a series of images from said tissue by following a trajectory in k-space, wherein said trajectory comprises a first sequence of steps, wherein said first sequence comprises a first collection step across the width of the first of said first pair of radial wedges; anda second collection step along the border of said first of said first pair of radial wedges, wherein the second collection step is oriented towards the center of k-space;further wherein said first and second collection steps are repeated until said trajectory reaches the center of k-space,wherein said trajectory further comprises a second sequence of steps wherein said second sequence comprises a first collection step across the width of the second of said first pair of radial wedges; anda second collection step along the border of said second of said first pair of radial wedges, wherein the second collection step is oriented away from the center of k-space;further wherein said trajectory comprises a third sequence of steps wherein said third sequence comprises a third collection step across the width of said second of said first pair of radial wedges and further across the first of a second pair of radial wedges, wherein said second pair of radial wedges is adjacent to said first pair of radial wedges; anda fourth collection step along the border of said second of said first pair of radial wedges or along the border of said first of said second pair of radial wedges, wherein the second collection step is oriented away from the center of k-space. 2. The method of claim 1, wherein the trajectory traverses the center of k-space at a background nulling point. 3. A method of acquiring magnetization prepared steady-state magnetic resonance imaging data from a tissue, comprising the steps of: (a) for each of a plurality of segments into which k-space has been segmented: (1) enhancing contrast in the tissue; and(2) acquiring image data from the tissue using a segmented steady state free precession sequence along a trajectory comprising points in k-space, wherein: the points of the trajectory are: confined to the segment of k-space or to the segment of k-space and an adjacent segment of k-space; andinclude points in k-space center and k-space periphery;image data is acquired at all points in the trajectory before any image data is acquired along another trajectory; andpoints nearest the k-space center are acquired at a time corresponding to a maximal contrast enhance... | - | A61B005/05 G01V003/00
US08099890 | B2 | 2012-01-24 | Closure mechanism for an identification medium adapted for receiving indicia forming material and dual closure means | 2011-02-28 | A closure mechanism receiving aperture presents a substantially contiguous, uninterrupted planar surface on a flexible medium, primarily for use with identification devices, i.e., wristbands. The aperture does not interrupt the printable area so that printer ink may be printed over the aperture without passing through. The inventive aperture also does not produce waster material or chads which can jam or foul-up a printer. An inventive wristband includes both snap closure and adhesive closure mechanisms that can be used at the same time. | 1. A dual-closure identification wristband, comprising: an elongated strap having first and second ends;a snap closure and an adhesive closure, both closures separately securing the first end of the strap to the second end of the strap, wherein the adhesive closure comprises a removable tab covering a pressure sensitive adhesive proximate to the first end of the strap, wherein removal of the tab exposes the pressure sensitive adhesive for attachment to the second end of the strap, wherein the snap closure comprises a securing aperture proximate to the first end of the strap, a plurality of sizing apertures proximate to the second end of the strap, and a mechanical retainer for engaging the securing aperture and one of the sizing apertures, wherein the securing aperture extends through the removable tab, the pressure sensitive adhesive and a media layer. 2. The dual-closure identification wristband of claim 1, wherein the elongated strap comprises a liner layer adhered to a media layer, and wherein the removable tab comprises a portion of the liner layer having a release layer thereunder. 3. The dual-closure identification wristband of claim 1, wherein the removable tab of the adhesive closure remains in place when the snap closure is utilized. | Precision Dynamics Corporation | A44C005/00 G09F003/14
US08755874 | B2 | 2014-06-17 | Implantable medical device and method for classifying arrhythmia events | 2008-05-07 | In an implantable medical device such as an implantable cardiac defibrillator, and a method for classifying arrhythmia events, IEGM signals are analyzed to detect an arrhythmia event and a respiratory pattern of the patient is sensed. At least one respiratory parameter reflecting characteristics of the respiratory pattern of the patient is determined based on the sensed respiratory pattern and a respiratory measure corresponding to a change of a rate of change of the at least one respiratory parameter is calculated. The detected arrhythmia event is classified based on the respiratory measure and the IEGM signals, wherein arrhythmia events that satisfy at least a first criterion is classified as an arrhythmia event requiring therapy. | 1. An implantable medical device to discriminate between ventricular tachycardias (VTs) and supraventricular tachycardias (SVTs), said implantable medical device being configured to be connected to at least one medical lead adapted to be placed at or within a heart of a patient, said medical lead including at least one electrode adapted to sense electrical activity of said heart and to produce at least one signal representative of said electrical activity, said implantable medical device comprising: an intra-cardiac electrogram (IEGM) analyzing circuit that receives said signal representative of said electrical activity and to analyse IEGMs of said heart to detect an arrhythmia event;a sensing circuit that senses a respiratory pattern of said patient and to produce at least one signal representative of said respiratory pattern;a respiratory parameter determining circuit that determines at least one respiratory parameter reflecting characteristics of the respiratory pattern of said patient based on said sensed respiratory pattern and that determines a respiratory measure corresponding to a change of a rate of change of said at least one respiratory parameter; andan arrhythmia classifying circuit that classifies the detected arrhythmia event as a VT or SVT based on said respiratory measure, including classifying arrhythmia events that satisfy at least a first criterion for VT as an arrhythmia event requiring therapy;wherein said respiratory parameter determining circuit determines a respiratory frequency and a tidal volume using said sensed respiratory pattern and that determines a change of a range of change of a weighted combination of said respiratory frequency and said tidal volume. 2. The implantable medical device according to claim 1, wherein said arrhythmia classifying circuit determines said at least first criterion to be satisfied when said respiratory measure is higher than a predetermined threshold value. 3. The implantable medical device according to claim 2, wherein said arrhythmia classifying circuit is configured to: upon receiving information from said arrhythmia detection circuit that an arrhythmia event has been detected, monitor said respiratory measure;compare the time points of the respective onset of said detected arrhythmia event and of a respiratory measure being higher than said predetermined threshold; anddetermine said at least first criterion to be satisfied if said respiratory measure is higher than a predetermined threshold value and an onset of said change occur within a predetermined period of time from the onset of said detected arrhythmia event. 4. The implantable medical device according to claim 1, further comprising: an impedance measuring circuit that, during impedance measuring sessions, measures at least one impedance signal between at least a first pair of electrodes being configured to be connected to said implantable medical device, said electrodes being located such that measured impedance signals sub... | St. Jude Medical AB | A61B005/04
US08423108 | B2 | 2013-04-16 | Device and system that identifies cardiovascular insufficiency | 2010-03-15 | A system and method for identifying volume status of a patient are disclosed. A pulse density signal is recorded from the patient. The pulse density signal is filtered to capture a respiration sampling period and a plurality of cardiac cycles occurring during the respiration sampling period. Mean pulse pressure and peak blood flow velocity for the respiration sampling period are calculated and are used as indices of volume status of the patient. | 1. A computer-assisted method for identifying a volume status of a patient, the method comprising: administering a cardiovascular preload to the patient; continuously recording, prior to, during, and after the step of administering the cardiovascular preload to the patient, a pulse density signal from a non-invasive transducer transmitting data from the patient;applying, using a processor, a first filter to the pulse density signal to identify a first respiration sampling period recorded prior to the step of administering the cardiovascular preload and a second respiration sampling period recorded after the step of administering the cardiovascular preload;applying, using the processor, a second filter to the pulse density signal to identify a first plurality of cardiac cycles occurring during the first respiration sampling period and a second plurality of cardiac cycles occurring during the second respiration sampling period; andcalculating, using the processor, a variation of the pulse density signal before and after the step of administering the cardiovascular preload, wherein the variation is used as an index of the volume status of the patient, wherein the variation is a quotient of a difference between a mean pulse density signal between the second and the first respiration sampling periods to a sum of the mean pulse density signal of the second and the first respiration sampling periods, wherein the mean pulse density signal for each respiration sampling period is a quotient of a sum of a difference between a maximum point on the pulse density signal and a minimum point on the pulse density signal for each cardiac cycle occurring during such respiration sampling period to a total number of cardiac cycles occurring during such respiration sampling period. 2. The method as in claim 1, further comprising calculating a peak blood flow velocity for at least one of the respiration sampling periods, wherein the peak blood flow velocity is a difference between a mean minimum point on the pulse density signal for the at least one of the respiration sampling periods and the mean pulse density signal for the at least one of the respiration sampling periods. 3. The method as in claim 2, further comprising calculating a change in peak blood flow velocity between the first and second respiration sampling periods. 4. The method as in claim 3, further comprising: classifying the change in peak blood flow velocity into a classification system; and translating the classified change in peak blood flow velocity to identify the volume status of the patient. 5. The method as in claim 1, further comprising: classifying the variation of the pulse density signal into a classification system; and translating the classified variation of the pulse density signal to identify the volume status of the patient. 6. The method as in claim 1, further comprising initiating a treatment protocol. 7. A system for identifying a volume status of a patient, the system compris... | Intelomed, Inc. | A61B005/00
US20110104067 | A1 | 2011-05-05 | METHODS FOR IN VIVO EVALUATION OF PULMONARY PHYSIOLOGY AND/OR FUNCTION USING NMR SIGNALS OF POLARIZED 129XE | 2011-01-10 | In certain embodiments, methods of the present invention obtain dynamic data sets of an NMR spectroscopy signal of polarized 129Xe in a selected structure, environment, or system. The signal data can be used to evaluate: (a) the physiology of a membrane or tissue; (b) the operational condition or function of a body system or portion thereof (when at rest or under stimulation); and/or (c) the efficacy of a therapeutic treatment used to treat a diagnosed disorder, disease, or condition. Thus, the present invention provides methods for screening and/or diagnosing a respiratory, cardiopulmonary disorder or disease such as chronic heart failure, and/or methods for monitoring the efficacy of therapeutics administered to subject to treat the disorder or disease. | 1.-67. (canceled) 68. An in vivo method for evaluating respiratory, pulmonary and/or cardiopulmonary physiology and/or function, comprising: (a) delivering polarized 129Xe in vivo to a subject such that the polarized 129Xe moves serially from an alveolar gas compartment through a tissue compartment with thickness LT, a capillary compartment with thickness LC comprising plasma and red blood cells, and is then taken away in the blood stream;(b) obtaining NMR spectroscopic signals of the polarized 129Xe in the subject over time at selected chemical shift frequencies to generate at least one dynamic data set of the NMR spectroscopic signal strength over time after successively destroying the polarization of the 129Xe repetitively at predetermined delays in the compartments; and(c) calculating a plurality of selected parameters associated with lung function and/or physiology according to a predetermined mathematical diffusion model of alveolar gas exchange using the NMR-derived dynamic data set. 69. A method according to claim 68, further comprising evaluating the calculated parameters to determine a pulmonary condition of a patient to thereby assess respiratory, pulmonary, and/or cardiopulmonary function or associated physiology. 70. A method according to claim 68, further comprising assessing both membrane diffusion and pulmonary perfusion based on the NMR data set. 71. A method according to claim 69, further comprising identifying whether the subject exhibits gas-exchange that is perfusion or diffusion limited. 72. A method according to claim 71, wherein the NMR signals include at least one signal associated with dissolved 129Xe, the dissolved-phase NMR signal having an associated signal build up curve with an initial exponential build-up portion and a dynamic equilibrium portion, wherein the build-up curve has an uptake time constant associated with the initial build-up portion of the NMR signal, the method further comprising taking the ratio of the uptake time constant and mean transit time to determine whether the gas exchange is perfusion or diffusion limited. 73. A method according to claim 68, wherein the NMR signals include signals at a frequency associated with dissolved-phase 129Xe, the dissolved-phase signals having signal uptake curves with an initial exponential build-up portion and a dynamic equilibrium portion, the dynamic equilibrium portion of the uptake curves having an associated asymptote slope and intercept, wherein the slopes and intercepts of the asymptotes of the dissolved-phase 129Xe NMR uptake curves are used to calculate mean transit time, perfusion and relative blood volume. 74. A method according to claim 68, wherein at least one of the NMR signals is a dissolved-phase 129Xe NMR signal that has a signal uptake curve with an initial exponential build-up portion and a dynamic equilibrium portion in a time period of interest, the signal having an associated asymptote slope and intercept after the initial exponential ... | MEDI-PHYSICS, INC. | A61B005/055
US08821418 | B2 | 2014-09-02 | Monitoring, predicting and treating clinical episodes | 2009-05-10 | Apparatus is provided, including a sensor configured to sense motion of a subject and generate a motion signal in response thereto. The apparatus includes an output unit and a control unit. The control unit is configured to: (a) analyze the motion signal, (b) at a first time, determine that the motion signal is indicative that the subject is likely sleeping, (c) at a second time following the first time, determine that the motion signal is indicative that the subject is likely no longer sleeping, and (d) in response thereto, drive the output unit to alert a clinician if the subject is at an increased risk for falling out of a bed. Other applications are also described. | 1. Apparatus comprising: a sensor configured to sense motion of a subject and generate a motion signal in response thereto;an output unit; anda control unit configured to: analyze the motion signal,at a first time, determine that the motion signal is indicative that the subject is likely to be sleeping,at a second time following the first time, determine that the motion signal is indicative that the subject is likely to no longer be sleeping,in response thereto, determine that the subject is at an increased risk for falling out of a bed, andin response thereto, drive the output unit to alert a clinician that the subject is at an increased risk for falling out of the bed,the control unit being configured to adapt a threshold for the alerting of the clinician based on an input selected from the group consisting of: an input indicative of a characteristic of a medication administered to the subject, an input indicative of the subject having undergone surgery, and an input indicative of a time of day,wherein the sensor is further configured to generate a signal indicative of whether the subject is in or out of bed, andwherein the control unit is further configured to (a) drive the output unit to generate an alert if the subject has been out of bed for at least a defined period of time, and (b) withhold driving the output unit to generate the alert if the subject has been out of bed for less than one minute. 2. The apparatus according to claim 1, wherein the control unit is configured to determine, at the first time, that the subject is likely to be sleeping, based on reduced motion of the subject. 3. The apparatus according to claim 2, wherein the control unit is configured to identify the reduced motion as being indicative that the subject is likely to be sleeping, based on previously-measured motion patterns of the subject. 4. The apparatus according to claim 1, wherein the control unit is configured to determine, at the first time, that the subject is likely to be sleeping, based on the motion signal and a time of day. 5. The apparatus according to claim 1, wherein the sensor is configured to sense motion of the subject without contacting the subject or clothes the subject is wearing. 6. Apparatus comprising: a sensor configured to sense motion of a subject and generate a motion signal in response thereto;an output unit; anda control unit configured to: analyze the motion signal,at a first time, determine that the motion signal is indicative that the subject is likely to be sleeping,at a second time following the first time, determine that the motion signal is indicative that the subject is likely to no longer be sleeping,in response thereto, determine that the subject is at an increased risk for falling out of a bed, andin response thereto, drive the output unit to alert a clinician that the subject is at an increased risk for falling out of the bed,the control unit being configured to adapt a threshold for the alerting of the clinician base... | Earlysense Ltd. | A61B005/103 A61B005/117 A61B005/00 A61B005/0205 A61B005/024 G08B021/04 G08B021/22 A61B005/145 A61B007/00 A61B005/021 A61B005/0402 A61B005/08 A61B005/11 A61B005/0452 A61B005/20 A61G007/015
US08784323 | B2 | 2014-07-22 | Methods and systems that use implanted posture sensor to monitor pulmonary edema | 2009-12-30 | In specific embodiments, a method to monitor pulmonary edema of a patient, comprises (a) detecting, using an implanted posture sensor, when a posture of the patient changes from a first predetermined posture to a second predetermined posture, (b) determining an amount of time it takes an impedance signal to achieve a steady state after the posture of the patient changes from the first predetermined posture to the second predetermined posture, where the impedance signal is obtained using implanted electrodes and is indicative of left atrial pressure and/or intra-thoracic fluid volume of the patient, and (c) monitoring the pulmonary edema of the patient based on the determined amount of time it takes the impedance signal to achieve the steady state after the posture of the patient changes from the first predetermined posture to the second pre-determined posture. | 1. A method to monitor pulmonary edema of a patient, comprising: (a) detecting, using an implanted posture sensor, when a posture of the patient changes from a first predetermined posture to a second predetermined posture;(b) determining, based on an impedance signal, an amount of time it takes the impedance signal to achieve a steady state after the posture of the patient changes from the first predetermined posture to the second predetermined posture, where the impedance signal is obtained using implanted electrodes and is indicative of at least one of left atrial pressure or intra-thoracic fluid volume of the patient, andwhere the amount of time it takes the impedance signal to achieve a steady state, after the posture of the patient changes from the first predetermined posture to the second predetermined posture, is indicative of a degree of pulmonary edema; and(c) monitoring the pulmonary edema of the patient based on the determined amount of time it takes the impedance signal to achieve the steady state after the posture of the patient changes from the first predetermined posture to the second predetermined posture. 2. The method of claim 1 wherein the first predetermined posture is an upright posture and the second predetermined posture is a supine posture. 3. The method of claim 1 wherein the first predetermined posture is a supine posture and the second predetermined posture is an upright posture. 4. The method of claim 1, wherein: steps (a) and (b) are repeated a period of time thereafter; andstep (c) includes identifying when a degree of the pulmonary edema of the patient, or a change therein, exceeds a threshold; andone or both of triggering an alert and adjusting therapy administered to the patient in response to identifying that the degree of the pulmonary edema, or change therein, exceeds the threshold. 5. The method of claim 4, wherein steps (a), (b) and (c) are performed by an implanted system including the implanted posture sensor and implanted electrodes; and wherein when an alert is triggered, the method further comprises: (d) transmitting the alert from the implanted system to a non-implanted system; and(e) at least one of displaying a visual warning or issuing an audible warning by way of the non-implanted system in response to the alert. 6. The method of claim 4, wherein steps (a), (b) and (c) are performed by an implanted system including the implanted posture sensor and implanted electrodes; and wherein when therapy administered to the patient is adjusted, the method further comprises: (d) adjusting medication based on a physician pre-determined formula; and(e) displaying a visual indication of the medication adjustment to the patient. 7. The method of claim 4, wherein steps (a), (b) and (c) are performed by an implanted system including the implanted posture sensor and implanted electrodes; and wherein when therapy administered to the patient is adjusted, the method further comprises: (d) adjusting electrical thera... | Pacesetter, Inc. | A61B005/0215 A61B005/00 A61B005/053 A61B005/11
US08591412 | B2 | 2013-11-26 | Method and system for preventing virus-related obesity and obesity related diseases | 2010-11-18 | A method for preventing obesity related to infection by an adipogenic adenovirus includes obtaining a sample from a person, assaying the sample to determine whether the person has been previously infected with an adipogenic adenovirus, and if the person has not been previously infected, providing the person with at least one sensor positioned to detect when a person's hand approaches a predetermined distance from the person's face. By warning the person of undesired hand-to-face contacts, the person is able to reduce the incidence of obesity related infections. Other embodiments are directed to a kit for preventing obesity caused by infection with an adipogenic, adenovirus, such kit including a container for assaying an agent indicating the presence of antibodies to Ad-36, and a sensor positioned on an item selected from the group consisting of one of a hat, a writing instrument, eye glasses, a belt, sunglasses, a bra, a shirt, and a tie. | 1. A method for preventing obesity related to infection by an adipogenic adenovirus, comprising: obtaining a sample from a person;assaying the sample to determine whether the person has been previously infected with an adipogenic adenovirus;if said assaying step indicates that such person has not been previously infected with an adipogenic adenovirus, providing to the person at least one sensor positioned to detect when a person's hand approaches a predetermined distance from the person's face;reducing the incidence of hand-to-face contacts by said person due to the person being warned by said at least one sensor of occasions when the person's hands approach said person's face, said at least one sensor incorporated into the bill of a baseball cap;providing a signal generating device operatively associated with said at least one sensor and with a semiconductor device programmed to select one of a plurality of predetermined distances of proximity so that a person's hand triggers the signal generating device to warn the person that their hand is in the proximity of their mouth, said signal generating device employing at least a vibratory signal; andcommunicating from the sensor to a recording device, via a transmitter, the occurrence of said vibratory signal, said recording device recording the occurrence of said vibratory signal to indicate the number of hand-to-face contacts by said person, thereby reducing the risk that obesity related infections will be acquired by said person. 2. The method of claim 1, wherein said at least one sensor having an IR emitter, an IR receiver, a signal generating device, and a semiconductor device programmed to select a predetermined distance of proximity so that a person's hand triggers the signal generating device to warn the person that their hand is in the proximity of their month. 3. The method of claim 1, wherein said step of assaying the sample comprises the steps of: screening the sample for the presence of antibodies specific to the adipogenic adenovirus in the sample; and determining the presence of antibodies specific to the adipogenic adenovirus in the sample. 4. The method of claim 1, wherein the adipogenic adenovirus is selected from the group consisting of adenovirus type 5, adenovirus type 36, and adenovirus type 37. 5. The method of claim 1, wherein the adipogenic adenovirus is selected from the group consisting of adenovirus type 5 and adenovirus type 36. 6. The method of claim 1, wherein the step of assaying comprises determining the presence of antibodies specific to one or more peptides related to the adenovirus. 7. The method of claim 1, wherein said assaying step is performed by using a method selected from the group consisting of serum neutralization assay and ELISA. 8. The method of claim 1, wherein the sample is selected from the group consisting of a biological sample, body fluid, a tissue sample, an organ sample, feces, blood, saliva, and any combination thereof. 9. The method as ... | Nohands, LLC | G08B021/22 G08B013/08 A61N001/36 G08B021/24
US08747316 | B2 | 2014-06-10 | In vivo component measurement method, data processing method for in vivo component measurement, in vivo component measurement apparatus and collection member | 2011-01-28 | An in vivo component measurement method allowing how long a high concentration state of a measurement target component continues in an organism to be grasped is provided. In this in vivo component measurement method, a value relating to an amount of a measurement target component in tissue fluid extracted for 60 minutes or more from an organism on which a treatment for enhancing extraction of tissue fluid has been made is acquired. | 1. An in vivo component measurement method, comprising: preparing tissue fluid extracted for an extract time of 60 minutes or more from an organism on which a treatment for enhancing extraction of tissue fluid has been made;acquiring, using a detection portion of an in vivo component measurement apparatus, a value relating to an amount of a measurement target component in the extracted tissue fluid;acquiring a value relating to an amount of an electrolyte contained in the extracted tissue fluid; andacquiring, using an analysis portion of the in vivo component measurement apparatus, a value corresponding to an area under the curve (AUC) of a blood concentration time curve of the measurement target component corresponding to the extraction time of 60 minutes or more based on the value relating to the amount of the measurement target component, the value relating to the amount of the electrolyte, and the extraction time of 60 minutes or more. 2. The in vivo component measurement method according to claim 1, wherein the treatment is made by forming fine pores on a skin of the organism, andthe extraction is performed through the skin where the fine pores are formed. 3. The in vivo component measurement method according to claim 1, wherein the extraction is performed by extracting tissue fluid to an extraction medium containing a hypertonic aqueous solution having a higher osmotic pressure than pure water. 4. The in vivo component measurement method according to claim 3, wherein the hypertonic aqueous solution contains an auxiliary component different from the measurement target component, the auxiliary component being contained with the measurement target component in the tissue fluid. 5. The in vivo component measurement method according to claim 4, wherein the auxiliary component is at least one selected from a group consisting of kalium chloride, glycine and urea. 6. The in vivo component measurement method according to claim 5, wherein the auxiliary component is contained in a concentration of 0.2 mmol/L or more in the hypertonic aqueous solution. 7. The in vivo component measurement method according to claim 1, wherein the extraction time is 120 minutes or more. 8. The in vivo component measurement method according to claim 1, wherein the extraction time is 180 minutes or more. 9. The in vivo component measurement method according to claim 1, wherein the measurement target component is glucose. 10. The in vivo component measurement method according to claim 1, further comprising acquiring a value obtained by dividing the value corresponding to the AUC of the measurement target component by the extraction time. 11. The in vivo component measurement method according to claim 1, further comprising informing a subject of an end of extraction when the extraction time of 60 minutes or more elapses from a start of extraction of the tissue fluid. 12. An in vivo component measurement apparatus, comprising: a set portion for setting a collection... | Sysmex Corporation | A61B005/00
US08965532 | B2 | 2015-02-24 | Capacitive medical electrode | 2011-03-07 | Medical electrodes in embodiments of the teachings may include one or more of the following features: (a) a metallic conductor, (b) the metallic conductor sandwiched between a first dielectric layer adjacent a top surface of the metallic conductor and a second dielectric layer located on a bottom surface of the metallic conductor, (c) a conducive gel coating on at least one of the first and second dielectric layers, (e) the metallic conductor, the dielectric layers, and the conductive gel being wrapped to form a multi-tiered electrode having a plurality of conductive surfaces, (f) an adhesive adhering the metallic conductor with the dielectric layers, (g) a tab connector to provide a connection to electrical monitoring equipment and (h) an attachment connector to provide electrical connection with a patient. | 1. A medical electrode for transmitting an electrical signal to a patient, the electrode comprising: a conductive ink layer having a top surface and a bottom surface;a first dielectric film having an inner facing surface and an outer facing surface, with the inner facing surface operatively connected to the top surface of the conductive ink layer; anda layer of conductive substance located on the outer facing surface of the first dielectric film; anda second dielectric film having an inner facing surface and an outer facing surface, with the inner facing surface operatively connected to the bottom surface of the conductive ink layer and to the inner facing surface of the first dielectric film;wherein the layer of conductive substance is galvanically isolated from the conductive ink layer;wherein the conductive ink layer and the layer of conductive substance form a capacitance; andwherein the first dielectric film has a thickness that is greater than the conductive ink layer. 2. The electrode of claim 1, wherein the conductive ink layer has a thickness of approximately 0.3 mil (0.007 mm). 3. The electrode of claim 1, wherein the first dielectric film has a thickness of approximately 0.5 mil (0.012 mm). 4. The electrode of claim 1, wherein the second dielectric film has a thickness of approximately 0.5 mil (0.012 mm). 5. The electrode of claim 1, wherein the conductive ink layer, the first dielectric film and the second dielectric film have a combined thickness of approximately 1.3 mil (0.033 mm). 6. The electrode of claim 1, wherein the first and second dielectric films are adhesively attached to each other. 7. The electrode of claim 6, wherein the layer of conductive substance comprises adhesive gel. 8. The electrode of claim 1, wherein the conductive ink comprises metallic material. 9. A medical electrode suitable for capacitive coupling to a patient, the medical electrode comprising: a dielectric layer having a first side and a second side;conductive ink carried by the first side of the dielectric layer, with the conductive ink comprising a plurality of similarly shaped, sequentially arranged sections; anda conductive gel connected to and carried by the second side of the dielectric layer opposite said plurality of sections of conductive ink;in which capacitances formed between each section of the conductive ink and conductive gel coating combine to form a total capacitance of the electrode; andwherein the total capacitance is greater than 7 nanofarads. 10. The electrode of claim 9, wherein the conductive ink comprises metallic material. 11. The electrode of claim 9, wherein each section has a length of approximately 1.5 inches. 12. The electrode of claim 9, wherein each section has an area of approximately 2 square inches. 13. The electrode of claim 9, wherein at least one section has a capacitance of approximately 7 nanofarads. 14. The electrode of claim 9, wherein the total capacitance has a range of between approximately 7 nano... | - | A61N001/00
US08583215 | B2 | 2013-11-12 | Reduction of catheter electrode loading | 2009-11-05 | A method for position sensing includes inserting a probe having a first probe-electrode and a second probe-electrode into a body of a subject, and coupling body-surface electrodes to a surface of the body. Currents passing between the first probe-electrode and the body-surface electrodes are measured, using first circuitry coupled to at least the first probe-electrode and having a first electrical ground, and position coordinates of the probe are determined responsively to the measured currents. Second circuitry, having a second electrical ground, is coupled to at least the second probe-electrode, and the first electrical ground is isolated from the second electrical ground. | 1. A method for position sensing, comprising: inserting a probe comprising a first probe-electrode and a second probe-electrode into a body of a subject;coupling body-surface electrodes to a surface of the body;measuring, using first circuitry coupled to at least the first probe-electrode and having a first electrical ground, currents passing between the first probe-electrode and the body-surface electrodes, the probe and the plurality of body surface electrodes defining respective paths and an impedance along each respective path;determining position coordinates of the probe responsively to the measured currents using positioning circuitry, having a second electrical ground, coupled to at least the second probe-electrode;isolating the first electrical ground from the second electrical ground; andproviding a predetermined inter-ground coupling impedance coupled between the first electrical ground and the second electrical ground, the inter-ground coupling impedance being higher than the impedance along each respective path; andmaintaining accuracy of the position coordinates determination by reducing parasitic capacitance between the positioning circuitry and one or more of the first electrical ground and the second electrical ground. 2. The method according to claim 1, wherein the value of the inter-ground coupling impedance is selected so as to maximize an accuracy of determining the position coordinates. 3. The method according to claim 2, wherein the inter-ground coupling impedance is between 500 and 5000 Ohm. 4. The method according to claim 1, wherein inserting the probe comprises passing the probe into a heart of the subject, and wherein coupling the second circuitry comprises measuring an electrical activity of the heart using at least the second probe-electrode. 5. The method according to claim 4, wherein both of the first and second probe-electrodes are coupled for use in both determining the position coordinates and measuring the electrical activity. 6. The method according to claim 1, wherein the first circuitry comprises a front-end comprising an isolation transformer having a primary winding and a secondary winding, which is coupled to at least the first probe-electrode, and wherein isolating the first electrical ground from the second electrical ground comprises coupling the secondary winding of the isolation transformer to the first electrical ground while the primary winding is coupled to the second electrical ground. 7. The method according to claim 1, wherein the probe has a front end, and wherein the front end of the probe measures the currents having an output impedance greater than 100,000 Ohm to transmit the currents through at least the first probe-electrode. 8. A medical system comprising: a probe adapted to be inserted into a body of a subject, the probe comprising a first probe-electrode and a second probe-electrode;a plurality of body-surface electrodes, which are adapted to be fixed to a surface of the body at ... | Biosense Webster (Israel) Ltd. | A61B005/05
US09675268 | B2 | 2017-06-13 | Detection and differentiation of sleep disordered breathing | 2010-11-05 | A vector method for monitoring a subject's sleep-disordered breathing utilizing a single, anatomy-attached (anatomically outside or implanted inside), three-orthogonal-axis accelerometer(s), including the steps of (1) collecting from a sleeping subject three-orthogonal-axis data relating to at least one of sound data, subject posture, subject activity, snoring, and respiration, and (2) following such collecting, processing and analyzing collected data to detect associated, disordered breathing including assessing the presence of at least one of (a) sleep-disordered breathing generally, (b) sleep apnea specifically, and (c) differentiation between central and obstructive sleep apnea. Further involved is the acquiring of ECG data, and that the mentioned processing and analyzing include recognition of such acquired ECG data. | 1. An external-only method for detecting and differentiating a sleeping subject's manifestations of sleep-disordered breathing comprising attaching a single, three-axis accelerometer to a subject's chest,acquiring data from the three axis accelerometer during a subject-recording period,identifying, based on accelerometer data, sleep times during the subject-recording period when the subject is asleep, and generating, based on accelerometer data, a wake summary,during the identified sleep times, utilizing solely the attached accelerometer, simultaneously collecting, externally-only, sleeping-subject-specific, three-axis, vector-mechanical accelerometer data,filtering the accelerometer data through five independent facet-specific signal filters outside the subject to generate components specific to (a) heart sounds, (b) body position, (c) body activity, (d) snoring, and (e) chest-movement-based respiration,collecting sleeping-subject-specific ECG data outside the subject, andgenerating a sleep summary based on the wake summary including processing and analyzing the components of the accelerometer data and the ECG data outside the subject, to detect and differentiate the presences of (a) sleep-disordered breathing generally, (b) sleep apnea generally, (c) differentiated central and obstructive sleep apnea specifically, and (d) hypopnea specifically. 2. The external-only method of claim 1, wherein attaching a single, three-axis accelerometer to a subject's chest is adjacent a V4, precordial ECG site. | Inovise Medical, Inc. | A61B005/08 A61B005/0456 A61B005/113 A61B005/00 A61B007/00 A61B005/11
US10194863 | B2 | 2019-02-05 | Integrated transmitter unit and sensor introducer mechanism and methods of use | 2011-02-07 | Method and apparatus for providing an integrated transmitter unit and sensor insertion mechanism is provided. | 1. A system, comprising: an integrated housing including sensor electronics operatively coupled to an analyte sensor, a portion of the analyte sensor extendable through the integrated housing at an insertion position selected from a plurality of insertion positions, wherein the insertion positions are each defined by a discrete aperture and spaced apart in a base portion of the integrated housing, and wherein each insertion position corresponds to a different insertion site across a skin surface indicated by the integrated housing,wherein the integrated housing is configured to mount across the skin surface at the plurality of insertion positions, andwherein a transmitter unit of the sensor electronics is configured to be selectively disabled after a predetermined time period has elapsed since coupling to the analyte sensor; anda sensor introducer assembly coupled to the integrated housing, the introducer assembly configured to position at least the portion of the analyte sensor under the skin surface at the selected insertion position. 2. The system of claim 1, wherein the sensor electronics includes the transmitter unit coupled to and in electrical contact with the analyte sensor, and configured to transmit one or more signals corresponding to a respective one or more signals received from the analyte sensor. 3. The system of claim 2, wherein the transmitter unit is configured to wirelessly transmit the one or more signals corresponding to the respective one or more signals received from the analyte sensor. 4. The system of claim 2, further including a receiver unit configured to receive the one or more signals transmitted by the transmitter unit. 5. The system of claim 4, wherein the receiver unit includes one of an infusion pump, a monitoring device, a personal digital assistant, a pager, or a mobile telephone. 6. The system of claim 2, wherein the one or more signals received from the analyte sensor substantially corresponds to a monitored analyte level under the skin surface. 7. The system of claim 1, wherein the introducer assembly is configured to be removably detached from the integrated housing. 8. The system of claim 1, wherein the analyte sensor is a glucose sensor. 9. The system of claim 1, wherein the predetermined time period comprises about 1 day or more. 10. The system of claim 1, wherein the introducer assembly is configured to be removed from the integrated housing after positioning of the at least the portion of the analyte sensor under the skin surface. 11. The system of claim 1, wherein the integrated housing is flexible to substantially conform to a shape of a body portion to which it is configured to be mounted. 12. The system of claim 1, wherein the analyte sensor comprises a plurality of electrodes including a working electrode, wherein the working electrode comprises an analyte-responsive enzyme and a mediator, wherein at least one of the analyte-responsive enzyme and the mediator is chemically bonded to a po... | ABBOTT DIABETES CARE INC. | A61B005/00 A61B005/145
US20110106200 | A1 | 2011-05-05 | STROKE RISK MONITORING SYSTEM INCLUDING IMPLANTABLE MEDICAL DEVICE | 2009-10-29 | One or more example techniques for monitoring the stroke risk of a patient via a system including an implantable medical device. In some examples, a method including monitoring at least one physiological parameter of a patient via an implantable medical device; determining whether each of a plurality stroke risk factors are present based at least in part on the at least one physiological parameters monitored via the implantable medical device; and generating a stroke risk score based on the stroke risk factors determined to be present, wherein the stroke risk score is reflective of the patient's risk of stroke. | 1. A method comprising: monitoring at least one physiological parameter of a patient via an implantable medical device;determining whether each of a plurality stroke risk factors are present based at least in part on the at least one physiological parameters monitored via the implantable medical device; andgenerating a stroke risk score based on the stroke risk factors determined to be present, wherein the stroke risk score is reflective of the patient's risk of stroke. 2. The method of claim 1, wherein the plurality of stroke risk factors include one or more of atrial fibrillation, hypertension, congestive heart failure, patient age above threshold, diabetes, prior stroke and prior transient ischemic attack. 3. The method of claim 1, further comprising: comparing the stroke risk score to a threshold value; andgenerating an indicator indicative of the stroke risk score based on the comparison. 4. The method of claim 3, further comprising presenting the indicator to a user via a user interface. 5. The method of claim 1, further comprising adjusting the delivery of electrical stimulation therapy to the patient based at least in part on the stroke risk score. 6. The method of claim 1, wherein the implantable medical device comprises an implantable cardiac device configured to deliver electrical stimulation to a heart of the patient. 7. The method of claim 1, wherein the one or more physiological parameters includes at least one of electrical activity of a heart of the patient, blood pressure of the patient, blood flow of the patient, blood sugar level of the patient, and intrathoracic impedance of the patient. 8. The method of claim 1, wherein the stroke risk score comprises a numerical value. 9. The method of claim 8, wherein the plurality of stroke risk factors are each associated with a stroke risk value used to generate the stroke risk score. 10. The method of claim 1, further comprising detecting the presence of atrial fibrillation, wherein the stroke risk score is generated based on both the stroke risk factors determined to be present and detecting the presence of atrial fibrillation. 11. The method of claim 1, wherein the stroke risk score is based on the presence of the two or more stroke risk factors. 12. A stroke risk monitoring system comprising: an implantable medical device including a sensing module configured to monitor at least one physiological parameter of a patient; anda processor configured to determine whether each of a plurality of stroke risk factors are present based at least in part the at least one physiological parameter, and generate a stroke risk score based on the stroke risk factors determined to be present,wherein the stroke risk score is reflective of the patient's risk of stroke. 13. The system of claim 12, wherein the plurality of stroke risk factors include one or more of atrial fibrillation, hypertension, congestive heart failure, patient age above threshold, diabetes, prior stroke and prior transien... | Medtronic, Inc. | A61N001/365 A61B005/0402
US20110124993 | A1 | 2011-05-26 | Analyte Sensors Comprising Self-Polymerizing Hydrogels | 2009-11-24 | Generally, embodiments of the present disclosure relate to analyte determining methods and devices (e.g., electrochemical analyte monitoring systems) that have improved uniformity of distribution of the sensing layer by inclusion of a self-polymerizing hydrogel, where the sensing layer is disposed proximate to a working electrode of in vivo and/or in vitro analyte sensors, e.g., continuous and/or automatic in vivo monitoring using analyte sensors and/or test strips. Also provided are systems and methods of using the, for example electrochemical, analyte sensors in analyte monitoring. | 1. An analyte sensor comprising: a working electrode;a counter electrode; anda sensing layer disposed proximate to the working electrode, wherein the sensing layer comprises a self-polymerizing hydrogel. 2. The analyte sensor of claim 1, wherein at least a portion of the analyte sensor is adapted to be subcutaneously positioned in a subject. 3. The analyte sensor of claim 1, wherein the sensing layer has a substantially arcuate profile as measured using a profilometer. 4. The analyte sensor of claim 1, further comprising a membrane disposed over the sensing layer. 5. The analyte sensor of claim 1, wherein the self-polymerizing hydrogel comprises a polyethylene glycol hydrogel or polyethylene glycol derivative hydrogel. 6. The analyte sensor of claim 1, wherein the self-polymerizing hydrogel comprises polyethylene glycol diacrylate. 7. The analyte sensor of claim 1, wherein the self-polymerizing hydrogel comprises an activating agent and an initiator. 8. The analyte sensor of claim 7, wherein the activating agent comprises ferrous gluconate. 9. The analyte sensor of claim 7, wherein the initiator comprises hydrogen peroxide. 10. The analyte sensor of claim 1, wherein the self-polymerizing hydrogel comprises a crosslinker. 11. The analyte sensor of claim 10, wherein the crosslinker comprises a 2-arm polyethylene glycol acrylate or a 4-arm polyethylene glycol acrylate. 12. The analyte sensor of claim 11, wherein the crosslinker has a molecular weight of 1,000 to 20,000 Da. 13. The analyte sensor of claim 1, wherein the sensing layer comprises a swelling modulator. 14. The analyte sensor of claim 13, wherein the swelling modulator comprises a 3-arm acrylate. 15. The analyte sensor of claim 13, wherein the swelling modulator comprises trimethylol propane triacrylate. 16. The analyte sensor of claim 1, wherein the self-polymerizing hydrogel is adapted to polymerize at room temperature. 17. The analyte sensor of claim 1, wherein the sensing layer comprises a glucose-responsive enzyme. 18. The analyte sensor of claim 17, wherein the sensing layer comprises a redox mediator. 19. The analyte sensor of claim 18, wherein the redox mediator comprises a ruthenium-containing complex or an osmium-containing complex. 20. The analyte sensor of claim 18, wherein the analyte sensor is a glucose sensor. 21. The analyte sensor of claim 18, wherein the analyte sensor is an in vivo sensor. 22. The analyte sensor of claim 18, wherein the analyte sensor is an in vitro sensor. 23.-68. (canceled) | Abbott Diabetes Care Inc. | A61B005/1468 G01N027/26
US20110152658 | A1 | 2011-06-23 | IDENTIFICATION OF ABERRANT MEASUREMENTS OF IN VIVO GLUCOSE CONCENTRATION USING TEMPERATURE | 2010-12-17 | Disclosed herein are methods and systems for generating an estimate of an in vivo analyte concentration and identifying whether the estimate is aberrant. In some embodiments, the system includes a sensor comprising an analyte sensor and a temperature sensing element, and a control unit programmed to identify changes in temperature that may indicate a non-physiologic condition (and result in an aberrant glucose measurement). In some embodiments, the methods include generating an estimate of analyte concentration at a particular time using the analyte sensor, and generating first and second signals indicative of temperature using the temperature sensing element. In some embodiments the methods include identifying the estimate of analyte concentration as aberrant if the magnitude of the difference between the first and second signals indicative of temperature exceeds a threshold value. | 1. A method of identifying whether an intravascular glucose measurement is aberrant, the method comprising: deploying a sensor within a blood vessel, wherein the sensor comprises along a distal region thereof, a temperature-sensing element adapted to measure the temperature of the blood contacting the distal region, and a glucose-sensing chemical indicator system adapted to measure the glucose concentration in the blood contacting the distal region;generating at a first time, a first temperature measurement (TO;generating at a second time, a second temperature measurement (T2) and a glucose measurement;comparing T1 and T2; andidentifying the glucose measurement as aberrant if the absolute magnitude of the difference between T1 and T2 exceeds a pre-selected threshold value. 2. The method of claim 1, further comprising alerting a user, alerting a monitor, and/or signaling a medication delivery device when an aberrant glucose measurement is identified. 3. The method of claim 1, wherein the threshold value is 0.1° C. 4. The method of claim 1, wherein the threshold value is 0.2° C. 5. The method of claim 1, wherein the threshold value is 0.3° C. 6. The method of claim 1, wherein the threshold value is 0.4° C. 7. A method of identifying whether an intravascular glucose measurement is aberrant at a given time, the method comprising: deploying a sensor within a blood vessel, wherein the sensor comprises along a distal region thereof, a temperature-sensing element adapted to measure the temperature of the blood contacting the distal region, and a glucose-sensing chemical indicator system adapted to measure the glucose concentration in the blood contacting the distal region;generating one or more temperature measurements prior to the given time;generating simultaneously, at the given time, a temperature measurement and a glucose measurement;comparing the temperature measurement generated at the given time with an average of the temperature measurements generated prior to the given time; andidentifying the glucose measurement generated at the given time as aberrant if the absolute magnitude of the difference between the temperature measurement generated at the given time and the average of the temperature measurements generated prior to the given time exceeds a pre-selected threshold value. 8. A system for measuring an in vivo glucose concentration and identifying whether the measurement is aberrant, the system comprising: a glucose sensor comprising a chemical indicator system configured to generate a signal indicative of the in vivo concentration of the glucose at a first time;a temperature sensing element configured to generate a plurality of signals comprising: a first signal indicative of temperature at a second time prior to the first time; anda second signal indicative of temperature at the first time;at least one sensor control unit, the at least one sensor control unit configured to control the operation of the glucose sensor and the tempera... | GLUMETRICS, INC. | A61B005/145
US20110112426 | A1 | 2011-05-12 | Brain Activity as a Marker of Disease | 2009-11-10 | A method of monitoring brain activity is provided, wherein the method includes receiving a signal associated with neuronal activity of a mammalian brain. The method also includes processing the signal using a linear, non-linear, or combination algorithm to extract a signal feature. A neuromarker may be determined based on an association between the signal feature and a library of features, wherein the library includes a plurality of signal features correlated with a plurality of disease states. | 1. A method of monitoring brain activity, comprising: receiving a signal of brain electrical activity from a patient;processing the signal using at least one processor implementing an algorithm to extract a signal feature based on a brain state; anddetermining a neuromarker using at least one processor, wherein the neuromarker is determined based on an association between the signal feature and a disease state. 2. The method of claim 1, wherein the neuromarker corresponds to a particular state of progression of disease 3. The method of claim 1, wherein the neuromarker corresponds to a particular disease stage. 4. The method of claim 1, wherein the neuromarker represents a discrete change of state in disease progression. 5. The method of claim 1, wherein the neuromarker represents an indicator of an appropriate time to apply treatment. 6. The method of claim 1, wherein the neuromarker represents an indicator of effectiveness of applied treatment. 7. The method of claim 1, wherein the algorithm is based on at least one of wavelet analysis, diffusion geometric analysis, fractal analysis, and spectral analysis. 8. The method of claim 1, wherein the neuromarker is correlated with a traditional disease state marker. 9. The method of claim 8, wherein the correlation uses a mutual information algorithm. 10. A method of monitoring brain activity, comprising: receiving a signal of brain electrical activity from a patient;processing the signal using at least one processor implementing a non-linear algorithm to extract a signal feature; anddetermining a neuromarker using at least one processor based on an association between the signal feature and a library of features stored in a memory, wherein the library includes a plurality of signal features correlated with a plurality of disease states. 11. The method of claim 10, wherein the neuromarker includes a discrete neuromarker. 12. The method of claim 10, wherein the signal includes an electro-encephalography signal. 13. The method of claim 10, wherein the non-linear algorithm is based on at least one of wavelet analysis, diffusion geometric analysis, fractal analysis, and spectral analysis. 14. The method of claim 10, further including pre-processing the signal, wherein pre-processing includes at least one of denoising, filtering, windowing, sampling, and digitizing. 15. The method of claim 10, wherein the library of features is at least partially determined using a genetic algorithm. 16. The method of claim 10, further including applying a treatment to the patient. 17. A system configured to display a neuromarker, comprising: a receiver configured to receive a signal associated with neuronal activity of a patient's brain;a processor configured to process the signal using a non-linear algorithm to extract a signal feature and determine a neuromarker based on an association between the signal feature and a library of features;a storage system configured to store the library of features, wherei... | BrainScope Company, Inc. | A61B005/0476
US07946994 | B2 | 2011-05-24 | Compact apparatus and methods for non-invasively measuring hemodynamic parameters | 2004-10-07 | Improved apparatus and methods for non-invasively assessing one or more hemodynamic parameters associated with the circulatory system of a living organism. In one aspect, the invention comprises spatially compact “bracelet” embodiment adapted to accurately place and maintain a sensor (e.g., tonometric pressure sensor) with respect to the anatomy of the subject, including an optional alignment apparatus which moveably captures the sensor to, inter alia, facilitate coupling thereof to an actuator used to position the sensor during measurements. The alignment apparatus also advantageously allows the sensor position to be maintained when the fixture is removed from the subject, such as during patient transport. A completely autonomous variant of the bracelet apparatus having internal power supply and wireless interfaces is also disclosed. Methods for positioning the alignment apparatus and sensor and providing treatment to the subject are also described. | 1. Physiologic sensing apparatus, comprising: at least one pressure sensor;a substantially arcuate-shaped support structure adapted to receive at least a portion of the anatomy of a subject therein, said support structure comprising: a first portion;a second portion; anda third portion coupled to said first and second portions and comprising an actuator mechanism, said first, second and third portions cooperating to surround a substantial fraction of the circumference of said anatomy; anda preload mechanism adapted to trigger locking of said substantially arcuate-shaped support structure when a predetermined level of compressive force is applied to said support structure;wherein said support structure is adapted to couple at least one force felt by said at least one sensor to said at least portion of said anatomy; andwherein upon actuation of said actuator mechanism, said first, second and third portions are kept in substantially constant relationship to one another. 2. The apparatus of claim 1, wherein said actuator is configured to controllably move said at least one pressure sensor. 3. The apparatus of claim 2, wherein said controllable movement of said at least one pressure sensor comprises applanating said at least one pressure sensor against the tissue of a limb of said subject proximate a blood vessel thereof. 4. The apparatus of claim 2, wherein said controllable movement of said at least one pressure sensor comprises movement in at least one of a lateral and/or proximal direction relative to a blood vessel of a limb of said anatomy. 5. The apparatus of claim 1, wherein said at least one pressure sensor is configured to generate first signals relating to pressure applied thereto. 6. The apparatus of claim 5, further comprising a signal interface adapted to transmit second signals to an external device, said second signals comprising scaled blood pressure signals generated based at least in part on said first signals. 7. The apparatus of claim 6, wherein said scaled blood pressure signals are scaled at least in part using at least one physiologic parameter associated with said subject. 8. An apparatus for performing physiologic sensing on a living subject, said apparatus comprising: at least one pressure sensor, said pressure sensor being configured to generate first signals relating to pressure applied thereto;a substantially U-shaped body adapted to receive at least a portion of the anatomy of said living subject therein; andan actuator apparatus coupled to said at least one pressure sensor and said body, said actuator being adapted to controllably move said at least one pressure sensor;wherein said apparatus is configured to lock said body and prevent rotation of one or more components thereof when a predetermined level of compressive force is applied thereto. 9. The apparatus of claim 8, wherein said body comprises first, second and third articulated portions, said first, second and third articulated portions cooperating to sur... | Tensys Medical, Inc. | A61B005/02
US07953477 | B2 | 2011-05-31 | Determination of sympathetic activity | 2009-02-06 | An apparatus, a method and a computer program for the determination of sympathetic activity are disclosed. The method comprises: receiving heart activity data of a person; determining a recovery heart rate change speed during a recovery heart rate period in the heart activity data; and determining a sympathetic activity state of the sympathetic nervous system of the person based on the determined recovery heart rate change speed. | 1. An apparatus comprising: an input interface configured to receive heart activity data of a person; anda processor configured to determine a recovery heart rate change speed during a recovery heart rate period in the heart activity data after a training period, andto determine a sympathetic activity state of the sympathetic nervous system of the person before the training period based on the determined recovery heart rate change speed. 2. The apparatus of claim 1, wherein the processor is further configured to store a plurality of predetermined sympathetic activity states, each predetermined sympathetic activity state being associated with a predetermined range of the recovery heart rate change speed, and to determine the sympathetic activity state by checking which predetermined range of the stored sympathetic activity state corresponds to the determined recovery heart rate change speed. 3. The apparatus of claim 1, wherein the processor is further configured to determine a plurality of recovery heart rate periods and a plurality of sympathetic activity states based on the plurality of recovery heart rate periods, and to determine the sympathetic activity state by utilizing a statistical analysis on the plurality of sympathetic activity states or by choosing the sympathetic activity state indicating the lowest sympathetic activity among the plurality of sympathetic activity states. 4. The apparatus of claim 1, wherein the processor is further configured to automatically detect (1) the recovery heart rate period from the heart activity data, and (2) detect the recovery heart rate period from the heart activity data on the basis of a recovery input received via a recovery input interface, the apparatus comprising the recovery input interface. 5. The apparatus of claim 4, wherein the processor is further configured to automatically detect the recovery heart rate period so that the recovery heart rate period starts after a training period fulfilling a predetermined condition for the heart rate. 6. The apparatus of claim 5, wherein the predetermined condition for the heart rate relates to the maximum heart rate of the person or to the rest heart rate of the person. 7. The apparatus of claim 5, wherein the processor is further configured to automatically detect the recovery heart rate period after the training period so that the recovery heart rate period starts after the training period when the heart rate falls at a higher rate than a predetermined speed or within a predetermined time limit below a predetermined limit. 8. The apparatus of claim 5, wherein the apparatus further comprises a motion input interface configured to receive motion data of the user, and the processor is further configured to detect the recovery heart rate period after the training period so that the recovery heart rate period starts after the training period when the motion data indicates that the person has stopped the training period. 9. The apparatus of claim 4,... | Polar Electro Oy | A61B005/04
US07959566 | B2 | 2011-06-14 | Methods for delivering a drug to a patient while restricting access to the drug by patients for whom the drug may be contraindicated | 2006-05-19 | Methods for delivering a drug to a patients in need of the drug, while restricting access to the drug by patients for whom the drug may be contraindicated are disclosed. The methods are of the type in which prescriptions for the drug are filled by a pharmacy only after a computer readable storage medium has been consulted to retrieve a prescription approval code. Embodiments are provided wherein the patients are assigned to risk groups based upon the risk that taking the drug will lead to an adverse side effect, and certain additional information, such as periodic surveys and diagnostic tests probative of the ongoing risk of the side effect developing are obtained before prescriptions for the drug are approved. | 1. A method of treating a male patient, suffering from erythema nodosum leprosum, with thalidomide, by permitting prescriptions for thalidomide to be filled by a pharmacy only after the pharmacy has retrieved a prescription approval code for the prescription, comprising the following steps: (a) via a non-transitory computer readable storage medium, registering a prescriber and the pharmacy with a distributor of thalidomide;(b) determining whether the patient is able to understand and carry out instructions;(c) upon determination that the patient is able to carry out the instructions, providing verbal and written warnings of the hazard of taking thalidomide and exposing fetus to the drug;(d) further providing verbal and written warnings of the risk of possible contraception failure and of the need to use barrier contraception when having sexual intercourse with women of child bearing potential;(e) obtaining acknowledgement of said warnings from the patient;(f) providing the prescription for thalidomide to the patient from the prescriber;(g) via a non-transitory computer readable storage medium, registering the patient with the distributor;(h) upon obtaining the acknowledgement and registrations, generating the prescription approval code, and via a non-transitory computer readable storage medium, retrieving the prescription approval code by the pharmacy before the prescription is filled; and(i) upon retrieving the prescription approval code, administering thalidomide to the patient. 2. The method of claim 1, wherein the acknowledgement requires the patient's acknowledgement of one or more of the following: (a) the understanding that thalidomide must not be taken if unprotected sex cannot be avoided;(b) the understanding of potential birth defects;(c) that the patient has been advised of the need for barrier contraception by the prescriber;(d) the obligation to inform the prescriber if the patient's sexual partner is suspected of becoming or being pregnant;(e) that thalidomide is solely for the use of the patient himself and must not be shared with any other person;(f) that the patient has read the information brochure or viewed the information film on thalidomide;(g) that the semen or blood must not be donated during the thalidomide treatment;(h) that all of the patient's inquiries regarding thalidomide treatment have been answered by the prescribing physician; or(i) the patient's understanding that participation in a survey and patient registry is required during the thalidomide treatment. 3. The method of claim 1 further comprising providing the patient, prior to generation of the approval code, with warnings of the side effects associated with administration of thalidomide, wherein said side effects are non-teratogenic side effects. 4. The method of claim 1 further comprising obtaining a written authorization by the prescriber prior to generation of the approval code. 5. The method of claim 1, wherein the acknowledgement is a written informe... | Celgene Corporation | A61B005/00
US08311618 | B2 | 2012-11-13 | Method and apparatus for quantitative assessment of cardiac electrical events | 2011-01-03 | Method and apparatus for computer enabled analysis of ECG (electro cardiographic) data by exploiting computerized three-dimensional spatial presentation of the measured data using vectors. A three-dimensional presentation of the human heart may be correlated with waveforms specific for standard ECG or derived ECG signals based on the dipole approximation of the heart electrical activity. Additional tools for analyzing ECG data are also provided which may be used to determine the time of cardiac electrical events, to select specific beats for automated cardiac interval determination, and to flag ECGs that have been evaluated by automated means but may benefit by human reading. | 1. A computer-assisted method for selecting ECG complexes for further evaluation, the method comprising: receiving a recorded patient ECG signal comprising a plurality of m ECG complexes; andusing a processor to evaluate ECG complexes of the plurality of m ECG complexes, and to identify a subset of n ECG complexes of the plurality for further evaluation, where n is greater than two and less than m, and wherein the n ECG complexes of the identified subset have a highest degree of similarity to one another relative to any other possible subset of n ECG complexes of the plurality of m ECG complexes. 2. The method of claim 1, wherein the identified subset of n ECG complexes is identified by applying an autocorrelation function to the m ECG complexes of the plurality of ECG complexes. 3. The method of 2, wherein the autocorrelation function is∫-∞∞⁢f⁡(t+τ)⁢f_⁡(t)⁢ⅆt⁢⁢or⁢⁢∫-∞∞⁢f⁡(t)⁢f_⁡(t-τ)⁢ⅆtwhere τ is an increment of time t, f is a function and f is the complex conjugate of function f. 4. The method of claim 1, wherein the identified subset of ECG complexes is identified from virtual ECG data derived from the recorded ECG signal. 5. The method of claim 4, wherein the virtual ECG data is vector magnitude data or normalized vector magnitude data. 6. The method of claim 1, wherein the identified subset of n ECG complexes is a first identified subset of n ECG complexes, the method further comprising: determining a first value of a complex dissimilarity factor for the first identified subset of n ECG complexes;identifying from the plurality of m ECG complexes a second subset of n ECG complexes, wherein the second identified subset has at least one ECG complex that than is not part of the first identified subset;determining a second value of a complex dissimilarity factor for the second identified subset; andif the second value is less than the first value, selecting the second identified subset for further evaluation instead of the first identified subset. 7. A computer-assisted method for selecting ECG complexes for further evaluation, the method comprising: receiving a recorded patient ECG signal comprising a plurality of m ECG complexes, each ECG complex of the plurality having a duration; andusing a processor to evaluate ECG complexes of the plurality of m ECG complexes, and to identify a subset of n ECG complexes for further evaluation, where n is greater than two and less than m, and wherein the n ECG complexes of the identified subset have a lowest standard deviation in duration out of all possible subsets of n ECG complexes of the plurality of m ECG complexes. 8. A computer-assisted method for selecting ECG complexes upon which to evaluate cardiac events, the method comprising: receiving a recorded patient ECG signal comprising a plurality of m ECG complexes;using a processor to evaluate a respective same electrical event in each ECG complex of the plurality of m ECG complexes; andusing a processor to identify a subset of n ECG complexes of... | New Cardio, Inc. | A61N005/02 A61N005/04
US07945307 | B2 | 2011-05-17 | Marker delivery system with obturator | 2006-08-04 | A marker delivery device is described which has an obturator with an elongated shaft, an inner lumen, a proximal end, and a substantially sealed distal end. One or more tissue markers are deployed within the inner lumen of the elongated shaft of the obturator. Preferably, the tissue marker(s) is disposed within an inner lumen of a marker delivery tube which is disposed within the inner lumen of the elongated shaft of the obturator. The marker delivery tube has an opening for discharging the tissue markers into a body (e.g. biopsy) cavity. The distal tip of the marker delivery tube is configured to penetrate the substantially sealed distal end of the obturator so that tissue markers can be delivered while the obturator is in place within the body. Preferably, the obturator includes a detectable element capable of producing a relatively significant image signature during MRI. | 1. An obturator, comprising: a. an elongated shaft comprising a cylindrical wall having a proximal end, a distal end and defining an internal lumen extending to the distal end and an obturator tip secured to the distal end; andb. a detectable element which is in contact with the mating surfaces of the terminal distal end of the cylindrical wall and the obturator tip, said detectable element comprising a material selected from the group consisting of metal, ceramic, metal-filled plastic, a radiopaque material, or a mineral-filled plastic and being configured to produce a significant magnetic resonance image signature. 2. The obturator of claim 1 wherein the distal end of the cylindrical wall has at least one aperture and the obturator tip has at least one raised tab configured to fit within said at least one aperture. 3. The obturator of claim 1 wherein the detectable element is a ring which is disposed at the junction between the obturator tip and the distal end of the cylindrical wall of the elongated shaft of the obturator. | Senorx, Inc. | A61B005/05
US07957793 | B2 | 2011-06-07 | Methods for identifying neuronal spikes | 2005-12-22 | A method for identifying neuronal spikes (extracellular action potentials) is described wherein measured microelectrode readings from tissue are reviewed to identify spikes (successive readings having prolonged rises and/or falls). The frequency of such spikes as a function of their amplitude assumes a bimodal distribution wherein higher amplitude spikes represent neuronal spikes (signal) and lower amplitude spikes represent noise, and thus the higher amplitude spikes can be assumed to be neuronal spikes. Neuronal spikes from the same neuron can then be assumed to have substantially the same waveform shape and period, with the only significant difference between them being the scaling of their amplitudes (i.e., the amplitudes of spikes from the same neuron tend to be proportionate at any given time along their period). Thus, by testing identified neuronal spikes for matching timing and for proportional amplitudes, the neuronal spikes may further be identified as coming from the same or different neurons. | 1. A method of identifying neuronal spikes comprising the steps of: a. obtaining amplitude readings sampled over time from an electrode situated within tissue;b. identifying series of successive readings having at least one of (1) continuously increasing amplitude, and(2) continuously decreasing amplitude, each such series of readings representing a spike candidate;c. identifying spikes within the spike candidates, the spikes being chosen from the spike candidates having greatest amplitude; andd. performing at least one of the steps of: (1) placing a therapeutic electrode within the tissue, and(2) delivering electrical stimulation from a therapeutic electrode to the tissue, in dependence on the identified spikes. 2. The method of claim 1 wherein the step of identifying spikes within the spike candidates includes the step of choosing as spikes at least some of the spike candidates having amplitudes greater than the most frequently occurring amplitudes of the spike candidates. 3. The method of claim 2 wherein the spike candidates chosen as spikes have amplitudes at least twice as great as the most frequently occurring amplitudes of the spike candidates. 4. The method of claim 1 wherein each series of successive readings includes at least three readings. 5. The method of claim 1 wherein the step of identifying spikes within the spike candidates includes the step of: a. identifying the lowest amplitude about which the greatest number of spike candidates occur;b. setting a spike threshold amplitude greater than this lowest amplitude;c. identifying as spikes at least some of the spike candidates having amplitudes greater than the spike threshold amplitude. 6. The method of claim 5 further comprising the step of assigning a start time to each identified spike, each spike having a start time corresponding to a zero amplitude for the spike occurring prior to the time when the spike attains the spike threshold amplitude. 7. The method of claim 6 where the start time corresponds to the second zero amplitude occurring prior to the time when the spike attains the spike threshold amplitude. 8. The method of claim 1 wherein the step of identifying spikes within the spike candidates further includes the steps of: a. identifying the two largest groups of spike candidates wherein: (1) each group has at least substantially similar amplitude; and(2) one group has larger amplitude than the other;b. choosing as spikes at least some of the spike candidates in the group having larger amplitude. 9. The method of claim 1 wherein the step of identifying spikes within the spike candidates further includes the steps of: a. identifying two groups of spike candidates wherein: (1) each group has high incidences of spike candidates with at least substantially similar amplitudes, and(2) one group has a median amplitude at least twice as great as the median amplitude of the other group; andb. choosing as spikes at least some of the spike candidates in the group having the ... | Wisconsin Alumni Research Foundation | A61B005/04
US07949545 | B1 | 2011-05-24 | Method and apparatus for providing a centralized medical record system | 2005-05-03 | An embodiment of the present invention provides a method of creating and maintaining a centralized medical record system. The method of this embodiment includes establishing, in a computer system, a record associated with a client. (The computer system is in communication with a network to which the client has access and is capable of storing records from a multiplicity of clients having medical information developed at a multiplicity of sources). The method also includes receiving medical data from a source pursuant to a written request by the client to the source for transfer of the medical data to the centralized medical record system. Also the method includes storing and logically associating a representation of the medical data with the record and providing the client with access to the record over the network such that the client may review the representation of the medical data. | 1. A method of creating and maintaining a centralized medical record system comprising: establishing, in a computer system, a record associated with an account of a client, wherein the client is a patient or a guardian of the patient, the computer system being in communication with a network to which the client has access;receiving a copy of at least one medical record or item of data, from one or more of a multiplicity of health care entity sources, sent to the centralized medical record system pursuant to a written request by the client to each of the health care entity sources;producing an electronic representation of each copy received from each of the multiplicity of health care entity sources;storing, in the centralized medical record system, the representation received and logically associating, with each representation, client information related to the client; andproviding the client with access, via a web page, to each of the representations over the network, each of the representations constituting an accessible entry in the web page, such that the client can review the representations, wherein the web page includes a facility by which the client can provide information describing each representation, such information being stored in logical association with such representation so as to permit subsequent access to and retrieval thereof via the web page, such information including a client-described label that is used in identifying the representation on the web page. 2. A method according to claim 1, wherein at least one health care entity source is a physician. 3. A method according to claim 1, wherein at least one health care entity source is a clinic. 4. A method according to claim 1, wherein at least one health care entity source is a hospital. 5. A method according to claim 1, wherein providing the client with access to each of the representations over the network includes providing the client with access to the representations via the Internet. 6. A method according to claim 1, further comprising: providing the client with access to a letter template for generating a letter to request that the copy of the medical records and items of data be sent from a source. 7. A method according to claim 6, wherein the letter template causes the letter to be imprinted with a machine readable code identifying at least the client. 8. A method according to claim 6, wherein the letter template causes the letter to be imprinted with a machine readable code identifying at least the health care entity source of the medical records or items of data. 9. A method according to claim 1, further comprising: providing the client with an interface that enables the client to accept or reject the association of the representations with the client. 10. A method according to claim 1, wherein the information includes a date associated with the medical record or item of data. 11. A method according to claim 1, wherein the information includes a name of ... | The Medical RecordBank, Inc. | A61B005/00 G06F019/00
US07963950 | B2 | 2011-06-21 | Self flushing luer activated blood sampling devices | 2007-11-13 | A blood sampling device for allowing a fluid pressure measurement and a fluid sample to be taken from a patient, having an inlet port configured to receive blood from the patient, an outlet port configured to be coupled to a monitoring channel having an infusion fluid, a sampling port configured for extraction of blood from the patient, a sampling channel configured to self-flush, after extraction, the blood with the infusion fluid, and an auxiliary channel for self-flushing the sampling channel. | 1. A patient fluid sampling device connected to a flushing fluid supply and a patient fluid supply, said patient fluid sampling device comprising: a monitoring channel configured to contain the flushing fluid or the patient fluid;a sampling port for providing access to samples of the patient fluid;an auxiliary channel having a first end and a second end, at least the first end connected in fluid communication with the monitoring channel; anda passageway having a first end, a second end and a passage extending between the first and second ends,said passageway having a sampling position wherein the passage connects the sampling port to the monitoring channel in fluid communication, andsaid passageway having a flushing position wherein the first end of the passageway aligns with the second end of the auxiliary channel and the second end of the passageway connects in fluid communication with the auxiliary channel so that the passageway and the auxiliary channel form a loop to enable direct flow of the flushing fluid through the entire passage simultaneously and in parallel with a flow of fluid through the monitoring channel. 2. A patient fluid sampling device of claim 1, wherein the passage is non-divergent for continuous, direct flow during flushing. 3. A patient fluid sampling device of claim 2, wherein the passage has no side-branches. 4. A patient fluid sampling device of claim 3, wherein the passage has no sharp bends. 5. A patient fluid sampling device of claim 4, wherein the passage is straight. 6. A patient fluid sampling device of claim 1, wherein the passageway is defined within a body and the first end and second end of the passageway define openings in an outer surface of the body. 7. A patient fluid sampling device of claim 6, wherein the body has no other openings connected in fluid communication with the passage of the passageway. 8. A patient fluid sampling device of claim 7, wherein the body is a paddle configured to switch the passageway between the sampling position and the flushing position. 9. A patient fluid sampling device of claim 7, wherein the body is a ball valve configured to switch the passageway between the sampling position and the flushing position. 10. A patient fluid sampling device of claim 1, wherein the sampling point is a diaphragm. 11. A patient fluid sampling device of claim 1, wherein the monitoring channel has a diameter approximately four times a diameter of the auxiliary channel. 12. A patient fluid sampling device of claim 10, wherein the diameter of the monitoring channel is approximately 0.2 inches and the diameter of the auxiliary channel is approximately 0.05 inches. 13. A patient fluid sampling device of claim 1, wherein the loop is continuous and non-divergent. 14. A patient fluid sampling device of claim 13, wherein the loop has a U-shape. 15. A patient fluid sampling device of claim 1, wherein the auxiliary channel and the passage have a same diameter. 16. A method of flushing a flui... | Edwards Lifesciences Corporation | A61M005/00 A61B005/00 B65D081/00 F15B013/00 F16K011/085 F16K025/00 F16K005/00
US08089363 | B2 | 2012-01-03 | Method and system for providing data management in data monitoring system | 2011-02-07 | Method and system for providing a fault tolerant data receiver unit configured with a partitioned or separate processing units, each configured to perform a predetermined and/or specific processing associated with the one or more substantially non-overlapping functions of the data monitoring and management system is provided. | 1. A method, comprising: monitoring an analyte level with an analyte sensor and generating one or more analyte related signals;receiving the one or more analyte related signals;generating a quiet host signal prior to receiving the one or more analyte related signals; andin response to the generated quiet host signal, completing a current processing cycle, if any, and entering an inactive mode during the reception of the one or more analyte related signals. 2. The method of claim 1 wherein receiving the one or more analyte related signals includes wirelessly receiving the one or more analyte related signals. 3. The method of claim 1 including transcutaneously positioning at least a portion of the analyte sensor in fluid contact with an analyte of a patient. 4. The method of claim 3 wherein the analyte includes one or more of glucose, blood, or oxygen. 5. The method of claim 1 further including receiving data associated with an analyte sample. 6. The method of claim 5 wherein the analyte sample is received from a glucose test strip. 7. The method of claim 5 further including processing the data associated with the analyte sample. 8. The method of claim 5 wherein the analyte sample includes a blood glucose measurement. 9. The method of claim 1 further including outputting one or more signals associated with a condition of a patient. 10. The method of claim 9 wherein the outputted one or more signals associated with the condition of the patient includes one or more of a directional arrow indicator, a color indicator, or a size indicator. 11. The method of claim 9 wherein outputting the one or more signals associated with the condition of the patient includes progressively increasing or decreasing an associated sound signal over a predetermined time period. 12. The method of claim 9 wherein outputting the one or more signals associated with the condition of the patient includes progressively increasing or decreasing an associated vibratory signal over a predetermined time period. 13. A method, comprising: performing, using a first processing unit, a first predetermined routine including generating a predetermined signal prior to a data reception period;receiving, using a second processing unit, the predetermined signal; andperforming, using the second processing unit, a second predetermined routine including completing a first processing cycle and subsequently entering an inactive mode if the predetermined signal is received during the first processing cycle, or entering into an active mode and determining whether a second processing cycle is to be performed during an impending data reception period if the predetermined signal is received other than during the first processing cycle. 14. The method of claim 13, wherein the second predetermined routine further includes, upon entering into the active mode: expediting the second processing cycle, generating a time flag associated with the expediting of the second processing cycle and enterin... | Abbott Diabetes Care Inc. | G08B023/00
US07967761 | B2 | 2011-06-28 | Sensor and guide wire assembly | 2006-12-01 | A sensor and guide wire assembly (21; 41; 61; 81) for intravascular measurements of a physiological variable in a living body includes a sensor element (22; 42; 62; 82) mounted at a distal sensor portion (23; 43; 63; 83) of a guide wire (24; 44; 64; 84), wherein the sensor portion exhibits a maximal cross-sectional dimension which is larger than a maximal cross-sectional dimension of guide wire portions (28, 30; 48, 50; 68, 70) located proximally of the sensor portion. | 1. A sensor and guide wire assembly for intravascular measurement of a physiological variable in a living body, comprising: a guide wire comprising an external contour, wherein a housing forms a portion of the external contour at a distal sensor portion of the guide wire; anda fluid pressure sensor element encompassed by the housing and at a radially internal location of the housing at the distal sensor portion,wherein the housing comprises a window for establishing fluid communication between the sensor element and a medium surrounding the sensor and guide wire assembly,wherein the external contour at the distal sensor portion exhibits a cross-section which is enlarged compared to a cross-section of the external contour proximal to the distal sensor portion that encompasses said fluid pressure sensor element. 2. A sensor and guide wire assembly according to claim 1, wherein the external contour at the distal sensor portion has a maximal cross-sectional dimension which is at least 10% larger than a maximal cross-sectional dimension of the external contour proximal to the distal sensor portion. 3. A sensor and guide wire assembly according to claim 2, wherein an external contour portion located distally of the distal sensor portion exhibits a maximal cross-sectional dimension which is smaller than said maximal cross-sectional dimension of said distal sensor portion. 4. A sensor and guide wire assembly according to claim 2, wherein an external contour portion located distally of the distal sensor portion exhibits a maximal cross-sectional dimension which is essentially equal to said maximal cross-sectional dimension of said distal sensor portion. 5. A sensor and guide wire assembly according to claim 1, wherein the external contour at the distal sensor portion has a circular cross-section. 6. A sensor and guide wire assembly according to claim 1, wherein the external contour at the distal sensor portion has a non-circular cross-section. 7. A sensor and guide wire assembly according to claim 2, wherein said maximal cross-sectional dimension of the external contour proximal to the distal sensor portion is 0.35 mm. 8. A sensor and guide wire assembly according to claim 2, wherein said maximal cross-sectional dimension of the external contour at the distal sensor portion is less than 2.0 mm. 9. A sensor and guide wire assembly according to claim 1, wherein said fluid pressure sensor element comprises circuitry for wireless communication of said measurement. 10. A sensor and guide wire assembly according to claim 1, wherein the guide wire further comprises a distal tip, a coil spring, a flexible distal tube, and a proximal tube, and wherein the housing, the distal tip, the coil spring, the flexible distal tube, and the proximal tube form the external contour. 11. A sensor and guide wire assembly according to claim 10, wherein the coil spring is attached to the distal tip, and extends to the housing, and wherein the housing is a jacket or sleev... | Radi Medical Systems AB | A61B005/00
US08452570 | B2 | 2013-05-28 | Systems and apparatuses for testing blood glucose measurement engines | 2009-12-18 | A system for testing a blood glucose measurement engine includes a host device emulator having a measurement engine port communicatively coupled to a communications port and electrically coupled to a power supply. A diagnostic computer may be communicatively coupled to the communications port of the host device emulator. The diagnostic computer may include a processor and a memory having computer readable and executable instructions. When the blood glucose measurement engine is communicatively coupled to the measurement engine port, the host device emulator simulates connection to a host device by facilitating the communication of signals between the diagnostic computer and the blood glucose measurement engine. The processor executes the computer readable and executable instructions to: transmit control signals and diagnostic signals to the measurement engine, receive and analyze data signals transmitted from the measurement engine, and monitor a glucose measurement process performed by the measurement engine. | 1. A system for testing a blood glucose measurement engine which performs a glucose measurement process that provides data signals in response to receiving control signals and diagnostic signals, said system comprising: a host device emulator comprising a measurement engine port communicatively coupled to a communications port and electrically coupled to a power supply; anda diagnostic computer communicatively coupled to the communications port of the host device emulator, the diagnostic computer comprising a processor and a memory having computer readable and executable instructions, and provides the control signals and the diagnostic signals to the blood glucose measurement engine, wherein: when the blood glucose measurement engine is communicatively coupled to the measurement engine port, the host device emulator simulates a connection to a host device by facilitating communication of the data signals, the control signals and the diagnostic signals between the diagnostic computer and the blood glucose measurement engine, andthe processor executes the computer readable and executable instructions to: transmit the control signals and the diagnostic signals from the diagnostic computer to the blood glucose measurement engine,receive and analyze the data signals transmitted from the blood glucose measurement engine,step the blood glucose measurement engine through the glucose measurement process,monitor a glucose measurement process performed by the blood glucose measurement engine, andpause the blood glucose measurement engine after an instruction of the glucose measurement process has been performed. 2. The system of claim 1 wherein the host device emulator further comprises a host device port communicatively coupled to the measurement engine port, the communications port and the power supply; and the computer readable and executable instructions are executed by the processor to monitor data signals and control signals exchanged between the blood glucose measurement engine communicatively coupled to the measurement engine port and the host device communicatively coupled to the host device port. 3. The system of claim 2 wherein, when a host device is coupled to the host device port, the host device is powered by the power supply. 4. The system of claim 1 wherein the computer readable and executable instructions are executed by the processor to calibrate the blood glucose measurement engine. 5. The system of claim 1 further comprising a wake request indicator, a measurement engine request indicator, and a measurement engine ready indicator, wherein: the wake request indicator is electrically coupled to the measurement engine port and provides a visual or audible signal when a wake request signal is transmitted to the measurement engine port;the measurement engine request indicator is electrically coupled to the measurement engine port and provides a visual or audible signal when an electronic signal is transmitted to or from the measurement en... | Roche Diagnostics Operations, Inc. | G06F019/00 A61B005/145
US08784334 | B2 | 2014-07-22 | Systems and methods for obtaining analytes from a body | 2010-12-03 | A method may include collecting at least one analyte from within a body, ejecting the collected at least one analyte from the body through at least one dermal layer of the body, and receiving the ejected at least one analyte outside the body. A system may include a means for collecting at least one analyte from within a body, a means for ejecting the collected at least one analyte from the body through at least one dermal layer of the body, and a means for receiving the ejected at least one analyte outside the body. | 1. A system, comprising: means for collecting at least one analyte from within a body;means for ejecting said at least one analyte from said body through at least one dermal layer of said body;means for receiving said at least one analyte outside said body, said receiving means including a means for charging an energy storage mechanism of said collecting means via an energy source located outside said body, located inside said body, or located both outside said body and inside said body. 2. The system of claim 1, further comprising: means for transferring said at least one analyte from said collecting means within said body to said ejecting means within said body. 3. The system of claim 2, wherein said collecting means is located distal to said ejecting means and said at least one analyte is transported to said ejecting means. 4. A system, comprising: means for collecting at least one analyte from within a body;means for ejecting said at least one analyte from said body through at least one dermal layer of said body;means for receiving said at least one analyte outside said body; andmeans for powering said collecting means via an energy source located outside said body, located inside said body, or located both outside said body and inside said body. 5. The system of claim 4, wherein said means for powering said collecting means via an energy source located outside said body, located inside said body, or located both outside said body and inside said body comprises: means for powering said collecting means via said means for receiving said at least one analyte outside said body. 6. The system of claim 4, further comprising: means for transmitting a location signal enabling said ejecting means to be located. 7. The system of claim 6, wherein said location signal comprises at least one of an electrical current, an electrical field, a magnetic flux, an optical signal, a radio frequency identification, an ultrasound, a vibration, an electromagnetic force, a force, or a pressure. 8. The system of claim 4, further comprising: means for sorting said at least one analyte. 9. The system of claim 4, further comprising: means for concentrating said at least one analyte. 10. The system of claim 4, further comprising: means for diluting said at least one analyte. 11. The system of claim 4, wherein said means for collecting at least one analyte from within a body comprises: means for collecting at least a portion of at least one of cells, proteins, bacteria, blood, a blood component, molecules, viruses, viral particles, pathogens, parasites, malarial parasites, oglionucleotides, lymph, a lymph component, or cerebral spinal fluid. 12. The system of claim 11, wherein said blood component comprises at least one of plasma or serum. 13. The system of claim 4, wherein said means for receiving said at least one analyte outside said body comprises: means for receiving said at least one analyte outside said body in a receiving device. 14. The system of cla... | The Invention Science Fund I, LLC | A61B005/00
US08594770 | B2 | 2013-11-26 | Multispectral detection and presentation of an object's characteristics | 2010-10-14 | An apparatus for capturing a multispectral image of an object is described. The apparatus includes one or more means for transmitting a beam of laser light at a first wavelength and a beam of laser light at one or more additional wavelengths different from the first wavelength. There is a means for causing the beams of laser light to travel in a coaxial path and a moving mirror. The beams of light bounce off the mirror thereby producing a two dimensional projection pattern. This pattern travels from the mirror along a first path to an object and wherein some of the laser light penetrates the object and travels to an internal structure of the object. The reflection of the laser light returns to a photo detector along a path different from said first path. | 1. An apparatus for capturing a multispectral image of a deep internal structure of an object, said apparatus comprising: means to transmit a beam of light comprising two or more different wavelengths, said two or more different wavelengths configured to determine one or more characteristics of the internal structure;means to scan said beam of light in a plurality of optical paths to produce a two dimensional scan pattern;an array of photo detectors distributed around the object; each said photo detector configured to detect said scanned two or more wavelengths of light after being differentially absorbed by the internal structure and the surrounding areas, and to output a signal of a contrast formed by the absorption spectrum of said two or more wavelengths by the internal structure, said means to transmit further configured to also transmit one or more visible wavelengths of light, by using said signal from said photodetectors, to selectively project the one or more characteristics of the internal structure onto the surface of the object;a memory;a comparator configured to receive the output of said photo detectors, and to provide one bit of data for each pixel of the photo detector output to indicate whether its brightness is too bright or too dark, said bit of pixel brightness data being stored in said memory and being incrementally updated for said selective projection, said stored pixel brightness data used to adjust power to said means to transmit light, on a pixel-by-pixel basis, for said selective projection to provide improved contrast sensitivity for said projected image. 2. The apparatus according to claim 1 wherein said wherein said array of photo detectors comprises a ring of said photodetectors uniformly distributed around the object. 3. The apparatus according to claim 2 wherein each said photo detector comprises a respective optical element configured to shape a field of view of said respective photo detector. 4. The apparatus according to claim 1 wherein said selective projection comprises said means to transmit being configured to implement one or more techniques to alternate between differential absorption of said two or more wavelengths and said projection of said one or more visible wavelength. 5. The apparatus according to claim 4 wherein said one or more techniques comprises: said means to transmit configured to alternately scan only said two or more wavelengths for absorption for a specific period of time, and to project said one or more wavelengths for another specific period of time;said means to transmit configured to alternately scan a line and to project a line;said means to transmit configured to scan multiple lines and to project one scan line;said means to transmit configured to scan a line and repeatedly project for multiple scan lines; andsaid means to transmit configured to scan for a full frame and to project for a full frame. 6. The apparatus according to claim 1 wherein said selective contrast sensitivi... | AccuVein, Inc. | A61B006/00
US08475371 | B2 | 2013-07-02 | Physiological monitoring garment | 2010-08-31 | The present invention is directed to systems and methods for monitoring characteristics of a subject. A system according to an exemplary embodiment of the invention includes a sensor subsystem including at least one respiratory sensor disposed proximate to the subject and configured to detect a respiratory characteristic of the subject, wherein the sensor subsystem is configured to generate and transmit at least one respiratory signal representing the respiratory characteristic, and at least one physiological sensor disposed proximate to the subject and configured to detect a physiological characteristic of the subject, wherein the sensor subsystem is configured to generate and transmit at least one physiological signal representing the physiological characteristic, and a processor subsystem in communication with the sensor subsystem, the processor subsystem being configured to receive at least one of the at least one respiratory signal and the at least one physiological signal. | 1. A fitness monitoring system for monitoring physiological parameters of a subject engaged in a physical activity, comprising: a sensor subsystem including a first sensor and a second sensor, wherein the first sensor comprises a multi-layered printed circuit comprising a conductive circular coil formed from a plurality of printed circuit boards;a microcontroller; anddriver circuitry configured to convert a square wave drive from the microcontroller into an efficient current drive for the coil,wherein each layer of the multi-layered printed circuit comprises a continuous turn of the coil on one of the printed circuit boards,wherein the turns of the coil are connected in series,wherein the first and second sensors are responsive to changes in a parameter, andwherein the sensor subsystem is configured to generate and transmit a signal representative of the parameter. 2. The system of claim 1, wherein: the conductive circular coil comprises a six-layer printed circuit board having eighteen turns on each layer. 3. The system of claim 2, wherein the coil is approximately two inches in diameter. 4. The system of claim 1, wherein the layers are connected together in series using printed circuit layer connectors. 5. The system of claim 1, further comprising a battery with a voltage range of approximately 2.8V to 3.7V, configured to supply power to the sensor subsystem. 6. The system of claim 1, wherein the driver circuitry is configured to provide monitoring of actual drive signals to the microcontroller. 7. The system of claim 1, further comprising: a processor subsystem in communication with the sensor subsystem and configured to receive the signal, wherein the processor subsystem is programmed and adapted to control the sensor subsystem and to process the signal, and wherein the processor subsystem includes a relationship for determining a respiratory characteristic from the signal, and wherein the processor subsystem is adapted to generate and transmit a respiratory parameter signal representing the respiratory parameter. 8. The system of claim 1, wherein the first and second sensors are responsive to changes in the distance therebetween, and wherein the signal represents a change in the distance between the first and second sensors. 9. The system of claim 1, further comprising a capacitor network configured to match the square wave drive to a voltage drive requirement of the coil. 10. The system of claim 1, wherein the sensor subsystem includes a device for monitoring the subject's movement. 11. The system of claim 10, wherein the sensor subsystem includes an inertial sensor. 12. The system of claim 10, wherein the sensor subsystem includes a GPS receiver. 13. The system of claim 1, wherein the conductive circular coil comprises a plurality of continuous turns on each layer of the multi-layered printed circuit. | adidas AG | A61B005/00
US08396267 | B2 | 2013-03-12 | Correcting subject motion in multiple resolutions in magnetic resonance imaging | 2009-11-04 | Systems, methods and articles of manufacture are disclosed for compensating for motion of a subject during an MRI scan of the subject. k-space data may be received from the MRI scan of the subject. A first graphical image may be generated from a first set of data elements from the k-space data. Similarly, a second graphical image may be generated from a second set of data elements from the k-space data. An offset in pixels may be determined by which to translate the second graphical image from the first graphical image to compensate for the motion. The k-space data may be modified at a sub-pixel offset relative to the determined offset. A motion-compensated graphical image of the subject may be generated from the modified k-space data. Doing so reduces the search space evaluated to sharpen images generated from the k-space data. | 1. A computer-implemented method for compensating for motion of a subject during a magnetic resonance imaging (MRI) scan of the subject, comprising: receiving k-space data from the MRI scan of the subject, wherein the k-space data comprises frequency domain data acquired from the MRI scan used to generate a graphical image of the subject, wherein the frequency domain representation includes a plurality of data elements, and wherein the plurality of data elements includes a first set of data elements acquired at a first point in time during the MRI scan and a second set of data elements acquired at a second point in time during the MRI scan;generating a first graphical image from the first set of data elements;generating a second graphical image from the second set of data elements;determining an offset in pixels by which to translate the second graphical image from the first graphical image to compensate for the motion based on evaluating, for each offset of a plurality of distinct offsets, a sharpness of a graphical image generated from the first graphical image and the second graphical image translated by the respective offset;modifying the k-space data based on the determined offset to compensate for the motion to a sub-pixel offset, relative to the determined offset; andgenerating a motion-compensated graphical image of the subject from the modified k-space data. 2. The computer-implemented method of claim 1, wherein the motion-compensated graphical image of the subject depicts a part of a human body. 3. The computer-implemented method of claim 1, wherein modifying the k-space data based on the determined offset to compensate for the motion to a sub-pixel offset relative to the determined offset comprises: determining, from the offset in pixels, a plurality of sub-pixel offsets relative to the offset in pixels by which to translate the second graphical image from the first graphical image to compensate for the motion;for each of the determined sub-pixel offsets: determining a complex number from the offset in pixels and the respective sub-pixel offset,multiplying the second set of data elements by the respective determined complex number,generate a graphical image from the first set of data elements and from the modified second set of data elements, andevaluating a sharpness of the generated graphical image; andselecting the sub-pixel offset based on the respective sharpness of each generated graphical image. 4. The computer-implemented method of claim 3, wherein the second set of data elements is multiplied by the respective determined complex number to modify a phase of at least one sinusoid represented by the second set of data elements. 5. The computer-implemented method of claim 3, wherein the complex number is determined using the Fourier shift theorem. 6. The computer-implemented method of claim 1, wherein generating the motion-compensated graphical image of the subject comprises performing a two-dimensional discrete Fourier trans... | International Business Machines Corporation | G06K009/00
US08355778 | B2 | 2013-01-15 | Trunk width measuring unit and visceral fat measuring device | 2011-02-08 | To provide a trunk width measuring unit capable of improving reliability of measurement precision, and a visceral fat measuring device. The trunk width measuring unit is provided with a first contact portion to be brought into contact with an upper surface of a trunk of a user in a supine position, a second contact portion to be brought into contact with one of side surfaces of the trunk, a third contact portion to be brought into contact with the other side surface of the trunk, and a trunk width calculating unit for calculating vertical width and horizontal width of the trunk from height from an upper surface of a bed to a contact position of the first contact portion and a distance between the second contact portion and the third contact portion. | 1. A trunk width measuring unit, comprising: a first contact portion to be brought into contact with an upper surface of a trunk of a user in a supine position;a second contact portion to be brought into contact with one of side surfaces of the trunk;a third contact portion to be brought into contact with the other side surface of the trunk;a unit main body in which the first contact portion is attached movably in a vertical direction, the second contact portion and the third contact portion are attached movably in a horizontal direction, and the second contact portion or the third contact portion is placed at a position where the second contact portion or the third contact portion is brought into contact with the side surface of the trunk of the user; anda trunk width calculating unit for calculating vertical width and horizontal width of the trunk, the trunk width calculating unit calculating the height from a plane on which the user lies to a contact position of the first contact portion based on a distance between the second contact portion and the third contact portion, whereinthe first contact portion is attached to the unit main body so as to be lowered by self-weight of the first contact portion and moved in the vertical direction relative to the unit main body in a state that the first contact portion is disposed on the upper surface of the trunk of the user, and the first contact portion is thus movable so as to follow a change in a shape of the trunk based on breathing of the user. 2. The trunk width measuring unit according to claim 1, further comprising: a first detector for detecting a lowering distance of the first contact portion from a first original position on an upper side of the trunk of the user to a contact point with the upper surface of the trunk;a second detector for detecting a moving distance of the second contact portion from a second original position in the horizontal direction; anda third detector for detecting a moving distance of the third contact portion from a third original position in the horizontal direction, whereinthe trunk width calculating unit calculates the height from the plane on which the user lies to the contact position of the first contact portion based on the moving distance of the first contact portion detected by the first detector, and calculates a distance between the second contact portion and the third contact portion based on the moving distance of the second contact portion detected by the second detector and the moving distance of the third contact portion detected by the third detector. 3. The trunk width measuring unit according to claim 2, wherein one of the second contact portion and the third contact portion is biased in the facing direction to the other of the second contact portion and the third contact portion and thus movable so as to follow the change in the shape of the trunk based on breathing of the user. 4. The trunk width measuring unit according to claim 2, including:... | Omron Healthcare Co., Ltd. | A61B005/00
US09375159 | B2 | 2016-06-28 | Device for indicator dilution measurements | 2008-10-15 | The invention relates to indicator dilution measurements of a central volume (V1) with a first site of injection (S1) upstream of the central volume (V1) a second site of detection (S2) of the diluted indicator downstream of the central volume (V1), wherein a first additional volume (V2) is defined between the first site (S1) and the central volume (V1) and a first additional branch (B2) is defined between the first site (S1) and the central volume (V1) and wherein a second additional volume (V3) is defined between the central volume (V1) and the second site (S2) and a second additional branch (B3) is defined between the central volume (V1) and the second site (S2) wherein a result of central volumetric parameters are corrected for the first and second additional volumes (V2, V3) and/or for the first and second additional branches (B2, B3). | 1. A method for correcting a dilution-based determination of an internal volume of a heart based on flow of a diluted indicator between an injection site upstream of the heart and a detection site downstream of the heart, the method comprising: injecting indicator at the injection site;diluting the injected indicator in blood flowing through the heart between the injection site and the detection site;detecting diluted indicator at the detection site; andoperating a processor configured to carry out the steps of: calculating the internal volume of the heart based on a detection of the diluted indicator at the detection site;estimating an additional volume through which the injected indicator flows, the estimated additional volume being estimated as substantially an entire internal volume of a blood vessel through which the injected indicator flows between the injection site and the entrance to the right atrium of the heart; andcalculating and applying a correction to calculation of the internal volume of the heart based on the estimated additional volume. 2. A method for correcting a dilution-based determination of an internal volume of a heart based on flow of a diluted indicator between an injection site upstream of the heart and a detection site downstream of the heart, the method comprising: injecting indicator at the injection site;diluting the injected indicator in blood flowing through the heart between the injection site and the detection site;detecting diluted indicator at the detection site; andoperating a processor configured to carry out the steps of: calculating the internal volume of the heart based on a detection of the diluted indicator at the detection site;estimating an additional volume through which the injected indicator flows, the estimated additional volume being estimated as substantially an entire internal volume of a blood vessel through which the injected indicator flows between the aortic valve and the detection site; andcalculating and applying a correction to the calculation of the internal volume of the heart based the estimated additional volume. 3. A method for correcting a dilution-based determination of an internal volume of a heart based on flow of a diluted indicator between an injection site upstream of the heart and a detection site downstream of the heart, the method comprising: injecting indicator at the injection site;diluting the injected indicator in blood flowing through the heart between the injection site and the detection site;detecting diluted indicator at the detection site; andoperating a processor configured to carry out the steps of: calculating the internal volume of the heart based on a detection of the diluted indicator at the detection site; andcalculating and applying a correction to the calculation of the internal volume of the heart based on an estimated flow of blood in a branch through which no indicator flows but through which blood flows upstream of the heart before entering the hea... | Edwards Lifesciences IPRM AG | A61B005/02 A61B005/029 A61B005/028 A61B005/0275
US09560984 | B2 | 2017-02-07 | Analysis of controlled and automatic attention for introduction of stimulus material | 2009-10-29 | A system analyzes neuro-response measurements including regional electroencephalography (EEG) measurements from subjects exposed to stimulus materials to determine locations in stimulus materials eliciting controlled attention and automatic attention. Additional stimulus materials are inserted into locations having salient attention attributes. In some examples, a challenging task is used to direct controlled attention onto a location and additional stimulus material is subtly presented in the location to benefit from automatic attention and salient attention measurements. | 1. A method comprising: identifying, by executing an instruction with a processor, a first change in amplitude and a second change in amplitude in first neuro-response data gathered from a user while the user is exposed to entertainment;synchronizing, by executing an instruction with the processor, the first neuro-response data with the entertainment;identifying, by executing an instruction with the processor, a first temporal location in the entertainment coinciding with the first change in amplitude and a second temporal location in the entertainment coinciding with the second change in amplitude;analyzing, by executing an instruction with the processor, the first neuro-response data to determine a first priming characteristic of the entertainment at the first temporal location based on one or more of (a) a first controlled attention level of the user at the first temporal location, (b) a first memory retention level of the user at the first temporal location, or (c) a first resonance level of the user at the first temporal location;analyzing, by executing an instruction with the processor, the first neuro-response data to determine a second priming characteristic of the entertainment at the second temporal location based on one or more of (a) a second controlled attention level of the user at the second temporal location, (b) a second memory retention level of the user at the second temporal location, or (c) a second resonance level of the user at the second temporal location;identifying, by executing an instruction with the processor, at least one of (1) a first retention characteristic of the entertainment at the first temporal location based on the first memory retention level or (2) a first resonance characteristic of the entertainment at the first temporal location based on the first resonance level;identifying, by executing an instruction with the processor, at least one of (1) a second retention characteristic of the entertainment at the second temporal location based on the second memory retention level or (2) a second resonance characteristic of the entertainment at the second temporal location based on the second resonance level;selecting, by executing an instruction with the processor, a third temporal location or a fourth temporal location in the entertainment as a first candidate temporal location for introduction of advertising material, the third temporal location occurring after the first temporal location and the fourth temporal location occurring after the second temporal location, the selecting of the first candidate temporal location including selecting the third temporal location based on (1) at least one of the first priming characteristic or the second priming characteristic and (2) when at least one of the first retention characteristic or the first resonance characteristic is increased relative to the respective second retention characteristic or the second resonance characteristic and selecting the fourth temporal loc... | The Nielsen Company (US), LLC | A61B005/04 A61B005/0484 G06Q030/02 A61B005/16
US20110125057 | A1 | 2011-05-26 | INTRAVENOUS CANNULA ASSEMBLY | 2007-09-28 | An intravenous cannula assembly comprises a longitudinally extending barrel (10, 210), a cannula hub (30, 230), a cannula (50, 250), a cannula guard (70, 270) and a deployment mechanism (90, 100, 290, 300). The cannula hub (30, 230) is mounted in the barrel (10, 210) and has a hub aperture (35, 235) extending between the hub leading end (31, 231) and hub trailing end (32, 232). The cannula hub (30, 230) is adapted to mount a tube at the hub trailing end (32, 232) in fluid communication with the hub aperture (35, 235). The cannula (50, 250) is mounted in, and in fluid communication with, the hub aperture (35, 235). The cannula (50, 250) extends from the hub leading end (31, 231) through the barrel leading end (11, 211). The cannula guard (70, 270) is concentrically mounted about the cannula hub (30, 230). The deployment mechanism (90, 100, 290, 300) is adapted to deploy the cannula guard (70, 270) from a retracted position substantially within the barrel (10, 210) to a deployed position extending from the barrel leading end (11, 211) and covering the portion of the cannula (50, 250) extending through the barrel leading end (11, 211). | 1. An intravenous cannula assembly comprising: a longitudinally extending barrel having a barrel leading end and a barrel trailing end;a cannula hub mounted in said barrel and having a hub leading end, a hub trailing end and a hub aperture extending between said hub leading end and said hub trailing end, said cannula hub being adapted to mount a tube at said hub trailing end in fluid communication with said hub aperture;a cannula mounted in, and in fluid communication with, said hub aperture, said cannula extending from said hub leading end through said barrel leading end;a cannula guard concentrically mounted about said cannula hub; anda deployment mechanism adapted to deploy said cannula guard from a retracted position at least substantially within said barrel to a deployed position extending from said barrel leading end and covering the portion of said cannula extending through said barrel leading end. 2. The assembly of claim 1, wherein said cannula guard engages said barrel adjacent said barrel leading end for all positions of said cannula guard between said retracted position and said deployed position. 3. The assembly of claim 2, wherein said barrel engages said cannula guard toward a trailing end of said cannula guard for all positions of said cannula guard between said retracted position and said deployed position. 4. The assembly of claim 2, wherein a leading portion of said cannula guard is provided with a series of longitudinally extending ribs spaced about a periphery of said leading portion of said cannula guard, said ribs engaging said barrel adjacent said barrel leading end. 5. The assembly of claim 1, wherein said deployment mechanism comprises: a spring biasing said cannula guard towards said deployed position; anda releasable locking arrangement adapted to retain said cannula guard in said retracted position until release of said releasable locking arrangement. 6. The assembly of claim 1, wherein said releasable locking arrangement comprises: a first interlocking structure provided on said cannula guard;a locking member mounted on said barrel and having a second interlocking structure;said locking member being displaceable between a locking position at which said first interlocking structure engages said second interlocking structure to thereby retain said cannula guard in said retracted position and an unlocking position at which said first interlocking structure disengages from said second interlocking structure, thereby releasing said cannula guard. 7. The assembly of claim 6, wherein said first interlocking structure is located within said barrel. 8. The assembly of claim 6, wherein said locking member is of a generally cylindrical form and is angularly displaceable about a longitudinal axis of said barrel between said locking position and said unlocking position. 9. The assembly of claim 6, wherein said locking member is mounted on said barrel adjacent said barrel trailing end. 10. The assembly of claim 9, wherein... | - | A61B005/15
US20110117566 | A1 | 2011-05-19 | Model System for Diagnosing Lipid Metabolism | 2009-06-10 | The invention relates to a population model for the analysis of blood lipoprotein physiology in a test subject comprising: a. a submodel for the production of blood lipoproteins; b. a submodel for the lipolysis of blood lipoproteins; c. a submodel for the reabsorption of blood lipoproteins; and d. a submodel relating blood lipoprotein particle size to biochemical composition, more specifically triglyceride content, there-by providing an analysis of the physiological processes underlying a steady state particle population distribution. Each submodel is given as function, using the size of the lipoprotein particle as the independent variable. | 1. A computer programme product comprising instructions for causing a processor to carry out the steps of calculating a population model for the analysis of blood lipoprotein physiology in a test subject comprising: a. a submodel for the production of blood lipoproteins;b. a submodel for the lipolysis of blood lipoproteins;c. a submodel for the reabsorption of blood lipoproteins; andd. a submodel relating blood lipoprotein particle size to biochemical composition, more specifically triglyceride content,wherein each submodel is given as function, using the size of the lipoprotein particle as the independent variable, thereby providing an analysis of the physiological processes underlying a steady state particle population distribution. 2. A computer programme product according to claim 1, wherein in the submodel for the lipolysis two models are contained, one for extra-hepatic tissue mediated lipolysis and one for hepatic lipolysis. 3. A computer programme product according to claim 1-2, wherein the submodel for the reabsorption is able to distinguish between apoB and apoE mediated reabsorption. 4. A computer programme product according to any of claims 1-3, wherein the population model further comprises one or more of the following submodels: a. a submodel relating blood lipoprotein particle size to total cholesterol content;b. a submodel relating blood lipoprotein particle size to free cholesterol content;c. a submodel relating blood lipoprotein particle size to cholesterol ester content;d. a submodel relating blood lipoprotein particle size to phospholipid content;e. a submodel relating blood lipoprotein particle size to total protein content. 5. A computer programme product according to claim 1-4 for calculating a population model for the presence of blood lipoproteins in a test subject wherein the total steady-state pool of lipoproteins in a diameter range [dadb] is given by:Qout([dadb])=∑di,j+12di,jr≤dbdi,j-12di,jr≥daQss(di,j)+Rwherein Qss(di,j) is the steady state pool of a cascade step at diameter di,j, and subclass resolution di,jr. R is the remainder for the boundary subclasses, which partially fall in the selected range:Rlow=(di,j+12di,jr)-dadi,jrQss(di,j)whereda∈[di,j-12di,jr,di,j+12di,jr]Rhigh=db-(di,j-12di,jr)di,jrQss(di,j)wheredb∈[di,j-12di,jr,di,j+12di,jr]R=Rhigh+Rlow 6. A computer programme product according to claim 5, wherein the steady-state pool in the cascade step Qssd0(di) at each size di,j is given by:Qss(di,j)=Jp(di,j)+Jl(di,j)+Jl,liver(di,j)kl(di,j)+kl,liver(di,j)+ku,liver(di,j)wherein di,j is the mean subclass particle diameter in the i-th step of a lipolysis cascade, starting from subclass j within the cascade step size range, herein Jp(di,j) is the particle influx resulting from production, Ji(di,j) is the particle influx resulting from extrahepatic lipolysis, Jl,liver(di,j) is the particle influx resulting from hepatic lipolysis, kl is the extrahepatic lipolysis rate, kl,liver is the hepa... | - | G06F017/10 G06G007/60 A61B005/00 C12Q001/68 C12Q001/44 G01N033/92 G01N033/68 G01N033/48
US20110105951 | A1 | 2011-05-05 | SYSTEMS AND METHODS FOR TREATING, SANITIZING, AND/OR SHIELDING THE SKIN OR DEVICES APPLIED TO THE SKIN | 2010-10-29 | The present invention generally relates, in certain aspects, to systems and methods for treating, sanitizing, and/or shielding blood on the surface of the skin or devices applied to the skin. In one set of embodiments, while blood is withdrawn from the skin of the subject, the blood is covered and/or shielded such that the blood is not seen by the subject. In another set of embodiments, any blood present on the skin may be altered or treated such that it no longer appears as blood, e.g., by exposing the blood to a chemical such as a peroxide that alters the color of the blood. In yet another set of embodiments, blood present on the skin is withdrawn or taken up by the device such that visible residual blood is not present in an amount visible to the unaided eye once the device is removed. In some cases, the same device may be used to sanitize the skin and deliver to and/or withdraw fluid from the skin and/or beneath the skin. The sanitization may be performed before, during, or after delivering to and/or withdrawing fluid from the skin and/or beneath the skin. Sanitization may be performed, for example, using a sanitizer such as hydrogen peroxide, ethyl alcohol, isopropyl alcohol, or the like. | 1. A device for extracting a fluid from the skin and/or from beneath the skin of a subject, the device comprising: a support structure for application to the skin of the subject;a fluid transporter associated with the support structure; anda chemical for altering, from the skin of the subject, residual blood and/or an appearance of residual blood such that when the support structure is removed from the skin of the subject, visible residual blood is not present in an amount that is visible to the unaided eye,wherein the support structure includes at least a cover between the subject's line of vision and any blood extracted by the device,whereby application, use, and removal of the device from the skin of the subject produces no blood visible by the unaided eye. 2. The device of claim 1, wherein the support structure further comprises an adhesive for adhesion of the support structure to the skin of the subject. 3. (canceled) 4. The device of claim 1, wherein the support structure further comprises a recess containing the fluid transporter. 5. The device of claim 1, wherein the support structure further comprises a microfluidic channel in fluidic communication with the fluid transporter. 6. The device of claim 1, wherein the fluid comprises blood. 7. The device of claim 1, wherein the fluid comprises interstitial fluid. 8. The device of claim 1, wherein the device further comprises a storage chamber in fluidic communication with the fluid transporter. 9. The device of claim 1, wherein the device is able to determine an analyte contained within the fluid withdrawn from the skin and/or from beneath the skin of the subject. 10-11. (canceled) 12. The device of claim 9, wherein the device is able to display a result of the determination of the analyte. 13. The device of claim 9, wherein the analyte is glucose. 14. The device of claim 1, wherein the fluid transporter comprises a needle. 15. The device of claim 14, wherein the needle is a microneedle. 16. The device of claim 1, wherein the fluid transporter comprises a solid microneedle. 17. The device of claim 1, wherein the fluid transporter comprises a hollow microneedle. 18-20. (canceled) 21. The device of claim 1, wherein the support structure is able to apply a pressure less than atmospheric pressure to the skin of the subject. 22. The device of claim 1, wherein the support structure comprises a self-contained vacuum chamber having a pressure less than atmospheric pressure. 23. (canceled) 24. The device of claim 1, wherein the fluid transporter is able to be activated by the subject. 25. The device of claim 1, wherein the fluid transporter is activated upon applying the device to the skin of the subject. 26-27. (canceled) 28. The device of claim 1, wherein the chemical comprises a penetration enhancer. 29. The device of claim 1, wherein the chemical comprises a peroxide. 30. The device of claim 1, wherein the chemical comprises an oxidizer. 31. The device of claim 1, wherein... | Seventh Sense Biosystems, Inc. | A61B005/00
US08419646 | B2 | 2013-04-16 | Blood pressure estimation apparatus and blood pressure estimation method | 2009-04-09 | To provide a non-invasive blood pressure estimation apparatus, which can accurately estimate systolic blood pressure from blood flow sound of a dialysis patient and continuously estimate systolic blood pressure by continuously picking up the blood flow sound. A blood pressure estimation apparatus creates a standard pulse curve by relating a start point and end point of a rising phase of a pulse wave to diastolic blood pressure and systolic blood pressure, respectively, creates a correspondence curve between blood flow sound power and estimated blood pressure by contrasting the standard pulse curve with the blood flow sound power curve obtained from blood flow sound at a shut site, with the two curves plotted on the same time axis, derives a systolic blood pressure estimation linear function from the correspondence curve, inputs a measured maximum value of the blood flow sound power into the linear function, and thereby estimates systolic blood pressure. | 1. A blood pressure estimation apparatus which estimates blood pressure of a subject using a standardization mode and an estimation mode, comprising: blood flow sound power waveform calculation means for determining a blood flow sound power waveform from blood flow sound of the subject in the standardization mode;blood pressure measuring means for measuring reference systolic blood pressure and reference diastolic blood pressure of the subject in the standardization mode;standard pulse curve storing means for storing a standard pulse curve whose time axis and pulse pressure axis have been standardized such that times and pulse pressures at a start point and an end point of a rising phase of a pulse wave measured in advance will be constant, in the standardization mode;first computing means for determining a correspondence curve which represents a relationship between blood flow sound power and estimated blood pressure from relationships among the blood flow sound power waveform, the reference systolic blood pressure, the reference diastolic blood pressure, and the standard pulse curve, in the standardization mode;blood flow sound power measuring means for continuously measuring the blood flow sound of the subject and the blood flow sound power in the estimation mode; andsecond computing means for continuously estimating blood pressure from the blood flow sound power measured continuously, using the correspondence curve in the estimation mode. 2. The blood pressure estimation apparatus according to claim 1, wherein the first computing means comprises: third computing means for creating a reference blood pressure curve which represents a relationship of blood pressure values to standardized time based on the standard pulse curve by bringing the reference diastolic blood pressure value and the reference systolic blood pressure value measured by the blood pressure measuring means into correspondence with a minimum value and a maximum value of standardized pulse pressure values of the standard pulse curve, respectively;fourth computing means for standardizing a time axis in a rising phase of the blood flow sound power waveform at the time corresponding to the time of measurement of the reference diastolic blood pressure value and the reference systolic blood pressure value and creating a reference blood flow sound power curve which represents a relationship of the blood flow sound power to the standardized time; andfifth computing means for creating the correspondence curve which represents a relationship between the blood flow sound power and the estimated blood pressure in a period of common standardized time based on a relationship between the reference blood pressure curve and the reference blood flow sound power curve in the period of common standardized time. 3. The blood pressure estimation apparatus according to claim 2, wherein the first computing means comprises sixth computing means for calculating a systolic blood pressure estimation lin... | Asahi Kasei Kabushiki Kaisha Nextier Corporation | A61B005/02
US08535215 | B2 | 2013-09-17 | Implantable heart monitoring device and method | 2008-08-29 | In an implantable heart monitoring device and method, particularly for monitoring diastolic dysfunction, a control circuit (a) detects the heart rate, (b) derives information correlated to the stroke volume of the heart at the detected heart rate, and (c) stores the detected heart rate and the derived information correlated to the stroke volume in a memory. The control circuit automatically implements (a), (b) and (c) at a number of different occasions for a number of different, naturally varying heart rates, so that the memory contains information indicating the stroke volume as a function of the heart rate. | 1. An implantable heart monitoring device comprising: a control circuit having an input that receives a signal from at least one electrode that senses a parameter related to cardiac activity;a memory in communication with said control circuit; andsaid control circuit being configured to (a) detect the heart rate of the subject from said signal (b) derive information correlated to the stroke volume of the heart at the detected heart rate, and (c) store the detected heart rate and the derived information correlated to the stroke volume in said memory, and said control circuit being configured to automatically execute (a), (b) and (c) at a plurality of different occasions for different, naturally varying heart rates, to enter information into said memory to produce a representation in said memory of the stroke volume as a function of the heart rate;wherein said control circuit is configured to operate said electrode to execute a measurement as to how an amount of blood in the left ventricle of the heart varies with time, and to derive said information correlated to the stroke volume of the heart from a result of said measurement;wherein said control circuit is configured to generate a curve, as said result of said measurement, indicating how said amount of blood in the left ventricle varies with time, and to derive said information correlated with the stroke volume from the curve or by an operation on said curve selected from the group consisting of filtering said curve, analyzing a template based on different curves for different heartbeats, and generating and analyzing a template based on filtering different curves for respectively different heartbeats; andwherein said control circuit is configured to derive said information correlated to the stroke volume from a characteristic of said curve, said filtered curve or said template, selected from the group consisting of an area and a peak-to-peak value. 2. An implantable heart monitoring device as claimed in claim 1 wherein said control circuit is configured to execute (a), (b) and (c) for at least three different heart rates. 3. An implantable heart monitoring device as claimed in claim 1 wherein said control circuit is configured to execute (a), (b) and (c) for different heart rates that differ from each other by at least 10% between a highest of said different heart rates and a lowest of said different heart rates. 4. An implantable heart monitoring device as claimed in claim 1 wherein said control circuit is configured to emit a warning signal if said information entered into said memory indicates that said stroke volume as a function of said heart rate satisfies a predetermined criterion. 5. An implantable heart monitoring device as claimed in claim 1 wherein said control circuit is configured to, for a predetermined heart rate, derive said information correlated to the stroke volume during a plurality of heartbeats. 6. An implantable heart monitoring device as claimed in claim 1 wherein sa... | St. Jude Medical AB | A61N001/00
US20110105931 | A1 | 2011-05-05 | System for Determining Patient Heart related Parameters for use in Heart Imaging | 2008-10-08 | A system uses integrated spatio-temporal analysis in X-ray angiography, for example, by using spatial information within each image frame and temporal information between image frames to provide robust and accurate estimation of stroke area and volume, two and three dimensional ejection fraction and to accommodate patient heart variation. A system determines patient heart related parameters for use in patient heart imaging examination. An image data processor processes data representing multiple cardiac images of a patient over multiple heart beat cycles of the patient to derive data representing a distribution curve of a heart section area over multiple heart beat cycle times and indicating heart section area change over a heart beat cycle. An area processor determines a heart section area in response to user command. Also a computation processor determines a heart function parameter in response to the determined heart section area and the indicated heart section area change. | 1. A system for determining patient heart related parameters for use in patient heart imaging examination, comprising: an image data processor for processing data representing a plurality of cardiac images of a patient over a plurality of heart beat cycles of said patient to derive data representing a distribution curve of a heart section area over a plurality of heart beat cycle times and indicating heart section area change over a heart beat cycle;an area processor for determining a heart section area in response to user command; anda computation processor for determining a heart function parameter in response to the determined heart section area and the indicated heart section area change. 2. A system according to claim 1, wherein said distribution curve indicates an end-diastolic (ED) location and end-systolic (ES) location andsaid heart section area change comprises a change in area of said heart section indicated by said distribution curve between said ED and ES locations andsaid heart function parameter is determined in response to a single determination of heart section area. 3. A system according to claim 2, wherein said image data processor automatically determines said heart section area change from data comprising said distribution curve. 4. A system according to claim 1, wherein said area processor determines said heart section area in response to user command and selection of at least one of, (a) at least a portion of the heart section area contour and (b) control points of the heart section area. 5. A system according to claim 1, wherein said computation processor determines a heart function parameter comprising a heart ejection fraction. 6. A system according to claim 5, wherein said heart ejection fraction comprises at least one of, (a) a two dimensional heart ejection fraction and (b) a three dimensional heart ejection fraction. 7. A system according to claim 1, wherein said computation processor determines a three dimensional heart function parameter comprising a heart ejection fraction in response to a two dimensional ejection fraction value and values of a left ventricle long apex to aortic axis and short apex to aortic axis. 8. A system according to claim 1, wherein said heart section area comprises a left ventricle area andsaid computation processor determines a heart function parameter comprising a heart ejection fraction in response to a single determination of left ventricle area corresponding to at least one of, (a) an end-diastolic (ED) point and (b) an end-systolic (ES) point. 9. A system according to claim 8, wherein said computation processor determines said heart ejection fraction, in response to combining a determined left ventricle area value and a left ventricle area change value. 10. A system according to claim 1, wherein said heart section comprises at least one of, (a) a Left Ventricle and (b) a Right Ventricle. 11. A system according to claim 1, wherein said computation processor determines a heart... | Siemens Medical Solutions USA, Inc. | A61B005/029 A61B005/02
US20110112415 | A1 | 2011-05-12 | SYSTEM AND METHOD FOR AUTOMATED EKG ANALYSIS | 2009-11-06 | Systems and methods for analyzing electronic cardiac signals for use in clinical diagnostics are described. Parameters pertinent to a first cardiac condition of a patient, such as determining an orientation of a vector related to the cardiac activity of said patient, and comparing the vector orientation relative to a centerpoint of a population distribution representative of a second cardiac condition, may be utilized. The second cardiac condition may be selected from the group consisting of benign early repolarization, left ventricular hypertrophy, and right bundle branch block. System and method embodiments are configured to assist in the analysis of details of EKG signals and vector cardiograms to determine how patients should be categorized into specific cardiac risk categories, such as an acute coronary syndrome category. | 1. A method for determining a parameter pertinent to a first cardiac condition of a patient, comprising: a. numerically analyzing a 3-D vector cardiogram derived from EKG data of a particular patient to determine an orientation of a vector related to the cardiac activity of said patient; andb. comparing the vector orientation relative to a centerpoint of a population distribution representative of a second cardiac condition. 2. The method of claim 1, wherein the vector orientation is the ST or T vector orientation. 3. The method of claim 1, wherein the second cardiac condition is selected from the group consisting of benign early repolarization, left ventricular hypertrophy, and right bundle branch block. 4. The method of claim 1, further comprising selecting a population from which the population distribution is derived, the population being selected based upon a factor included in the group consisting of age, gender, race, residency, citizenship, occupation, and profession. 5. The method of claim 1, further comprising utilizing said comparing as one factor in multifactorial analysis to detect a first cardiac condition of the patient. 6. The method of claim 1, wherein the centerpoint of the population distribution is determined utilizing a mean calculation. 7. The method of claim 1, wherein numerically analyzing comprises analyzing the rate of change of a vector magnitude associated with said 3-D vector cardiogram. 8. A method for detecting a first cardiac condition, comprising: a. providing a quantity of 12-lead EKG data from a patient;b. constructing a three-dimensional representation of cardiac activity from the data;c. numerically analyzing a 3-D vector cardiogram derived from the data to determine an orientation of a vector related to the cardiac activity of said patient;d. comparing the vector orientation relative to a centerpoint of a population distribution representative of a second cardiac condition; ande. automatically drawing one or more conclusions regarding the first cardiac condition of the patient based at least in part upon the comparing. 9. The method of claim 8, wherein the vector position is the ST or T vector position. 10. The method of claim 8, wherein the second cardiac condition is selected from the group consisting of benign early repolarization, left ventricular hypertrophy, and right bundle branch block. 11. The method of claim 8, further comprising selecting a population from which the population distribution is derived, the population being selected based upon a factor included in the group consisting of age, gender, race, residency, citizenship, occupation, and profession. 12. The method of claim 9, further comprising utilizing said comparing as one factor in multifactorial analysis to detect a first cardiac condition of the patient. 13. The method of claim 12, wherein the first cardiac condition comprises a specific categorization of acute coronary syndrome for the patient relative to other patients. 1... | NEWCARDIO, INC | A61B005/0402
US20110137147 | A1 | 2011-06-09 | MINIMUM TIME FEEDBACK CONTROL OF EFFICACY AND SAFETY OF THERMAL THERAPIES | 2006-10-10 | A thermal treatment control system including an imaging device for specifying the geometry and/or location of the treatment target, a thermal energy element for applying a thermal treatment for the heating or cooling of a target tissue for therapeutic purposes, a thermal energy detecting element for detecting a measured tissue response to the thermal treatment and a feedback controller for a real-time modification of the intensity and spatial distribution of the thermal dose in order to achieve therapeutic efficacy over a minimum or reduced treatment time while satisfying treatment constraints imposed to limit damage to normal tissues. | 1. A thermal treatment control system, comprising: an imaging device for specifying a treatment target's geometry and/or location;a thermal energy element for applying a thermal treatment wherein the thermal treatment comprises heating and/or cooling of a target for therapeutic purposes;a thermal energy element for applying cooling and/or heating of normal tissues to prevent or minimize normal tissue damage outside the treatment target;a thermal energy detecting element for detecting a measured tissue response to the thermal treatment; anda feedback controller for a real-time modification of the intensity and spatial distribution of a thermal energy created by the thermal energy element in order to achieve the efficacy objectives of the therapy in a minimum or reduced time while simultaneously satisfying normal tissue safety constraints;wherein the measured tissue response to the thermal treatment is used as feedback by the feedback controller and the real-time modification is made to the operation of the thermal energy element in reaction to the measured tissue response. 2. The thermal treatment control system of claim 1, wherein a prescribed target treatment efficacy in terms of a thermal dose, a temperature distribution, a temperature change, or other treatment parameters is specified for achieving a desired treatment outcome. 3. The thermal treatment control system of claim 1, wherein treatment conditions are specified for ensuring the safety of the tissues outside a treatment target and wherein the treatment parameter is in the form of temperature, temperature change, thermal dose, or limits on thermal response of the normal tissues outside the treatment target. 4. The thermal treatment control system of claim 1, wherein a hardware constraint of the thermal treatment system or thermal element is specified. 5. The thermal treatment control system of claim 1, wherein the thermal energy element for applying a thermal treatment comprises a noninvasive power delivery element. 6. The thermal energy element of claim 5, further comprising single or multiple transducers or transducer arrays for ultrasound heating, radio frequency heating, and/or microwave heating. 7. The thermal treatment control system of claim 1, wherein the thermal energy element for applying a thermal treatment comprises an invasive power delivery element. 8. The thermal energy element of claim 7, further comprising an interstitial microwave, radio frequency and/or optical needles and applicators. 9. The thermal energy element of claim 7, further comprising an interstitial ultrasound element. 10. The thermal treatment control system of claim 1, wherein the thermal energy detecting element for detecting a measured tissue response to the thermal treatment comprises a noninvasive thermal energy detecting element. 11. The thermal treatment control system of claim 10, further comprising means for taking a magnetic resonance temperature measurement. 12. The thermal treatment... | UNIVERSITY OF UTAH RESEARCH FOUNDATION | A61B005/055 A61B018/00
US20110118597 | A1 | 2011-05-19 | SYSTEM AND METHOD FOR SURGICAL INSTRUMENT DISABLEMENT VIA IMAGE-GUIDED POSITION FEEDBACK | 2011-01-25 | A system for selectively disabling a surgical instrument operating in a surgical site of a region of interest of a living subject. In one embodiment, the system includes means for noninvasively placing a number of fiducial markers in an anatomic space of the region of interest of the living subject, means for pre-operatively measuring a location of each fiducial marker in the anatomic space, an imaging acquisition device for pre-operatively acquiring an image volume from the region of interest of the living subject, a probe for intra-operatively monitoring a location of the surgical instrument in the anatomic space, and a controller configured to perform the steps of identifying a centroid of each fiducial marker in the image volume, registering the identified centroid of each fiducial marker in the image volume to the measured location of the corresponding fiducial marker in the anatomic space to determine a registration transformation, mapping the monitored location of the surgical instrument in the anatomic space onto a corresponding location in the image volume by an inverse of the registration transformation, and generating a signal to disable the surgical instrument when the mapped location of the surgical instrument is substantially close to a boundary of the surgical site of the region of interest in the image volume. | 1-19. (canceled) 20. A system for selectively disabling a surgical instrument operating in a surgical site of a region of interest of a living subject, wherein the surgical instrument has a distal end portion and is operable by a surgeon during a surgery, comprising: (a) a probe for intra-operatively monitoring a physiological condition of the living subject in the region of interest during the surgery; and(b) a controller operatively coupled with the surgical instrument and the probe, respectively, and configured to perform during the surgery the steps of: (i) determining a variation of the physiological condition of the living subject from an initial state of the living subject in the region of interest; and(ii) generating a signal to disable the surgical instrument when the variation of the physiological condition is greater then a predetermined value, wherein the probe comprises at least one of a needle electrode placed in the region of interest of the living subject, a Doppler vascular probe, a temperature sensing device, an oxygenation detector, and a tissue boundary detector. 21. The system of claim 20, further comprising at least one of a display and an audio device for visually and/or audibly displaying the intra-operatively monitored physiological condition of the living subject in the region of interest, respectively. 22. (canceled) 23. The system of claim 20, wherein the physiological condition comprises an electromyogram of electrical activities of muscles of the region of interest of the living subject. 24. (canceled) 25. The system of claim 20, wherein the physiological condition comprises a flow rate of blood vessels of the region of interest of the living subject. 26. (canceled) 27. The system of claim 20, wherein the physiological condition comprises a temperature of tissues of the region of interest of the living subject. 28. (canceled) 29. The system of claim 20, wherein the physiological condition comprises a venous oxygenation saturation and an arterial oxygenation saturation of the region of interest of the living subject. 30. (canceled) 31. The system of claim 20, wherein the tissue boundary detector comprises at least one of an ultrasound sensing device, an infrared radiation sensing device, an ultraviolet radiation sensing device, a fluorescent radiation sensing device, and a video frame device. 32. The system of claim 31, wherein the physiological condition comprises a tissue type identification of tissues at the distal end portion of the surgical instrument in the region of interest of the living subject. 33. The system of claim 32, wherein the physiological condition comprises a volume of suction from the region of interest of the living subject. 34. A method for selectively disabling a surgical instrument operating in a surgical site of a region of interest of a living subject, wherein the surgical instrument has a distal end portion and is operable by a surgeon during a surgery, comprising the steps o... | Vanderbilt University | A61B005/05
US20110112390 | A1 | 2011-05-12 | APPARATUS AND METHOD FOR BLOOD CLOTTING DIAGNOSTICS | 2009-04-28 | In order to be able still more precisely to determine the blood clotting activity in a patient according to the invention there is provided a method of extracorporeal determination of the blood clotting activity of a blood sample of an individual, in which the clotting time which elapses until the blood sample reaches a given degree of clotting or the degree of clotting reached by the blood sample within a given time is measured or in which in a corresponding manner the fibrinolysis time or the degree of fibrinolysis of a blood sample which has previously clotted to a certain degree is measured, the method being characterised in that measurement of the clotting/fibrinolysis time or the degree of clotting/fibrinolysis is effected with the exclusion of air. In addition according to the invention there is proposed a corresponding apparatus for the extracorporeal determination of the blood clotting activity wherein the apparatus has a measurement chamber for measuring the clotting time, the degree of clotting, the fibrinolysis time or the degree of fibrinolysis of a blood sample, the apparatus being characterised in that the measurement chamber is of such a configuration that measurement of the clotting/fibrinolysis time or the degree of clotting/fibrinolysis is effected with the exclusion of air. | 1. A method of extracorporeal determination of the blood clotting activity of a blood sample of an individual, in which the clotting time which elapses until the blood sample reaches a given degree of clotting or the degree of clotting reached by the blood sample within a given time is measured or in which in a corresponding manner the fibrinolysis time or the degree of fibrinolysis of a blood sample which has previously clotted to a certain degree is measured, wherein the pH-value is determined for a blood sample aliquot derived from the blood sample, measurement of the clotting/fibrinolysis time or the degree of clotting/fibrinolysis is effected with the exclusion of air and during the entire measurement of the clotting/fibrinolysis time or the degree of clotting/fibrinolysis the predetermined pH-value of the blood sample is maintained, which was determined for the derived blood sample aliquot. 2. A method according to claim 1 wherein the addition of additives to the blood sample is effected pH-neutral by the added substances either being so selected that they have no influence on the pH-value of the blood sample or in the case of a plurality of substances they are so selected that the influence of the individual substances on the pH-value of the blood sample is neutralised. 3. A method according to claim 1 wherein the volume of the total of the added reagents includes at most 5% of the volume of the blood sample in order not to influence the original pH-value of the sample by dilution. 4. A method according to claim 1 wherein all added reagents are unbuffered in order not to influence the original pH-value of the sample by buffer substances. 5. A method according to claim 1 wherein extraction of the blood sample from the body of the individual, storage of the blood sample until the measurement operation and transfer of the blood sample into an apparatus for measuring the clotting/fibrinolysis time and/or the degree of clotting/fibrinolysis are effected with the exclusion of air. 6. A method according to claim 1 wherein the blood sample is taken with a gas-tight syringe or with a glass capillary which is coated on the inside with lithium heparinate. 7. A method according to claim 1 wherein measurement is effected without contact of the blood sample with a gas phase or in an artificial atmosphere, wherein for example carbogen is used as the artificial atmosphere. 8. A method according to claim 1 wherein a barrier gas which is markedly heavier than air and lighter than blood such as for example sulphur hexafluoride (SF6) is used to overlay the blood sample arranged in the measurement chamber so that the blood sample cannot come into contact with the gas phase above the barrier gas. 9. Apparatus for the extracorporeal determination of the blood clotting activity of a blood sample of an individual, wherein the apparatus has a measurement chamber for measuring the clotting time which elapses until the blood sample reaches a certain degree of... | - | A61B005/00
US07959576 | B2 | 2011-06-14 | Apparatus for detecting diastolic heart failure | 2004-05-06 | An implantable medical apparatus for detecting diastolic heart failure, DHF, has a DHF determining device for determining at least one DHF parameter for detecting a DHF state of the heart of a patient. The DHF includes circuitry for determining, as the DHF parameter, the time duration of a predetermined phase of diastole. A pacemaker has such an apparatus and a control unit that optimizes pacing therapy and pacemaker settings depending on the determined time duration. A corresponding method of detecting diastolic heart failure, DHF, includes determining at least one DHF parameter for detecting a DHF state of the heart of a patient. As the DHF parameter, the time duration of a predetermined phase of diastole is determined. | 1. An implantable medical apparatus for detecting diastolic heart failure (DHF) comprising: a sensor configured to interact with a heart to obtain information associated with functioning of the heart; anda DHF determining device supplied with said information that detects a DHF state of the heart from said information by extracting from said information a time duration between predetermined diastolic events in a diastolic phase of the heart representing only a portion of said diastolic phase, and determining that said DHF state exists, and emitting an output signal indicating said DHF state when said time duration exceeds an upper limit value or is below a lower limit value. 2. An apparatus as claimed in claim 1 wherein said DHF determining device comprises a comparator that compares said time duration with said upper limit value and said lower limit value to obtain a comparison result, said comparison result being indicative of said DHF state. 3. An apparatus as claimed in claim 1 wherein said DHF determining device comprises a calculating unit that calculates, from said information from said sensor, said time duration, as a time from an occurrence of peak blood flow velocity through the mitral valve of the heart to a time of occurrence of zero blood flow velocity through the mitral valve of the heart. 4. An apparatus as claimed in claim 3 wherein said calculating unit determines said time duration by extrapolating said mitral blood flow velocity to zero, if an actual occurrence of zero blood flow velocity through the mitral valve does not occur before an atrial contraction of the heart. 5. An apparatus as claimed in claim 4 wherein said calculating unit extrapolates the mitral blood flow velocity to zero by determining a time derivative of blood flow velocity through the mitral valve shortly after said occurrence of said peak blood flow velocity through the mitral valve. 6. An apparatus as claimed in claim 3 wherein said sensor senses an IEGM signal from the heart, and wherein said calculating unit calculates the time of occurrence of said peak blood flow velocity through the mitral valve to the time of occurrence of zero blood flow velocity through the mitral valve from said IEGM. 7. An apparatus as claimed in claim 3 wherein said sensor is an impedance sensor that senses an impedance of the heart, and wherein said calculating unit calculates the time from the occurrence of said peak blood flow velocity through the mitral valve to zero blood flow velocity through the mitral valve from said impedance. 8. An apparatus as claimed in claim 3 wherein said sensor is a sound sensor that detects a sound signal associated with said functioning of the heart, and wherein said calculating unit calculates the time from the occurrence of said peak blood flow velocity through the mitral valve to zero blood flow velocity through the mitral valve from said sound signal. 9. An apparatus as claimed in claim 3 wherein said sensor is an accelerometer that ... | St. Jude Medical AB | A61B005/021
US08577451 | B2 | 2013-11-05 | System and methods for neurologic monitoring and improving classification and treatment of neurologic states | 2009-12-16 | Methods and systems for assessing brain activity and collecting information related to a patient's condition and brain electrical activity are provided. | 1. A method for monitoring or evaluating a neurologic state of a patient, comprising: collecting data related to brain electrical activity of the patient at a first location using at least one electrode;transferring the data related to brain electrical activity to a memory unit located at a second location that is different from the first location;providing a processor configured to classify the neurological state of the patient based on the data related to brain electrical activity and data stored in a database;receiving data related to non-electrical diagnostic information related to the neurological state of the patient based on at least one evaluation technique;updating the database to include the data related to brain electrical activity and the non-electrical diagnostic information related to the neurological state; andcorrelating the data related to brain electrical activity and the data related to non-electrical diagnostic information to allow development of one or more diagnostic and treatment algorithm. 2. The method of claim 1, further including repeating the method with at least one additional patient using the updated database to classify the neurological state of the additional patient. 3. The method of claim 1, further including repeating the method, wherein collecting data related to brain electrical activity is performed with two or more patients at at least two or more different locations. 4. The method of claim 1, further including, upon classifying the neurological state, notifying a predetermined person that a classification has been provided so that the predetermined person can evaluate the patient and provide information related to the neurological state of the patient based on at least one evaluation technique not including brain electrical activity data. 5. The method of claim 1, wherein the classification includes at least one of a psychiatric disorder, an organic neurological disorder, a structural neurological disorder, a metabolic disorder, an intoxicated state, and a normal state. 6. The method of claim 1, wherein the classification includes at least one of an ischemic brain state and a hemorrhagic state. 7. The method of claim 1, wherein the classification includes an abnormal intracranial pressure. 8. The method of claim 1, further including providing a treatment recommendation based on data stored in the database and the classification of the neurologic state of the patient. 9. A method for monitoring or evaluating a neurologic state of a patient, comprising: collecting data related to brain electrical activity of a patient at a first location using at least one electrode;transferring the data related to brain electrical activity to a memory unit located at a second location that is different from the first location;classifying, using a processor, the neurological state of the patient based on the data related to brain electrical activity and data stored in a first database;providing an assessment of the n... | BrainScope Company, Inc. | A61B005/04
US08542017 | B2 | 2013-09-24 | System and method for measuring the shape of an organ of a patient using a magnetic induction radio sensor integrated in a stretchable strap | 2009-12-21 | A system and method for measuring the shape of an object using a magnetic induction radio sensor involves at least partially enclosing the object with a magnetic loop antenna of the magnetic induction radio sensor, where the inductance of the magnetic loop antenna depends on the shape of the object, and providing a particular capacitance at an antenna matching circuit coupled to the magnetic loop antenna in response to the inductance of the magnetic loop antenna such that the magnetic loop antenna and the antenna matching circuit form a resonant circuit and the resonant circuit has a fixed resonant frequency, where the particular capacitance is used to measure the shape of the object. | 1. A magnetic induction radio sensor configured for measuring the shape of an object, the magnetic induction radio sensor comprising: a magnetic loop antenna configured to at least partially enclose the object, wherein the inductance of the magnetic loop antenna depends on the shape of the object, wherein the magnetic loop antenna is configured to deform as a result of expansion and/or contraction of the object;an antenna matching circuit coupled to the magnetic loop antenna, wherein the antenna matching circuit comprises an adjustable capacitance module configured to provide a particular capacitance such that the magnetic loop antenna and the antenna matching circuit form a resonant circuit and the resonant circuit has a fixed resonant frequency; anda measuring unit configured to generate a measurement value using the particular capacitance of the adjustable capacitance module, the measurement value representing a measurement of the shape of the object,wherein the object is an organ of a patient, and wherein the magnetic loop antenna is integrated in a stretchable strap that at least partially encloses the organ of the patient. 2. The magnetic induction radio sensor of claim 1, wherein the adjustable capacitance module comprises an array of switchable capacitors in order to provide the particular capacitance. 3. The magnetic induction radio sensor of claim 2, wherein the array of switchable capacitors comprises capacitor branches that are connected in parallel with the magnetic loop antenna, and wherein each of the capacitor branches comprises a switch and a capacitor. 4. The magnetic induction radio sensor of claim 3, wherein capacitances of the capacitors in the capacitor branches have a numeric relationship of factors of two. 5. The magnetic induction radio sensor of claim 2, wherein capacitances of the capacitors in the capacitor branches are equal. 6. The magnetic induction radio sensor of claim 1, comprising multiple magnetic loop antennas configured to at least partially enclose the object, wherein the inductances of the magnetic loop antennas depend on the shape of the object, the magnetic induction radio sensor further comprising an array of switches configured for selecting the magnetic loop antennas. 7. The magnetic induction radio sensor of claim 1 further comprising a transmitter configured to transmit the generated measurement value to a remote device using the magnetic loop antenna. 8. The magnetic induction radio sensor of claim 7 further comprising a scheduling unit configured to repetitively schedule measurement of the shape of the object and transmission of the generated measurement value to the remote device according to predefined slots in a time division multiplexing scheme. 9. The magnetic induction radio sensor of claim 8, wherein the scheduling unit is further configured to schedule the measurement of the shape of the object after reception of a synchronization word from the remote device. 10. The magnetic induc... | NXP B.V. | G01R033/36 A61B005/055
US08491474 | B2 | 2013-07-23 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2011-01-27 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. A system for processing a data stream obtained from a glucose sensor, the system comprising: a signal processing module configured to monitor a data stream from a continuous glucose sensor over a period of time;a detection module configured to determine a severity of signal artifacts in the data stream based on an amplitude and a duration of the signal artifacts;a signal estimation module configured to selectively apply one or more coefficients or parameters to a signal estimation algorithm based on the severity of the signal artifacts and to apply the signal estimation algorithm to the data stream to generate an estimated signal. 2. The system of claim 1, wherein the signal estimation algorithm comprises a recursive filter. 3. The system of claim 2, wherein the signal estimation algorithm comprises an infinite impulse response filter. 4. The system of claim 1, wherein the signal estimation algorithm comprises a non-recursive filter. 5. The system of claim 4, wherein the signal estimation algorithm comprises an finite impulse response filter. 6. The system of claim 1, wherein the severity of signal artifacts comprise a numeric value. 7. The system of claim 6, wherein signal estimation module is configured to calculate the one or more coefficients or parameters from the numeric value. 8. The system of claim 1, further comprising an output module configured to transmit or display a representation of the estimated signal. 9. A method for processing a data stream obtained from a glucose sensor, the method comprising: monitoring a data stream from a continuous glucose sensor over a period of time;determining, using processing electronics, a severity of signal artifacts in the data stream based on an amplitude and a duration of the signal artifacts;selectively applying one or more coefficients or parameters to a signal estimation algorithm based on the severity of the signal artifacts; andapplying the signal estimation algorithm to the data stream to generate an estimated signal. 10. The method of claim 9, wherein the signal estimation algorithm comprises a recursive filter. 11. The method of claim 10, wherein the signal estimation algorithm comprises an infinite impulse response filter. 12. The method of claim 10, wherein the signal estimation algorithm comprises an finite impulse response filter. 13. The method of claim 9, wherein the signal estimation algorithm comprises a non-recursive filter. 14. The method of claim 9, wherein the severity of signal artifacts comprise a numeric value. 15. The method of claim 14, wherein selectively applying comprises calculating the one or more coefficients or parameters from the numeric value. 16. The method of claim 9, further comprising transmitting or displaying a representation of the estimated signal. | DexCom, Inc. | A61B005/00 A61B005/05
US08684939 | B2 | 2014-04-01 | Blood pressure measurement device | 2011-02-24 | A blood pressure measurement device is mounted with a radio-controlled clock function in a main body without lowering reception performance of the standard radio wave of the radio-controlled clock. A substrate including a sensor mounting surface and a sensor non-mounting surface is mounted with a pressure sensor. An antenna for receiving the standard radio wave including time information of the radio-controlled clock is mounted. The antenna includes a bar-shaped magnetic body core and a coil wound around the magnetic body core. A pump is arranged such that an axis line direction of the motor and an extending direction of the magnetic body core are substantially orthogonal. The substrate and the substrate are arranged so that the sensor non-mounting surface of the substrate and the antenna mounting surface of the sensor face each other. | 1. A blood pressure measurement device comprising: a cuff, which is configured to be attached to a blood pressure measurement site of a person to be measured and which includes a gas bag filled with gas;a pump for transferring gas to the gas bag;a flow rate control valve for controlling a gas flow rate discharged from the gas bag;a pressure sensor for detecting pressure in the gas bag;a first substrate including a sensor mounting surface and a sensor non-mounting surface, and having the pressure sensor mounted on the sensor mounting surface;a radio-controlled clock, including an antenna for receiving a standard radio wave including time information, for timing a current time; anda second substrate including an antenna mounting surface and an antenna non-mounting surface, and having the antenna mounted on the antenna mounting surface;wherein the pump includes a motor that rotates about an axis line and operates the pump,wherein the antenna includes a bar-shaped magnetic body core and a coil wound around the magnetic body core,wherein the pump is arranged such that an axis line direction of the motor and an extending direction of the magnetic body core are substantially orthogonal; andwherein the first substrate and the second substrate are arranged so that the sensor non-mounting surface and the antenna mounting surface face each other, or so that the sensor mounting surface and the antenna non-mounting surface face each other. 2. The blood pressure measurement device according to claim 1, wherein the second substrate is arranged parallel to an installing surface where the blood pressure measurement device is installed. 3. The blood pressure measurement device according to claim 1, wherein the first substrate is arranged inclined with respect to the second substrate. 4. The blood pressure measurement device according to claim 1, wherein the first substrate and the second substrate are arranged in parallel. 5. The blood pressure measurement device according to claim 1, wherein the flow rate control valve is an electromagnetic drive valve; andwherein the device is arranged such that a direction of a magnetic field generated by the electromagnetic drive valve and the extending direction of the magnetic body core are orthogonal. 6. The blood pressure measurement device according to claim 1, wherein the second substrate is interposed between the flow rate control valve and the antenna; andwherein the flow rate control valve is installed so that the antenna non-mounting surface and the flow rate control valve face each other. | OMRON HEALTHCARE Co., Ltd. | A61B005/02
US08750952 | B2 | 2014-06-10 | Apparatus and method for dating a body sample | 2009-04-21 | A method and apparatus for dating a body sample, for example a sample of body fluid involves taking a series of spectroscopic measurements of the sample, each measurement in the series including at least two predetermined positions in the spectrum. The positions have spectral characteristics corresponding to two or more predetermined substances present in the sample that have a time varying relationship with each other. The measurements in the series are spaced in time. A concentration of each of the substances present in the sample is then determined from each of the spectroscopic measurements at each point in time. Next, a ratio of the concentrations of the two predetermined substances at each point in time is determined and then the ratios of the concentrations of the two predetermined substances over time are analyzed to estimate when the concentrations of the two substances were at a limit of their concentrations, thereby providing an indication of the age of the sample. The sample may be blood, which may be either within the body (e.g. as a bruise), or external to the body. | 1. A method of dating a body sample comprising: taking a spectroscopic measurement of a host material;taking a series of spectroscopic measurements of the sample, each measurement in the series of spectroscopic measurements including at least two predetermined positions in a spectrum, the predetermined positions having spectral characteristics corresponding to two predetermined substances present in the sample that have a time varying relationship with each other, the measurements in the series of spectroscopic measurements being spaced in time;determining a concentration of each of the two predetermined substances present in the sample from each measurement in the series of the spectroscopic measurements at each point in time;determining a ratio of the concentrations of the two predetermined substances at each point in time;analyzing the ratios of the concentrations of the two predetermined substances over time to estimate when the concentrations of the two predetermined substances were at a limit of their concentrations, thereby providing an indication of the age of the sample; andwherein determining the concentrations of the two predetermined substances from the series of spectroscopic measurements comprises: dividing a measured absorbance spectrum obtained from the series of spectroscopic measurements of the sample by an assumed optical path length to provide a corrected absorbance spectrum;fitting measured spectra of the two predetermined substances to the corrected absorbance spectrum to obtain a residue, andfitting the spectroscopic measurement of the host material to the residue to obtain relative concentrations of the two predetermined substances within the sample. 2. A method of dating a body sample according to claim 1, wherein the sample is either inside or outside the body. 3. A method of dating a body sample according to claim 1, wherein the sample is skin. 4. A method of dating a body sample according to claim 1, wherein the sample is a body fluid. 5. A method of dating a body sample according to claim 4, wherein the body fluid is blood. 6. A method of dating a body sample according to claim 5, wherein one of the two predetermined substances is hemoglobin and the other substance is met-hemoglobin. 7. A method of dating a body sample according to claim 1, wherein the series of spectroscopic measurements comprises reflectance, Raman or fluorescence spectroscopic measurements. 8. A method of dating a body sample according to claim 1, wherein one of the two predetermined substances is hemoglobin and the other of the two predetermined substances is bilirubin. 9. A method of dating a body sample according claim 1, wherein the series of spectroscopic measurements comprise interferometric spectroscopic measurements made at a series of depths in the body. 10. A method of dating a body sample according to claim 1, wherein the series of spectroscopic measurements include measurements made at several different lateral positions on the... | Academisch Medisch Centrum | A61B005/1455
US07945865 | B2 | 2011-05-17 | Information processing system, information processing apparatus, and method | 2006-10-06 | A degree of confidence is objectively evaluated without permitting a user's subjective evaluation, and a content to be output is determined based on this evaluation.;An information processing system comprising: an input section for receiving an input from a user; a signal detection section for measuring a signal concerning an event-related potential of electroencephalograms of the user; a determination section for determining a degree of confidence of the user with respect to the input based on an amount of negative shift in the event-related potential during a predetermined period after the input is received; a control section for determining a content to be presented to the user based on the degree of confidence; and an output section for presenting the determined content to the user. | 1. An information processing system comprising: an input section for receiving an input from a user;a signal detection section for measuring a signal concerning an event-related potential of electroencephalograms of the user;a determination section for determining a degree of confidence of the user with respect to the input based on an amount of negative shift in voltage in the measured event-related potential of electroencephalograms of the user during a predetermined period, the predetermined period being subsequent to a receipt of the input from the user;a control section for determining a content to be presented based on the degree of confidence; andan output section for presenting the determined content. 2. The information processing system of claim 1, wherein, the output section presents a question to the user;as the input, the input section receives an answer of the user to the question which is output by the output section;based on the degree of confidence, the control section determines to present information related to the question or related to a correct answer thereto; andthe output section presents the information after lapse of the predetermined period. 3. The information processing system of claim 1, wherein the determination section determines the degree of confidence of the user by defining as the predetermined period a period within about 2000 milliseconds starting from a point in time of receiving the input. 4. The information processing system of claim 2, further comprising a database storing information related to the question and related to the correct answer, wherein, the control section determines correctness of the answer of the user by referring to the database, and based on the degree of confidence of the user and a determination result of correctness, selects the information related to the question or related to the correct answer thereto. 5. The information processing system of claim 4, wherein, the database stores information of a plurality of hints related to the correct answer; andthe control section selects from among the hints based on the degree of confidence of the user and the determination result of correctness. 6. The information processing system of claim 4, wherein the control section outputs information representing a degree of understanding of the user, based on the degree of confidence of the user and the determination result of correctness. 7. The information processing system of claim 1, wherein, the amount of negative shift in voltage in the event-related potential has a corresponding relationship with a gradient of a waveform of negative shift in voltage; andthe determination section retains a threshold which falls between a gradient value of negative shift in voltage corresponding to ‘confident’ and a gradient value of negative shift in voltage corresponding to ‘unconfident’, andby comparing a gradient of a waveform of the acquired event-related potential against the threshold, makes a ‘confi... | Panasonic Corporation | G06F017/30 A61B005/00
US08326395 | B2 | 2012-12-04 | Electrode for electroretinographic use and method of application | 2009-11-19 | A flexible loop electrode for electroretinographic (ERG) recordings of extra-corneal signals with a protective bead and a U-shaped electrode form that is particularly adapted for use with spectacles or a fixed mount. Means also employed to detect electrode integrity. | 1. A conductive electrode assembly for the measurement of corneal electroretinographic (ERG) signals, the electrode assembly comprising: a bead shaped electrode adapted to contact a subject's eye; anda flexible wire supporting the bead shaped electrode, wherein the wire is configured to be mounted to enable the bead shaped electrode to make contact with a junction between a cornea and sclera of the subject's eye. 2. The electrode assembly of claim 1 wherein the wire is constructed of a super-elastic alloy, conductive elastomer, conductive plastic, conductive glass capillary, or a conductive glass fiber. 3. The electrode assembly of claim 1 wherein the electrode is fabricated from a silver alloy, a gold alloy, a platinum alloy, a titanium alloy, or a base metal alloy with a protective coating. 4. The electrode assembly of claim 1 wherein the electrode is fabricated from a conductive elastomer, a conductive plastic, or a conductive glass. 5. The electrode assembly of claim 1 wherein the electrode is fabricated from an insulating material with a conductive coating. 6. The electrode assembly of claim 1 wherein the electrode is fixed to the flexible wire. 7. The electrode assembly of claim 1 wherein the electrode is free to move on the flexible wire, or is constrained to only rotate. 8. The electrode assembly of claim 1 wherein the flexible wire is shaped into a U-form with the electrode mounted at the bottom of the U shape. 9. The electrode assembly of claim 1 wherein the wire is adapted to be mounted to a fixture at an angle so as to align the electrode against the subject's eye but not obscuring a pupil of the subject's eye or an ERG illumination light source. 10. The electrode assembly of claim 1 wherein the electrode assembly is constructed to sense current, voltage, resistance or impedance. 11. An electrode assembly for measuring electroretinographic (ERG) signals from a subject's eye, the electrode assembly comprising: an electrode adapted to contact a subject's eye; anda wire supporting the electrode between first and second ends of the wire, wherein the first and second ends of the wire are configured to be mounted to enable the electrode to make contact with a junction between a cornea and sclera of the subject's eye. 12. The electrode assembly of claim 11 wherein the wire is constructed of a flexible material. 13. The electrode assembly of claim 11 wherein the electrode is bead shaped. 14. The electrode assembly of claim 11 wherein the electrode is fixed to the wire. 15. The electrode assembly of claim 11 wherein the electrode is adapted to rotate about the wire and/or move along the wire. 16. The electrode assembly of claim 11 wherein the wire is adapted to be mounted to a fixture. 17. A method of applying an electrode to a patient's eye, the method comprising: providing an electrode assembly for contacting a subject's eye, the electrode assembly including an electrode and a wire supporting the electrode between first and s... |  | A61B005/04
US07945458 | B1 | 2011-05-17 | Care funding and care planning system | 2008-04-29 | A care funding and care planning system (10) generates a care plan option report (192) for a care receiver. The system (10) includes a client computer (14) having an input data map (16) for receiving and storing care-receiver data (13) including predetermined, critical categories of care-receiver information. A system processor (36) processes the care-receiver data (13) through knowledge management software (12) to evaluate and select from the care-receiver data (13) at least physical functional status (72), cognitive and sensory status (74), prospective functional status (76), living environment status (84), and long-term care resource status abstractions (88). A meta needs-resource weighting engine (56) assigns care-receiver specific values (160A) to the care-receiver data abstractions (72). A dynamic data base (48) allocates values to the weighted care-receiver data abstractions (72). A report generator (58) produces a care plan options report (192) for the care receiver. | 1. A care funding and care planning system (10) for generating a care plan option report (192) for a care receiver, the system (10) comprising: a. a client computer (14) for receiving and storing care-receiver data, the client computer (14) including a user interface data map (16) configured to select predetermined care-receiver data (13) from a plurality of critical categories of care-receiver information, the care-receiver data (13) being received and stored by the client computer (14), wherein the critical categories of care-receiver information include the care-receiver's current physical functional status (72), current cognitive and sensory functional status (74), prospective functional status including physical, cognitive and sensory information (76), living environmental status abstractions (84), and long term care resource abstractions (88);b. a system processor (36) in communication with the care-receiver data (13) stored within the client computer (14), the system processor being configured to process the critical categories (72, 74, 76, 84, 88) of care-receiver information of the care-receiver data (13) through knowledge management software (12), the system processor (36) including data abstraction processing capability (24) configured to select and evaluate the care-receiver data (13) critical categories (72, 74, 76, 84, 88) of information to produce a current care needs budget (96) from the care-receiver's current physical functional status (72) and current cognitive and sensory functional status (74), a prospective care needs budget (114) from the prospective functional status (76), a home modification budget (118) from the living environmental status abstractions (84), and a summary of overall resources (150) from the long terra care resource abstractions (88), the budgets (96, 114, 118) and summary of overall resources (150) being referred to as care-receiver data abstractions (96, 114, 118, 150);c. a data base (48) in communication with the system processor (36) for providing values to the critical categories of information (72, 74, 76, 84, 88);d. a meta needs resource weighting engine (56) in communication with the system processor (36) and configured to weigh and process the care-receiver data abstractions (96, 114, 118, 150) to produce a summary care planning budget (160A, 160B, 160C, 160D, 160E, 160F) from the care-receiver data abstractions (96, 114, 118, 150) by processing the care-receiver data abstractions (96, 114, 118, 150) through at least one of Analysis A (164), Analysis B (172), Analysis C (182), Analysis D (184), Analysis E (210), and Analysis F (218); and,e. a report generating output (58) in communication with the system processor (36) and configured to produce the care plan options report (192) for the care receiver from the summary care planning budget (160A, 160B, 160C, 160D, 160E, 160F). 2. The care funding and care planning system (10) of claim 1, wherein the long-term care resource abstractions (88) includ... | - | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US08306597 | B2 | 2012-11-06 | Physiological signal sensing device | 2010-03-10 | A physiological signal sensing device for examination of human is provided. The physiological signal sensing device includes a light emitting fiber and a light receiving fiber. The light emitting fiber includes a plurality of light emitting portions, wherein the light emitting fiber provides a plurality of sensing beams, and the sensing beams are respectively emitted through the light emitting portions. The light receiving fiber includes a plurality of light receiving portions. The light receiving fiber corresponds to the light emitting fiber. The sensing beams are emitted through the light emitting portions, reflected or refracted by the human. And then the sensing beams are received by the light receiving portions. | 1. A physiological signal sensing module for examination of a human, comprising: a plurality of physiological signal sensing devices, wherein the physiological signal sensing devices are arranged along a first direction and a second direction, and the physiological signal sensing devices are interlaced to be a reticular structure, wherein each of the physiological signal sensing devices comprises: a light emitting fiber, comprising a plurality of light emitting portions, wherein the light emitting fiber provides a plurality of sensing beams, and the sensing beams are respectively emitted through the light emitting portions;a light receiving fiber, comprising a plurality of light receiving portions, wherein the light receiving fiber corresponds to the light emitting fiber, and the sensing beams are emitted through the light emitting portions, reflected or refracted by the human, and are received by the light receiving portions; anda soft carrier, wherein the light receiving fiber and the light emitting fiber are disposed on the soft carrier side by side. 2. The physiological signal sensing module as claimed in claim 1, wherein the light receiving fiber and the light emitting fiber are partially embedded in the soft carrier. 3. The physiological signal sensing module as claimed in claim 2, wherein the light emitting fiber has a first protrusion, the first protrusion protrudes on a surface of the soft carrier, and the light emitting portion is formed in the first protrusion. 4. The physiological signal sensing module as claimed in claim 3, wherein the light receiving fiber has a second protrusion, the second protrusion protrudes on the surface of the soft carrier, and the light receiving portion is formed in the second protrusion. 5. The physiological signal sensing device as claimed in claim 1, wherein the soft carrier is made of polyurethane, polyimide or silica gel. 6. The physiological signal sensing module as claimed in claim 1, each physiological signal sensing device further comprising a total reflection layer, wherein the total reflection layer is formed on a surface of the soft carrier contacting the light receiving fiber and the light emitting fiber. 7. The physiological signal sensing module as claimed in claim 1, wherein the light emitting fiber has a first fiber core and a first fiber cover, the first fiber cover covers the first fiber core, the light emitting portions are formed on the first fiber cover, a sensing light travels in the first core, and the sensing light is emitted through the light emitting portions to be the sensing beams. 8. The physiological signal sensing module as claimed in claim 1, wherein the light receiving fiber has a second fiber core and a second fiber cover, the second fiber cover covers the second fiber core, the light receiving portions are formed on the second fiber cover, and the sensing beams enter the second core through the light receiving portions and travel therein. 9. The physiological sign... | Industrial Technology Research Institute | A61B005/1455 A61B006/00
US07951080 | B2 | 2011-05-31 | On-body medical device securement | 2009-10-30 | Devices and methods for maintaining a medical device on-body are provided. Embodiments include medical device securement systems having first and second on-body securement elements. Also provided are systems and kits for use maintaining a medical device on-body. | 1. A method for maintaining a control unit of an analyte monitoring system on a body part of a subject using an on-body securement element comprising an adhesive surface and a non-adhesive surface that is opposite to the adhesive surface, the on-body securement element being protected by a first removable layer adjacent the adhesive surface and a second removable layer adjacent the non-adhesive surface, the method comprising: removably attaching the analyte monitoring system control unit to a body part of a patient;after attaching the control unit to the body part, positioning the on-body securement element about the control unit so that the adhesive surface contacts the control unit; andremoving the first and second removable layers from the on-body securement element. 2. The method of claim 1, further comprising attaching the on-body securement element to a mounting unit. 3. The method of claim 1, further comprising controlling a continuous analyte monitoring device with the control unit. 4. The method of claim 3, wherein the analyte is glucose. 5. The method of claim 3, wherein the continuous analyte monitoring device is a glucose sensor, wherein at least a portion of the sensor is positionable under the skin of the subject. 6. The method of claim 1, further comprising maintaining rigidity of the on-body securement element by at least one of the removable layers. 7. The method of claim 1, wherein the on-body securement element further comprises a hole therein, and wherein positioning of the on-body securement element about the control unit comprises positioning the securement element hole around the control unit. 8. The method of claim 1, wherein at least the first removable layer comprises a plurality of release liners. 9. The method of claim 8, further comprising removing at least a first release liner of the first removable layer after attaching the control unit to the body part and prior to positioning the on-body securement element about the control unit. 10. The method of claim 9, further comprising removing at least a second release liner of the first removable layer after positioning the on-body securement element about the control unit. 11. The method of claim 1, wherein upon positioning the on-body securement element about the control unit, the adhesive surface faces the body part and the non-adhesive surface faces away from the body part. 12. A method for maintaining a control unit of an analyte monitoring system on a body part of a subject, the method comprising: attaching a first adhesive surface of a first securement element to a body part;attaching a second adhesive surface of the first securement element to the control unit, wherein the second adhesive surface is opposite the first adhesive surface;removing at least one removable liner from an adhesive surface of a second securement element, wherein the second securement element further comprises a hole therein; andpositioning the second securement element about the ... | Abbott Diabetes Care Inc. | A61B005/00
US07938783 | B2 | 2011-05-10 | Medical body core thermometer | 2004-06-18 | A medical thermometer with a miniature sensor that touches the patient's skin by a spring-loaded probe. The device computes a deep body temperature by using data of probe housing temperature and accounting for multiple responses of skin contact temperature sensor before and after touching the skin. A motion detector is employed to turn power on automatically. | 1. A non-invasive contact thermometer for intermittent determination of body core temperature of a patient comprising: a housing containing a probe;a guard being part of said probe;a first temperature sensor shielded by said guard, said guard being configured to slide within said probe and relative to said first temperature sensor after said probe is pressed against skin of the patient to bring said first temperature sensor in thermal contact with the skin;a thermally conductive plate attached to said first temperature sensor;a microcontroller for receiving an electrical signal from said first temperature sensor and for computing the body core temperature of the patient; anda display communicating with said microcontroller for presenting the body core temperature computed by said microcontroller. 2. A thermometer of claim 1 wherein said first temperature sensor is positioned on a movable shaft fabricated of a thermally insulating material. 3. A thermometer of claim 1 further comprising a second temperature sensor that is thermally insulated from the first temperature sensor. 4. A method of performing deep body temperature measurement from skin of a patient using a contact thermometer having a probe with a guard and a first temperature sensor shielded by the guard, the guard being movable relative to the first temperature sensor, the method comprising: measuring temperature of the first temperature sensor before the first temperature sensor is brought in thermal contact with the skin of the patient,pressing the guard against the skin to slide the guard within the probe and thereby bring the first temperature sensor in thermal contact with the skin,measuring temperature of the first temperature sensor after the first temperature sensor is brought in thermal contact with the skin,computing a deep body temperature of the patient by accounting for the temperatures measured by the first temperature sensor prior to and after contacting the skin and by experimentally determined constants, anddisplaying the deep body temperature. 5. A method of performing deep body temperature measurement of claim 4, further comprising: measuring temperature of the first temperature sensor at least two times after the first temperature sensor is brought in thermal contact with the skin of the patient. 6. A method of performing deep body temperature measurement of claim 4, further comprising: elevating the temperature of the first temperature sensor prior to pressing the guard against the skin of the patient. 7. A thermometer of claim 1, further comprising: a power supply communicating with said microcontroller and said first temperature sensor. 8. A thermometer of claim 1, wherein said plate is constructed from metal. 9. A thermometer of claim 1, wherein said microcontroller computes the patient's body core temperature from at least two sequential responses of said first temperature sensor. 10. A thermometer of claim 1, further comprising: a heating element therm... | Advanced Monitors Corporation | A61B005/00
US08246553 | B2 | 2012-08-21 | Muscle hardness meter | 2007-08-23 | A muscle hardness meter according to the present invention, which measures the hardness of muscle tissue of a living organism, includes: an auxiliary part which touches a periphery of a section of the living organism to be measure so as to apply pressure to the periphery; a touching part which touches the section to be measured so as to apply pressure to the section to be measured; a first pressure sensor which measures pressures that the touching part and the auxiliary part sustain, respectively, from the section to be measured and the periphery, and outputs a first measurement result; a second pressure sensor which measures a pressure that the touching part sustains from the section to be measure, and outputs a second measurement result; a notification part which notifies the second measurement result; and a control part which determines whether or not the first measurement result teaches a reference pressure value input in advance and notifies the second measurement result to the notification part at a time when it is determined that the first measurement result reaches the reference value. | 1. A muscle hardness meter which measures hardness of muscle tissue of a living organism, the muscle hardness meter comprising: an auxiliary part which is configured to touch a periphery of a section of the living organism to be measured so as to apply pressure to the periphery;a touching part which is configured to touch the section to be measured so as to apply pressure to the section to be measured;a first pressure sensor which, measures both of a first pressure that the touching part sustains from the section to be measured and a second pressure that the auxiliary part sustains from the periphery, and outputs a total of the first pressure and the second pressure as a first measurement result;a second pressure sensor which measures the first pressure that the touching part sustains from the section to be measured, and outputs the first pressure as a second measurement result;a notification part which notifies the second measurement result; anda control part which determines whether or not the first measurement result reaches a reference pressure value input in advance and notifies the second measurement result to the notification part at a time when it is determined that the first measurement result reaches the reference value. 2. A muscle hardness meter according to claim 1, further comprising: an input section which inputs the reference pressure value in advance, wherein the input section includes a plurality of pressure value selecting sections in which a plurality of pressure values determined in stages in advance are set, wherein the pressure value set in one of the plurality of pressure value selecting sections is set as the reference pressure value. 3. A muscle hardness meter according to claim 2, wherein the plurality of pressure value selecting sections are at least three. 4. A muscle hardness meter according to claim 2, further comprising a change section which changes the pressure values set for each of the plurality of pressure value selecting sections. 5. A muscle hardness meter according to claim 4, wherein the change section includes a plurality of ratio selecting parts in which change ratios determined in stages in advance are set, wherein the pressure value that is changed according to the change ratio set by one which is selected from the plurality of ratio selecting parts is set as the reference pressure value. 6. A muscle hardness meter according to claim 1, wherein the control part calculates an average value of a plurality of the second measurement results for a plurality of measurements. | Ito Co., Ltd. | A61B019/00 A61B005/00 A61B005/103 A61B005/117 G01N003/48 G01N003/42
US08755855 | B2 | 2014-06-17 | Method and device for measuring dissolved substances in human or animal intraocular fluid | 2009-03-02 | Non-invasive optical measurement of glucose and other dissolved substances in human or animal intraocular fluid. A method takes advantage of the fact that the wave dependence of optical activity is fundamentally different from corneal birefringence. The optical activity of substances dissolved in the intraocular fluid, such as glucose, lactate, ascorbic acid or amino acids, is scaled as a first approximation with the reciprocal value of the wavelength square. In contrast, corneal birefringence is scaled with the reciprocal value of the wavelength and therefore behaves considerably different from the optical activity. For the method according to the invention, a physical model is used, which describes the influence on the polarization of measurement radiation by the components of the eye, particularly by the intraocular fluid and the cornea. | 1. A method for determining concentrations of optically active substances in the aqueous humour, comprising radiating polarized measurement radiation into the eye,polarizing reflected measurement radiation reflected by the lens of the eye linearly, circularly or elliptically with a polarizer,quantitatively recording the reflected measurement radiation with a spectrometric array,comparing a measured wavelength dependency of a recorded radiation power with a simulated wavelength dependency of a radiation power simulated in a model calculation,varying the unknown concentration values of the optically active substances via their physiologically possible range in the model calculation, andidentifying concentration values in which the smallest deviations between the measured and the simulated wavelength dependencies of the radiation power occur as measurement concentration values. 2. The method according to claim 1, further comprising selecting the optically active substance from the group consisting of glucose, lactate and ascorbic acid. 3. The method according to claim 1, further comprising measuring reference radiation which is split off between a radiation source and the polarizer simultaneously or quasi-simultaneously and spectrometrically resolved and using a spectral intensity correction of the measurement values of the spectrometric array. 4. The method according to claim 1, further comprising setting a linear or approximately linear polarization of the measurement radiation in front of the polarizer by application of a variable or fixed phase shifter. 5. The method according to claim 1, wherein the polarizer is a linear polarizer and further comprising setting the polarization direction of the linear polarizer such that the polarization direction of the linear polarizer forms an angle of between 20 and 70 degrees, with a main axis of a polarization ellipse of the measurement radiation striking the polarizer. 6. The method according to claim 1, wherein the polarizer is a linear polarizer and further comprising setting the polarization direction of the linear polarizer such that the polarization direction of the linear polarizer forms an angle of between 40 and 50 degrees, with the main axis of the polarization ellipse of the measurement radiation striking the polarizer. 7. The method according to claim 1, further comprising choosing an angle of incidence of the measurement radiation on a cornea individually or for the majority of eyes such that corneal birefringence and thus a value θm become minimal. 8. The method according to claim 1, further comprising presenting the eye with a spatially varying fixation target from a luminous display in an alignment device by which the eye is animated in a targeted manner to perform artificial saccades to influence a value Δθ in a targeted manner. 9. The method according to claim 1, further comprising preparing a polarization state of the measurement radiation valid for an LED centroid wavelength in... | Carl Zeiss AG | A61B005/1455
US08581731 | B2 | 2013-11-12 | Circuits, systems, and methods for monitoring and reporting foot impact, foot placement, shoe life, and other running/walking characteristics | 2011-02-16 | Various embodiments of the present invention provide an electronic virtual running/walking coach that notifies a person as to how hard the person's feet have been impacting a surface, notifies the person as to whether the person's feet are impacting a surface at a proper relative distance from the body and/or at a proper side-to-side angle, and/or notifies the person as to how much cushioning life remains in a pair of shoes. When the notifications are provided to the person during the running/walking activity, the person can respond by taking immediate actions to adjust the foot placement relative to the body and/or adjust the level of foot impact to an acceptable level by varying the speed, step distance, and/or posture while running/walking and/or by selecting among different available surfaces on which to run/walk. A person may also more accurately determine when a particular pair of shoes should be replaced because they no longer provide a sufficient level of cushioning to compensate for the levels of foot impact that are occurring. | 1. A foot impact monitoring system comprising: an impact measurement circuit configured to measure impact from a foot repetitively striking a surface while a person is walking/running; andan impact alert circuit that is configured to regulate a background sound component in response to measured impacts, and to combine the background sound component with a music component generated by a music player to generate a combined signal that is played to the person to audibly indicate to the person how much impact occurred from the foot striking the surface. 2. The foot impact monitoring system of claim 1, wherein: the impact measurement circuit comprises an accelerometer configured to measure acceleration and/or rate of change of acceleration as the measured impact from the foot striking the surface while the person is walking/running. 3. The foot impact monitoring system of claim 1, wherein: the impact measurement circuit comprises a force transducer configured to measure force as the measured impact from the foot striking the surface while the person is walking/running. 4. The foot impact monitoring system of claim 1, wherein: the impact alert circuit is configured to regulate the background sound component to generate an indication for the person of a peak magnitude of the measured impact. 5. The foot impact monitoring system of claim 1, wherein: the impact alert circuit is further configured to generate a baseline threshold level in response to measured impacts, and to respond to a presently measured impact exceeding the baseline threshold level by causing the background sound component signal to be audible within the combined signal played through the music player to the person to provide an audible warning to the person when the measured impact from the foot striking the surface has become excessive. 6. The foot impact monitoring system of claim 1, wherein: the impact alert circuit is further configured to: generate at least one baseline threshold level in response to measured impacts;increase loudness of the background sound component relative to the music component of the combined signal being played to the person in response to a presently measured impact exceeding one of the at least one baseline threshold level by an increased amount to audibly indicate to the person when the presently measured impact from the foot striking the surface has increased; anddecrease loudness of the background sound component relative the music component of the combined signal being played to the person in response to a presently measured impact exceeding one of the at least one baseline threshold level by a decreased amount to audibly indicate to the person when the presently measured impact from the foot striking the surface has decreased. 7. A foot impact monitoring system comprising: an impact measurement circuit configured to measure impact from a foot repetitively striking a surface while a person is walking/running;a sound generation device that is con... | - | G08B023/00
US09987452 | B2 | 2018-06-05 | Ventilation analysis and monitoring | 2009-06-30 | A ventilation analysis system comprising an interface module adapted to receive a carbon dioxide (CO2) signal and a breath flow dynamics signal, and control logic adapted to produce a ventilation indicator, based on a mutual analysis of the CO2 signal and the breath flow dynamics signal. The ventilation indicator may be, for example, an estimated CO2 waveform, an End-Tidal CO2 (ETCO2) value, and/or a minute CO2 value. | 1. A method for ventilation analysis, the method comprising: receiving discrete partial pressure (PCO2) signals from a CO2 sensor and breath flow dynamics (BFD) signals from a BFD sensor; wherein said CO2 sensor has a response time of more than 100 ms;constructing, using a processor, an estimated CO2 waveform by way of mutual analysis of the PCO2 and BFD signals, wherein said estimated CO2 waveform depicts partial pressure of CO2 in exhaled breath as a function of time, and wherein constructing the estimated CO2 waveform comprises determining a start point and an end point of exhalation and a duration of said exhalation based on said BFD signals;determining, based on said CO2 signals, an average PCO2 value obtained during the exhalation determined based on said BFD signals;constructing an estimated PCO2 waveform based on a shape of a common CO2 waveform, the determined said start point of exhalation, the determined said end point of exhalation, the determined said duration of the exhalation, and on the determined average PCO2 value; andrelaying the estimated PCO2 waveform to a clinician by displaying on a display. 2. The method according to claim 1, further comprising calculating an estimated EtCO2 value. 3. The method according to claim 1, further comprising calculating an estimated minute CO2 value. 4. The method according to claim 1, wherein the receiving of the CO2 signal comprises wirelessly receiving the CO2 signal. 5. The method according to claim 1, wherein the CO2 signal comprises discrete PCO2 readings taken every 500 ms or more. 6. The method according to claim 1, wherein the CO2 signal comprises discrete PCO2 readings taken every 900 ms or more. 7. The method according to claim 1, wherein the receiving of the breath flow dynamics signal comprises wirelessly receiving the breath flow dynamics signal. 8. The method according to claim 1, wherein the breath flow dynamics signal comprises a flow measurement. 9. The method according to claim 1, wherein the breath flow dynamics signal comprises an acoustic measurement. 10. The method according to claim 1, wherein the breath flow dynamics signal comprises a thermal measurement. 11. The method according to claim 1, wherein the breath flow dynamics signal comprises a chest movement measurement. 12. The method according to claim 1, wherein the breath flow dynamics signal comprises a computer-aided video analysis. 13. The method according to claim 1, wherein the breath flow dynamics signal comprises a Doppler radar signal. 14. The method according to claim 1, further comprising receiving an SpO2 signal. 15. A ventilation analysis system comprising: a CO2 sensor configured to measure discrete partial pressure (PCO2) in exhaled breath of a subject; said CO2 sensor having a response time of more than 100 ms;a breath flow dynamics (BFD) sensor configured to measure the BFD of the subject;an interface module configured to receive discrete partial pressure PCO2 signals from the CO2 sensor... | ORIDION MEDICAL 1987 LTD. | A61B005/08 A61M016/10 A61B005/083 A61B005/087 A61M016/00
US09649041 | B2 | 2017-05-16 | Blood flow perfusion analyzing apparatus, blood flow perfusion analyzing method, fluid analyzing apparatus and fluid analyzing | 2010-11-03 | According to one embodiment, a blood flow perfusion analyzing apparatus includes tissue and arterial TCC calculation units, first and second making units, a deconvolution unit and a blood flow information calculation unit. The tissue TCC calculation unit obtains tissue TCCs. The arterial TCC calculation unit calculates an arterial TCC. The first making unit makes first sets, starting from a first time in which elements corresponding to the tissue are arrayed one-dimensionally, based on the tissue TCCs. The second making unit makes a second set, starting from a second time after the first time, in which elements corresponding to the artery are arrayed two-dimensionally, based on the arterial TCC. The deconvolution unit calculates transfer functions of the tissue based on the first and second sets. The blood flow information calculation unit calculates information on a blood flow perfusion based on the transfer functions. | 1. A blood flow perfusion analyzing apparatus comprising: at least one computer having a central processing unit (CPU) coupled to memory, display and input/output ports, said computer being configured by stored executable program instructions to measure a variation in pixel value between time phases at each pixel based on contrast image data of an object to obtain time concentration curves each showing a time variation in amount corresponding to a concentration of a contrast medium in a tissue, the contrast image data corresponding to the time phases and being acquired by a medical imaging apparatus;a define an arterial region based on the time concentration curves corresponding to the tissue and to calculate a time concentration curve corresponding to an artery;make the time concentration curves corresponding to the tissue discrete by setting a first time as a first starting point time of sampling points for a first discretization of time concentration curves corresponding to the tissue to make first data sets, each of the first data sets having one-dimensionally arrayed elements including values corresponding to concentrations of the contrast medium at different times in the tissue;make the time concentration curve corresponding to the artery discrete by setting a second time as a second starting point time of sampling points for a second discretization of a time concentration curve corresponding to the artery to make a second data set in which elements including values corresponding to concentrations of the contrast medium at different times in the artery are arrayed two-dimensionally, the second starting point time being after the first starting point time;calculate transfer functions of the tissue by deconvolution with regard to each of the first data sets and the second data set;calculate information on blood flow perfusion based on the transfer functions wherein the information on blood flow perfusion includes at least one of blood flow, mean transit time, or blood volume; anddisplay the information on blood flow perfusion as a diagnostic aid. 2. The blood flow perfusion analyzing apparatus of claim 1, wherein the first time is a time for starting an acquisition of the contrast image data or a predetermined time between the time for starting the acquisition and the second time while the second time is a time at which the time concentration curve corresponding to the artery begins to rise. 3. The blood flow perfusion analyzing apparatus of claim 1, wherein each of the first data set is a set made by adding a number H of first elements to be disposed ahead of a set of one-dimensionally arrayed values in a time direction while the second data set is a set made by adding a second H number of elements to be disposed after a set of two-dimensionally arrayed values in the time direction, the one-dimensionally arrayed values corresponding to a number L concentrations of the contrast medium at the different times in the tissue, the two-dimensiona... | TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/02 A61B005/026 A61B005/027 A61B006/00 A61B005/00 A61B006/03 G01R033/563
US10688278 | B2 | 2020-06-23 | Catheter with pressure measuring tip | 2009-11-30 | A medical probe, consisting of a flexible insertion tube, having a distal end for insertion into a body cavity of a patient, and a distal tip, which is disposed at the distal end of the flexible insertion tube is configured to be brought into contact with tissue in the body cavity. The probe also includes a coupling member, which couples the distal tip to the distal end of the insertion tube and which consists of a tubular piece of an elastic material having a plurality of intertwined helical cuts therethrough along a portion of a length of the piece. | 1. A medical probe, comprising: a flexible insertion tube, having a distal end for insertion into a body cavity of a patient;a distal tip having an electrode fixedly mounted thereon, which is disposed at the distal end of the flexible insertion tube and is configured to be brought into contact with tissue in the body cavity;a flexible coupling member separate from the electrode, which fixedly couples the distal tip to the distal end of the flexible insertion tube, the flexible coupling member having a two-piece construction, comprising a flexible tubular piece and a separate tubular part, the flexible tubular piece including an elastic member having a plurality of intertwined, single helices formed by a plurality of multiple start, longitudinally adjacent helical cuts therethrough along a portion of a length of the flexible tubular piece, wherein the intertwined, single helices subtend an angle between 360 degrees and 720 degrees about a central axis of the flexible tubular piece, the separate tubular part is configured to fit over at least a portion of the flexible tubular piece and to be fixedly connected to the flexible tubular piece and to the distal end of the flexible insertion tube, the separate tubular part and the flexible tubular piece having a common outer circumference;a magnetic field generator within the tubular part configured to generate a magnetic field; anda position sensor within the tubular piece, the position sensor configured to sense a position of the distal tip relative to the distal end of the insertion tube, the position changing in response to deformation of the coupling member and wherein the position sensor is configured to generate a signal in response to the magnetic field, wherein the signal is indicative of the relative position of the distal tip to the distal end. 2. The probe according to claim 1, wherein the plurality of intertwined helical cuts comprises n cuts, where n is an integer greater than 1, and wherein the cuts are configured so that the tubular piece has n-fold rotational symmetry about an axis of the flexible tubular piece. 3. The probe according to claim 1, wherein at least one of the helical cuts subtends an angle of 450° about the central axis of the flexible tubular piece. 4. The probe according to claim 1, wherein the flexible coupling member is configured to bend in response to pressure exerted on the distal tip when the distal tip engages the tissue, and wherein at least one of the helical cuts has a width chosen so as to provide a predetermined angular limit on initial bending of the flexible coupling member. 5. The probe according to claim 1, wherein the flexible tubular piece comprises a stem, wherein the stem has the same wall thickness as the flexible tubular piece and wherein the fixed connection comprises one or more welds between the stem and the separate tubular part. 6. The probe according to claim 1, wherein the flexible insertion tube, the distal tip and the flexible coupling... | Biosense Webster (Israel), Ltd. | A61M025/00 A61B018/14 A61B018/00 A61B090/00 A61B034/20
US11490826 | B2 | 2022-11-08 | Apparatus, systems and methods for measuring flow and pressure within a vessel | 2010-07-14 | Exemplary embodiments of apparatus, system and method can be provided to measure a flow of fluid within an anatomical structure. For example, it is possible to use at least one first probe arrangement structured to be insertable into a vessel and configured to direct at least one radiation to at least one portion of the anatomical structure. Further, it is possible to provide at least one second arrangement which configured to detect an interference between a first radiation provided from the fluid via the probe arrangement and second a second radiation provided from a reference path as a function of wavelength thereof. Further, at least one third arrangement can be provided which is configured to determine at least one characteristic of the fluid as a function of the interference. | 1. An apparatus to determine at least one characteristic of a flowing fluid within an anatomical structure, comprising: at least one probe coupled to a radiation source, the at least one probe comprising distal optics and being structured to be insertable into a vessel,the at least one probe configured to: direct, using the distal optics, at least one radiation from the radiation source to at least one portion of the anatomical structure at a longitudinal position along the anatomical structure, andreceive, using the distal optics, a first radiation from the flowing fluid within the anatomical structure at the longitudinal position based on directing the at least one radiation from the radiation source to the at least one portion of the anatomical structure, the at least one probe and the distal optics being configured to remain stationary during operation of the apparatus such that the distal optics continuously direct the at least one radiation toward the at least one portion of the anatomical structure and receive returned radiation based on the directed at least one radiation;a reference path coupled to the radiation source and to a reference structure, the at least one radiation from the radiation source being directed to the reference structure via the reference path,a second radiation being received from the reference structure via the reference path based on the at least one radiation from the radiation source being directed to the reference structure;at least one detector to detect an interference signal between the first radiation provided from the flowing fluid via the at least one probe and the second radiation provided from the reference structure via the reference path as a function of at least one wavelength of at least one of the first radiation or the second radiation, the interference signal identifying a scattering substance at a plurality of depths within the flowing fluid at the longitudinal position; anda processor configured to determine, based on data obtained while the at least one probe and the distal optics remain stationary, the at least one characteristic of the flowing fluid using speckle decorrelation analysis of the scattering substance at the plurality of depths within the flowing fluid based on an intensity change in the interference signal detected from the flowing fluid at at least two different points in time. 2. The apparatus according to claim 1, wherein the at least one characteristic comprises particular parameters which include at least one of: viscosity, coronary flow reserve, fractional flow reserve, coronary flow velocity reserve, average peak velocity, maximum peak velocity, or peak velocity of the flowing fluid within the vessel. 3. The apparatus according to claim 2, wherein the at least one characteristic comprises a multi-dimensional distribution of the particular parameters. 4. The apparatus according to claim 1, wherein the processor is configured to determine the at least one characteristic ... | The General Hospital Corporation | A61B005/026 A61B005/02 A61B005/00 A61B005/021 A61B005/024 A61B005/0215
US20110130685 | A1 | 2011-06-02 | METHOD AND DEVICE FOR REAL TIME MECHANICAL IMAGING OF PROSTATE | 2011-01-12 | The present invention relates to a method for real time mechanical imaging of a prostate with a transrectal probe. In the method, generating a composite two- and three-dimensional prostate mechanical image from a plurality of partial mechanical images extracted from pressure response data and a probe orientation data starts with examining the prostate by pressing a probe head pressure sensor array against it at various overlapping locations. Merging of partial mechanical images together is accomplished by analyzing an overlap between each subsequent and previous partial mechanical image. Finding the prostate is assisted with a supplemental pressure response data indicating the location of a sphincter known to be about 4-5 cm away from the prostate. Data processing is improved by including probe orientation data to further increase the accuracy and sensitivity of the method. The probe is equipped with a two-dimensional head pressure sensor array, a supplemental shaft sensor array and orientation tracking sensors including a three-axis magnetic sensor and a two-axis accelerometer sensor for calculating elevation, rotation and azimuth angles of the probe. | 1. A method for mechanical imaging of a palpable organ through a natural body opening comprising the steps of: (a) providing a probe equipped with a two-dimensional pressure sensor array adapted to obtain pressure response data when pressed against said palpable organ,(b) conducting examination of said palpable organ by inserting said probe through said natural body opening and pressing said probe against said palpable organ at various locations about thereof to obtain said pressure response data, each subsequent location overlapping a previous location,(c) obtaining a plurality of partial mechanical images from said pressure response data, each subsequent and previous partial mechanical image corresponding to the respective location of pressing of said probe against said palpable organ,(d) comparing each subsequent partial mechanical image with the previous partial mechanical image to find an overlap therebetween, and(e) constructing a composite mechanical image of said palpable organ from said plurality of partial mechanical images using overlaps between each subsequent and previous partial mechanical image to merge them together,whereby the entire mechanical image of said palpable organ is obtained with said two-dimensional pressure sensor array irrespective of the movements of said organ during examination. 2. The method as in claim 1, wherein said palpable organ is a human prostate gland and said natural body opening is a rectum. 3. The method as in claim 1, wherein said probe is further equipped with a supplemental pressure sensor array located in a known geometrical relationship to said pressure sensor array along said probe, said step (b) further including obtaining supplemental pressure response data from said second pressure sensor array. 4. The method as in claim 3, wherein said step (b) further including identifying a preliminary position of said palpable organ at a predetermined distance from said supplemental reference organ as detected from said supplemental pressure response data. 5. The method as in claim 1, wherein said probe is further equipped with orientation tracking means, said step (c) further including obtaining orientation data, and said step (e) further including constructing said composite image of said palpable organ using said orientation data and said overlaps to merge said partial images together. 6. The method as in claim 1, wherein said step (e) includes constructing at least one two-dimensional composite image of said palpable organ. 7. The method as in claim 1, wherein said step (e) includes constructing a three-dimensional composite image of said palpable organ. 8. The method as in claim 1 further including a step (f) of displaying said composite image. 9. The method as in claim 1 further including a step (f) of calculating geometrical parameters of said palpable organ and mechanical parameters of its inner structure. 10. The method as in claim 9 further including a step (g) of displaying said composi... | ARTANN LABORATORIES, INC. | A61B005/103
US20110137201 | A1 | 2011-06-09 | ORAL THERMOMETER WITH CURVED PROBE | 2010-07-23 | An oral thermometer including a probe for insertion into the mouth of a user. The probe includes and an elongate first stem, a joint stem extending the first stem and disposed at a first angle to the first stem, and an elongate second stem extending from the joint stem and disposed at a second angle to the joint stem. The first angle open down and the second angle open up. The second stem includes a probe tip at a distal end thereof that is configured to be inserted into the mouth of a user. In operation the thermometer is operable to heat the probe tip to a pre-warmed temperature, detect an elevated temperature higher than the pre-warmed temperature when inserted in a user's mouth, and initiate a temperature reading upon detecting the elevated temperature. | 1. An oral thermometer comprising: a probe configured to be at least partially inserted into the mouth of a user, the probe including: an elongated first stem having a distal end and a proximal end;a joint stem having a distal end and a proximal end, the joint stem extending from the distal end of the first stem and disposed at a first angle to the first stem about a first bend; andan elongated second stem having a distal end and a proximal end, the second stem extending from the distal end of the joint stem and disposed at a second angle to the joint stem about a second bend,wherein the first angle opens generally in a downward direction and the second angle opens generally in an upward direction. 2. The oral thermometer of claim 1 further comprising a temperature sensor disposed at the probe tip. 3. The oral thermometer of claim 1 further comprising a heater operable to heat the probe tip. 4. The oral thermometer of claim 1 wherein each of the first and second angle is between 45 and 135 degrees. 5. The oral thermometer of claim 2 wherein the probe tip includes a metal cup and wherein the temperature sensor is disposed inside the metal cup. 6. The oral thermometer of claim 1 wherein the second stem has a length of at least 10 mm. 7. The oral themiometer of claim 6 wherein the second stem has a length of about 25 mm. 8. The oral thermometer of claim 1 wherein the joint stem has a length of at least 8 mm. 9. The oral the mometer of claim 8 wherein the joint stem has a length of about 15 mm. 10. The oral thermometer of claim 1 further comprising a housing, wherein the probe is pivotally attached to the housing. 11. The oral theuiiometer of claim 10 further comprising a switch operable to activate the thermometer in response to pivoting between the probe and the housing. 12. The oral thermometer of claim 3 further comprising an electronic control unit operable to heat the probe tip to a pre-warmed temperature tH that is lower than a lowest patient temperature tp−min; detect an elevated temperature higher than the pre-warmed temperature by a predetermined threshold using a temperature sensor disposed in proximity to the probe tip; andinitiate a temperature reading upon detecting the elevated temperature. 13. The oral thermometer of claim 12 wherein the electronic control unit is further operable to maintain the probe tip at the pre-warmed temperature for at least a set amount of time until the elevated temperature is detected. 14. The oral thermometer of claim 12 wherein the electronic control unit is further operable to turn off the heater in response to the probe tip reaching the elevated temperature. 15. The oral thermometer of claim 1 further comprising a low thermal conductivity insert disposed between the second bend and the probe tip. 16. The oral thermometer of claim 1 wherein the probe is generally z-shaped. 17. A method of measuring a temperature of a user with an oral thermometer comprising: providing a thermometer with a ... | - | A61B005/01
US20110130796 | A1 | 2011-06-02 | Thermostimulation methods using multilayer pads with integrated temperature regulation | 2009-11-27 | A therapeutic method. The inventive method includes the steps of: applying a pad to a patient having a biometric sensor adapted to feedback a signal; coupling said feedback signal to an inline control system; coupling said inline control system to a console; and regulating energy applied to said pad by the inline control system based on the output of said sensor coupled to said inline control system and an output of said console. In the illustrative embodiment, the therapeutic method is thermostimulation and includes the steps of applying heat and stimulation via the pad, sensing temperature at the pad with an embedded temperature sensor and regulating the heat current at the pad via the inline control system in response to the output of the temperature sensor. In an alternative embodiment, the embedded sensor is a galvanic skin response sensor for measuring skin conductivity and having output is used by the inline controller in the treatment of dermatological conditions, relaxation treatment, desensitization training or other purposes. In another embodiment, the embedded sensor is electromyography sensor for measuring muscular electrical activity for the treatment of patients for weakness, impaired muscle strength, or gait analysis. In yet another embodiment, the embedded sensor is a pulse sensor. In this case, the inline controller uses the data to measure the patient's heart rate. In any case, useful data with respect to the performance of and conditions at each pad is displayed either on the inline controller and/or at the console. In addition, the inline controllers may be adapted to communicate with each other and/or with other external devices. | 1. A therapeutic method including the steps of: applying a pad to a patient having a biometric sensor adapted to feedback a signal;coupling said feedback signal to an inline control system;coupling said inline control system to a console; andregulating energy applied to said pad by the inline control system based on the output of said sensor coupled to said inline control system and an output of said console. 2. The invention of claim 1 wherein said therapeutic method includes the steps of applying heat via said pad. 3. The invention of claim 2 wherein said sensor is a temperature sensor. 4. The invention of claim 3 wherein said therapeutic method includes the step of applying electrical stimulation via said pad. 5. The invention of claim 1 wherein said therapeutic method includes the step of applying electrical stimulation via said pad. 6. The invention of claim 1 further including the step of measuring skin conductivity 7. The invention of claim 6 wherein said step of measuring skin conductivity includes the step of measuring skin conductivity with a galvanic skin response sensor embedded within said pad. 8. The invention of claim 7 further including the step of inputting the galvanic skin response sensor output to the inline control system. 9. The invention of claim 8 further including the step using galvanic skin response sensor output in the treatment of dermatological conditions. 10. The invention of claim 9 further including the step using galvanic skin response sensor output in a relaxation treatment. 11. The invention of claim 10 further including the step using galvanic skin response sensor output for desensitization training. 12. The invention of claim 1 further including the step of measuring muscular electrical activity. 13. The invention of claim 12 wherein the step of measuring muscular activity includes the step of measuring muscular activity with an electromyography sensor embedded within said pad. 14. The invention of claim 13 further including the step of coupling the output of the electromyography sensor to the inline control system. 15. The invention of claim 14 wherein said inline control system adjusts the heat and/or stimulation levels based on the output from the electromyography sensor. 16. The invention of claim 15 wherein the inline control unit provides data to the console based on the output from the electromyography sensor. 17. The invention of claim 16 wherein the output from the electromyography sensor is displayed on a display of the inline control system. 18. The invention of claim 17 wherein the output from the electromyography sensor is displayed on the console. 19. The invention of claim 18 further including the step of using the wherein the output from the electromyography sensor in the treatment of patients for weakness, impaired muscle strength, or gait analysis. 20. The invention of claim 1 further including the step of measuring a patient's heart rate. 21. The invention of claim 20 wh... | - | A61F007/08 A61N001/00 A61B005/053
US20110144476 | A1 | 2011-06-16 | Integrated Surgical Sampling Probe | 2009-08-17 | The present invention is directed to devices for collecting tissue samples from patients. The devices create tissue fragments which are then collected by aspiration. The devices include signal emitters that allow the exact position where a biopsy sample was obtained to be determined. In addition the invention includes methods for obtaining tissue samples using these devices. | 1. A medical device comprising: a) a hand held support;b) an elongated metal rod extending from and attached to said hand-held support, wherein said rod terminates at its distal end in a tip that either itself vibrates in response to ultrasonic energy or comprises a separate component which vibrates in response to ultrasonic energy;c) means for supplying ultrasonic energy to said tip or said separate component;d) means for supplying irrigating fluid to said tip;e) means for aspirating tissue fragments created at said tip;f) means for stereotactically determining position of said tip.g) means for collecting tissue samples. 2. The medical device of claim 1, wherein said means for stereotactically determining the position of the tip, comprises one or more electromagnetic emitters located on said device that transmit signals to a separate electromagnetic receiver or infrared energy detectors for optical tracking. 3. The medical device of claim 1 wherein said rod has a hollow core that provides a fluid passageway running along the length of said rod to said tip. 4. The medical device of claim 1, wherein said aspirating means creates suction in said hollow core of said rod to aspirate material through said tip in the direction of said hand held support. 5. The medical device of claim 4, wherein said fluid passageway terminates at its proximal end in a tissue collection container. 6. The medical device of claim 5, wherein said tissue collection container is chilled or comprises a fluid for fixing tissue samples. 7. The medical device of claim 1, wherein said means for supplying ultrasonic energy operates at a frequency of 15-100 kHz. 8. The medical device of claim 1, wherein said means for supplying irrigating fluid comprises a tubular channel running through said rod, wherein said channel is connected at its proximal end to a reservoir of irrigating fluid and wherein said channel terminates at the tip of said rod in an opening through which said irrigating fluid can flow. 9. The medical device of claim 1, wherein said hand held support comprises an actuator switch which, when activated, permits the transmission of ultrasonic energy to the tip of said rod or to said separate component which vibrates in response to ultrasonic energy and which does not permit said transmission when not activated. 10. The medical device of claim 9, wherein, said means for stereotactically determining the position of the distal end of said rod, comprises electromagnetic emitters located on said device that transmit signals to a separate electromagnetic receiver, wherein said electromagnetic receiver communicates signal data to a computer programmed to analyze position or infrared energy detectors for optical tracking. 11. A method of collecting a tissue sample from a biopsy site of a patient comprising: a) inserting the tip of the medical device of claim 1 into said patient so that the distal end of said rod is positioned at said biopsy site;b) transmitting ultras... | THE BRIGHAM AND WOMEN'S HOSPITAL, INC. | A61B010/02 A61B005/055 A61B006/03
US07945316 | B2 | 2011-05-17 | System and method for monitoring or treating nervous system disorders | 2005-12-19 | A medical device system for comparing a cardiopulmonary signal to a brain signal. In one embodiment of the invention, a medical device system is provided that includes a brain monitoring element, respiratory monitoring element and a processor. The processor is configured to receive a brain signal from the brain monitoring element and a respiratory signal from the respiratory monitoring element. The processor is further configured to compare the brain signal to the respiratory signal. Methods of comparing a brain signal to a cardiopulmonary signal are also provided. | 1. A medical device system comprising: (a) a brain monitoring element for sensing activity of the brain and outputting a brain signal;(b) a respiratory monitoring element for sensing activity of the respiratory system and outputting a respiratory signal; and(c) a processor in communication with the brain monitoring element and the respiratory monitoring element, the processor configured to: (i) receive the brain signal and obtain information identifying one or more neurological events in the brain signal;(ii) receive the respiratory signal and obtain information identifying one or more respiratory events in the respiratory signal; and(iii) compare the brain signal to the respiratory signal,wherein the processor being configured to compare the brain signal to the respiratory signal comprises the processor being configured to perform the following: (a) for each of the one or more neurological events, determining whether said neurological event is within a specified time period of one of the one or more respiratory events; and(b) for each of the one or more respiratory events, determining whether said respiratory event is within a specified time period of one of the one or more neurological events. 2. The medical device system of claim 1 wherein the medical device system includes a user interface for display of the brain signal and the respiratory signal, the user interface in communication with the processor, and wherein the processor being configured to obtain information identifying one or more neurological events in the brain signal comprises the processor being configured to receive a start point and an end point of each neurological event from a user via the user interface, and wherein the processor being configured to obtain information identifying one or more respiratory events in the respiratory signal comprises the processor being configured to receive a start point and an end point of each respiratory event from a user via the user interface. 3. The medical device system of claim 1 wherein the processor being configured to obtain information identifying one or more neurological events in the brain signal comprises the processor being configured to detect one or more neurological events from the brain signal, and wherein the processor being configured to obtain information identifying one or more respiratory events in the respiratory signal comprises the processor being configured to detect one or more respiratory events from the respiratory signal. 4. The medical device system of claim 1 wherein the processor being configured to compare the brain signal to the respiratory signal further comprises the processor being configured to perform the following: (a) determining the number of neurological events within the specified time period of a respiratory event;(b) determining the number of neurological events that are not within the specified time period of a respiratory event; and(c) determining the number of respiratory events that are n... | Medtronic, Inc. | A61B005/00
US07963927 | B2 | 2011-06-21 | Electromyography system | 2004-04-21 | An electromyography system and related method for performing automated screw test procedures, involving automatically determining an onset neuro-muscular response to the application of an electrical stimulus to a portion of pediculur bone and communicating to a user an onset electrical stimulus level which causes the onset neuro-muscular response. | 1. A method of determining structural integrity of a bone within the spine of a patient using a neurophysiology system, the bone having a first aspect and a second aspect, said second aspect separated from said first aspect by a width and located adjacent to a spinal nerve, said method comprising: applying an electrical stimulus to said first aspect of said bone;electrically monitoring a plurality of leg muscle myotome locations via a plurality of electromyographic (EMG) sensor electrodes coupled to the leg muscle myotome locations, at least one of the leg muscle myotome locations being associated with said spinal nerve;detecting an onset neuro-muscular response having an amplitude greater than a predetermined positive value in response to the application of said electrical stimulus to said first aspect of said bone by automatically increasing said electrical stimulus in constant increments until said onset neuro-muscular response is detected by one or more of the EMG sensor electrodes, wherein said automatic increasing is controlled by said neurophysiology system; anddisplaying on an integrated display device of said neurophysiology system while viewed by a surgeon operating on the patient's spine an onset electrical stimulus current level which causes said onset neuro-muscular response. 2. The method of claim 1, wherein the electrical stimulus is emitted from an electrode disposed on the distal end of an instrument inserted through a cannula extending toward said bone. 3. The method of claim 1, wherein automatically increasing said electrical stimulus in constant increments comprises applying a plurality of electrical stimulus pulses. 4. The method of claim 3, wherein the plurality of electrical stimulus pulses comprises current pulses that automatically increase in constant increments over time until said onset neuromuscular response is determined. 5. The method of claim 1, wherein the electrical stimulus is increased by constant current increments from within the range of 0.5 to 4 milliamps. 6. The method of claim 1, wherein the electrical stimulus is increased within a current range from 0.5 to 32.0 milliamps. 7. The method of claim 1, wherein said spinal nerve exits from successive vertebrae within the patient's spine. 8. The method of claim 1, wherein said onset neuro-muscular response is an electromyography response from a muscle coupled to said spinal nerve. 9. The method of claim 1, wherein electrically monitoring said plurality of leg muscle myotome locations is performed through the use of needle electrodes electrically coupled to leg muscle myotomes. 10. The method of claim 1 and further, comprising audibly indicating to said surgeon an intensity level representing said onset electrical stimulus current level causing said onset neuro-muscular response for said spinal nerve. 11. The method of claim 10, wherein audibly indicating comprises sounding an alarm if said onset neuro-muscular response is detected at a predetermined ... | NuVasive, Inc. | A61B005/04 A61B005/05
US07949166 | B2 | 2011-05-24 | Diagnosis support system | 2007-04-12 | A region of interest is set by region of interest setting means in each of medical images obtained in chronological order. Similar case search means searches for a case information set including a case image having a region of similar pictorial characteristics to each of the regions of interest. When the corresponding case information set has been extracted through the search, case information display means displays test result information sets including not only the case images of similar pictorial characteristics but also all case images, test purposes, findings, and the like in the case information sets in chronological order of tests. | 1. A diagnosis support system that supports diagnosis based on interpretation of a given medical image by searching case information sets registered with a database for case information sets including case images of similar pictorial characteristics to the given medical image and by displaying the case information sets including the case images, the system comprising: region of interest setting means for setting a region of interest in each of medical images obtained in chronological order, by displaying the medical images on a monitor screen and by receiving a region specifying operation on at least one of the medical images;similar case search means for extracting through search of the database a case information set including case images each having a region representing pictorial characteristics that are similar to pictorial characteristics of a corresponding one of the regions of interest having been set by the region of interest setting means;case information display means for displaying all the case images belonging to the case information set extracted by the similar case search means on the monitor screen in chronological order; andfurther comprising corresponding region estimating means for estimating in a second one of the medical images a region representing a target that is the same as a target represented by a region specified in a first one of the medical images according to coordinate information identifying the region specified in the first medical image, whereinthe region of interest setting means receives specification of the region of interest by the region specifying operation only in one of the medical images and sets the region of interest in each of the medical images by causing the corresponding region estimating means to estimate the region of interest in each of the remaining images by supplying the coordinate information identifying the specified region thereto. 2. The diagnosis support system according to claim 1, wherein the similar case search means carries out processing for searching for case images each having a region representing pictorial characteristics that are similar to the pictorial characteristics of a corresponding one of the regions of interest having been set by the region of interest setting means;processing for identifying a group of case images of one and the same patient each having the region representing the pictorial characteristics that are similar to the pictorial characteristics of the corresponding region of interest having been set by the region of interest setting means among the case images having been detected through the search for the case images; andprocessing for supplying a case information set including the case images belonging to the identified group to the case information display means as the case information set having been extracted through the search of the database. 3. The diagnosis support system according to claim 1, wherein the similar case search means carries out pr... | Fujifilm Corporation | G06K009/00 A61B005/00 A61B005/05
US08690820 | B2 | 2014-04-08 | Automatic insulin pumps using recursive multivariable models and adaptive control algorithms | 2010-10-05 | A method and device for monitoring or treating patient glucose levels. The device includes a glucose sensor for measuring a glucose level of a patient, a physiological status monitoring system for measuring at least one physical or metabolic variable of the patient, and an automatic controller in communication with the glucose sensor and the physiological status monitoring system. The controller includes a prediction module for automatically predicting a future glucose level using data measured by the glucose sensor and the physiological sensor. | 1. A method of predicting glucose levels in a patient, the method comprising: automatically measuring a glucose level in the patient to obtain an actual glucose measurement;automatically monitoring a physiological variable of the patient to obtain a measured physiological variable;automatically modeling a glucose concentration of the patient as a function of the actual glucose measurement and the measured physiological variable;automatically estimating a future glucose level for the patient from the model of the glucose concentration of the patient; andautomatically updating the model of the glucose concentration as a function of further glucose measurements and further physiological variables. 2. The method according to claim 1, further comprising automatically recursively updating the model of the glucose concentration with each of the further glucose measurements. 3. The method of claim 1, wherein the updating occurs upon determining a difference between the estimated future glucose level for a timeframe and a further actual glucose level at the timeframe. 4. The method of claim 1, wherein the updating of the model comprises modifying model parameters upon comparing the estimated future glucose for a timeframe to the further physiological variables measured during the timeframe. 5. The method according to claim 1, further comprising periodically measuring a glucose level in the patient with a continuous glucose monitor worn by the patient. 6. The method according to claim 1, further comprising monitoring for physiological variables of the patient using at least one sensor worn by the patient that measures movement, skin temperature, dissipated heat from the body, galvanic skin response, or combinations thereof, and further comprising deriving at least one physiological variable selected from sleep, total energy expenditure, stress, physical activity, or combinations thereof. 7. The method according to claim 1, further comprising activating an alarm when an estimated future glucose level is below a predetermined glucose level. 8. The method according to claim 1, further comprising: determining changes in glucose concentration model parameters; andmodifying the glucose concentration model in view of the new parameters. 9. The method according to claim 1, wherein the method is a closed-loop insulin therapy without manual inputs by the patient. 10. The method according to claim 1, further comprising determining a difference between an estimated future glucose level for a timeframe and an actual glucose measurement for the timeframe, and modifying the model of the glucose concentration as a function of the difference. 11. The method according to claim 1, further comprising measuring the glucose level with a subcutaneous continuous glucose monitor and compensating for time delays between blood glucose levels and subcutaneous glucose levels. 12. The method according to claim 1, further comprising storing the actual glucose measurements and... | Illinois Institute of Technology | A61B005/00
US08600469 | B2 | 2013-12-03 | Medical sensor and technique for using the same | 2011-02-07 | A medical sensor may be adapted to account for factors that cause irregularities in pulse oximetry measurements or other spectrophotemetric measurements. Sensors are provided with surface features that reduce the amount of outside light or shunted light that impinge the detecting elements of the sensor. The sensor is adapted to reduce the effect of outside light or shunted light on pulse oximetry measurements. | 1. A sensor comprising: a conformable sensor body;an emitter disposed on a tissue-contacting surface of the sensor body;a detector disposed on the tissue-contacting surface of the sensor body; anda patterned region disposed on the tissue-contacting surface of the sensor body, wherein the patterned region is configured to reduce an incidence of shunted or ambient light reaching the detector and wherein the patterned region protrudes less than about 0.001 mm from the tissue-contacting surface of the sensor body. 2. The sensor, as set forth in claim 1, wherein the sensor is in a reflectance configuration. 3. The sensor, as set forth in claim 1, wherein the sensor is in a transmission configuration. 4. The sensor, as set forth in claim 1, wherein the sensor body is adapted for placement on a patient's digit. 5. The sensor, as set forth in claim 1, wherein the sensor body is adapted for placement on a patient's forehead. 6. The sensor, as set forth in claim 1, wherein the sensor comprises a pulse oximetry sensor. 7. The sensor, as set forth in claim 1, wherein the patterned region comprises a single material applied in varying concentration to the tissue-contacting surface of the sensor. 8. A sensor comprising: a sensor body;an emitter disposed on the sensor body;a detector disposed on the sensor body; anda patterned region disposed on a tissue-contacting surface of the sensor body, wherein the patterned region comprises portions of a light absorbing material that are configured to reduce an incidence of shunted or ambient light reaching the detector and wherein the portions of the light absorbing material defining the patterned region are microscopic in scale. 9. The sensor, as set forth in claim 8, wherein the sensor is in a reflectance configuration. 10. The sensor, as set forth in claim 8, wherein the sensor is in a transmission configuration. 11. The sensor, as set forth in claim 8, wherein the sensor body is adapted for placement on a patient's digit. 12. The sensor, as set forth in claim 8, wherein the sensor body is adapted for placement on a patient's forehead. 13. The sensor, as set forth in claim 8, wherein the sensor comprises a pulse oximetry sensor. 14. The sensor, as set forth in claim 8, wherein the sensor body is conformable. 15. A sensor comprising: a sensor body;an emitter disposed on the sensor body;a detector disposed on the sensor body; andat least one protrusion integrally formed with the sensor body and protruding non-orthogonally from a tissue-contacting surface of the sensor body, wherein the at least one protrusion comprises an irregular profile on a surface and wherein the at least one protrusion is configured to reduce an incidence of shunted or ambient light reaching the detector. 16. The sensor, as set forth in claim 15, wherein the at least one protrusion substantially surrounds the detector. 17. The sensor, as set forth in claim 15, wherein the at least one protrusion protrudes at least about 0.001 mm f... | Covidien LP | A61B005/1455
US20110112411 | A1 | 2011-05-12 | BLOOD PRESSURE MONITOR | 2007-05-30 | The present invention relates to an apparatus and method for facilitating blood pressure measurement. An additional fluid reservoir is arranged to be connected in fluid communication with the blood pressure cuff. The fluid reservoir provides an additional, compliant volume to the fluid volume of the cuff, facilitating blood pressure measurement, particularly for small volume cuffs. | 1. An apparatus for facilitating blood pressure measurement, including a fluid reservoir arranged to be connected in fluid communication with a blood pressure cuff, the fluid reservoir being arranged to provide an additional fluid volume to the fluid volume of the cuff during measurement of blood pressure of a patient. 2. An apparatus in accordance with claim 1, wherein the fluid reservoir is arranged to change in volume when a change in pressure is applied. 3. An apparatus in accordance with claim 2, a fluid reservoir being arranged to add little or no pressure back into the system when it changes in volume. 4. An apparatus in accordance with claim 1, 2 or 3, wherein the fluid reservoir has a compliance. 5. An apparatus in accordance with claim 4, wherein the compliance is arranged such that oscillations of pressure from the blood pressure cuff are not significantly damped by the fluid reservoir. 6. An apparatus in accordance with claim 4 or claim 5, wherein the compliance is arranged such that the pressure on the fluid reservoir during any blood pressure measurement does not exceed the fluid reservoir's Critical Opening Pressure. 7. An apparatus in accordance with claim 4, 5 or 6, wherein the fluid reservoir includes a fluid container having at least one expandable wall providing compliance. 8. An apparatus in accordance with claim 7, wherein the fluid container is of flexible plastic material. 9. A system for measuring blood pressure, including an apparatus in accordance with any one of the preceding claims and a blood pressure cuff, the fluid reservoir being arranged to be connected in fluid communication with the blood pressure cuff. 10. An automated system for measuring blood pressure, including a controller for controlling the provision of fluid under pressure to a blood pressure cuff, a display for providing a reading of blood pressure value, and an apparatus in accordance with any one of claims 1 to 6, the fluid reservoir being arranged to be connected in fluid communication with the blood pressure cuff. 11. A method of measuring blood pressure using a blood pressure cuff, including the step of providing an additional fluid volume in fluid connection with the blood pressure cuff. 12. A method in accordance with claim 11, wherein the additional fluid volume is arranged to change the effective volume of the blood pressure cuff when a change of pressure is applied. 13. A method in accordance with claim 12, wherein when the additional fluid volume changes the effective volume of the cuff, it is arranged to add little or no pressure back into the system. 14. A method in accordance with claim 11, 12 or 13, wherein the additional fluid volume has a compliance. 15. A method in accordance with claim 14, wherein the compliance is sufficient such that oscillations of pressure from the blood pressure cuff are not significantly damped. 16. A method in accordance with claim 14 or claim 15, wherein the compliance is such that the pressur... | - | A61B005/022
US20110105856 | A1 | 2011-05-05 | Diagnostic testing | 2010-10-13 | One or more living entities are evaluated for health conditions by taking samples, especially samples of volatile compounds, derived from, e.g., the breath, urine, feces, saliva, skin exudates, etc., from the living entities, optionally with concentration of the samples; analyzing the samples using a multivariate analytical test; comparing the patterns present in the results of said test or tests, either; a. on an individual entity at a previous time to the patterns present in the results of said test or tests on the individual entity at the present time, or b/ comparing the patterns present in the results of the test on the said living entities to the patterns present in the results of tests on populations of similar entities having a normal health condition and/or to the patterns in the results of tests on populations of similar entities having a given disease or abnormal health condition, the results from the entities having normal health and each of the populations of entities having a given disease or abnormal condition being treated as reference sets for their respective specific disease or abnormal condition as compared to the reference set for entities having normal health. | 1. A method for determining the condition of a living entity wherein at least one material given off from said living entity is analyzed with at least one multivariate analytical test method to provide test results and those results are then compared either to the test results from the same test method or methods on either the same entity at a time when the entity's condition is known, or to the results from the same test method or methods on one more reference individuals or reference populations of similar entities having a known condition. 2. The method of claim 1 which comprises a step in which at least one volatile material is collected at least once, in a standard way, from at least one individual living entity and analyzed using said multivariate analytical test method to provide test results that are used to create at least one reference set representative of the condition of said living entity. 3. The method of claim 1 which comprises at least one of: (1) the step of developing one or more data bases, each of said data bases comprising one or more reference sets, and each of said reference sets comprising the test results from at least one multivariate analytical test method on materials given off by at least one individual from a population of individuals having similar or identical basic characteristics and a shared health condition, the reference set being used to represent all individuals having similar basic characteristics and the same said health condition;(2) the step of using data mining software to establish a pattern in the test results in at least one of the reference sets of (1) where the pattern is representative of a population sharing a known condition;(3) the step in which an individual is tested using the same multivariate analytical test method used to create the reference set of (1) and the individual's test results are compared to at least one of the said reference sets of (1);(4) the step of (3) wherein said individual's test results are compared to the results from at least one reference set derived from at least one healthy population having similar genetic backgrounds, allergies, family histories of susceptibility to diseases, diets, and other relevant characteristics, to determine if the individual shares the same inherited basic characteristics as one of the populations;(5) the step of using one of the methods herein as an initial step in the diagnosis of the condition of an individual;(6) the step of using at least one pattern found by at least one of the methods herein to determine what body processes are related to a disease;(7) the step of using at least one of the methods herein to monitor the progress of an individual's condition the step of using one of the methods herein to determine when changes occur in a microorganism or virus;(8) the step of providing the data found using at least one of the methods herein in digital form for storage, retrieval, and/or analysis;(9) the step in which one or more vo... | - | A61B005/00
US08550997 | B2 | 2013-10-08 | Sensor and monitor system | 2009-12-31 | A monitor system is disclosed. The monitor system includes a sensor with a sensor port. The monitor system further includes a recorder with a recorder port within a recording housing. The recorder port interfaces with the sensor port to receive signals from the sensor port. A recorder clock is defined within the recorder housing, with a recorder processor to store signals from the sensor. The recorder includes a data port to interface with a dock receiver. The monitor system includes a dock remotely located from the sensor and the recorder. The dock receiver couples the recorder to the dock. The monitor system further includes a data processor to analyze the sensor signals from the recorder. The data processor includes memory and a clock. Further included with the data processor are program instructions to assign the time and date of the sensor signals. | 1. A monitor system to monitor a characteristic of a user, the system comprising: a sensor to produce a sensor signal on a continuous basis, the sensor signal being indicative of a glucose characteristic measured in the user when placed in subcutaneous tissue, the sensor having a sensor port;a recorder that includes a recorder housing, a battery being contained within the recorder housing,a recorder port defined to interface with the sensor port to receive the produced signals from the sensor,a recorder clock defined within the recorder housing to assign a time to the signals from the sensor,a recorder processor that includes a recorder memory coupled to the recorder port to store the produced signals from the sensor,a data port defined to interface with a dock receiver,a dock remotely located from the sensor and the recorder, the dock includes the dock receiver to physically couple the recorder to the dock via the data port, a dock processor coupled to the dock receiver, and a data processor defined to analyze the signals from the sensor that were stored in the recorder, the data processor includes: a data processor memory to store data from the recorder and a data processor clock,wherein the data processor is programmed to assign the time and date of the signals from the sensor by comparing the time and date on the data processor clock with the time assigned to the signals from the sensor by the recorder clock,wherein the sensor port and the data port are combined into one multifunction port. 2. A monitor system according to claim 1, wherein the data processor is further programmed to convert the signals from the sensor to glucose concentration values experienced by the user. 3. A monitor system according to claim 1, further comprising: a meter having an internal meter clock to provide reference values to the data processor, the reference values including a timestamp from the internal meter clock, and the data processor includes a program to assign time and date to the reference values by comparing the time and date on the processor clock with the timestamp of the reference values from the meter. 4. A monitor system according to claim 1, wherein the dock further includes a dock port. 5. A monitor system according to claim 4, wherein the data processor further includes a data processor port. 6. A monitor system according to claim 5, further comprising a cable defined to interface between the dock port and the data processor port. 7. A monitor system according to claim 1, such that when the recorder is coupled to the dock, the dock is programmed to: (a) recharge the battery of the recorder, (b) perform a diagnostic test of the recorder, and (c) provide actionable feedback based on results of the diagnostic test. 8. A monitor system according to claim 7, wherein the actionable feedback is provided using at least one light integrated into the dock. 9. A monitor system according to claim 7, wherein the actionable feedback is provided using a... | Medtronic Minimed, Inc. | A61B005/00 A61B005/05
US08460220 | B2 | 2013-06-11 | System and method for monitoring the gait characteristics of a group of individuals | 2009-12-18 | A system is provided for monitoring gait characteristics of a group of enrollees in a monitoring region of a public area. The system includes an identifying device to enroll the group of enrollees into the system by measuring identifying information. The system further includes a gait monitoring device to measure a sample gait characteristic during an enrollment phase, and to measure a gait characteristic of each enrollee subsequent to the enrollment phase. The system further includes an alerting device to receive an alert signal from the gait monitoring device, when the measured gait characteristic of a respective enrollee exceeds a safety threshold. The system further includes a display device to issue a trend report of the measured gait characteristic of a selectable enrollee for a selectable time period. | 1. A system for monitoring gait characteristics of one or more enrollees in a monitoring region of a public area, said system comprising: an identifying device comprising a video camera positioned in the monitoring region of the public area, wherein the identifying device is configured to: automatically prompt a potential enrollee having entered the monitoring region of the public area to provide consent for monitoring the enrollee in the monitoring region, andenroll the enrollee on receiving the enrollee's consent, said enrolling including automatically prompting the enrollee to move to a designated location in the monitoring region of the public area to acquire identifying information of the enrollee during an enrollment phase and automatically prompting the enrollee to move in monitoring region within a field of view of the video camera to measure a sample gait characteristic of the enrollee during the enrollment phase; andacquire identifying information of one or more individuals having entered the monitoring region during a monitoring phase subsequent to the enrollment phase;a processor configured to identify the enrollee from the one or more individuals in the monitoring region by comparing the acquired identifying information of the one or more individuals in the monitoring region with the acquired identifying information of the enrollee during the enrollment phase to determine if one of the individuals is the enrollee; anda gait monitoring device configured to measure the sample gait characteristic of the enrollee during the enrollment phase; said gait monitoring device being further configured to measure a gait characteristic of the enrollee during the monitoring phase, if the processor identifies that one of the individuals is the enrollee;wherein said processor is further configured to compare the measured gait characteristic of the enrollee with at least one of a previously-measured gait characteristic and the sample gait characteristic of the enrollee, to determine if the measured gait characteristic exceeds a safety threshold. 2. The system of claim 1, further comprising: an alerting device configured to receive an alert signal from said gait monitoring device, said alert signal based on the measured gait characteristic of the enrollee having exceeded the safety threshold; anda display device configured to receive said measured gait characteristic data and the enrollee's identifying information of said measured gait characteristic from said gait monitoring device; said display device further configured to issue a trend report of the measured gait characteristic of a selectable enrollee for a selectable time period. 3. The system of claim 2, wherein the identifying device includes a video camera; wherein during said enrollment phase, said enrollee is prompted to position the enrollee's face in a field of view of the video camera to capture an image of the face as said identifying information; wherein said identifying information is... | General Electric Company | A61B005/00
US07959565 | B2 | 2011-06-14 | Method and apparatus for estimating a likelihood of shoulder dystocia | 2004-11-30 | A method and an apparatus for estimating a level of risk of shoulder dystocia with neonatal injury associated to an obstetrics patient are provided. A set of information data elements associated to an obstetrics patient is received including information derived from a maternal weight component, a maternal height component and a fetal weight component. The set of information data elements is processing to derive a ranking data element associated to the obstetrics patient. The ranking data element conveys a level of risk of shoulder dystocia associated to the obstetrics patient. The ranking data element is then released. | 1. An apparatus suitable for estimating a level of risk of shoulder dystocia associated to an obstetrics patient, said apparatus comprising: a) an input for receiving a set of information data elements associated to an obstetrics patient, the set of information data elements including information derived from: (i) a maternal weight component;(ii) a maternal height component; and(iii)a fetal weight component;b) a processing unit in communication with said input, said processing unit being operative for deriving a ranking data element conveying a level of risk of shoulder dystocia associated to the obstetrics patient, said ranking data element depending at least in part on the fetal weight component and on a ratio between the maternal weight component and the maternal height component;c) an output for releasing said ranking data element. 2. An apparatus as defined in claim 1, wherein said ranking data element is indicative of a likelihood score. 3. An apparatus as defined in claim 1, wherein said ranking data element is indicative of a risk level selected from a set of risk levels. 4. An apparatus as defined in claim 1, wherein said processing unit is further operative for generating a false detection data element conveying a false positive detection rate associated to the ranking data element. 5. An apparatus as defined in claim 4, wherein said processing unit is further operative for generating a successful detection data element conveying a successful true positive detection rate associated to the ranking data element. 6. An apparatus as defined in claim 5, wherein said output is adapted for releasing said successful detection data element and said false detection data element. 7. An apparatus as defined in claim 1, wherein said processing unit is further operative for: generating at least one recommendation data element associated to the obstetrics patient. 8. An apparatus as defined in claim 7, wherein said processing unit is operative for processing said ranking data element to generate the at least one recommendation data element associated to the obstetrics patient. 9. An apparatus as defined in claim 7, wherein said processing unit is operative for generating the at least one recommendation data element in part based on data conveying a policy associated to shoulder dystocia. 10. An apparatus as defined in claim 9, said processing unit being operative for processing the ranking data element based on data conveying a policy associated to shoulder dystocia to derive said at least one recommendation data element. 11. An apparatus as defined in claim 10, wherein said processing unit is operative for receiving data conveying the policy associated to shoulder dystocia. 12. An apparatus as defined in claim 11, wherein said data conveying the policy data associated to shoulder dystocia includes a set of risk levels. 13. An apparatus as defined in claim 12, wherein said data conveying the policy data associated to shoulder dystocia incl... | Perigen, Inc. | A61B005/00 G01N033/48
US08188332 | B2 | 2012-05-29 | Disposable cover for use in an incubator for premature infants | 2006-12-22 | The invention refers to a disposable cover for use in an incubator for premature infants comprising a moisture-pervious surface layer and a preferably moisture-impervious back sheet, wherein the total weight of the cover before use is communicated to the caregiver by indication on said cover, a label accompanying said cover or on the package for said cover. The cover may further be provided with a wetness indicator. | 1. A packaged disposable cover for an incubator for premature infants, the packaged disposable cover comprising a package and at least one disposable cover, the package enclosing the disposable cover prior to use, the disposable cover being dimensioned and configured for use in an incubator for premature infants, the disposable cover comprising a moisture-pervious surface layer intended to face the premature infant during use and a back sheet, wherein the total weight of said disposable cover before use is communicated to a caregiver by indication on said cover, wherein said disposable cover has a width of 5-40 cm and a length of 10-60 cm; said disposable cover has a defined weight within a tolerance of ±1 g;the total weight of said disposable cover is printed, marked, or embossed onto said cover; andthe total weight is the actual weight of the individual disposable cover. 2. The disposable cover according to claim 1, wherein the total weight of said disposable cover before use is printed onto said cover. 3. The disposable cover according to claim 1, wherein the total weight of said disposable cover before use is embossed onto said cover. 4. The disposable cover according to claim 1, wherein the disposable cover comprises a wetness indicator. 5. The disposable cover according to claim 4, wherein said wetness indicator comprises a colorant changing colour upon wetting. 6. The disposable cover according to claim 5, wherein said colorant is a pH-indicator. 7. The disposable cover according to claim 1, wherein the disposable cover comprises a sensor indicating the presence of proteins in the urine, blood in the urine or unwanted bacteria in the urine. 8. The disposable cover according to claim 1, wherein said cover is a laminate of the moisture pervious surface layer and the back sheet, wherein the moisture-pervious surface layer is an apertured plastic film or a nonwoven material and the back sheet is moisture-impervious. 9. The disposable cover according to claim 1, wherein said cover is a laminate comprising the moisture pervious surface layer and the back sheet, wherein the moisture-pervious surface layer is of synthetic fibers or an apertured plastic film and the back sheet is moisture-impervious. 10. The disposable cover according to claim 1, wherein said disposable cover has a defined weight within a tolerance of ±0.5 g. 11. The disposable cover according to claim 1, wherein said disposable cover has a length of 12-50 cm and a width of 6-30 cm. 12. The disposable cover according to claim 1, wherein said disposable cover has a length of 15-45 cm and a width of 7-25 cm. 13. The disposable cover according to claim 1, wherein the disposable cover comprises a sensor indicating the presence of albumin, in the urine. 14. The disposable cover according to claim 1, wherein the disposable cover comprises a sensor indicating the presence of blood in the urine. 15. A method of monitoring basal functions of a premature infant, comprising: ind... | SCA Hygiene Products AB | A61F013/15 A61B005/22 B65B061/26
US07963925 | B1 | 2011-06-21 | Method and apparatus for defining the effect of atrial arrhythmias on cardiac performance and directing therapy using a plurality of intrinsically and extrinsically derived signals | 2007-03-15 | This invention describes methods and algorithms for processing a plurality of relevant signals/data intrinsic to a patient and/or derived from external diagnostic equipment for management of atrial arrhythmias. The intrinsic signals are acquired from intracardiac leads/sensors and analogous extrinsic data obtained from imaging equipment and patient demographics. These data are input into software algorithms that use digital signal processing to output informational data of clinical and technical relevance after comparisons are made to patients with access to this technology whose outcome under varying treatments is known. These combined data are used to define prognosis, make treatment suggestions, direct programming of cardiac devices and digitally convert intrinsically and extrinsically derived indices into a common metric. In a preferred embodiment, the intrinsically and extrinsically acquired data is utilized in the design of catheters and software algorithms for performing intracardiac procedures such as ablation of atrial arrhythmias. | 1. A cardiac system comprising: a catheter having a distal end and a proximal end, said distal end including a sensor configured to contact a portion of a cardiac tissue when said catheter is inserted into said patient, said sensor being configured to sense real time motion of said portion, said real time motion having a real time motion parameter and to generate a corresponding sensor signal having a sensor signal characteristic with a magnitude dependent on said real time motion parameter, said proximal end having a component providing at least one of a vibratory and motion sensation to the hand of an operator holding the handle; anda processor receiving said sensor signal and transforming said sensor signal into respective haptic representation of the motion of the tissue by driving said component, said haptic representation having an amplitude dependent on the magnitude of said real time parameter. 2. The cardiac system of claim 1 further comprising an external sensor monitoring the cardiac tissue externally and a display generating a corresponding real time image of the position of said distal end. 3. The system of claim 2 wherein said external sensor monitors the cardiac tissue using at least one of ultrasonic, electromagnetic, magnetic, radiographic waves, electrical, impedance data. 4. The system of claim 1 wherein said processor and said component cooperate to generate said haptic representation of at least one cardiac parameter selected from the group consisting of cardiac wall displacement magnitude, cardiac wall velocity, cardiac wall acceleration, blood flow, cardiac acoustic signals, cardiac rotation/torsion, cardiac tissue vibration, frequency of motion, cardiac arrhythmia, cardiac structure and cardiac geometry. 5. The system of claim 1 wherein said internal sensor generates sensor signals indicative of at least one of tissue motion at various frequencies and a motion vector having a magnitude and an angle. 6. The cardiac system of claim 1 wherein said processor is configured to drive said component to provide one of said translation, rotation and vibration wherein said translation, rotation or vibration is proportional to the translation, rotation and vibration of the respective tissue. 7. The cardiac system of claim 1 wherein said haptic representation is indicative of the frequency of the motion of said cardiac tissue. 8. A cardiac system comprising: a catheter having a distal end and a proximal end, said distal end including a sensor configured to contact a portion of a cardiac tissue when said catheter is inserted into said patient, said sensor is configured to generate sensor signals indicative of a sensed parameter of the real time motion of said portion, said sensed parameter having a real time sensed amplitude, said proximal end having a component providing at least one of a vibratory and motion sensation to the hand of an operator holding the handle, said sensation having a sensation parameter; anda processor rece... | - | A61B005/02
US08228181 | B2 | 2012-07-24 | Physiological trend monitor | 2011-01-31 | A physiological trend monitor has a sensor signal responsive to multiple wavelengths of light transmitted into a tissue site. The transmitted light is detected after attenuation by pulsatile blood flow within the tissue site. A processor has an input responsive to the sensor signal and a physiological parameter output. Features are extracted from the physiological parameter output. Criteria are applied to the features. An alarm output is generated when the criteria are satisfied. | 1. An oximeter, comprising: a plurality of processing modules, each responsive to a noninvasive sensor signal indicative of light attenuated by body tissue, at least one processing module having different characteristics from another processing module, each of said plurality of processing modules outputting a parallel measurement,an alarm;a memory storing criteria indicative of alarm conditions, said criteria including at least a duration of at least one of said parallel measurements crossing a threshold;an alarm indicator receiving said parallel measurements from said plurality of processing modules, extracting features from said parallel measurements, comparing said features to said criteria, and when said features match said criteria, triggering said alarm on said oximeter. 2. The oximeter according to claim 1 wherein said criteria including at least a trend of said parallel measurements and a behavior of a slope of said trend. 3. The oximeter according to claim 1 wherein said criteria including at least a pattern. 4. The oximeter according to claim 1 wherein said criteria including at least a number of times said measurements crossed said threshold. 5. The oximeter according to claim 1 wherein at least one of said parallel measurements comprises a relatively slow measurement and at least one of said parallel measurements comprises a relatively fast measurement. 6. The oximeter according to claim 5, wherein said alarm indicator extracts threshold features from said slow measurement and at least one waveform feature from said fast measurement and said alarm indicator triggers said alarm based on a combination of said threshold and said waveform features. 7. The oximeter according to claim 5, wherein said alarm indicator extracts threshold features from said slow measurement and triggers said alarm based on said threshold features, said alarm indicator also extracting pattern features from said fast measurement and triggers said alarm based on said pattern features. 8. A method of determining when to alert a caregiver during oximeter-based patient monitoring, the method comprising: providing a plurality of electronic processing modules, at least one processing module having different characteristics from another processing module;providing an alarm;providing a memory;providing an alarm indicator;receiving a noninvasive sensor signal indicative of light attenuated by body tissue;electronically processing said sensor signal with said plurality of processing modules;outputting a parallel measurement from each of said plurality of processing modules;storing criteria indicative of alarm conditions, said criteria including at least a duration of at least one of said parallel measurements crossing a threshold;extracting features from said parallel measurements;comparing said features to said criteria; andwhen said features match said criteria, triggering said alarm. 9. The method according to claim 8 wherein storing said criteria comprises stor... | Masimo Corporation | G08B029/00
US10285597 | B2 | 2019-05-14 | Process and device for monitoring a patient during anesthesia and for determining the combined effect of a plurality of anesthetics | 2011-01-13 | A process is provided for monitoring a patient being anesthetized, as well as a process for determining a combined effect of different anesthetics used. Devices for carrying out the processes are also provided. | 1. An anesthetic process comprising the steps of: administrating a plurality of active anesthetic ingredients to a patient undergoing a painful procedure;determining a concentration of each of the active anesthetic ingredients at a site of action;determining an action of each of the active anesthetic ingredients from the determined concentrations of each active anesthetic ingredient at the site of action;determining an interaction relationship between each of the active anesthetic ingredients;combining the actions of the plurality of active anesthetic ingredients according to the interaction relationship to obtain a combined potency of all the plurality of active anesthetic ingredients;converting said combined potency to a Noxious Stimulus Response Index (NSRI), said NSRI being a one-dimensional parameter and having a full scale range with a first endpoint of said range indicating an alert and nonanesthetized patient, and with a second endpoint of said range indicating a deepest anesthesia;displaying the NSRI to an operator controlling said administering of the plurality of active anesthetic ingredients. 2. An anesthetic process in accordance with claim 1, wherein: the interaction relationship is calculated using a mathematical formula based on the individual potencies of each of the active anesthetic ingredients;the NSRI is calculated using a mathematical formula based on the combined potency. 3. An anesthetic process in accordance with claim 1, further comprising: grouping the active anesthetic ingredients into different classes based on action mechanisms;combining the actions for each of the active anesthetic ingredients in each class to determine a class potency;said determining of the interaction relationship is determined between the different classes of active anesthetic ingredients;said combining of the actions includes combining the class potencies according to the interaction relationship to obtain the combined potency of all the plurality of active anesthetic ingredients. 4. An anesthetic process in accordance with claim 3, wherein: said step of determining an interaction relationship is performed using a mathematical formula. 5. An anesthetic process in accordance with claim 1, wherein: said combined potency of the active anesthetic ingredients are based on a mean anesthetic efficacy of the active ingredients used. 6. An anesthetic process in accordance with claim 1, wherein: said step of determining the interaction includes determining a synergistic interaction between the active anesthetic ingredients used for the anesthesia by means of extrapolation. 7. An anesthetic process in accordance with claim 1, wherein: said combined potency of the anesthetic actions of the active ingredients administered is determined based on a mean anesthetic efficacy of the active ingredients used. 8. An anesthetic process in accordance with claim 1, further comprising the step of: stating the combined potency as a value in a range from 0 to 100.... | DRÄGERWERK AG & CO. KGAA INSELSPITAL BERN, DIREKTION LEHRE UND FORSCHUNG | A61B005/00 A61B005/024 A61B005/083 A61M016/01 A61M016/10 A61B005/0205
US20110130647 | A1 | 2011-06-02 | System And Method Using Coupler-Resonators For Electron Paramagnetic Resonance Spectroscopy | 2009-07-30 | A coupler-resonator for electron paramagnetic resonance (EPR) spectroscopy in subjects has a wire loop formed into a coupling loop, a central transmission portion, and sensor loops. The sensor loops hold EPR sensor materials and are coated with biocompatible plastic. The coupler-resonator is implanted in a subject, the subject in a nonuniform magnetic field with a pickup coil for RF response measurement apparatus near the subject's skin and inductively coupled to the coupling loop. Resonances are measured at multiple sensor loops distinguished by sweeping magnetic field or radio frequency. A biopsy sampler has an outer needle with sensor loop and a central sampling needle with cavity for biopsy samples and EPR sensor material. A device for EPR of fingernails has sensor loops in a partial glove for holding loops next to fingertips. A device for EPR of teeth has sensor loops in plastic chips that can be held between the teeth. | 1. An implantable resonator and coupling device for electron paramagnetic resonance spectroscopy in living mammals or subjects comprising: a conductive wire loop of non-ferrous metal formed into a coupling loop, a central parallel-conductor transmission-line portion, and a sensor loop; andan electron paramagnetic resonance sensor material disposed within the sensor loop;wherein the conductive wire loop is coated with a biocompatible plastic. 2. The implantable resonator and coupling device of claim 1 wherein the transmission line portion is twisted to form a twisted-pair transmission line portion. 3. The implantable resonator and coupling device of claim 1 wherein the transmission line portion comprises parallel wire. 4. The implantable resonator and coupling device of claim 1 wherein the conductive wire loop is further formed into a second sensor loop and a second transmission line portion. 5. The implantable resonator and coupling device of claim 4 wherein the conductive wire loop is further formed into a third sensor loop and a third transmission line portion. 6. The implantable resonator and coupling device of claim 4 wherein the conductive wire loop is further formed into a third transmission line portion, the third transmission line portion disposed between the coupling loop and an origin of the first and second transmission line portions. 7. The implantable resonator and coupling device of claim 4 wherein the coupling loop has a diameter of approximately between one-half centimeter and one and a half centimeters, and the sensor loop a diameter between one-half and one millimeter. 8. A system for measuring parameters in a living mammal or subject comprising: a magnet for providing a nonuniform magnetic field in the mammal or subject;a coupling device comprising a conductive wire loop of non-ferrous metal formed into a coupling loop, a twisted central portion, and at least a first and second sensor loop, and having an electron paramagnetic resonance sensor material disposed within each sensor loop; andapparatus for measuring a radio frequency response coupled to a pickup coil, the pickup coil disposed near a skin surface of the mammal or subject and inductively coupled to the coupling loop, the apparatus for measuring a radio frequency response being capable of measuring a response of the electron paramagnetic resonance sensor material;wherein the electron paramagnetic resonance sensor material is sensitive to a parameter of interest, and wherein the system is capable of measuring resonance for each of the first and second sensor loops individually by sweeping a system parameter selected from the group consisting of a strength of the magnetic field and a frequency of the radio frequency response. 9. The system of claim 8 wherein the electron paramagnetic resonance sensor material is selected from the group consisting of lithium Phtalocyanine, India ink, coals, charcoals, nitroxides, dithiocarbamates, nitrone compounds and nitroso com... | The Trustees of Dartmouth College noneprofit corporation | A61B005/055 G01R033/44
US20110137213 | A1 | 2011-06-09 | METHOD AND SYSTEM FOR THERAPEUTIC EXERGAMING | 2010-11-30 | A computer program product comprising a non-transitory computer readable medium for storing or recording instructions in machine readable form. The instructions, when executed in a motion sensor enabled smart phone, record target data associated with a response of the motion sensor to a desired sequence of movements of a user performing a therapeutic exercise while wearing the smart phone, map the target data to run-time target parameters of a software application, receive performance data associated with a response of the motion sensor to a subsequent sequence of movements of the user performing the therapeutic exercise while wearing the smart phone, map the performance data to the target parameters of the software application to determine the operation of the software application and provide information for display on a remote monitor visible to the user indicating the quality of the user's performance of the exercise. | 1. A computer program product comprising a non-transitory computer readable medium for storing or recording instructions in machine readable form, which when executed in a motion sensor enabled smart phone, is arranged to: record target data associated with a response of the motion sensor to a desired sequence of movements of a user performing a therapeutic exercise while wearing the smart phone;derive run-time target parameters of a software application from the target data;receive performance data associated with a response of the motion sensor to a subsequent sequence of movements of the user performing the therapeutic exercise while wearing the smart phone;map the performance data to the target parameters of the software application to determine the operation of the software application; andprovide information for display on a remote monitor visible to the user indicating the quality of the user's performance of the exercise. 2. The computer program product of claim 1, comprising a module executable on the smart phone for recording the target data; and a separate module executable on the smart phone for receiving the performance data and providing the information for display. 3. The computer program product of claim 1, comprising a module executable on the smart phone for streaming sensor data to a remote server or local computer with application software responsible for both recording target movement data and receiving performance data and providing information for display. 4. The computer program product of claim 1, wherein said software application is a gaming application. 5. The computer program product of claim 1, further arranged to compare said performance data with said target data to determine a performance level associated with the user's ability to perform the exercise. 6. The computer program product of claim 1, further arranged to record input data for indicating software application preferences, wherein said input data comprises at least one of gaming environment, exercise choice, duration, and difficulty level. 7. The computer program product of claim 1, wherein the target data is recorded while the user is subjected to a patient analysis test by a practitioner of the therapeutic exercise. 8. The computer program product of claim 7, wherein the patient analysis test is a star excursion balance test. 9. The computer program product of claim 1, further arranged to receive a configuration input, wherein the configuration input dictates a response of the sensor to movement. 10. The computer program product of claim 9, wherein the configuration input relates to a particular game control of the software application. 11. The computer program product of claim 1, wherein the mapping of the target data to run-time target parameters of the software application corresponds with a mapping of the recorded movements of the user into a particular action in the application software. 12. A smart phone including a motion sensor and a c... | UNIVERSITY COLLEGE DUBLIN, NATIONAL UNIVERSITY OF IRELAND, DUBLIN | A61B005/11 H04W088/02
US20110144516 | A1 | 2011-06-16 | RELATIVE CONTRIBUTION OF THORACIC MUSCLES TO BREATHING | 2009-05-15 | A method and apparatus for evaluating the relative contribution of the diaphragm versus other thoracic muscles to breathing by obtaining measurements of parameters that correlate with changes in thoracic cavity and intra-abdominal cavity pressures over identical time increments and organizing the measurement data in manner that reveals whether the pressures changes are characteristic of the pattern of contemporaneous pressure changes in those cavities that accompany contraction and/or relaxation of the diaphragm muscles or whether the pressure changes are characteristic of the pattern of contemporaneous pressure changes in those cavities that accompany contraction and/or relaxation of the intercostal muscles. | 1. A method for evaluating the relative contribution of the diaphragm versus other thoracic muscles to breathing: obtaining measurements of parameters that correlate with changes in thoracic cavity and intra-abdominal cavity pressures over identical time increments; andorganizing the measurement of data in manner that outputs at least one of the following: a) whether the pressures changes are characteristic of the pattern of contemporaneous pressure changes in those cavities that accompany contraction and/or relaxation of the diaphragm muscles;b) whether the pressure changes are characteristic of the contemporaneous pattern of pressure changes in those cavities that accompany contraction and/or relaxation of the thoracic muscles. 2. A method according to claim 1, wherein the change in thoracic cavity pressure is represented and thereby quantified by a volumetric change in the thoracic cavity. 3. A method according to claim 2, wherein the volumetric change in the thoracic cavity is measured by the tidal volume of a patient's breath. 4. A method according to claim 1, wherein the intra-thoracic pressure is indirectly measured in real time by monitoring intra-esophageal pressure in real time. 5. A method according to claim 4, wherein the intra-esophageal pressure is measured using an intra-gastric tube fitted with a pressure transducer. 6. A method according to claim 1, wherein the intra-abdominal pressure is obtained by measuring intra-gastric pressure in real time. 7. A method according to claim 6, wherein intra-gastric pressure is measured using an intra-gastric tube. 8. A method according to claim 7, wherein intra-gastric pressure is measured using an intra-gastric tube of the type normally used to feed a ventilated patient. 9. A method according to claim 7, wherein the proximal end of the tube is attached to a pressure transducer connected to a microprocessor which records the pressure as a function of time. 10-11. (canceled) 12. A method according to claim 3, wherein machine intelligence is used to plot a graph of tidal volume versus intra-gastric pressure on a breath by breath basis to record distinct tracings that can be differentiated by a technician when the relative contribution of the diaphragm and intercostal muscles to breathing is dominated by one or the other group of muscles, optionally the tracings are loop-like tracings that can be discriminated by eve. 13. (canceled) 14. An apparatus for recording the relative contribution of the diaphragm versus other thoracic muscles to breathing, the apparatus comprising: a means for monitoring (optionally a measurement device) in a series of breaths or common increments in real time:the value of a first variable correlated with intra-thoracic pressurethe value of a second variable correlated with intra-abdominal pressuremachine intelligence outputting the relative direction of flux of the monitored variables. 15. An apparatus according to claim 14, wherein the output is generated ... | - | A61B005/08
US20110127177 | A1 | 2011-06-02 | CONTAINER FOR RECEIVING LAB SAMPLES, AND USE OF A DEVICE FOR TAKING SAMPLES | 2009-07-28 | Container for receiving lab samples, whereby the samples are held in the container at a distance from the internal wall and in a defined height range, and whereby a drying device is provided inside the container at least in said height range and surrounds the sample on the side, at least partially, and includes a drying agent, and is constructed such that gas exchange between sample and drying agent is feasible. | 1. A container for receiving lab samples, whereby the samples are held in the container at a distance from an internal wall and in a defined height range, wherein a drying device, which includes a drying agent, is provided inside the container at least in said height range and surrounding the sample at least partially on the side, such that a gas exchange between sample and drying agent is possible. 2. The container according to claim 1, wherein the drying device is a plastic molded part that contains the drying agent. 3. The container according to claim 1, wherein the drying device is an insert comprising a wall region that forms an external compartment and an internal compartment in the container, whereby the external compartment is intended to be filled with drying agent and perforations enabling gas exchange between sample and drying agent are provided in the wall region. 4. The container according to claim 3, wherein a coating is provided, that covers the wall region on the inside and/or outside and defines a boundary layer that allows water molecules in the gas phase to pass through it, but prevents larger particles from passing through it. 5. The container according to claim 1, wherein the drying device and the drying agent contained therein jointly are provided in the form of a single component that can be inserted into the container. 6. The container according to claim 5, wherein the drying device is a drying agent that is made into a desired shape. 7. The container according to claim 5, wherein the drying device includes an external wall region an internal wall region, and a bottom region, whereby the two wall regions and form an external compartment that can be closed by means of a ring-shaped closure and can be filled with drying agent. 8. A device for sample-taking, consisting of a rod that is held by a container closure and whose free end is cut off obliquely such that an edge is formed and which is provided with a flocking at its free end. 9. A method for producing the container according to claim 2, comprising producing the drying device and the container in a two-step injection molding procedure. 10. The container according to claim 6, wherein the drying agent is made into the desired shape by means of compacting, sintering or extruding. | PRIONICS AG | B65D081/26 A61B005/00 B29C045/14
US20110130636 | A1 | 2011-06-02 | Systems, Methods and Devices for the Rapid Assessment and Deployment of Appropriate Modular Aid Solutions in Response to Disasters. | 2010-08-27 | The invention relates generally to systems, devices and methods for global disaster response, more particularly to the rapid detection, qualified assessment and monitoring of disasters and electronic triage of victims, communication, alert and evacuation systems, provision of suitable modular sensing or medical aid solutions, and their rapid deployment via delivery platforms such as disaster messaging formats and resources on client mobile phone applications or physically via remote operated vehicles (unmanned aerial sea or land systems) or targeted air delivery. | 1. A disaster response system (DRS) comprising a communication and monitoring environment (CME) in optional combination with modular aid solutions (MAS) or a deployment system (DS); where said communication and monitoring environment comprises communication infrastructure capable of data exchange such as via pre-caching, real-time or delayed transmission, from and between central command or distributed information resources and a plurality of client devices in the field comprising at least some of sensor devices, wearable monitoring units, mobile phone or computing devices. 2. A disaster response system according to claim 1 where said data exchange makes use of a compact short message service format, where said format is further codified by means of a disaster messaging standard (DMS) being a means of key codes that can be expanded by reference to previously cached or stored local data resources. 3. A disaster response system according to claims 1 and 2 where said code expansion is by means of a Disaster Markup Language (DML) which enables a mark up or visualization of key features and resources. 4. A disaster response system according to claims 1 and 2 where said data resources may include selectively include information maps, disaster resource ontologies, evacuation routes, emergency resources, medical advice, body maps for electronic triage assessment, damage assessment and reporting resources. 5. A disaster response system according to claims 1 and 2 where said disaster messaging standard (DMS) may be used via a software application on said client device for capturing field data or a user driven data entry, which may selectively include dynamic monitoring of an environmental variable or vital sign, categorizing and facilitating coding and prioritization of a disaster need, coding according to a disaster resource ontology, electronic medical triage of a victim needs, coding and compressing disaster site assessment data such as photographs audio or text reports, coding of medical data or medical resource usage. 6. A disaster response system according to claims 1 and 2 where said disaster messaging standard (DMS) may be used by an incident command resource to selectively broadcast a set of locally specific instructions to said client devices in the field, or to authenticate and authorize information access from a previously cached or local data resource, or to prioritize information allowed to be sent or received across a limited or damaged communication infrastructure. 7. A disaster response system according to claims 1 and 2 where said previously stored local data may selectively be information previously stored on the client device, geo-local information locally cached on the client device on entering a zone or building, data available in an external local disaster data resource point, data relating to a building or transport system, data relating to hazardous material, data relating to a specific disaster category, data relating to a d... | - | A61B005/00 G05B023/02 G06F015/16
US07962201 | B2 | 2011-06-14 | Methods of generating a virtual lead associated with a physiological recording | 2005-09-07 | A method of creating a virtual lead associated with a physiological recording is provided. The method can include obtaining a first physiological signal from a first lead having a first angle and a second physiological signal from a second lead having a second angle, the first lead and the second lead being associated with a subject. The method can also include transforming the first and second physiological signals into a vector representation of the first and second physiological signals and transforming the vector representation to a virtual physiological signal representing a virtual lead having a virtual lead angle. | 1. A method of creating a virtual lead associated with an electrocardiogram signal, comprising the steps of: obtaining a first electrocardiogram signal from a first lead having a first angle and a second electrocardiogram signal from a second lead having a second angle, the first lead and the second lead being associated with a subject;transforming the first and second electrocardiogram signals into a vectorcardiogram;moving average filtering the vectorcardiogram in polar coordinate space to reduce signal artifacts prior to transforming the vectorcardiogram into a virtual electrocardiogram signal; andtransforming the vectorcardiogram into the virtual electrocardiogram signal representing a virtual lead having a virtual lead angle. 2. The method of claim 1, wherein the step of transforming the vectorcardiogram occurs during an overlapping period of time with respect to the step of obtaining the first and the second electrocardiogram signals. 3. The method of claim 1, wherein the virtual lead angle is the same as either the first angle or the second angle. 4. The method of claim 1, wherein the virtual lead angle is different from both the first angle and the second angle. 5. The method of claim 1, wherein the step of transforming the vectorcardiogram to the virtual electrocardiogram signal representing the virtual lead further includes a step of altering a virtual peak magnitude of the virtual electrocardiogram signal to approximate or equal a peak magnitude of either the first electrocardiogram signal or the second electrocardiogram signal. 6. The method of claim 1, wherein either of the first or second electrocardiogram signals has been lost, degraded, or altered prior to the step of transforming the vectorcardiogram into the virtual electrocardiogram signal. 7. The method of claim 1, wherein the signal artifact is a member selected from the group consisting of electrical noise, thermal noise, movement artifacts, breathing artifacts, and combinations thereof. 8. A method of diagnosing and/or monitoring a cardiac condition, comprising: obtaining a first electrocardiogram signal from a first lead having a first angle and a second electrocardiogram signal from a second lead having a second angle, the first lead and the second lead being associated with a subject; determining a monitoring angle;transforming the first and second electrocardiogram signals into a vectorcardiogram;moving average filtering the vectorcardiogram in polar coordinate space to reduce signal artifacts prior to transforming the vectorcardiogram into a virtual electrocardiogram signal; andtransforming the vectorcardiogram into the virtual electrocardiogram signal representing a virtual lead having the monitoring angle. 9. The method of claim 8, wherein the monitoring angle corresponds to an injury current angle. 10. The method of claim 8, wherein the monitoring angle is different from both the first angle and the second angle. 11. The method of claim 8, wherein the moni... | Hewlett Packard Development Company, L.P. | A61B005/0402
US08292810 | B2 | 2012-10-23 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2011-01-27 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. A system for processing data signals obtained from a glucose sensor, the system comprising: a receiving module that receives a signal indicative of an output current from a glucose sensor over a period of time; anda processor module comprising programming that evaluates a spectral density of the signal to detect transient non-glucose related signal artifacts and replaces at least some of the signal artifacts with estimated glucose values. 2. The system of claim 1, further comprising a glucose sensor. 3. The system of claim 1, wherein programming that replaces at least some of the signal artifacts is configured to output the estimated glucose values in response to detection of the transient non-glucose related signal artifacts, wherein the estimated glucose values comprise filtered, calibrated glucose values. 4. The system of claim 3, wherein the processor module further comprises programming configured to output unfiltered, calibrated glucose values when transient non-glucose related signal artifacts are not detected. | DexCom, Inc. | A61B005/00
US07960710 | B2 | 2011-06-14 | Particle radiation therapy equipment | 2006-06-16 | Particle radiation therapy equipment arranged to apply a charged particle beam in a predetermined direction to a region of application within an imaging volume, comprising a charged particle beam source arranged to direct a charged particle beam in the predetermined direction, further comprising magnetic field generation means for generating a magnetic field in the region of application at the same time that the charged particle beam is applied, wherein the magnetic field generation means is arranged to provide access to the region of application for the charged particle beam, and to provide a homogeneous magnetic field in the region of application of the charged particle beam, said magnetic field being directed substantially in the predetermined direction. | 1. Particle radiation therapy equipment arranged to apply a charged particle beam, in a predetermined direction, to a region of application, said radiation therapy equipment comprising: a charged particle beam source arranged to direct a charged particle beam in the predetermined direction; andmagnetic field generation means for generating a magnetic field in an imaging volume which includes the region of application, at the same time that the charged particle beam is applied; wherein,the magnetic field generation means is arranged to provide access to the region of application, and to provide a homogeneous magnetic field in the region of application, said magnetic field being directed substantially in the predetermined direction;the magnetic field generation means comprises magnetic coils, each comprising windings of electrically conductive material;the magnetic coils are arranged symmetrically about a midplane (XY) and about a plane of reflection (YZ) that is perpendicular to the midplane (XY), such that the magnetic coils have a further plane of symmetry (XZ), which is perpendicular to both the midplane (XY) and the plane of reflection (YZ);each of the coils is wound around an axis that is perpendicular to the midplane (XY); andthe windings are configured such that, in operation, current flow is symmetrical about the plane of reflection (YZ) and anti-symmetrical about the midplane (XY), to produce a resultant field at the center of the magnetic field generation means which is perpendicular to the plane of reflection (YZ), parallel to the midplane (XY), and directed substantially in the predetermined direction. 2. The particle radiation therapy equipment according to claim 1, wherein: the magnetic coils are arranged in a group of four sets of coils; andthe sets of coils are arranged symmetrically about the midplane (XY) and the plane of reflection (YZ), such that the magnetic coils are symmetrical about the further plane of symmetry (XZ). 3. The particle radiation therapy equipment according to claim 1, where the magnetic coils comprise D coils. 4. The particle radiation therapy equipment according to claim 1, wherein the magnetic coils comprise banana coils. 5. The particle radiation therapy equipment according to claim 1, wherein the magnetic coils comprise sub-coils. 6. The particle radiation therapy equipment according to claim 5, wherein the sub-coils are provided in one of a nested arrangement and an overlapping arrangement. 7. The particle radiation therapy equipment according claim 1, wherein the coils are non-planar. 8. The particle radiation therapy equipment according to claim 1, wherein the magnetic field generation means is adapted for use in a magnetic resonance imaging system. 9. The particle radiation therapy equipment according to claim 1, further comprising a cryostat unit and a gradient magnet assembly, wherein: the magnetic field generation means is situated within the cryostat that provides an operating temperature ... | Siemens Plc | A61B005/05
US20110152957 | A1 | 2011-06-23 | CHAOS-BASED DETECTION OF ATRIAL FIBRILLATION USING AN IMPLANTABLE MEDICAL DEVICE | 2009-12-21 | Techniques are provided for detecting atrial fibrillation (AF) based on variations in ventricular intervals detected by a pacemaker, implantable cardioverter-defibrillator (ICD) or implantable cardiac monitor (ICM). In one example, ventricular beats are detected and intervals between the ventricular beats are measured, such as RR intervals. Irregular ventricular beats are identified, including ectopic beats, bigeminal beats, and the like. The degree of variability within the ventricular intervals is then determined while excluding any intervals associated with irregular beats. AF is then detected based on the degree of variability. That is, AF is detected based on variability occurring within ventricular intervals after ectopic beats and other irregular beats have been eliminated, thus mitigating detection problems that might arise if the variability were instead calculated based on all ventricular beat intervals. Techniques are also described herein for distinguishing AF from sinus tachycardia, which can also cause a high degree of variability in RR intervals. | 1. A method for use in an implantable medical device comprising: detecting ventricular beats and measuring intervals between the ventricular beats;identifying irregular ventricular beats;evaluating a degree of variability within the ventricular intervals while excluding intervals associated with the irregular beats; anddetecting a supraventricular arrhythmia based on the degree of variability. 2. The method of claim 1 wherein detecting ventricular beats is performed to detect R-waves. 3. The method of claim 2 wherein measuring intervals between the ventricular beats includes measuring RR intervals. 4. The method of claim 1 wherein identifying irregular ventricular beats includes identifying one or more of ectopic beats, bigeminal beats and trigeminal beats, premature atrial contractions (PACs), premature ventricular contractions (PVCs), and junctional beats. 5. The method of claim 1 wherein evaluating a degree of variability includes evaluating a degree of chaos within the ventricular intervals while excluding intervals associated with the irregular beats. 6. The method of claim 5 wherein evaluating the degree of chaos includes: determining dRR values based on RR intervals of the ventricular beats while excluding RR intervals associated with the irregular beats;determining an average dRR value;identifying dRR values deviating significantly from the average; andgenerating a numerical value representative of the number of dRR values that deviate significantly from the dRR average, the numerical value being representative of the degree of chaos. 7. The method of claim 6 wherein detecting a supraventricular arrhythmia includes: comparing the numerical value representative of the degree of chaos against a threshold indicative of supraventricular arrhythmia; andgenerating a signal indicative of an episode of supraventricular arrhythmia if the numerical value representative of the degree of chaos exceeds the threshold. 8. The method of claim 7 further including distinguishing atrial fibrillation (AF) from other supraventricular arrhythmias by: determining whether the episode of supraventricular arrhythmia occurred substantially suddenly; andidentifying the episode as begin AF if the episode of occurred substantially suddenly and identifying the episode as being a non-AF supraventricular arrhythmias otherwise. 9. The method of claim 8 wherein non-AF supraventricular arrhythmias include sinus tachycardia (ST). 10. The method of claim 1 further including distinguishing various cardiac rhythms based on the ventricular intervals. 11. The method of claim 10 wherein distinguishing cardiac rhythms based on the ventricular intervals includes: identifying atrial flutter and junctional rhythms based on substantially tightly varying ventricular intervals;identifying premature atrial contractions (PACs) and premature ventricular contractions (PVCs) based on sudden changes in the ventricular intervals;identifying bigeminal and trigeminal rhythms based on su... | - | A61N001/365 A61B005/046 A61N001/39
US20110118612 | A1 | 2011-05-19 | Valved Catheter with Integrated Pressure Measurement Capabilities | 2009-11-18 | Valve assemblies and valved catheters including pressure sensors, and related methods for treating patients, are disclosed. The valve assembly includes a valve member extending across a first lumen and includes a planar flexible member with an internal slit acting as a valve. The valve assembly further includes a pressure sensor located distally from the valve member and in fluid communication with the first lumen. | 1. A valve assembly for use with a catheter comprising: a proximal end, a distal end in contact with the catheter, and a first lumen extending between said proximal and distal ends, said first lumen being in fluid communication with the catheter,a valve member extending across said first lumen, said valve member comprising: a planar flexible member comprising first and second valve portions separated from one another by an internal slit, the first and second valve portions moving, when subjected to a fluid pressure of at least a predetermined threshold level, to a first open position so that material may flow distally through said valve member, through said first lumen, and into the catheter,the first and second valve portions remaining substantially closed at all times when a fluid pressure exerted thereagainst is less than the threshold level to substantially prevent flow through said valve member; anda pressure sensor located distally from said valve member. 2. The valve assembly of claim 1, further comprising a valve housing and a sensor housing, wherein said valve housing contains said first lumen and said valve member, and said sensor housing comprises said pressure sensor and a second lumen extending between said pressure sensor and said first lumen. 3. The valve assembly of claim 2, wherein said sensor housing is rotatable about said valve housing. 4. The valve assembly of claim 3, wherein when said sensor housing is rotated about said valve housing, it is moveable from a position in which said second lumen is in fluid communication with said first lumen to a position in which said second lumen is not in fluid communication with said first lumen. 5. The valve assembly of claim 4, wherein said sensor housing comprises a first contacting element and said valve housing comprises a second contacting element, said first and second contacting elements being configured to make contact during the rotation of said sensor housing about said valve housing to indicate that said sensor housing is in said position in which said second lumen is in fluid communication with said first lumen. 6. The valve assembly of claim 2, further comprising a pressure seal extending across said second lumen. 7. The valve assembly of claim 1, wherein said distal end of said assembly is insertable into the catheter. 8. The valve assembly of claim 1, further comprising an electrical connection between said pressure sensor and a voltage source. 9. The valve assembly of claim 1, further comprising a data connection between said pressure sensor and a user interface. 10. A valve assembly for use with a catheter comprising: a valve housing and a sensor housing;said valve housing comprising: a proximal end, a distal end in contact with the catheter, and a first lumen extending between said proximal and distal ends, said first lumen being in fluid communication with the catheter,a valve member extending across said first lumen, said valve member comprising a planar flex... | Navilyst Medical, Inc. | A61B005/0215 A61M039/22
US20110152635 | A1 | 2011-06-23 | Motivational Profiling for Behavioral Change Technologies: A State-Trait Approach | 2009-12-23 | An electronic system for health based behavior modification may receive real time data to assess behavioral state changes. The real time data may be used as an input to a database that matches states and trait-based typologies to determine motivational profiles. A rules database may be used to develop recommended motivational coaching changes based on the real time inputs. | 1. A method comprising: electronically collecting state based information relevant to motivation in real time;electronically identifying a motivational profile based on an electronically stored mapping of possible motivational-based states mapped to trait-based typologies; andrecommending an action through a user interface of an electronic device using said motivational profile. 2. The method of claim 1 including developing a table that maps states to trait-based typologies. 3. The method of claim 1 including providing an initial policy for recommended motivational action and modifying that policy based on real time information received thereafter. 4. The method of claim 1 including collecting information about real time states and forwarding said information to a server over the Internet. 5. The method of claim 1 including collecting real time biological state information including at least one of the following of blood pressure, pulse rate, weight, speed, distance or performance. 6. The method of claim 1 including receiving electronic questionnaire responses to identify trait-based typologies. 7. The method of claim 1 including electronically providing motivational coaching for behavior modification that changes based on the user's current motivational state. 8. The method of claim 7 including enabling an electronic device to match states with trait-based typologies and to determine motivational profiles and thereby to select one of a plurality of recommended behavioral modification coaching techniques and forwarding the selected technique to the user over an input/output interface. 9. A computer readable medium storing instructions that are executed by a computer to: collect state based information relevant to motivation in real time;identify a motivational profile based on motivation based states mapped to trait-based typologies; andrecommend an action through a user interface using said motivational profile. 10. The medium of claim 9 further storing instructions to develop a table that maps states to trait-based typologies. 11. The medium of claim 9 further storing instructions to modify an initial assessment of trait-based typology based on real time information. 12. The medium of claim 9 further storing instructions to collect information about real time motivational states and to forward this information to a server over the Internet. 13. The medium of claim 9 further storing instructions to collect biological metrics in real time. 14. The medium of claim 9 further storing instructions to generate an electronic questionnaire to collect information about the user's current state. 15. An apparatus comprising: a controller;a biometric sensor coupled to said controller; anda storage coupled to said controller storing software to electronically collect state based information from said sensor in real time relevant to motivation and to recommend a course of action by mapping a trait-based typology to the user's current motivationa... | - | A61B005/00
US20110160600 | A1 | 2011-06-30 | ERECTILE FUNCTION INDEX MEASURING AND ANALYZING SYSTEM AND MEASURING AND ANALYZING METHOD THEREOF | 2010-12-22 | An erectile function index measuring and analyzing system includes a measuring device and an analyzing device. The measuring device is configured to output a standard pulse signal based upon a penile artery pulse of a subject that is measured while a penis of the subject is in a relaxed state, a first comparative pulse signal based upon a penile artery pulse that is measured after releasing the penis from external pressure larger than a systolic pressure, and a second comparative pulse signal based upon a penile artery pulse that is measured while the penis is in an erect state. The measuring device is further configured to measure heartbeat for outputting an electrocardiogram signal. According to the above signals, the analyzing device is operable to obtain a vasodilatation index, a pulse wave velocity and an erectile index to serve as erectile function indices used for determining erectile function of the subject. | 1. An erectile function index measuring and analyzing method, comprising the steps of: a) measuring a penile artery pulse of a subject to obtain a standard pulse signal while a penis of the subject is in a relaxed state;b) obtaining at least one of a first comparative pulse signal based upon a penile artery pulse of the subject that is measured after releasing the penis of the subject from external pressure larger than a systolic pressure of the subject, and a second comparative pulse signal based upon a penile artery pulse of the subject that is measured while the penis of the subject is in an erect state;c) processing the standard pulse signal and said at least one of the first and second comparative pulse signals to obtain a digital standard pulse signal and at least one of a first digital comparative pulse signal and a second digital comparative pulse signal, respectively; andd) according to signal parameters of the digital standard pulse signal and said at least one of the first and second digital comparative pulse signals, obtaining at least one of a vasodilatation index with reference to the first digital comparative pulse signal and an erectile index with reference to the second digital comparative pulse signal to serve as an erectile function index used for determining erectile function of the subject. 2. The erectile function index measuring and analyzing method as claimed in claim 1, wherein, in step d), the vasodilatation index is obtained according to area values with reference to the digital standard pulse signal and the first digital comparative pulse signal. 3. The erectile function index measuring and analyzing method as claimed in claim 2, wherein the vasodilatation index is calculated as a quotient of a maximum area value with reference to the first digital comparative pulse signal divided by an average area value with reference to the digital standard pulse signal. 4. The erectile function index measuring and analyzing method as claimed in claim 1, wherein, in step d), the vasodilatation index is obtained according to amplitude values with reference to the digital standard pulse signal and the first digital comparative pulse signal. 5. The erectile function index measuring and analyzing method as claimed in claim 4, wherein the vasodilatation index is calculated as a quotient of a maximum amplitude value with reference to the first digital comparative pulse signal divided by an average amplitude value with reference to the digital standard pulse signal. 6. The erectile function index measuring and analyzing method as claimed in claim 1, wherein, in step d), the vasodilatation index is obtained according to increasing slopes and decreasing slopes relative to waveforms of the digital standard pulse signal and the first digital comparative pulse signal under a reactive hyperemia period, the increasing and decreasing slopes of the waveforms being obtained by analyzing connection between peaks of the waveforms using Hilbert-Hu... | - | A61B005/02
US07942827 | B2 | 2011-05-17 | Minimally invasive allergy testing system | 2006-07-11 | An allergy testing system includes encapsulated allergens, a microneedle array, and an activation system coupled to the microneedle array and the encapsulated allergens such that the encapsulated allergens are moved into contact with a subject as the microneedle array is moved from a resting position to a penetrating position. A method for determining a degree of reaction to one or more allergens by a patient in a minimally invasive manner is also disclosed. Penetration of one or more microneedles into a skin of the patient is caused. Each of the penetrations into the skin is exposed with an allergen from each of the one or more microneedles. One or more images are captured of each of the penetrations into the skin. Each of the captured images are analyzed to assess the degree of reaction to the specific allergen. Allergic reactivity data is output for at least one of the allergens. | 1. An allergy testing system, comprising: encapsulated allergens;a microneedle array;an activation system coupled to the microneedle array and the encapsulated allergens such that the encapsulated allergens are moved into contact with a subject as the microneedle array is moved from a resting position to a penetrating position;at least one imaging system configured to capture a plurality of images of penetration sites by the microneedle array;at least one image sensor;an imaging optics system which focuses the images of penetration sites on the at least one image sensor;a second image sensor;a second imaging optics system which focuses the images of the penetration sites on the second image sensor, wherein the image sensors are used to determine a topographic profile of an allergic reaction; andan analyzer coupled to the at least one imaging system and configured to analyze the plurality of images to determine a time rate of change of an allergy response to one or more of the allergens. 2. The allergy testing system of claim 1, wherein the microneedle array pierces the encapsulated allergens while moving to the penetrating position. 3. The allergy testing system of claim 1, wherein the encapsulated allergens are coupled between the activation system and the microneedle array, and the encapsulated allergens are released by a pressure exerted on the encapsulated allergens by the activation system as the microneedle array is moved to the penetrating position. 4. The allergy testing system of claim 1, wherein at least one microneedle in the microneedle array comprises a hollow needle. 5. The allergy testing system of claim 1, wherein at least one microneedle in the microneedle array comprises a grooved needle. 6. The allergy testing system of claim 1, wherein at least one microneedle in the microneedle array comprises a corrugated needle. 7. The allergy testing system of claim 1, wherein the microneedle array comprises at least one needle of a first penetration depth and at least one needle of a second penetration depth which is different from the first penetration depth. 8. The allergy testing system of claim 1, wherein the microneedle array comprises at least one needle with a cross-section that is selected from the group consisting of: square, rectangular, triangular, and circular. 9. The allergy testing system of claim 1, wherein the microneedle array comprises at least one needle with a varying cross-section. 10. The allergy testing system of claim 1, wherein the microneedle array comprises silicon. 11. The allergy testing system of claim 1, wherein the microneedle array comprises glass. 12. The allergy testing system of claim 1, wherein the microneedle array comprises quartz. 13. The allergy testing system of claim 1, wherein the microneedle array comprises a substantially transparent material. 14. The allergy testing system of claim 1, further comprising a package that at least partially encloses the microneedle array and the enca... | Infotonics Technology Center, Inc. | A61B005/00
US20110118583 | A1 | 2011-05-19 | DEVICE FOR DIAGNOSING TISSUE INJURY | 2008-05-07 | The device for diagnosing tissue injury of the invention has a catheter insertable into the body and a radical sensor provided in the catheter, and is characterized in that the radical sensor has a sensor electrode capable of measuring superoxide anion radicals provided at a tip end of the catheter, a lead wire connector for a sensor provided at a basal portion of the catheter, and a lead wire for a sensor for connecting the sensor electrode portion to the lead wire connector for a sensor. With the device, in vivo free radicals typified by superoxide anions in systemic organ or tissue injury caused by cerebral ischemia reperfusion injury, severe infection or sepsis can be promptly and quantitatively monitored, and whether the in vivo tissue conditions are good or not can be accurately diagnosed. | 1. A device for diagnosing tissue injury comprising a catheter insertable into the body, and a radical sensor provided in the catheter, wherein said radical sensor has a sensor electrode capable of measuring superoxide anion radicals provided at a tip end of said catheter, a lead wire connector for a sensor provided at a basal portion of said catheter, and a lead wire for a sensor for connecting said sensor electrode portion to the lead wire connector for a sensor. 2. The device for diagnosing tissue injury according to claim 1, further comprising comparative means for comparing a concentration of superoxide anion radicals in blood measured by said sensor electrode with a predetermined threshold to distinguish a tissue injury-based value and a healthy man-based value. 3. The device for diagnosing tissue injury according to claim 1 or 2, wherein the sensor electrode capable of measuring superoxide anion radicals uses metal porphyrin complex polymerized coating. 4. The device for diagnosing tissue injury according to claim 1 or 2, wherein transfusion supply means is provided in said catheter, and said transfusion supply means has tip end members provided at a tip end of said catheter and having a hole capable of discharging transfusion, transfusion lines each of which is connected to each of these tip end members, and transfusion connectors each of which is coupled to each of these transfusion lines and provided at a basal portion of said catheter. 5. The device for diagnosing tissue injury according to claim 4, wherein at least one of said transfusion lines is coupled to a pressure transducer for venous pressure measurement. 6. The device for diagnosing tissue injury according to claim 4, wherein a hole capable of discharging transfusion is also provided at a middle portion of the catheter. 7. The device for diagnosing tissue injury according to claim 1 or 2, wherein the tissue injury is cerebral ischemia reperfusion injury. 8. The device for diagnosing tissue injury according to claim 1 or 2, wherein the tissue injury is cerebral ischemia reperfusion injury, and further is cerebral ischemia reperfusion injury exacerbated by hyperglycemia (or diabetes or other disease caused by hyperglycemia). 9. The device for diagnosing tissue injury according to claim 1 or 2, wherein the tissue injury is cerebral ischemia reperfusion injury and its treatment is drug therapy or low-temperature therapy. 10. The device for diagnosing tissue injury according to claim 1 or 2, wherein the tissue injury is cerebral ischemia reperfusion injury and its treatment is high-concentration oxygen therapy. 11. The device for diagnosing tissue injury according to claim 1 or 2, wherein the tissue injury is systemic organ or tissue injury caused by severe infection or sepsis. | - | A61B005/04
US07953609 | B2 | 2011-05-31 | Disease and case management system | 2007-12-10 | Disclosed herein is a computer implemented method and system that addresses the need for an online disease based and case based management of medical care. An online disease and case management system is provided. The information of the patients is retrieved from multiple information sources. The patient information is used to identify one of a disease condition in a patient population and a patient with a disease condition. Further, the patients are enrolled. The enrolled patients are educated about the disease condition for self monitoring of the disease condition. The status of the self monitored disease condition is recorded in a disease log. Visits to a provider are scheduled for medical treatment of the enrolled patients. The enrolled patients are monitored for medical progress based on the medical treatment provided. Further, the disease and case management of medical care are evaluated and reported based on clinical and financial outcomes. | 1. A computer implemented method of disease based management of medical care provided to a patient population, comprising the steps of: providing an online disease and case management system for the management of medical care of said disease;providing a plurality of online information sources accessible by the online case management system, wherein said online information sources comprise information entered by a patient and an existing provider for medical treatment, wherein said online information sources include member information, provider information, medical claims, pharmacy claims, referrals and authorizations information, information from labs and images, and patient surveys;retrieving patient information of said patient population from said plurality of online information sources by said online disease and case management system;identifying a disease condition of the patient population from said retrieved patient information;enrolling patients of the patient population with said disease condition for the disease based management of medical care;assigning a disease manager to each of said enrolled patients using said online disease and case management system, wherein said disease manager manages said enrolled patient and provides personalized support for said enrolled patient;educating said enrolled patients about the disease condition for self monitoring of the disease condition;recording status of said self monitored disease condition in a disease log of personal health record on the online disease and case management system by the enrolled patients;scheduling visit to a provider for medical treatment of the enrolled patients by a scheduling engine of the online disease and case management system, wherein said scheduling engine schedules said visit based on said recorded status;synchronizing schedules of the enrolled patients with schedules of one or more of said disease manager and said provider for medical treatment;generating a plurality of reports of the disease based management of medical care provided to the patient population, wherein said generated reports comprise utilization metrics of the disease based management of medical care, further wherein said utilization metrics are used for rewarding said provider, said disease manager, and the enrolled patients; andmonitoring medical progress of the enrolled patients based on said medical treatment;whereby the disease based management of medical care using the online disease and case management system improves quality of the medical care and reduces medical care costs. 2. The computer implemented method of claim 1, wherein the enrolled patients are managed by a disease manager, wherein said disease manager educates the enrolled patients about the disease condition. 3. The computer implemented method of claim 1, wherein said plurality of online information sources includes one or more of claim data, health risk questionnaires, and utilization analytics data. 4. The computer implem... | - | G06Q010/00 G06Q050/00 G06Q099/00 A61B005/00 G06F019/00
US08617066 | B2 | 2013-12-31 | Automated interactive drug testing system | 2010-12-10 | An automated interactive drug testing system designed for testing for substance abuse that provides an immediate test result by visual analysis. The system includes at least one remote device and a central computer. Each remote device as well as the central computer include means for receiving a video feed over a network connection such that a testing subject located at a remote device can use its video camera to bi-directionally communicate with a person located at the central computer. In application, the testing subject faces the video camera so that the operator at the central computer can witness the testing and testing results. Preferably, the testing can include an eye scan and a saliva swab drug test. It is important for purposes of testing integrity, that during the entire testing process, the test specimen is never removed from view of the camera and the operator stationed at the central computer such that the operator at the central computer can view the testing samples to determine whether the test is negative or non-negative. | 1. A method for testing for substance abuse that provides an immediate test result by visual analysis, said method comprising the following steps: receiving a video feed over a network connection between a testing subject at a remote device and an operator at a central computer, each said remote device and central computer includes an imaging device to bi-directionally communicate, and wherein said testing subject is facing the imaging device at the remote device during the testing method so that the operator at the central computer can witness each step of the testing;performing an eye scan test comprising the following steps:capturing a current electronic photograph of the testing subject's right eye,capturing a current electronic photograph of the testing subject's left eye,visually comparing the current electronic photograph of the testing subject's right eye with a baseline electronic photograph of the subject's right eye stored in the central computer, wherein said baseline electronic photograph of the subject's right eye is stored at time of the subject's enrollment,visually comparing the current electronic photograph of the testing subject's left eye with a baseline electronic photograph of the subject's left eye stored in the central computer, wherein said baseline electronic photograph of the subject's left eye is stored at the time of the subject's enrollment,determining from said visual comparisons whether the eye scan test is negative or non-negative and in the event said eye scan test is negative then reporting the results of the eye scan test; and, in the event said eye scan test is non-negative then performing a swab test comprising the following steps in the sequence set forth:removing one test swab, said test swab including a swab holder having a controlled fluid,wiping said test swab under the testing subject's tongue to obtain a sampling,inserting the test swab with the sampling into the swab holder,said operator viewing said sampling from the imaging device at the central computer to determine whether the swab test is negative or non-negative,reporting the results from the swab test. 2. The method of claim 1, wherein the step of reporting the results from the eye scan test includes the step of generating a report indicating the test results and reporting the results to a client. 3. The method of claim 1, wherein the step of reporting the results from the swab test includes the step of generating a report indicating the test results and reporting the results to a client. 4. A method for testing for substance abuse that provides an immediate test result by visual analysis, said method comprising the following steps: receiving a video feed over a network connection between a testing subject at a remote device and an operator at a central computer, each said remote device and central computer including a video camera to bi-directionally communicate, and wherein said testing subject is facing the video camera at the remote devi... | - | A61B005/00
US08827910 | B2 | 2014-09-09 | Magnetically guided catheter with flexible tip | 2008-07-03 | A catheter includes a flexible tubing having a proximal end and a distal end. The catheter also includes an electrode assembly attached to the distal end of the flexible tubing and including a first magnet therein. The electrode assembly including an electrically conductive tip electrode and an electrically nonconductive coupler which is connected between the tip electrode and the distal end of the flexible tubing. The electrically conductive tip electrode comprises a hollow elongated tip electrode including a sidewall provided with one or more elongated gaps extending through the sidewall. The one or more elongated gaps providing flexibility in the sidewall for bending movement of the tip electrode relative to a longitudinal axis of the catheter body. The first magnet is responsive to an external magnetic field to selectively position and guide the electrode assembly within a body of a patient. | 1. A catheter comprising: a flexible tubing having a proximal end and a distal end;an electrode assembly attached to the distal end of the flexible tubing and including a first magnet therein, the electrode assembly including an electrically conductive tip electrode and an electrically nonconductive coupler which is connected between the tip electrode and the distal end of the flexible tubing;wherein the electrically conductive tip electrode comprises a hollow elongated tip electrode including a sidewall provided with one or more elongated gaps extending through the sidewall, the one or more elongated gaps providing flexibility in the sidewall for bending movement of the tip electrode relative to a longitudinal axis of the flexible tubing, wherein the electrically conductive tip electrode comprises a flexibility that enables it to radially deform such that a cross sectional shape of the tip electrode is changed when the electrode impinges on a tissue surface; andwherein the first magnet is responsive to an external magnetic field to selectively position and guide the electrode assembly within a body of a patient. 2. A catheter in accordance with claim 1, further comprising a second magnet spaced from the electrode assembly along a longitudinal axis of the tubing; wherein the first magnet and the second magnet are responsive to an external magnetic field to selectively position and guide the electrode assembly within a body of a patient. 3. A catheter in accordance with claim 2, further comprising a lumen passing through the first magnet and the second magnet, and in fluid communication with the electrode assembly, wherein the one or more elongated gaps allow fluid flow therethrough. 4. A catheter in accordance with claim 2, further comprising a third magnet separated from the second magnet by a first distance along the longitudinal axis of the flexible tubing, wherein the second magnet is spaced from the electrode assembly by a second distance along the longitudinal axis of the flexible tubing, the first distance being greater than the second distance. 5. A catheter in accordance with claim 1, wherein the electrically nonconductive coupler and the tip electrode are coupled by an interlocking connection. 6. A catheter in accordance with claim 5, wherein the interlocking connection is formed by coupling an annular projection on the tip electrode and a groove on the electrically nonconductive coupler. 7. A catheter in accordance with claim 5, wherein the interlocking connection is a snap-fit connection. 8. A catheter in accordance with claim 1, wherein the first magnet is disposed in a cavity formed at least partially inside the electrically nonconductive coupler. 9. A catheter in accordance with claim 1, wherein the electrode assembly further includes a band electrode on an external surface of the coupler, the band electrode being spaced from the tip electrode. 10. A catheter in accordance with claim 1, wherein a portion of the flexible tu... | St. Jude Medical, Atrial Fibrillation Divsion, Inc. | A61B008/14 A61B005/04 A61M031/00 A61B018/18 A61B018/14 A61B018/00 A61B017/00
US08744875 | B2 | 2014-06-03 | Systems and methods for synchronizing data of a patient monitor and a portable sensor module | 2009-12-23 | A patient monitor is synchronized with a portable sensor module by detecting a first coupling of the portable sensor module to the patient monitor. In response to the first coupling, the portable sensor module and the patient monitor each store configuration settings and record patient data acquired through the portable sensor module. After the portable sensor module is decoupled from the patient monitor, the portable sensor module continues to store patient data and the configuration settings stored in the patient monitor and the portable sensor module are allowed to change relative to one another. Upon detecting a second coupling of the portable sensor module to the patient monitor, the portable sensor module and the patient monitor resynchronize with one another to remove any differences in the configuration settings and to provide copies of any patient data missing from either the portable sensor module and the patient monitor. | 1. A patient monitor system, comprising: a patient monitor comprising: a sensor module interface to receive and process patient data according to a first set of configuration settings;a first display device to display the patient data according to the first set of configuration settings; anda first memory device to store the patient data and the first set of configuration settings; anda portable sensor module configured to selectively couple with and decouple from the sensor module interface of the patient monitor, the portable sensor module configured to receive patient data signals directly from one or more sensors attached to a patient, the portable sensor module comprising: a patient data acquisition unit for receiving and processing the patient data signals from the one or more sensors and to communicate the processed patient data signals as patient data to the sensor module interface of the patient monitor; anda second memory device to store the patient data and a second set of configuration settings,wherein in response to coupling the portable sensor module to the patient monitor, the portable sensor module and the patient monitor communicate with each other to synchronize their respective patient data and to selectively synchronize the first set of configuration settings with the second set of configuration settings, wherein the first set and second set of configuration settings define one or more display settings;wherein the first memory device comprises a first circular buffer for storing the patient data and the second memory device comprises a second circular buffer for storing the patient data, wherein a circular buffer overwrites newest data over oldest data when the buffer is full, and wherein the patient monitor is configured to synchronize the respective patient data by:identifying one or more memory blocks in the second circular buffer which includes patient data that is missing from the first circular buffer;identifying one or more memory blocks in the first circular buffer having a time correspondence with the one or more memory blocks identified in the second circular buffer; andcopying the patient data in the identified one or more memory blocks in the second circular buffer to the identified one or more memory blocks in the first circular buffer. 2. The system of claim 1, wherein the portable sensor module further comprises: a second display device; anda user interface,wherein the second display device and the user interface allow the portable sensor module to operate in a monitor mode that allows a user to view at least some of the second set of configuration settings using the second display device. 3. The system of claim 2, wherein the portable sensor module operating in the monitor mode is further configured to provide an alarm based on the patient data signals and to allow the user to silence the alarm. 4. The system of claim 1, wherein the patient monitor comprises a portable patient monitor. 5. The system of clai... | Mindray DS USA, Inc. | G06Q050/22 G06Q050/24
US08942775 | B2 | 2015-01-27 | Handheld apparatus and method for the automated application of cosmetics and other substances | 2011-02-02 | An applicator head is provided for a reflectance modifying agent (RMA) applicator is moved across the skin by means of a floating ring having dispersed raised contact points to maintain a proper distance from the surface to be treated, reduce the influence of outside light during scanning, and limit smudging during deposition. During an application session, software on the computer uses a camera to sense aspects of color and texture on human features, calculates cosmetic enhancements, and uses the printer head to apply RMA precisely to the features to create those enhancements. Skin landmarks are used for registration. The head uses differential lighting by providing a sequence of directional lighting, with some exposures left dark to adjust for ambient light leakage. The exposures are co-synchronized in stacks, where each stack is a grouping of data about a particular instant of time during the scanning. | 1. A handheld device, comprising: an applicator for applying one or more reflectance modifying agent (RMAs), the applicator being configured to reduce outside light from entering around a base of the applicator that comprises edges in contact with an area of skin;a plurality of light sources, each light source being positioned within the applicator and above the base of the applicator, in a respective direction to selectively illuminate the area of skin;a camera for generating images of the area of skin; anda computing environment coupled to the applicator and the camera, the computing environment being configured to: select an aim color,selectively activate light sources of the plurality of light sources to provide a lighting cycle comprising a plurality of lighting modes, the plurality of lighting modes comprising a first lighting mode comprising deactivation of all of the plurality of light sources and a second-lighting mode comprising activation of at least one of the plurality of light sources,selectively activate the camera to generate a plurality of images of the area of skin, each image corresponding to reflective properties of the area of skin associated with a respective lighting mode of the plurality of lighting modes while the applicator reduces outside light from illuminating the area of skin,process the plurality of images to determine an amount of ambient light leaking under edges of the applicator during the first lighting mode and subtracting the amount of ambient light from a subset of the plurality of images corresponding to the second lighting mode,calculate cosmetic enhancements to the area of skin based on the processing,determine a position of the handheld device relative to the area of skin based on the processing, andtransmit a signal to the applicator for applying the one or more RMAs to the area of skin based on the cosmetic enhancements to achieve the aim color. 2. The handheld device of claim 1, wherein the computing environment is configured such that the plurality of lighting modes comprises: a first lighting mode, in which one or more light sources of the plurality of light sources corresponding to a first directional orientation are turned on,a second lighting mode, in which one or more light sources of the plurality of light sources corresponding to a second directional orientation are turned on, anda third lighting mode, in which one or more light sources of the plurality of light sources corresponding to a third directional orientation are turned on. 3. The handheld device of claim 1, wherein the first lighting mode is used to identify at least one of a camera saturation and a level of a sensor noise. 4. The handheld device of claim 1, further comprising an applicator head housing the applicator. 5. The handheld device of claim 4, wherein the computing environment is configured to process the plurality of images, calculate the cosmetic enhancements, determine the position, and transmit the signal to the app... | TCMS Transparent Beauty LLC | A61B005/00 A45D044/00 A45D034/04
US08369925 | B2 | 2013-02-05 | Method to derive anatomical and/or pathological structures from data of imaging technologies | 2010-10-13 | A method to derive anatomical structures from non-invasive imaging technologies is provided. Non-invasive imaging technologies are computed tomography (CT), magnetic resonance imaging (MRI), positron emission tomography (PET), surface scans and others. Imaging data obtained by a non-invasive technology is used to create a surface contour. The imaging data is transmitted to a rapid prototyping apparatus as readable data. With this readable data the rapid prototyping apparatus establishes a positive or negative surface model, such as a fixation device, a mask, or a prosthesis, or other instruments for medical purposes. | 1. A method to derive an artificial anatomical structure from imaging data of an anatomical structure obtained by a non-invasive imaging technology, the non-invasive imaging technology including computed tomography (CT), magnetic resonance imaging (MRI), optical imaging, ultrasound or laser 3-D scanning, the method comprising: a) obtaining the imaging data of the anatomical structure via the non-invasive imaging technology;b) performing digital segmentation, isolation, or both of the imaging data of the anatomical structure;c) performing data transformation for 3-D reconstruction of an inner surface or an outer surface of the anatomical structure or performing reconstruction of the anatomical structure by mirroring data of the anatomical structure; andd) generating the artificial anatomical structure by a rapid prototyping technique, and upon rapid prototyping the artificial anatomical structure, hollow cavities are created into which a radiation influencing material including lead, copper, or bisphosphates are inserted to influence radiation beam distribution upon radio therapy treatment,wherein according to method step c) the artificial anatomical structure is generated in at least two parts for radiation therapy treatment, andwherein the artificial anatomical structure is generated with a variation in wall thickness to protect healthy or sensitive tissue from radiation beams and to focus radiation beams on affected tissue. 2. The method according to claim 1, wherein by adding or subtracting data a complementation of data for 3-D remodeling is performed. 3. The method according to claim 1, wherein the artificial anatomical structure comprises a partial or complete off-set, wherein the partial or complete off-set is a positive or a negative off-set. 4. The method according to claim 1, wherein the artificial anatomical structure comprises foramina for ears, nose and eyes and other parts of the human body. 5. The method according to claim 1, wherein a digital complementation of data into a fixation geometry is performed to allow fixation of separate parts of the artificial anatomical structure to a treatment area. 6. The method according to claim 1, wherein upon generating the artificial anatomical structure, a manual complementation of a fixation geometry is performed during or after a printing process to improve positioning of the artificial anatomical structure in a treatment position. 7. The method according to claim 1, wherein according to step d) the artificial anatomical structure is generated by the rapid prototyping technique based on a layer-by-layer technique. 8. The method according to claim 1, wherein according to step d) the artificial anatomical structure is generated by use of the rapid prototyping technique based on a 3-D printing technique. 9. The method according to claim 1, wherein the artificial anatomical structure obtained according to step d) is treated by powder, glue or a binding material to improve mechanical sta... | DKFZ Deutsches Krebforschungszentrum Stiftung des oeffentlichen Rechts Ruprecht-Karls-Universitaet Heidelberg | A61B005/05
US08483818 | B2 | 2013-07-09 | Enhancements to the detection of pulmonary edema when using transthoracic impedance | 2011-02-23 | This patent document discusses, among other things, systems, devices, and methods for enhancing detection of pulmonary edema using, in addition to thoracic impedance, one or a combination of: physiologic information about a subject, at least one statistical parameter, a user-programmable detection level, at least one parameter associated with a previous pulmonary edema event, and patient symptom information about the subject. In one example, a (base) thoracic impedance threshold is modified to an adjusted thoracic impedance threshold. The adjusted thoracic impedance threshold provides an increased sensitivity of pulmonary edema detection as compared to the base thoracic impedance threshold. In another example, an alert is provided to a subject, a caregiver, or other user based on a pulmonary edema indication determined by the present systems, devices, and methods. In a further example, a therapy (provided to the subject) is adjusted or initiated in response to the pulmonary edema indication. | 1. A system comprising: an implantable thoracic impedance measurement circuit configured to sense a thoracic impedance signal from a subject;a parameter collection device configured to sense or receive physiologic information about the subject, the physiologic information including heart wall motion, the parameter collection device including a heart wall dyssynchrony measurement device, wherein the heart wall dyssynchrony measurement device is configured to sense a first heart wall motion signal from a first cardiac site and a second heart wall motion signal from a second cardiac site, and wherein the parameter collection device is configured to provide a heart wall dyssynchrony signal indicative of the heart wall motion from a phase difference between the first heart wall motion signal and the second heart wall motion signal; anda pulmonary edema indicator coupled with the thoracic impedance measurement circuit to accept the thoracic impedance signal and coupled with the parameter collection device to accept the heart wall dyssynchrony signal, wherein the pulmonary edema indicator is configured to provide a pulmonary edema indication using the thoracic impedance signal and the heart wall dyssynchrony signal. 2. The system of claim 1, comprising a posture sensor adapted to sense a posture signal, the posture sensor being coupled with the pulmonary edema indicator, wherein the pulmonary edema indicator is configured to provide a pulmonary edema indication using the posture signal. 3. The system of claim 2, comprising a posture compensation module configured to compute a posture-compensated thoracic impedance signal from the thoracic impedance signal using the posture signal, wherein the pulmonary edema indicator is configured to determine the pulmonary edema indication using, at least in part, the posture-compensated thoracic impedance signal. 4. The system of claim 1, wherein: the pulmonary edema indicator determines the pulmonary edema indication when the thoracic impedance signal is less than a thoracic impedance threshold; andthe pulmonary edema indicator includes a threshold adjustment module adapted to adjust the thoracic impedance threshold using, at least in part, the physiologic information driven by the presence or absence of pulmonary edema. 5. The system of claim 1, comprising a frequency selective filter circuit configured to extract from the thoracic impedance signal a near-DC thoracic impedance signal. 6. The system of claim 1, comprising: a blood impedance measurement circuit configured to sense a blood impedance signal from the subject; anda blood impedance correction module coupled with the blood impedance measurement circuit, wherein the blood impedance correction module is configured to extract from the thoracic impedance signal a blood resistivity-influenced component using the blood impedance signal, and wherein the pulmonary edema indicator is adapted to determine the pulmonary edema indication using, at least in part, ... | Cardiac Pacemakers, Inc. | A61B005/05 A61B005/04 A61N001/00
US08475768 | B2 | 2013-07-02 | Paramagnetic chelates thereof and their use as contrast agents in magnetic resonance imaging (MRI) | 2008-08-12 | The present invention relates to chelators, in particular to chelators which are capable of forming complexes, i.e. paramagnetic chelates, with paramagnetic metal ions. The invention also relates to said paramagnetic chelates, said paramagnetic chelates linked to other molecules and their use as contrast agents in magnetic resonance imaging (MRI). | 1. Compound of formula (I):whereinR1 is H, NO2, NH2 or NHC(═O)R2, wherein R2 is C1-C6-alkyl, C6-C15-aryl, C7-C22-arylalkyl, O—C1-C6-alkyl, O—C7-C22-arylalkyl, (CH2)n—(C6H4)—NCS, (CH2)m—C≡CH or (CH2)m—N3 wherein n is 0 to 6 and m is 1 to 6;W and the bonds represented as dotted lines are present or absent and when present, W is N; and wherein when W is not present, R1 is NHC(O)R2, wherein R2 is C7-C22-arylalkyl, O—C1-C6-alkyl, O—C7-C22-arylalkyl, (CH2)n—(C6H4)—NCS, (CH2)m—C≡CH or (CH2)m—N3;X is a chelator moiety X consisting of a 6-membered aromatic or partially saturated ring system containing up to three heteroatoms selected from nitrogen and oxygen and having a hydroxyl group as a first substituent bound to a first atom in said ring system, and a hydroxyl group or an oxygen atom doubly bound to a second atom in said ring system wherein said first and second atom are adjacent atoms and wherein said first and second substituents are in ring positions such that X is capable of forming a complex with a paramagnetic metal ion; and wherein X is optionally substituted by up to three additional substituents, R, where each R is independently a hydrophilic group which renders the compound of formula (I) soluble in aqueous solutions. 2. Compound according to claim 1 further comprising a paramagnetic metal ion M to form a compound of formula (II):whereinR1, W and X are as defined in claim 1 and M is a paramagnetic metal ion. 3. Compound according to claim 1 wherein X is derived from hydroxypyrones, dihydroxypyridines, hydroxypyrimidones, hydroxypyridones, hydroxy-pyridinones and dihydroxyphenols. 4. Compound according to claim 2 wherein M is a paramagnetic metal ion of Mn, Fe, La, Co, Ni, Eu, Gd, Dy, Tm or Yb. 5. Compound according to claim 1 wherein R1 is H or NHC(═O)R2, wherein R2 is C1-C6-alkyl, C6-C15-aryl, C7-C22-arylalkyl, O—C1-C6-alkyl or O—C7-C22-arylalkyl. 6. Compound of formula (I):whereinR1 is NHC(═O)R2, wherein R2 is (CH2)n—(C6H4)—NCS, (CH2)m—C≡CH or (CH2)m—N3 wherein n is 0 to 6 and m is 1 to 6;W and the bonds represented as dotted lines are present or absent and when present, W is N; andX is a chelator moiety X consisting of a 6-membered aromatic or partially saturated ring system containing up to three heteroatoms selected from nitrogen and oxygen and having a hydroxyl group as a first substituent bound to a first atom in said ring system, and a hydroxyl group or an oxygen atom doubly bound to a second atom in said ring system wherein said first and second atom are adjacent atoms and wherein said first and second substituents are in ring positions such that X is capable of forming a complex with a paramagnetic metal ion; and wherein X is optionally substituted by up to three additional substituents, R, where each R is independently a hydrophilic group which renders the compound of formula (I) soluble in aqueous solutions. 7. Compound according to claim 1 linked to another molecule via the R1-group. 8. Compound according to claim 7 wher... | GE Healthcare AS | A61B005/055 C08G018/62 C07D213/62
US09259189 | B2 | 2016-02-16 | Magnetic resonance imaging apparatus | 2011-03-04 | A magnetic resonance imaging apparatus performs myocardial perfusion imaging of an object. An imaging unit acquires image data by imaging a heart of the object in synchronism with a biological signal from the object. An image generating unit generates an image concerning the heart of the object based on the image data. The imaging unit applies a probe pulse for detecting body motion of the object before imaging of the heart, and applies a spatial nonselective saturation pulse before application of the probe pulse, and a local selective pulse for flipping back a flip angle of the spatial nonselective saturation pulse with regard to a region to which the probe pulse is applied. | 1. A magnetic resonance imaging (MRI) apparatus which performs myocardial perfusion imaging of an object, the magnetic resonance imaging (MRI) apparatus comprising: an MRI gantry having components including static and gradient magnetic field generators, at least one radio/frequency (RF) coil coupled to an imaging region, RF receiver and transmitter circuits and control circuits connected to control said components, said control circuits including at least one central processing unit and associated memory configured to effectacquisition of image data by imaging a heart of the object in synchronism with a biological signal from the object; andgenerating an image concerning the heart of the object based on the image data,wherein the control circuits are configured to: apply an RF probe pulse to a first region including the object for detecting body motion of the object before imaging of the heart, and to apply, before application of the RF probe pulse, (a) a spatially nonselective RF saturation pulse, and (b) a local spatially selective RF pulse for flipping back to a second region including the first region after the spatially nonselective RF saturation pulse is applied; andapply an RF excitation pulse for MR imaging the heart while adjusting an excited slice position of the heart in accordance with an object's body motion detected by the RF probe pulse. 2. The magnetic resonance imaging (MRI) apparatus according to claim 1, wherein the control circuits are configured to control an imaging position of the heart of the object in accordance with the body motion of the object detected by the RF probe pulse. 3. The magnetic resonance imaging (MRI) apparatus according to claim 1, wherein the control circuits are configured to apply the spatially non-selective RF saturation pulse, the local spatially selective RF pulse, the RF probe pulse, a fat suppression RF pulse, a magnetization stabilization dummy pulse, and a radio-frequency excitation pulse for MR imaging the heart in the order named. 4. The magnetic resonance imaging (MRI) apparatus according to claim 1, wherein the biological signal comprises one of (a) an ECG waveform and (b) a pulse waveform. 5. The magnetic resonance imaging (MRI) apparatus according to claim 1, wherein: the control circuits are configured to control at least one of a time from application of the local spatially selective RF pulse to application of an excitation radio-frequency pulse in the MR imaging process, a flip angle of the local spatially selective RF pulse, and a flip angle of the excitation radio-frequency pulse in the MR imaging process, in accordance with a variation amount concerning a biological signal. 6. The magnetic resonance imaging (MRI) apparatus according to claim 5, wherein the first region includes a heart of the object and the biological signal is an ECG waveform of the object. 7. The magnetic resonance imaging (MRI) apparatus according to claim 5, wherein the first region includes an artery of th... | KABUSHIKI KAISHA TOSHIBA TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/05 A61B005/00 A61B005/055 G01R033/567 G01R033/48 G01R033/483 G01R033/56 G01R033/561 G01R033/563 G01R033/565
US20110118555 | A1 | 2011-05-19 | SYSTEM AND METHODS FOR SCREENING, TREATING, AND MONITORING PSYCHOLOGICAL CONDITIONS | 2010-04-29 | A system for and method of remote monitoring, screening, assessment and treatment of patients having a mental health illness such as post-traumatic stress disorders or other traumatic stress injury and co-occurring symptomatology. Patients in constant communication with one or more healthcare professionals through a wireless network, complete executable programs on their patient handheld electronic devices and transmit the results of the executable programs to their supervising mental health professional. The mental health professionals review and analyze the collected patient data to make clinical assessments of the patients' mental health status. | 1. A computer-implemented system for monitoring post-traumatic stress disorder in a patient in need, comprising: an electronic, portable patient device having at least an input mechanism to enter patient information, a display to review the patient information and communication apparatus; anda computer-implemented processor accessible by one or more mental health professionals and located remote from the patient device, the computer-implemented processor configured to transmit one or more executable programs in a patient clinical module to the patient device to receive executed programs from the patient,wherein the patient clinical modules provide therapies for mental health care for post-traumatic stress disorder or symptoms thereof, selected from the group comprising: a therapy module to monitor and detect a post-traumatic stress disorder symptoms; an informational module to deliver educational material about post-traumatic stress disorder; andwherein the system distributes one or more of these modules to the patient device and can connect patients to health care providers. 2. The system of claim 1, wherein the patient is an adult associated with the military. 3. The system of claim 1, wherein the patient's post-traumatic stress disorder corresponds to a military-related exposure. 4. The system of claim 1, wherein the patient clinical module is transmitted from the computer-implemented processor to the patient device through the communication apparatus. 5. The system of claim 4, wherein contents of the patient clinical module is customized to each patient device. 6. The system of claim 1, wherein the patient clinical module is installed through a removable memory. 7. The system of claim 1, wherein herein the communication means is wireless. 8. The system of claim 1, wherein the communication apparatus of patient device comprises an alert function to alert the healthcare worker of symptoms exceeding predetermined threshold levels corresponding to post-traumatic stress disorder. 9. The system of claim 1, wherein the computer-implemented processor communicates with one or more military databases with patient information. 10. The system of claim 1, the computer-implemented processor communicates with a third-party associated with patient mental healthcare through a communication network, wherein the computer-implemented processor comprises an alert function to electronically alert a first responder party through the communication network when the patient exceeds at least one predetermined threshold diagnostic measurement level corresponding to post-traumatic stress disorder or is deemed by the mental health professional to be having a mental health emergency. 11. The system of claim 1, wherein the handheld device further comprises a security mechanism, wherein patient information cannot be accessed upon activation of the security mechanism. 12. The system of claim 1, wherein a monitoring program comprises: a self-report module, wherein a... | - | A61B005/00
US20110160616 | A1 | 2011-06-30 | SYSTEM AND METHOD FOR ORTHOPEDIC LOAD AND LOCATION SENSING | 2010-06-29 | A sensing insert device (100) is disclosed for measuring a parameter of the muscular-skeletal system. The sensing insert device (100) can be temporary or permanent. The sensing module (200) is a self-contained encapsulated measurement device having at least one contacting surface that couples to the muscular-skeletal system. The sensing module (200) comprises one or more sensing assemblages, electronic circuitry (307), an antenna (2302), and communication circuitry (320). The sensing assemblages are between a top plate (1502) and a bottom plate (1504) in a sensing platform (121). The sensing assemblages measure the parameter and comprise a load disc (2004) and a piezo-resistive sensor (2002). Three sensing assemblages are coupled at predetermined positions to the top plate (1502). The sensing module (200) can measure a location where the parameter is applied to the top plate (1502). | 1. A measurement system for measuring a parameter of the muscular-skeletal system comprising: a sensing platform comprising: a contacting surface; andat least three sensing assemblages coupled to the contacting surface at predetermined locations, where each sensing assemblage provides a voltage or current magnitude that corresponds to the parameter applied to the contacting surface, and where a location where the parameter is applied to the contacting surface is determined from the parameter measurements of the at least three sensing assemblages. 2. The measurement system of claim 1 where the sensing platform is in a sensing module and where the sensing module comprises: at least one printed circuit board with electronic circuitry mounted thereto;flexible interconnect coupling the electronic circuitry and the at least three sensing assemblages; andan encapsulating enclosure enclosing the sensing platform, printed circuit board, and flexible interconnect. 3. The measurement system of claim 2 where the printed circuit board underlies the sensing platform in the encapsulating enclosure. 4. The measurement system of claim 2 where the encapsulating enclosure comprises: a housing; anda cap overlying an opening of the housing where the cap is sealed to the housing, where the cap couples to the contacting surface of the sensing platform, and where the cap moves in relation to the housing when a force, pressure, or load is applied thereto. 5. The measurement system of claim 4 where the sensing platform is supported by at least one support structure of the housing. 6. The measurement system of claim 1 further including:: electronic circuitry comprising: operational circuitry operatively coupled to the at least three sensing assemblages;power management circuitry operatively coupled to the operational circuitry;communication circuitry coupled to the operational circuitry to wirelessly transmit measured parameter data; anda power source operatively coupled to the electronic circuitry. 7. The measurement system of claim 6 where the electronic circuitry is an application specific integrated circuit. 8. The measurement system of claim 1 where a sensing assemblage further includes a piezo-resistive sensor. 9. The measurement system of claim 1 where a sensing assemblage further includes one of a MEMS sensor, a strain gauge, or a mechanical sensor. 10. The measurement system of claim 1 further including an insert that includes the sensing platform where the insert is used intra-operatively and where the insert is substantially equal in dimensions to a final insert. 11. The measurement system of claim 1 further including a final insert where the final insert includes the sensing platform. 12. A measurement system for measuring a force, pressure, or load applied by the muscular-skeletal system where a prosthetic component includes a sensing module and where the sensing module includes an elastic member to return a sensing platform to a predetermined posit... | OrthoSensor | A61B005/103 G01R029/22
US20110112394 | A1 | 2011-05-12 | NEUROMODULATION OF DEEP-BRAIN TARGETS USING FOCUSED ULTRASOUND | 2010-11-05 | Disclosed are methods and systems for non-invasive deep brain or superficial neuromodulation for up-regulation or down-regulation using ultrasound impacting one or multiple points in a neural circuit to produce Long-Term Potentiation (LTP) or Long-Term Depression (LTD) to treat indications such as neurologic and psychiatric conditions. Ultrasound transducers are positioned by spinning them around the head on a track, as well as individually rotated or not, with control of direction of the energy emission, intensity, frequency, and phase/intensity relationships to targeting and accomplishing up-regulation and/or down-regulation. Alternatively the ultrasound transducers may be at fixed locations on the track. Use of ancillary monitoring or imaging to provide is optional. | 1. A method of neuromodulating one or a plurality of deep-brain targets using ultrasound stimulation, the method comprising: aiming one or a plurality of ultrasound transducers at one or a plurality of deep-brain targets,applying power to each of the ultrasound transducers via a control circuit thereby neuromodulating the activity of the deep brain target region,moving one or a plurality of transducers around a track surrounding the mammal's head. 2. The method of claim 1, further comprising identifying a deep-brain target. 3. The method of claim 1, further comprising where neuromodulation of a plurality of targets is selected from the group consisting of up-regulating all neuronal targets, down-regulating all neuronal targets, up-regulating one or a plurality of neuronal targets and down-regulating the other targets. 4. The method of claim 1, wherein the step of aiming comprising orienting the ultrasound transducer and focusing the ultrasound so that it hits the target. 5. The method of claim 1, wherein the acoustic ultrasound frequency is in the range of 0.3 MHz to 0.8 MHz. 6. The method of claim 1, where in the power applied is selected from group consisting of less than 180 mW/cm2 and greater than 180 mW/cm2 but less than that causing tissue damage. 7. The method of claim 1, wherein a stimulation frequency of 300 Hz or lower is applied for inhibition of neural activity. 8. The method of claim 1, wherein the stimulation frequency is in the range of 500 Hz to 5 MHz for excitation. 9. The method of claim 1, wherein the focus area of the pulsed ultrasound is selected from the group consisting of 0.5 to 500 mm in diameter and 500 to 1500 mm in diameter. 10. The method of claim 1, wherein the number of ultrasound transducers is between 1 and 25. 11. The method of claim 1, wherein a disorder is treated by neuromodulation, wherein the target brain regions are selected from the group consisting of NeoCortex, any of the subregions of the Pre-Frontal Cortex, Orbito-Frontal Cortex (OFC), Cingulate Genu, subregions of the Cingulate Gyrus, Insula, Amygdala, subregions of the Internal Capsule, Nucleus Accumbens, Hippocampus, Temporal Lobes, Globus Pallidus, subregions of the Thalamus, subregions of the Hypothalamus, Cerebellum, Brainstem, Pons, or any of the tracts between the brain targets. 12. The method of claim 1, wherein the disorder treated is selected from the group consisting of: addiction, Alzheimer's Disease, Anorgasmia, Attention Deficit Hyperactivity Disorder, Huntington's Chorea, Impulse Control Disorder, autism, OCD, Social Anxiety Disorder, Parkinson's Disease, Post-Traumatic Stress Disorder, depression, bipolar disorder, pain, insomnia, spinal cord injuries, neuromuscular disorders, tinnitus, panic disorder, Tourette's Syndrome, amelioration of brain cancers, dystonia, obesity, stuttering, ticks, head trauma, stroke, and epilepsy. 13. The method of claim 1 where the ultrasound is applied for the purpose selected from the group co... | - | A61B005/055 A61N007/00 A61N001/05 A61N002/00 A61N005/06
US07966055 | B2 | 2011-06-21 | Magnetic resonance system and method for cardiac imaging | 2007-06-01 | In a method for representation of the heart in a magnetic resonance system at least one MR overview image of the heart is acquired. An image plane with a predetermined position relative to the magnetic resonance system is selected for this overview image. The acquired MR overview image is displayed and a number of marking points are established in the displayed overview image. Further image planes for representation of the heart are calculated using some of the established marking points. Further MR images are acquired in the calculated image planes. | 1. A method for representing a heart of a subject by magnetic resonance (MR), using an MR system, comprising the steps of: acquiring at least one MR overview image of a heart in an image plane having a predetermined position relative to the MR system;visually displaying the acquired MR overview image;allowing user interaction with the displayed MR overview image to manually mark five punctiform marking points at respective anatomical points in the displayed MR overview image;automatically electronically calculating three image planes each containing the long axis of the heart using only the five manually marked marking points; andacquiring further MR images respectively in the calculated three image planes. 2. A method as claimed in claim 1 comprising selecting said predetermined position of the image plane of the MR overview image as the pseudo-short axis slice, which contains the right and left ventricles, both atria, and the large vessels of the heart. 3. A method as claimed in claim 1 comprising establishing said five manually marked marking points on the displayed MR overview image by manually marking at least two points on a straight line defined by the base of the left ventricle and the apex of the heart. 4. A method as claimed in claim 1 comprising automatically electronically calculating image planes for respective images of the short axis of the heart from the image planes for the long axis of the heart, as planes that are substantially perpendicular to the long axis of the heart. 5. A method as claimed in claim 1 wherein said MR system comprises a magnet having a center, and comprising moving the subject to cause a predetermined region of the body of the subject to be located at the center of the magnet for acquiring said MR overview image, with said MR overview image being in an image plane having a predetermined angle and a predetermined position relative to said center of said magnet. 6. A method as claimed in claim 1 comprising marking additional manually marked marking points in the displayed overview image respectively designating intersection points of the right ventricle and the left ventricle, and automatically electronically generating a heart model from said five manually marked marking points and the manually marked intersection points. 7. A method as claimed in claim 1 comprising automatically electronically determining a quantitative function of the left ventricle of the heart using said five manually marked marking points on the displayed overview image. 8. A method as claimed in claim 1 comprising automatically electronically segmenting respective anatomical regions of the heart using said five manually marked marking points on the displayed overview MR image. 9. A method as claimed in claim 1 comprising automatically electronically segmenting images of the pseudo-short axis slice, which contains the right and left ventricles, both atria, and the large vessels of the heart, using said five manually marked markin... | Siemens Aktiengesellschaft | A61B005/055
US07949165 | B2 | 2011-05-24 | Method for determining a temporal profile of a probability distribution with local three-dimensional resolution for the presence of a substance in a vascular system | 2008-03-27 | The invention relates to a method for determining a probability distribution with local three-dimensional resolution for a substance in a vascular system. An acquisition time is assigned to a group of x-ray images showing a presence distribution with local 2D resolution for a substance in the vascular system at the acquisition time. A first preliminary presence probability distribution is determined using the x-ray images assigned to the acquisition time and a volume data set of the vascular system. A second preliminary presence probability distribution is determined using the presence probability distribution for the preceding acquisition time and a migration probability distribution with the substance migrating from first vascular locations to second vascular locations. A high individual probability to vascular locations is assigned if both preliminary presence probability distributions have a high probability value for the respective vascular location. Otherwise a lower and greater than zero individual probability is assigned. | 1. A method for determining a presence probability distribution with local three-dimensional resolution for a substance in a vascular system of a patient, comprising: predetermining a volume data set and a temporal sequence of x-ray images of the vascular system;assigning acquisition times to the x-ray images showing a presence distribution with local two-dimensional resolution for the substance in the vascular system at the acquisition times;determining a first preliminary presence probability distribution for a first acquisition time based on the x-ray images assigned to the first acquisition time and the volume data set with the substance at a first vascular location;determining a second preliminary presence probability distribution for a second acquisition time based on the first preliminary probability distribution and a migration probability distribution with the substance migrating to a second vascular location at the second acquisition time; andassigning a high probability to the first and the second vascular locations if both the first and the second preliminary presence probability distributions have a high probability value or assigning a low probability to the first or the second vascular location if the respective first or the second preliminary presence probability distribution has a low but greater than zero probability value. 2. The method as claimed in claim 1, wherein the presence probability distribution is determined by multiplying the first and the second preliminary presence probability distributions. 3. The method as claimed in claim 1, wherein the first preliminary presence probability distribution is determined by steps comprising: locating substance locations in the vascular system in the x-ray images;defining projection lines in the volume data set for the substance locations;selecting nearest projection lines from the projection lines for vascular locations in the x-ray images;calculating distances between the nearest projection lines and the vascular locations; anddetermining the first preliminary presence probability distribution as a function of the distances. 4. The method as claimed in claim 3, wherein the first preliminary presence probability distribution is above an upper limit probability if the greatest of the distances is below a maximum distance, orwherein the first preliminary presence probability distribution is between a lower limit probability that is greater than zero and the upper limit probability. 5. The method as claimed in claim 1, wherein the first preliminary presence probability distribution is determined by steps comprising: locating substance locations in the vascular system in the x-ray images;selecting nearest substance locations from the substance locations for vascular locations in the x-ray images;defining mapping locations for the vascular locations in the x-ray images;calculating distances between the mapping locations and the nearest substance locations; anddetermining the first p... | Siemens Aktiengesellschaft | G06K009/00 A61B006/00 A61B005/02 G06F017/18
US07942823 | B2 | 2011-05-17 | Methods and apparatus for stroke patient treatment | 2006-01-25 | Methods and apparatus for assessing the condition of and treating patients for stroke during the delivery of continuous positive airway pressure (CPAP) are disclosed. By determining central and obstructive apnea occurrences or the percentile of pressure delivered to the patient from patient airflow, stroke indicators may be calculated and analyzed to provide information on the type of stroke a patient has suffered and the degree of a patient's neuro-recovery. The apparatus may be programmed with a particular protocol intended to identify between alternative forms of CPAP treatment appropriate for the patient based upon an assessment of the patient's condition. The device can select and recommend between the various treatment forms, identify billing reimbursement codes and generate data to track patient response to stroke treatments. | 1. A method for evaluating the condition of a patient after stroke comprising the steps of: delivering breathable gas at a pressure above atmospheric to a patient after stroke;measuring the airflow of said patient; anddetermining, using a computer, a stroke indicator from said airflow, said stroke indicator representing information about the patient's stroke condition. 2. The method of claim 1 wherein said determining step includes analyzing said airflow to determine an occurrence of central apnea and said stroke indicator is calculated as a function of said occurrence of central apnea. 3. The method of claim 2 wherein said determining step includes analyzing said airflow to determine an occurrence of obstructive apnea and said stroke indicator is calculated as a function of said occurrence of obstructive apnea. 4. The method of claim 1 further comprising the step of evaluating changes in said stroke indicator to assess the efficacy of an administered drug. 5. The method of claim 4 wherein said step of evaluating said stroke indicator includes storing said stroke indicator in a database of patient information. 6. The method of claim 5 wherein said stroke indicator is a function of a number of occurrences of central apnea. 7. The method of claim 6 wherein said stroke indicator is a function of a number of occurrences of obstructive apnea. 8. A method for evaluating the condition of a patient after stroke in accordance with claim 1 further comprising the steps of: determining the number of central and obstructive apneas which the patient has during a first time period; andcalculating a stroke indicator index from a comparison of the number of central and obstructive apneas which the patient has during the first time period, said stroke indicator representing information about the patient's rehabilitation from stroke. 9. An apparatus for evaluating the condition of a patient after stroke comprising: a source for delivering breathable gas at a pressure above atmospheric to a patient after stroke;a sensor for measuring the airflow of said patient; andmeans for determining a stroke indicator from said airflow, said stroke indicator representing information about the patient's stroke condition. 10. An apparatus for evaluating the condition of a patient after stroke in accordance with claim 9 wherein said determining means analyzes said airflow to determine occurrences of central apneas and said stroke indicator is determined as a function of said occurrences of central apneas. 11. An apparatus for evaluating the condition of a patient after stroke in accordance with claim 10 wherein said determining means analyzes said airflow to determine occurrences of obstructive apneas and said stroke indicator is determined as a function of said occurrences of obstructive apneas. 12. An apparatus for evaluating the condition of a patient after stroke in accordance with claim 9 wherein said determining means further evaluates changes in said stroke indica... | Resmed Limited | A61B005/08
US08261611 | B2 | 2012-09-11 | Method for analyzing gait pattern | 2010-12-15 | A method for analyzing a gait pattern includes: measuring, by a plurality of force sensing resistor (FSR) sensors, foot pressure values, and outputting the measured foot pressure values, respectively; searching for a maximum pressure local area in which the sum of the output values from the FSR sensors included in each of a plurality of pressure local areas is maximized; calculating a center of pressure (COP) with respect to the detected maximum pressure local area; and analyzing a gait pattern by adding the calculated COP to the trajectory of COPs. | 1. A method for analyzing a gait pattern, the method comprising: measuring, by a plurality of force sensing resistor (FSR) sensors, foot pressure values, and outputting the measured foot pressure values, respectively;searching for a maximum pressure local area at which the sum of the output values from the FSR sensors included in each of a plurality of pressure local areas is maximized;calculating a center of pressure (COP) with respect to the detected maximum pressure local area; andanalyzing a gait pattern by adding the calculated COP to the trajectory of COPs. 2. The method of claim 1, further comprising: removing an island as an error from the output values of the plurality of FSR sensors. 3. The method of claim 2, wherein the output values of the plurality of FSR sensors are values obtained by sampling the foot pressure values by more than the total number of rows of left and right FSR sensors. 4. The method of claim 1, wherein the pressure local areas are an aggregate of the FSR sensors each connected with an output value greater than a pre-set reference value. 5. The method of claim 1, wherein the detecting of the maximum pressure local area comprises: calculating the sum of the output values from the FSR sensors included in each of the pressure local areas;determining the calculated sum of the output values as a total pressure value representing each of the pressure local area; anddetecting a pressure local area having a maximum total pressure value among the plurality of pressure local areas, as the maximum pressure local area. | Electronics and Telecommunications Research Institute | A61B005/103
US08840551 | B2 | 2014-09-23 | Tethering capsule system | 2010-12-20 | A capsule system comprising a first capsule (16) configured to be ingested by a subject, a second capsule (18) configured to be ingested by the subject separately from the first capsule, the first capsule having an electromagnetic element (29, 30) and a sensor (22, 23) for sensing a physiological parameter of a gastrointestinal tract of the subject, the second capsule having a magnetic element (29, 30) and a sensor (22, 23) for sensing a physiological parameter of the gastrointestinal tract of the subject, the first capsule and the second capsule having an active state wherein a magnetic force is selectively provided between the first and second capsules attracting the first and second capsules while the capsules are within the gastrointestinal tract of the subject and having an inactive state wherein the magnetic force is absent, and a control (26/32) configured to selectively place the capsules in the active state from the inactive state while the capsules are within the gastrointestinal tract of the subject. | 1. A capsule system comprising: a first capsule configured to be ingested by a subject;a second capsule configured to be ingested by said subject separately from said first capsule;said first capsule having an electromagnetic element and a sensor for sensing a physiological parameter of a gastrointestinal tract of said subject;said second capsule having a magnetic element and a sensor for sensing a physiological parameter of said gastrointestinal tract of said subject;each said first capsule and said second capsule having an active state wherein a magnetic force is selectively provided between said first and second capsules attracting said first and second capsules while said capsules are within said gastrointestinal tract of said subject and having an inactive state wherein said attracting magnetic force is absent; anda control configured to selectively place each of said capsules in said active state from said inactive state while said capsules are within said gastrointestinal tract of said subject. 2. The system set forth in claim 1, wherein said magnetic element is an electromagnet. 3. The system set forth in claim 1, wherein said control comprises a microprocessor and a switch connected to said electromagnet. 4. The system set forth in claim 1, wherein said control comprises an actuator configured to move said magnetic element from a first position to a second position closer to an outside surface of said capsule than said first position. 5. The system set forth in claim 1, wherein said first capsule comprises a receiver configured to receive a signal from exterior of said gastrointestinal tract of said subject and wherein said control places said capsule in said active state as a function of said signal. 6. The system set forth in claim 1, wherein said sensor of said first and/or second capsule provides a signal as a function of said physiological parameter of said gastrointestinal tract of said subject and said control places said capsules in said active state as a function of said signal. 7. The system set forth in claim 6, wherein said control places said capsules in said active state when said sensor senses said parameter above or below a selected amount. 8. The system set forth in claim 1, wherein said control is configured to selectively place said capsules in said inactive state from said active state while said capsules are within said gastrointestinal tract of said subject. 9. The system set forth in claim 8, wherein said first capsule comprises a receiver configured to receive a signal from exterior of said gastrointestinal tract of said subject and wherein said control places said capsules in said inactive state as a function of said signal. 10. The system set forth in claim 8, wherein said sensor of said first and/or second capsule provides a signal as a function of said physiological parameter of said gastrointestinal tract of said subject and said control places said capsules in said inactive state as a function of sa... | Given Imaging, Inc. | A61B005/07 A61B005/145 A61B005/06 A61B001/00 A61B005/01 A61B001/04 A61B005/03
US08588883 | B2 | 2013-11-19 | Electrode assemblies and bruxism monitoring apparatus | 2008-09-22 | The invention relates to an apparatus for monitoring muscle activity related to bruxism, said apparatus comprising:—an electrode assembly for providing signals indicative of said muscle activity, means for processing said signals in order to detect said bruxism, and means for providing a feedback signal in response to detecting said bruxism; wherein said electrode assembly comprises three electrodes in a fixed spatial relationship one to another, each electrode having a contact area for electrical connection with the skin which is spaced from the contact area of each other electrode in the electrode assembly by at least 2 mm, the maximum distance from an edge of one electrode contact area to the furthest edge of the furthest away of the other electrode contact areas being not greater than 60 mm. | 1. Apparatus for monitoring muscle activity related to bruxism, said apparatus comprising: an electrode assembly for providing signals indicative of said muscle activity,means for processing said signals in order to detect said bruxism, andmeans for providing a feedback signal to said electrode assembly in response to detecting said bruxism;wherein said electrode assembly comprises three electrodes in a fixed spatial relationship one to another, each electrode having a contact area for electrical connection with the skin which is spaced from the contact area of each other electrode in the electrode assembly by at least 2 mm, the maximum distance from an edge of one electrode contact area to the furthest edge of the furthest away of the other electrode contact areas being not greater than 60 mm,wherein the electrodes are arranged such that centres of said electrodes lie on a line which is an arc of a circle having a radius of from 60-200 mm,wherein the contact areas of the electrodes extend along or on said line a distance of from 3 to 7 mm from one side to an opposite side of the electrode contact area, andwherein the contact areas of the electrodes extend transversely of said line by a distance of from 3 to 7 mm from one side to an opposite side of the electrode contact area. 2. Apparatus as claimed in claim 1, wherein the maximum distance along said line from an edge of a first said electrode contact area to an opposite edge of the furthest away of the other two electrode contact areas is from 20 to 50 mm. 3. Apparatus as claimed in claim 1, wherein the contact areas of the electrodes are arranged in a triangular arrangement in which the longest side of the triangle so defined is not more than 40 mm. 4. Apparatus as claimed in claim 1, wherein said means for processing said signals in order to detect said bruxism measures the amplitude of EMG signals received from said electrode assembly at one or more predetermined frequencies. 5. Apparatus as claimed in claim 4, wherein said apparatus is operable in a set-up mode and in a use mode and in said set-up mode is configured to accept an input defining said one or more predetermined frequencies and in said use mode detects bruxism episodes by measuring the amplitude of EMG signals at said one or more frequencies received from said electrode assembly. 6. Apparatus as claimed in claim 5, wherein said apparatus is operable in said set-up mode to measure a first power spectrum of signals received from said electrode assembly when a user carries out a first exercise simulating bruxism muscle activity and to measure a second and/or a third power spectrum of signals received from said electrode assembly when a user carries out a second and/or a third exercise simulating muscle activity from which bruxism muscle activity is to be distinguished. 7. Apparatus as claimed in claim 6, wherein the second power spectrum of signals received from said electrode assembly is measured when a user carries out norm... | Sunstar Suisse SA | A61B005/0488 A61B005/0492
US07963916 | B2 | 2011-06-21 | Method and apparatus for estimating a likelihood of shoulder dystocia | 2005-10-28 | A system and apparatus implementing a graphical user interface module for displaying information conveying a level of risk of delivery with shoulder dystocia associated to an obstetrics patient. The graphical user interface module displays a set of user modifiable information fields for allowing a user to enter a set of patient information data elements associated to the obstetrics patient. The graphical user interface module also displays a control allowing a user to cause the set of information data elements to be transmitted to a processing unit adapted to derive patient risk data at least in part on the basis of the set of information data elements, the patient risk data conveying a level of risk of delivery with shoulder dystocia. The graphical user interface module then displays the patient risk data with reference to risk assessment information. | 1. A computer readable storage medium storing a program element suitable for execution by a processing unit, said program element implementing a graphical user interface module for displaying information conveying a level of risk of delivery with shoulder dystocia associated with an obstetrics patient, said graphical user interface module being configured for: a) displaying a set of user modifiable information fields for allowing a user to enter a set of patient information data elements associated with the obstetrics patient, said set of patient information data elements including: i. a maternal weight component;ii. a maternal height component; andiii. a fetal weight component;b) displaying a control allowing the user to cause said set of information data elements to be transmitted to a remote processing unit, and causing the remote processing unit to derive patient risk data at least in part based on said set of information data elements, the patient risk data conveying a level of risk of delivery with shoulder dystocia, said patient risk data being conditioned at least in part based on the fetal weight component and on a ratio between the maternal weight component and the maternal height component;c) receiving the patient risk data from the remote processing unit;d) displaying information derived from the patient risk data for viewing by the user. 2. A computer readable storage medium as defined in claim 1, wherein said patient risk data includes a ranking data element conveying a level of risk of delivery with shoulder dystocia associated with the obstetrics patient. 3. A computer readable storage medium as defined in claim 2, wherein said ranking data element is a likelihood score. 4. A computer readable storage medium as defined in claim 1, wherein said graphical user interface module is configured for displaying information derived from the patient risk data concurrently with risk assessment information, said risk assessment information conveying a reference range of levels of risk associated with women who delivered without shoulder dystocia. 5. A computer readable storage medium as defined in claim 4, wherein said reference range of levels of risk is a first reference range of levels of risk, said risk assessment information further conveying a second reference range of levels of risk associated with women who delivered with shoulder dystocia. 6. A computer readable storage medium as defined in claim 1, wherein said graphical user interface module is configured for displaying information derived from the patient risk data concurrently with risk assessment information, said risk assessment information conveying a reference range of levels of risk associated with women who delivered with shoulder dystocia. 7. A computer readable storage medium as defined in claim 1, wherein said graphical user interface module is configured for displaying information derived from the patient risk data in graphical format. 8. A computer readable stor... | Perigen, Inc. | A61B005/00
US09220412 | B2 | 2015-12-29 | Method and apparatus for analysis of turbid media via single-element detection using structured illumination | 2010-11-19 | Method and apparatus for obtaining qualitative and quantitative analysis of the optical properties or structures of tissue or turbid medium at one or more wavelengths via 1) detection at a single spatial location on the surface of a turbid medium (such as tissue) under two or more structured light conditions or 2) detection at two or more spatial locations on the surface under a single structured light condition. | 1. An apparatus for determining surface or subsurface optical properties or structures of a sample of turbid media over an area of the sample comprising: a source to expose an area of the sample to a plurality of structured illuminations each having one or more spatial features, wherein the one or more spatial features of individual ones of the plurality of structured illuminations differing from the one or more spatial features of other individual ones of the plurality of structured illuminations, and wherein at least one of the plurality of structured illuminations having a plurality of spatial features comprising one or more peaks and one or more valleys;a detection system comprising an optical relay device and a CCD coupled to the optical relay device, wherein the detection system is configured to collect an optical signal with no spatial dependence emitted from a collection area at a single spatial location on the surface of the sample within the area of the sample illuminated by the plurality of structured illuminations, wherein the collection area is smaller in size than the area of the sample exposed to the plurality of structured illuminations and having a dimension less than or equal to the length of an individual spatial feature of the plurality of spatial features of the at least one of the plurality of structured illuminations; anda signal processor coupled to the detection system and configured to reconstruct optical data from the optical signals with no spatial dependence collected from the collection area at the single spatial location by the detection system and corresponding to the plurality of structured illuminations, wherein the collected optical signals from which optical data is reconstructed includes only one optical signal per each of the plurality of structured illuminations. 2. The apparatus of claim 1 wherein the detection system further comprises a tunable spectrograph optically coupled between the optical relay device and the CCD. 3. The apparatus of claim 1 wherein the optical relay device is a lens relay system. 4. The apparatus of claim 1 wherein the optical relay device is an optical fiber. 5. A method of determining surface or subsurface optical properties or structures of a sample of turbid media over an area of the sample comprising: exposing an area of the sample to a plurality of structured illuminations each having one or more spatial features, wherein the one or more spatial features of individual ones of the plurality of structured illuminations differing from the one or more spatial features of other individual ones of the plurality of structured illuminations, and wherein at least one of the plurality of structured illuminations having a plurality of spatial features comprising one or more peaks and one or more valleys;collecting an optical signal with no spatial dependence emitted from the sample per each of the plurality of structured illuminations by collecting optical signals with a detection sy... | MODULATED IMAGING INC. | A61B006/00 A61B005/00
US08989837 | B2 | 2015-03-24 | Methods and systems for determining fluid content of tissue | 2010-04-14 | Embodiments of the subject application include a diagnostic apparatus including one or more antennas disposed on a thorax of a living body to direct radio frequency (RF) electromagnetic waves through tissue and output signals responsively to the waves that have passed through the tissue. The apparatus may also include processing circuitry configured to process the signals over time so as to measure one or more RF path characteristics of the RF electromagnetic waves. The RF path length may be defined by a length of time required for the RF waves to pass through the thorax to an antenna and/or pass to tissue and reflect therefrom to an antenna, based on the path characteristic to assess a fluid content of the tissue. | 1. A diagnostic apparatus, comprising: one or more antennas configured to be disposed on a thorax of a body of a patient so as to direct radio frequency (RF) electromagnetic waves through a lung in the body and to output signals responsively to the RF electromagnetic waves that have passed through the lung; andprocessing circuitry is configured to: process the signals to determine an effective RF path length of the RF electromagnetic waves,wherein the effective RF path length is defined by the length of time required for the RF electromagnetic waves to pass through at least one of the thorax to an antenna and pass through to the lung and reflect therefrom to an antenna,compare the effective RF path length to one or more stored RF path length values corresponding to identified levels of fluid content of lungs, anddetermine fluid content of the lung based on the comparison. 2. The apparatus according to claim 1, wherein the processing circuitry is further configured to measure a change in the effective RF path length over one or more respiratory cycles of the lung, and to assess the fluid content responsively to the change. 3. The apparatus according to claim 1, wherein the processing circuitry is further configured to receive a measure of a physical distance traversed by the RF electromagnetic waves through the thorax and compare the effective RF path length to the physical distance. 4. The apparatus according to claim 1, wherein the one or more antennas comprise a transmitting antenna at a first location on a first side of the thorax, which transmits the RF electromagnetic waves through the lung, and a receiving antenna, which receives the RF electromagnetic waves that have passed through the lung at a second location on a second side of the thorax, opposite the first side, and wherein the physical distance is measured between the first and second locations. 5. The apparatus according to claim 1, wherein the one or more antennas comprise at least one antenna that is configured to direct the RF electromagnetic waves through the lung toward a heart in the body, and to output the signals responsively to the RF electromagnetic waves reflected from the heart. 6. The apparatus according to claim 5, further comprising an ultrasonic transducer, which is adjacent to the at least one antenna and is configured to direct ultrasonic waves toward the heart and receive the ultrasonic waves reflected from the heart so as to provide a measure of the physical distance. 7. The apparatus according to claim 1, wherein the stored values comprise RF path lengths for at least one of healthy lungs, edematous lungs, and previous effective RF path length measurements made on the patient. 8. A method for diagnosis, comprising: directing radio frequency (RF) electromagnetic waves from one or more antennas disposed on a thorax of a body of a patient so that the RF electromagnetic waves pass through or reflect off a lung in the body;outputting signals responsively to th... | Kyma Medical Technologies Ltd. | A61B005/05 A61B005/11 A61B005/0402 A61B005/00 A61N001/365
US09119554 | B2 | 2015-09-01 | Pharma-informatics system | 2010-11-18 | Compositions, systems and methods that allow for the detection of the actual physical delivery of a pharmaceutical agent to a body are provided. Embodiments of the compositions include an identifier and an active agent. The invention finds use in a variety of different applications, including but not limited to, monitoring of therapeutic regimen compliance, tracking the history of pharmaceutical agents, etc. | 1. An ingestible device comprising: a support;circuitry associated with the support;a first material electrically coupled to the circuitry and associated with the support; anda second material electrically coupled to the circuitry and associated with the support, wherein the first material is electrically isolated from the second material, wherein the first and second materials are selected to provide a voltage potential difference when in contact with an electrically conductive fluid, and wherein the circuitry is configured to modulate current flow through the electrically conductive fluid to generate an encoded monophasic signal which is conductively transmitted to a remote receiver; andwherein the circuitry comprises a switch; andwherein the circuitry is configured such that when the switch is closed, a current is produced between the first material and the second material. 2. The device according to claim 1, wherein the support comprises a semiconductor. 3. The device according to claim 1, wherein at least one of the first material and the second material are configured to be consumed and dissolved by the electrically conductive fluid. 4. The device according to claim 3, wherein the first material comprises a metal is selected from the group consisting of magnesium, zinc, sodium, lithium and iron. 5. The device according to claim 4, wherein the first material is doped with a non-metal. 6. The device according to claim 1, wherein the second material is a salt. 7. The device according to claim 6, wherein the salt is selected from the group consisting of copper salts and iron salts. 8. The device according to claim 1, wherein the first material is a metal and the second material is a salt. 9. The device according to claim 1, wherein the device is smaller than 10 mm3. 10. The device according to claim 9, wherein the device is smaller than 1 mm3. 11. The device according to claim 1, wherein the first and second materials are electrode layers positioned on opposing sides of the support. 12. An ingestible composition comprising: (a) a device comprising: (i) a support;(ii) circuitry associated with the support;(iii) a first material electrically coupled to the circuitry; and(iv) a second material electrically coupled to the circuitry, wherein the first material is electrically isolated from the second material, wherein the first and second materials are selected to provide a voltage potential difference when in contact with an electrically conductive fluid, and wherein the circuitry is configured to modulate current flow through the electrically conductive fluid to generate an encoded monophasic signal which is conductively transmitted to a remote receiver; andwherein the circuitry comprises a switch; andwherein the circuitry is configured such that when the switch is closed, a current is produced between the first material and the second material; and(b) a pharmaceutically acceptable carrier. 13. The composition according to claim 12, wh... | Proteus Digital Health, Inc. | H04Q005/22 A61B005/07 A61J003/00 H01Q001/27 G06K007/10 A61B005/00
US08900157 | B2 | 2014-12-02 | Electronic sphygmomanometer and blood pressure measurement control method | 2009-08-14 | A blood pressure measurement control method for controlling the measurement of a blood pressure according to an oscillometric method has the steps of determining whether to employ a depressurization measurement method or a pressurization measurement method to measure the blood pressure according to a person to be measured, performing a measurement process of the blood pressure in a depressurization process when determined to measure the blood pressure with the depressurization measurement method, performing a measurement process of the blood pressure in a pressurization process when determined to measure the blood pressure with the pressurization measurement method, and outputting the measured blood pressure value. | 1. An electronic sphygmomanometer for measuring a blood pressure according to an oscillometric method, the electronic sphygmomanometer comprising: a cuff to be wrapped around a predetermined measurement site;an adjustment unit for adjusting a pressure in the cuff by pressurization and depressurization;a pressure detection unit for detecting a cuff pressure representing the pressure in the cuff; anda control unit for performing a control for measuring the blood pressure;wherein the control unit includes:a determination unit for selecting a depressurization measurement method or a pressurization measurement method to measure the blood pressure according to a person to be measured,a first measurement processing unit for measuring blood pressure using a depressurization process when the determination unit selects the depressurization measurement method, anda second measurement processing unit for measuring blood pressure using a pressurization process when the determination unit selects the pressurization measurement method,wherein during the pressurization process, the blood pressure is determined prior to any cuff depressurization,wherein the electronic sphygmomanometer further comprises an input accepting unit for accepting input information from a user,wherein the input information indicates whether an arm of the person to be measured is thick,wherein the determination unit selects the depressurization or pressurization measurement method based on the input information,wherein the input accepting unit accepts information on whether or not an arm of the person to be measured is thick as the input information, andwherein the determination unit selects the pressurization measurement method if the input information indicates that the arm of the person to be measured is thick. 2. The electronic sphygmomanometer according to claim 1, further comprising: a display unit for displaying the measurement result of the blood pressure,wherein the display unit further displays information for specifying the method selected for the measurement of the depressurization measurement method and the pressurization measurement method. 3. The electronic sphygmomanometer according to claim 1, wherein the first measurement processing unit transitions the performing of the measurement process to a depressurization control at slow speed when the systolic blood pressure is estimated in the pressurization process at a speed faster than the depressurization process from previous measurements. 4. The electronic sphygmomanometer according to claim 1, wherein the second measurement processing unit rapidly depressurizes when the systolic blood pressure is detected in the pressurization process at slow speed. 5. A method for measuring a blood pressure according to an oscillometric method using an electronic sphygmomanometer, comprising: a wrapping step of wrapping a cuff of the electronic sphygmomanometer around a predetermined measurement site;an adjustment step of adjusting v... | OMRON HEALTHCARE Co., Ltd. | A61B005/02 A61B005/0225 A61B005/021 A61B005/022
US08271072 | B2 | 2012-09-18 | Detecting worsening heart failure | 2010-04-27 | A method comprises monitoring a heart rate, a respiration rate and an activity level of a patient, comparing the monitored heart rate, respiration rate and activity level to a predetermined threshold zone which is a function of heart rate, respiration rate and activity level, determining the patient is experiencing worsening heart failure when the monitored heart rate, respiration rate and activity level are outside the predetermined threshold zone; and after determining the patient is experiencing worsening heart failure when the monitored heart rate, respiration rate and activity level are outside the predetermined threshold zone, issuing an alert to indicate that the patient is experiencing worsening heart failure. | 1. A method comprising: monitoring a heart rate, a respiration rate and an activity level of a patient;comparing the monitored heart rate, respiration rate and activity level to a predetermined threshold zone which is a function of heart rate, respiration rate and activity level;determining the patient is experiencing worsening heart failure when the monitored heart rate, respiration rate and activity level are outside the predetermined threshold zone; andafter determining the patient is experiencing worsening heart failure when the monitored heart rate, respiration rate and activity level are outside the predetermined threshold zone, issuing an alert to indicate that the patient is experiencing worsening heart failure. 2. The method of claim 1, further comprising: measuring baseline heart rates and baseline respiration rates at a multitude of activity levels of the patient; andcalculating the predetermined threshold zone as a function of the measured baseline heart rates and baseline respiration rates for the multitude of activity levels. 3. The method of claim 2, wherein calculating the predetermined threshold zone comprises: recording a series of data points corresponding to short term average monitored heart rates, respiration rates and activity levels of the patient; anddefining the predetermined threshold zone according to the series of data points, wherein the predetermined threshold zone includes a majority of the data points in the series of data points. 4. The method of claim 3, wherein the predetermined threshold zone includes a substantially all of the data points in the series of data points. 5. The method of claim 3, wherein the measuring baseline heart rates and baseline respiration rates comprises recording the series of data points during a plurality of activity levels. 6. The method of claim 5, wherein the plurality of activity levels includes a sleeping activity level and an activity level consistent with exercise. 7. The method of claim 2, wherein the method of claim 2 is performed by an implantable medical device implanted within the patient, wherein the implantable medical device includes: one or more accelerometers that outputs acceleration information indicative of the activity level of the patient;one or more electrodes that outputs information indicative of the heart rate of the patient and the respiration rate of the patient; anda diagnostic unit configured to: compare the heart rate, respiration rate and activity level to the predetermined threshold zone, anddetermine the patient is experiencing worsening heart failure when the monitored heart rate, respiration rate and activity level are outside the predetermined threshold zone. 8. The method of claim 1, wherein the predetermined threshold zone is calculated specifically for the patient. 9. The method of claim 1, wherein the method of claim 1 is performed by an implantable medical device implanted within the patient. 10. A medical system for monitoring a cond... | Medtronic, Inc. | A61B005/0205
US07970461 | B2 | 2011-06-28 | Method and apparatus for determining conditions of a biological tissue | 2005-06-20 | In one method, one or more excitation signals with the same or different frequencies are applied to a biological object such as a tissue, simultaneously or consequently. Response signals are then cross-correlated with delayed excitation signals. Cross-correlation products are then auto-correlated. Cross-correlation products correspond to conditions of the tissue and auto-correlation product corresponds to changes in the conditions. Measuring electrical characteristics at low, intermediate and high frequency is also disclosed. At low frequency, the current flows mostly through the extracellular liquid of tissue. At high frequency, the current passes through the cell membranes freely enough to dominate the overall impedance. At both frequencies, the delay is less than 1/30 of the period of the respective signal. The intermediate frequency between the low frequency and the high frequency carries information about quick changes in the condition of the tissue. The delay in one example is from about 1/30 to ¼ of the period of the intermediate signal. | 1. A method for determining conditions of a bio-object, comprising a biological tissue by measuring electrical characteristics of the bio-object, the method comprising: generating a first excitation signal and generating a first reference signal, whereas said first reference signal is delayed compared to said first excitation signal by a first delay time and said first excitation signal having a periodic waveform with a first predetermined frequency;introducing said first excitation signal into the bio-object and receiving a first response signal from the bio-object;cross-correlating said first response signal and said first reference signal, thereby producing a first cross-correlation product, said first cross-correlation product being a real number; andcalculating a first set of indices characterizing conditions of the bio-object, based on said first cross-correlation product and said first delay time, wherein said first delay time is selected so as to maximize said first cross-correlation product for said first predetermined frequency. 2. A method as in claim 1, wherein said first excitation signal is a sine wave and said first predetermined frequency is selected low enough so that a current caused by said first excitation signal flows mostly through an extracellular liquid of the biological tissue and said first cross-correlation product corresponding to said first excitation signal reflects the amount of extracellular liquid in the biological tissue. 3. A method according to claim 2, wherein the first predetermined frequency is between about 10 to about 1000 Hz and is chosen so that the maximal value of said cross-correlation product is obtained at said first delay time being less than about 1/30 of the period of said first excitation signal. 4. A method as in claim 1, wherein said first excitation signal is a sine wave and said first predetermined frequency that is selected so that a current caused by said first excitation signal is capable to flow through cell membranes of the biological tissue so that said first cross-correlation product corresponding to said first excitation signal reflects the total amount of liquid in intracellular and extracellular region of the biological tissue. 5. A method according to claim 4, wherein the first predetermined frequency is between about approximately 10 kHz to about approximately 10 MHz and is chosen so that the maximal value of said first cross-correlation product is obtained at said first delay time being less than about approximately 1/30 of the period of said first excitation signal. 6. A method as in claim 1, additionally comprising generating a second excitation signal and a second reference signal, wherein said second reference signal is delayed compared to said second excitation signal by a second delay time, and introducing said second excitation signal into the biological tissue and receiving a second response signal from the biological tissue, cross-correlating said second response s... | - | A61B005/05
US07970194 | B2 | 2011-06-28 | Image processing apparatus, magnetic resonance imaging apparatus and image processing method | 2007-05-17 | An image processing apparatus includes a storage unit, a specifying unit, a calculation unit and a display unit. The storage unit stores diffusion weighted image data. The specifying unit specifies a calculation target region on the diffusion weighted image data. The calculation unit calculates at least one of a diffusion coefficient and a fractional anisotropy serving an index of diffusion anisotropy with regard to the calculation target region based on the diffusion weighted image data. The display unit displays at least one of the diffusion coefficient and the fractional anisotropy calculated by the calculation unit. | 1. An image processing apparatus comprising: a storage unit configured to store diffusion weighted image data;a specifying unit configured to specify a calculation target region encompassing only a subset of the diffusion weighted image data;a calculation unit configured to calculate at least one of diffusion coefficient and fractional anisotropy data serving as an index of diffusion anisotropy within the calculation target region based on the diffusion weighted image data; anda display unit configured to display at least one of the diffusion coefficient and the fractional anisotropy data calculated by said calculation unit. 2. An image processing apparatus according to claim 1, wherein said calculation unit is configured to perform projection processing with regard to the calculation target region, andsaid display unit is configured to display an image subjected to the projection processing as either of an apparent diffusion coefficient image and a fractional anisotropy image. 3. An image processing apparatus according to claim 1, wherein said specifying unit is configured to specify a range of either of pieces of diffusion weighted image data and diffusion weighted volume data beyond a threshold as the calculation target region. 4. An image processing apparatus according to claim 1, wherein said specifying unit is configured to specify a range beyond a threshold of a projected image obtained by projection processing of diffusion weighted volume data as the calculation target region. 5. An image processing apparatus according to claim 1, wherein said specifying unit is configured to specify a region having possibility of cancer as the calculation target region. 6. An image processing apparatus according to claim 1, wherein said storage unit is configured to store plural types of diffusion weighted image data of which at least one of corresponding b values each meaning a gradient magnetic field factor and corresponding directions of applying diffusion gradient magnetic field pulses are mutually different, andsaid calculation unit is configured to calculate at least one of the diffusion coefficient and the fractional anisotropy data based on the plural types of diffusion weighted image data. 7. An image processing apparatus according to claim 1, wherein said calculation unit is configured to generate information regarding to an organ based on at least one of transverse relaxation time weighted image data of which corresponding b value meaning a gradient magnetic field factor is able to be assumed to be substantially zero and image data acquired by imaging different from diffusion weighted imaging, andsaid display unit is configured to combine the information regarding to the organ with at least one of the diffusion coefficient and the fractional anisotropy data for displaying. 8. An image processing apparatus according to claim 1, further comprising a designating unit for marking to one of a diffusion weighted image, an apparent diffusion c... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | G06K009/00 G01V003/00 A61B005/05
US09603549 | B2 | 2017-03-28 | Implantable device with post-MRI cardiac signal sensing adjustment | 2010-11-01 | An implantable device, such as a pacer, defibrillator, or other cardiac rhythm management device, can include an automatic testing and adjusting of cardiac signal sensing after an MRI scanning procedure, such as to accommodate an MRI-induced physiological change in sensed cardiac depolarization amplitude or in lead impedance such as at an electrode/tissue interface. | 1. An apparatus comprising: an ambulatory or implantable medical device comprising: a cardiac signal sensing circuit; anda controller circuit, coupled to the cardiac signal sensing circuit, the controller configured to receive an indication that a magnetic resonance imaging (MRI) scanning procedure has completed and, in response, to automatically test a cardiac signal sensing of the cardiac signal sensing circuit, and to automatically adjust and then automatically use the adjusted cardiac signal sensing parameter when a physiological change in depolarization amplitude or lead impedance causes a post-MRI scanning procedure loss of cardiac depolarization sensing. 2. The apparatus of claim 1, wherein the implantable medical device comprises an MR-mode, and wherein the indication that the MRI scanning procedure has completed comprises a transition from the MR-mode. 3. The apparatus of claim 1, wherein the implantable medical device comprises an MRI detector circuit, coupled to the controller circuit, the MRI detector circuit configured to detect whether an MRI scanner device is present near the implantable medical device, and wherein the MRI detector circuit is configured to use information about whether the MRI scanner is present near the implantable medical device to provide the indication that the MRI scanning procedure has completed. 4. The apparatus of claim 1, wherein the indication that the MRI scanning procedure has completed is based on information about a detected fault condition associated with an MRI scanning procedure. 5. The apparatus of claim 1, comprising a communication module, coupled to the controller circuit, wherein the communication module is configured to communicate from the implantable medical device a report indicative of a cardiac signal sensing parameter that was automatically altered in response to post-MRI scanning procedure testing of the cardiac signal sensing. 6. The apparatus of claim 1, comprising a communication module, coupled to the controller circuit, wherein the communication module is configured to communicate from the implantable medical device an alert indicative of loss of cardiac signal sensing after the MRI scanning procedure. 7. The apparatus of claim 1, wherein the controller is configured to repeatedly automatically test the cardiac signal sensing over a period of time associated with post-MRI scanning procedure to test physiological variability of a cardiac signal amplitude or lead impedance. 8. An apparatus comprising: an ambulatory or implantable medical device comprising: a cardiac signal sensing circuit; anda controller circuit, coupled to the cardiac signal sensing circuit, the controller configured to receive an indication that a magnetic resonance imaging (MRI) scanning procedure has completed and, in response, to automatically test a cardiac signal sensing of the cardiac signal sensing circuit, and to automatically adjust and then automatically use the adjusted cardiac signal sensing pa... | Cardiac Pacemakers, Inc. | A61B005/0402 A61B005/06 A61N001/37 A61N001/08
US20110152647 | A1 | 2011-06-23 | DIODE LASER DEVICE FOR THE NON-INVASIVE MEASUREMENT OF GLYCAEMIA | 2011-01-27 | A device has two diode laser sources in a range between 500 and 1000 nm, with a power in the range between 0.01 to 100 mW. Rays from the sources are conveyed through an optical condenser and, starting up a key, are emitted either on a nail or on the skin or even on a free blood sample. A photodiodic sensor or a CPU read the energy subtracted by glycate haemoglobin and free plasmic glucose. This value is converted into the immediate glycaemia value and appears on a device display. | 1. A device to determine the glycaemia value without pricking the skin and without taking a blood sample, comprising: two laser diode sources for emitting respective laser rays, said two laser sources being operating in a frequency range between 500 and 1000 nm, with a power in the range between 0.01 to 100 mW;a photodiode sensor for receiving the laser rays emitted by the two laser diodes;a four-partitioned diaphragm for housing a finger of a patient; anda display,wherein the use and working of said device include the following steps:the patient puts a finger into the device, while operating a lever that opens the four-partitioned diaphragm;upon releasing the lever, a spring lets the diaphragm close, thus determining the centred arrangement of the finger, do no matter its dimensions, on the focus of the two laser rays emitted by the two laser diodes and directed towards the photodiodes of said sensor;the patient switches the device on, whereby the passing of the finger sets off an optical barrier that controls the starting of the two laser diodes;the photodiodes of the sensor read the two values of the laser radiation resulting from the selective absorption of the laser rays, emitted by the two laser diodes, by the glycate haemoglobin and the free plasmic glucose present in the blood of the patient; andan analogic stage forwards said values to a processor that, after their stabilization, digitises, normalises and extracts a high number of them, strikes an average, and, through a particular algorithm, obtains the instant value of glycaemia, that then appears on the device display,wherein in the selective absorption of the rays, emitted by said two laser diodes, by the glycate haemoglobin and the free plasmic glucose, the glycate haemoglobin works as a control parameter, so as to obtain absolutely accurate glycaemia values, andwherein said two values of the laser radiation resulting from the selective absorption of glycate haemoglobin and free plasmic glucose are calculated in connection with each other, whereby the measurement of the glycaemia is absolutely accurate notwithstanding the thickness and the colour of the skin and the thickness of the patient's finger. 2. The device according to claim 1, wherein it only needs a few seconds to determine the instantaneous value of glycaemia; in an absolutely painless and non-invading method, whereby it can be used to get several glycaemia measurements, even in a very short time range, to monitor the trend through time, to check the effectiveness of a therapy, to build glycaemic curves of newborns, children, adults and elderly people,wherein the values are stored in sequence up to 5000 assessments and the memory can be displayed through two slide keys, andwherein the device has a USB gate, a PC connection cable and is provided with a software to visualize and process the data. 3. The device according to claim 2, wherein it is easily activated by pressing the switch-on button, and wherein it is enough t... | PIGNOLO S.P.A. | A61B005/1455
US20110118587 | A1 | 2011-05-19 | OPTIMIZATION OF RF TRANSMIT AND GRADIENT MAGNETIC FIELD IMAGING USING RADIO FREQUENCY AND GRADIENT COILS | 2009-03-04 | The invention provides various systems, machine readable programs and methods for performing imaging using a MR scanner. The MR scanner includes at least one local radio-frequency transmit coil and at least one local gradient coil. The local radio-frequency transmit coil(s) and local gradient coil(s) cooperate to define an imaging volume. The MR scanner further includes a control system for performing an imaging operation on a patient's anatomy disposed within the imaging volume. The control system permits a selective simultaneous increase of the gradient magnetic field strength and peak B1 field strength by substantially the same factor, f. | 1. A system for performing imaging comprising a MR scanner, the MR scanner including: a) at least one local radio-frequency transmit coil;b) at least one local gradient coil, wherein the at least one local radio-frequency transmit coil and at least one local gradient coil cooperate to define an imaging volume;c) a control system for performing an imaging operation on a patient's anatomy disposed within the imaging volume, wherein the control system permits a selective simultaneous increase of the gradient magnetic field strength and peak B1 field strength by substantially the same factor, f. 2. The system of claim 1, wherein the control system reduces the duration of the B1 transmit pulse by the factor, f when the control system increases the gradient magnetic field strength and peak B1 field strength by the factor, f. 3. The system of claim 1, wherein the frequency response and the spatial response of the excited spins of the region being imaged is substantially preserved as a result of simultaneously increasing the gradient magnetic field strength and peak B1 field strength in substantially the same proportion. 4. The system of claim 1, wherein the flip angle resulting from the transmit pulse is substantially unchanged when increasing the gradient and peak B1 field strength. 5. The system of claim 1, wherein the excitation properties of the combined RF transmit and gradient pulse are preserved when increasing the gradient and peak B1 field strength. 6. The system of claim 1, wherein the control system is adapted and configured to increase the gradient strength in excess of 30 mT/m. 7. The system of claim 1, wherein the control system is adapted and configured to increase the gradient strength at a rate in excess of 100 T/m/sec. 8. The system of claim 1, wherein the control system is adapted and configured to create an estimate of a physiological limit of a region to be imaged. 9. The system of claim 8, wherein the estimate is based at least in part on a predetermined value. 10. The system of claim 9, wherein the estimate is based at least in part on the weight of a patient. 11. The system of claim 10, wherein the estimate is based at least in part on a predetermined value established for the particular anatomy being imaged. 12. The system of claim 8, wherein an estimate of the mass to be imaged is created based on one or more factors selected from the group consisting of (i) the weight of the patient, (ii) the anatomy being imaged, and (iii) the size of the imaging coil. 13. The system of claim 8, wherein the factor f is maximized for a given physiological limit. 14. The system of claim 8, wherein the physiological limit relates to SAR arising from application of the B1 field. 15. The system of claim 1, wherein the B1 field is maximized to not exceed a limit of 20 Watts/kg. 16. The system of claim 8, wherein the physiological limit relates to peripheral nerve stimulation arising from electric fields induced in the patient's anato... | - | A61B005/055
US20110105884 | A1 | 2011-05-05 | MRI INVOLVING DYNAMIC PROFILE SHARING SUCH AS KEYHOLE AND MOTION CORRECTION | 2008-08-20 | The invention relates to a device and to a method for magnetic resonance imaging (MRI) of a body. It is an object of the invention to provide a technique that enables dynamic profile sharing with significantly reduced motion artifacts. The method of the (such as keyhole) invention comprises the following steps: a) acquiring an MR data set (21) from an incomplete first part of k-space (C) by subjecting the body to an imaging sequence of RF pulses and switched magnetic field gradients; b) reconstructing an incomplete MR image (31) from the MR data set (21) acquired in step a) and deriving image transformation parameters describing motion of the body from the reconstructed incomplete MR image (31), -c) acquiring an MR data set (24) from an incomplete second part of k-space (P), which second part (P) is different from the first part (C) sampled in step a); d) applying a motion correction (40) to at least one of the MR data sets (21, 24) acquired in steps a) and c) according to the image transformation parameters derived in step b); e) reconstructing a final MR image (81) from a combination (71) of the motion-corrected MR data sets (61, 64). | 1. Device for magnetic resonance imaging of a body (7) placed in an examination volume, the device (1) comprising: means (2) for establishing a substantially homogeneous main magnetic field in the examination volume;means (3, 4, 5) for generating switched magnetic field gradients superimposed upon the main magnetic field;means (6) for radiating RF pulses towards the body (7;control means (12) for controlling the generation of the magnetic field gradients and the RF pulses;means (10) for receiving and sampling magnetic resonance signals; andreconstruction means (14) for forming MR images from the signal samples; the device (1) being arranged to:a) acquire an MR data set from an incomplete first part of k-space by subjecting the body (7) to an imaging sequence of RF pulses and switched magnetic field gradients;b) reconstruct an incomplete MR image from the MR data set acquired in step a) and derive image transformation parameters describing motion of the body (7) from the reconstructed incomplete MR image;c) acquire an MR data set from an incomplete second part of k-space, which second part is different from the first part sampled in step a);d) apply a motion correction to at least one of the MR data sets acquired in steps a) and c) according to the image transformation parameters derived in step b);e) reconstruct a final MR image from a combination of the motion-corrected MR data sets. 2. Device of claim 1, wherein the incomplete first part of k-space sampled in step a) is a central part of k-space and wherein the incomplete second part of k-space sampled in step c) is a peripheral part of k-space. 3. Device of claim 1 or 2, wherein the device (1) is further arranged to: repeat steps a) and b) in order to acquire a plurality of MR data sets from the first part of k-space successively in time;reconstruct an incomplete MR image from each partially acquired MR data set immediately after its acquisition; andderive image transformation parameters from each reconstructed incomplete MR image. 4. Device of any one of claims 1-3, wherein the device (1) is further arranged to use the image transformation parameters derived from an incomplete MR image to apply a motion correction to the imaging parameters of the imaging sequence used for the acquisition of a subsequent MR data set. 5. Method for MR imaging of a body of a patient placed in an examination volume of an MR device, the method comprising the following steps: a) acquiring an MR data set (21) from an incomplete first part of k-space (C) by subjecting the body to an imaging sequence of RF pulses and switched magnetic field gradients;b) reconstructing an incomplete MR image (31) from the MR data set (21) acquired in step a) and deriving image transformation parameters describing motion of the body from the reconstructed incomplete MR image (31);c) acquiring an MR data set (24) from an incomplete second part of k-space (P), which second part (P) is different from the first part (C) sampled in step ... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/055 G06K009/00
US20110112417 | A1 | 2011-05-12 | METHOD AND APPARATUS FOR POST-PROCESSING OF EPISODES DETECTED BY A MEDICAL DEVICE | 2010-10-18 | A method and system for determining undersensing during post-processing of sensing data generated by a medical device that includes transmitting a plurality of stored sensing data generated by the medical device to an access device, the stored sensing data including sensed atrial events and sensed ventricular events. The access device determines, in response to the transmitted data, instances where the medical device identified a cardiac event being detected in response to the sensing data, and determines whether one of a predetermined number of undersensing criteria have been met in response to the transmitted data. | 1. A method of determining undersensing during post-processing of sensing data generated by and stored within an implantable medical device, comprising: receiving the sensing data from the implantable medical device with an external access device, wherein the sensing data includes sensed atrial events and sensed ventricular events;determining, with the external access device, in response to the received data, instances where the implantable medical device identified a cardiac event being detected in response to the sensing data; anddetermining, with the external access device, whether one of a predetermined number of undersensing criteria have been met in response to the received data, wherein determining whether one of a predetermined number of undersensing criteria have been met comprises calculating a median AA interval in a sliding window of a predetermined number of intervals, and comparing at least one interval within the sliding window to the median AA interval. 2. The method of claim 1, wherein determining whether one of the predetermined number of undersensing criteria have been met further comprises: determining whether the median AA interval is greater than a predetermined threshold. 3. The method of claim 2, wherein the predetermined threshold is 1300 milliseconds. 4. The method of claim 1, wherein determining whether one of the predetermined number of undersensing criteria have been met further comprises determining whether the median AA interval is greater than a first predetermined threshold, and a current atrial interval is greater than the median AA interval by a second predetermined threshold. 5. The method of claim 4, wherein the first predetermined threshold is 350 milliseconds and the second predetermined threshold is 1.7 times the median AA interval or more. 6. The method of claim 1, wherein determining whether one of the predetermined number of undersensing criteria have been met further comprises determining whether the median AA interval is within a predetermined range, and a current atrial interval is greater than the median AA interval by a predetermined threshold. 7. The method of claim 6, wherein the predetermined range comprises a range greater than 200 milliseconds and less than or equal to 359 milliseconds, and wherein the predetermined threshold is 1.9 times the median AA interval or more. 8. The method of claim 6, wherein the predetermined range comprises a range greater than 0 milliseconds and less than or equal to 200 milliseconds, and wherein the predetermined threshold is 3.25 times the median AA interval or more. 9. A system for post-processing of sensing data associated with identification of a cardiac event, comprising: an implantable medical device generating and storing a plurality of sensing data, wherein the sensing data includes sensed atrial events and sensed ventricular events;an access device located externally from the medical device; andan interface for receiving the sensing data from the... | Medtronic, Inc. | A61B005/0452
US20110144541 | A1 | 2011-06-16 | BIOMETRIC DATA-MEASURING INSTRUMENT, BIOMETRIC DATA-MEASURING SYSTEM, MUSCLE STRENGTH METER, AND MUSCLE STRENGTH-MEASURING SYSTEM | 2010-12-09 | A biometric data-measuring instrument that measures data relating to a living organism by applying pressure to the living organism. The instrument includes: a casing, an auxiliary contacting part that extends from the casing, the auxiliary contacting part being contacted against a vicinity of a point to be measured on the living organism and applying pressure to the vicinity of the point to be measured, a main contacting part that, in a state where the auxiliary contacting part is applying pressure to the vicinity of the point to be measured, is contacted against the point to be measured and applies pressure to the point to be measured in the direction in which the auxiliary contacting part is applying pressure to the vicinity of the point to be measured, a pressure sensor that is provided inside the casing and measures a pressure that the main contacting part receives from the point to be measured, and a biometric data display unit that is provided on the casing and displays the measured biometric data. A tip of the auxiliary contacting part extends outward from a base side thereof. | 1. A biometric data-measuring instrument that measures data relating to a living organism by applying pressure to said living organism, comprising: a casing;an auxiliary contacting part that extends from said casing, the auxiliary contacting part being contacted against a vicinity of a point to be measured on said living organism and applying pressure to the vicinity of said point to be measured;a main contacting part that, in a state where said auxiliary contacting part is applying pressure to the vicinity of said point to be measured, is contacted against said point to be measured and applies pressure to said point to be measured in the direction in which said auxiliary contacting part is applying pressure to the vicinity of said point to be measured;a pressure sensor that is provided inside said casing and measures a pressure that said main contacting part receives from said point to be measured;a biometric data display unit that is provided on said casing and displays said measured biometric data;a tip of said auxiliary contacting part extending outward from a base side thereof. 2. The biometric data-measuring instrument according to claim 1, wherein said biometric data display unit is facing in the pressure-receiving direction in which said auxiliary contacting part receives pressure from said living organism. 3. The biometric data-measuring instrument according to claim 1, wherein said casing has a grasp part extending in a direction intersecting said pressure-applying direction, and said biometric data display unit is facing in an opposite direction to a direction in which the grasp part extends. 4. The biometric data-measuring instrument according to claim 1, wherein said auxiliary contacting part comprises: a detachable extending part that is detachably provided at a tip of said auxiliary contacting part; anda locking mechanism at a tip of this detachable extending part. 5. The biometric data-measuring instrument according to claim 1, comprising: a securing mechanism that secures said auxiliary contacting part such that the tip of said main contacting part is positioned further to said pressure-applying direction side than the tip of said auxiliary contacting part; anda switch that, when switched on, makes said biometric data display unit display said measured biometric data. 6. A biometric data-measuring system comprising: the biometric data-measuring instrument according to claim 1 comprising a radio communication unit; anda printer that performs a radio communication with said radio communication unit to print said measured biometric data. 7. A muscle strength meter comprising: an attachment that contacts against a living organism; anda muscle strength meter main body that is attached to said attachment, detects pressure from said living organism via said attachment, and measures the muscle strength of said living organism;said attachment comprising:an contacting seat part with an extending contacting face that contacts against ... | ITO CO., LTD. | A61B005/103
US20110124984 | A1 | 2011-05-26 | Device for Aliquoting and Filtering Blood | 2009-03-24 | Device for aliquoting and filtering blood, comprising: an element (1, 10) for collecting blood (S); a capillary conduit (2) which has a capillary force greater than that of said collecting element and is in fluid communication therewith; and a dispensing element (3) which has a capillary force greater than that of the capillary conduit and is in fluid communication therewith; characterized in that said dispensing element comprises a filtering device (3) for separating blood cells from plasma. | 1. A device for aliquoting and filtering blood, the device comprising: a collector element for collecting blood;a capillary duct presenting a capillary force stronger than that of said collector element and in fluid flow communication therewith; anda dispenser element presenting a capillary force stronger than that of the capillary duct and in fluid communication therewith;the device being characterized in that the dispenser element includes a filter device for separating blood corpuscles from plasma. 2. A device according to claim 1, wherein said collector element is a well or depression. 3. A device according to claim 1, wherein said collector element includes an array of microneedles and a reservoir located beneath the array. 4. A device according to claim 1, wherein an adhesive zone is provided on a surface of said device in register with said collector element so as to provide leaktight contact between said collector element and a zone of a patient's skin. 5. A device according to claim 1, wherein said collector element and said capillary duct are adapted to contain a volume of blood lying in the range 10 μL to 1 mL, and preferably in the range 50 μL to 500 μL. 6. A device according to claim 1, including visible marks making it possible, merely by inspection, to estimate the volume of blood present in said capillary duct. 7. A device according to claim 1, wherein said collector element and/or said capillary duct contain at least one chemical substance selected from: anticoagulating agents for blood;agglutinating agents for blood corpuscles; andcoagulating agents for blood. 8. A device according to claim 7, wherein at least one of the inside walls of said capillary duct is coated in a hydrophilic gel that progressively releases a blood anticoagulating agent while the duct is being filled with blood for aliquoting and filtering. 9. A device according to claim 1, wherein said collector element and said capillary duct are integrated in a common element of plane structure. 10. A device according to claim 9, wherein said element of plane structure is transparent. 11. A device according to claim 9, wherein said element of plane structure is made by molding or machining a plastics material. 12. A device according to claim 1, wherein said collector element and said capillary duct are incorporated in a surface of a single element of plane structure, said filter device comprising a membrane placed on an opposite face of said element, the capillary duct and the filter device being put into fluid flow communication via a hole passing through the element. 13. A device according to claim 1, wherein said filter device is a paper filter. 14. A device according to claim 1, wherein said filter device is adapted to become saturated in corpuscles after filtering a volume of blood that depends on the hematocrit content of said blood and that is in any event less than the content of said collector element plus said capillary duct; and wherein visible ... | - | A61B005/157 B01L003/02
US08952333 | B2 | 2015-02-10 | Methods for improved single photon emission computed tomography using exact and stable region of interest reconstructions | 2010-11-02 | The present invention provides systems, methods, and devices for improved computed tomography (CT) and, more specifically, to methods for improved single photon computed tomography (SPECT) using exact and stable region of interest (ROI) reconstructions. This technology can be extended across all tomographic modalities. Embodiments provide a method and a system for reconstructing an image from projection data provided by a single photon emission computed tomography scanner comprising: identifying a region of interest in an object; defining an attenuation coefficient and object boundary; computing the generalized Hilbert transform of the data through the defined region of interest and a known subregion; and reconstructing the image with improved temporal resolution at lower radiation doses, wherein the reconstructing comprises performing a reconstruction method that yields an exact and stable reconstruction. Embodiments also provide a method and a system for reconstructing an image from projection data provided by a single photon emission computed tomography scanner comprising: identifying a region of interest in an object; defining an attenuation coefficient and object boundary; and reconstructing the images by minimizing the high order total variation while minimizing the data discrepancy. | 1. A theoretically exact local reconstruction method comprising reconstructing a region of interest (ROI) of an object into an image from local single photon emission computed tomography (SPECT) projection data of the ROI by modeling the projection data as an attenuated Radon transform comprising formula:Po⁡(θ,s)=∫-∞∞⁢f⁡(s⁢⁢θ+t⁢⁢θ⊥)⁢ⅇ-∫t∞⁢μ⁡(s⁢⁢θ+t~⁢⁢θ⊥)⁢ⅆt~⁢ⅆt,where Po represents the original projection data, f(x) is the image to be reconstructed, θ=(cos θ, sin (θ)), θ⊥=(−sin θ, cos (θ)), and μ(x) is the attenuation coefficient map. 2. The method of claim 1, further comprising: defining a compact support and the corresponding attenuation coefficient; and measuring the attenuated Radon transform of the projection data through the ROI. 3. The method of claim 2, further comprising: modeling the region of interest as piecewise polynomial; and reconstructing the image by performing a high order TV minimization while minimizing the data discrepancy. 4. The method of claim 3, wherein the high order TV is minimized by steepest descent searching method. 5. The method of claim 3, wherein the high order TV is minimized by soft-threshold filtering method. 6. The method of claim 2, further comprising: identifying a known subregion inside the region of interest; backprojecting the derivative data to a PI line passing through the known subregion; and reconstructing the image by inverting the generalized truncated Hilbert transformgw⁡(u)=-g⁡(u)2⁢⁢π=PV⁢∫csbcse⁢chμ0⁡(u-u~)⁢f⁡(u~)⁢ⅆu~,u∈(cvb,cve). 7. The method of claim 6, wherein the generalized truncated Hilbert transform is inverted by singular value decomposition method. 8. The method of claim 6, wherein the generalized truncated Hilbert transform is inverted by projection onto convex set method. 9. The method of claim 1, further comprising scanning an object using a SPECT scanner to acquire projection data relating to the object. 10. The method of claim 1, wherein the single photon emission computed tomography projection data are uniformly attenuated local projections. 11. The method of claim 1, wherein the single photon emission computed tomography projection data are generalized attenuated local projections. 12. The method of claim 1, wherein the attenuation coefficient μ(x) is known on the whole compact support. 13. The method of claim 1, wherein the attenuation coefficient μ(x) is reconstructed from global un-truncated transmission projections. 14. The method of claim 1, wherein the attenuation coefficient μ(x) is reconstructed from local truncated transmission projections passing through the same region of interest. 15. The method of claim 14, wherein the attenuation coefficient μ(x) is known inside a subregion of the region of interest. 16. The method of claim 14, wherein the attenuation coefficient μ(x) is piecewise polynomial inside the region of interest. 17. The method of claim 1, wherein the projection data Po(θ,s) is acquired with gamma camera. 18. The method of claim 1, wherein t... | Virginia Tech Intellectual Properties, Inc. | G06T011/00 A61B006/03
US08849458 | B2 | 2014-09-30 | Collection device with selective display of test results, method and computer program product thereof | 2010-06-18 | A collection device with a selective display of test results and method thereof are disclosed. A structured collection procedure defining data collection times and the associated context of the collection also defines what information regarding the results of the collection may be viewable by a user performing the structured collection procedure on the device. In this manner, the patient can be monitored according to the structured collection procedure while preventing the patient from modifying his or her behavior based on collection results. | 1. A portable, handheld collection device for the management of diabetes comprising: a display;a user interface;a measurement engine to measure a biomarker value;a non-transitory memory containing a data file and a structured collection procedure, the structured collection procedure having parameters defining: a schedule of collection events having one or more collection events which provide a request for a measurement of the biomarker value under defined conditions which are related to the biomarker value as contextual information associated with the biomarker value, anda view flag associated with each of the one or more collection events, the view flag having a viewable value and a non-viewable value;a processor connected to the display, the user interface, the measurement engine, and the memory; anda power supply for powering the device;wherein the processor is programmed to: read the structured collection procedure from the memory,permit an authorized user to select at least one criterion under which a data record of the data file for the structured collection procedure is designated as viewable,run the structured collection procedure,send automatically the request of each of the one or more collection events according to the structured collection procedure,store automatically biomarker value(s) measured by the measurement engine or information entered via the user interface in an associated data record of the data file in the memory in response to sent request,store and link automatically contextual information associated with each biomarker value or the information entered in the associated data record of the data file to form contextualized biomarker data,designate automatically the associated data record in the memory containing the contextualized biomarker data as viewable if the view flag associated with the sent request in the structured collection procedure has the viewable value or the at least one criterion is met and as not viewable if the view flag has the non-viewable value or the at least one criterion is not met, andprovide to the display only biomarker values associated to data records containing the contextualized biomarker data if the data records are designated in the memory as viewable. 2. The portable collection device according to claim 1, wherein the structured collection procedure further comprises entry criterion defining one or more conditions which when satisfied start the schedule of collection events, and exit criterion defining one or more conditions which when satisfied end the structured collection procedure, and wherein the software having instructions that when executed by the processor further causes the processor to run automatically the structured collection procedure when the entry criterion is satisfied, and end automatically the structured collection procedure when the exit criterion is satisfied. 3. The portable collection device according to claim 1, wherein the at least one viewable criterion is at... | Roche Diagnostics Operations, Inc. | G05B021/00 G06F019/00 G06Q050/22 A61B005/00 A61B005/145
US07959581 | B2 | 2011-06-14 | Test magazine and method for processing the same | 2006-10-02 | A test magazine with two film strips (12, 14), which are joined to one another in a sandwich-like manner, can be wound up, and between which holding cells (16) for test elements are kept free. The test magazine also comprises a multitude of test elements (18, 20) each having a puncturing unit (18) for inserting into body tissue and a test unit (20) for being subjected to body fluid. To this end, the puncturing units (18) and test units (20) are placed in separate holding cells (16) whereby separating them from one another. | 1. A test magazine comprising: two spoolable foils joined together in a layered manner and between which holding cells are formed for holding test elements;a plurality of test elements which each comprise a lancing unit for puncturing body tissue and a test unit to which body fluid can be appliedwherein the lancing units and test units are arranged separately from one another in separate holding cells;wherein the holding cells are separately sealed from one another;wherein the holding cells containing the test units do not contain the lancing units; andwherein the lancing units have a capillary structure that is formed by a semi-open channel to collect body fluid. 2. The test magazine of claim 1, characterized in that the lancing units and test units are arranged alternately in a tape direction or adjacent to one another in pairs at right angles to the tape direction. 3. The test magazine of claim 1, characterized in that the lancing units in the holding cells were sterilized by irradiation through a mask in order to screen the test units. 4. The test magazine of claim 1, characterized in that the foils are joined to a test tape over their length and that cells on the test tape are exposed by perforations in the foils to form the test units. 5. The test magazine of claim 1, characterized in that a desiccant for the test units is stored in the holding cells. 6. A test magazine comprising: two spoolable foil tapes joined together in a layered manner and between which holding cells are formed for holding test elements;and a plurality of test elements arranged in the holding cells, wherein the test elements each comprise a lancing unit for puncturing body tissue and a test unit to which body fluid can be applied;wherein at least the lancing units can be removed by transport means from their respective holding cell and can be moved into a working position which is completely separate from the foil tapes;wherein the holding cells are separately sealed from one another;wherein the holding cells containing the test units do not contain the lancing units; andwherein the lancing units have a capillary structure that is formed by a semi-open channel to collect body fluid. 7. The test magazine of claim 6, further comprising: the transport means, wherein the transport means comprises a tape tensioning device to pull apart the foil tapes in different directions. 8. The test magazine of claim 7, characterized in that the tape tensioning device has two deflecting cylinders that can rotate in opposite directions or are stationary and are arranged laterally spaced apart from one another at a dispensing site, wherein the tape tensioning device has take-up spools for receiving the foil tape located downstream from the deflecting cylinders. 9. The test magazine of claim 6, further comprising: the transport means, wherein the transport means are in operative connection with a test unit. 10. The test magazine of claim 6, further comprising: the transport means... | Roche Diagnostics Operations, Inc. | A61B005/00 A61B017/14 A61B017/32 B65D081/00 G01N021/00 G01N031/00 G01N033/00 G01N035/00
US08332010 | B2 | 2012-12-11 | Method for non-contrast enhanced magnetic resonance angiography | 2011-02-14 | A method for non-contrast enhanced magnetic resonance angiography (“MRA”) that has a short scan time and is insensitive to patient motion is provided. More particularly, the method provides significant arterial conspicuity and substantial venous signal suppression. A two-dimensional single shot acquisition is employed and timed to occur a specific time period after the occurrence of an R-wave in a contemporaneously recorded electrocardiogram. In this manner, k-space data is acquired that is substantially insensitive to variations in arterial flow velocity, or heart rate, and that further substantially suppresses unwanted venous signal in a prescribed imaging slice. | 1. A method for non-contrast-enhanced magnetic resonance angiography to produce an image depicting vasculature of a subject with a magnetic resonance imaging (MRI) system, the steps comprising: a) acquiring a signal indicative of a cardiac phase of the subject;b) performing a pulse sequence that directs the MRI system to: i) apply a radio frequency (RF) saturation pulse to a prescribed imaging region such that MR signals therein are substantially suppressed;ii) select a time interval that begins at a time point at which the RF saturation pulse is applied in step b)i) and ends at a time point at which spins associated with unsaturated arterial blood has maximally flowed into the prescribed imaging region;iii) wait the time interval selected in step ii);iv) acquire k-space data from the subject as a single shot in coordination with an application of a series of RF excitation pulses, wherein the k-space data is acquired during a time period that is sufficiently short such that the k-space data is acquired while only arterial blood is flowing into the prescribed imaging region, thereby mitigating variations in flow velocity during acquisition of the k-space data; andc) reconstructing, from the acquired k-space data, an image depicting vasculature of the subject. 2. The method as recited in claim 1 in which the pulse sequence performed in step b) further directs the MRI system to: v) apply an RF saturation pulse to a prescribed slab that is outside of the imaging region, such that MR signals indicative of venous blood that flows into the imaging region are substantially suppressed. 3. The method as recited in claim 2 in which step b)v) is performed at least one of before step b)i) and after step b)i). 4. The method as recited in claim 2 in which step b)iv) is performed before the time interval specified in step b)iii). 5. The method as recited in claim 4 in which the pulse sequence performed in step b) further directs the MRI system to: vi) apply another RF saturation pulse to the prescribed slab that is outside of the imaging region, such that MR signals indicative of venous blood that flows into the imaging region are substantially suppressed;wherein step b)vi) is performed after the time interval specified in step b)iii) and before k-space data is acquired in step b)iv). 6. The method as recited in claim 2 in which step b)v) is performed after the time interval specified in step b)iii) and before k-space data is acquired in step b)iv). 7. The method as recited in claim 1 in which the pulse sequence performed in step b) further directs the MRI system to: v) acquire additional k-space data during the time interval specified in step b)iii);and further comprising:d) reconstructing another image from the additional k-space data acquired in step b)v). 8. The method as recited in claim 7 further comprising: e) producing a difference image by subtracting the image reconstructed in step c) and the another image reconstructed in step d). 9. The metho... | Ghost Medical Technologies, Ltd | A61B005/05
US08103342 | B2 | 2012-01-24 | Therapy-selection methods for implantable heart monitors | 2011-02-18 | An implantable heart-monitoring device comprising one or more leads for sensing electrical signals of a patient's heart a therapy circuit for delivering pharmaceuticals to the patient; and a monitoring circuit coupled to the one or more leads and to the therapy circuit, the monitoring circuit for monitoring heart activity of the patient through one or more of the leads, the monitoring circuit operable to: determine a minimum interval from a set of two or more intervals based on sensed heart activity, determine a range parameter based on a difference between a first subset and second subset of the intervals, determine a dispersion parameter quantifying dispersion of a subset of the intervals, and to determine an assessment value for the set of intervals, based on the minimum interval, the range parameter, the dispersion index. | 1. An implantable medical device comprising a processor-readable medium, the processor-readable medium comprising instructions that when performed by a processor, cause the implantable medical device to: obtain information indicative of cardiac electrical activity including at least a portion of a tachyarrhythmia;using the obtained information, determine intervals between successive cardiac electrical events comprising the cardiac electrical activity;determine a first indication of spread of durations of the intervals, using information about a first set of the intervals;determine a second indication of spread of durations of the intervals, using information about a different second set of the intervals; andusing at least the first and second indications of spread, classify the tachyarrhythmia. 2. The implantable medical device of claim 1, wherein the processor-readable medium comprises instructions that when performed by the processor, cause the implantable medical device to determine a minimum interval duration included in the intervals. 3. The implantable medical device of claim 2, wherein the processor-readable medium comprises instructions that when performed by the processor, cause the implantable medical device to: determine a value using information about the first indication of spread, the second indication of spread, and the minimum interval duration;compare the value to a specified threshold; andusing information about the comparison, classify the tachyarrhythmia. 4. The implantable medical device of claim 3, wherein the specified threshold is specified as a function of one or more of the first indication of spread, the second indication of spread, or the minimum interval duration. 5. The implantable medical device of claim 1, wherein the instructions to determine the first indication of spread of durations of the intervals, using information about the first set of the intervals, include instructions to discard, for the purpose of determining the first indication of spread, one or more of the shortest-duration or longest-duration intervals included in the intervals, to provide the first set of the intervals. 6. The implantable medical device of claim 5, wherein the instructions to determine the first indication of spread of durations of the intervals, using information about the first set of the intervals includes instructions to: determine a first central tendency of a subset of the shortest-duration intervals included the first set of the intervals;determine a second central tendency of a subset of the longest-duration intervals included in the first set of the intervals; anddetermine a relative indication of information about the first and second central tendencies to provide the first indication of spread. 7. The implantable medical device of claim 6, wherein the instructions to determine the first and second central tendencies include instructions to determine respective first and second average interval durations of the res... | Cardiac Pacemakers, Inc. | A61N001/00
US08744562 | B2 | 2014-06-03 | Method for monitoring or treating nervous system disorders | 2010-12-10 | A medical device system that includes a brain monitoring element, cardiac monitoring element and a processor. The processor is configured to receive a brain signal from the brain monitoring element and a cardiac signal from the cardiac monitoring element. The processor is further configured to compare the brain signal to the cardiac signal. A method of comparing a brain signal to a cardiac signal is also provided. | 1. A method for comparing a brain signal to a cardiac signal for use in detecting neurological events, the method comprising: receiving a brain signal with a medical device system comprising processing circuitry for a period of time;receiving a cardiac signal with the medical device system for a period of time;obtaining information with the processing circuitry identifying one or more neurological events in the brain signal;obtaining information with the processing circuitry identifying one or more cardiac events in the cardiac signal; andcomparing the brain signal and the cardiac signal with the processing circuitry to determine a temporal relationship between the one or more neurological events and the one or more cardiac events for determining whether neurological event detection is improved by using the cardiac signal, the comparing comprising categorizing with the processing circuitry each of the identified neurological events as cardiac matched when there is a cardiac event within a specified time period of the neurological event; andcomputing with the processing circuitry a rate of concordance between the neurological events and the cardiac events based on the number of cardiac matched events and the number of neurological events. 2. The method of claim 1, wherein computing the rate of concordance comprises determining with the processing circuitry the number of cardiac matched events divided by the number of neurological events. 3. The method of claim 1, wherein obtaining information identifying the one or more neurological events in the brain signal comprises applying with the processing circuitry a neurological event detection algorithm to the brain signal and wherein obtaining information identifying the one or more cardiac events in the cardiac signal comprises applying with the processing circuitry a cardiac event detection algorithm to the cardiac signal. 4. The method of claim 1, further comprising for each neurological event categorized as cardiac matched, determining with the processing circuitry the temporal order of the neurological event and the cardiac event that is within a specified time period of the neurological event. 5. The method of claim 4, further comprising categorizing with the processing circuitry each of the cardiac events as cardiac unmatched when there is no neurological event within a specified time period of the cardiac event. 6. The method of claim 5, wherein determining whether the neurological event detection is improved comprises considering the rate of concordance, the temporal order of the cardiac matched events, and the number of cardiac unmatched events. 7. The method of claim 4, further comprising activating with the processing circuitry a cardiac event detection algorithm for use in neurological event detection based on a determination that the neurological event detection is improved by using the cardiac signal. 8. The method of claim 7, wherein the neurological event to be detected by the c... | Medtronic, Inc. | A61B005/04
US08998841 | B2 | 2015-04-07 | Monitoring conditions of implantable medical fluid delivery device | 2010-12-07 | A programmer device includes an interface that communicates with an implantable fluid delivery device and a user interface that allows a user to troubleshoot a catheter connected to the fluid delivery device by measuring a pressure associated with the catheter. A processor may compare the measured pressure waveform to a previously-acquired waveform or the user interface may display the measured pressure waveform for the user to compare the displayed pressure waveform to a previously-acquired baseline waveform. The comparison between the two waveforms may be based on the pressure decay time. A historical log of measured pressure decays associated with the catheter may be maintained by the programmer device or the implantable fluid delivery device and displayed for the user to determine whether the performance of the catheter is deteriorating. | 1. A method of troubleshooting a catheter associated with an implantable medical fluid delivery device, the method comprising: delivering a fluid from the implantable medical fluid delivery device through the catheter to a patient;measuring a pressure associated with the catheter during the delivery of the fluid to the patient;defining a current pressure waveform based on the measured pressure; anddetermining a characteristic of the catheter based on a comparison of at least a first waveform portion of the current pressure waveform to at least a second waveform portion of a previously-acquired pressure waveform associated with the catheter during delivery of the fluid. 2. The method of claim 1, further comprising: obtaining a posture associated with the patient when the pressure associated with the catheter is measured; andassociating the determined characteristic of the catheter with the obtained posture. 3. The method of claim 1, wherein measuring the pressure comprises determining a time for the measured pressure to fall from a maximum pressure level to a signature pressure level, wherein the signature pressure level is a pressure when a pump stroke of the implantable medical fluid delivery device is not present. 4. The method of claim 1, further comprising identifying one or more characteristics indicative of one or more catheter malfunctions based on the determined characteristic of the catheter. 5. The method of claim 1, further comprising storing at least one of: the determined characteristic of the catheter or the current pressure waveform. 6. The method of claim 1, wherein the previously-acquired pressure waveform corresponds to one or more waveforms corresponding to one or more malfunctions associated with the catheter. 7. The method of claim 1, wherein the current pressure waveform comprises a pressure decay profile for a current pump stroke of the implantable medical fluid delivery device, and the previously-acquired pressure waveform comprises a pressure decay profile for a previous pump stroke of the implantable medical fluid delivery device. 8. The method of claim 1, wherein the first waveform portion of the current waveform comprises first variations in pressure over a first time, and wherein the second waveform portion of the previously-acquired waveform comprises second variations in pressure over a second time. 9. The method of claim 1, wherein determining the characteristic of the catheter comprises determining a patency of the catheter based on the comparison of at least the first waveform portion of the current pressure waveform to at least the second waveform portion of the previously-acquired pressure waveform associated with the catheter during delivery of the fluid. 10. The method of claim 1, wherein determining the characteristic of the catheter comprises calculating a pressure decay time based on a maximum pressure value associated with the measured pressure. 11. The method of claim 10, wherein the decay time... | Medtronic, Inc. | A61M031/00 A61M005/142 A61M005/168 G06F019/00
US08764656 | B2 | 2014-07-01 | Sensing device of emotion signal and method thereof | 2010-12-02 | The present invention relates to a sensing device of an emotion signal and a method thereof capable of recognizing and analyzing the change in emotion by collecting at least one of the bio signals and peripheral environmental signals, guiding the state of emotion to a user, and sharing emotion information with authorized persons and communicating it between them. The sensing device of the emotion signal according to the embodiment of the present invention includes a sensing signal processor that senses a plurality of bio signals and peripheral environment signals of a user; an emotion signal processor that generates emotion signals representing the emotional state of the user for each of the sensed bio signals and collects the emotion signals to generate an emotion index; and an emotion signal communication unit that optionally transmits the emotion signals and the emotion index to external authorized devices. | 1. An emotion sensing device, the device comprising: a sensing signal processor configured to sense a plurality of bio signals of a user and peripheral environment signals of a user;an emotion signal processor configured to generate an emotion signal associated with an emotional state of the user for each of the sensed bio signals, and collect the emotion signals to generate an emotion index; andan emotion signal communication unit configured to transmit any of the emotion signals and the emotion index to an external authorized device,wherein the emotion signal processor is configured to generate error recognition information for the emotion signals using the peripheral environment signals, and use the error recognition information to recognize an error for the emotion signals and perform error compensation. 2. The sensing device according to claim 1, wherein the emotion signal processor uses the emotion signals and the sensed peripheral environment signals to generate the emotion index. 3. The sensing device according to claim 1, wherein the sensing signal processor includes at least one of a PPG sensor, a GSR sensor, a temperature/humidity sensor, an acceleration sensor, a voice sensor, and an infrared sensor. 4. The sensing device according to claim 1, wherein the emotion signal processor is configured to use an emotion changing threshold value to identify an emotion change of the user, and generate an emotion signal corresponding to a sensed bio signal by using the emotion changing threshold value to extract the emotion signal from the bio signal. 5. The sensing device according to claim 1, wherein the emotion signal processor is configured to perform a multimodal convergence emotion inference using the emotion signals to determine the emotion index. 6. The sensing device according to claim 1, wherein the emotion signal communication unit is configured to perform a privacy security process before transmitting the emotion signals or the emotion index to the external authorized device. 7. The sensing device according to claim 1, wherein the sensing signal processor includes: a sensor unit including a plurality of sensors configured to sense the bio signals and the peripheral environment signals;an amplifier configured to amplify signals input from the plurality of sensors;a noise filtering unit configured to remove external noise from the amplified signals;an analog/digital I/F unit configured to convert the output signals from the noise filtering unit into digital signals; anda digital signal output unit configured to output the digital signals. 8. The sensing device according to claim 7, further comprising a hardware (H/W) controller configured to provide a platform for the sensing signal processor, the emotion signal processor, and the emotion signal communication unit, and wherein the H/W controller includes:a sensor analog interface unit configured to receive analog signals from the plurality of sensors;a main signal processor confi... | Electronics and Telecommunications Research Institute | A61B005/00 G06F003/01
US08870784 | B2 | 2014-10-28 | Precision of xenon content measurement in a ventilatory anesthesia apparatus | 2009-05-12 | The invention relates to an apparatus for the ventilatory anaesthesia of a patient by administration of a gas containing gaseous xenon that comprises a main gas circuit (CP) and a means for determining the xenon content (S6, M1) adapted for and capable of determining the gaseous xenon content in the main circuit (CP), and comprising at least one hot wire sensor (S6-E) including at least one electrically conducting wire (FC1) in direct contact with at least a portion of the xenon-containing gaseous flow. At least one hot wire sensor (S6-E) including at least one conducting wire (FC1) is arranged on a gas-supply main line (LP) including a bypass line (BP) fluidically connected to said gas-supply main line (LP) upstream and downstream from said at least one conducting wire (FC1) and taking into account the gas flow direction in said main line (LP), and at least a first electrovalve (EV1) is provided at the upstream connection of the bypass line (BP) to said main line (LP) in order to direct the xenon-containing gaseous flow either towards the main line (LP) on which said at least one conducting wire (FC1) is provided, or towards the bypass line (BP). | 1. An apparatus for ventilatory anesthesia of a patient by administration of a gas containing gaseous xenon, said apparatus comprising: a main gas circuit (CP) in the form of an open or closed circuit having an inhalation branch (16) for supplying a gaseous mixture containing xenon to the patient and an exhalation branch (18) for conveying the gaseous mixture containing xenon exhaled by the patient, andmeans (S6, M1) for determining the concentration of xenon, which are designed and able to determine the gaseous xenon content in at least part of the main circuit (CP), said means (S6, M1) for determining the concentration of xenon comprising a first hot-wire sensor (S6-E) having a first electrically conductive wire (FC1) in direct contact with at least part of the gaseous flow containing the xenon, wherein:the first hot-wire sensor (S6-E) having the first conductive wire (FC1) further comprises a gas-supply main line (LP) having a first bypass line (BP) connected fluidically to said gas-supply main line (LP) upstream and downstream from said first conductive wire (FC1), as seen in the direction of circulation of the gas in the main line (LP), andat least a first solenoid valve (EV1) is arranged in the area of the upstream connection of the first bypass line (BP) to said main line (LP) in such a way as to direct the gaseous flow containing the xenon either to the main line (LP), on which said first conductive wire (FC1) is arranged, or to the first bypass line (BP). 2. The apparatus of claim 1, wherein the gas-supply main line (LP) and the first bypass line (BP) are ramifications of a second bypass line (S1, S3, S5) communicating fluidically with the main circuit (CP), and in that the first solenoid valve (EV1) is arranged at the intersection of said second bypass line (S1, S3, S5), the gas-supply main line (LP) and the first bypass line (BP), and the connection of the second bypass line (S1, S3, S5) to the main circuit (CP) is made on the inhalation branch (16) and/or on the exhalation branch (18) and/or at a site located in immediate proximity to the patient's mouth. 3. The apparatus of claim 2, wherein the connection of the second bypass line (S1, S3, S5) to the main circuit (CP) is made in the area of a connection site (17) between the inhalation branch (16) and the exhalation branch (18) of said main circuit (CP). 4. The apparatus of claim 3, wherein at least a second solenoid valve (EV2) is arranged in the area of the downstream connection of the first bypass line (BP) to said main line (LP). 5. The apparatus of claim 4, wherein at least a second sensor with conductive wire (FC2) is arranged on the first bypass line (BP) and measures the xenon concentration, during the exhalation phase, while the first conductive wire (FC1) measures the concentration during the inhalation phase, or vice versa. 6. The apparatus of claim 5, wherein the first and second solenoid valves (EV1, EV2) are controlled by control means (3) or by an independent inte... | L'Air Liquide, Societe Anonyme pour l'Etude et l'Exploitation des Procedes Georges Claude | A61B005/08 A61M016/01 A61M016/10 A61M016/20 A61M016/12 G01N033/00 A61M016/00 A61M016/22 A61M016/08
US07962195 | B2 | 2011-06-14 | Model-based correction of position measurements | 2006-06-01 | Visualization of a probe when impedance-based measurement technology is being used is improved by stabilizing a displayed image of the probe or catheter. Using a model of reasonable probe shapes and a matching algorithm, an erroneous probe image is adjusted so that it assumes a realistic shape on a display. A range of positional variations is also incorporated in the model. When an apparent probe position exceeds a permissible range of motion, the probe image is constrained to a realistic position. | 1. A method of determining a position of a probe that has been inserted into a body of a subject, comprising the steps of: providing a processing unit and a probe model;determining an apparent position of said probe in said body with the processing unit;using the processing unit and the probe model to establish that a first displacement between said apparent position and a prior position of said probe at a known time corresponds to a first rate of motion that exceeds a predetermined limit;using the processing unit to adjust said apparent position to a new position such that a second displacement between said new position and said prior position corresponds to a second rate of motion that is less than said predetermined limit; anddisplaying said new position. 2. The method according to claim 1, wherein said apparent position is determined by measuring impedance between said probe and a plurality of locations that are remote from said probe. 3. The method according to claim 2, wherein measuring said impedance comprises performing the steps of: passing electrical currents through said body between at least one electrode disposed on said probe and said plurality of locations; andmeasuring respective characteristics of currents passing through said plurality of said locations. 4. The method according to claim 1, wherein said step of adjusting said apparent position is performed using a Kalman filter. 5. An apparatus for position sensing, comprising: a probe, comprising at least one probe electrode, which is adapted to be inserted into a body of a subject;a plurality of body surface electrodes, which are adapted to be fixed to a surface of said body at respective locations;a display; anda controller comprising a processing unit and a probe model and, which is adapted to be coupled to said probe and to said body surface electrodes so as to pass electrical currents through said body between said probe electrode and said body surface electrodes, and to determine position coordinates of said probe by measuring respective characteristics of said currents passing through said body surface electrodes, said controller being operative for:determining an apparent position of said probe in said body with the processing unit;using the processing unit and the probe model to establish that a first displacement between said apparent position and a prior position of said probe at a known time corresponds to a first rate of motion that exceeds a predetermined limit;using the processing unit to adjust said apparent position to a new position such that a second displacement between said new position and said prior position corresponds to a second rate of motion that is less than said predetermined limit; anddisplaying said new position on said display. 6. The apparatus according to claim 5, wherein said apparent position is determined by measuring impedance between said probe and said locations. 7. The apparatus according to claim 6, wherein measuring said impedanc... | Biosense Webster, Inc. | A61B005/05
US07949389 | B2 | 2011-05-24 | Fetal ECG monitoring | 2009-04-15 | A method for fetal monitoring includes acquiring electrical signals from a set of electrodes, for example, a set of surface electrodes applied to a maternal abdominal region. The electrical signals are analyzed, including by performing a morphological analysis of fetal electrocardiogram signals. A clinical indicator is then determined from a result of performing the morphological analysis. | 1. A method for fetal monitoring comprising: acquiring electrical signals from a plurality of electrodes, including a plurality of electrodes applied to a maternal abdominal region;analyzing the electrical signals, including performing a morphological analysis of fetal electrocardiogram signals;determining a clinical indicator of a pathologic condition from a result of performing the morphological analysis. 2. The method of claim 1 wherein performing the morphological analysis includes determining a quantitative measure of morphological variation. 3. The method of claim 2 wherein determining the quantitative measure of morphological variation includes characterizing segments of signals determined from the acquired electrical signals according to a group of classes, and determining a measure of variation in sequences of segment classifications. 4. The method of claim 2 wherein the quantitative measure of morphological variation includes an entropy of a sequence of segment classifications. 5. The method of claim 1 wherein determining a clinical indicator includes determining an indicator of a fetal condition. 6. The method of claim 1 wherein determining a clinical indicator includes determining an indicator of an inflammation condition. 7. The method of claim 1 wherein determining a clinical indicator includes determining an indicator of at least one of chorioamnionitis, preeclampsia, inflammation, and infection. 8. A fetal monitoring system comprising: a data acquisition system for acquiring surface signals including signals representing surface measurements of cardiac activity;a signal analyzer coupled to the data acquisition system, configured to analyze the acquired surface signals to generate an output having at least a clinical indicator characterizing a pathologic condition, the signal analyzer including: a signal processor for extracting fetal electrocardiogram signals from the acquired surface signals; anda clinical condition detector configured to perform a morphological analysis of the extracted fetal electrocardiogram signals, and based on a result of the morphological analysis, determining the clinical indicator; andan output system for presenting a representation of the clinical indicator. 9. The fetal monitoring system of claim 8, wherein the data acquisition system includes an electrode array having at least a plurality of electrodes attachable to a maternal abdominal region. 10. The fetal monitoring system of claim 9, wherein the electrode array further includes a second plurality of electrodes attachable to a maternal lumbar region. 11. The fetal monitoring system claim 10, wherein the electrode array further includes a third plurality of electrodes attachable to a maternal side region. 12. The fetal monitoring system of claim 11, wherein the pluralities of electrodes are positioned in a pre-determined arrangement on a garment. 13. The fetal monitoring system of claim 8, wherein the output system includes a display uni... | Tufts Medical Center, Inc. Massachusetts Institute of Technology E-Trolz, Inc. | A61B005/0444
US07938787 | B2 | 2011-05-10 | Method and apparatus for penetrating tissue | 2006-09-29 | A tissue penetrating system has a housing member. A plurality of penetrating members are positioned in the housing member. A tissue stabilizing member is coupled to the housing. A penetrating member sensor is coupled to the plurality of penetrating members. The penetrating member sensor is configured to provide information relative to a depth of penetration of a penetrating member through a skin surface. | 1. A system for withdrawing body fluid from a body part via a withdrawal site located on the surface of the body part, the system comprising: a housing member;a plurality of penetrating members positioned in the housing member;a plurality of sample chambers, each of a sample chamber associated with a penetrating member, each of a sample chamber including at least one non-optical analyte detecting member coupled to a sample chamber;an electronic penetrating member driver coupled to and for driving each of the plurality of penetrating members;a penetrating member transport device for moving each of the plurality of penetrating members into a position for alignment with the penetrating member driver; anda processor configured to control the penetrating member driver and the penetrating member transport device to drive the plurality of penetrating members and determine that a penetrating member has contacted a skin surface prior to a lancing event and in response the processor adjusting a speed of the penetrating member or the power delivered to the penetrating member for skin penetration to maintain a desired penetration velocity of the penetrating member, the processor determining whether a distal end of the penetrating member has reached a brake depth to achieve a desired final penetration depth of the penetrating member, wherein the processor uses a determination of skin contact by the penetrating member as a consideration of skin tenting by the penetrating member, wherein skin tenting is used as a consideration factor for brake depth, wherein the processor determines whether a distal end of the penetrating member has reached a brake depth based on an amount of said skin tenting, wherein the brake depth is pre-determined and programmed into the processor or the processor dynamically determines the brake depth during actuation and the processor causes a braking force to be applied to penetrating, wherein the processor is configured to provide a time and position of tissue contact by the penetrating member. 2. The system of claim 1, further comprising: a position or contact sensor. 3. The system of claim 1, wherein the processor is adapted to trigger at least one of the following reactions following choices of user dependent on the value of at least one detected variable: blocking or releasing of the at least one penetrating member; setting a puncture depth of the penetrating member selected for creating the opening in the body; setting the geometric shape of a press-on area of the system for withdrawing body fluid; setting the cross section of an opening in a press-on area of the system for withdrawing body fluid, through which opening the penetrating member can penetrate into the surface of the body part; selecting one of at least two press-on areas included in the system for withdrawing body fluid; selecting one of at least two penetrating members contained in the system for withdrawing body fluid; setting the duration of the withdrawal proced... | Pelikan Technologies, Inc. | A61B005/00
US07959568 | B2 | 2011-06-14 | Advanced patient management for identifying, displaying and assisting with correlating health-related data | 2008-12-22 | Systems, devices and methods for identifying, displaying and assisting in correlating health-related data. One aspect is a programmable device having machine executable instructions for performing a method for use in correlating data to manage a patient's health. In various embodiments, data related to the patient's health is acquired. The acquired data includes one or more trended health-related parameters, and at least one of one or more predetermined health-related events and one or more health-related alerts. Correlations are identified. The one or more trended health-related parameters and the at least one of the one or more predetermined health-related events and the one or more health-related alerts establish a causal relationship between the acquired data and the patient's health. | 1. A device for monitoring a patient's health condition, comprising: at least one data input to receive data, including one or more trended parameters, and at least one of one or more predetermined events and one or more alerts; anda correlation module to: communicate with the at least one data input;correlate the one or more trended parameters, the one or more predetermined events, and the one or more alerts; andestablish a correlative relationship; wherein the correlation module automatically triggers an implantable medical device (IMD) therapy change based on the correlative relationship. 2. The device of claim 1, wherein the correlation module to establish a correlative relationship includes a correlation module to establish a causal relationship. 3. The device of claim 2, wherein the correlation module to establish a causal relationship includes a correlation module to establish a causal relationship between data and the patient's health. 4. The device of claim 3, wherein the correlation module displays the data in a manner suitable to visibly identify the causal relationship with respect to the patient's health. 5. The device of claim 4, wherein the correlation module displays an indicator along with the data in a manner suitable to visibly identify the causal relationship with respect to the patient's health. 6. The device of claim 5, wherein the correlation module to trigger an implantable medical device (IMD) therapy change based on the causal relationship with respect to the patient's health. 7. The device of claim 3, wherein one or more correlation algorithms are used by the correlation module to automatically identify the causal relationship with respect to the patient's health. 8. The device of claim 3, wherein the correlation module triggers an action based on the causal relationship with respect to the patient's health. 9. The device of claim 1, wherein the at least one data input includes a first data input to receive one or more trended parameters, a second data input to receive one or more predetermined events, and a third data input to receive one or more alerts. 10. The device of claim 1, wherein the device includes a wellness monitoring device (WMD). 11. The device of claim 1, wherein the device includes an implantable medical device (IMD). 12. A system for correlating data for use in managing a patient's health, comprising: acquisition means for acquiring health-related data, including means for acquiring one or more trended parameters, and at least one of one or more predetermined events and one or more physician alerts; andcorrelation means for correlating the one or more trended parameters, the one or more predetermined events and the one or more physician alerts to establish a correlative relationship; wherein the correlation means automatically triggers an implantable medical device (IMD) therapy change based on the correlative relationship. 13. The system of claim 12, wherein the correlation means includes ... | Cardiac Pacemakers, Inc. | A61B005/00
US07945309 | B2 | 2011-05-17 | Dynamic metal immunity | 2002-11-22 | The disclosure concerns position-sensing apparatus, having radiators which generate electromagnetic energy fields and a position sensor which generates sensor signals responsive to the energy fields. Reference elements are placed at respective positions near the sensor to generate reference signals responsive to the energy fields. And, a control unit is used to calculate a position of the sensor based on sensor signals and reference element errors in order to account for the effects of interfering metal objects. | 1. Position-sensing apparatus, comprising: a set of radiators, which are adapted to be placed at respective positions in a vicinity of a body of a subject and to generate electromagnetic energy fields;a position sensor, which is adapted to placed in the body of the subject and to generate sensor signals responsive to the energy fields;one or more reference elements, which are adapted to be placed at fixed respective positions in a vicinity of the sensor and to generate reference signals responsive to the energy fields; anda control unit, which is adapted to:determine and store in memory, for each of the reference elements, respective undistorted reference parameters as undistorted magnetic field values, responsive to the positions of the reference elements,receive the sensor signals and the reference signals,calculate a reference element error for each reference element, responsive to an interaction of a metal article with the energy fields, responsive to the stored undistorted magnetic field values of the reference element, and responsive to the reference signal generated by the reference element, andcalculate a position of the sensor, responsive to the sensor signals and the reference element errors. 2. Apparatus according to claim 1, wherein at least one of the one or more reference elements is adapted to be placed outside of the body of the subject. 3. Apparatus according to claim 1, wherein at least one of the one or more reference elements is adapted to be placed at a fixed, known position relative to the set of radiators during operation of the apparatus. 4. Apparatus according to claim 1, wherein the control unit is adapted to calculate corrected sensor signals responsive to the reference element errors and responsive to the generated sensor signals, and to calculate the position of the sensor responsive to the corrected sensor signals. 5. Apparatus according to claim 1, wherein the control unit is adapted to designate the reference element error, for at least one of the reference elements, to be responsive to a measured magnetic field value at the at least one of the reference elements, which value is responsive to the interaction of the metal article with the energy fields. 6. Apparatus according to claim 1, wherein the control unit is adapted to calculate the undistorted reference parameter of at least one of the reference elements responsive to a relative position of the at least one of the reference elements with respect to the set of radiators. 7. Apparatus according to claim 6, wherein the control unit is adapted to perform the calculation of the reference parameter substantially independently of the interaction of the metal article with the energy fields. 8. Apparatus according to claim 6, wherein the control unit is adapted to calculate the undistorted reference parameter of the at least one of the reference elements responsive to a calculation of a magnetic field value at the at least one of the reference elements, which ... | Biosense, Inc. | A61B005/05
US08444654 | B2 | 2013-05-21 | Method for detecting body parameters | 2011-01-27 | A method for detecting biometric parameters includes the steps of performing a bone graft procedure on a joint, providing a biometric sensor at the joint, and measuring a biometric parameter at the joint with the sensor. The sensor is capable of measuring parameters in an adjacent surrounding including pressure, tension, shear, relative position, and vascular flow. Data relating to the biometric parameter is transmitted to an external source and the data are analyzed to evaluate a biometric condition of the joint. A set of the sensors can be placed on the joint. | 1. A method for detecting biometric parameters, which comprises: performing a surgical procedure on at least one joint;implanting at least one biometric transceiver at the at least one joint;transmitting a first energy wave from the transceiver into a procedure area including at least one of the joint and an area adjacent the joint;quantitatively assessing the behavior of the energy wave with the transceiver;after transmitting the first energy wave: transmitting a second energy wave from the transceiver into the procedure area; andquantitatively assessing the behavior of the second energy wave;wherein at least one of the first and second energy waves are pulsed during transmission in a vibratory manner to stimulate the procedure area in accordance with at least one detected parameter of the procedure area; anddetermining a current status of the at least one parameter of the procedure area selected from the group consisting of pressure, tension, shear, relative position, bone density, fluid viscosity, temperature, strain, angular deformity, vibration, venous flow, lymphatic flow, load, torque, distance, tilt, shape, elasticity, motion, bearing wear, subsidence, bone integration, change in viscosity, particulate matter, kinematics, stability, and vascular flow with the transceiver based upon a comparison between the assessed behavior of the first and second energy waves. 2. The method according to claim 1, which further comprises: transmitting data relating to the at least one biometric parameter to an external source; andanalyzing the data to evaluate a biometric condition of the at least one joint. 3. The method according to claim 2, which further comprises, based upon the evaluation of the biometric condition of the at least one joint, changing a currently ongoing interoperative procedure at the procedure area. 4. The method according to claim 2, which further comprises, based upon the evaluation of the biometric condition of the at least one joint, chronically monitoring the anatomic condition relating to the procedure area. 5. The method according to claim 1, which further comprises providing a set of the transceivers on the at least one joint. 6. The method according to claim 1, which further comprises providing energy from outside the procedure area to the transceiver to power the transceiver and, thereby, create the energy wave and quantitatively assess the behavior of the energy wave. 7. The method according to claim 6, which further comprises providing the energy through at least one of an electromagnetic couple, a magnetic couple, a capacitive couple, an inductive couple, a sonic couple, an ultrasonic couple, a fiber optic couple, an optical couple, and an infrared couple. 8. The method according to claim 1, which further comprises: providing a set of the biometric transceivers at the procedure area;transmitting an energy wave from the transceivers into the procedure area;quantitatively assessing the behavior of the energy wave wi... | - | A61B017/58 A61B017/60 A61F002/00
US09113797 | B2 | 2015-08-25 | Blood vessel characteristics measuring apparatus and blood vessel characteristics measuring method | 2009-07-03 | The blood vessel characteristics measuring apparatus 100 includes a blood flow measuring unit 20 held at a position opposed to a skin surface 10 of a measurement region of a subject, an optical sensor unit 30 housed in the blood flow measuring unit 20, an electrocardiograph 40, and a controller 50. The controller 50 includes a blood flow measuring part 60, a blood vessel displacement deriving part 70, and a blood vessel condition deriving part 80. The blood flow measuring part 60 measures the displacements of a blood vessel and tissues around the blood vessel caused by a blood flow based on light intensity at the time of receiving light emitted from a light-emitting part 32 of the sensor unit 30 in light-receiving parts 34 and 36. The blood vessel displacement deriving part 70 derives the wall of the blood vessel 12 based on the displacements of the blood vessel and tissues around the blood vessel. The blood vessel condition deriving part 80 determines pulse wave velocity at each measurement position based on a phase difference between the waveform of the electrocardiographic signal of the electrocardiograph 40 and a detection signal waveform obtained from the light-receiving parts 34 and 36, and derives the condition of the displacement of the wall of the blood vessel 12 from the pulse wave velocity. | 1. A blood vessel characteristics measuring apparatus, comprising: a sensor unit provided at a position opposed to a measurement region of a subject, the sensor unit including a light-emitting part configured to emit light onto the measurement region, a light-receiving part configured to contactlessly receive the light propagated through the measurement region and output a detection signal corresponding to a light intensity of the received light, and a radio communication part configured to convert the detection signal into a radio signal and transmit the radio signal; anda controller configured to receive the radio signal transmitted from the radio communication part of the sensor unit, the controller including a blood flow measuring part configured to measure a displacement of a blood vessel and a displacement of a tissue around the blood vessel caused by a blood flow in the measurement region based on the light intensity read from the received radio signal;a blood vessel displacement deriving part configured to derive a displacement of a wall of the blood vessel based on the displacement of the blood vessel and the displacement of the tissue around the blood vessel measured by the blood flow measuring part;an electrocardiographic measurement part configured to measure an electrocardiographic signal of the subject;a blood vessel condition deriving part configured to determine a pulse wave velocity using a difference between a waveform of the electrocardiographic signal and one of a detection signal waveform obtained from the light-receiving part and a waveform of the derived displacement of the wall of the blood vessel, and to derive a condition of the wall of the blood vessel at each of measurement positions by deriving blood vessel characteristics corresponding to the determined pulse wave velocity from a database, wherein the derived condition of the wall of the blood vessel is a degree of arteriosclerosis, and the derived blood vessel characteristics include a modulus of elasticity of the blood vessel, an amount of plaque inside the blood vessel, and the degree of arteriosclerosis; anda control part configured to generate a first image based on the measured displacement of the blood vessel and the measured displacement of the tissue around the blood vessel and a second image based on the derived condition of the wall of the blood vessel, and display the first image and the second image on a display. 2. The blood vessel characteristics measuring apparatus as claimed in claim 1, wherein the blood vessel condition deriving part is configured to derive the condition of the wall of the blood vessel at each of the measurement positions based on a phase difference between the waveform of the electrocardiographic signal and the detection signal waveform obtained from the light-receiving part. 3. The blood vessel characteristics measuring apparatus as claimed in claim 1, wherein the blood flow measuring part is configured to optically measure a bl... | University of Tsukuba | A61B005/02 A61B005/026 A61B005/0285 A61B005/00 A61B005/021 A61B005/0295
US20110144515 | A1 | 2011-06-16 | SYSTEMS AND METHODS FOR PROVIDING ENVIRONMENT MONITORING | 2009-02-16 | The present invention describes systems and methods for providing portable environment monitoring systems. An exemplary embodiment of the present invention provides a portable environment monitoring system comprising a sensor enabled to sense an airborne analyte. The portable environment monitoring system also includes a microprocessor in communication with the sensor and enabled to process information received from the sensor. Additionally, the portable environment monitoring system includes a memory device in communication with the microprocessor and enabled to store information received from the microprocessor. Furthermore, a user is enabled to ambulate with the portable environment monitoring system. | 1. A portable environment monitoring system comprising: a sensor enabled to sense an airborne analyte;a respiratory monitoring device;a microprocessor in communication with the sensor and enabled to process information received from the sensor and information received from the respiratory monitoring device;a memory device in communication with the microprocessor and enabled to store information received from the microprocessor; andwherein a user is enabled to ambulate with the system. 2. (canceled) 3. The portable environment monitoring system of claim 1, wherein the respiratory monitoring device is enabled to provide data regarding a pulmonary function of the user. 4. The portable environment monitoring system of claim 3, wherein the respiratory monitoring device is enabled to identify a reduction in the pulmonary function of the user. 5. The portable environment monitoring system of claim 1, wherein the sensor, the microprocessor, and the memory device are contained within a housing and, together with the housing, have a total specific weight of less than five pounds. 6. (canceled) 7. The portable environment monitoring system of claim 6, wherein the housing containing the sensor, the microprocessor, and the memory device has a specific total weight of less than one pound. 8. The portable environment monitoring system of claim 1, wherein the sensor is a volatile organic compound sensor. 9. The portable environment monitoring system of claim 8, further comprising a carbon dioxide sensor, nitrogen dioxide sensor, an ozone sensor, and a formaldehyde sensor. 10. A method for determining a respiratory deficiency trigger comprising: providing a user with a portable environment monitoring system comprising a sensor, a microprocessor, a memory device, and a respiratory monitoring device;collecting a plurality of data from the sensor and the respiratory monitoring device with the user in a plurality of environments;analyzing the plurality of data received from the sensor and the respiratory monitoring device; anddetermining whether a relationship exists between the exposure of the user to an airborne analyte and a deficient pulmonary function by the user. 11. The method for determining an asthmatic trigger of claim 10, wherein the user is enabled to ambulate with the portable environment monitoring system. 12. The method for determining an asthmatic trigger of claim 10, wherein the deficient pulmonary function by the user corresponds to asthmatic reaction. 13. The method for determining an asthmatic trigger of claim 12, further comprising establishing an asthmatic trigger based on the determination of the relationship between the exposure of the user to the airborne analyte and the deficient pulmonary function by the user. 14. The method for determining an asthmatic trigger of claim 10, wherein the collecting of a plurality of data from the sensor is automatically performed by the portable environment monitoring system. 15. The method for ... | Georgia Tech Research Corporation | A61B005/08
US07966062 | B2 | 2011-06-21 | Multi-color overlay system for processing and displaying electrocardiac signals | 2006-05-12 | In one aspect, a method for displaying closely proximate cardiac signals is provided and can include the steps of: (a) identifying one or more overlapping portions of a template signal and a data signal; (b) processing the overlapping portion so as to have a first color when displayed on a display; (c) processing the non-overlapping portion of the template signal to have a second color; (d) processing the non-overlapping portion of the data signal to have a third color; and (e) displaying the processed signals. | 1. A method for displaying closely proximate cardiac signals comprising the steps of: identifying one or more overlapping portions of a template signal and a data signal, wherein the template signal and the data signal each comprise a segment of a electrocardiogram (ECG) waveform signal;processing the overlapping portion so as to have a first color when displayed on a display;processing the non-overlapping portion of the template signal to have a second color;processing the non-overlapping portion of the data signal to have a third color, wherein the first color is different than the second and third colors; anddisplaying the processed signals in the first, second and third colors to show the degree of alignment therebetween,wherein the identifying step identifies the overlapping portions as including (a) perfect matches between the template signal and the data signal, and (b) not perfect matches of closely-approximating, additional portions in which a degree of alignment between the template signal and the data signal exceeds a predetermined threshold value. 2. The method of claim 1, including the additional step of deriving the data signal from an average of segments of ECG waveform signals from multiple beats of a heart. 3. The method of claim 1, wherein the overlapping portion is defined as portions of the template signal and data signal that perfectly match one another. 4. The method of claim 1, wherein the threshold value is greater than 95% correlation. 5. The method of claim 1, wherein the first, second and third colors are readily distinguishable from one another. 6. The method of claim 1, including the additional steps of capturing ECG waveform signals and selecting first and second ECG waveform signals from respective first and second heart beats that occur immediately after one another, wherein the first ECG waveform signal represents the data signal and the second ECG waveform signal represents the template signal. 7. The method of claim 1, further including the steps of: calculating a quantitative indicator of a correlation coefficient calculation that is used to identify a best fit between the template signal and the data signal, andoutputting the quantitative indicator. 8. The method of claim 7, including the additional step of defining the quantitative indicator as either: a composite average of coefficients by a calculating step using multiple leads that capture at least the data signal; ora bar graph showing percentage of fit. 9. The method of claim 1, further including the step of: performing a correlation coefficient calculation to identify a best fit between the template signal and the data signal. 10. The method of claim 1, wherein processing the template signal and data signal includes the steps of: selecting a QRS-T segment of a reference ECG signal;permitting a user to mark a begin point and an end point of the selected QRS-T segment;defining the template signal as being a waveform segment between the marked beg... | C.R. Bard, Inc. | A61B005/044
US07969142 | B2 | 2011-06-28 | System and method for magnetic tracking of a sensor having an asymmetric magnetic core | 2006-07-12 | A system and method for determining the location of a remote object connected to an induction coil using a magnetic tracking sensor. The system and method include locating a magnetic core asymmetrically disposed within the induction coil located near the remote object, and operably connecting a single DC electrical circuit o ends defining the induction coil. The DC electrical circuit provides a DC current to the induction coil while the induction coil is disposed in an external AC magnetic field. The DC current adjusts the level of saturation of the magnetic core, and hence varies a response signal of the induction coil disposed in the external AC magnetic field to provide magnetic tracking of the induction coil. | 1. A magnetic tracking sensor for determining the location of a remote object, comprising: an induction coil connectable to the remote object;a magnetic core asymmetrically disposed within the induction coil; anda DC electrical circuit operably connected to ends defining the induction coil, the DC electrical circuit providing a DC current to the induction coil to adjust an influence the magnetic core has on a response signal of the induction coil disposed in an external AC magnetic field and provide magnetic tracking of the remote object based on the response signal. 2. The magnetic tracking sensor of claim 1, wherein the magnetic core is a soft magnetic wire noncoaxially disposed within the induction coil. 3. The magnetic tracking sensor of claim 2, wherein the soft magnetic wire is defined by opposing ends disposed against opposing sides defining an interior of the induction coil. 4. The magnetic tracking sensor of claim 2, wherein the magnetic core includes a plurality of soft magnetic wires substantially parallel to one another and noncoaxially disposed within the induction coil. 5. The magnetic tracking sensor of claim 1, further comprising a plurality of induction coils operably connected to the DC electrical circuit. 6. The magnetic tracking sensor of claim 1, wherein the induction coil includes a series of collinear induction coils each having a noncoaxially aligned soft magnetic wire as the magnetic core, the series of collinear induction coils operably connected to the DC electrical circuit. 7. The magnetic tracking sensor of claim 1, wherein the magnetic core includes three soft magnetic wires substantially parallel to one another and noncoaxially disposed within the induction coil. 8. The magnetic tracking sensor of claim 1, wherein the DC electrical circuit includes a DC power supply and a signal response acquisition device. 9. The magnetic tracking sensor of claim 1, wherein the magnetic core is diagonally disposed within the induction coil to cancel axial symmetry of the induction coil. 10. A system for determining the location of a remote object, comprising: an induction coil connectable to the remote object;a magnetic core asymmetrically disposed within the induction coil;a DC electrical circuit operably connected to ends defining the induction coil, the DC electrical circuit providing a DC current to the induction coil to adjust an influence the magnetic core has on a response signal of the induction coil disposed in an external AC magnetic field;a data acquisition system adapted to receive a response signal from the induction coil; anda monitor operably coupled to the data acquisition system,wherein magnetic tracking of the induction coil is viewable on the monitor. 11. The system of claim 10, wherein the magnetic core is a soft magnetic wire noncoaxially disposed within the induction coil. 12. The system of claim 11, wherein the soft magnetic wire is defined by opposing ends disposed against opposing sides defining ... | Koninklijke Philips Electronics N.V. | G01B007/14 A61B005/05
US20110124986 | A1 | 2011-05-26 | Microchip, Blood Analysis System and Blood Analysis Method | 2009-07-06 | A microchip to be installed in a blood analysis system for measuring properties of blood, provided with a plurality of channels through which blood passes, wherein only each of first subset channels among the plurality of channels is provided with a first plurality of barriers for locally changing a direction of blood flow, in a blood flow direction on a wall face demarcating the channels, and among the first plurality of barriers, barriers which are adjacent to each other in the blood flow direction differ in shape from each other. | 1. A microchip to be equipped in a blood analysis system for measuring properties of blood, comprising: a plurality of channels through which blood passes, wherein only each of first subset channels among the plurality of channels is provided with a first plurality of barriers for locally changing a direction of blood flow, in a blood flow direction on a wall face demarcating the channels, and among the first plurality of barriers, barriers which are adjacent to each other in the blood flow direction differ in shape from each other. 2. A microchip to be equipped in a blood analysis system for measuring properties of blood, comprising: a plurality of channels through which blood passes, wherein only each of first subset channels among the plurality of channels is provided with a second plurality of barriers for locally changing a direction of blood flow, in a blood flow direction on a wall face demarcating the channels, and the second plurality of barriers have random shapes. 3. The microchip of claim 1, wherein each of other subset channels than the first subset channels among the plurality of channels is provided with a third plurality of barriers for locally changing a direction of blood flow, in a blood flow direction on a wall face demarcating the channels with a constant interval, and the third plurality of barriers have a same shape. 4. The microchip of claim 1, wherein the plurality of barriers are provided so as to protrude inside the channels. 5. The microchip of claim 1, wherein an inner wall portion of the plurality of channels is configured to have a cross section area becoming wider or becoming narrower from upstream toward downstream of the blood flow. 6. The microchip of claim 1, wherein, the wall face demarcating the channels among the plurality of channels, on which the plurality of barriers are not provided, is formed flat. 7. The microchip of claim 1, wherein the wall face demarcating the plurality of channels has a circular section. 8. A blood analysis system comprising: the microchip of claim 1;an imaging device to capture an image of blood flow in the plurality of channels of the microchip; andan analyzing section to analyze the image captured by the imaging device and calculates a blood property. 9. The blood analysis system of claim 8, wherein the analyzing section calculates the blood property by analyzing the captured image at a prescribed channel among the plurality of channels. 10. The blood analysis system of claim 8, wherein the analyzing section compares each blood properties by analyzing each of the captured images at a channel of the first subset channels and at a channel of other channels than the first subset channels among the plurality of channels respectively. 11. The blood analysis system of claim 8, wherein the imaging device captures images of blood flow at a channel of the first subset channels and at the other channel among the plurality of channels. 12. A blood analysis method utilizing the m... | - | A61B005/00
US07949382 | B2 | 2011-05-24 | Devices, systems, methods and tools for continuous glucose monitoring | 2010-07-27 | One aspect of the invention provides a glucose monitor having a plurality of tissue piercing elements, each tissue piercing element having a distal opening, a proximal opening and interior space extending between the distal and proximal openings; a sensing area in fluid communication with the proximal openings of the tissue piercing elements; sensing fluid extending from the sensing area into substantially the entire interior space of the tissue piercing elements; and a glucose sensor adapted to detect a concentration of glucose in the sensing fluid within the sensing area. Another aspect of the invention provides a method of in vivo monitoring of an individual's interstitial fluid glucose concentration including the following steps: inserting distal ends of a plurality of tissue piercing elements through a stratum corneum area of the individual's skin, the tissue piercing elements each having a distal opening, a proximal opening, an interior space extending between the distal and proximal openings, and a sensing fluid filling substantially the entire interior space; and sensing a glucose concentration of the sensing fluid. | 1. A method of in vivo monitoring of an individual's interstitial fluid comprising: creating a plurality of fluid paths through a stratum corneum area of the individual's skin, the fluid paths each comprising a distal end in fluid communication with interstitial fluid of the individual, a proximal end in analyte communication with a sensing area located outside of the patient's body, and an interior space extending between the distal and proximal ends of the fluid path; moving sensing fluid out of a sensing fluid reservoir and into the sensing area;holding the sensing fluid in the sensing area in a non-flowing manner;allowing at least one analyte to passively diffuse from the patient's interstitial fluid through the fluid paths and into the sensing area;sensing a concentration of the at least one analyte in the sensing fluid using a sensor located at least partially in the sensing area during the holding step. 2. The method of claim 1 wherein fluid in the fluid paths is not moving when the at least one analyte is passively diffusing there-through. 3. The method of claim 1 wherein there is no fluid movement between the distal ends of the fluid paths and the sensor when the at least one analyte is passively diffusing through the fluid paths and into the sensing area. 4. The method of claim 1 wherein no interstitial fluid is extracted from the individual during the diffusion and sensing steps. 5. The method of claim 1 further comprising displaying analyte concentration information remote from the stratum corneum area of the individual's skin. 6. The method of claim 5 further comprising wirelessly communicating analyte concentration information to a display. 7. The method of claim 1 wherein the sensing step is performed by a sensor in fluid communication with the sensing area and the interior spaces of the fluid paths, the method further comprising calibrating the sensor by moving sensing fluid into the sensing area. 8. The method of claim 7 further comprising moving sensing fluid out of the sensing area as sensing fluid is moved into the sensing area. 9. The method of claim 1 further comprising attaching a device comprising the sensor area and the sensor to the individual with adhesive. 10. The method of claim 1 further comprising providing a membrane located between the distal ends of the fluid paths and the sensor. 11. The method of claim 1 wherein creating the plurality of fluid paths through the stratum corneum area of the individual's skin comprises applying a microneedle array to the stratum corneum area. 12. The method of claim 1 wherein the proximal ends of the fluid paths are in fluid communication with the sensing area. 13. The method of claim 1 wherein the step of moving sensing fluid comprises operating an actuator. 14. The method of claim 1 wherein the step of moving sensing fluid comprises moving sensing fluid having a glucose concentration of between about 0 mg/dl and about 400 mg/dl. 15. The method of claim 1 wherein th... | ArKal Medical, Inc. | A61B005/05
US07957786 | B2 | 2011-06-07 | Methods and apparatus for communicating signals between portions of an apparatus in relative movement to one another | 2005-08-15 | A data transmission system that includes a transmitter; a receiver; and a magnetically coupled rotary transformer having a first portion and a second portion in moveable relationship with one another. The transmitter is electrically and mechanically coupled to the first portion of the rotary transformer, and the receiver is electrically and mechanically coupled to the second portion of the rotary transformer. The transmitter and the receiver are configured to wirelessly communicate data across the rotary transformer while the first and the second portions of the rotary transformer are in relative rotary motion. | 1. An imaging system for acquiring an image of an object, the imaging system comprising: a rotating gantry;a stationary electronic system and a rotating electronic system that rotates with the rotating gantry;a magnetically coupled rotary transformer with a magnetic core having both a stationary portion and a rotating portion, the magnetic core including a data winding carried in a first groove and a power winding carried in the first groove;wherein the rotating electronic system is electrically and mechanically coupled to the rotating portion of the rotary transformer, and the stationary electronic system is electrically and mechanically coupled to the stationary portion of the rotary transformer, and wherein the rotating electronic system and stationary electronic system are configured to wirelessly communicate data modulated on a carrier frequency at a frequency different than that of electrical power transmitted across the rotary transformer. 2. A system in accordance with claim 1 further comprising a radio frequency modulator electrically and mechanically coupled to the rotating portion of the rotary transformer and a radio frequency demodulator electrically and mechanically coupled to the stationary portion of the rotary transformer, wherein the radio frequency modulator is configured to modulate data from the rotating electronic system onto a carrier frequency and the radio frequency demodulator is configured to demodulate carrier frequency to recover the data for the stationary electronic system. 3. A system in accordance with claim 2 having a separate data winding and a separate power winding in both the stationary portion and the rotating portion of the rotary transformer. 4. A system in accordance with claim 3 configured to transmit power between the power windings at a low frequency and to transmit data at a high carrier frequency. 5. A system in accordance with claim 4 wherein the frequency for transmitting power between the power windings is 200 kHz or less, and the frequency for transmitting data is 1 MHz or more. 6. A system in accordance with claim 4 further comprising at least one filter configured to remove or reduce a power component on the data windings. 7. A system in accordance with claim 1 further comprising a power supply, and further configured to couple power across the rotary transformer. 8. A system in accordance with claim 1 configured to wirelessly communicate the data across an air gap in the rotary transformer. 9. A system in accordance with claim 1 configured to phase a carrier with the data to reduce waveform jitter. 10. A system in accordance with claim 1 wherein the imaging system is a computed tomographic imaging system. 11. A data transmission system comprising: a transmitter;a receiver; anda magnetically coupled rotary transformer having a first portion and a second portion in moveable relationship with one another, the first portion including a first magnetic core, the first magnetic core includi... | General Electric Company | A61B005/055 H05G001/10 H05G001/02 H01F021/02
US07966056 | B2 | 2011-06-21 | Magnetic resonance imaging method and compounds for use in the method | 2003-12-16 | The present invention relates to a magnetic resonance imaging (MRI) method, in particular to a MRI method enabling early detection of myocardial ischemia and to compounds for use as MR contrast agents in the method. | 1. A method of magnetic resonance (MR) imaging to identify diseased areas comprising areas suffering from myocardial ischemia, said method comprising a) administering to a human or non-human animal body a contrast agent comprising (i) at least one complex of the formula Mnm HPTA Zo,wherein m is 1 or 2, o is 0 to 2, Z is hydrogen or a pharmaceutical acceptable counterion, HPTA is 1,7-dicarboxy-2,6-bis(carboxymethyl)-4-hydroxy-2,6-diaza)-heptane,wherein said complex has a thermodynamic formation constant between 103 and 1016 and wherein said complex shows a dissociation of greater than 50% and a half-life of less than 1 minute under physiological conditions; and(ii) 0 to 2 mol Ca2+per mol Mn2+;b) exposing said body to a regime of physical and/or pharmacological stress before or simultaneously to the contrast agent administration;c) collecting MR imaging data and;d) providing MR images of an area of interest; and identifying the areas suffering from myocardial ischemia using the MR images. 2. The method according to claim 1, wherein the contrast agent comprises a mixture said complex. 3. The method according to claim 1, wherein the thermodynamic formation constant k is between 105 and 1010 . 4. The method according to claim 1, wherein the contrast agent comprises 0.1 to 1.75 mol Ca2+per mol Mn2+. 5. The method according to claim 1, wherein the contrast agent comprises 0.5 to 1 mol Ca2+per mol Mn2+. 6. The method according to claim 1 where Mn is Mn2+, o is 0 or 2 and Z is hydrogen, sodium or calcium. 7. The method according to claim 1, wherein the stress is a physical stress. 8. The method according to claim 1, wherein the contrast agent is administered at peak stress. 9. The method according to claim 3, wherein the thermodynamic formation constant k is between 107 and 109.5. 10. The method according to claim 1, wherein the area of interest is the myocardium. 11. A mixture comprising a complex having a thermodynamic formation constant between 103 and 1016 which comprises Mn2+and a complexing moiety, wherein said complexing moiety is HPTA and said complex shows a dissociation of greater than 50% and a half-life of less than 1 minute under physiologic conditions and 0.1 to 2 mol Ca2+per mol Mn2+for the manufacture of a contrast agent for use in MR imaging of the myocardium. 12. The mixture according to claim 11, wherein the mixture comprises 0.5 to 1 mol Ca2+per mol Mn2+. 13. Mixture for use in the MR imaging detection of myocardial ischemia and infarction wherein said mixture comprises (i) at least one complex of the formula MnmHPTA Zo,wherein m is 1 or 2, o is 0 to 2, Z is hydrogen or a pharmaceutical acceptable counterion,wherein said complex has a thermodynamic formation constant between 103 and 1016 and wherein said complex shows a dissociation of greater than 50% and a half-life of less than 1 min under physiological conditions; and(ii) 0 to 2 mol Ca2+per mol Mn2+. | GE Healthcare AS | A61B005/05 A61B005/055 A61B005/103
US07962189 | B2 | 2011-06-14 | Non-invasive living body measuring device | 2006-07-12 | A non-invasive living body measuring device is described, a representative one of which includes: a non-invasive living body measuring device for measuring the components contained in blood, comprising: an imaging part for imaging a living body; a display part; and a display control part for generating a blood vessel image showing blood vessels within a living body image obtained by the imaging part, and displaying the blood vessel image on the display part, wherein an index representing a suitable region for imaging by the imaging part is displayed on the display part. | 1. A non-invasive living body measuring device for measuring components contained in a blood, comprising: a device body; anda mounting part that is configured to be movably mountable on a wrist of a living body and holds the device body,wherein the device body comprises: an imaging part including a lens and an image pickup element, and configured to image the wrist of the living body by the image pickup element through the lens;a display part; anda controller that displays an image of the wrist obtained by the imaging part on the display part, generates a blood vessel image showing a blood vessel within the obtained image of the wrist, and displays the generated blood vessel image and two index lines on the display part, a region between the two index lines representing a suitable region for imaging a blood vessel by the imaging part,wherein the mounting part holds the device body so as to be rotatable around an optical axis of the lens, andwherein the controller measures the components in a blood by analyzing a blood vessel image of the wrist obtained by the imaging part. 2. The non-invasive living body measuring device according to claim 1, wherein the blood vessel image is a blood vessel pattern generated based on a pixel of a lowest brightness in brightness distribution determined in a fixed direction from a plurality of specific pixels in the image of the wrist. 3. The non-invasive living body measuring device according to claim 1, wherein the controller determines whether or not the blood vessel image is positioned within the two index lines. 4. The non-invasive living body measuring device according to claim 3, further comprising an execution key for starting a measurement,wherein the execution key is valid when the controller determines that the blood vessel image is positioned within the two index lines. 5. The non-invasive living body measuring device according to claim 3, further comprising: an alerting part;wherein when the controller determines that the blood vessel image is positioned within the two index lines, the alerting part alerts the user that the blood vessel is positioned within the region suitable for image the blood vessel. 6. The non-invasive living body measuring device according to claim 3, further comprising: an alerting part;wherein when the controller determines that the blood vessel image is not positioned within the two index lines, the alerting part alerts the user to a direction in which to move the device to bring the blood vessel within the two index lines. 7. The non-invasive living body measuring device according to claim 3, wherein: the controller changes a color of the two index lines when the controller determines that the blood vessel image is positioned within the two index lines. 8. The non-invasive living body measuring device according to claim 1, wherein when the blood vessel image is not positioned within the two index lines, the controller displays on the display part a direction in which t... | Sysmex Corporation | A61B005/00 A61B006/00
US07949179 | B2 | 2011-05-24 | Shape measurement device and method thereof | 2004-09-15 | The present invention is a system for determining the three dimensional shape of an object, having a first measuring means for measuring a first distance and direction from a fixed first observation point to a first group of measurement points on an object, and a second measuring means for measuring a second distance and direction from a movable second observation point to a second group of measurement points on the aforementioned object, at least one point among said second group of measurement points being the same measurement point as at least one point among the aforementioned first group of measurement points, and said second group of measurement points including at least one measurement point not included in the aforementioned first group of measurement points, and further being provided with a calculating means for calculating the traveling velocity vector of the second observation point from the measurement results concerning the aforementioned same measurement point due to the first and second measuring means, correcting the aforementioned second measurement results based upon said traveling velocity vector, and calculating the three dimensional coordinates of the measurement points of the first group and the second group, and a displaying means that displays an image of the object based upon the aforementioned three dimensional coordinates. | 1. A system for determining the three dimensional shape of an object, comprising: a first measuring device for measuring a first distance and direction from a fixed first observation point to a first group of measurement points on the object, thereby obtaining first measurement results;a second measuring device for measuring a second distance and direction from a movable second observation point to a second group of measurement points on the object, thereby obtaining second measurement results, wherein at least one point among said second group of measurement points is a common measurement point, the common measurement point being the same measurement point as at least one point among the first group of measurement points, said second group of measurement points including at least one measurement point not included in the first group of measurement points;a computer for: (i) calculating a traveling velocity vector of the second observation point from the measurement results concerning the common measurement point made by the first and second measuring devices,(ii) correcting the second measurement results based upon said traveling velocity vector, and(iii) calculating the three dimensional coordinates of the measurement points of the first group and the second group of measurement points; anda display that displays an image of the object based upon the three dimensional coordinates. 2. A system as described in claim 1, characterized in that the second measuring device is provided with a scanner unit comprising a laser radar unit for ranging each point, a four-faceted polygon minor for performing horizontal scanning, and a planar swing mirror for performing vertical scanning. 3. A system as described in claim 2, characterized in that the second measuring device is provided with a controller unit which internally houses a radar unit control portion, control portions for two minors, and an interface portion for sending measurement results to a measuring computer. 4. A system as described in claim 3, characterized in that the second measuring device is provided with a computer having a recording medium that can store measurement results on said recording medium and that can control the a scanner unit and controller unit. 5. A system as described in claim 1, wherein the traveling velocity vector is determined by the conjugate gradient method, with the assumption that the traveling velocity vector is time-independent, using the measurement results concerning the common measurement point, and wherein the measurement result from the second observation point is corrected using said traveling velocity vector. 6. A system as described in claim 5, characterized in that the translational motion vector that minimizes the following equation is determined by the conjugate gradient method:[Equation⁢⁢1]E⁡(p)=1N⁡(M-1)⁢∑iN⁢⁢∑jM⁢⁢ρ⁡(zij⁡(p))⁢⁢where⁢⁢p=(m,q)⁢⁢zij⁡(p)=R⁡(q)⁢g⁡(v)⁢i+m-yij2⁢⁢ρ⁡(zij⁡(p))=log⁡(1+12⁢zij⁢i⁡(p)2)N: number of points of measured dataM: n... | The University of Tokyo Japan Science and Technology Agency | G06K009/46 A61B005/103 G01R019/00
US08688185 | B2 | 2014-04-01 | Device for measuring blood, tissue, and skin parameters | 2009-12-18 | The invention relates to a device and a method for measuring blood, tissue, or skin parameters, in particular the oxygen saturation in blood, by attaching one or more sensors to body parts such as fingers, earlobes, toes, hand, or foot. The invention relates further to a method for the preparation of said device. | 1. A device for measuring blood, tissue, or skin parameters, including oxygen saturation in blood, through attachment to bodily organs, said device comprising: a flexible fixing strap (2) having a length defined by a longitudinal axis, wherein said flexible fixing strap is comprised of a tension-proof upper part (2.1) having a hook-and-loop fastening member (3) arranged thereto at one end and a plastic foam-like, skid-resistant lower part (2.2) congruent in size and shape with said upper part;a foldable capsule band (7) comprised of mating first and second halves and a self-adherent inner surface, wherein neither said first half nor said second half is congruent in size with either the upper part or the lower part of said flexible fixing strap;a transmitting optics (4), a detector (5), and a connecting cable (6) mounted to the first half of said capsule band by means of said self-adherent inner surface and then covered by the second half of said capsule band to yield a sealed, moisture-tight sensor unit, wherein said sensor unit has a length defined by a longitudinal axis whereby the length of said sensor unit is less than the length of said flexible fixing strap, further wherein said sensor unit is arranged between said upper and lower parts of said fixing strap; anda strain relief band;wrapped around said connecting cable so as to further secure said sensor unit to said fixing strap. 2. The device according to claim 1, characterized in that the strain relief band is self-adhering. 3. The device according to claim 1, characterized in that a lower surface of said upper part (2.1) is connected to an upper surface of said lower part (2.2) by a layer of hook-and-loop material. 4. The device according to claim 1, characterized in that a lower surface of said upper part (2.1) is connected to an upper surface of said lower part (2.2) by a layer of adhesive. 5. The device according to claim 1, characterized in that the capsule band (7) and the lower part (2.2) are provided with windows (2.3) in the area of the transmitting optics (4) and the detector (5). 6. The device according to claim 1, characterized in that the upper part (2.1) has an upper surface formed of a hook-and-loop material. 7. The device according to claim 1, characterized in that the connecting cable (6) is disposed between the upper and lower parts of said fixing strap and thus the strain relief band (8) is wrapped around both the connecting cable (6) and the upper part (2.1) and the lower part (2.2) of the fixing strap (2). 8. The device according to claim 1, characterized in that the foldable capsule band (7) is composed of a mainly white moisture-impermeable material. 9. A method for measuring blood, tissue, or skin parameters, including oxygen saturation in blood, characterized in that a device according to claim 1 is attached to a body part selected from among fingers, toes, earlobes, hands, and feet. 10. A method for the preparation of a device for measuring blood, tissue... | EnviteC-Wismar GmbH | A61B005/00
US08446273 | B2 | 2013-05-21 | Environmental risk management system and method | 2009-08-14 | A system and method for directing and monitoring the whereabouts of persons within an environmentally hazardous area includes equipping each person with devices for monitoring personal physiological conditions (56), equipment conditions (64), topographical locations (34), environmental conditions (57) and other pertinent data. The data are individually encoded and uploaded to a processing center (36) where they are analyzed (58,60,61) in order to ensure compliance with health or environmental norms and safety regulations, generate hot-spot mapping (53) and to issue real-time risk avoidance behavior directives (62), worker, mechanism and material traffic directions, warnings (59), permissions or interdictions. | 1. A method for managing and monitoring the safe circulation of individuals within a hazardous area which comprises: providing at least one environmental condition monitor for measuring an environmental parameter associated with a hazard at a site;providing at least one physiological condition monitor to at least one of said individuals, namely a worker; said physiological condition monitor being adapted to measure a physiological parameter associated with exposure of said worker to said hazard;providing at least one piece of protective equipment to said worker, said equipment being adapted to diminish the physiological effect of said hazard upon said worker;scheduling said worker to enter said site containing said hazard;maintaining a historical database of monitored data comprising: historical measurements of said physiological parameter associated with said worker;historical measurements of said environmental parameter associated with said hazard at said site; and,historical measurements of an equipment parameter associated with said piece of protective equipment; and,analyzing said historical database of monitored data to obtain: a current effectiveness condition of said piece of equipment;a current physiological condition for said worker;a current hazard level condition in said site;assessing said obtained conditions to determine a risk;generating a real-time directive to said worker in response to said risk. 2. The method of claim 1, wherein said generating comprises generating and issuing predictive warnings. 3. The method of claim 1, wherein said generating comprises generating and issuing worker behavior directions. 4. The method of claim 1, wherein said generating comprises generating and issuing worker, mechanism and material traffic directions. 5. The method of claim 1, wherein said historical measurements of physiological parameter and said historical measurements of said environmental parameter comprise readings taken by a personalized condition monitor. 6. The method of claim 1, wherein: said hazard is dust;said environmental parameter is airborne dust concentration;said physiological condition monitor is a breathing monitor;said physiological parameter is respiration rate;said protective equipment is an air purifying mask;said equipment parameter is filter core effectiveness; and,said directive is an order to rest for a period. 7. The method of claim 1, wherein said providing at least one piece of protective equipment comprises providing at least one equipment condition monitor associated with said at least one piece of protective equipment. 8. The method of claim 1, wherein said method further comprises providing at least one location monitor for tracking a location of said worker. 9. The method of claim 1, which further comprises operating a processing center programmed for said maintaining, said analyzing, said assessing, and said generating. 10. The method of claim 1, which further comprises providing a network of wir... | Honeywell International Inc. | G08B001/08 G08B023/00
US08351573 | B2 | 2013-01-08 | Device and method for positioning a target volume in a radiation therapy apparatus | 2010-12-06 | A device is provided for positioning a target volume, such as a phantom or a patient, in a radiation therapy apparatus. The apparatus directs a radiation beam towards the target. The apparatus can include a target support whereon the target is immobilized, a two dimensional radiation detector fixed with fixations means in a known geometric relationship to the target support, the radiation detector being capable of detecting the position of intersection of the radiation beam with the detector, and correcting means for correcting the relative position of the beam and the target support, based on the detected intersection position. | 1. A radiation therapy apparatus configured to align a radiation beam with a target receiving therapy, the radiation therapy apparatus comprising: a target support configured to immobilize the target;a fixation device connected to the target support;a radiation detector connected to the fixation device so as to be positioned in front of the target, the radiation detector configured to determine an intersection position of the radiation beam with the radiation detector in front of the target; anda correcting device configured to correct a relative position of the radiation beam and the target support, based on the intersection positiion determined by the radiation detector. 2. The apparatus of claim 1 wherein the radiation detector has a longitudinal axis and the fixation device holds the longitudinal axis of radiation detector parallel to the target support. 3. A radiation therapy apparatus configured to align a radiation beam with a target receiving therapy, the radiation therapy apparatus comprising: a target support configured to immobilize the target;a fixation device connected to the target support;a radiation detector connected to the fixation device so as to be positioned in front of the target, the radiation detector having a planar detection surface which is configured to determine in front of the target an intersection position of the radiation beam along X, Y axes of the planar detection surface of the radiation detector; anda correcting device configured to correct a relative position of the radiation beam and the target support, based on the intersection position determined on the X, Y axes of the planar detection surface of the radiation detector. 4. The apparatus of claim 3 wherein the planar surface of the radiation detector has a longitudinal axis and the fixation device holds the longitudinal axis of radiation detector parallel to the target support. 5. The apparatus of claim 4 wherein the fixation device is configured to hold the planar detection surface perpendicular to the radiation beam. 6. A radiation therapy apparatus configured to align a radiation beam with a target receiving therapy, the radiation therapy apparatus comprising: a target support configured to immobilize the target;a radiation detector comprising a radiation detector plane which intersects the radiation beam on the radiation detector plane in front of the target, the radiation detector configured to determine an intersection position of the radiation beam on the radiation detector plane in front of the target;a fixation device configured to hold the radiation detector in proximity to the target support and in front of the target and hold the radiation detector plane perpendicular to the radiation beam; anda correcting device configured to correct a relative position of the radiation beam and the target support, based on the determined intersection position on the radiation detector plane. 7. The apparatus of claim 6 wherein the radiation detector pla... | Ion Beam Applications S.A. | G21K005/08
US08515551 | B2 | 2013-08-20 | Diagnostic method and apparatus | 2010-11-10 | An electrical interferential device comprises a circuit for delivering electrical interferential energy into the body of a patient. A sensor detects a function of the autonomic nervous system of the patient and provides an output indicative of the response of the autonomic nervous system to the electrical interferential energy. A treatment regimen is selected which uses a combination of carrier and beat frequencies and electrode placement pattern that produce a desired response in the autonomic nervous system of the patient. In one embodiment of the invention, a diagnostic tool is used to determine which combination of carrier and beat frequencies are desirable and treatment is provided by a second electrical interferential device. In another embodiment of the invention, the same diagnostic tool or a portion of the same diagnostic tool may be used to treat the patient. | 1. A diagnostic system for identifying a preferred treatment for a patient comprising at least two pairs of electrodes; each of said pairs adapted to be temporarily attached to opposite limbs of a patient, respectively, at various selected locations and an electrical circuit connected to the electrodes for producing treatments comprising conducting a plurality of carrier frequencies and a plurality of beat frequencies through said electrodes for delivering therapeutic electrical interferential energy to the patient at said selected locations; said treatments producing a variety of responses of the autonomic nervous system wherein sympathetic and parasympathetic activity occurs in response to said treatments; a sensor for monitoring the autonomic nervous system of the patient for detecting said sympathetic and parasympathetic nervous system activity and providing a series of outputs indicative of said activity; and a recorder for recording the outputs whereby the electrodes may be repositioned and the carrier frequency can be changed and the beat frequency can be changed to produce a series of said treatments, at least some of which produce indications of autonomic nervous system activity that promotes augmenting the activity of the parasympathetic nervous system and decreasing the effects of the activity of the sympathetic nervous system thereby identifying an effective set of treatments to be applied to a patient. 2. The diagnostic system of claim 1 wherein the sensor is a heart rate sensor. 3. The diagnostic system of claim 1 wherein the sensor is a blood pressure monitor. 4. The diagnostic system of claim 1 wherein the sensor is a skin resistivity sensor. 5. The diagnostic system of claim 1 wherein the sensor is an iris pupil diameter measuring device. 6. The diagnostic system of claim 1 wherein the recorder comprises a data processor having a calculating device for manipulating the sensor outputs and providing a data processor output representative of the reaction of the patient to the electrical interferential energy. 7. The diagnostic system of claim 1 wherein the recorder comprises a data processor having a calculating device for manipulating the sensor outputs and providing a data processor output representative of sympathetic nervous system activity and parasympathetic nervous system activity. 8. The diagnostic system of claim 1 wherein the recorder comprises a data processor having a calculating device for manipulating the sensor outputs and providing a data processor output representing a ratio of sympathetic nervous system activity and parasympathetic nervous system activity. 9. The diagnostic system of claim 1 wherein the circuit includes a subcircuit for adjusting the carrier frequency and the system determines which combination of carrier and beat frequencies that produce a desired reaction of the autonomic nervous system. 10. The diagnostic system of claim 1 wherein the system comprises a data processor having a calculati... | - | A61N001/00
US09044170 | B2 | 2015-06-02 | Ballistocardiogram analysis method and device | 2009-07-06 | The present invention discloses a method and signal processing device (620) for analyzing a ballistocardiogram (BCG). The actual energy content of the cardiac part of the BCG in a predefined time window is determined and compared with a reference value. A difference between the actual energy content and the reference energy content exceeding a predefined threshold is indicative of the presence of an arrhythmia in the evaluated part of the BCG because the force exerted by an arrhythmic ventricular contraction, if at all present, is typically substantially smaller than the force exerted by a ventricular contraction of a regular heartbeat. This insight facilitates a more accurate detection of the presence of arrhythmias in a BCG compared to BCG detection schemes based on a heart rate analysis. | 1. A method of analyzing a ballistocardiogram to detect an arrhythmia comprising: selecting a first time window of interest of the ballistocardiogram, said first time window comprising a cardiac component signal incorporating a plurality of heart beats;defining a reference energy content including one of: determining a heart rate from said first time window and calculating the reference energy content by integrating an area under the cardiac component signal of a part of the cardiac component signal incorporating a single heart beat and multiplying the calculated energy content by the determined heart rate and first time window duration,ORdefining a further time window of the ballistocardiogram by displacing the first time window by a predefined amount of time and determining the reference energy content by integrating an area under the cardiac component signal over the further time window;calculating an actual energy content including one of: determining a heart rate from said first time window and calculating the actual energy content by integrating an area under the cardiac component signal of a part of the cardiac component signal incorporating a single heart beat and multiplying the calculated energy content by the determined heart rate and first time window duration,ORdefining a further time window of the ballistocardiogram by displacing the first time window by a predefined amount of time and determining the actual energy content by integrating an area under the cardiac component signal over the further time window,comparing the calculated actual energy content with the determined reference energy content; andgenerating an arrhythmia signal when a difference between the actual energy content and the reference energy content exceeds a predefined threshold. 2. The method according to claim 1, further comprising: generating the cardiac component signal by separating the cardiac component signal from a respiratory component signal of the ballistocardiogram. 3. The method according to claim 1, further comprising: determining a respiration pattern during said time window; andmodulating the cardiac component signal in accordance with the determined respiration pattern. 4. The method according to claim 3, wherein said modulating comprises increasing the amplitude of the signal during an inhalation process and reducing the amplitude of the signal during an exhalation process. 5. A signal processing device configured to: select a time window of interest of a ballistocardiogram, said time window of interest comprising a cardiac component signal incorporating a plurality of heart beats;determine a heart rate from said time window of interest;define a reference energy content by integrating an area under a part of the cardiac component signal incorporating a single regular heart beat and multiplying the calculated energy content by the determined heart rate and a duration of the time window;determine an actual energy content by integrating an area ... | Koninklijke Philips N.V. | A61B005/04 A61B005/11 A61B005/024 A61B005/113 A61B005/00 A61B005/0452
US08517938 | B2 | 2013-08-27 | Medical system having consumables monitoring | 2010-04-26 | The invention proposes a medical system which comprises a central control unit which is set up to display at least one diagnostic measured value using a display element. The medical system also comprises at least one invasive unit, wherein the invasive unit has at least one invasive consumable. The invasive consumable is set up to invasively intervene in a tissue of a patient. The invasive unit has at least one contactlessly readable electronic identifier for storing at least one piece of information. The central control unit is set up to electronically read the at least one piece of information in the electronic identifier. The central control unit and the invasive unit are functionally independent of one another such that the respective function of the central control unit and of the invasive unit can be performed independently of the presence of the respective other unit. | 1. A medical system comprising a central control unit configured to display a diagnostic measured value using a display element and an invasive unit comprising an invasive consumable, wherein the invasive consumable is configured to invasively intervene in a tissue of a patient, wherein the invasive unit has a contactlessly readable electronic identifier for storing a piece of information, wherein the central control unit is set up to electronically read said piece of information in the electronic identifier and perform an action responsive to said piece of information, wherein the central control unit and the invasive unit are functionally independent of one another such that the central control unit has a function which can be used by the user without the presence of the invasive unit and conversely the function of the invasive unit can be used by the user without the need for the central control unit to be present for this purpose, the respective function of the central control unit and of the invasive unit being capable of being performed independently of the presence of the respective other unit, wherein the central control unit has a measuring unit which is set up to use at least one test element to ascertain a concentration of an analyte in a body fluid, wherein the measuring unit is selected from the group consisting of an optical measuring unit and an electrochemical measuring unit. 2. The medical system as claimed in claim 1, wherein the invasive unit has a function selected from the group consisting of: a lancet function; a dosage function; an analytical function which is functionally independent of the measurement function of the central control unit; and a diagnostic function which is functionally independent of the measurement function of the central control unit. 3. The medical system as claimed in claim 2, wherein the invasive unit is configured to perform exclusively mechanical functions. 4. The medical system as claimed in claim 1, wherein the invasive unit further comprises a unit selected from the group consisting of: a peripheral device in the form of a lancing device, wherein the lancing device is designed to perforate a skin part of a patient using an invasive consumable in the form of a lancet; a peripheral device in the form of a medication pen, wherein the medication pen is designed to inject a dose of a medicament into a patient using an invasive consumable in the form of a cannula; a peripheral device in the form of a medication pump, wherein the medication pump is designed to inject a dose of a medicament into a patient using an invasive consumable in the form of a catheter; and a peripheral device in the form of a measuring unit, wherein the measuring unit is designed to use an invasive consumable comprising a subcutaneous sensor implanted in a body tissue of a patient, to ascertain a concentration of an analyte in a body fluid. 5. The medical system as claimed in claim 1, wherein the central control unit is set ... | Roche Diagnostics Operations, Inc. | A61B005/00 C12Q001/54 G08B001/08
US09341614 | B2 | 2016-05-17 | Apparatus and methods for taking blood glucose measurements and recommending insulin doses | 2010-11-03 | The present disclosure related to an apparatus that may be used for taking blood glucose measurements and providing individualized insulin dose recommendations wherein the apparatus is easy to use and facilitates improved diabetes control in patients. Also disclosed are related methods. | 1. An apparatus comprising; a test strip port for receiving a test strip;at least a first a computer-readable memory;a display screen;a processor operatively connected to the at least first computer-readable memory and the display screen, the processor configured to: A. cause the successive display of information corresponding to at least the following: (i) a patient's current blood glucose level measurement as determined from a sample of the patient's blood provided on a test strip;(ii) an event associated with the said current blood glucose level measurement; and(iii) a recommended insulin dose; andB. cause the continuous and simultaneous display of indicia corresponding to at least the patient's current blood glucose level measurement, an event associated said current blood glucose level measurement, and a recommended insulin dosage while simultaneously causing the display of said successive information; anda plurality of buttons that may be user-activated operatively connected to the processor, the plurality of buttons including a first button operative to enable a user to selectively cycle through the successive information displays on the display screen, and second and third buttons operative to enable a user to selectively alter the information displayed in one or more of the successive information displays on the display screen; andwherein the continuously displayed indicia corresponding to each of the successive information displays are positioned on the display screen in an alignment with the position of the first button, and the successively displayed information is positioned on the display screen in an alignment with the second and third buttons. 2. The apparatus of claim 1, wherein the processor is configured to cause the successive display of information corresponding to at least the following: (i) a patient's current blood glucose level measurement as determined from a sample of the patient's blood provided on a test strip;(ii) an event associated with the said current blood glucose level measurement;(iii) an estimate for the number of carbohydrates associated with the said event; and(iv) a recommended insulin dose. 3. The apparatus of claim 1, wherein the successive display on the display screen of the information display corresponding to an event associated with the said current blood glucose level measurement automatically succeeds the display of the information display corresponding to a patient's current blood glucose level measurement if the first button is not user-actuated within a predetermined period of time. 4. The apparatus of claim 1, wherein the apparatus is programmed to enable a user to selectively override the recommended insulin dose displayed on the display screen using one or more of the plurality of buttons. 5. The apparatus of claim 1, wherein the apparatus comprises a labeling area on which personalized identifying indicia may be provided. 6. The apparatus of claim 1, wherein the estimation of the numbe... | Hygieia, Inc. | G01N031/22 G01N033/487 G01N033/52 G01N027/327 G01N021/84
US20110137152 | A1 | 2011-06-09 | SYSTEM AND METHOD FOR COOLING COMPONENTS OF A SURGICAL NAVIGATION SYSTEM | 2009-12-03 | In one embodiment, a surgical navigation system is provided with a transmitter assembly that includes at least one coil for creating an electromagnetic field in the vicinity of a surgical operation. The surgical navigation system also includes a cooling unit thermally coupled to the transmitter assembly for dissipating heat generated by the transmitter assembly and a receiver assembly that includes at least one coil configured to generate signals induced by the field. The surgical navigation system further includes electronic circuitry coupled to and communicating with the transmitter assembly and the receiver assembly for calculating at least a position of the receiver coil. | 1. A surgical navigation system comprising: a transmitter assembly comprising at least one coil for creating an electromagnetic field in the vicinity of a surgical operation;a cooling unit thermally coupled to the transmitter assembly for dissipating heat generated by the transmitter assembly;a receiver assembly comprising at least one coil configured to generate signals induced by the field; andelectronic circuitry coupled to and communicating with the transmitter assembly and the receiver assembly for calculating at least a position of the receiver coil. 2. The surgical navigation system of claim 1, wherein the cooling unit comprises a plurality of semiconductor thermocouples. 3. The surgical navigation system of claim 2, wherein the semiconductor thermocouples are powered by a controller to generate both an endothermic end and an exothermic end on each of the semiconductor thermocouples. 4. The surgical navigation system of claim 3, wherein the endothermic ends of the semiconductor thermocouples are connected to the transmitter assembly and the exothermic ends of the semiconductor thermocouples are connected to a heat sink. 5. The surgical navigation system of claim 1, wherein the cooling unit comprises a heat sink that is air cooled or water-cooled. 6. The surgical navigation system of claim 1, further comprising a temperature sensor. 7. The surgical navigation system of claim 6, wherein the temperature sensor is electronically coupled to and communicating with a controller configured for regulating the temperature of the transmitter assembly. 8. The surgical navigation system of claim 7, wherein the controller is configured to regulate the temperature of the transmitter assembly independent of user input. 9. The surgical navigation system of claim 6, wherein the temperature sensor is coupled to circuitry comprising an output device configured to indicate the temperature of the transmitter assembly. 10. The surgical navigation system of claim 1, wherein the transmitter coil and the receiver coil are both operable to generate and sense an electromagnetic field. 11. The surgical navigation system of claim 1, wherein the transmitter assembly is coupled to a surgical table. 12. The surgical navigation system of claim 11, wherein the transmitter assembly is secured within the surgical table beneath a patient support. 13. The surgical navigation system of claim 11, wherein heat generated by the transmitter assembly is dissipated away from the transmitter assembly to outer edges of the surgical table. 14. The surgical navigation system of claim 1, wherein the transmitter assembly comprises a first layer that includes the at least one coil and a second layer that includes the cooling unit. 15. The surgical navigation system of claim 1, wherein the transmitter assembly comprises an electromagnetic shielding layer. 16. A surgical navigation system comprising: an imaging system operable to acquire image data;a display configured to displa... | General Electric Company | A61B005/05
US20110117020 | A1 | 2011-05-19 | IMAGING DENDRIMER NANOPROBES AND USES THEREOF | 2009-03-02 | This invention relates to imaging nanoprobe and methods of use thereof. Specifically, the invention relates to long-lived oxygen-insensitive nanoprobe comprising a lumisescent moeity with a long excited state lifetime, embedded in a dendrimer, wherein said dendrimer is internally cross-linked and has hydrophilic peripheral layer. | 1. A quencher-insensitive nanoprobe comprising: a lumisescent moeity with a long excited state lifetime, embedded in a dendrimer, having a hydrophilic peripheral layer, wherein emission lifetime of the luminescent moiety is longer than the diffusion time of the quencher to the luminescent core. 2. A quencher-insensitive nanoprobe comprising: a lumisescent moeity with a long excited state lifetime, embedded in a dendrimer, wherein emission lifetime of the luminescent moiety is longer than the diffusion time of the quencher to the luminescent core. 3. The nanoprobe of claim 2, wherein said dendrimer has a hydrophilic peripheral layer. 4. The nanoprobe of claim 1 or 2, wherein said luminescent moiety is a luminescent metalloporphyrin represented by any one of the compounds of Formula I-V. 5. The nanoprobe of claim 1 or 3, wherein said peripheral layer is poly(ethyleneglycol) (PEG), poly(lactide-co-glycolide) acid (PLGA), poly(L-lactide) acid (PLLA), poly(D-lactide) acid (PDLA), is polyvinylalcohol (PVOH) or polysorbate. 6. The nanoprobe of claim 1 or 2 wherein said dendrimer is internally cross-linked to prevent diffusional quenching by oxygen. 7. The nanoprobe of claim 1 or 3, wherein said polymeric layer is optionally connected to an agent capable of binding to a pre-determined target. 8. The nanoprobe of claim 1 or 3, having the general formula: Cπ-(AGn-R)m where: C is a metalloporphirine core;π is the number of π-extensions of the metalloporphirine core;AG is the dendritic aryl-glycine skeleton;n is the generation number;R is the peripheral unit; andm is the number of dendritic wedges attached to the core. 9. The nanoprobe of claim 8, wherein said metal of said metalloporphirin is palladium or platinum ions. 10. The nanoprobe of claim 8, wherein said n is between 1 and 5. 11. The nanoprobe of claim 8, wherein said m is between 6 and 8. 12. The nanoprobe of claim 8, wherein π is between 0 and 3. 13. The nanoprobe of claim 2, wherein the dendrimer has a hydroxyl terminal group. 14. A method of providing imaging contrast to a tissue of a subject, comprising the step of administering to the tissue the nanoprobe of claim 1; exposing the tissue to an electromagnetic radiation, thereby exciting the nanoprobe; and using an imaging device, imaging the tissue. 15. The method of claim 14, whereby excitation is in a form of a light pulse, which is shorter than, or equals to the luminescence lifetime of the luminescent moiety and the step of imaging is performed after a delay following the excitation pulse, but prior to the end of the luminescence decay, thus eliminating the background signal and increasing the signal-to-noise ratio. 16. A method for an in-vivo imaging of a tumor neovasculature in a subject comprising (i) administering a quencher-insensitive nanoprobe comprising: a lumisescent moeity with a long excited m state lifetime, embedded in a dendrimer, optionally having a hydrophilic peripheral layer, wherein emission lifetime of the... | THE TRUSTEES OF THE UNIVERSITY OF PENNSYLVANIA | A61K049/00 A61B005/05
US20110124975 | A1 | 2011-05-26 | Method for Medical Diagnosis Utilizing PDA Software Robots | 2011-02-04 | A method for diagnosing problems using piece removal techniques and which can include establishing a health status variable, collecting and storing information from a user in the health status variable tracking the health status variable, comparing the health status variable to a knowledge base, using the health status variable to reduce the size of the knowledge base, leveraging patterns and past history to formulate a search strategy, using the search strategy to collect further diagnostic information from the user, updating the health status variable, comparing the updated health status variable to the reduced knowledge base, diagnosing a problem based on the further diagnostic information, and selecting treatment for the problem. | 1. A method for diagnosing problems using piece removal techniques, comprising: establishing a health status variable, the health status variable containing a plurality of elements corresponding to body components;collecting information from a user regarding the status of the user's body components;storing the information from the user by altering the state of the corresponding elements in the health status variable;tracking the state of the plurality of elements in the health status variable to determine patterns and past history;comparing the health status variable to a knowledge base containing other health status variables;reducing the size of the knowledge base by removing health status variables and objects based on the collected information, patterns, and past history;leveraging patterns and past history associated with the health status variables remaining in the knowledge base to formulate a present search strategy;collecting further diagnostic information from the user consistent with the present search strategy;updating the health status variable associated with the user by altering the state of the corresponding elements in the health status variable;comparing the updated health status variable to the reduced knowledge base;diagnosing a problem based on the further diagnostic information; andselecting treatment for the problem based on the further diagnostic information, patterns, and past history 2. The method of claim 1, wherein the health status variable is a binary string. 3. The method of claim 2, wherein the binary string contains approximately 200 bits. 4. The method of claim 2, wherein each bit of the binary string corresponds to a body component of the user. 5. The method of claim 4, wherein collecting information from the user further comprises changing the bit in the health status variable corresponding to each body component in a way that reflects the state of that body component. 6. The method of claim 1, further comprising storing the health status variables associated with each diagnosis in order to optimize subsequent diagnoses of the same type. 7. A method for diagnosing problems using equilibrium-reestablishing techniques, comprising: establishing a health status variable, the health status variable containing a plurality of elements corresponding to body components;collecting information from a user regarding the status of the user's body components;storing the information from the user by altering the state of the corresponding elements in the health status variable;tracking the state of the plurality of elements in the health status variable to determine patterns and past history;identifying differences between previous versions of the user's health status variable and the user's current health status variable;collecting further diagnostic information from the user regarding those differences;updating the health status variable associated with the user by altering the state of the corresponding elements in t... | - | A61B005/00
US20110152632 | A1 | 2011-06-23 | Universal Body Sensor Network | 2009-08-06 | A sensor module (2A00, 2B00, 2C00) for an on-body sensor (1A00, 1B00) is provided. The sensor (1A00, 1B00) is intended for connecting to the sensor module (2A00, 2B00, 2C00) and for transmitting an analog value of a patient (18). The sensor module (2A00, 2B00, 2C00) comprises a converter means, a storage means, a determining means, and a transmitting means. The converter means (37) converts the analog value of the patient (18) to its digital format whilst the storage means (40, 41) stores the digital value. The determining means (35) determines a parameter of the patient (18) using a value of the digital value to and the transmitting means (39) sends data regarding the parameter. | 1. A sensor module for a patient on-body sensor, the sensor module comprising a device for providing electrical power to a patient on-body sensor,a converter means for converting an analog patient on-body sensor signal to a digital signal,a storage means for storing a patient on-body sensor identifier information and for storing the digital signal,a determining means for determining a parameter of the patient using a value of the digital signal, anda transmitting means for sending data regarding the parameter and the patient on-body sensor identifier information. 2. A sensor module of claim 1, wherein the transmitting means is provided for sending a patient identifier information. 3. (canceled) 4. (canceled) 5. (canceled) 6. A sensor module of claim 1, wherein the sensor module further comprises a clock for recording a timing of the analog signal. 7. (canceled) 8. A sensor module of claim 1, wherein the transmitting means comprises a means for wired transmission of the data regarding the parameter. 9. A sensor module of claim 1, wherein the transmitting means comprises a means for wireless transmission of the data regarding the parameter. 10. A monitoring device of a Body Area Network comprising a sensor module of claim 1 anda sensor that is connected to the sensor module,wherein the sensor receives electrical power from the sensor module and wherein a storage means of the sensor module comprises a distinguishing sensor identifier information. 11. A sensor hub for receiving a pre-determined signal from at least one monitoring device of claim 10, processing the pre-determined signal, and transmitting the processed signal to a wrist monitor. 12. A sensor hub of claim 11 wherein the pre-determined signal comprises data regarding a parameter of a patient. 13. A Body Area Network that comprises at least one monitoring device claim 10, anda sensor hub of claim 11 that is connected to the at least one monitoring device.wherein a sensor module of the monitoring device sends a distinguishing sensor identifier information for a sensor of the monitoring device to the sensor hub. 14. A Body Area Network of claim 13, wherein the sensor hub comprises a summary means for providing a periodic summary of data regarding a parameter of a patient. 15. A Body Area Network of claim 14, wherein the sensor hub comprises a sending means for sending the periodic summary. 16. A method of operating a sensor module for an on-body sensor, the method comprising providing electrical power to the sensor,the sensor transmitting an analog signal of a patient to the sensor module,converting the analog signal to a digital signal,storing the digital signal,determining a parameter of the patient using a value of the digital signal,incorporating data regarding a sensor identifier into data regarding the parameter of the patient, andsending the data regarding the parameter of the patient. 17. (canceled) 18. (canceled) 19. (canceled) 20. A method of claim 19, wherein t... | E-VITAE PTE. LTD. | A61B005/00
US20110105890 | A1 | 2011-05-05 | MRI OPERATING METHOD | 2010-09-03 | A magnetic-resonance imager is operated by, within the time of a R-R interval of the heart, carrying out a preparation sequence for suppressing signal contributions from the blood is carried out, in particular by a saturation sequence. At least the first refocusing pulse is generated simultaneously with a layer-selective gradient magnetic field that acts orthogonally to the layer-selective gradient magnetic field at the time of the generation of the RF excitation pulse. In addition the measuring value acquisition and image generation takes place by means of subsampling the data space and/or partially sampling the data space. | 1. A method of operating a magnetic-resonance imager for spatially resolved spin resonance measurement on an object, in particular a living object, arranged in a static magnetic field B0, where an alignment of the spins of the object and a longitudinal net magnetization Mz along the magnetic field direction Z is generated, wherein by means of at least one radio-frequency excitation pulse in resonance a spin-flip about a desired flip angle is generated at which a transverse magnetization component Mxy is generated or changed whose T2* relaxation is acquired for carrying out susceptibility-weighted measurements, for which purpose a metrological recording of a plurality of spin-echo signals from a desired volume element of the object takes place, which signals are generated after a RF excitation pulse by a sequence of radio-frequency refocusing pulses that are equidistant from one another, wherein prior to their acquisition a susceptibility dependency is imprinted on the echo signals by an additional evolution time inserted between the radio-frequency excitation pulse and the first radio-frequency refocusing pulse for development of the dephasing generated by inhomogeneities of the magnetic field, of the transverse magnetization Mxy generated by the RF excitation pulse, and wherein the location of the metrologically acquired volume element of the object is determined by gradient magnetic fields that superpose the homogenous magnetic field B0 at least temporarily, and the start and/or preparation sequences of a measurement is/are synchronized to one or more detected forms of physiological movement, wherein within the time of a R-R interval of the heart a) a preparation sequence for suppressing signal contributions from the blood is carried out, in particular by a saturation sequence, andb) at least the first refocusing pulse (3) is generated simultaneously with a layer-selective gradient magnetic field (4) that acts orthogonally to the layer-selective gradient magnetic field (2) at the time of the generation of the RF excitation pulse (1), andc) the measuring value acquisition and image generation takes place by means of subsampling the data space and/or partially sampling the data space. 2. The method according to claim 1, wherein the layer-selective gradient magnetic fields (4), which are generated at the same time as the refocusing pulses (3), are arranged orthogonally to one another during at least two sequential refocusing pulses (3). 3. The method according to claim 1, wherein a diffusion weighting is performed by generating at least one bipolar gradient magnetic field (5) prior to a refocusing pulse or between the refocusing pulses (3), or two unipolar gradient magnetic fields (6) around a refocusing pulse (3). 4. The method according to claim 3, wherein the generation of the bipolar gradient magnetic field (5) takes place within the evolution time that follows after the RF excitation pulse (1). 5. The method according to claim 1, wherein ... | - | A61B005/055
US07935051 | B2 | 2011-05-03 | Surgical access system and related methods | 2009-04-22 | A system for accessing a surgical target site and related methods, involving an initial distraction system for creating an initial distraction corridor, and an assembly capable of distracting from the initial distraction corridor to a secondary distraction corridor and thereafter sequentially receiving a plurality of retractor blades for retracting from the secondary distraction corridor to thereby create an operative corridor to the surgical target site, both of which may be equipped with one or more electrodes for use in detecting the existence of (and optionally the distance and/or direction to) neural structures before, during, and after the establishment of an operative corridor to a surgical target site. | 1. A system for accessing a surgical target site, comprising: a distraction assembly to create a tissue distraction corridor in a lateral, trans-psoas path to a targeted spinal site, wherein said distraction assembly includes an elongate inner element and first and second dilators, the elongate inner element being positionable in a lumen of the first dilator and at least one of the elongate inner element and the first dilator includes at least one stimulation electrode, and the second dilator being advanceable over the first dilator to the targeted spinal site;a retraction assembly to retract tissue from said tissue distraction corridor to thereby form an operative corridor in the lateral, trans-psoas path to a targeted spinal site, the retraction assembly being guidable along at least one dilator of the distraction assembly to said targeted spinal site, wherein the retraction assembly includes a plurality of retractor blades for engaging the tissue to maintain the operative corridor in the lateral, trans-psoas path to a targeted spinal site; anda control unit to transmit an electrical stimulation signal to said at least one stimulation electrode and to receive an electromyographic (EMG) sensor signal of a neuromuscular response of a nerve depolarized by said stimulation signal, the control unit including a display device with a graphical user interface that displays information indicative of at least one of nerve proximity and nerve direction in response to the EMG sensor signal. 2. The system of claim 1, wherein said elongate inner element has a first stimulation electrode proximate to a distal tip of the elongate inner element. 3. The system of claim 2, wherein the elongate inner element comprises a K-wire that is slidably received by the first dilator. 4. The system of claim 2, wherein said first dilator includes at least one stimulation electrode proximate to a distal tip of the first dilator. 5. The system of claim 4, wherein said second dilator includes at least one stimulation electrode proximate to a distal tip of the second dilator. 6. The system of claim 5, wherein at least one of said retractor blades includes a stimulation electrode proximate to a distal tip of the retractor blade. 7. The system of claim 1, further comprising at least one EMG sensor electrode to sense the neuromuscular response of a leg muscle coupled to said depolarized nerve, the EMG sensor electrode being electrically coupled to the control unit so as to communicate the EMG sensor signal to the control unit. 8. The system of claim 1, wherein each retractor blade of the plurality of retractor blades are sequentially advanced along the lateral, trans-psoas path to the targeted spinal site. 9. The system of claim 1, further comprising a handle coupled to at least one of elongate inner element and the first dilator, the handle having at least one button for initiating the electrical stimulation signal transmitted to said at least one stimulation electrode. 10... | NuVasive, Inc. | A61B001/32 A61B005/04
US07942828 | B2 | 2011-05-17 | Method for determining fluctuation in attentional state and overall attentional state | 2004-02-03 | The invention provides methods for determining shifts in the attentional state of a subject. These methods are useful for diagnosing subjects with a psychological or behavioral disorder. The invention also features methods for determining the effect of a therapy on the overall attentional state and shifts in the attentional state of a subject. | 1. A method for diagnosing a psychological, neurological, or behavioral disorder in a subject, said method comprising: (a) providing from said subject scoring data that has been collected by (i) presenting to said subject a sequence comprising target and nontarget stimuli during a continuous testing period; and (ii) scoring the response of said subject on the number of targets responded to and the number of nontargets responded to;(b) on the basis of said scoring data, (i) dividing said sequence into three or more epochs and classifying the attentional state of said subject during each epoch, (ii) calculating the number of shifts in the attentional state of said subject during said continuous testing period, (iii) identifying epochs in which said subject is in an impulsive attentional state, and (iv) calculating the amount of time said subject is in an impulsive attentional state during said continuous testing period; and(c) on the basis of the results of step (b), diagnosing said subject. 2. A method for diagnosing a psychological, neurological, or behavioral disorder in a subject, said method comprising: (a) presenting to said subject a sequence comprising target and nontarget stimuli during a continuous testing period, recording the responses of said subject to produce subject data, and transmitting said subject data to a computer for analysis;(b) receiving the results of said analysis, wherein said analysis comprises (i) scoring the response of said subject on the number of targets responded to and the number of nontargets responded to; and (ii) on the basis of said scoring data, (b1) dividing said sequence into three or more epochs and classifying the attentional state of said subject during each epoch, (b2) calculating the number of shifts in the attentional state of said subject during said continuous testing period, (b3) identifying epochs in which said subject is in an impulsive attentional state, and (b4) calculating the amount of time said subject is in an impulsive attentional state during said continuous testing period; and(c) on the basis of the analysis of step (b), diagnosing said subject. 3. The method of claim 1 or 2, wherein said stimuli are visual symbols selected from numbers, letters, shapes, or a combination thereof. 4. The method of claim 1 or 2, wherein said stimuli are auditory. 5. The method of claim 1 or 2, wherein said disorder is depression, an anxiety disorder, schizophrenia, drug addiction, an eating disorder, attention deficit disorder, attention deficit hyperactivity disorder, a learning disorder, Alzheimer's disease, dementia, epilepsy, stroke, or traumatic brain injury. 6. The method of claim 5, wherein said disorder is attention deficit hyperactivity disorder. 7. The method of claim 1 or 2, wherein step (c) further comprises comparing the results of step (b) in subjects diagnosed with a disorder and subjects not diagnosed with said disorder. 8. The method of claim 1 or 2, wherein said subject is involve... | The McLean Hospital Corporation | A61B005/00
US07957791 | B2 | 2011-06-07 | Multi-beat integration for cardiac mapping | 2008-06-13 | A non-contact cardiac mapping method is disclosed that includes: (i) inserting a catheter into a heart cavity having an endocardium surface, the catheter including multiple, spatially distributed electrodes; (ii) measuring signals at the catheter electrodes in response to electrical activity in the heart cavity with the catheter spaced from the endocardium surface; and (iii) determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface. Related systems and computer programs are also disclosed. | 1. A method comprising: inserting a catheter into a heart cavity having an endocardium surface, the catheter comprising multiple, spatially distributed electrodes;measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;integrating information derived from the signals for common phases of the electrical heart beat cycles;determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface; anddisplaying at least a portion of the determined physiological information;wherein the signals are measured during multiple different portions of a heart beat cycle and wherein the physiological information is determined, at least in part, by accounting for changes in the location of the catheter electrodes with respect to the endocardium surface during the multiple different portions of a heart beat cycle. 2. The method of claim 1, further comprising: for each of the multiple different portions of a heart beat cycle, determining a position of the catheter electrodes with respect to the endocardium surface and measuring signals at the catheter electrodes in response to electrical activity in the heart cavity,wherein the determination of the physiological information at the multiple locations of the endocardium surface is based further on the positions of the catheter electrodes and the measured signals at the different catheter positions. 3. The method of claim 1, wherein accounting for changes in the location of the catheter electrodes with respect to the endocardium surface during the multiple different portions of a heart beat cycle comprises accounting for changes in the shape of the endocardium surface during the multiple different portions of a heart beat cycle. 4. The method of claim 1, wherein integrating information derived from the signals for common phases of the electrical heart beat cycles comprises separately integrating information derived from the signals for systole and end diastole phases of the heart beat cycle. 5. The method of claim 1, wherein the determined physiological information includes at least one of: electrical potential values at the multiple locations of endocardium surface at each of one or more different phases of the heart beat cycle,a maximum voltage amplitude for one or more heart beat cycles at different ones of the endocardium surface locations,a frequency representation of electrical activity at multiple locations of the endocardium surface during the heart beat cycle, wherein the frequency representation is computed based on electrical potential values at the multiple locations of endocardium surface at different phases of the heart beat cycle, andan activation time for each of different locations of the endocardium surface. 6. The method of claim 1, wherein displaying at least a portion of the determined physiological information comprises displaying se... | Rhythmin Medical, Inc. | A61B005/02
US07951137 | B2 | 2011-05-31 | Method and system for terminating an atrial arrhythmia | 2005-10-06 | A method and system is provided for responding, from internally within a patient, to an atrial arrhythmia in a heart including measuring from within the patient at least one electrocardiogram characteristic indicative of the atrial arrhythmia, and controlling from within the patient drug therapy delivery to the patient responsive to measuring the at least one electrocardiogram characteristic. Drug therapy is initiated to the patient responsive to measuring the at least one electrocardiogram characteristic. According to one aspect of the present invention, the drug therapy is staged within the patient prior to measuring the at least one electrocardiogram characteristic. According to another example embodiment, the heart is paced from within the patient at a predefined rate responsive to measuring the at least one electrocardiogram characteristic, pacing occurring alone, or in combination with drug therapy. | 1. An implantable system for delivering a therapy to a patient having a heart, comprising: an inplantable recorder and monitor adapted to monitor an electrical signal generated by the heart of the patient and identify a characteristic of the electrical signal generated by the heart of the patient indicative of an atrial arrhythmia a charateristic of the electrical signal generated by the heart predictive of a ventricular proarrhythmia; anda drug pump, operatively coupled to the implantable recorder and monitor, adapted to deliver a drug therapy to the patient based, at least in part, on the characteristic of the electrical signal indicative of the atrial arrhythmia and the characteristic of the electrical signal predictive of the ventricular proarrhythmia. 2. The implantable system of claim 1, wherein the drug pump is adapted to deliver the drug therapy to the patient based, at least in part, on the characteristic of the electrical signal indicative of the atrial arrhythmia and if the characteristic of the electrical signal predictive of the proarrhythmia indicates an unlikelihood of a future proarrhythmia. 3. The implantable system of claim 1 wherein the drug therapy comprises an anti-arrhythmia drug. 4. The implantable system of claim 3 wherein the anti-arrhythmia drug comprises Ibutilide. 5. The implantable system of claim 1 wherein the electrical signal generated by the heart defines, at least in part, a Q-T interval, and wherein the characteristic of the electrical signal predictive of the ventricular proarrhythmia is a time duration of the Q-T interval. 6. The implantable system of claim 5, wherein the time duration of the Q-T interval is outside of a selected range. 7. The implantable system of claim 5 wherein the time duration of the Q-T interval is at least 480 milliseconds. 8. The implantable system of claim 1 wherein the electrical signal generated by the heart defines, at least in part, a Q-T interval, and wherein the characteristic of the electrical signal predictive of the ventricular proarrhythmia is variability of time durations of individual ones of a plurality of Q-T intervals. 9. An implantable system for delivering a therapy to a patient having a heart, comprising: an implantable recorder and monitor adapted to monitor an electrical signal generated by the heart of the patient and identify a characteristic of the electrical signal generated by the heart of the patient indicative of an atrial arrhythmia and a characteristic of the electrical signal generated by the heart indicative of a ventricular proarrhythmia;a drug pump, operatively coupled to the implantable recorder and monitor, adapted to deliver a drug therapy to the patient based, at least in part, on the characteristic of the electrical signal indicative of the atrial arrhythmia; andan implantable pacing device, operatively coupled to the implantable recorder and monitor adapted to deliver a pacing therapy to the heart of the patient based, at least in part, o... | Medtronic, Inc. | A61K009/22 A61N001/00 A61B005/04
US07970453 | B2 | 2011-06-28 | Reception coil maintaining element for an MR-imaging system | 2006-04-28 | A reception coil maintaining element for an MR-imaging system includes two phase-controlled multi-element arrays of an RF multi-channel surface reception coil. Both of these phase-controlled, multi-element arrays have a reception surface including a center having a longitudinal axis that crosses that center. Each of these multi-element arrays is arranged on an arm which can pivot about a common pivotable axis. A vertical first plane is parallel to this pivotable axis. A vertical left plane extends along the longitudinal axis of the center of the left reception surface of the left multi-element array. A vertical right plane extends along a longitudinal axis of the center of the right reception surface of the right multi-element array. Both of these vertical left and right planes have approximately the same spacing with respect to the vertical first plane of the pivotable axis where the reception coil maintaining element in accordance with the present invention is in its operational state. | 1. A receiving coil support comprising: a first phased multi-element coil array in a first radiofrequency multi-channel support receiving coil;a first housing for said first phased multi-element coil array;a second phased multi-element coil array in a second radiofrequency multi-channel support receiving coil;a second housing for said second phased multi-element coil array, said first and second coil arrays being adapted to enable a parallel imaging process for an magnetic resonance imaging system of an magnetic resonance tomograph and with an operational main magnetic field direction;a first receiving surface for said first phased multi-element coil array and having a first center through which a first longitudinal axis passes;a second receiving surface for said second phased multi-element coil array and having a second center through which a second longitudinal axis passes;a receiving coil support holder including having a first arm free end and a second arm having a second arm free end;a first pivot joint securing said first housing to said free end of said first arm of said receiving coil support holder; anda second pivot joint securing said second housing to said free end of said second arm of said receiving coil support holder. 2. The receiving coil support of claim 1 wherein each of said first and second arms of said receiving coil support element is a leaf spring. 3. The receiving coil support of claim 2 wherein each of said leaf springs is flexible to bending. 4. The receiving coil support of claim 1 wherein said first arm and said second arm each has an arm beginning which first arm beginning is remote from each said first arm pivot joint connection to said first housing, and which second arm beginning is remote from said second arm pivot joint connection to said second housing, said first arm beginning and said second arm beginning being positively connected to each other. 5. The receiving coil support of claim 4 wherein said positive connection is a pivot point. 6. The receiving coil support of claim 1 wherein said first housing and said second housing are positioned with said first receiving surface and said second receiving surface being opposite to each other. 7. The receiving coil support of claim 1 further including a base plate adapted to be positioned in the interior of an magnetic resonance tomograph, said base plate including mounting means for said receiving coil support element. 8. The receiving coil support of claim 7 wherein said mounting means include at least one of straight guides and sliding key connections. 9. The receiving coil support of claim 8 further including a stationary part of said at least one of said straight guide and said sliding leg connection on said base plate. 10. The receiving coil support of claim 9 wherein said stationary part of said at least one of said straight guide and said sliding leg connection is oriented one of transverse to and parallel to said operational main magnetic field di... | - | A61B005/05
US09060807 | B2 | 2015-06-23 | Light marker for indwelt use in patient's body and for identification of lesions inside patients tubular organ using the light marker | 2011-03-08 | A lesion identification system for surgical operations has a light marker adapted for placement in the vicinity of a lesion inside an organ of an organism, and a location-identifying device detecting a light emitted from the light marker at an outside of the organ for identifying a location of the lesion. The marker includes a light emitter and an engagement member associated with the emitter to be engageable with a wall of the organ. The engagement member includes a clip. The light emitter emits light with a wavelength in a near-infrared range. The device includes an endoscope comprising an inserter section having an image pickup element picking up a reflected light, involving the light from the light marker, at an outside of the organ, and an image pickup unit allowing the reflected light, picked by the image pickup element, to be generated as an image for display on a monitor. | 1. A method of identifying a lesion which is present inside a tubular organ of a patient for surgical operation, the method comprising the steps of: delivering, using applicator, a light marker into a tubular organ of a body of a patient through a natural opening of the patient and introducing, from inside a tubular wall of the tubular organ, the light marker to a site underneath a mucous membrane, the site being formed by partially incising the mucous membrane of the tubular wall from inside the tubular organ within the body, the site being in the vicinity of the lesion, the introduced light marker being indwelt in the vicinity of the lesion of the tubular organ, wherein the light marker is composed of a plurality of beady fluorescent balls formed to emit light whose wavelengths belong, at least partly, to a range of wavelengths of near-infrared range; andinserting an endoscope into the body of the patient and locating a tip of the endoscope outside the tubular organ within the body, the tip of the endoscope facing the lesion via the wall of the tubular organ;making the endoscope radiate visible light toward the wall of the tubular organ from outside the wall;picking up, using the endoscope, an image which is based on both reflected light of the light radiated from the endoscope and the light emitted from the fluorescent substance of the indwelt light marker and transmitted through the wall at an outside position of the wall of the tubular organ; andidentifying a location of the indwelt light marker based on a position of the light on the image by processing both the visible reflected light and the light emitted from the fluorescent substance of the indwelt light marker into two types of image signals, respectively, and displaying on a monitor the two types of image signals as an image in which an image processed based on one of the two types of image signals processed with the light emitted from the fluorescent substance of the indwelt light marker is superposed on an image processed based on the other of the two types of image signals processed with the visible reflected light, the indwelt light marker positionally indicating an area of the tubular organ in which the lesion is present. 2. The method of claim 1, wherein the beady fluorescent balls each containing the fluorescent substance and emitting the light. 3. The method of claim 1, wherein the applicator contains therein the beady fluorescent balls beforehand and discharging the beady fluorescent balls at the site. 4. The method of claim 1, comprising a step of excising, after the surgical operation, the lesion of the tubular organ together with the indwelt light marker. 5. The method of claim 4, wherein the delivering step includes a step of forming the site by partially incising the mucous membrane of the wall of the tubular organ from inside the tubular organ within the body before delivering the light marker to the site. 6. The method of claim 1, wherein the delivering step inclu... | OLYMPUS CORPORATION | A61B005/06 A61B019/00
US20110112424 | A1 | 2011-05-12 | DETERMINING FUNCTIONAL RESIDUAL LUNG CAPACITY | 2009-04-23 | Determining functional residual lung capacity (FRC) by changing a subject's inspirium FiO2 by a predetermined amount, and a) for each breath in a series of breaths subsequent to changing the FiO2, determining expiratory tidal volume, determining expiratory fractional N2 tidal volume, multiplying the expiratory tidal volume by an absolute difference between the expiratory fractional N2 tidal volume of the breath and that of an immediately preceding breath for a first multiplication result, dividing the first multiplication result by the sum of the differences for a first division result, and multiplying the fractional N2 tidal volume by the sum of the first division results of the breaths for a second multiplication result, and b) dividing the sum of the second multiplication results of the breaths by the absolute difference between the fractional N2 tidal volume of the first and last breaths to produce a measurement of the subject's FRC. | 1. A method for determining the functional residual lung capacity of a subject, the method comprising: changing the FiO2 of a subject's inspirium by a predetermined amount;for each breath in a series of breaths of said subject subsequent to changing said FiO2, determining an expiratory tidal volume measurement value of said breath,determining an expiratory fractional N2 tidal volume measurement value of said breath,multiplying said expiratory tidal volume measurement value of said breath by an absolute difference between said expiratory fractional N2 tidal volume measurement value of said breath and that of a breath immediately preceding said breath, thereby yielding a first multiplication result,dividing said first multiplication result by the sum of said absolute differences of each of said breaths, thereby yielding a first division result, andmultiplying said expiratory fractional N2 tidal volume measurement value of said breath by the sum of said first division results of each of said breaths, thereby yielding a second multiplication result; anddividing the sum of said second multiplication results of each of said breaths by the absolute difference between said expiratory fractional N2 tidal volume measurement values of the first and last breaths in said series of breaths, thereby producing a functional residual lung capacity measurement of said subject. 2. A method according to claim 1 wherein said changing step comprises increasing said FiO2. 3. A method according to claim 1 wherein said changing step comprises decreasing said FiO2. 4. A method according to claim 1 wherein said changing step comprises changing said FiO2 by an amount that is within the range of about 20% to about 25% of total inspired volume of said subject. 5. A method according to claim 1 wherein said changing step comprises changing said FiO2 in accordance with a single step function. 6. A method according to claim 1 and further comprising: determining a fractional expiratory CO2 tidal volume of expirium of said subject; anddetermining a fractional expiratory O2 tidal volume of expirium of said subject,wherein said step of determining said fractional expiratory N2 tidal volume comprises determining said fractional expiratory N2 tidal volume as a function of said O2 and CO2 fractional expiratory O2 tidal volumes. 7. A method according to claim 6 wherein said step of determining said fractional expiratory CO2 tidal volume comprises determining prior to changing said FiO2. 8. A method according to claim 6 wherein said step of determining said fractional expiratory CO2 tidal volume comprises determining said fractional expiratory CO2 tidal volume separately for each of said breath in said series of breaths. 9. A method according to claim 1 wherein said determining steps comprise determining until any of said expiratory fractional tidal volumes reaches a steady state. 10. A method according to claim 9 wherein said determining steps comprise determining until consecuti... | - | A61B005/091 A61B005/08 A61M016/00
US07963922 | B2 | 2011-06-21 | Volume depletion detection | 2006-04-28 | Detection of volume depletion, particularly after an incidence of volume overload is disclosed. Various methods, systems, and devices are disclosed that sense and analyze a physiological parameter related to a patient's fluid level in order to warn patients of potentially dangerous volume depletion conditions while minimizing false notifications. | 1. A computer-readable medium programmed with instructions for performing a method in an implantable medical device, the medium comprising instructions for causing a programmable processor to: intermittently measure a physiological parameter related to a patient's fluid level to determine measured values;calculate long term averages and short term averages of the measured values; andcompare the short term averages to the long term averages to detect the presence of a volume depletion condition, the detection including calculating a cumulative sum of the difference between the short term averages and the long term averages. 2. A computer-readable medium according to claim 1, further comprising instructions to compare the short term averages and the long term averages to detect the presence of a volume overload condition, the detection of the volume depletion occurring after detecting the presence of the volume overload. 3. A computer-readable medium according to claim 1, wherein the detecting of the volume depletion condition includes comparing the cumulative sum to a predetermined threshold value. 4. A computer-readable medium according to claim 3, further including instructions to reset the cumulative sum to zero when the values of a predetermined number of the short term averages exceed the values of associated long term averages. 5. A computer-readable medium according to claim 1, further including instructions to stop the detection of the volume depletion condition when the values of a predetermined number of the short term averages exceed the values of associated long term averages. 6. A computer-readable medium according to claim 5, wherein the predetermined number of the short term averages are consecutive. 7. A computer-readable medium according to 1, wherein the intermittent measurements include periodic measurements. 8. A system for detecting volume depletion in a patient comprising: an implantable sensor for measuring a physiological parameter related to a patient's fluid level and producing a representative signal thereof;a processor within an implantable medical device operatively connected to the sensor and receiving the sensor signal, the processor calculating long and short term averages of the sensor signal and comparing the long and short term averages to detect the presence of a volume overload condition, the processor comparing the long and short term averages to detect the presence of a volume depletion condition after detection of the presence of the volume overload condition; andan indicator operatively connected to the processor that indicates the detection of the volume depletion condition. 9. A system according to claim 8, wherein the implantable sensor includes a blood pressure sensor that measures intracardiac blood pressure. 10. A system according to claim 9, wherein the blood pressure sensor measures one of at least right ventricular systolic pressure (RVSP) and right ventricular diastolic pressures (RVDP). | Medtronic, Inc. | A61B005/02
US07935055 | B2 | 2011-05-03 | System and method of measuring disease severity of a patient before, during and after treatment | 2003-09-19 | A system for obtaining serial biochemical, anatomical or physiological in vivo measurements of disease from one or more medical images of a patient before, during and after treatment, and measuring extent and severity of the disease is provided. First anatomical and functional image data sets are acquired, and form a first co-registered composite image data set. At least a volume of interest (VOI) within the first co-registered composite image data set is identified. The first co-registered composite image data set including the VOI is qualitatively and quantitatively analyzed to determine extent and severity of the disease. Second anatomical and functional image data sets are acquired, and form a second co-registered composite image data set. A global, rigid registration is performed on the first and second anatomical image data sets, such that the first and second functional image data sets are also globally registered. | 1. A system for measuring extent and severity of disease in a patient, comprising: an imaging device for acquiring a plurality of anatomical image data sets and a plurality of functional image data sets of the patient, said plurality of anatomical image data sets and said plurality of functional image data sets form a plurality of co-registered composite image data sets;a user device for identifying at least a volume of interest (VOI) within a co-registered composite image data set of said plurality of co-registered composite image data sets, and identifying at least a VOI within a globally registered image data set of a plurality of globally registered image data sets using the identified VOI within said co-registered composite image data set; anda computer system for performing the following steps:qualitatively and quantitatively analyzing a first co-registered composite image data set of said plurality of co-registered composite image data sets including the identified VOI to determine extent and severity of the disease;performing a global, rigid registration of a first anatomical image data set and a second anatomical image data set of said plurality of anatomical image data sets, such that a first functional image data set and a second functional image data set of said plurality of functional image data sets are also globally registered;performing a local, non-rigid registration of the VOI within said first co-registered composite image data set and the VOI within the globally registered image data set, thereby producing a first co-registered serial image data set, said local, non-rigid registration being more precise than said global, rigid registration; andqualitatively and quantitatively analyzing said first co-registered serial image data set including the VOIs to determine severity of the disease and/or response to treatment of the patient. 2. The system of claim 1, wherein said imaging device includes a computed tomography (CT) scanner and a single photon emission computed tomography (SPECT) scanner. 3. The system of claim 2, wherein said CT scanner is housed in a single CT gantry, and said SPECT scanner is housed in a single SPECT gantry. 4. The system of claim 3, wherein a single patient bed is movable between said CT gantry and said SPECT gantry. 5. The system of claim 3, wherein said CT scanner and said SPECT scanner is housed in a combined CT and SPECT gantry. 6. The system of claim 5, wherein a single patient bed is movable between said CT and SPECT gantry. 7. The system of claim 1, wherein said first anatomical image data set and said first functional image data set are acquired before treatment of the patient, and said second anatomical image data set and said second functional image data set are acquired during treatment of the patient. 8. The system of claim 1, wherein said first anatomical image data set and said first functional image data set are acquired before treatment of the patient, and said second anatomical i... | Siemens Medical Solutions USA, Inc. | A61B005/00
US07970448 | B2 | 2011-06-28 | Device and method for determining analyte levels | 2010-04-19 | Devices and methods for determining analyte levels are described. The devices and methods allow for the implantation of analyte-monitoring devices, such as glucose monitoring devices, that result in the delivery of a dependable flow of blood to deliver sample to the implanted device. The devices comprise a unique microarchitectural arrangement in the sensor region that allows accurate data to be obtained over long periods of time. | 1. A system for processing sensor data from a continuous analyte sensor, comprising: a sensor comprising an electrode configured for implantation in a subcutaneous tissue of a host, wherein the sensor configured to continuously measure an analyte concentration, whereby sensor data associated with the subcutaneous analyte concentration is generated; andelectronic circuitry operably connected to the sensor and configured to calibrate the sensor data during a useful life of the system using pre-implant calibration information;wherein the system is configured to provide a substantially accurate measurement of the analyte concentration in the host such that at least 95% of calibrated sensor analyte values are within 25% of actual analyte values as determined by analysis of blood obtained during the useful life of the system. 2. The system of claim 1, wherein the pre-implant calibration information comprises a value obtained from an in vitro calibration. 3. The system of claim 2, wherein the in vitro calibration is obtained by measuring at least two different analyte concentrations. 4. The system of claim 2, wherein the in vitro calibration is obtained at about 37 degrees C. 5. The system of claim 1, wherein the analyte is glucose, and wherein the sensor is configured to respond substantially linearly to changes in glucose concentration at glucose concentrations ranging from about 40 mg/dL to about 500 mg/dL. 6. The system of claim 1, wherein the system is configured to provide a stable sensor offset during the useful life of the sensor. 7. The system of claim 1, wherein the electronic circuitry is configured to determine trends in analyte concentration. 8. The system of claim 1, wherein the electronic circuitry is configured to physically couple the sensor. 9. The system of claim 1, wherein the electronic circuitry is configured to couple, via radio telemetry, the sensor. 10. The system of claim 1, wherein the analyte is glucose. 11. The system of claim 1, wherein the useful life of the system is more than about 3 days of implantation. 12. The system of claim 1, wherein the useful life of the system is more than about 5 days of implantation. 13. The system of claim 1, wherein the useful life of the system is more than about 7 days of implantation. 14. The system of claim 1, wherein the useful life of the system is more than about one month of implantation. 15. The system of claim 1, wherein the electronic circuitry is configured to perform recalibration during the useful life of the sensor. 16. The system of claim 15, wherein the electronic circuitry is configured to perform recalibration throughout the useful life and the sensor. 17. The system of claim 16, wherein the recalibration is performed periodically. 18. The system of claim 16, wherein the recalibration is based on reference data. 19. The system of claim 1, wherein the electronic circuitry is configured to perform a calibration check during the useful life of the sensor. 2... | DexCom, Inc. | A61B005/05
US07959585 | B2 | 2011-06-14 | Device and method for the time-based analysis of the agitation of a bedridden patient | 2007-09-11 | Device for the time-based analysis of the agitation of a bedridden patient includes elements (4) for acquiring the patient's shouts and elements (5) for acquiring the patient's movements, the acquisition elements being connected to storage elements (9) suitable for storing the acquired data accompanied by a time stamping of the latter, the storage elements being connected to analysis and presentation elements (10) suitable for providing care staff with a representation of the frequency and duration of the patient's shouts and movements which are representative of the patient's agitation over a determined time span, characterized in that the elements for acquiring shouts include a bandpass filter (6) suitable for isolating the shouts from speech and ambient noise. | 1. A device for a time-based analysis of the agitation of a bedridden patient comprising: means (4) for acquiring data related to the patient's shouts; andmeans (3) for acquiring data related to the patient's movements, the acquisition means being connected to storage means (9) suitable for storing the acquired data accompanied by time-stamping (8) of said acquired data, connected to analysis and presentation means (10, 11) suitable for providing care staff with a representation of frequency and duration of the patient's shouts and movements which are representative of the patient's agitation over a determined time span,wherein the means (4) for acquiring data related to the shouts comprise a band-pass filter (6) suitable for isolating shouts from speech and ambient noise, the band-pass filter comprising pass-band selection means between two predetermined bandwidths, said two predetermined bandwidths having a bandwidth of approximately 1500 Hz centered on 2500 Hz for the first and a bandwidth of approximately 900 Hz centered on 1500 Hz for the second, andthe means for acquiring data related to the movements comprise a sensor of the vibrations of a patient's bed and an intensity filter adapted so that only the vibrations having an intensity above a predetermined threshold are stored in the storage means. 2. The analysis device according to claim 1, wherein the band-pass filter is a digital filter based on a Z-transform technique. 3. A method for a time-based analysis of the agitation of a bedridden patient comprising the steps of: acquisition (24, 26) of data related to the patient's shouts;acquisition (20, 22) of data related to the patient's movements,storage (28) of the data acquired accompanied by time-stamping of said acquired data; andanalysis and presentation (30) of the time-stamped data for providing care staff with a representation of frequency and duration of the patient's shouts and movements which are representative of the patient's agitation over a determined time period,wherein the acquisition of data related to the shouts comprises a band-pass filtering substep suitable for isolating shouts from speech and ambient noise, the band-pass filtering comprising a preliminary step of pass-band selection between two predetermined bandwidths, said two predetermined bandwidths having a bandwidth of approximately 1500 Hz centered on 2500 Hz for the first and a bandwidth of approximately 900 Hz centered on 1500 Hz for the second, andwherein the step of acquisition of data related to the movements comprises a sub-step of acquisition of vibrations of a patient's bed, followed by an intensity filtering sub-step adapted such that only the vibrations having an intensity above a predetermined threshold are stored. | INSERM (Institut National de la Santé et de la Recherche Médicale) | A61B005/103 A61B005/117
US07951094 | B2 | 2011-05-31 | Shapeable intraluminal device and method therefor | 2007-12-04 | A hollow medical device includes a coil member having a distal end, a proximal end, a plurality of turns extending therebetween, and a lumen extending through the turns. The coil member further includes at least one weld extending from one turn to an adjacent turn. | 1. A hollow medical device comprising a coil member comprising a distal end, a proximal end, a plurality of turns extending therebetween, and a lumen extending through said turns, said coil member comprising a plurality of intersecting welds extending from one turn to an adjacent turn to provide stiffness to said coil, and plurality of non-intersecting welds extending from one turn to an adjacent turn to from at least one ductile member, said coil member configured to be plastically deformed by bending said ductile member at the plurality of non-intersecting welds, such that said ductile member retains a bent shape of said coil member, said lumen configured to house at least one of an optical fiber, a balloon inflation tube, a drug delivery tube, and a reagent delivery tube. 2. A hollow medical device in accordance with claim 1 wherein said device comprises one of a catheter and a guide wire. 3. A hollow medical device in accordance with claim 1 wherein said coil member comprises at least one of a helical coil of wire and a wound ribbon. 4. A hollow medical device in accordance with claim 1 wherein said coil member comprises metal. 5. A hollow medical device in accordance with claim 4 wherein said metal is at least one of steel and nitinol. 6. A hollow medical device in accordance with claim 1 wherein said coil member includes an axis extending from said proximal end to said distal end, said plurality of non-intersecting welds substantially parallel to said axis. 7. A hollow medical device in accordance with claim 1 wherein at least one of said plurality of non-intersecting welds connects at least three adjacent turns. 8. A hollow medical device in accordance with claim 1 wherein said device further comprises a distal tip, said distal end connected to said distal tip. 9. A method for shaping an intraluminal device having a coiled member configured to house at least one of an optical fiber, a balloon inflation tube, a drug delivery tube, and a reagent delivery tube, said method comprising: forming the coil member including a distal end, a proximal end, a plurality of turns extending between the proximal end and the distal end, and a lumen extending though the turns; welding a plurality of intersecting welds that extend from one turn to an adjacent turn to provide stiffness to said coil; welding a plurality of non-intersecting welds that extend from one turn to an adjacent turn to form a ductile member; and plastically bending the ductile member at the plurality of non-intersecting welds into a bent shape suitable to traverse a body lumen, such that the ductile member retains the bent shape of the coil member. 10. A method in accordance with claim 9 wherein plastically bending the coil member comprises plastically deforming the coil member at the plurality of non-intersecting welds. 11. A method in accordance with claim 9 wherein welding comprises at least one of welding and soldering the at least one turn of the coil member to an adjacen... | The Spectranetics Corporation | A61B005/00 A61M025/00
US08346338 | B2 | 2013-01-01 | System and methods for replacing signal artifacts in a glucose sensor data stream | 2011-01-27 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. A system for processing data signals obtained from a glucose sensor, the system comprising: a receiving module that receives a signal from a glucose sensor over a period of time;a pressure detector configured to generate a signal indicative of a measured pressure within the glucose sensor; anda processor module comprising programming that when executed by a processor configures the processor module to detect transient non-glucose related signal artifacts based at least in part on the signal from the glucose sensor and the signal from the pressure detector and output to an external device or display estimated glucose values in response to the detection of the detected transient non-glucose related signal artifacts. 2. The system of claim 1, further comprising a glucose sensor. 3. The system of claim 1, wherein the programming configured to output estimated glucose values during transient non-glucose related signal artifacts is configured to display data representative of the estimated glucose values, including at least one of a numeric representation of the estimated glucose values, an indication of directional trend of the estimated glucose values, and a graphical representation of the estimated glucose values in response to detection of transient non-glucose related signal artifacts, wherein the estimated glucose values comprise filtered, calibrated glucose values. 4. The system of claim 3, wherein the processor module further comprises programming configured to output unfiltered, calibrated glucose values when transient non-glucose related signal artifacts are not detected. 5. The system of claim 1, wherein the glucose sensor is an intravenous blood glucose sensor, and wherein the detected transient non-glucose related signal artifact is produced due to an oxygen deficiency resulting from the intravenously placed sensor resting against a wall of a blood vessel. 6. The system of claim 1, wherein the pressure detector is a microelectromechanical system. 7. The system of claim 1, further comprising a sensor housing that houses the processor module, wherein the pressure detector comprises a strain gauge configured to monitor a stress on the sensor housing. 8. A method for processing data signals obtained from a glucose sensor, the method comprising: receiving, using receiving module, a signal from a glucose sensor over a period of time;generating, using a pressure detector, a signal indicative of a measured pressure within the glucose sensor;detecting, using a processor module, transient non-glucose related signal artifacts based at least in part on the signal from the glucose sensor and the signal from the pressure detector; andoutputting, using the processor module, to an external device or display estimated glucose values in response to the detection of the detected transient non-glucose related signal artifacts. 9. The method of claim 8, further comprising generating, using the glucose sensor, the signal from the glucose sensor. 10.... | DexCom, Inc. | A61B005/05 A61B005/00
US07949388 | B1 | 2011-05-24 | Methods and systems to characterize ST segment variation over time | 2007-03-16 | A system and method are provided for characterizing ST segment (STS) variations of a patient. The method includes collecting STS variations for multiple data collection periods, calculating multiple statistical parameters based on the collected STS variations, and constructing a STS variation trend. The statistical parameters are associated with a corresponding data collection period. The STS variation trend is based on the multiple statistical parameters over multiple data collection periods. | 1. A method for characterizing ST segment (STS) variations of a patient, comprising: collecting and storing STS variations for multiple data collection periods, each of the data collection periods separated by a cycle;calculating multiple statistical parameters based on the STS variations collected, wherein each statistical parameter is associated with a corresponding data collection period;sorting the calculated statistical parameters for the STS variations into a plurality of heart rate bins based on at least two or more of a normal resting heart rate bin, a normal active heart rate bin, an elevated heart rate bin, a fast heart rate bin, an arrhythmic heart rate bin, and a bradycardia heart rate bin for each of the multiple data collection periods; andconstructing a STS variation trend of the sorted STS variations as a function of the sorted heart rate bins for each of the multiple data collection periods. 2. The method according to claim 1, wherein statistical parameters comprise at least one of an average, an average deviation, a standard deviation, a regression analysis, and combinations thereof. 3. The method according to claim 1 wherein statistical parameters comprises regression analysis comprises at least one of a simple linear regression, a least squares linear regression, a robust regression, a multiple linear regression, a nonlinear regression, a maximum likelihood estimation, and combinations thereof. 4. The method according to claim 1, wherein the data collection period comprises at least of a one day period, a three day period, a seven day period, a ten day period and a fourteen day period. 5. The method according to claim 1, wherein start times for successive data collection periods are separated by a cycle, the cycle being based on a period of time elapsed since implant of an implantable medical device. 6. The method according to claim 1, further comprising presenting to a user as at least one of a display, a pictorial representation, a symbolic representation, a graph, a bar graph, a chart, a histogram, a pie chart, a Venn diagram, and combinations thereof. 7. The method according to claim 1, wherein the STS variation trend indicates an abnormal physiology. 8. The method according to claim 7, wherein the abnormal physiology comprises at least one of an ischemia, a myocardial infarction, a post-myocardial infarction, a silent myocardial infarct, an arrhythmia, a fibrillation, a heart block or a congestive heart failure. 9. A system for characterizing ST segment (STS) variations of a patient, comprising: a memory module to store a plurality of cardiac signals having ST segments collected during multiple data collection periods, each of the data collection periods separated by a cycle, the memory module storing the values of STS variations; anda processor module configured to:calculate multiple statistical parameters based on the STS variations recorded,sort the calculated statistical parameters for the STS variations into ... | Pacesetter, Inc. | A61B005/04
US08355774 | B2 | 2013-01-15 | System and method to evaluate electrode position and spacing | 2009-10-30 | A method and apparatus for tracking and illustrating the location of leads positioned within the volume is disclosed. For example, the lead electrodes can be positioned within a heart of a patient that can be tracked over time. The lead electrodes can be tracked with an electrode potential or bioimpedance tracking system to determine the position of the lead electrodes. A method and apparatus is disclosed to analyze the position information for analyzing the selected position of the lead electrodes. | 1. A method of determining movement of a portion of a subject, comprising: positioning a first implantable lead having a first implantable electrode into a structure of the subject and positioning the first implantable electrode at a first location;positioning a second implantable lead having a second implantable electrode into a structure of the subject and positioning the second implantable electrode at a second location;positioning a third electrode on or in the subject;injecting at least two axes of current into a volume of the subject;determining a first impedance using the positioned first implantable electrode and the third electrode;determining a second impedance using the positioned second implantable electrode and the third electrode;executing instructions with a processor external to the subject to determine a plurality of sequential first positions of the first implantable electrode and a plurality of sequential second positions of the second implantable electrode based on the determined first impedance and the determined second impedance using the positioned first implantable electrode, the positioned second implantable electrode, and the third electrode;calculating a unit vector;projecting a first displacement vector representing motion of the positioned first implantable electrode onto the unit vector and a second displacement vector representing motion of the positioned second implantable electrode onto the unit vector;illustrating a plurality of projected positions on a plane of the subject over time; anddetermining motion of the structure of the subject based on the determined plurality of sequential first positions of the first implantable electrode and the plurality of sequential second positions of the second implantable electrode. 2. The method of claim 1, wherein selecting the unit vector includes rotating an initial unit vector around one of two injected axes relative to the selected unit vector. 3. The method of claim 2, further comprising: injecting three axes into the volume of the subject;wherein selecting the unit vector includes rotating the initial unit vector around two of the three injected axes relative to the selected unit vector. 4. The method of claim 1, wherein rotating a component of a first vector and a second vector includes executing instructions with the processor to mathematically compute the rotations. 5. The method of claim 1, wherein illustrating the plurality of positions on the plane over time includes illustrating a line. 6. The method of claim 1, further comprising: implanting an implantable medical device into the subject to at least one of cardiovert the heart of the subject, pace the heart of the subject, defibrillate the heart of the subject or combinations thereof. 7. The method of claim 6, further comprising: implanting the first implantable electrode and the second implantable electrode; andconnecting the first implantable electrode and the second implantable electrode to the implant... | Medtronic, Inc. | A61B005/05 A61B019/00 A61N001/00
US08437836 | B2 | 2013-05-07 | Reverse data reconstruction for optimal time sampling of counts in physiological list-mode nuclear imaging | 2009-05-29 | A diagnostic imaging device includes detector elements (16) for detecting y-rays indicative of nuclear decay events. Pairs of concurrently detected γ-rays define lines of response (LORs) which are collected, time stamped, and compiled in list-mode. In tissue perfusion studies, it is beneficial to use the data that concurrently maximizes contrast and signal-to-noise ratio in the reconstructed images. Using the list-mode data, events in an adjustable temporal window (33) are reconstructed and the reconstructed images are analyzed to determine a figure of merit based on contrast and signal-to-noise properties of the image. By iteratively adjusting the temporal window, extending its start point (36) backwards in time, and repeating the reconstructing, analyzing, and adjusting steps, an image with an optimal figure of merit is obtained. | 1. A diagnostic imaging apparatus comprising: a detector array for receiving events from an imaging region;a triggering processor for assigning a time stamp to the received events;a reconstruction processor that reconstructs an adjustable portion of the events that fall in a temporal window into an image representation;an analyzer that analyzes the reconstructed image representation to determine a parameter indicative of image quality and adjusts the adjustable portion of the verified events by adjusting the temporal window in accordance with the determined parameter, the reconstruction processor reconstructing the adjusted portion of the verified events in the adjusted temporal window into a subsequent image representation;wherein the reconstruction processor and the analyzer iteratively repeat the analyzing, adjusting the temporal window, and reconstructing until the determined parameter meets a preselected criteria. 2. The diagnostic imaging apparatus as set forth in claim 1, wherein the determined parameter is a figure of merit based on a selected parameter of the image representation. 3. The diagnostic imaging apparatus as set forth in claim 1, further including: an event storage buffer that stores verified events in a list by time of receipt of the event. 4. The diagnostic imaging apparatus as set forth in claim 1, further including: a display configured for displaying at least one of the reconstructed image representations and an indication of the temporal window in which the reconstructed events fall to a user. 5. A method of diagnostic imaging comprising: collecting a set of data points indicative of received nuclear decay events;time stamping the received events;reconstructing an image representation from an adjustable portion of the collected data points that fall in a temporal window;analyzing the image representation to determine a value of a parameter indicative of image representation quality;adjusting the adjustable portion of the collected data points which are used in reconstructing a subsequent image representation by adjusting the temporal window in accordance with the determined parameters. 6. The method as set forth in claim 5, wherein the parameter determining step includes: determining a figure of merit for the image representation. 7. The method as set forth in claim 6, further including: comparing the figure of merit for the subsequent image representation to the figure of merit for the previous image representation, and if the figure of merit for the subsequent image representation is greater than or equal to the figure of merit for the previous image representation, repeating further adjusting the portion of the collected data points that are used to reconstruct a next image representation. 8. The method as set forth in claim 5, wherein the portion of the collected data points used in reconstructing the subsequent image representation is adjusted based on the determined value and is different from the portion of... | Koninklijke Philips Electronics N.V. | A61B006/00
US08529463 | B2 | 2013-09-10 | Systems and methods for testing vestibular and oculomotor function | 2009-04-14 | The invention relates to systems and methods for testing vestibular and oculomotor function. One aspect of the invention provides a vestibular and oculomotor function testing device including a track supported by a plurality of bearings, an engine configured to selectively displace the track, and a head coupling component coupled to the track. The head coupling component is configured to convey a movement generated by the engine to a subject's head in one or more axes. | 1. A vestibular and oculomotor function testing device comprising: a curved track supported by a plurality of bearings;an engine configured to selectively displace the track;a bite block coupled to the track, the bite block configured to convey a movement generated by the engine to a subject's head so as to rotate the subject's head about an axis outside the device and through the subject's cranio-cervical junction region;one or more angle-adjustable connectors adapted and configured to facilitate rapid reorientation of the track to align its center of rotation with axes that are outside the device and through the cranio-cervical junction region of the subject's head and substantially parallel to axes of one or more semicircular canals in the subject's inner ears; anda sensor adapted and configured to measure and report motion of the engine, the track, and/or the head. 2. The vestibular and oculomotor function testing device of claim 1, further comprising: an oculographic device configured to measure the subject's vestibular and oculomotor function by capturing and comparing motion of the subject's eyes and head in one or more dimensions. 3. The vestibular and oculomotor function testing device of claim 2, wherein the oculographic device is a video-oculographic device. 4. The vestibular and oculomotor function testing device of claim 3, wherein the video-oculographic device comprises: one or more cameras; andan analysis module in communication with the camera for analyzing movement of the subject's eyes. 5. The vestibular and oculomotor function testing device of claim 4, wherein the video-oculographic device further comprises: one or more markers for affixation to the subject's head;wherein the analysis module is configured to correct partial decoupling of the subject's head from the camera through analysis of motion of the one or more markers. 6. The vestibular and oculomotor function testing device of claim 2, wherein the oculographic device is a search coil. 7. The vestibular and oculomotor function testing device of claim 2, wherein the oculographic device is an electro-oculographic device. 8. The vestibular and oculomotor function testing device of claim 1, further comprising: a device configured to record the subject's myogenic potentials or electro-oculographic potentials. 9. The vestibular and oculomotor function testing device of claim 1, wherein the engine is selected from the group consisting of: an electric motor, a hydraulic piston, and a pneumatic piston. 10. The vestibular and oculomotor function testing device of claim 9, wherein the electric motor is one selected from the group consisting of: a stepper motor, a servomotor, and a linear actuator. 11. A method for eliciting vestibular function in a subject, the method comprising: providing the vestibular and oculomotor function testing device of claim 1;positioning the subject in contact with the bite block;instructing the subject to visually fixate or track one or more ... | The Johns Hopkins University | H04R025/00 A61B005/04 A61B013/00 A61B007/00 A61B005/103 A61B005/117 A61B003/10 A61B003/02
US09274126 | B2 | 2016-03-01 | Risk factors and prediction of myocardial infarction | 2010-11-15 | Biomarkers and methods are disclosed for diagnosing the risk of a myocardial infarction in an individual by measuring the levels of a set of biomarkers in a sample from an individual. A risk score is calculated for the individual by weighting the measured levels of the biomarkers. The risk score then is used to identify whether the individual is likely to experience a myocardial infarction. In addition, kits are disclosed that include a set of reagents for specifically measuring biomarker levels in a sample from an individual. | 1. A method of diagnosing the risk of a future myocardial infarction in an individual, the method comprising: measuring the level of each biomarker of a set of biomarkers in a sample from an individual, wherein each biomarker of the set of biomarkers is measured using a respective antibody that specifically measures the biomarker and the set of biomarkers comprises: (i) carcinoembryonic antigen, beta-2 microglobulin and transferrin; and(ii) at least one of N-terminal pro B-type natriuretic peptide, alpha-1 antitrypsin, and C-reactive protein;calculating a risk score for the individual by weighting the measured levels of the biomarkers, wherein calculating a risk score comprises:transforming logarithmically the measured levels of the biomarkers to generate a transformed value for each measured biomarker;multiplying the transformed value of each biomarker by a biomarker constant to generate a multiplied value for each biomarker; andsumming the multiplied value of each biomarker to generate the risk score; andusing the risk score to identify a likelihood that the individual will experience a myocardial infarction in the future. 2. The method of claim 1, wherein the set of biomarkers comprises N-terminal pro B-type natriuretic peptide. 3. The method of claim 1, wherein the set of biomarkers comprises alpha-1 antitrypsin. 4. The method of claim 3, wherein the set of biomarkers comprises N-terminal pro B-type natriuretic peptide. 5. The method of claim 3, wherein the set of biomarkers further comprises apolipoprotein A1 and apolipoprotein B. 6. The method of claim 5, wherein the set of biomarkers further comprises alpha fetoprotein. 7. The method of claim 6, wherein the set of biomarkers further comprises vascular endothelial growth factor. 8. The method of claim 1, wherein the set of biomarkers comprises C-reactive protein. 9. The method of claim 8, wherein the set of biomarkers further comprises N-terminal pro B-type natriuretic peptide, apolipoprotein A1, and apolipoprotein B. 10. The method of claim 1, comprising the step of transmitting, displaying, storing, or printing; or outputting to a user interface device, a computer readable storage medium, a local computer system or a remote computer system, information related to the likelihood of myocardial infarction in the individual. 11. The method of claim 10, wherein the information is the risk score or an equivalent thereof. 12. The method of claim 1, comprising recommending, authorizing, or administering treatment if the individual is identified as having an increased likelihood of a future myocardial infarction. 13. The method of claim 1, comprising identifying the individual as having an increased likelihood of having a future myocardial infarction event if the risk score is greater than a reference risk score, and identifying the individual as having a decreased likelihood of having a future myocardial infarction event if the risk score is less than the reference risk score. 14. Th... | BG Medicine, Inc. | G01N033/68 A61B005/00
US20110160566 | A1 | 2011-06-30 | MRI AND ULTRASOUND GUIDED TREATMENT ON A PATIENT | 2010-12-20 | A method is provided for guiding a procedure on a region of interest in a body part of patient, where the body part moves within the patient by breathing, cardiac or other action. The procedure includes radiation therapy, or guidance of a probe for example for biopsy or brachytherapy. The method includes using an MRI system to obtain an MR image of the body part, during the procedure on the body part of the patient, obtaining real time ultrasound images of the body part of the patient as the body part moves within the body of the patient, registering the MR image of the body part with the ultrasound image of the body part so as to locate the region of interest in the ultrasound image of the body part and using the registered images to target the action to the region of interest as the body part moves. | 1. A method for targeting an action on a region of interest in a body part of a body of a patient, where the body part moves within the body of the patient during functioning of the patient, the method comprising: using an MR imaging system to obtain at least one MR image of the body part within the patient so as to locate the region of interest within the body part;during the action on the body part of the patient, obtaining real time ultrasound images of the body part of the patient as the body part moves within the body of the patient;registering the MR image of the body part with the ultrasound image of the body part so as to locate the region of interest in the ultrasound image of the body part;and using the registered images to target the action to the region of interest as the body part moves. 2. The method according to claim 1 wherein the ultrasound images are obtained using an ultrasound imaging probe. 3. The method according to claim 2 wherein the MR image of the body part is registered with the ultrasound image of the body part by imaging the probe in the MR image of the body part and including on the probe components which are located in the MR image so as to locate the probe in the MR image. 4. The method according to claim 3 wherein the probe carries at least one RF coil arrangement for use in obtaining the MR image and wherein the RF coil arrangement is used as a probe component for locating the probe in the MR image. 5. The method according to claim 2 wherein the probe carries at least one RF coil arrangement for use in obtaining the MR image. 6. The method according to claim 1 wherein the body part is an organ of the body visible in the ultrasound image. 7. A method for targeting radiation therapy on a region of interest in a body part of a body of a patient, where the body part moves within the body of the patient during the radiation therapy, the method comprising: using an MR imaging system to obtain at least one MR image of the body part within the patient so as to locate the region of interest within the body part;during the radiation therapy, obtaining real time ultrasound images of the body part of the patient as the body part moves within the body of the patient;registering the MR image of the body part with the ultrasound image of the body part so as to locate the region of interest in the ultrasound image of the body part;and using the registered images to target the radiation therapy to the region of interest as the body part moves. 8. The method according to claim 7 wherein the ultrasound images are obtained using an ultrasound imaging probe located adjacent or within the body of the patient. 9. The method according to claim 8 wherein the MR image of the body part is registered with the ultrasound image of the body part by imaging the probe in the MR image of the body part and including on the probe components which are located in the MR image so as to locate the probe in the MR image. 10. The method accordin... | - | A61B005/055
US20110152698 | A1 | 2011-06-23 | METHOD AND APPARATUS FOR BLOOD PRESSURE WAVEFORM BASELINE ESTIMATION AND REMOVAL | 2009-12-23 | An implantable medical device system including an implantable blood pressure sensor extracts a baseline signal from the sensed blood pressure signal and subtracts the extracted baseline signal from the sensed blood pressure signal to obtain a corrected pressure monitoring signal. The corrected pressure signal is monitored to detect a cardiac-related condition. | 1. A method for removing baseline wander from a blood pressure signal acquired by an implantable medical device, the method comprising: sensing a blood pressure signal using an implantable pressure sensor; extracting a baseline signal from the sensed blood pressure signal;subtracting the extracted baseline signal from the sensed blood pressure signal to obtain a corrected pressure monitoring signal; anddetecting a cardiac-related condition in response to the corrected pressure monitoring signal. 2. The method of claim 1 further comprising: computing a variability of a plurality of pressure waveforms from the corrected pressure monitoring signal;comparing the variability to an allowable threshold of variability; andadjusting the baseline signal extraction if the measured variability exceeds the allowable threshold. 3. The method of claim 1 further comprising detecting a non-cardiac related condition in response to the extracted baseline signal. 4. The method of claim 3 further comprising storing the detected cardiac-related and non-cardiac related conditions. 5. The method of claim 4 further comprising adjusting a therapy based on the detected cardiac-related and non-cardiac related conditions. 6. The method of claim 1 wherein extracting the baseline signal comprises: identifying a baseline point on the pressure signal for each of a plurality of cardiac cycles;computing a baseline value of the pressure signal for each of the identified baseline points; andinterpolating the baseline signal between the identified baseline points using the computed baseline values. 7. The method of claim 6 wherein identifying a baseline point comprises: computing the first time derivative of the pressure signal;summing a plurality of consecutive sample points of the first time derivative;comparing the sum of the plurality of consecutive sample points to a threshold;selecting a sample point of the sensed pressure signal corresponding in time to one of the plurality of consecutive sample points as the baseline point in response to the sum being less than the threshold. 8. The method of claim 7 wherein computing the baseline value comprises using a sample point of the sensed pressure signal corresponding in time to at least one of the plurality of consecutive sample points to compute the baseline value. 9. The method of claim 6 wherein identifying the baseline point comprises: detecting a cardiac electrical event;selecting a sample point of the sensed pressure signal corresponding to the detected cardiac electrical event. 10. The method of claim 6 wherein computing the baseline value comprises: detecting a cardiac electrical event;selecting a plurality of sample points of the sensed pressure signal corresponding to the cardiac electrical event;computing the baseline value using at least one of the plurality of sample points of the sensed pressure signal correlated in time to the cardiac electrical event. 11. An implantable medical device system, comprising: an... | - | A61B005/0215
US20110118554 | A1 | 2011-05-19 | DEVICE-ASSISTED SOCIAL NETWORKING FOR HEALTH MANAGEMENT | 2009-11-17 | Device-assisted social networking for health management, in one aspect, collects health related information associated with one or more users, performs analysis and publishes data resulting from the analyzing to a social networking website associated with one or more users, wherein the data may be shared among the users in a group for peer encouragement in habits that would improve the users' health conditions. | 1. A method for device-assisted social networking for health management, comprising: collecting at a server executing one or more computer processors, health related information associated with a user;analyzing, by a computer processor, the health related information associated with a user; andpublishing data resulting from the analyzing to a social networking website associated with the user. 2. The method of claim 1, wherein the user is located remotely from the computer processor. 3. The method of claim 1, wherein the health related information associated with a user is collected automatically from biometric devices, the biometric devices enabled to measure the health related information associated with the user. 4. The method of claim 3, wherein the biometric devices are located remotely from the computer processor. 5. The method of claim 1, wherein the health related information associated with a user is submitted by the user by the user entering the health related information on a user device and transmitting the health related information to the computer processor. 6. The method of claim 1, wherein the publishing includes presenting the data in a graph format showing changes to health condition of the user over a predetermined period of time. 7. The method of claim 1, wherein the publishing includes providing one or more recommendations for improving health conditions of the user. 8. The method of claim 1, wherein the publishing includes sharing the data with a selected group of users in the social networking website. 9. A system for device-assisted social networking for health management, comprising: a processor;a server module operable to execute on the processor, and further operable to collect health related information associated with a user, and analyze the health related information associated with a user, the server module further operable to publish data resulting from the analyzing to a social networking website associated with the user. 10. The system of claim 9, further including: a social networking web site operable to allow one or more users to interact in a group and share data relating to health of the one or more users. 11. The system of claim 9, wherein the server module is operable to present the analyzed health related information on the social networking website to provide peer encouragement in adjusting user's health behavior. 12. A computer readable storage medium storing a program of instructions executable by a machine to perform a method of device-assisted social networking for health management, comprising: collecting at a server executing one or more computer processors, health related information associated with a user;analyzing, by a computer processor, the health related information associated with a user; andpublishing data resulting from the analyzing to a social networking website associated with the user. 13. The computer readable storage medium of claim 12, wherein the user is located remote... | COMPUTER ASSOCIATES THINK, INC. | A61B005/00 G06Q099/00 G06Q050/00
US07938785 | B2 | 2011-05-10 | Fusion-based spatio-temporal feature detection for robust classification of instantaneous changes in pupil response as a correlate of cognitive response | 2009-01-23 | A computationally efficient and robust approach for monitoring the instantaneous pupil response as a correlate to significant cognitive response to relevant stimuli derives data samples d(n) of the pupil diameter (area), v(n) of pupil velocity and a(n) of pupil acceleration from the pupillary response and segments the data samples into a sequence of time-shifted windows. A feature extractor extracts a plurality of spatio-temporal pupil features from the data samples from the response and baseline periods in each window. A classifier trained to detect patterns of the extracted spatio-temporal pupil features for relevant stimuli generates an output indicative of the occurrence of absence of a significant cognitive response in the subject to a relevant stimulus. | 1. A method for detecting relevant stimuli using pupillary response as a correlate to cognitive response, comprising: providing a classifier, said classifier trained to detect patterns of spatio-temporal pupil features and generate an output indicative of the occurrence or absence of a relevant stimulus that evokes both a cognitive response and a pupillary response in a subject;using a sensor to measure the pupillary response of a subject subjected to stimuli in an environment, wherein the occurrence of a relevant stimulus is initially unknown to the classifier;deriving data samples d(n) of the pupil diameter from the pupillary response;segmenting the data samples d(n) into a sequence of time-shifted windows, each window including a response period and a baseline period;extracting a plurality of spatio-temporal pupil features from the data samples d(n) in said response and baseline periods in each said window; andfor each said window, presenting the extracted pupil features to the classifier to generate an output indicative of the occurrence or absence of a relevant stimulus. 2. The method of claim 1, wherein said plurality of pupil features includes at least one comparative pupil feature extracted from samples d(n) in both the response and baseline periods. 3. The method of claim 2, wherein said at least one comparative pupil feature is a difference between a first diameter, velocity or acceleration feature computed from samples d(n) in the response period and a respective second diameter, velocity or acceleration feature computed from samples d(n) in the baseline period. 4. The method of claim 2, wherein said plurality of pupil features includes at least one absolute pupil feature extracted from samples d(n) in only the response period. 5. The method of claim 4, wherein at least one of the absolute pupil features is a diameter, velocity or acceleration feature. 6. The method of claim 4, wherein at least one of the absolute or pupil features is a velocity or acceleration feature. 7. The method of claim 1, further comprising deriving data samples v(n) of the velocity and a(n) of the acceleration, said pupil features including a diameter, a velocity and an acceleration feature computed from the d(n), v(n) and a(n) samples in the response and baseline periods. 8. The method of claim 1, further comprising synthesizing the one or more temporal outputs indicative of the occurrence of a significant cognitive response to a particular relevant stimulus to reinforce or reject the output and, if reinforced, to resolve the time-stamp of the stimulus. 9. The method of claim 1, wherein the output is provided as a feature to the classifier for the next time-shifted window. 10. The method of claim 1, further comprising measuring light conditions, extracting a light feature indicative of illumination changes and inputting the light feature to the classifier to discount pupillary responses due to illumination changes. 11. The method of claim 1, wherein ... | Teledyne Scientific & Imaging, LLC | A61B013/00 A61B005/103 A61B005/117
US07935061 | B1 | 2011-05-03 | Method and apparatus for monitoring physiological conditions | 2006-05-09 | The invention concerns a physiological monitoring device that, in one embodiment, comprises a capacitive sensor. The sensor comprises a first plate, a second plate and a dielectric medium coupled to the first plate and second plate. A power source charges the sensor to generate an initial capacitance. A band couples the sensor to a blood vessel of a patient. For example, the blood vessel can be located in the umbilical cord of a human baby. An impulse from the blood vessel exerts pressure against the second plate of the sensor and consequently generates a first capacitance. A controller monitors the patient's physiological measurement by detecting the generation of the first capacitance. | 1. A method for determining a physiological measurement comprising the following steps: bounding a capacitive sensor to a pulsating blood vessel of an umbilical cord of a patient;said sensor including a first plate, a second plate, and a dielectric medium operatively coupled to said first plate and said second plate;said dielectric medium including a material that enables said first plate and said second plate to move relative to one another when pressure from said pulsating blood vessel is exerted on said second plate;detecting a first capacitance generated in response to a first pressure from a first pulse from said pulsating blood vessel; wherein said first pressure is exerted against said second plate of said sensor; anddetecting a physiological measurement of said patient, by a controller, in response to detecting said first capacitance. 2. The method of claim 1, wherein said step of detecting said physiological measurement includes detecting a blood pressure of said patient. 3. The method of claim 2, wherein said step of detecting said blood pressure is performed for more than 1 hour in a substantially continuous manner. 4. The method of claim 2, wherein said step of detecting said blood pressure further comprises the following steps: detecting a second capacitance generated in response to a second pressure from a second pulse from said pulsating blood vessel, wherein said second pressure is exerted against said second plate of said sensor; anddetecting said blood pressure using a moving window comprised of said first capacitance and said second capacitance. 5. The method of claim 1, wherein said step of operatively coupling said capacitive sensor to said pulsating blood vessel is performed non-invasively. 6. The method of claim 1, wherein said step of detecting said physiological measurement includes detecting a pulse rate of said patient. 7. The method of claim 6, further comprising the following steps: detecting a second capacitance generated in response to a second pressure from a second pulse from said pulsating blood vessel, wherein said second pressure is exerted against said second plate of said sensor; anddetecting said pulse rate in response to detecting said first capacitance and detecting said second capacitance. 8. The method of claim 1, wherein said pulsating blood vessel is substantially located in a part of the patient's body chosen from the group consisting of said patient's head, arm, torso and leg. 9. One or more program storage media readable by a machine and containing instructions for performing the following steps of the method contained in claim 1: detecting a first capacitance generated in response to a first pressure from a first pulse from said pulsating blood vessel; wherein said first pressure is exerted against said second plate of said sensor; anddetecting a physiological measurement of said patient in response to detecting said first capacitance. 10. The method of claim 1, further comprising the foll... | - | A61B005/02 A61B005/021 A61B005/026
US07949167 | B2 | 2011-05-24 | Automatic learning of image features to predict disease | 2009-04-22 | A method for training a computer system for automatic detection of regions of interest includes receiving patient records. For each of the received patient records a text field and a medical image are identified from within the patient record and the medical image is automatically segmented to identify a structure of interest. The text field is searched for one or more keywords indicative of a particular abnormality associated with the structure of interest. The medical image is added to a grouping representing the particular abnormality when the text field indicates that the patient has the particular abnormality and the medical image is added to a grouping representing the absence of the particular abnormality when the text field does not indicate that the patient has the particular abnormality. The groupings of medical images are used to automatically train a computer system for the subsequent detection of the particular abnormality. | 1. A method for training a computer system for automatic detection of regions of interest in medical image data using a computer-based image processing device, comprising: receiving a plurality of patient records from an electronic medical records database, and for each of the received patient records: identifying a text field and a medical image from within the patient record;automatically segmenting the medical image to identify a structure of interest;searching the text field for one or more keywords indicative of a particular abnormality associated with the structure of interest;determining whether the text field indicates that the patient has the particular abnormality; andadding the medical image to a grouping of medical images representing the particular abnormality when it is determined that the text field indicates that the patient has the particular abnormality and adding the medical image to a grouping of medical images representing the absence of the particular abnormality when it is determined that the text field does not indicate that the patient has the particular abnormality, andusing the grouping of medical images representing the particular abnormality and the grouping of medical images representing the absence of the particular abnormality to automatically train a computer system for the subsequent detection of the particular abnormality. 2. The method of claim 1, wherein for all records, the medical image is a CT scan, an MRI, a PET or SPECT scan, photograph or an ultrasound image. 3. The method of claim 1, wherein for each record, the text field is from a radiology report corresponding to the medical image. 4. The method of claim 1, wherein the structure of interest is a particular organ. 5. The method of claim 4, wherein the particular abnormality is a particular disease, injury or defect of the particular organ. 6. The method of claim 1, wherein determining whether the text field indicates that the patient has the particular abnormality includes searching the text field for keywords that are indicative of a positive diagnosis for the particular abnormality. 7. The method of claim 1, wherein training the computer system for the subsequent detection of the particular abnormality includes training one or more classifiers based on the grouping of medical images representing the particular abnormality and the grouping of medical images representing the absence of the particular abnormality. 8. The method of claim 1, wherein using the grouping of medical images representing the particular abnormality and the grouping of medical images representing the absence of the particular abnormality to automatically train for the subsequent detection of the particular abnormality includes using the grouping of medical images representing the particular abnormality and the grouping of medical images representing the absence of the particular abnormality as training data for a machine-learning operation. 9. A method for automatic dete... | Siemens Medical Solutions USA, Inc. Siemens Aktiengesellschaft | G06K009/00 A61B006/00 A61B005/00
US08053204 | B2 | 2011-11-08 | Beta-2 microglobulin as a biomarker for peripheral artery disease | 2010-12-20 | The present invention provides β2 microglobulin as a biomarker for qualifying or assessing peripheral artery disease in a subject. | 1. A method for qualifying peripheral artery disease status in a subject comprising: (a) measuring at least one biomarker in a biological sample from the subject, wherein said at least one biomarker is selected from the group consisting of β2-microglobulin, lysozyme and cystatin C; and(b) further measuring one or more of the following criteria of said subject: C-reactive protein levels, total cholesterol levels, triglyceride levels, low density lipoprotein levels, high density lipoprotein levels, homocysteine levels, interleukin levels, fibrinogen levels, lipoprotein A levels, 8-iso-prostaglandin F 2alpha (8-iso-PGF 2alpha) levels, and soluble Fas levels; and(c) correlating said measurements (a) and (b) with peripheral artery disease versus non-peripheral artery disease. | Vermillion, Inc. Board of Trustees of the Leland Stanford Junior University | G01N033/53 G01N033/543
US07969161 | B2 | 2011-06-28 | Optimal drive frequency selection in electrical tomography | 2009-12-01 | Electrical tomography drive frequency selection systems and methods are disclosed. One aspect of the present invention pertains to a system for optimally selecting a drive frequency of an electrical tomography which comprises a sensor electrode stably associated with a tissue site within an internal organ of a subject for generating an induced signal based on a noise signal over a range of frequency bands, wherein an electrical field for the electrical tomography is turned off. In addition, the system comprises a noise processing module for isolating the induced signal for each frequency band over the range of frequency bands. Furthermore, the system comprises a frequency select module for selecting a drive frequency of the electrical field for the electrical tomography by comparing the induced signal for each frequency band over the range of frequency bands. | 1. A system for optimally selecting a drive frequency of an electrical tomography, comprising: a sensor electrode implanted at a tissue site of a subject to generate an induced signal based on a noise signal over a range of frequency bands, wherein an electrical field for the electrical tomography is turned off;a noise processing module to separate the induced signal for each frequency band over the range of frequency bands; anda frequency select module to select a drive frequency of the electrical field for the electrical tomography by comparing the induced signal for each frequency band over the range of frequency bands. 2. The system of claim 1, wherein the range of frequency bands is between 50 KHz and 150 KHz. 3. The system of claim 1, wherein a bandwidth of each frequency band comprises 100 Hz. 4. The system of claim 1, wherein the sensor electrode is coupled to the noise processing module via an electrical lead. 5. The system of claim 1, wherein the noise processing module comprises a spectrum analyzer. 6. The system of claim 5, wherein the spectrum analyzer is operable for performing a Fast Fourier Transform of the induced signal. 7. The system of claim 1, wherein the selecting the drive frequency of the electrical field comprises selecting a frequency band with a minimal instantaneous amplitude among the frequency bands. 8. The system of claim 1, wherein the sensor electrode comprises a quadrature amplitude modulator. 9. The system of claim 8, wherein the noise processing module comprises a quadrature amplitude demodulator. 10. The system of claim 9, wherein the quadrature amplitude demodulator comprises a first mixer and a first low pass filter coupled in series to demodulate an in-phase modulated signal (I) and a second mixer and a second low pass filter to demodulate a quadrature-phase modulate signal (Q). 11. The system of claim 10, wherein the noise processing module further comprises a frequency oscillator coupled to the first mixer and the second mixer to sweep across the range of frequency bands. 12. The system of claim 11, wherein the selecting the drive frequency of the electrical field comprises selecting a frequency band with a minimal instantaneous amplitude among the frequency bands. 13. The system of claim 1, wherein the noise processing module comprises a plurality of band pass filters for processing the range of frequency bands. 14. The system of claim 13, wherein the selecting the drive frequency of the electrical field comprises selecting a frequency band with a minimal instantaneous amplitude among respective outputs of the plurality of band pass filters. 15. The system of claim 1, further comprising an electrical field generator module to generate the electrical field with the drive frequency. 16. The system of claim 1, wherein the electrical field is based on a sinusoidal wave. 17. The system of claim 1, wherein the electrical field is based on a spread spectrum signal. 18. A method for optimally se... | Proteus Bomedical, Inc. | G01R029/26 A61B005/05
US07953471 | B2 | 2011-05-31 | Method and apparatus for implantation between two vertebral bodies | 2009-07-27 | A method and system to assist in a planning and navigation of procedure. Generally, the system allows for image acquisition of a selected area of the anatomy. The images may then be used to mark various points to determine true anatomical definitions and planes. The definitions may assist in positioning a prosthesis. | 1. A method of selecting a position for a prosthesis to be implanted in an anatomy, comprising: obtaining image data of the anatomy that includes at least a portion of a spine having at least two substantially adjacent vertebrae;displaying the obtained image data on an image device viewable by a user;defining a plurality of points on the displayed image data corresponding at least to each of a first vertebra and a second vertebra of the at least two substantially adjacent vertebrae, the plurality of defined points including a posterior point on a spinous process, a lateral point on a transverse process, a lateral point on a vertebral body, or combinations thereof;determining a midline of the first vertebra and the second vertebra based at least on the defined plurality of points;determining a midline of the portion of the spine based on the determined midline of the first vertebra and the determined midline of the second vertebra;displaying the determined spinal midline on the displayed image data relative to at least the portion of the spine;tracking a sensor coupled to an instrument;determining a position of the instrument relative to at least one of the defined plurality of points of the first and second vertebrae and the determined spinal midline based on tracking of the sensor;determining an incision point relative to at least one of the defined plurality of points of the first and second vertebrae and the determined spinal midline;navigating a surgical procedure through an incision and relative to the defined plurality of points and the determined spinal midline; andnavigating the prosthesis through the incision and relative to at least the determined spinal midline to the selected position, wherein said prosthesis navigation substantially assists in positioning the prosthesis in the selected position. 2. The method of claim 1, wherein defining a plurality of defined points includes defining the plurality of points relative to the displayed image data, and storing the plurality of points relative to the image data in a storage device. 3. The method of claim 1, wherein defining a plurality of points on the displayed image data includes manipulating a user input on an image device to select certain points on the displayed image data. 4. The method of claim 1, wherein defining a plurality of points of the first and second vertebrae further includes defining lateral most points of the transverse process, defining a posterior most point of the spinous process, defining two lateral most points of the vertebral body, or combinations thereof. 5. The method of claim 1, wherein determining a midline of the portion of the spine further includes: determining an orientation of the obtained image data; anddisplaying the determined spinal midline obliquely relative to the obtained image data. 6. The method of claim 1, wherein the determined spinal midline includes a substantially weight bearing spinal midline. 7. The method of claim 1, further compr... | Medtronic Navigation, Inc. | A61B005/05
US20110152640 | A1 | 2011-06-23 | METHOD AND SYSTEM FOR UTILIZING HEALTH CARE PRODUCTS | 2009-12-17 | In varying embodiments, a method and system of utilizing health care products is disclosed. The method and system employ a uniquely configured device for measuring a biometric component of the human body. Based on this measured biometric component, health care products can be selectively used that are best suited for the measured biometric component. Consequently, health care products can be utilized in a more precise and accurate fashion. An aspect of the invention is a method for utilizing health care products. The method comprises using a sensory device to measure a biometric component and applying one of a plurality of health care products based on the measured biometric component. | 1. A method for utilizing health care products comprising: using a sensory device to measure a biometric component;applying one of a plurality of health care products based on the measured biometric component. 2. The method of claim 1 wherein applying one of a plurality of health care products further comprises employing a symbolic schema to associate the measured biometric component with one of the plurality of health care products. 3. The method of claim 2 wherein the symbolic schema associates a plurality of indicators with a plurality of measured biometric components. 4. The method of claim 3 wherein each of the plurality of indicators is a distinct color. 5. The method of claim 3 wherein each of the plurality of indicators is a distinct symbol. 6. The method of claim 3 wherein the biometric component comprises a pH level of skin. 7. The method of claim 3 wherein the biometric component comprises a pH level of hair. 8. The method of claim 3 wherein the plurality of health care products comprise at least one of creams, soaps, cosmetics and shampoos. 9. A system for utilizing health care products comprising: a sensory device for measuring a biometric component;a plurality of health care products; anda symbolic schema for associating the measured biometric component with each of the plurality of health care products. 10. The system of claim 9 wherein the symbolic schema associates a plurality of indicators with a plurality of measured biometric components. 11. The system of claim 10 wherein each of the plurality of indicators is a distinct color. 12. The system of claim 10 wherein each of the plurality of indicators is a distinct symbol. 13. The system of claim 10 wherein the biometric component comprises a pH level of skin. 14. The system of claim 10 wherein the biometric component comprises a pH level of hair. 15. The system of claim 10 wherein the plurality of health care products comprise at least one of creams, soaps, cosmetics and shampoos. | SENSORDERM | A61M035/00 A61B005/00
US20110118576 | A1 | 2011-05-19 | Noninvasive fetal blood oxygen monitoring system and associated method | 2008-08-28 | A method for non-invasively measuring the oxygen saturation of an in utero fetus's blood using near-infrared spectroscopy. Exemplary methods include placement of a sensor on the outside of the uterus approximate the placenta. Other exemplary methods include inserting a probe into the uterus. Sensors may be positioned approximate a particular portion of the fetus, such as the brain or kidney, to measure the oxygen saturation within the particular portion of the fetus. | 1. A method of measuring oxygen concentration within fetal blood comprising: placing a sensor approximate an outside of a wall of a uterus, the sensor being adapted to be operatively coupled to a near-infrared spectroscopy device; andmeasuring a saturation of oxygen in blood of a fetus present within the uterus using the near-infrared spectroscopy device and the sensor. 2. The method of claim 1, wherein the step of placing the sensor includes placing the sensor approximate the outside of the wall of the uterus generally opposing a placenta present within the uterus; and wherein the step of measuring the saturation of oxygen includes measuring a saturation of oxygen in the fetus's blood present within the placenta. 3. The method of claim 1, wherein the sensor is a miniature sensor. 4. The method of claim 3, further comprising, prior to the step of placing the sensor, creating a minimally invasive incision and inserting the sensor through the minimally invasive incision. 5. The method of claim 4, further comprising, prior to the step of placing the sensor, visualizing the uterus using a laparoscope. 6. The method of claim 1, wherein the sensor and the near-infrared spectroscopy device are adapted to be operatively connected via a wireless data link. 7. The method of claim 1, wherein the step of measuring the saturation of oxygen includes continuously measuring the saturation of oxygen during at least a portion of a therapeutic procedure. 8. The method of claim 7, wherein the therapeutic procedure includes placing the fetus on cardiopulmonary bypass. 9. The method of claim 1, wherein the step of placing the sensor includes placing the sensor approximate the outside of the wall of the uterus generally near at least one of a brain and a kidney of the fetus; and wherein the step of measuring the saturation of oxygen includes measuring a saturation of oxygen in the fetus's blood present within at least one of the brain and the kidney. 10. A method of measuring fetal blood oxygen concentration comprising: inserting a probe into a uterus, the probe including a sensor adapted to be operatively coupled to a near-infrared spectroscopy device; andmeasuring a saturation of oxygen in blood of a fetus present within the uterus using the near-infrared spectroscopy device and the sensor. 11. The method of claim 10, wherein the step of inserting the probe includes placing the sensor approximate a placenta present within the uterus; and wherein the step of measuring the saturation of oxygen includes measuring a saturation of oxygen in the fetus's blood present within the placenta. 12. The method of claim 10, wherein the sensor is a miniature sensor. 13. The method of claim 12, further comprising, prior to the step of inserting the probe, creating a minimally invasive incision; wherein the step of inserting the probe includes inserting the probe through the minimally invasive incision. 14. The method of claim 13, further comprising, prior to placing the ... | - | A61B005/1455
US08731667 | B2 | 2014-05-20 | Pacing output configuration selection for cardiac resynchronization therapy patients | 2011-01-31 | Cardiac therapy systems include multiple electrodes respectively positionable at multiple left ventricular electrode sites. A pulse generator is coupled to the electrodes and configured to deliver a cardiac resynchronization therapy (CRT). A processor is configured to measure, for each left ventricular electrode site, a timing interval between first and second cardiac signal features associated with left ventricular depolarization. The timing interval is associated with a degree of responsiveness of each left ventricular electrode site to CRT. The processor is configured to determine a pacing output configuration that provides improved patient responsiveness to CRT based on the timing interval measurements and to select at least one left ventricular electrode site from the plurality of left ventricular electrode sites based on the timing interval measurements. The processor may be configured to monitor for a change in hemodynamic status of the patient based on a change in the timing interval. | 1. A method, comprising: sensing, from within a patient, cardiac electrical signals from a plurality of disparate left ventricular electrode sites;measuring, for each left ventricular electrode site, a timing interval between first and second cardiac electrical signal features associated with ventricular depolarization, the timing interval comprising a Q1-LV timing interval or a Q-LV timing interval and associated with a degree of responsiveness of the left ventricular electrode sites to cardiac resynchronization pacing therapy (CRT);detecting a change in the timing interval for each left ventricular electrode site; andmonitoring for a change in hemodynamic status of the patient based on the detected timing interval change. 2. The method of claim 1, wherein monitoring for the change in hemodynamic status comprises tracking progression or regression of a heart failure status. 3. The method of claim 1, further comprising generating an alert signal responsive to the change in the patient's hemodynamic status. 4. The method of claim 1, further comprising modifying selection of one or more of the left ventricular pacing sites responsive to the change in the patient's hemodynamic status. 5. The method of claim 1, further comprising modifying one or more pacing delays responsive to the change in the patient's hemodynamic status. 6. The method of claim 1, further comprising modifying an atrioventricular (AV) delay parameter responsive to the change in the patient's hemodynamic status. 7. The method of claim 1, further comprising modifying delivery of CRT in response to the change in the patient's hemodynamic status. 8. The method of claim 1, wherein detecting the change in the timing interval comprises comparing a change in each timing interval to a predetermined threshold. 9. The method of claim 1, wherein detecting the change in the timing interval comprises comparing a change in each timing interval to a threshold based on a Q1-LV timing interval or a threshold based on a Q-LV timing interval. 10. The method of claim 1, wherein detecting the change in the timing interval comprises comparing a change in each timing interval to one or more thresholds based on sensed intervals for various cardiac electrode sites other than Q1-LV and Q-LV timing intervals. 11. A system, comprising: a plurality of electrodes positionable at a plurality of disparate sites of the left ventricle and configured for sensing cardiac electrical signals at the disparate left ventricular sites;a pulse generator coupled to the electrodes and configured to deliver at least a cardiac resynchronization therapy (CRT); anda processor coupled to the electrodes and the pulse generator, the processor configured to measure, for each left ventricular electrode site, a timing interval between first and second cardiac electrical signal features associated with ventricular depolarization, the timing interval comprising a Q1-LV timing interval or a Q-LV timing interval and associated wi... | Cardiac Pacemakers, Inc. | A61N001/365
US08372147 | B2 | 2013-02-12 | Method for detecting body parameters | 2011-01-27 | A method for detecting biometric parameters includes the steps of performing a bone graft procedure on a hip, providing a biometric sensor at the hip, and measuring a biometric parameter at the hip with the sensor. The sensor is capable of measuring parameters in an adjacent surrounding including pressure, tension, shear, relative position, and vascular flow. Data relating to the biometric parameter is transmitted to an external source and the data are analyzed to evaluate a biometric condition of the hip. A set of the sensors can be placed on the hip. | 1. A method for detecting biometric parameters, which comprises: performing a surgical procedure on at least one hip;implanting at least one biometric transceiver at the at least one hip;transmitting a first energy wave from the transceiver into a procedure area including at least one of the hip and an area adjacent the hip;quantitatively assessing the behavior of the energy wave with the transceiver;after transmitting the first energy wave: transmitting a second energy wave from the transceiver into the procedure area; andquantitatively assessing the behavior of the second energy wave;wherein at least one of the first and second energy waves are pulsed during transmission in a vibratory manner to stimulate the procedure area in accordance with at least one detected parameter of the procedure area; anddetermining a current status of the at least one parameter of the procedure area selected from the group consisting of pressure, tension, shear, relative position, bone density, fluid viscosity, temperature, strain, angular deformity, vibration, venous flow, lymphatic flow, load, torque, distance, tilt, shape, elasticity, motion, bearing wear, subsidence, bone integration, change in viscosity, particulate matter, kinematics, stability, and vascular flow with the transceiver based upon a comparison between the assessed behavior of the first and second energy waves. 2. The method according to claim 1, which further comprises: transmitting data relating to the at least one biometric parameter to an external source; andanalyzing the data to evaluate a biometric condition of the at least one hip. 3. The method according to claim 2, which further comprises, based upon the evaluation of the biometric condition of the at least one hip, changing a currently ongoing interoperative procedure at the procedure area. 4. The method according to claim 2, which further comprises, based upon the evaluation of the biometric condition of the at least one hip, chronically monitoring the anatomic condition relating to the procedure area. 5. The method according to claim 1, which further comprises providing a set of the transceivers on the at least one hip. 6. The method according to claim 1, which further comprises providing energy from outside the procedure area to the transceiver to power the transceiver and, thereby, create the energy wave and quantitatively assess the behavior of the energy wave. 7. The method according to claim 6, which further comprises providing the energy through at least one of an electromagnetic couple, a magnetic couple, a capacitive couple, an inductive couple, a sonic couple, an ultrasonic couple, a fiber optic couple, an optical couple, and an infrared couple. 8. The method according to claim 1, which further comprises: providing a set of the biometric transceivers at the procedure area;transmitting an energy wave from the transceivers into the procedure area;quantitatively assessing the behavior of the energy wave with at least one ... | - | A61F002/44
US20110118619 | A1 | 2011-05-19 | METHODS AND APPARATUS FOR MONITORING CONSCIOUSNESS | 2010-08-09 | The systems of the present invention provide improved accuracy in monitoring, analysing, detecting, predicting and/or providing alerts and alarms associated with depth of anaesthesia, depth of consciousness, hypnotic state, sedation depth, fatigue or vigilance of a subject, with as few as 3 surface electrodes. The systems incorporate real-time phase, amplitude and frequency analysis of a subject's electro-encephalogram. The systems weight outputs of various types of analyses to produce an integrated analysis or display for precise indication or alert to users of the systems including anaesthetists, nurses and other medical personnel, transport drivers and machine workers. The systems weight the outputs of one or more analysis algorithms including combinations of simultaneous, real-time R&K analysis, AEP spectral analysis-SEF-MF, Bi-coherence analysis, initial wave analysis, auditory response, arousal analysis, body movement analysis, 95% spectral edge analysis and anaesthetic phase and spectral energy variance measurement in association with a subject's state of consciousness. | 1. A method of monitoring consciousness of a sentient subject and automatically detecting whether the subject is in a transition from a conscious state to a less conscious state or vice versa, by reducing effects of frequency-based changes in neurological data from the subject, said method including: obtaining an electro-encephalogram signal from the subject;performing a frequency-based analysis on the electro-encephalogram signal to obtain a frequency-based signal;performing a phase-based analysis on the electro-encephalogram signal to obtain a phase-based signal;detecting by comparing the frequency-based signal and the phase-based signal whether the subject is in transition from said conscious state to said less conscious state or vice versa; andproviding a warning signal when said subject is in said transition to said conscious state. 2. The method according to claim 1, wherein said frequency-based analysis includes depth of sleep analysis, and wherein said phase-based analysis includes at least one of optimized bicoherence, bispectrum or triple product analysis. 3. The method according to claim 2, wherein said depth of sleep analysis includes real-time optimized scoring of human sleep physiology. 4. The method according to claim 1, wherein said step of detecting is augmented with optimized audio evoked potential. 5. The method according to claim 2, wherein said step of detecting is augmented with optimized audio evoked potential. 6. The method according to claim 3, wherein said step of detecting is augmented with optimized audio evoked potential. 7. The method according to claim 1, including means for adapting the or each analysis to parameters specific to said subject, including body mass index, age and sex of said subject. 8. The method according to claim 2, including means for adapting the or each analysis to parameters specific to said subject, including body mass index, age and sex of said subject. 9. The method according to claim 3, including means for adapting the or each analysis to parameters specific to said subject, including body mass index, age and sex of said subject. 10. The method according to claim 4, including means for adapting the or each analysis to parameters specific to said subject, including body mass index, age and sex of said subject. | - | A61B005/0476 A61B005/0484
US07945308 | B2 | 2011-05-17 | Systems, methods and apparatus for an endo-rectal receive-only probe | 2005-12-27 | Systems, methods and apparatus are provided through which a compact pod insertable into the rectum for Magnetic Resonance Imaging/Magnetic Resonance Spectroscopy (MRI/MRS) examination of the prostate and containing two receive coils, each connected to transmit blocking and pre-tuned trap circuitry which can be superimposed within the pod in close proximity, can, without either circuit interfering with the other, efficiently gather, for imaging and tissue analysis, radio frequency signals emanating from magnetically disturbed nuclei in prostate tissue. | 1. An apparatus to receive radio frequency signals characteristic of a targeted specimen within a patient, the apparatus comprising: a housing designed to be inserted into a patient; andat least two radio frequency receiver coils, each of which (i) is positioned inside the housing, (ii) has an operating radio frequency, and (iii) includes: a first circuit proximal to the radio frequency receiver coil operable to block an external transmitter only during a transmit cycle at an operating radio frequency of the circuit anda second circuit operable to provide a passive block at the operating radio frequency of other ones of the radio frequency coils;wherein each circuit is operable to seek only a single frequency,wherein the radio frequency receiver coils are superimposed upon each other,wherein the first circuit and the second circuit respectively associated with each of the at least two radio frequency receiver coils provide at least two active traps and two passive traps,wherein the radio frequency signals received by the at least two receiver coils are generated by different target nuclei associated with the specimen. 2. The apparatus of claim 1, wherein the circuit of each radio frequency receiver coil receives radio frequency signals from human tissue. 3. The apparatus of claim 2, wherein the circuit of each radio frequency receiver coil sends radio frequency signals to an imaging device. 4. The apparatus of claim 3 wherein the circuit of one radio frequency receiver coil is pre-tuned to receive only radio frequency signals characteristic of magnetically disturbed carbon-13 nuclei. 5. The apparatus of claim 3 wherein the circuit of one radio frequency receiver coil is pre-tuned to receive only radio frequency signals characteristic of magnetically disturbed hydrogen nuclei. 6. The apparatus of claim 3 wherein the circuit of each radio frequency receiver coil sends radio frequency signals to an imaging device that is part of an MRI/MRS medical diagnostic system. 7. The apparatus of claim 3, wherein the circuit of one radio frequency receiver coil is pre-tuned to receive only radio frequency signals characteristic of proton Larmor frequency. 8. The apparatus of claim 3, wherein the circuit of one radio frequency receiver coil is pre-tuned to receive only radio frequency signals characteristic of proton Larmor frequency. 9. The apparatus of claim 2, wherein the circuit of each radio frequency receiver coil receives radio frequency signals from human tissue that is prostate tissue. 10. The apparatus of claim 1, wherein the circuit of at least one radio frequency receiver coil receives radio frequency signals from human tissue containing carbon-13. 11. The apparatus of claim 1 wherein the apparatus further comprises a circuit that is pre-tuned to receive only the radio frequency characteristic of magnetically disturbed phosphorous nuclei. 12. The apparatus of claim 1 wherein the insertable housing further comprises: a re-usable portion. ... | General Electric Company | A61B005/05
US08161816 | B2 | 2012-04-24 | Hearing test method and apparatus | 2009-11-03 | A method and apparatus for administering a hearing test that emits a test sound with changing volume with respect to elapsed time. The volume changes according to a pre-determined relationship between the volume and elapsed time. The hearing test utilizes the elapsed time until a subject signals the test sound is audible, and the pre-determined relationship, to determine the subject sound pressure level threshold for the test sound. | 1. A hearing test method comprising: reading a pre-generated, read-only, test file via a computer processor from a test file storage location, the test file comprising: computer-executable test instructions for causing a sound generating system to generate a continuous test sound at changing test volume amplitude where the test volume amplitude changes according to a pre-defined relationship between the test volume amplitude and elapsed time, the test sound being a single tone at a single audible frequency;executing the computer-executable test instructions via the computer processor and causing the sound generating system to generate the test sound;detecting, via the computer processor, a subject signal that indicates test sound audibility;measuring, via the computer processor, a test elapsed time from executing the test instructions until detecting the subject signal;transforming the test elapsed time, via the computer processor, into a subject sound pressure level threshold at the single audible frequency using the pre-defined direct relationship between the test volume amplitude and the test elapsed time; anddisplaying the subject sound pressure level threshold via a graphic display device. 2. The hearing test method of claim 1 wherein the test volume amplitude changes by increasing according to a pre-defined direct relationship between the test volume amplitude and elapsed time. 3. The hearing test method of claim 2 wherein the pre-defined direct relationship between the test volume amplitude and elapsed time comprises: the increasing test volume amplitude increasing linearly with elapsed time from a starting test volume amplitude to a final test volume amplitude. 4. The hearing test method of claim 3 wherein transforming the test elapsed time into a subject sound pressure level comprises: multiplying a test volume amplitude linear increase rate per second by the test elapsed time measured in seconds. 5. The hearing test method of claim 2 wherein the final test volume amplitude causes the sound generating system to generate the test sound at a final test sound pressure level that is between 60 decibels and 90 decibels. 6. The hearing test method of claim 2 wherein displaying the subject sound pressure level threshold further comprises: comparing the subject sound pressure level threshold to a representative sound pressure level threshold of a theoretical average subject hearing the single tone at the single audible frequency. 7. The hearing test method of claim 2 wherein the increasing test volume amplitude is a continuously increasing test volume amplitude. 8. A hearing test method comprising: reading a pre-generated, read-only, first test file from a first test file storage location, the first test file comprising: computer-executable first instructions for causing a sound generating system to generate a continuous first test sound at an increasing first test volume amplitude where the first test volume amplitude increases according... |  | A61B005/12 H04R029/00
US08538513 | B2 | 2013-09-17 | Stimulation electrode selection | 2009-12-16 | Bioelectrical signals may be sensed within a brain of a patient with a plurality of sense electrode combinations. A stimulation electrode combination for delivering stimulation to the patient to manage a patient condition can be selected based on a frequency domain characteristic of the sensed bioelectrical signals. In some examples, a stimulation electrode combination is selected based on a determination of which of the sense electrodes are located closest to a target tissue site, as indicated by the one or more sense electrodes that sensed a bioelectrical brain signal with a relatively highest value of the frequency domain characteristic. In some examples, determining which of the sense electrodes are located closest to the target tissue site may include executing an algorithm using relative values of the frequency domain characteristic. | 1. A system comprising: a plurality of electrodes;a sensing module that senses a plurality of bioelectrical brain signals, each of the plurality of bioelectrical brain signals being sensed via at least one respective electrode of the plurality of electrodes; anda processor that determines a frequency domain characteristic for each of the plurality of bioelectrical signals; determines a plurality of relative values of the frequency domain characteristic, wherein each of the plurality of relative values is based on at least two of the frequency domain characteristics; determines whether a first relative value of the plurality of relative values is greater than or less than at least one other relative value of the plurality of relative values; in response to determining that the first relative value is greater than or less than the at least one other relative value, determines whether a second relative value of the plurality of relative values is greater than or less than another relative value of the plurality of relative values; determines whether a target tissue site is positioned between two electrodes of the plurality of electrodes based on determining whether the first relative value is greater than or less than at least one other relative value and based on determining whether the second relative value is greater than or less than another relative value; and selects at least one electrode of the plurality of electrodes for delivering stimulation to the patient based on determining whether the target tissue site is positioned between the two electrodes. 2. The system of claim 1, wherein the frequency domain characteristic comprises a power level of the respective bioelectrical signal in a selected frequency band. 3. The system of claim 2, wherein the selected frequency band comprises at least one of a beta band or a gamma band. 4. The system of claim 1, wherein the processor determines a greatest relative value of the frequency domain characteristic and selects the at least one electrode of the plurality of electrodes for delivering stimulation to the patient based on determining the greatest relative value of the frequency domain characteristic. 5. The system of claim 1, wherein the sensing module senses the plurality of bioelectrical signals with a plurality of unipolar electrode combinations. 6. The system of claim 5, wherein the processor determines a magnitude of a difference in the frequency domain characteristics of at least two bioelectrical signals in order to determine the plurality of relative values of the frequency domain characteristic. 7. The system of claim 1, wherein the sensing module senses the plurality of bioelectrical signals with a plurality of bipolar electrode combinations. 8. The system of claim 7, wherein the frequency domain characteristic of each bioelectrical signal indicates the relative values. 9. The system of claim 1, wherein the bioelectrical signals comprise a first set of bioelectrical signals, wher... | Medtronic, Inc. | A61B005/04
US08529476 | B2 | 2013-09-10 | Catheter with strain gauge sensor | 2009-12-28 | A medical probe, including a flexible insertion tube, having a distal end for insertion into a body cavity of a patient and which is configured to be brought into contact with tissue in the body cavity. The probe further includes a sensor tube of an elastic material, contained inside the distal end of the insertion tube and configured to deform in response to forces exerted by the tissue on the distal end. The probe also includes a plurality of strain gauges fixedly attached to a surface of the sensor tube at different, respective locations and configured to generate respective signals in response to deformations of the sensor tube. | 1. A medical probe, comprising: a flexible insertion tube, having a distal end for insertion into a body cavity of a patient and which is configured to be brought into contact with tissue in the body cavity;a generally cylindrical thin-walled sensor tube comprising an elastic material, contained inside the distal end of the insertion tube and configured to deform in response to forces exerted by the tissue on the distal end wherein the sensor tube is fixedly attached around the circumference of each end of the sensor tube to the inside the distal end to the insertion tube; anda plurality of strain gauges fixedly attached to a surface of the sensor tube at different, respective locations and configured to generate respective signals in response to deformations of the sensor tube. 2. The probe according to claim 1, wherein the surface of the sensor tube comprises an outer curved surface of the sensor tube. 3. The probe according to claim 1, wherein the respective signals are configured to generate a magnitude and a direction of a pressure on a termination of the distal end. 4. The probe according to claim 1, wherein the respective signals are configured to generate a magnitude and a direction of a deflection of a termination of the distal end. 5. The probe according to claim 1, wherein the plurality of strain gauges are symmetrically disposed with respect to an axis of the sensor tube. 6. The probe according to claim 1, wherein the plurality of strain gauges are positioned on a circumference of the surface of the sensor tube, and wherein the circumference is centrally located with respect to the sensor tube. 7. The probe according to claim 1, wherein the plurality of strain gauges have respective gauge directions, and wherein the gauges are fixedly attached to the surface of the sensor tube so that the respective gauge directions are parallel to an axis of the sensor tube. 8. A medical probe, comprising: a flexible insertion tube, having a distal end for insertion into a body cavity of a patient and which is configured to be brought into contact with tissue in the body cavity;a generally cylindrical thin-walled sensor tube comprising an elastic material, contained inside the distal end of the insertion tube and configured to deform in response to forces exerted by the tissue on the distal end;a plurality of strain gauges fixedly attached to a surface of the sensor tube at different, respective locations and configured to generate respective signals in response to deformations of the sensor tube; and at least one temperature-compensating strain gauge configured to generate signals to compensate for changes in temperature of the plurality of strain gauges wherein the at least one temperature-compensating strain gauge is mounted on a block having a block-thermal mass, and wherein the sensor tube has a sensor-tube-thermal mass equal to the block-thermal mass. 9. The probe according to claim 8, wherein the at least one temperature-compensating s... | Biosense Webster (Israel), Ltd. | A61B005/00
US08870785 | B2 | 2014-10-28 | Contactless respiration monitoring of a patient | 2009-05-04 | Respiration of a patient is detected by emitting an electromagnetic signal towards the patient; receiving a reflected electromagnetic signal reflected from the patient; converting the reflected electromagnetic signal, yielding a first signal; phase-shifting the reflected electromagnetic signal and converting the phase-shifted reflected electromagnetic signal, yielding a second signal; determining a first vector being defined by the time derivatives of the first signal and the second signal, for a common first point in time; determining a second vector being defined by the time derivatives of the first signal and the second signal, for a common second point in time; and calculating the scalar product of the normalized first vector and the normalized second vector as an indicator value for a change from expiration to inspiration of the patient or vice versa. A change from expiration to inspiration of the patient or vice versa is preferably indicated if the indicator value is below a threshold value, preferably below a value of 0. In this way, respiration is monitored contactlessly and remotely based on the Doppler radar principle which is reliable and easy to handle. | 1. A method for detection of respiration of a patient comprising the following steps: emitting using a two-channel Doppler radar sensor an electromagnetic signal towards the patient;receiving a reflected electromagnetic signal reflected from the patient using the two-channel Doppler radar sensor;converting using the two-channel Doppler radar sensor the reflected electromagnetic signal to a first signal;phase-shifting using the two-channel Doppler radar sensor the reflected electromagnetic signal and converting the phase-shifted reflected electromagnetic signal to a second signal;determining using a calculating device a first vector by taking time derivatives of the first signal and the second signal, for a common first point in time;determining using the calculating device a second vector by taking the time derivatives of the first signal and the second signal, for a common second point in time;normalizing using the calculating device the first and second vectors;calculating using the calculating device the scalar product of the normalized first vector and the normalized second vector as an indicator value for a change from expiration to inspiration of the patient or vice versa; andcontrolling a display to display a respiration value based on the indicated changed from expiration to inspiration of the patient or vice versa. 2. The method according to claim 1, further including: determining the first point in time and the second point in time for building the first vector and the second vector by determining, respectively, characteristic points in time defined by specific criteria fulfilled simultaneously in the first signal and the second signal. 3. The method according to claim 2, wherein the specific criteria is a detected zero-crossing of the time derivative of the first signal or the second signal, respectively. 4. The method according to claim 1, wherein an indicator value is compared within a predefined threshold value. 5. The method according to claim 4, wherein the predefined threshold value is 0. 6. The method according to claim 4, wherein the change from expiration to inspiration of the patient or vice versa is indicated using the display if the indicator value is below the predefined threshold value. 7. The method according to claim 1, wherein the first vector, the second vector and the scalar product of the normalized first vector and the normalized second vector as an indicator value for the change from expiration to inspiration of the patient or vice versa are consecutively determined, preferably in predefined time periods. 8. The method according to claim 1, wherein calculating includes calculating a vorticity {right arrow over (V)} for the first point in time T1 and the second point in time T2, respectively, by {right arrow over (V)}1=−{right arrow over (r)}1×({right arrow over (M)}1(T1)−{right arrow over (r)}1) and {right arrow over (V)}2=−{right arrow over (r)}2×({right arrow over (M)}2(T2)−{right arrow over (r)}2), where... | Koninklijke Philips N.V. | A61B005/08 A61B005/00 G01S013/88 A61B005/05 A61B005/113 G01S013/58 G01S007/35
US08588893 | B2 | 2013-11-19 | System and method for tracking healing progress of tissue | 2011-01-04 | A system and method for analyzing tissue healing may include capturing an image including a tissue site. A clinician may be enabled to define multiple regions of the image of the tissue site, where at least two of the regions of the image define tissue types at different stages of tissue healing. An area may be calculated of each of the at least two regions of the image defining tissue type at different stages of tissue healing. The calculated areas may be displayed to a clinician. In one embodiment, an image color adjustment may be determined by adjusting the image of the reference color marker to the predetermined color. The image of the tissue site may be normalized by applying the image color adjustment to generate a normalized image. | 1. A method for analyzing tissue healing, said method comprising: capturing an image including a tissue site;enabling a clinician to define multiple regions of the image of the tissue site, at least two of the regions of the image defining tissue types at different stages of tissue healing while in those healing stages;automatically comparing a first color associated with a first tissue type to imaged tissue at the tissue site to identify tissue of the first tissue type at the tissue site;automatically comparing a second color associated with a second tissue type to the tissue at the tissue site to identify tissue of the second tissue type at the tissue site;calculating areas of each of the first and second tissue types; anddisplaying the calculated areas of the first and second tissue types. 2. The method according to claim 1, wherein capturing further includes capturing a reference color marker, and further comprising: adjusting the image of the reference color marker to a predetermined color; andadjusting the image of the tissue site the same as the adjustment to the image of the reference color marker to color normalize the image of the tissue site. 3. The method according to claim 1, further comprising displaying indicia on the image of the tissue site to indicate the at least two regions of the image defining tissue type at different stages of tissue healing. 4. The method according to claim 1, further comprising: automatically comparing a third color associated with a third tissue type to the tissue at the tissue site to identify tissue of the third tissue type at the tissue site;calculating an area of the third tissue type; anddisplaying the area of the third tissue type. 5. The method according to claim 4, further comprising: automatically comparing a fourth color associated with a fourth tissue type to the tissue at the tissue site to identify tissue of the fourth tissue type at the tissue site;calculating an area of the fourth tissue type; anddisplaying the area of the fourth tissue type. 6. The method according to claim 1, further comprising: receiving a tissue site depth;calculating a volume of the tissue site based on the tissue site depth; anddisplaying the calculated volume of the tissue site. 7. The method according to claim 1, wherein capturing the image includes receiving digital data from a camera external of a system configured to perform treatment on the tissue site. 8. The method according to claim 1, wherein enabling the clinician to define multiple regions includes enabling the clinician to trace perimeters around the multiple regions on the captured image. 9. The method according to claim 8, wherein tracing the perimeters around the multiple regions includes tracing perimeters around one region that is a sub-region within another region. 10. A system for analyzing tissue healing, said system comprising: a memory configured to store data;an input/output unit configured to communicate data remotely to and from the... | KCI Licensing, Inc. | A61B006/00 G06K009/00
US08655440 | B2 | 2014-02-18 | System and method of speech sound intelligibility assessment, and program thereof | 2011-03-01 | A speech sound intelligibility assessment system includes: a biological signal measurement section for measuring an electroencephalogram signal of a user; a presented-speech sound control section for determining a speech sound to be presented by referring to a speech sound database retaining a plurality of monosyllabic speech sounds; an audio output section for presenting the speech sound determined by the presented-speech sound control section as an audio; a characteristic component detection section for utilizing the electroencephalogram signal of the user measured by the biological signal measurement section to determine presence or absence of a characteristic component of an event-related potential at 800 ms±100 ms from a point of presenting the audio; and a speech sound intelligibility assessment section for, based on a result of determination by the characteristic component detection section, determining whether the user has aurally comprehended the speech sound or not. | 1. A speech sound intelligibility assessment system, comprising: a biological signal measurement section for measuring an electroencephalogram signal of a user;a presented-speech sound control section for determining a speech sound to be presented by referring to a speech sound database retaining a plurality of monosyllabic speech sounds;an audio output section for presenting the speech sound determined by the presented-speech sound control section as an audio;a characteristic component detection section for utilizing the electroencephalogram signal of the user measured by the biological signal measurement section, determining presence or absence of a characteristic component of an event-related potential at 800 ms±100 ms from a point of presenting the audio; anda speech sound intelligibility assessment section for, based on a result of determination by the characteristic component detection section, determining whether the user has aurally comprehended the speech sound or not. 2. The speech sound intelligibility assessment system of claim 1, wherein, the event-related potential is acquired by utilizing an electrode position Pz according to the International 10-20 system, and the characteristic component detection section determines that a characteristic component exists in the event-related potential when a component equal to or greater than a predetermined value is present in the event-related potential; andif the characteristic component detection section determines that the characteristic component does not exist in the event-related potential, the speech sound intelligibility assessment section determines that the user has aurally comprehended the speech sound, andif the characteristic component detection section determines that the characteristic component exists in the event-related potential, the speech sound intelligibility assessment section determines that the user has not aurally comprehended the speech sound. 3. The speech sound intelligibility assessment system of claim 1, wherein, the event-related potential is acquired by utilizing an electrode position Cz according to the International 10-20 system, and the characteristic component detection section determines that a characteristic component is present in the event-related potential when a component equal to or less than a predetermined value is present in the event-related potential; andif the characteristic component detection section determines that the characteristic component does not exist in the event-related potential, the speech sound intelligibility assessment section determines that the user has aurally comprehended the speech sound, andif the characteristic component detection section determines that the characteristic component exists in the event-related potential, the speech sound intelligibility assessment section determines that the user has not aurally comprehended the speech sound. 4. The speech sound intelligibility assessment system of claim 1, wherein, in... | Panasonic Corporation | A61B005/04
US08884754 | B2 | 2014-11-11 | Monitoring vital parameters of a patient using a body sensor network | 2009-08-18 | The invention relates to a method of monitoring a plurality of vital parameters of a patient 10 using a body sensor network with a set of on-body sensors 1 and at least one off-body monitoring device 2, the method comprising the following steps: with each on-body sensor 1, sensing a vital parameter and transmitting data related to the sensed vital parameter to the off-body monitoring device 2, and for at least one of the on-body sensors 1, determining the inclination of this on-body sensor 1 relative to the off-body monitoring device 2. In this way, a reliable and easy to use possibility for monitoring vital parameters of a patient 10 using a body sensor network is provided that minimizes the performance problem produced by RF attenuation caused by the body of the patient 10. | 1. A method of monitoring vital sign data of a patient using a body sensor network with a set of on-body vital sign sensors and at least one off-body monitoring device, the method comprising the following steps: with each on-body sensor, sensing a vital sign of the patient and transmitting vital sign data indicative of the sensed vital sign of the patient to the off-body monitoring device,wherein the on-body sensors communicate wirelessly with the off-body monitoring device and with one another,wherein at least one of the on-body sensors determines an inclination of the at least one on-body sensor relative to the off-body monitoring device,wherein the at least one on-body sensor determines the on-body location of the at least one on-body sensor,wherein the at least one on-body sensor determines a direct data delivery route between at least one on-body sensor for which the inclination and the on-body location has been determined and the off-body monitoring device,determining a reliability value indicative of data packet error rates along a direct data delivery route on the basis of the inclination and the on-body location of the at least one on-body sensor, androuting the vital sign data of the at least one on-body sensor via another on-body sensor to the off-body monitoring device if the reliability value of the direct data delivery route between the at least one on-body sensor and the off-body monitoring device is less than a predefined value. 2. The method according to claim 1, wherein the step of routing the vital sign data of the at least one on-body sensor via another on-body sensor to the off-body monitoring device is responsive to the reliability value of the direct data delivery route via of the another on-body sensor being better than the reliability value of the direct data delivery route of the at least one on-body sensor. 3. The method according to claim 1, further comprising the step of choosing the another on-body sensor out of the set of other on-body sensors based on an expected remaining operating time of the other on-body sensors, respectively. 4. The method according to claim 1, further comprising the step of choosing the another on-body sensor out of the set of other on-body sensors based on the number of hops of the respective data delivery route via the respective other on-body sensors. 5. The method according to claim 1, further comprising the step of monitoring the packet error rates and storing the reliability value of at least two different direct data delivery routes in relation to the patient's posture detected by the inclination of an on-body sensor. 6. The method according to claim 1, further comprising the step of choosing the another on-body sensor out of the set of other on-body sensors based on a stored performance of a respective data delivery route comprising the respective another on-body sensor. 7. A body sensor network for monitoring vital signs of a patient, with a set of on-body sensors and at least ... | Koninklijke Philips N.V. | G08B001/08 H04W040/12 G06F019/00 A61B005/00
US08374686 | B2 | 2013-02-12 | Continuous monitoring of risk burden for sudden cardiac death risk stratification | 2010-06-03 | System and method for assessing a likelihood of a patient to experience a cardiac arrhythmia. A biological sensor is configured to sense biological parameters of the patient. A processor is coupled to the biological sensor and is configured to determine the likelihood of the patient experiencing a cardiac arrhythmia based, at least in part, on a combination of the biological parameters, the combination dynamically weighting each of the plurality of biological parameters based on another one of the plurality of biological parameters. | 1. A system for assessing a likelihood of a patient to experience a cardiac arrhythmia, comprising: a biological sensor configured to sense a plurality of biological parameters of said patient;a processor operatively coupled to said of biological sensor and configured to determine said likelihood of said patient experiencing a cardiac arrhythmia based, at least in part, on a combination of said plurality of biological parameters, said combination dynamically weighting each of said plurality of biological parameters based on another one of said plurality of biological parameters. 2. The system of claim 1 wherein at least one of said weightings of said plurality of biological parameters is different from at least one other one of said weightings. 3. The system of claim 1 wherein said processor determines said likelihood of said patient to experience a cardiac arrhythmia based, at least in part, on a quantitative analysis using a number of said plurality of biological parameters exceeding a corresponding number of predetermined thresholds. 4. The system of claim 1 wherein each individual one of said plurality of biological parameters corresponds to a qualitative value, and wherein said processor determines said likelihood of said patient to experience a cardiac arrhythmia based, at least in part, on a qualitative analysis using a total of said qualitative values of said plurality of biological parameters. 5. The system of claim 1 wherein each of said plurality of biological parameters corresponds to one of a plurality of groups, and wherein said processor is configured to determine said likelihood of said patient to experience a cardiac arrhythmia is based, at least in part on at least one biological parameter from each of said plurality of groups. 6. The system of claim 5 wherein one of said plurality of groups is a genetic information group. 7. The system of claim 5 wherein one of said plurality of groups incorporates ones of said plurality of biological parameters which indicate a condition of a substrate of a heart of said patient. 8. The system of claim 5 wherein one of said plurality of groups incorporates ones of said plurality of biological parameters which indicate a condition of an autonomic system of said patient. 9. The system of claim 5 wherein one of said plurality of groups incorporates ones of said plurality of biological parameters which indicate a burden of an arrhythmia on said patient. 10. The system of claim 1: wherein each of said plurality of biological parameters corresponds to one of a plurality of groups;wherein one of said plurality of groups incorporates ones of said plurality of biological parameters which indicate a condition of a substrate of a heart of said patient;wherein one of said plurality of groups incorporates ones of said plurality of biological parameters which indicate a condition of an autonomic system of said patient;wherein one of said plurality of groups incorporates ones of said plurality of bi... | Medtronic, Inc. | A61B005/04
US08922209 | B2 | 2014-12-30 | System and method for magnetic resonance imaging | 2009-03-11 | A magnetic resonance imaging method comprises performing imaging where more than one polarizing magnetic field strength is used during scanning and processing at least one image resulting from the scanning to yield an enhanced contrast image. | 1. A magnetic resonance imaging (MRI) method comprising: imaging a subject that has been administered with a contrast agent or endogenous substance using an MRI machine, said contrast agent or endogenous substance binding to target molecules in said subject and demonstrating a magnetic field-dependent variation in one or more MRI properties, wherein during said imaging: the strength of the static polarizing magnetic field of said MRI machine is shifted in different directions so that said subject is exposed to a variable magnetic field that highlights the magnetic field-dependent variation in the one of more MRI properties of said contrast agent or endogenous substance,an image in the presence of said static polarizing magnetic field following a shift in the strength thereof in one direction is acquired and then,an image in the presence of said static polarizing magnetic field following a shift in the strength thereof in the opposite direction is acquired; andsubtracting the acquired images to yield an enhanced contrast image. 2. The method of claim 1 wherein during said imaging the strength of the static polarizing magnetic field of said MRI machine is shifted during relaxation portions of pulse sequences. 3. The method of claim 2 wherein the timing, magnitude and duration of each static polarizing magnetic field shift is chosen to maximize image contrast. 4. The method of claim 2 wherein the static polarizing magnetic field is shifted in each direction by substantially the same amount. 5. The method of claim 2 wherein during said imaging an electromagnet is conditioned to generate magnetic field pulses of opposite polarities to shift the static polarizing magnetic field in opposite directions. 6. The method of claim 5 wherein during said imaging said electromagnet is removably inserted into said MRI machine. 7. The method of claim 6 wherein during said imaging said electromagnet is positioned axially within a bore of said MRI machine. 8. The method of claim 6 wherein during said imaging said electromagnet is positioned either beneath or beside a patient positioned within a bore of said MRI machine. 9. A magnetic resonance imaging (MRI) method comprising: scanning a subject that has been administered with either a contrast agent or an endogenous substrate using an MRI machine and generating an image for each scan, wherein said contrast agent or endogenous substance binds to target molecules in said subject and demonstrates a magnetic field-dependent variation in one or more MRI properties and wherein said scanning comprises shifting the strength of the static polarizing magnetic field of said MRI machine in different directions so that said subject is exposed to a variable magnetic field that highlights the magnetic field-dependent variation in the one of more MRI properties of said contrast agent or endogenous substance, and acquiring an image in the presence of said static polarizing magnetic field following a shift in the strength the... | The University of Western Ontario | G01V003/00 G01R033/44 G01R033/50 G01R033/56 G01R033/381 G01R033/563
US09915608 | B2 | 2018-03-13 | Optical sensor for determining the concentration of an analyte | 2008-10-31 | A method and apparatus for non-invasively determining a concentration of glucose in a subject using optical excitation and detection is provided. The method includes emitting an exciter beam (B1) to irradiate a portion (130) of tissue of the subject, causing physical and chemical changes in the surface, and causing an initial back scattering (D1) of light. The method further includes periodically emitting a probe beam (B2) which irradiates the portion of tissue and causes periodic back scatterings (D2) of light. The initial and periodic back scatterings are detected and converted into electrical signals of at least the amplitude, frequency or decay time of the physical and chemical changes, the back scatterings being modulated by the physical and chemical changes. By differentiating over time at least one of the amplitude, frequency or decay time of the physical and chemical changes, the concentration of glucose may be determined. | 1. An apparatus for determining a concentration of an analyte in tissue, comprising: at least one radiation source operative to emit a first radiation beam and at least one second radiation beam, the first radiation beam irradiating a testing area of tissue and cause a first scattering of radiation, and the at least one second radiation beam having a lower intensity than the first radiation beam and periodically irradiating the testing area of tissue causing periodic second scatterings of radiation;at least one detector comprising a first radiation detector and a second radiation detector, the first radiation detector and the second radiation detector being configured to detect the first scattering of radiation and the second scatterings of radiation and the radiation detectors convert the detected scatterings into electrical signals, wherein the first radiation detector and the second radiation detector are located at different distances from the testing area of tissue enabling a relative amplitude between the detected scatterings to be calculated and used as an offset; anda processor for determining the concentration of the analyte based on said electrical signals. 2. The apparatus according to claim 1, wherein: the first and second radiation beams have predetermined wavelengths selected from a characteristic absorption band of the analyte in a predetermined medium. 3. The apparatus according to claim 1, wherein: the first and second radiation beams have predetermined wavelengths, wherein the predetermined wavelength of the first radiation beam corresponds to a peak of an absorption band of the analyte in a predetermined medium, and wherein the first radiation beam interacts with the tissue causing at least one transient process dependent on the concentration of the analyte. 4. The apparatus according to claim 1, wherein the first and second radiation beams have the same wavelength. 5. The apparatus according to claim 1, wherein the first and second radiation beams have different wavelengths. 6. The apparatus according to claim 5, wherein one of the first and second radiation beams each has a wavelength that is greater than an absorption band peak of the analyte in a predetermined medium, and one of the first and second radiation beams has a wavelength that is less than the absorption band peak of the analyte in a predetermined medium. 7. The apparatus according to claim 1, wherein the first radiation beam excites at least one periodic or non-periodic transient process in the tissue, and wherein at least one of the first and second scatterings of radiation are at least partially modulated by the at least one transient process. 8. The apparatus according to claim 7, wherein the at least one transient process includes photo-acoustic oscillations. 9. The apparatus according to claim 7, wherein the electrical signals represent at least one of the amplitude, frequency, or decay time of the at least one transient process. 10. The apparatus a... | BioSensor, Inc. | A61B006/00 G01N021/49 A61B005/00 A61B005/145 A61B005/1455 G01N021/17
US10271751 | B2 | 2019-04-30 | Monitoring the blood pressure of a patient | 2009-08-12 | A method for monitoring the blood pressure of a patient, comprising the following steps: determining a pulse arrival time signal from the patient 2 based on the pulse wave velocity method; determining an accelerometer signal from the patient 2; and triggering an additional measure or deriving a blood pressure value, taking into account the pulse arrival time signal and a DC component of the accelerometer signal. In this way, a possibility for monitoring the blood pressure of a patient is provided with which false alarms and/or unnecessary additional cuff-based blood pressure measurements can be avoided. | 1. A method for monitoring blood pressure of a patient, comprising the following steps: (a) determining a pulse arrival time signal from the patient with a pulse wave velocity unit based on a pulse wave velocity method;(b) determining an accelerometer signal from the patient with an accelerometer sensor after determining the pulse arrival time signal;(c) detecting a DC component of the accelerometer signal indicative of a posture of the patient with a monitoring device;(d) normalizing the pulse arrival time signal with a reference pulse arrival time value with the monitoring device by dividing the determined pulse arrival time signal by a reference pulse arrival time value to generate a normalized pulse arrival time signal;(e) comparing a value of the normalized pulse arrival time signal with predefined upper and lower pulse arrival time threshold values with the monitoring device;(f) determining if the value of the normalized pulse arrival time signal exceeds the upper reference threshold value or under-runs the lower reference threshold value with the monitoring device;(g) checking for a change in the posture of the patient based on a change of the DC component of the accelerometer signal;at least one of:(h) repeating steps (a)-(g) when the value of the normalized pulse arrival time signal exceeds or under-runs the predefined pulse arrival time threshold values;(i) in response to the normalized pulse arrival time signal exceeding or under-running the predefined pulse arrival time threshold values and the DC component change not being detected, at least one of: controlling an alarm to generate an alarm signal; andcontrolling a blood pressure cuff to measure the blood pressure of the patient, and(j) in response to the normalized pulse arrival time signal exceeding or under-running the predefined pulse arrival time threshold values and the DC component change being detected, updating the reference pulse arrival time value; and(k) deriving and displaying a blood pressure value with the monitoring device from the normalized pulse arrival time signal when the value of the normalized pulse arrival time signal does not exceed or under-run the predefined pulse arrival time threshold values. 2. The method according to claim 1, wherein the reference pulse arrival time threshold value is dependent on a heart rate. 3. The method according to claim 1, further including: with the monitoring device, in response to the normalized pulse arrival time signal exceeding or under-running the pulse arrival time threshold values and the posture change being detected, controlling the blood pressure cuff to make a cuff-based measurement of the blood pressure value of the patient. 4. The method according to claim 1, further including, with the monitoring device, at least one of: inputting the pulse arrival time signal and the DC component of the accelerometer signal into a look-up table which was pre-calibrated with blood pressure values to give the blood pressure valu... | KONINKLIJKE PHILIPS N.V. | A61B005/0285 A61B005/021 A61B005/022 A61B005/00 A61B005/0295
US07952692 | B2 | 2011-05-31 | Method and apparatus for determination of analyte concentration | 2006-12-12 | A method and system are presented for use in determination of the concentration of an analyte in a subject's medium. The medium is irradiated with at least two radiation components to produce detectable radiation responses of the medium thereto. These at least two radiation components are selected to have different mean wavelengths and such that the spectral bandwidth of at least one of said at least two radiation components is characterized by relatively high variability of the extinction coefficient of the analyte of interest across said spectral bandwidth. This enables analysis of data indicative of detected radiation responses of the medium to said at least two radiation components and determination of the concentration of said analyte. | 1. A method for use in measurements of the concentration of an analyte in a medium having a certain extinction coefficient spectrum, said method comprising: providing a selected light source unit capable of generating at least first and second radiation components, each having a selected mean wavelength and a selected spectral bandwidth such that the mean wavelength and spectral bandwidth of said first radiation component correspond to a first spectral region of the extinction coefficient spectrum of said analyte where variation of an extinction coefficient, Δβ1, per volume fraction of the analyte across the spectral bandwidth of the first radiation component is higher than in other spectral regions of the extinction coefficient spectrum, and the mean wavelength and second spectral bandwidth of said second radiation component correspond to a second spectral region of the extinction coefficient spectrum where variation of an extinction coefficient, Δβ2 per volume fraction of the analyte across the spectral bandwidth of the second radiation component is smaller than in other spectral regions of the extinction coefficient spectrum, and at least one of said at least two radiation components contains a plurality of wavelengths with radiation intensities that are in a fixed relation to each other during periods of said measurements;irradiating the medium with said at least first and second radiation components during a certain measurement time to thereby produce detectable radiation responses of the medium thereto,analyzing data indicative of detected radiation responses of the medium to said at least first and second radiation components to determine the concentration of said analyte. 2. The method of claim 1, comprising applying at least two different sources of radiation to the medium to irradiate it by said at least two radiation components, respectively. 3. The method of claim 1, wherein said selecting of the mean wavelength of the spectral bandwidth for at least one of said at least two radiation components to be used in the measurements comprises analyzing data indicative of the extinction coefficient spectrum of the analyte of interest. 4. The method of claim 1, wherein the mean wavelength and the spectral bandwidth of at least one of said at least two radiation components are selected such that the variability of the extinction coefficient of said analyte across this spectral bandwidth is one of global and local maxima in absolute value. 5. The method of claim 1, wherein the mean wavelength of said at least two radiation components contains red-near infrared optical spectral regions. 6. The method of claim 1, wherein the mean wavelength of said at least two radiation components is within about 400 nm-2300 nm. 7. The method of claim 1, wherein the radiation responses comprise intensities of at least one of transmitted and reflected radiation through and from an irradiated part of the medium, respectively, for said at least two radiation c... | Orsense Ltd. | G01N033/48 G01N021/00 A61B005/00
US07966077 | B2 | 2011-06-21 | Electrode assembly for a stimulating medical device | 2006-09-29 | An electrode assembly comprising a low-profile, low-volume elongate electrode carrier and a corresponding guide tube for introducing the carrier into the cochlea to place electrodes disposed at the distal end of the carrier at desired locations along the tonotopically-mapped cochlea. The electrode assembly facilitates intra- and extra-cochlea atraumatic implantation of the unobtrusive electrode carrier of the present invention thereby minimizing adverse impact to natural auditory functioning. For example, the electrode assembly may be utilized to implant the low-profile, low-volume elongate electrode carrier into the scala tympani without damaging the delicate structures of the cochlea and without interfering with the natural hydrodynamic nature of the cochlea such as the natural flow of perilymph in the cochlea canals. In one particular embodiment, the carrier is pre-curved to attain a perimodiolar position in the scala tympani to facilitate accurate delivery of electrical stimulation with a minimum stimulation current and power consumption. | 1. An electrode assembly for implanting electrodes in a cochlea of a recipient, comprising: a flexible, elongate electrode carrier configured to have a spiral configuration such that bias forces are exerted by said electrode carrier when arranged in a non-spiral position; andan elongate guide tube having interior dimensions defining a lumen for slidingly retaining the carrier, and exterior dimensions that enable the guide tube to be inserted into the recipient's cochlea, the guide tube comprising a rigid tube body and a distal end region sufficiently flexible to bend in response to said bias forces applied to the distal end region by the electrode carrier when the electrode carrier is positioned within the lumen, such that the distal end region curves to cause the electrode carrier to follow a trajectory not parallel with a longitudinal axis of the rigid tube body as the carrier exits the guide tube. 2. The electrode assembly of claim 1, wherein a first longitudinal imaginary plane extends through the guide tube and a second longitudinal imaginary plane extends through the guide tube and substantially orthogonal to the first plane, wherein the guide tube further comprises: one or more longitudinally-spaced slots substantially aligned along the first plane and disposed along a side of the elongate tube, wherein the longitudinally-space slots are configured to allow the guide tube to bend in a direction within the first plane. 3. The electrode assembly of claim 1, wherein the distal end region is formed separately from the rigid tube body and subsequently fixed to the distal end of the rigid tube body. 4. The electrode assembly of claim 1, wherein the guide tube comprises two or more layers, and further wherein the distal end region is formed by removing one or more layers from the guide tube. 5. The electrode assembly of claim 4, wherein the one or more layers of the guide tube are configured to be removable by chemical removal. 6. The electrode assembly of claim 4, wherein the one or more layers of the guide tube are configured to be removable by laser removal. 7. The electrode assembly of claim 1, wherein the distal end region is configured as a fraction of a tube. 8. The electrode assembly of claim 1, wherein the surface of the distal end region is configured as a corrugated or convoluted surface. 9. The electrode assembly of claim 1, wherein the surface of the distal end region comprises one or more slits configured to allow the distal end region to bend. 10. The electrode assembly of claim 1, wherein the rigid guide tube comprises a braid structure configured to provide rigidity to the rigid guide tube. 11. The electrode assembly of claim 1, wherein the carrier is pre-curved to attain a perimodiolar position in the scala tympani of the cochlea when implanted. 12. The electrode assembly of claim 1, wherein the guide tube and carrier are configured to implant the carrier into the scala tympani without substantial damage to the cochle... | Cochlear Limited | A61N001/00 A61B005/04
US20110118627 | A1 | 2011-05-19 | DISPOSABLE PATIENT INTERFACE FOR INTRADUCTAL FLUID ASPIRATION SYSTEM | 2011-01-26 | Disclosed is a disposable patient interface for an intraductal fluid aspiration system. The interface has a flexible tubular distal member for contacting the patient, and a proximal support portion forming a vacuum chamber therein. The support portion has a vacuum aperture for communication with a source of vacuum, and a retention structure for releasably mounting within a hand piece of an intraductal fluid aspiration system. The flexible member preferably has a low modulus of elasticity and a high tear strength, such that it can be stretched to fit over a rigid support without rupture. | 1. A one time use disposable patient interface for releasable connection with an intraductal fluid aspiration device, comprising: a support, having a proximal end, a distal end and a side wall defining a cavity therein;a flexible patient interface membrane, extending distally from the support, said flexible patient interface membrane having a proximal end operably connected to the distal end of the support and an unsupported distal end, the unsupported distal end configured and sufficiently flexible to be collapsible under its own weight; anda vacuum port on a proximal end of a tubular body, the vacuum port extending proximally with respect to the proximal end of the support. 2. A one time use disposable patient interface as in claim 1, wherein the support comprises a tubular side wall having sufficient rigidity to resist collapse upon application of a 150 mm Hg vacuum within the cavity. 3. A one time use disposable patient interface as in claim 1, wherein the support comprises a tubular side wall having sufficient rigidity to resist collapse upon application of a 250 mm Hg vacuum within the cavity. 4. A one time use disposable patient interface as in claim 1, wherein the unstressed diameter of the distal end of the flexible patient interface membrane is within the range of from about 2 inches to about 4 inches. 5. A one time use disposable patient interface as in claim 1, wherein the diameter of the distal end of the flexible patient interface membrane may be stretched to at least 150% of its unstressed diameter without rupturing the membrane. 6. A one time use disposable patient interface as in claim 1, wherein the diameter of the distal end of the flexible patient interface membrane may be stretched to at least 200% of its unstressed diameter without rupturing the membrane. 7. A one time use disposable patient interface, comprising: a flexible tubular membrane, having a proximal end, an unsupported distal end and a passageway extending therethrough, the unsupported distal end configured and sufficiently flexible to be collapsible under its own weight, wherein the proximal end has a smaller cross section than the distal end; a proximal cap having a proximal end and a distal end, the distal end of the proximal cap operably connected to the proximal end of the flexible tubular membrane; anda vacuum port on the cap extending proximally from the proximal end of the cap for connection to a source of vacuum, wherein the proximal cap comprises a first material, and the flexible tubular membrane comprises a second, different material. 8. A one time use disposable patient interface as in claim 7, wherein an unstressed diameter of the distal end of the flexible tubular membrane is within the range of from about 2 inches to about 4 inches. 9. A one time use disposable patient interface as in claim 7, wherein an unstressed diameter of the distal end of the flexible tubular membrane is no more than about 4 inches. 10. A one time use disposable pati... | NEOMATRIX, LLC | A61B005/00
US07946997 | B2 | 2011-05-24 | Measurement system to measure a physiological condition in a body | 2007-02-16 | Measurement system comprising a sensor wire provided, at its distal end, with a physiological condition sensor to measure a physiological condition inside a patient, and to provide measured data to an external device, the measurement system comprises a transceiver unit adapted to be connected to the proximal end of the sensor wire, and a communication unit arranged in connection with the external device. The transceiver unit is adapted to communicate, by a communication signal, with the communication unit, in order to transfer measured data to the external device. The communication signal, including the measured data, is generated by the transceiver unit and transferred as an output signal and the communication unit is arranged to be connected to a standard input/output connector of the external device and to communicate with the external device in accordance with an established standard, or in accordance with relevant parts of an established standard, e.g. BP22 or USB. The measurement system further comprises a physical optical communication link arranged between the transceiver unit and the communication unit, wherein the communication signal is an optical signal transferred by the optical communication link. The transceiver unit also comprises an energy means adapted to energize the sensor, the transceiver unit and also the optical communication link. | 1. A measurement system comprising: a sensor wire provided, at its distal end, with a physiological condition sensor configured to measure a physiological condition inside a patient, and to provide measured data to an external device,a transceiver unit configured to be connected to a proximal end of the sensor wire,a communication unit connected with the external device,wherein said transceiver unit is configured to communicate, by a communication signal, with said communication unit, to transfer measured data to the external device, wherein the communication signal, including the measured data, is generated by the transceiver unit and is transferred as an output signal,wherein said communication unit is configured to be connected to a standard input/output connector of the external device and to communicate with the external device in accordance with an established standard, or in accordance with relevant parts of an established standard, anda physical optical communication link arranged between said transceiver unit and said communication unit,wherein said communication signal is an optical signal transferred by said optical communication link, and that said transceiver unit comprises an energy device configured to energize said sensor, said transceiver unit and a distal part of said optical communication link,wherein said transceiver unit comprises a sensor signal adapting circuitry configured to filter, process, and format the signal received from the sensor wire,wherein said sensor signal adapting circuitry comprises a programmable sensor conditioning device, a calibration unit, the energy device, and an output amplifying unit. 2. A measurement system according to claim 1, wherein said optical communication link comprises a first communication device configured to be electrically connected to the transceiver unit to handle the communication to and from the transceiver unit, to convert an electrical signal from the transceiver unit to an optical signal and vice versa. 3. A measurement system according to claim 2, wherein the optical communication link further comprises a second communication device configured to be electrically connected to the communication unit to handle the communication to and from the communication unit, to convert an optical signal to an electrical signal and vice versa. 4. A measurement system according to claim 2, wherein said energy device in the transceiver unit is configured to energize the first communication device of the optical communication link. 5. A measurement system according to claim 3, wherein the second communication device of the optical communication link is energized by the communication unit. 6. A measurement system according to claim 1, wherein the measurement system is configured such that a data stream is in the form of data packets. 7. A measurement system according to claim 1, wherein said transceiver unit and communication unit are provided with a device configured to perform bi-directio... | RADI Medical Systems AB | A61B005/00 A61B005/02 A61M025/00
US20110104076 | A1 | 2011-05-05 | Compositions, Methods For Preparing Amino Acids And Nuclear Magnetic Resonance Spectroscopy | 2011-01-11 | The present invention relates to amino acids, complexes, and compounds comprising deuterium and tritium isotopes preferably alpha deuterated amino acids, polypeptides, antibodies, derivatives and saccharide-amino acid complexes and conjugates. In some embodiments, the invention relates to methods of using compounds comprising deuterium for imaging biochemical concentrations and distributions in mammalian tissues using nuclear magnetic resonance spectroscopy. In some embodiments, the invention relates to the used of said amino acids derivatives and complexes in boron neutron capture therapy. In some embodiments, the present invention relates to the preparation of amino acids, polypeptides, antibodies, derivatives and saccharide complexes/conjugates comprising heavy hydrogen isotopes. In some embodiments, the invention relates to racemizing amino acids starting from compositions of any optical purity. In further embodiments, the invention relates to the preparation of amino acids and their N-acyl counterparts with deuterium incorporated at the alpha carbon. | 1. A method comprising a) providing i) a subject comprising mammalian tissue,ii) a compound comprising a deuterium atom, andiii) a nuclear magnetic resonance spectrometer configured to detect said deuterium atom;b) administering said compound to said subject; andc) detecting said deuterium atom with said nuclear magnetic resonance spectrometer in said tissue. 2. The method of claim 1, wherein said compound comprises an alpha-deutero-amino acid. 3. The method of claim 1, wherein said compound further comprises boron. 4. The method of claim 2, wherein said alpha-deutero-amino acid is alpha-deutero-boronophenylalanine. 5. The method of claim 4, wherein said alpha-deutero-boronophenylalanine is an alpha-deutero-boronophenylalanine saccharide conjugate. 6. The method of claim 5, wherein said alpha-deutero-boronophenylalanine saccharide conjugate is selected from the group consisting of: alpha-deutero p-boronophenylalanine -α,α-glucooctitol;alpha-deutero p-boronophenylalanine—volemitol;alpha-deutero p-boronophenylalanine—persietol;alpha-deutero p-boronophenylalanine -α-glucoheptitol;alpha-deutero p-boronophenylalanine—lactitol;alpha-deutero p-boronophenylalanine—cellobiitol;alpha-deutero p-boronophenylalanine—isomaltitol;alpha-deutero p-boronophenylalanine—maltitol;alpha-deutero p-boronophenylalanine—aminosorbitol;alpha-deutero p-boronophenylalanine—aminodulcitol;alpha-deutero p-boronophenylalanine—dulcitol;alpha-deutero p-boronophenylalanine—mannitol;alpha-deutero p-boronophenylalanine—sorbitol;alpha-deutero p-boronophenylalanine—rhamnitol; andalpha-deutero p-boronophenylalanine—xylitol. 7. The method of claim 1, further comprising performing boron neuron capture therapy by exposing said subject to neutron irradiation. 8. A purified isotopic composition of a compound having the following structure: 9. A composition comprising a boronophenylalanine saccharide conjugate. 10. The composition of claim 9, wherein said boronophenylalanine saccharide conjugate is selected from the group consisting of: p-boronophenylalanine -α,α-glucooctitol;p-boronophenylalanine—volemitol;p-boronophenylalanine—persietol;p-boronophenylalanine -α-glucoheptitol;p-boronophenylalanine—lactitol;p-boronophenylalanine—cellobiitol;p-boronophenylalanin—isomaltitol;p-boronophenylalanine—maltitol;p-boronophenylalanine—aminosorbitol;p-boronophenylalanine—aminodulcitol;p-boronophenylalanine—dulcitol;p-boronophenylalanine—mannitol;p-boronophenylalanine—sorbitol;p-boronophenylalanine—rhamnitol; andp-boronophenylalanine—xylitol. 11. The composition of claim 9, wherein said boronophenylalanine saccharide conjugate is an alpha-deuterated boronophenylalanine. 12. The composition of claim 11 wherein said alpha-deuterated boronophenylalanine is selected from the group consisting of: alpha-deuterated p-boronophenylalanine -α,α-glucooctitol;alpha-deuterated p-boronophenylalanine—volemitol;alpha-deuterated p-boronophenylalanine—persietolalpha-deuterated p-boronophenylalanine -α-glucohepti... | Glyconix Corporation | A61B005/055 C07C229/28
US20110125064 | A1 | 2011-05-26 | SMART CLOTHING FOR MOTION PHYSIOLOGICAL MEASUREMENT AND DYNAMICAL STABLE APPARATUS THEREOF | 2010-09-23 | A dynamical stable apparatus of a smart clothing for motion physiological measurement includes at least one physiological sensing module. The physiological sensing module includes a carrier, at least one physiological sensor and a first tension adjusting component. The carrier has a surface for contacting with a body of a user, and the surface is coated with a slide resistant material. The physiological sensor is disposed on the surface of the carrier and configured for measuring a motion physiological signal of the user. The first tension adjusting component is disposed on two opposite ends of the carrier and configured for adjusting a tension of the carrier so as to fix the carrier on the body of the user. When the user wears the smart clothing for motion physiological measurement to measure a motion physiological signal, a stable motion physiological signal can be obtained. | 1. A dynamical stable apparatus of a smart clothing for motion physiological measurement, comprising: at least one physiological sensing module, each physiological sensing module comprising: a carrier having a surface for contacting with a body of a user, the surface being coated with a slide resistant material;at least one physiological sensor disposed on the surface of the carrier and configured for measuring a motion physiological signal of the user; anda first tension adjusting component disposed on two opposite ends of the carrier and configured for adjusting a tension of the carrier to fix the carrier on the body of the user. 2. The dynamical stable apparatus as claimed in claim 1, wherein the carrier is a flexible belt textile and is suitable to surround the body of the user, the slide resistant material is latex, a coating thickness of the slide resistant material is ranged from 0.01 millimeters to 3 millimeters, a coating manner for the slide resistant material is point-shaped, linear, surface-shaped or porous, the slide resistant material has a plurality of micro-vents, and the first tension adjusting component is a Velcro belt or a buckling element. 3. The dynamical stable apparatus as claimed in claim 1, wherein the surface includes a first area and a second area located on an area except the first area, the at least one physiological sensor is disposed on the first area, and the slide resistant material is coated on the first area. 4. The dynamical stable apparatus as claimed in claim 1, wherein the surface includes a first area and a second area located on an area except the first area, the at least one physiological sensor is disposed on the first area, the slide resistant material comprises a first slide resistant material coated on the first area and a second slide resistant material coated on the second area, and a friction force of the first area is larger or smaller than that of the second area. 5. The dynamical stable apparatus as claimed in claim 4, wherein the first slide resistant material and the second slide resistant material are the same, and a coating thickness of the first slide resistant material coated on the first area is larger or smaller than that of the second slide resistant material coated on the second area. 6. The dynamical stable apparatus as claimed in claim 4, wherein the first slide resistant material and the second slide resistant material are the same, and a coating density of the first slide resistant material coated on the first area is larger or smaller than that of the second slide resistant material coated on the second area. 7. The dynamical stable apparatus as claimed in claim 4, wherein the first slide resistant material and the second slide resistant material are different, and a friction force of the first slide resistant material coated on the first area is larger or smaller than that of the second slide resistant material coated on the second area. 8. The dynamical stable apparatus ... | Feng Chia University | A61B005/103
US07937135 | B2 | 2011-05-03 | Method and apparatus for adjusting a sensing parameter | 2007-05-01 | A method and apparatus for determining oversensing in a medical device that includes sensing cardiac signals, determining an oversensing characteristic associated with cardiac signals sensed during a predetermined sensing window, identifying oversensing in response to the oversensing characteristic, determining, in response to oversensing being identified, an adjusting characteristic associated with cardiac signals sensed during the predetermined sensing window, and updating a sensing parameter in response to the determined adjusting characteristic. | 1. A method of determining oversensing in a medical device, comprising: sensing cardiac signals;determining an oversensing characteristic associated with cardiac signals sensed during a predetermined sensing window, wherein the predetermined sensing window comprises a predetermined duration of time unrelated to one or more beat-to-beat durations;identifying oversensing in response to the oversensing characteristic;determining, in response to oversensing being identified, an adjusting characteristic associated with cardiac signals sensed during the predetermined sensing window; andupdating a sensing parameter in response to the determined adjusting characteristic. 2. The method of claim 1, wherein the cardiac signals are sensed from a non-transvenous location. 3. The method of claim 1, wherein determining an oversensing characteristic includes determining a first oversensing characteristic and a second oversensing characteristic, the first oversensing characteristic corresponding to a morphology of the sensed cardiac signals sensed during a predetermined sensing window, and the second oversensing characteristic corresponding to RR-intervals associated with the sensed cardiac signals. 4. The method of claim 1, wherein determining the adjusting characteristic comprises: defining a plurality of sub-segments associated with the sensing window; anddetermining the adjusting characteristic associated with the cardiac signals within each sub-segment of the plurality of sub-segments. 5. The method of claim 4, wherein determining the adjusting characteristic comprises: determining a sub-segment parameter for each of the plurality of sub-segments;comparing the sub-segment parameter with a sub-segment parameter threshold; anddetermining whether to update the sensing parameter in response to the comparing. 6. The method of claim 5, further comprising: determining a maximum amplitude and an absolute minimum amplitude associated with the sensing window; andcomparing the maximum amplitude and the absolute minimum amplitude, wherein the sub-segment parameter and the sub-segment parameter threshold are determined in response to the comparing of the maximum amplitude and the absolute minimum amplitude. 7. The method of claim 5, further comprising: determining a local parameter for each sub-segment of the plurality of sub-segments; andcomparing the local parameter with the sub-segment parameter. 8. The method of claim 7, wherein the sensing parameter is updated in response to the sub-segment parameter and one or more local parameters determined for only a predetermined number of sub-segments of the plurality of sub-segments. 9. The method of claim 1, further comprising: defining a plurality of sub-segments associated with the sensing window;determining intervals associated with each sub-segment of the plurality of sub-segments;determining a rate associated with each sub-segment of the plurality of sub-segments in response to the determined intervals; andadjus... | Medtronic, Inc. | A61B005/0402
US07942826 | B1 | 2011-05-17 | Insulated pedicle access system and related methods | 2006-06-06 | A pedicle access system including a cannula, a stylet, and a removable T-handle. The pedicle access system may be used to percutaneously approach the pedicle, initiate pilot hole formation, and conduct a stimulation signal to the target site for the purposes of performing a pedicle integrity assessment during the pilot hole formation. To do this, the cannula and stylet are locked in combination and inserted through an operating corridor to the pedicle target site, using the T-handle to facilitate easy movement and positioning of the cannula/stylet combination. A stimulation signal may be applied during pilot hole formation to conduct the pedicle integrity assessment. In a significant aspect, the T-handle may be detached from the cannula/stylet combination to facilitate the use of various surgical tools as necessary. | 1. A system for use in a medical procedure, comprising: an access needle, including:an elongated needle element having a shaped tip advanceable into a bony structure of a patient's spine;an elongated cannula lockingly mateable with said needle element and having a lumen to slidably receive said needle element therethrough such that said shaped tip of said needle element extends distally beyond a distal end of said cannula;an insulative sheath disposed about the exterior of said cannula, wherein said insulative sheath is coupled to an attachment element situated near a proximal end of said cannula, the attachment element including an exterior surface defining a lumen enclosing a distal portion of the cannula and a distal portion of the insulative sheath, the attachment element including at least one longitudinally running slot and the proximal portion of the insulative sheath including at least one tab extending into the slot;a locking mechanism that selectively locks said needle element and said cannula together when said needle element is fully received in said cannula; anda handle attachable to at least one of said needle element and said cannula;wherein said access needle has a contact element coupleable to a source of electrical stimulation and in electrical communication with at least one of said shaped tip and said distal end of said cannula such that a stimulation signal is deliverable to a target region of said bony structure to perform a pedicle integrity assessment when said access needle is in contact with said bony structure. 2. The system of claim 1, wherein said handle is removable. 3. The system of claim 1, wherein said contact element is a conductive surface on at least one of said needle element, cannula, and handle, said contact element being engageable by a stimulation clip comprising said stimulation source. 4. The system of claim 1, wherein said insulative sheath is slidable along the exterior of said cannula between a first position and a second position. 5. The system of claim 1, wherein a distal end of said insulative sheath is adapted to engage a surface of said bony structure when said needle element and cannula are advanced into said bony structure such that said insulative sheath slides relative to said cannula as said needle element and cannula are advanced deeper into said bony structure exposing the exterior of said cannula to the interior of the bony structure. 6. The system of claim 1, wherein said insulative sheath is biased to a first position in which the exterior of the cannula is least exposed. 7. The system of claim 1, wherein said locking mechanism comprises a lock collar. 8. The system of claim 1, wherein said handle is adapted to matingly engage an instrument in a fixed position relative to said handle. 9. The system of claim 8, wherein said instrument is at least one of a clip to couple said access needle to said source of electrical stimulation and a sensor for monitoring the angular orientation... | NuVasive, Inc. | A61B005/05
US08229551 | B2 | 2012-07-24 | Method of presenting electrocardiographic data | 2009-11-24 | A system and method of presenting physiological data in a component ring. The method includes receiving a plurality of leads of physiological data. Morphology features are extracted from the plurality of leads. The morphology features are graphed by presenting a component ring associated with each of the morphology features. The system includes at least two sensors configured to collect physiological data. A processor is configured to extract amplitudes of morphology features. A graphical user interface is configured to graphically display the morphology features in a plurality of component rings. | 1. A method of presenting physiological data, the method comprising: receiving a plurality of leads of physiological data, the leads taken with a plurality of sensors from at least one plane through a patient;extracting morphology features from each of the plurality of leads; andgraphing the morphology features from each of the plurality of leads on a graphical user interface by presenting a component ring with a ring height associated with each of the morphology features and graphing each of the extracted morphology features in each of the plurality of leads in the associated component ring. 2. The method of claim 1, wherein the physiological data is electrocardiographic (ECG) data and the morphology features are an amplitude and duration of each of the P wave, QRS wave, ST segment, and T wave. 3. The method of claim 2, wherein the morphology features further include an amplitude and duration of the U wave. 4. The method of claim 1, further comprising dynamically modifying the height of each of the component rings to be the highest absolute value of the amplitude of the morphology features of each ring. 5. The method of claim 1, wherein the physiological data is ECG data and the at least one plane is a frontal plane and the plurality of leads are I, II, III, aVR, aVL, and aVF. 6. The method of claim 1, wherein the physiological data is ECG data and the at least one plane is a horizontal plane and the plurality of leads comprises at least V1, V2, V3, V4, V5, and V6. 7. The method of claim 1, wherein the at least one plane is a three-dimensional sphere. 8. The method of claim 7, wherein the three-dimensional sphere is extrapolated from leads in the frontal plane and the horizontal plane. 9. The method of claim 1, further comprising: indicating areas wherein the amplitude of a morphology feature is positive; andindicating areas wherein the amplitude of a morphology feature is negative. 10. The method of claim 9, wherein the absolute value of the amplitude of the morphology feature is presented in the component rings and the polarity of the amplitude is visually indicated. 11. The method of claim 9, wherein the indicating steps are performed using arrows within the component rings. 12. A computer readable medium comprising computer readable code embodying a computer program that upon execution by a processor performs the steps of: extracting morphology features for each of a plurality of electrocardiographic (ECG) leads;forming a first plurality of component rings on a graphical user interface, each component ring in the plurality being associated with one of the extracted morphology features, each component ring having an area and representing an amplitude of the associated morphology feature with respect to each of the leads;indicating in the associated component ring, the amplitudes of each of the morphology features extracted from each of the ECG leads. 13. The computer readable medium of claim 12, further comprising presenting the c... | General Electric Company | A61B005/0402
US08285490 | B2 | 2012-10-09 | Method and system for obtaining improved data perfusion measurements | 2010-11-29 | A method of deriving blood perfusion indices for a region of interest (ROI) of a subject, comprising the steps of administering a contrast agent to the subject during a dynamic imaging scan, converting signal intensity data from raw images of the scan into contrast agent concentration data, deriving parameters from the contrast agent concentration data using at least one transport function that accounts for delay and dispersion of the contrast agent, and calculating the blood perfusion indices from the derived parameters. | 1. A method of estimating blood perfusion indices, for a region of interest (ROI) by operating a computer program on intensity data input to a computer comprising: a. providing a contrast agent to an artery, wherein said artery is upstream from a region of interest (ROI), wherein said ROI comprises tissue;b. measuring intensity data to provide a tissue contrast agent curve C(t) in said tissue, wherein said C(t) is a known intensity profile measured against time;c. measuring an arterial input function AIFa(t) in a vessel leading to said ROI, wherein said AIFa(t) is known intensity profile measured against time;d. estimating a tissue blood flow Ft and a tissue impulse residue function R(t) by deconvolving C(t)=(Ft/kH)AIFa(t)R(t), wherein kH=(1−Ha)/(1−Ht) is a correction constant using different values of an arterial hematocrit Ha and a tissue hematocrit Ht, wherein said Ha has a value of about 0.45, wherein said Ht has a value of about 0.25, wherein when said AIFa(t) in a major artery comprises no delay and no dispersion, said C(t) can be derived by a convolution of said AIFa(t) and said R(t);e. determining a blood flow rate Ft and a tissue impulse residue function Re(t) using said C(t) and a tissue input function AIFt(t) by deconvolving the relation C(t)=(Ft/kH)AIFt(t)Re(t), wherein is the convolution operator, wherein said AIFt(t) comprises AIFt(t)=AIFa(t)ha(t), wherein ha(t) is a vascular transport function, wherein for t&lt;t1, ha(t)=0 and fort≥t1,ha⁡(t)=1σ1⁢ⅇ-(t-t1)/σ1, wherein the initial value of a delay time t1 can be determined by the maximum position of said R(t), wherein σ1=(t1)(β1)/(1−β1), wherein β1 is a known relative dispersion value having a range from 0 to 1;f. measuring an arterial input function AIFa(t) in said artery at a position that is upstream from said tissue;g. simulating a tissue contrast agent concentration curve Cs(t) using a tissue transport function hs(t), whereinCs⁡(t)=(Ft/kH)⁢AIFt⁡(t)⊗Rs⁡(t)⁢⁢and⁢⁢Rs⁡(t)=1-∫0t⁢hs⁡(τ)⁢⁢ⅆτ, wherein, saidhs⁡(t)=1σ2⁢ⅇ-(t-t2)/σ2 for t≧t2, and said hs(t)=0 for t&lt;t2, wherein t is time, wherein the intial value of a dispersion time σ2 can be dertermined by 1/PH and the initial vaule of a minimum transit time t2 can be determined by MTT−σ2, wherein PH is a peak height and MTT is a mean transit time of a tissue transport function he(t) in said tissue, wherein he(t) is derived from said Re(t), wherein he(t)=−dRe(t)/dt;h. using an iterative least squares method to fit the said simulated Cs(t) to said measured tissue curve C(t) with four adjustable parameters (Ft, t1, t2 and σ2);i. using a model-free deconvolution method to estimate starting values of said adjustable parameters for said iterative least squares fitting process in order to derive optimized values of said adjustable parameters; andj. calculating perfusion indices from the said optimized values of the adjustable parameters. 2. The method of claim 1, wherein said method further comprises using an adiabatic approximation to a hom... | Apollo Medical Imaging Technology Pty Ltd | G01N033/48
US07937136 | B2 | 2011-05-03 | Cardiac mapping, including resolution map | 2009-07-23 | A non-contact cardiac mapping method is disclosed that includes: (i) inserting a catheter into a heart cavity having an endocardium surface, the catheter including multiple, spatially distributed electrodes; (ii) measuring signals at the catheter electrodes in response to electrical activity in the heart cavity with the catheter spaced from the endocardium surface; and (iii) determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface. Related systems and computer programs are also disclosed. | 1. A method comprising: inserting a catheter into a heart cavity having an endocardium surface, the catheter comprising multiple, spatially distributed electrodes;measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the catheter electrodes with respect to the endocardium surface, wherein determining the physiological information comprises determining the physiological information based at least in part on a mathematical operator approximating Laplace's equation;computing values indicative of a degree of spatial resolution of the determined physiological information for at least some of the locations of the endocardium surface, the degree of spatial resolution being based at least in part on a distance between the catheter electrodes and the endocardium surface; anddisplaying information related to at least some of the determined physiological information and information related to at least some of the computed resolution values. 2. The method of claim 1, further comprising: moving the catheter to each of multiple, different positions in the heart cavity; andfor each of the different catheter positions, determining the positions of the catheter electrodes with respect to the endocardium surface and measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;wherein the determination of the physiological information at the multiple locations of the endocardium surface is based further on the positions of the catheter electrodes and the measured signals at the different catheter positions. 3. The method of claim 2, wherein determining the position of the catheter electrodes comprises measuring information about at least one of a position and orientation of the catheter within the heart cavity. 4. The method claim 3, wherein determining the position of the catheter electrodes is based further on information about the distribution of the electrodes on the catheter. 5. The method of claim 4, wherein measuring information about at least one of a position and orientation of the catheter within the heart cavity comprises measuring the position of one or more catheter electrodes within the heart cavity. 6. The method of claim 2, wherein determining the position of the catheter electrodes comprises directly measuring the position of each catheter electrode within the heart cavity. 7. The method of claim 1, wherein the computed values are derived, at least in part, from a transformation function relating the physiological information at the multiple locations of the endocardium surface to the signals measured by the catheter electrodes. 8. The method of claim 1, wherein displaying at least a portion of the endocardium surface to include information related to at least some of the determined physiological information further comprises ... | Rhythmia Medical, Inc. | A61B005/02
US07942817 | B2 | 2011-05-17 | Patient monitoring and treatment medical signal interface system | 2008-12-09 | A medical signal interface device bidirectionally conveys signals between a patient and patient monitoring devices. The device comprises a bidirectional electrical signal interface that receives and buffers patient parameter monitoring signals received from a patient via patient attached leads and outputs treatment related signals used in applying invasive or non-invasive treatment to a patient. A bidirectional electrical signal processor operates in response to commands received from a control processor and is coupled to the electrical signal interface, for processing received patient parameter monitoring signals using filtering and amplification to provide processed patient monitoring signals for output to at least one patient monitoring device. The bidirectional electrical signal processor processes the treatment related signals for output by buffering the treatment related signals for output to a patient. A control processor provides data representing the commands in response to at least one of, (a) predetermined configuration data and (b) deriving data representing the commands from data entered by a user via a displayed user interface image. | 1. A medical signal interface device for bidirectionally conveying signals between a patient and patient monitoring and treatment devices, comprising: a bidirectional electrical signal interface for, receiving and buffering patient parameter monitoring signals received from a patient via patient attached leads andconcurrently outputting treatment related signals used in applying invasive or non-invasive treatment to a patient and providing electrical isolation of said patient parameter monitoring signals received from said patient via patient attached leads from treatment related signals used in applying invasive or non-invasive treatment output to a patient;a bidirectional electrical signal processor, responsive to commands received from a control processor and coupled to said electrical signal interface, for processing received patient parameter monitoring signals using filtering and amplification to provide processed patient monitoring signals for output to at least one patient monitoring device and for processing the treatment related signals for output by buffering the treatment related signals for output to a patient; anda control processor for providing data representing the commands in response to at least one of, (a) predetermined configuration data and (b) deriving data representing the commands from data entered by a user via a displayed user interface image. 2. A system according to claim 1, wherein said control processor adaptively configures said bidirectional electrical signal processor by providing data representing the commands, the commands being determined in response to data identifying said at least one patient monitoring device andsaid bidirectional electrical signal interface receives, buffers and analog to digitally converts patient parameter monitoring signals received from a patient via patient attached leads. 3. A system according to claim 2, wherein said control processor adaptively configures said bidirectional electrical signal processor by providing data representing the commands, the commands being determined in response to data identifying a device used in applying said invasive or non-invasive treatment to said patient andsaid bidirectional electrical signal interface performs digital to analog conversion in providing said treatment related signals used in applying invasive or non-invasive treatment to said patient. 4. A system according to claim 1, wherein said bidirectional electrical signal processor, provide electrical isolation of signals provided to, and received from, said patientsaid bidirectional electrical signal interface provides separate isolated safety protection processing of, (a) said patient parameter monitoring signals received from said patient via patient attached leads and(b) treatment related signals used in applying invasive or non-invasive treatment output to a patient. 5. A system according to claim 4, wherein said separate isolated safety protection processing comprises voltage clamp... | Siemens Medical Solutions USA, Inc. | A61B005/00 A61B005/04 A61N001/37
US08396538 | B2 | 2013-03-12 | Method and apparatus for improving signal to noise ratio of ECG signals to facilitate cardiac beat detection | 2010-11-24 | Improving the signal-to-noise ratio of electrocardiogram (ECG) measurements facilitates cardiac beat detection in a human or animal patient. ECG signals measured either on the surface of the skin or subcutaneously from pairs of leads may be combined to calculate a differential signal. The measured signal may also be averaged to produce a second estimate. A point-by-point product of the differential signal and the averaged signal is generated if the sample pair has the same polarity. The product signals feature enhanced signal components and reduced noise components, thus improving the signal-to-noise ratio of the to respective input channels. Product signals are then subjected to peak detection through a conventional auto-sensing approach. Preliminary event detection results from the conditional product signals may then be aggregated, and final sense markers for ECG beat detection may be generated by means of a voting algorithm. | 1. A method of enhancing the signal-to-noise ratio (SNR) of measured electrocardiogram (ECG) signals, the method comprising: providing at least three cardiac input signals derived from the measured ECG signals;forming a first estimate from each of at least three pairs of input signals;forming a second estimate from each of the input signals; andobtaining conditional product (CP) signals by generating, from each first estimate and each second estimate, a conditional product wherein the obtained CP signals exhibit increased signal components and reduced noise components, thus producing an enhanced ratio of signal to noise components (SNR). 2. The method of claim 1, wherein the first estimate is a differential signal and the second estimate is an averaged signal formed by using a weighted moving average filter. 3. The method of claim 1, wherein the measured ECG signals are of a type obtained from surface measurements. 4. The method of claim 1, wherein the measured ECG signals are of a type obtained from subcutaneous measurements. 5. The method of claim 1, wherein the enhanced SNR facilitates cardiac beat detection when the obtained conditional product signals are subject to peak detection using an auto-sensing approach. 6. The method of claim 2, wherein the conditional product depends on the relative polarities of the differential and averaged signals, such that if the differential and averaged signals have the same polarity, then the conditional product is obtained by calculating a sample-by sample-product of the differential signal and the average signal. 7. The method of claim 2, wherein the conditional product depends on the relative polarities of the differential and averaged signals, such that if the differential and averaged signals have different polarity, then the conditional product is set to zero. 8. A cardiac device for use in detecting heartbeats, to which at least three external sensing electrodes are attached to provide at least three analog voltage input signals, the device comprising an electrocardiogram (ECG) sensing unit for use in producing digital signals from the analog voltage input signals, wherein the sensing unit includes: a. a multiplexer;b. signal processing units; andc. a controller configured to enhance the signal to noise ratio (SNR) of the digital signals by generating conditional product signals from differential signals and weighted average signals, thereby facilitating heartbeat detection. 9. The cardiac device of claim 8, wherein if the differential signals and average signals have the same polarity, then the conditional product is obtained by calculating a sample-by sample-product of the differential signal and the average signal. 10. The cardiac device of claim 8, wherein if the differential signals and average signals have different polarity, then the conditional product is set to zero. 11. The cardiac device of claim 8, wherein the signal processing units are selected from the group consisting of dif... | Biotronik CRM Patent AG | A61B005/04
US08512367 | B2 | 2013-08-20 | Blood sampling device with dual-link drive mechanism | 2010-02-25 | A housing and an internal mechanism including a blade, a dual-link drive mechanism, and a trigger. The blade is driven by the drive mechanism through a rotating and translating slicing motion that defines a sampling sequence, from a first retracted position to an extended position to a second retracted position. The dual-link drive mechanism includes a blade link arm coupled to the blade, a rotary drive link arm that drives the blade arm, a drive spring that drives the a rotary drive arm, and a cam-and-follower mechanism that guides the blade arm as it is driven by the rotary drive arm. The trigger is operable to release the dual-link drive mechanism to propel the blade through the sampling sequence. And a sterility cap protects the blade and prevents operation of the device prior to removal. | 1. A fluid-sampling device for penetrating skin, comprising: a housing having a blade opening;a blade having a sharp edge with a tip that travels along a travel path from a retracted position shielded within the housing to an extended position extending through the blade opening to penetrate the skin;a dual-link drive mechanism including a blade link arm coupled to the blade, a rotary drive link arm that is pivotally coupled to and drives the blade arm, and a cam-and-follower guide mechanism that guides movement of the blade arm, wherein the drive arm and the cam-and-follower mechanism cooperatively drive and guide the blade arm to propel the blade tip along the travel path in a rotational and translating motion, wherein the cam-and-follower mechanism includes a non-linear cam surface and a cam follower that is guided by the non-linear cam surface; anda trigger operable to actuate the drive mechanism to propel the blade tip along the travel path;wherein as the drive arm rotationally drives the blade arm, the blade arm rotates about the cam follower as the cam follower is translationally guided along the non-linear cam surface in a non-linear path to convert the rotary motion of the rotary drive arm to the rotational and translating motion of the blade tip. 2. The fluid-sampling device of claim 1, wherein the non-linear cam surface is defined by a channel formed in the housing and the cam follower is provided by a pin extending from the blade arm and slidingly received in the channel. 3. The fluid-sampling device of claim 1, wherein the cam follower is intermediately positioned on the blade arm between the blade and the pivotal coupling to the drive arm. 4. The fluid-sampling device of claim 1, wherein the non-linear cam surface is generally vertical and curved so that the blade travel path is non-linear and non-symmetrical. 5. The fluid-sampling device of claim 4, wherein the non-linear cam surface is curved so that the blade travel path is generally triangular. 6. The fluid-sampling device of claim 5, wherein the non-linear cam surface is curved so that the generally triangular blade travel path has a descent segment and an ascent segment that is steeper than the descent segment. 7. The fluid-sampling device of claim 1, wherein the drive mechanism further comprises a drive spring that biases the rotary drive arm in a rotary drive direction. 8. The fluid-sampling device of claim 1, wherein the drive mechanism defines a catch surface and the trigger defines a catch surface that engages the drive mechanism catch surface to retain the drive arm in a ready position. 9. A fluid-sampling device for penetrating skin, comprising: a housing having a blade opening;a blade having a sharp edge with a tip that travels along a travel path from a first retracted position shielded within the housing, through an extended position extending through the blade opening to penetrate the skin, and to a second retracted position shielded within the housing;a dua... | Facet Technologies, LLC | A61B005/00
US08398558 | B2 | 2013-03-19 | Automatic transmission power adjustment for Doppler radar | 2009-07-03 | The invention relates to a Doppler radar heart sensor comprising a radar signal emitter (102); a radar signal receiver (103); a signal pattern matcher (106); and a transmission power adjuster (108). The signal pattern matcher (106) is connected to the radar signal receiver (102) and arranged to analyze a received radar signal (105), to compare the received radar signal with signal patterns, and to issue a corresponding indication signal upon detection of a sufficiently high similarity between the received radar signal and at least one of the signal patterns. A transmission power adjuster (109) is connected to the radar signal emitter (102) and the signal pattern matcher (106), the transmission power adjuster being arranged to gradually modify a transmission power delivered to the radar signal emitter (102) until the signal pattern matcher issues the indication signal. The invention also relates to a corresponding method and a computer program product. | 1. Doppler radar heart sensor comprising: a radar signal emitter;a radar signal receiver;a signal pattern matcher connected to the radar signal receiver and arranged to analyze a received radar signal, to compare the received radar signal with signal patterns, and to issue a corresponding indication signal upon detection of a sufficiently high similarity between the received radar signal and at least one of the signal patterns;a transmission power adjuster connected to the radar signal emitter and the signal pattern matcher, the transmission power adjuster being arranged to gradually modify a transmission power delivered to the radar signal emitter until the signal pattern matcher issues the indication signal. 2. Doppler radar sensor according to claim 1, wherein the signal pattern matcher compares properties of the received radar signal with properties of typical heart beat signals. 3. Doppler radar heart sensor according to claim 1, wherein the signal patterns correspond to Doppler radar sensed signals of heart activity. 4. Doppler radar heart sensor according to claim 1, wherein the signal patterns comprise Doppler radar sensed signals of known calibration movements in a coverage area of the radar signal receiver. 5. Doppler radar heart sensor according to claim 4, further comprising a moving object for generating the known calibration movements. 6. Doppler radar heart sensor according to claim 1, further comprising a memory for calibration signal quality measurements of a prior calibration session, wherein the signal pattern matcher determines a current signal quality and the transmission power adjuster stops modifying the transmission power when the current signal quality measurement is equal to or better than the calibration signal measurement. 7. Doppler radar heart sensor according to claim 1, further comprising an electrocardiogram signal interface connected to the signal pattern matcher, wherein the signal pattern matcher compares the received radar signal also to a signal provided via the electrocardiogram interface. 8. Method for adjusting a transmission power of a Doppler radar heart sensor, the method comprising: a) sending Doppler radar signals;b) receiving a received radar signal;c) analyzing the received radar signal;d) comparing the received radar signal with signal patterns;e) if said comparing reveals a sufficiently high similarity between the received radar signal and at least one of the signal patterns terminating the method for adjusting, or else adjusting the transmission power by gradually modifying the transmission power and repeating actions a) to e). 9. Method according to claim 8, wherein adjusting the transmission power comprises gradually increasing the transmission power from a low level to a level at which the sufficiently high similarity is detected. 10. Method according to claim 8, wherein comparing the received radar signal with signal patterns comprises comparing properties of the received radar signa... | Koninklijke Philips Electronics N.V. | A61B005/04 A61B008/00 A61B005/05
US08380293 | B2 | 2013-02-19 | Cardiac activation sequence monitoring and tracking | 2011-02-15 | Cardiac monitoring and/or stimulation methods and systems provide monitoring, diagnosis, and defibrillation and/or pacing therapies. A signal processor receives a plurality of composite signals associated with a plurality of sources, performs a source separation, and produces one or more cardiac signal vectors associated with all or a portion of one or more cardiac activation sequences based on the source separation. A method of signal separation involves detecting a change in a characteristic of the cardiac signal vector relative to a baseline. One or more vectors and/or activation sequences may be selected, and information associated with the vectors and/or activation sequences may be stored and tracked. | 1. A processor-implemented method, comprising: sensing a plurality of composite cardiac signals over time at a plurality of patient-internal locations;performing a plurality of source separations using sensed composite cardiac signals over time;producing a cardiac signal vector indicative of a cardiac activation sequence of an area of the patient's heart based on each of the plurality of source separations;monitoring a characteristic of the cardiac signal vectors relative to a baseline established for the characteristic;detecting a change in the characteristic relative to the baseline; andidentifying an anomalous cardiac condition based on the detected change;wherein the processor implements at least some of the method. 2. The method according to claim 1, wherein performing at least some of the plurality of source separations is initiated periodically or in response to occurrence of a predetermined event. 3. The method according to claim 1, wherein performing at least some of the plurality of source separations is initiated manually. 4. The method according to claim 1, wherein performing at least some of the plurality of source separations is initiated in response to a change in a therapy delivered to the patient or a change in patient posture. 5. The method according to claim 1, comprising establishing the baseline of the cardiac signal vector characteristic by performing an initial source separation using sensed composite cardiac signals. 6. The method according to claim 1, comprising establishing the baseline of the cardiac signal vector characteristic using clinical data. 7. The method according to claim 1, comprising tracking changes of the cardiac signal vector characteristic relative to the baseline over time. 8. The method according to claim 1, comprising tracking acute and chronic changes of the cardiac signal vector characteristic relative to the baseline over time. 9. The method according to claim 1, comprising trending changes of the cardiac signal vector characteristic relative to the baseline over time. 10. The method according to claim 1, comprising one or more of initiating, adjusting, optimizing, and terminating therapy delivery to the patient in response to detecting the change in the characteristic relative to the baseline. 11. The method according to claim 1, wherein the cardiac activation sequence comprises one or both of a cardiac depolarization sequence and a cardiac repolarization sequence. 12. An apparatus, comprising: a plurality of implantable electrodes configured for sensing composite cardiac signals over time;a housing configured for implantation in a patient; anda processor provided in the housing and coupled to the plurality of implantable electrodes, the processor configured to perform a plurality of source separations using sensed composite cardiac signals over time, produce a cardiac signal vector indicative of a cardiac activation sequence of an area of the patient's heart based on each of the plural... | Cardiac Pacemakers, Inc. | A61B005/0428 A61B005/0452
US08636671 | B2 | 2014-01-28 | Wearable respiration measurement apparatus | 2010-09-09 | A wearable respiration measurement apparatus including: an eyeglass frame having a nose flap which is formed in such a manner as to cover a user's nose; and an induced current generation device attached to the nose flap and generating an induced current corresponding to an amount of air inhaled or exhaled through the user's nose. | 1. A wearable respiration measurement apparatus comprising: an eyeglass frame including a nose flap that covers a user's nose; andan induced current generation device attached to the nose flap and generating an induced current corresponding to an amount of air inhaled through the user's nose,wherein the induced current generation device comprises:a body having provided therein a structure defining an exhalation path and a structure defining an inhalation path housed within the structure defining the exhalation path, the structure defining the inhalation path including an inhalation valve plate, and the structure defining the exhalation path including an exhalation valve plate;wherein the exhalation valve plate is provided at an upper portion of the exhalation path and is configured to be open during exhalation and closed during inhalation,wherein the inhalation valve plate is provided at an upper portion of the inhalation path and is configured to be open during inhalation and closed during exhalation,wherein the exhalation valve plate is provided over the inhalation valve plate, andwherein a path extends from the inhalation valve plate to the exhalation valve plate to form a hole in the exhalation valve plate so that inhaled air passes through the path and the hole. 2. The wearable respiration measurement apparatus of claim 1, wherein the induced current generation device comprises: a propeller installed in the inhalation path so as to be rotated by air inhaled through the inhalation path; anda plurality of coils installed on an inner surface of the structure defining the inhalation path, the plurality of coils generating the induced current-,wherein the induced current corresponds to alternating magnetic fields caused by rotation of the propeller. 3. The wearable respiration measurement apparatus of claim 2, wherein the propeller comprises a plurality of blades having a magnetic property, wherein the plurality of blades includes odd-numbered blades and even-numbered blades, andwherein the odd-numbered blades have an opposite magnetic polarity to the even-numbered blades. 4. The wearable respiration measurement apparatus of claim 2, wherein the plurality of coils includes odd-numbered coils and even-numbered coils, and wherein the odd-numbered coils have an opposite polarity to the even-numbered coils. 5. The wearable respiration measurement apparatus of claim 2, wherein the plurality of coils are coupled in parallel to each other. 6. The wearable respiration measurement apparatus of claim 2, wherein the body further comprises: a frame fixing a position of the inhalation path inside of the exhalation path and fixing a position of the propeller. 7. The wearable respiration measurement apparatus of claim 1, wherein the nose flap is implemented in such a manner as to cover only the nose or cover both the nose and the mouth. 8. The wearable respiration measurement apparatus of claim 7, further comprising a mouthpiece for inducing respiration ... | Electronics and Telecommunications Research Institute | A61B005/08
US09445745 | B2 | 2016-09-20 | Tool shape estimation | 2010-10-18 | A medical tool includes a location sensor at a distal end. While the tool is being maneuvered to a desired region of interest, a control unit records a first location reading obtained according to an output from the sensor when requested by a user, indicative of when a patient's anatomy constrains free movement of the tool in at least one degree of freedom. Once the tool has reached the region of interest, the control unit records a second location reading according to the sensor output indicative of a current location. A relaxation shape of the tool forms a shape constraint on the tool shape reconstruction. The control unit determines a reconstruction representing the current shape of the tool based on the positional constraints and the shape constraint. The reconstruction is graphically superimposed on an image of the region of interest. | 1. An apparatus for determining a shape of a medical tool in a region of interest in a patient's body, comprising: a medical positioning system configured to produce location readings with respect to points within said region of interest in accordance with an output of a location sensor connected to said tool moving within said region of interest, each location reading being associated with a three-dimensional reference coordinate system; anda control unit including a main processor and an associated memory wherein said control unit is coupled to said medical positioning system, said control unit being configured to execute pre-programmed instructions stored in said associated memory to determine a reconstruction representing the shape of said tool as a function of (1) a positional constraint that includes a first location corresponding to a first location reading produced by the medical positioning system, said first location reading indicating a constricting point along a tool path where said patient's body anatomically constrains movement of said tool in at least one degree of freedom, and (2) a shape constraint corresponding to a relaxation shape of said tool, wherein said control unit is configured through a user interface to receive an input from a user to record said first location reading produced by said medical positioning system, and wherein said signal to record said first location reading is received at a time when the medical tool is constrained by said patient's body. 2. The apparatus of claim 1 further including a superimposing processor configured to graphically superimpose said reconstruction on an image of said region of interest acquired at an earlier time. 3. The apparatus of claim 1 wherein said reconstruction representing the shape of the tool comprises a function or a plurality of functions. 4. The apparatus of claim 1 wherein said reconstruction representing the shape of the tool comprises a spline. 5. The apparatus of claim 1 wherein said positional constraint further includes a second location in said coordinate system corresponding to a second location reading, said second location indicating the current location of said tool in said region of interest. 6. The apparatus of claim 5 wherein said location sensor in said tool is a first location sensor, said tool further including a second location sensor wherein said positional constraint further includes a third location corresponding to a third location reading, said third location indicating the location of said second location sensor. 7. The apparatus of claim 6 wherein at least one of said first, second and third locations comprise a respective position and orientation in said coordinate system. 8. The apparatus of claim 1 wherein said relaxation shape is determined in accordance with a model, said model approximating the relaxation shape of said tool in three-dimensional space. 9. The apparatus of claim 8 wherein said model comprises a polynomial expression.... | MediGuide Ltd. | A61B005/00 A61B005/06 A61B008/08
US10624553 | B2 | 2020-04-21 | Probe data mapping using contact information | 2009-12-08 | A method of mapping includes receiving inputs measured by a probe at respective locations inside a body cavity of a subject. At each of the respective locations, a respective contact quality between the probe and a tissue in the body cavity is measured. The inputs for which the respective contact quality is outside a defined range are rejected, and a map of the body cavity is created using the inputs that are not rejected. | 1. A method of mapping a body cavity of a subject using a probe having a distal tip and being operatively connected to a processor, comprising: contacting the distal tip of the probe with tissue at a location in the cavity of the subject causing the distal tip to deflect, the probe having a distal end configured for insertion into the body cavity of the subject and further comprising a contact force sensor including a position sensor, a first magnetic field transmitter, and a resilient member comprising a spring coil with a helical cut in the distal end of the probe, the distal end having the distal tip and an electrode, the electrode configured to acquire physiological electrical activity in tissue, the coil spring configured to permit deflection of the distal tip relative to the distal end;generating one or more first signals with the position sensor and the first magnetic field transmitter of the contact force sensor indicative of a deflection of the distal tip resulting from a contact pressure exerted on the tissue by the distal tip, and transmitting the one or more first signals to the processor;generating one or more second signals with the position sensor and a second magnetic field transmitter, the one or more second signals distinguishable from the one or more first signals and indicative of location and orientation coordinates of the distal end of the probe, and transmitting said one or more second signals to the processor;in the processor, measuring the deflection of the distal tip at the location and calculating a respective contact quality between the probe and the tissue in the body cavity at the location using the one or more first signals;rejecting the one or more first signals for which the contact quality is outside a defined range;in the processor, automatically collecting a map point only when the calculated respective contact quality is within the defined range, the map point including coordinate reading from the position sensor responsive to the second magnetic field transmitter and electrical signals from the electrode that are indicative of the physiological electrical activity in the tissue; creating or updating a map of the body cavity using the map point for the one or more first signals that are not rejected; and wherein the processor further configured to display the map of the body cavity. 2. The method according to claim 1, wherein the body cavity comprises a chamber of a heart. 3. The method according to claim 1, wherein measuring the respective contact quality comprises measuring an electrical impedance between the probe and the tissue. 4. The method according to claim 1, wherein creating the map comprises adding tags to the map indicating the respective contact quality at one or more of the locations. 5. The method according to claim 1, wherein measuring the respective contact quality comprises measuring a pressure exerted on the distal end of the probe. 6. The method according to claim 5, wherein rejecting ... | Biosense Webster (Israel), Ltd. | A61B005/042 A61B005/00 A61B090/00 A61B005/06 A61B034/20
US07946995 | B1 | 2011-05-24 | Analyzing circadian variations of a hemodynamic parameter to determine an adverse cardiac condition | 2006-11-09 | A system and method of determining the status of an adverse cardiac condition of a medical patient based on circadian variation of one or more hemodynamic parameters are provided. In some embodiments, the system and method calculate a first average value of a series of first values during a first time period, a second average value of a series of second values during a second time period, and a difference between the first average value and the second average value. The method provides an indication of an adverse cardiac condition when the difference is less than a predetermined threshold. | 1. A method of determining the status of an adverse cardiac condition of a medical patient, the method comprising: determining a series of first values of a hemodynamic parameter during a first time period from a sensor implanted within the medical patient, wherein the first time period is during the daytime;determining a series of second values of the hemodynamic parameter during a second time period from the sensor, wherein the second time period is during the nighttime;calculating, with a signal processor, a first average value of the series of first values and a second average value of the series of second values;calculating a difference between the first average value and the second average value; andproviding an indication of an adverse cardiac condition when the difference is less than a first predetermined threshold value;wherein the hemodynamic parameter comprises a left atrial pressure of the patient's heart. 2. The method of claim 1, wherein the first and second periods correspond to non-overlapping portions of a circadian rhythm. 3. The method of claim 1, wherein the adverse cardiac condition comprises a heart failure exacerbation. 4. The method of claim 1, wherein the adverse cardiac condition comprises an acute cardiac event. 5. The method of claim 1, wherein the adverse cardiac condition is heart failure. 6. The method of claim 1, wherein the adverse cardiac condition is pulmonary edema. 7. The method of claim 1, wherein the first predetermined threshold value is in a range of about 30-50% of a historical variation of the patient. 8. The method of claim 1, wherein the first predetermined threshold value is in a range of about 50-70% of a historical variation of the patient. 9. The method of claim 1, further comprising calculating a mean value of the first average value and the second average value; and providing an indication of an adverse cardiac condition when the difference is less than a first predetermined threshold value and the mean value is greater than a second predetermined threshold. 10. An apparatus for treating cardiovascular disease in a medical patient, the apparatus comprising: a housing configured to be implanted within a medical patient;an implantable sensor, operable to generate a sensor signal indicative of a hemodynamic parameter of the medical patient;at least one implantable lead having a proximal end and a distal end, the implantable lead coupled to the implantable housing at the proximal end and further coupled to the sensor;a signal processor located within the housing, the signal processor operable to monitor circadian variations in the hemodynamic parameter of a medical patient by comparing a first data set taken during the daytime with a second set taken during the nighttime and determine an occurrence of an adverse cardiac condition based upon a change in the circadian variations over a predetermined time period;wherein the hemodynamic parameter comprises a left atrial pressure of the patient'... | Pacesetter, Inc. | A61B005/02
US07948362 | B2 | 2011-05-24 | Implantable medical transceiver device | 2007-06-27 | An implantable medical device having a transceiver for transmitting/receiving wirelessly transmitted data, which is turned off or switched into an energy-saving rest state between individual data transmissions via the transceiver and having a waking unit which is implemented to switch the transceiver by a waking signal from its turned-off state or its rest state into its fully operational state, the waking unit having a low-power receiver and a waking control unit of which the low-power receiver is implemented to monitor multiple predefined frequency ranges in such a way that in case of a transmission of sufficient signal strength in one of the frequency ranges, it generates an output signal and outputs it to the waking control unit and of which the waking control unit is implemented to analyze output signals of the low-power receiver and output a waking signal to the transceiver which switches it on or to fully operational. | 1. An implantable medical device (20) comprising: a transceiver (203) that comprises a first transmitter and first receiver configured to transmit and receive wirelessly transmitted data respectively, wherein said transceiver is turned off or switched into an energy-saving rest state between individual data transmissionsa waking unit (202), which is configured to switch said transceiver (203) from its turned-off state or its rest state into its fully operational state by a waking signal;wherein said waking unit comprises a second receiver independent of said first receiver wherein said second receiver is a low-power receiver (210) that requires less energy as said transceiver in said transceiver's fully operational state anda waking control unit (206), of which said low-power receiver (210) is configured to monitor multiple predefined frequency ranges in such a way that, in case of a transmission of sufficient signal strength in one or more frequency ranges said low-power receiver (210) generates an output signal and outputs said output signal to said waking control unit (206), and of which said waking control unit (206) is configured to analyze output signals of said low-power receiver (210) and to output said waking signal to said transceiver (203), which switches said transceiver (203) on or to fully operational, if a predefined condition is met or multiple predefined conditions are met;wherein said waking control unit (206) is configured to output said waking signal to said transceiver (203) if a condition is met that said low-power receiver (210) outputs a sequence of output signals which identify that said low-power receiver (210) has detected a series of data transmissions of sufficient signal strength in various frequency ranges which correspond to a predefined series/sequence of frequencies; and,a transmission control unit (204), wherein said transmission control unit (204) has or is connected to a random generator (205) and is configured, after lapse of a waiting time after said transceiver (203) is turned on by said waking control unit (206), to transmit a response signal and to define a time of a transmission beginning after said transceiver (203) is turned on by said waking control unit (206) for this purpose in such a way that said time of said transmission beginning corresponds to an ending time of said waiting time, which begins with said waking signal and has a duration which corresponds to a product ZZ·SD of a random number ZZ generated by said random generator (205) and at least one predetermined average transmission duration SD. 2. The implantable medical device (20) according to claim 1, wherein said low-power receiver (210) has multiple bandpass filters (211′, 211″, 211′″, 211″″), each having an assigned signal detector (212′, 212″, 212′″, 212″″), which work together in such a way that a particular signal detector (212′, 212″, 212′″, 212″″) outputs a signal if said low-power receiver (210) receives a transmission having suf... | Biotronik CRM Patent AG | H04Q005/22 G08B013/14 H04B001/00 H04B001/10 A61N001/00 A61B005/00
US07949380 | B2 | 2011-05-24 | Oximeter sensor with digital memory encoding sensor data | 2006-06-01 | Embodiments of the present invention include systems and methods that relate to a sensor with memory. Specifically, one embodiment includes a sensor, comprising a light emitting element configured to emit light, a light detecting element configured to detect the light, and a memory storing data corresponding to algorithms used by an oximeter monitor, the memory providing access to the oximeter monitor to read the data. | 1. An oximeter sensor, comprising: a light emitting element configured to emit light;a light detecting element configured to detect the light; anda memory storing data comprising updated algorithms for use by an oximeter monitor to update functionality of the oximeter monitor, the memory providing access to the oximeter monitor to read the data. 2. The oximeter sensor of claim 1, wherein the data comprises executable software code configured to upgrade software on the monitor that employs the algorithms. 3. The oximeter sensor of claim 1, wherein the data is digitally encoded. 4. A method of manufacturing an oximeter sensor, comprising: providing an access point configured to facilitate communication with an oximeter monitor;providing a light emitting element configured to emit light;providing a light detecting element configured to detect the light; andproviding data in a memory of the sensor, the data comprising updated algorithms for use by the oximeter monitor to update functionality of the oximeter monitor. 5. A method of operation of an oximeter sensor, comprising: emitting a light into a patient with a light emitting element;detecting the light with a detecting element; andproviding data in a memory of the sensor to an oximeter monitor, the data comprising updated algorithms for use by the oximeter monitor to update functionality of the oximeter monitor. 6. An oximeter system, comprising: an oximeter sensor, comprising: a light emitting element configured to project light into a patient;a light detector configured to detect the light; anda memory storing data comprising updated algorithms for interpreting detected light; andan oximeter monitor, comprising: an output circuit configured to provide signals to the oximeter sensor; anda read circuit configured to read the data comprising updated algorithms from the memory of the oximeter sensor. 7. An oximeter monitor, comprising: a drive circuit configured to provide signals to circuitry of an oximeter sensor that is coupleable to a patient;a circuit configured to receive data from a memory of the oximeter sensor, the data comprising updated algorithms for interpreting detected light; anda processor configured to interpret light detected by the oximeter sensor using the data. 8. A method of oximeter monitor operation, comprising: providing signals to circuitry of an oximeter sensor that is coupleable to a patient;receiving signals corresponding to light detected by the oximeter sensor;receiving data from a memory of the oximeter sensor, the data comprising updated algorithms for interpreting the signals; andinterpreting the signals using the updated algorithms. | Mallinckrodt Inc. | A61B005/1455
US07963929 | B2 | 2011-06-21 | Devices and systems for sizing of a gastric pouch | 2008-09-16 | Devices, methods, and systems for measuring the size of gastric pouches. At least some embodiments include a catheter comprising an outer elongate tube, an inner elongate tube disposed within the outer elongate tube, a proximal balloon attached to the outer elongate tube, a distal balloon having a first end attached to the outer elongate tube and a second end attached to the inner elongate tube, at least two detection electrodes positioned on the outer elongate tube relative to at least two excitation electrodes, and a pressure transducer positioned on the outer elongate tube. | 1. A catheter for measuring a size of a gastric pouch, comprising: an outer elongate tube having a proximal end and a distal end;an inner elongate tube having a proximal end and a distal end, the inner elongate tube positioned at least partially within the outer elongate tube such that the inner elongate tube is capable of sliding within the outer elongate tube;a proximal balloon surrounding at least a portion of the outer elongate tube near the distal end of the outer elongate tube, the proximal balloon having a first end attached to the outer elongate tube and a second end attached to the outer elongate tube;a distal balloon surrounding a portion of the outer elongate tube distal to the proximal balloon, the distal balloon having a first end attached to the outer elongate tube and a second end attached to the inner elongate tube;at least two excitation electrodes attached to the portion of the outer elongate tube surrounded by the distal balloon, the at least two excitation electrodes capable of attachment to a power source;at least two detection electrodes attached to the outer elongate tube positioned relative to the at least two excitation electrodes, the at least two detection electrodes capable of attachment to a processor;a first pressure transducer attached to the portion of the outer elongate tube surrounded by the proximal balloon; anda second pressure transducer attached to the portion of the outer elongate tube surrounded by the distal balloon. 2. The catheter of claim 1, wherein the at least two detection electrodes comprise three or more detection electrodes. 3. The catheter of claim 1, further comprising: a first fluid passageway for injecting fluid into the proximal balloon; anda second fluid passageway for injecting fluid into the distal balloon. 4. The catheter of claim 1, wherein the inner elongate tube comprises at least one gradation at or near the proximal end of the inner elongate tube. 5. The catheter of claim 4, wherein the at least one gradation comprises at least two gradations positioned on the inner elongate tube a predetermined distance apart. 6. The catheter of claim 5, wherein the at least two gradations comprise notches on the inner elongate tube, and wherein the outer elongate tube comprises a latch at or near the proximal end of the outer elongate tube, the latch configured to engage one or more of the notches such that the inner elongate tube and the outer elongate tube are reversibly fastened. 7. The catheter of claim 6, wherein the outer elongate tube comprises a flexible portion near the distal end of the outer elongate tube, the flexible portion capable of being bent into a curve, and wherein the inner elongate tube is flexible such that the inner elongate tube is capable of sliding within the flexible portion of the outer elongate tube when the flexible portion is bent into a curve. 8. The catheter of claim 6, wherein the outer elongate tube comprises a curved portion near the distal end of the out... | DTherapeutics, LLC | A61B005/00 A61B005/103 A61B005/117 A61M029/00
US20110144462 | A1 | 2011-06-16 | MINIATURIZED MULTI-SPECTRAL IMAGER FOR REAL-TIME TISSUE OXYGENATION MEASUREMENT | 2009-03-19 | A portable multi-spectral imaging system and device is disclosed. The system includes at least one image acquisition device for acquiring an image from a subject, a filtering device to filter the light received by the image acquisition device, a processor for processing the image acquired by the image acquisition device, and a display. There is software running on the processor that determines oxygenation values of the subject based on the processed image. | 1. A portable multi-spectral imaging system, comprising: a plurality of image acquisition devices;wherein each respective image acquisition device in said plurality of image acquisition devices has a corresponding filtering device to filter out corresponding respective specific wavelengths of the light received by the respective image acquisition device, thereby acquiring a multi-spectral image from a subject, the multi-spectral image comprising an image from each of the plurality image acquisition devices, in the wavelengths of visible, near infrared, or far infrared;a processor for processing the multi-spectral image acquired by the plurality of image acquisition devices, thereby forming a processed image;a software application executing on the processor for determining oxygenation values of the subject based on the processed image; anda display for displaying at least one of the processed image and the oxygenation values. 2. The portable multi-spectral imaging system of claim 1, wherein the oxygenation values are based on (a) at least one of oxyhemoglobin, deoxyhemoglobin, and oxygen saturation levels of the subject measured from the processed image, or (b) a component analysis. 3. The portable multi-spectral imaging system of claim 1, wherein the software application filters the multi-spectral image to reduce noise. 4. The portable multi-spectral imaging system of claim 1, wherein multiple multi-spectral images are acquired of the subject by the plurality of image acquisition devices, and wherein optionally the software application corrects at least one of the multi-spectral images in the multiple multi-spectral images to account for motion of the subject. 5. The portable multi-spectral imaging system of claim 1, wherein the software application further comprises an imaging acquisition parameter dynamics module for (i) on-the-fly adjustment of an image sensor parameter and (ii) controlling a lighting condition present during an image acquisition phase. 6. The portable multi-spectral imaging system of claim 1, wherein the software application eliminates extraneous objects from the multi-spectral image. 7. The portable multi-spectral imaging system of claim 1, wherein the software application compares the data of the multi-spectral image to stored data. 8. The portable multi-spectral imaging system of claim 1, further comprising an illumination source, wherein the illumination source is in electrical communication with the software application. 9. The portable multi-spectral imaging system of claim 1, wherein the system communicates with an analysis device. 10. The portable multi-spectral imaging system of claim 9, wherein a communication between the system and the analysis device is one of wired and wireless. 11. The portable multi-spectral imaging system of claim 9, wherein the analysis device performs at least one of the following actions: comparison of multiple multi-spectral images, storage of acquired multi-spectral images, and ... | - | A61B005/1455
US20110105857 | A1 | 2011-05-05 | IMPRESSION DEGREE EXTRACTION APPARATUS AND IMPRESSION DEGREE EXTRACTION METHOD | 2009-04-14 | An impression degree extraction apparatus which precisely extracts an impression degree without imposing a strain on a user in particular. A content editing apparatus (100) comprises a measured emotion property acquiring section (341) which acquires measured emotion properties which show an emotion having occurred in the user in a measurement period, and an impression degree calculating part (340) which calculates the impression degree being a degree which shows how strong the user was impressed in the measurement period by comparing reference emotion properties which shows an emotion having occurred in the user in a reference period and the measured emotion properties. The impression degree calculating part (340) calculates the impression degree to be higher with the increase of the difference between the first emotion properties and the second emotion properties with the second emotion properties as the reference. | 1. An impression degree extraction apparatus comprising: a first emotion characteristic acquisition section that acquires a first emotion characteristic indicating a characteristic of an emotion that has occurred in a user in a first period; andan impression degree calculation section that calculates an impression degree that is a degree indicating intensity of an impression received by the user in the first period by means of a comparison of a second emotion characteristic indicating a characteristic of an emotion that has occurred in the user in a second period different from the first period with the first emotion characteristic. 2. The impression degree extraction apparatus according to claim 1, wherein the impression degree calculation section calculates the impression degree as higher the greater a difference between the first emotion characteristic and the second emotion characteristic as a reference. 3. The impression degree extraction apparatus according to claim 1, further comprising a content editing section that performs content editing based on the impression degree. 4. The impression degree extraction apparatus according to claim 1, further comprising: a biological information measurement section that measures biological information of the user; anda second emotion characteristic acquisition section that acquires the second emotion characteristic, wherein:the first emotion characteristic acquisition section acquires the first emotion characteristic from the biological information; andthe second emotion characteristic acquisition section acquires the second emotion characteristic from the biological information. 5. The impression degree extraction apparatus according to claim 1, wherein the second emotion characteristic and the first emotion characteristic are at least one of an emotion measured value indicating intensity of an emotion including arousal and valence of an emotion, an emotion amount obtained by time integration of the emotion measured value, and emotion transition information including a direction or velocity of a change of the emotion measured value. 6. The impression degree extraction apparatus according to claim 1, wherein the second period is a period in which a user is in a normal state, or a period in which external environment information is obtained that is identical to external environment information obtained in the first period. 7. The impression degree extraction apparatus according to claim 4, wherein the biological information is at least one of heart rate, pulse, body temperature, facial myoelectrical signal, voice, brainwave, electrical skin resistance, skin conductance, skin temperature, electrocardiographic frequency, and facial image, of a user. 8. The impression degree extraction apparatus according to claim 3, wherein: the content is video content recorded in the first period; andthe editing is processing whereby a summary video is generated by extracting a scene for which an impression degr... | PANASONIC CORPORATION | A61B005/00
US20110144531 | A1 | 2011-06-16 | Gastric Band System Injector with Accelerometer | 2009-12-16 | A pressure sensor may be used to determine the pressure of fluid within an implantable gastric restriction device. An accelerometer may be used to measure movement of the patient in which the gastric restriction device is implanted, such that the movement may be accounted for when evaluating data from the pressure sensing system. For instance, movement data from the accelerometer may be used to identify pressure data that should be ignored. A display may annotate a pressure graph to show changes in pressure associated with patient movement; or may otherwise influence the display of pressure data based on patient movement. The pressure sensor and accelerometer may be incorporated into an adapter placed between a syringe and needle that are used to adjust the fluid pressure in the gastric restriction device. Upon examination of pressure data in relation to movement data, the implantable gastric restriction may be adjusted as needed. | 1. An apparatus comprising: (a) a body defining a hollow interior portion;(b) a first port in selective fluid communication with the hollow interior portion of the body;(c) a second port in fluid communication with the hollow interior portion of the body;(d) a first sensing device in fluid communication with the hollow interior portion, wherein the first sensing device is configured to sense pressure of fluid within the hollow interior portion; and(e) a second sensing device coupled with the body, wherein the second sensing device is configured to sense at least one of movement of the body, orientation of the body, or a physical attribute of a patient. 2. The apparatus of claim 1, wherein the first sensing device comprises a pressure sensor located within the hollow interior portion. 3. The apparatus of claim 1, wherein the second sensing device comprises an accelerometer. 4. The apparatus of claim 1, wherein the second sensing device is located in the body. 5. The apparatus of claim 1, further comprising a connector, wherein the connector comprises an interface portion configured to removably connect with the body, wherein the connector is further configured to couple with an external display device. 6. The apparatus of claim 5, wherein the connector comprises a cable coupled with the interface portion, wherein the cable is configured to communicate data from the first sensing device and the second sensing device. 7. The apparatus of claim 6, further comprising an external display device coupled with the cable, wherein the external display device includes a logic configured to simultaneously display data from the first sensing device and data from the second sensing device as a function of time. 8. The apparatus of claim 5, wherein the second sensing device is located within the interface portion of the connector. 9. The apparatus of claim 5, wherein the first sensing device is located within the interface portion of the connector. 10. The apparatus of claim 1, wherein the first port is configured to couple with a syringe, wherein the second port is configured to couple with a needle. 11. A system for measuring pressure within an implantable restriction device, the system comprising: (a) a pressure sensor, wherein the pressure sensor is configured to measure pressure data associated with the implantable restriction device;(b) an accelerometer or gravitometer configured to measure at least one of movement of the implantable restriction device, orientation of the implantable restriction device, or a physical attribute of a patient; and(c) a processing device configured to receive and process data associated with the pressure sensor and the accelerometer or gravitometer, wherein the processing device is further configured indicate movement or orientation of the implantable restriction device corresponding with the pressure data. 12. The system of claim 11, wherein the processing device comprises a display device having a display configur... | - | A61B005/00 A61F002/00
US20110137184 | A1 | 2011-06-09 | PRESSURE SENSING | 2009-08-18 | A high electron mobility transistor (HEMT) is disclosed capable of performing as a pressure sensor. In one embodiment, the subject pressure sensor can be used for the detection of body fluid pressure. A piezoelectric, biocompatible film can be used to provide a pressure sensing functionalized gate surface for the HEMT. Embodiments of the disclosed sensor can be integrated with a wireless transmitter for constant pressure monitoring. | 1. A pressure sensor, comprising: a high electron mobility transistor (HEMT) comprising a piezoelectric biocompatible film on a gate region. 2. The pressure sensor according to claim 1, wherein the gate region has an area in the range of about a few hundred microns square to less than one micron square. 3. The pressure sensor according to claim 1, wherein the gate region has an area of about one hundred microns square. 4. The pressure sensor according to claim 1, wherein the gate region has an area of less than one micron square. 5. The pressure sensor according to claim 1, wherein the piezoelectric thin film comprises polyvinylidene fluoride (PVDF). 6. The pressure sensor according to claim 1, wherein the piezoelectric thin film comprises a metal oxide having piezoelectric properties. 7. The pressure sensor according to claim 1, further comprising: a wireless transmitter for constant pressure monitoring electrically connected to a source or drain of the HEMT. 8. The pressure sensor according to claim 1, wherein the HEMT comprises an AlGaN/GaN HEMT. 9. The pressure sensor according to claim 1, wherein the HEMT comprises an AlGaAs/GaAs HEMT. 10. The pressure sensor according to claim 1, wherein the HEMT comprises an InGaP/GaAs HEMT. 11. The pressure sensor according to claim 1, wherein the HEMT comprises an InAlAs/InGaAs HEMT. 12. A method of detecting body fluid pressure, comprising: providing a high electron mobility transistor (HEMT) pressure sensor comprising a piezoelectric biocompatible film in or near a fluid vessel, wherein the piezoelectric biocompatible film changes mechanical displacement caused by fluid in the fluid vessel into electrical signals and the HEMT amplifies the electrical signals. 13. The method according to claim 12, wherein the piezoelectric biocompatible film comprises polyvinylidene fluoride (PVDF). 14. The method according to claim 12, wherein the piezoelectric biocompatible film comprises a metal oxide having piezoelectric properties. 15. The method according to claim 12, wherein the HEMT comprises an AlGaN/GaN HEMT. 16. The method according to claim 12, wherein the HEMT comprises an AlGaAs/GaAs HEMT. 17. The method according to claim 12, wherein the HEMT comprises an InGaP/GaAs HEMT. 18. The method according to claim 12, wherein the HEMT comprises an InAlAs/InGaAs HEMT. 19. The method according to claim 12, further comprising wirelessly transmitting the amplified electrical signals from the HEMT to a display. 20. The method according to claim 12, wherein providing the HEMT in or near a fluid vessel comprises mounting the HEMT on a head of intravenous therapy. | - | A61B005/021 H01L029/778 H01L029/205
US20110105853 | A1 | 2011-05-05 | SYSTEMS AND METHODS FOR HEALTHCARE DELIVERY, OBSERVATION, AND COMMUNICATION BETWEEN A DE-CENTRALIZED HEALTHCARE SYSTEM AND A PATIENT LIVING AT HOME | 2010-09-10 | The disclosure describes de-centralized healthcare systems and methods of using the same to provide healthcare to a subject living at home. A de-centralized healthcare system of the invention provides at or above hospital-level care through the implementation of care pathway protocols for effective treatment of designated health conditions. A subject is monitored by a central station, which is staffed by healthcare practitioners and providers. A dispatch system allows healthcare providers to visit the subject for appointments and deliveries. The de-centralized healthcare system provides for the subject's laboratory, nutrition, and hygiene needs. | 1. A de-centralized healthcare system comprising: a) a central station staffed by at least one healthcare practitioner;b) a plurality of care pathway protocols, wherein each care pathway protocol comprises a protocol for providing healthcare for a selected designated health condition to a subject in need or want thereof; andc) a communication system for communication between the subject and the healthcare practitioner,wherein the subject resides at a location remote from the central station and is registered with the de-centralized healthcare system. 2. The de-centralized healthcare system of claim 1, wherein the central station is effective to monitor the health, status, and location of the subject. 3. The de-centralized healthcare system of claim 1, wherein the central station further comprises a record-keeping system, a computing system, and a library of reference materials, wherein the record-keeping system comprises a record for each subject registered with the de-centralized healthcare system. 4. The de-centralized healthcare system of claim 1, wherein the designated health condition comprises one or more of: heart transplant or implant of heart assist system; craniotomy and endovascular intracranial procedures; extracranial procedures; multiple sclerosis and cerebellar ataxia; intracranial hemorrhage or cerebral infarction; transient ischemia; nonspecific cerebrovascular disorders; seizures; headaches; disequilibrium (vertigo); major chest procedures; other respiratory system operating room procedures; pulmonary embolism; respiratory infections and inflammations; respiratory neoplasms; pulmonary edema and respiratory failure; chronic obstructive pulmonary disease; simple pneumonia and pleurisy; bronchitis and asthma; respiratory signs and symptoms; respiratory system diagnosis with ventilator support; deep venous thrombosis, chronic heart failure; cardiac valve and other major cardiothoracic procedures; cardiac defibrillator implant with cardiac catheterization with acute myocardial infarction/heart failure/shock; cardiac defibrillator implant without cardiac catheterization with major coronary bypass with percutaneous transluminal coronary angioplasty; major cardiovascular procedures or thoracic aorta; permanent cardiac pacemaker implant; percutaneous cardiovascular procedure with drug-eluting stent or other circulatory system operating room procedures; other vascular procedures; cardiac pacemaker revision except device replacement; other circulatory system operating room procedures; acute myocardial infarction; circulatory disorders except acute myocardial infarction, with cardiac catheterization; heart failure and shock; deep vein thrombophlebitis; dehydration; diversion diagnosis; end-stage disease states; external feeding; New York Heart Association class IV heart failure; severe/symptomatic valvular heart disease in individuals who are poor surgical candidates; Global Initiative for Chronic Obstructive Lung Disease classification ... | - | A61B005/00
US20110153197 | A1 | 2011-06-23 | INSOLE TYPE NAVIGATION APPARATUS AND OPERATION METHOD THEREOF | 2010-08-20 | An insole type navigation apparatus for indicating the direction of a user by interworking with a navigation device providing a road guidance service, includes: a GPS module mounted on the user's shoes and providing the user's location information to the navigation device; a plurality of vibration motors mounted in the insole of the shoes and indicating the recommended direction of the user's travel through vibrations; and a controller selectively driving vibration motors corresponding to each direction to inform the user about the direction of a destination, when the navigation device provides the road guidance service. | 1. An insole type navigation apparatus for indicating the direction of a user by interworking with a navigation device providing a road guidance service, the apparatus comprising: a GPS module mounted on the user's shoes and providing the user's location information to the navigation device;a plurality of vibration motors mounted in the insole of the shoes and indicating the recommended direction of the user's travel through vibrations; anda controller selectively driving vibration motors corresponding to each direction to inform the user about the direction of a destination, when the navigation device provides the road guidance service. 2. The apparatus of claim 1, further comprising: a plurality of pressure sensors mounted in the insole of the shoes to sense the user's foot pressure data,wherein the controller detects the user's current condition by using the user's foot pressure data and controls the operation of the plurality of vibration motors according to the user's current condition. 3. The apparatus of claim 2, wherein the controller collects user's foot pressure data by using the plurality of pressure sensors, divides the bottom of the shoes into four areas according to the locations of the plurality of pressure sensors to obtain an average pressure value of each area, and recognizes the user's current condition as whether or not the user is in motion according to a change of sizes of the average pressure values of each area. 4. The apparatus of claim 3, wherein when the user is in motion, the controller controls the plurality of vibration motors to operate. 5. The apparatus of claim 3, wherein when the user is not moving, the controller discriminates the user's current condition into standing, sitting, and raising his foot. 6. The apparatus of claim 5, wherein when the user's current condition standing or raising his foot, the controller controls the plurality of vibration motors so as to operate. 7. The apparatus of claim 5, wherein when the user's current condition is sitting, the controller controls the plurality of vibration motors so as not to operate. 8. The apparatus of claim 1, wherein the plurality of vibration motors are mounted in four directions including up, down, left, and right of the insole of the shoes, and indicates eight directions including up, down, left, right, left up, left down, right up, and right down. 9. An operation method of an insole type navigation apparatus for guiding a user along a road by interworking with a navigation device providing a road guidance service, the method comprising: providing the road guidance service based on the user's location information; andwhen the navigation device provides the road guidance service, selectively driving a plurality of vibration motors mounted in an insole of the user's shoes to inform the user about the recommended direction of travel to a destination. 10. The method of claim 9, wherein, informing the user about the recommended direction is indicating ... | Electronics and Telecommunications Research Institute | G01C021/34 G01S019/42 A61B005/103
US20110124996 | A1 | 2011-05-26 | DIABETES HEALTH MANAGEMENT SYSTEMS AND METHODS | 2009-11-20 | Diabetes health management systems for use in portable devices having and methods thereof having a user interface, a processor, a memory, and a communication circuit are disclosed. In one embodiment, a diabetes health management system has program code further including a communications module, a data module, a therapy module, and an analysis module. The communications module wirelessly couples the portable device to a plurality of user devices. The data module receives and stores into the memory blood glucose measurement values, insulin dosage data, and health data entries. The therapy module determines a therapy advice message based at least in part on the received blood glucose measurement values, the received insulin dosage data, and displays the therapy advice message on the user interface. The analysis module displays on the user interface a graphical representation of selected blood glucose measurement values, selected insulin dosage data, selected health data entries, or combinations thereof. | 1. A diabetes health management system for use in a portable device having a user interface, a processor, a memory, and a communication circuit, the diabetes health management system comprising program code further comprising a communications module, a data module, a therapy module, and an analysis module, wherein: the communications module causes the processor to control the communication circuit to wirelessly couple the portable device to a plurality of user devices, and display information provided by the plurality of user devices on the user interface;the data module causes the processor to receive and store blood glucose measurement values, insulin dosage data, and health data entries into the memory;the therapy module causes the processor to determine a therapy advice message based at least in part on the blood glucose measurement values, the insulin dosage data, and the health data entries stored in the memory, and display the therapy advice message on the user interface; andthe analysis module causes the processor to generate and display to a user on the user interface a graphical representation of selected blood glucose measurement values, selected insulin dosage data, selected health data entries, or combinations thereof. 2. The diabetes health management system of claim 1 wherein the communications module further causes the processor to transmit to a health network server the blood glucose measurement values, the insulin dosage data, the health data entries, or combinations thereof via the communication circuit. 3. The diabetes health management system of claim 1 wherein the therapy module causes the processor to generate an emergency alarm based at least in part on the blood glucose measurement values. 4. The diabetes health management system of claim 3 wherein the communications module further causes the processor to wirelessly transmit the emergency alarm to a caregiver via the communication circuit. 5. The diabetes health management system of claim 1 wherein the therapy advice message comprises a basal rate recommendation and a bolus delivery recommendation. 6. The diabetes health management system of claim 1 wherein: the plurality of user devices comprises an insulin pump; andthe communications module further causes the processor to control the communication circuit to transmit wireless command signals to the insulin pump such that the portable device is operable to remotely control the insulin pump. 7. The diabetes health management system of claim 6 wherein the wireless command signals correspond to a basal rate profile, a basal interval, a basal rate lag time, bolus delivery velocity, bolus initiation, or combinations thereof. 8. The diabetes health management system of claim 6 wherein the therapy module comprises a plurality of user definable and selectable basal rate profiles, and the communications module causes the processor to control the communication circuit to transmit the wireless command signals to the insulin ... | ROCHE DIAGNOSTICS OPERATIONS, INC. | A61B005/145 A61B005/00 A61M005/142
US07962187 | B2 | 2011-06-14 | Optical imaging and oximetry of tissue | 2003-03-13 | Systems and methods are disclosed for detecting at least one region of a sample having an absorption level different from a background level of absorption in the sample by obtaining thicknesses of the sample and intensities of light transmitted through the sample at a plurality of locations. The system includes glass plates (10) for compressing the tissue, distance sensors (20, 30), illuminations fibers (40) connected to a light source (70), and collection fibers (50) connected to spectrograph (110). Spatial second derivatives are calculated from products of the thicknesses of the sample and the intensities of the transmitted light for the locations. The data points are compared to detect the region of the sample having an absorption level different from the background level of absorption within the sample. The new systems and method can be used to optically image, detect, and characterize tissue, lesions, such as cancer. | 1. A method of determining whether a tumor in a tissue sample is malignant or benign, the method comprising: (a) selecting two wavelengths of light to minimize a difference between relative changes in intensity of light transmitted through the tumor for the two wavelengths, wherein the relative changes in intensity of the light for each wavelength are measured relative to a background intensity of light transmitted through the sample;(b) obtaining measures of background reduced scattering coefficients of the sample at the two wavelengths;(c) calculating an absolute oxygenation level of the tumor in the sample by using the relative changes in intensity of the light for the two wavelengths and the measures of the background reduced scattering coefficients of the sample for the two wavelengths; and(d) determining whether the tumor is malignant or benign based on the value of the calculated absolute oxygenation level. 2. The method of claim 1, wherein the tissue sample is selected from the group consisting of breast tissue, brain tissue, and muscle tissue. 3. The method of claim 1, wherein the two wavelengths of light are in the near infrared spectrum. 4. The method of claim 1, wherein the difference between the relative changes in intensity is zero. 5. The method of claim 1, wherein the absolute oxygenation level of the tumor is calculated using a formulaSO2=ɛHb⁡(λ2)-ɛHb⁡(λ1)⁢μs⁢⁢0′⁡(λ1)⁢Δ⁢⁢I/I0⁢❘max(λ2)μs⁢⁢0′⁡(λ2)⁢Δ⁢⁢I/I⁢⁢❘max(λ1)[ɛHb⁡(λ2)-ɛHbO2⁡(λ2)]+ [ɛHbO2⁡(λ1)-ɛHb⁡(λ1)]⁢μs⁢⁢0′⁡(λ1)⁢Δ⁢⁢I/I0⁢❘max(λ2)μs⁢⁢0′⁡(λ2)⁢Δ⁢⁢I/I⁢⁢❘max(λ1)⁢,wherein: SO2 is the oxygenation level of the region;ΔI/I0¦max(λ1) is a maximum value of a relative change in intensity at a first wavelength in the pair of wavelengths;ΔI/I0¦max(λ2) is a maximum value of a relative change in intensity at a second wavelength in the pair of wavelengths;μs⁢⁢0′⁡(λ1)μs⁢⁢0′⁡(λ2)is a ratio of the background reduced scattering coefficient at the first wavelength to the background reduced scattering coefficient at the second wavelength; εHb(λ1) and εHb(λ2) are molar extinction coefficients for deoxy-hemoglobin at the first and second wavelengths; andεHbO2(λ1) and εHbO2(λ2) are molar extinction coefficients for oxy-hemoglobin at the first and second wavelengths. 6. The method of claim 1, further comprising, prior to selecting the two wavelengths of light: (e) illuminating the sample with a plurality of wavelengths of light; and(f) detecting light transmitted through the sample at a plurality of locations,wherein the two wavelengths of light are selected from among wavelengths of the light transmitted through the sample. 7. The method of claim 6, further comprising (g) displaying an image of the absolute oxygenation level at the plurality of locations within the tissue sample. 8. The method of claim 1, further comprising (e) storing values of the absolute oxygenation level of the tumor in a computer-readable medium. 9. A method of determining whether a tumor in a tissue sample is malignant o... | Tufts University | A61B005/00
US20110152645 | A1 | 2011-06-23 | PULSE OXIMETRY SENSOR ADAPTER | 2010-11-29 | An adapter allows the interconnection of a sensor originating from one manufacturer to be coupled with conventionally incompatible monitors originating from other manufacturers to form a properly functioning pulse oximetry system. The adapter matches a sensor driver in a monitor to the current requirements and light source configuration of a sensor. The adapter also matches a sensor's light detector signal level to the dynamic range requirements of a monitor preamplifier. Further, the adapter provides compatible sensor calibration, sensor type and security information to a monitor. The adapter may have a self-contained power source or it may derive power from the monitor, allowing both passive and active adapter components. The adapter is particular suited as an adapter cable, replacing a conventional patient cable or sensor cable as the interconnection between a sensor to a monitor in a pulse oximetry system. | 1. An adapter capable of modifying information exchanged between an optical sensor and a patient monitor, said optical sensor configured to output a signal indicative of light attenuated by body tissue, said patient monitor configured to determine measurement values for one or more physiological parameters of said body tissue responsive to said output signal of said optical sensor, said adapter comprising: an input from said sensor, said input including information stored on an information element associated with said sensor;a translator configured to receive said information stored on said information element and output modified information different from said information stored on said information element; andan output to said monitor, said output including said modified information, wherein said modified information being configured to replace information expected by said monitor thereby enabling said monitor to provide said measurement values. 2. The adapter of claim 1 wherein the information stored on said information element comprises an indication of a type of sensor. 3. The adapter of claim 1 wherein the information stored on said information element comprises an indication of sensor origin. 4. The adapter of claim 1 wherein the information stored on said information element comprises calibration information for light sources of said sensor. 5. The adapter of claim 1 wherein the information stored on said information element comprises information indicative of sensor use. 6. The adapter of claim 1 wherein the information stored on said information element comprises sensor identification. 7. The adapter of claim 6 wherein the information stored on said information element comprises a unique sensor identification. 8. A method of communicating modified sensor information to a monitor, said sensor configured to output a signal indicative of light attenuated by body tissue, said monitor configured to process said output signal to determine measurement values for one or more physiological parameters of said body tissue, said method comprising: attaching a monitor to a first end of an adapter cable;attaching a sensor to a second end of the adapter cable;receiving through said adapter cable originally stored information stored on said sensor;modifying said information; andproviding said modified information to a reader expecting to read said originally stored information on said sensor, said modified information changing the behavior of said monitor from behavior corresponding to having received said originally stored information. 9. The method of claim 8 wherein the originally stored information stored on said sensor comprises an indication of a type of sensor. 10. The method of claim 8 wherein the originally stored information stored on said sensor comprises an indication of sensor origin. 11. The method of claim 8 wherein the originally stored information stored on said sensor comprises calibration information for light sources of sa... | - | A61B005/1455
US20110131854 | A1 | 2011-06-09 | SHEET OF WRISTBANDS | 2009-12-09 | A sheet of wristbands comprising a plurality of wristband straps, each including opposed side edges. A side edge of each strap is coupled to a side edge of an adjacent strap to form a series of side-by-side straps. A first strap defines a first side edge of the series and a second strap defines a second side edge of the series. A closure is positioned on each strap. A first tear strip is coupled to the first side edge, and a second tear strip is coupled to the second side edge. The first tear strip has a first length, and the second tear strip has a second length substantially equal to the first length. It is preferred that the first and second lengths are substantially equal to a length of the wristband straps. At least one wristband strap has an end edge that is exposed. Each tear strip can also have an end edge that is exposed. Preferably, the tear strips each have a width that is narrower than the width of the wristband straps. Each strap can also have tamper cuts, but it is preferred that the strips do not have tamper cuts. | 1. A sheet of wristbands comprising: a plurality of wristband straps, each strap including opposed side edges, wherein a side edge of each strap is detachably coupled to a side edge of an adjacent strap to form a series of side-by-side straps, wherein a first strap defines a first side edge of the series of straps and a second strap defines a second side edge of the series of straps;a closure positioned on each strap; anda first tear strip positioned adjacent to the first strap and detachably coupled to the first side edge, the first tear strip having a first length; anda second tear strip positioned adjacent to the second strap and detachably coupled to the second side edge, the second tear strip having a second length substantially equal to the first length, and wherein the first and second lengths are substantially the same as a length of at least one of the wristband straps. 2. A sheet of wristbands as defined in claim 1, wherein all of the wristband straps have substantially the same length. 3. A sheet of wristbands as defined in claim 1, wherein at least one wristband strap has an end edge that is exposed. 4. A sheet of wristbands as defined in claim 3, wherein each wristband strap has an edge that is exposed. 5. A sheet of wristbands as defined in claim 4, wherein the tear strips each have end edges that are exposed. 6. A sheet of wristbands as defined in claim 1, wherein the tear strips each have a width that is narrower than a width of at least one of the wristband straps. 7. A sheet of wristbands as defined in claim 6, wherein the width of the tear strips is less than 75% of the width of wristband straps 8. A sheet of wristbands as defined in claim 7, wherein the width of the tear strips is at least 50% of the width of wristband straps. 9. A sheet of wristbands as defined in claim 1, wherein a detachable coupling between straps comprises a perforation. 10. A sheet of wristbands as defined in claim 1, wherein a detachable coupling between straps and strips comprises a perforation. 11. A sheet of wristbands as defined in claim 1, wherein each strap includes tamper cuts, but strips do not include tamper cuts. 12. A sheet of wristbands comprising: a plurality of wristband straps, each strap including opposed side edges, wherein a side edge of each strap is detachably coupled to a side edge of an adjacent strap to form a series of side-by-side straps, wherein a first strap defines a first side edge of the series of straps and a second strap defines a second side edge of the series of straps, and wherein at least one wristband strap has an end edge that is exposed;a closure positioned on each strap; anda first tear strip positioned adjacent to the first strap and detachably coupled to the first side edge; anda second tear strip positioned adjacent to the second strap and detachably coupled to the second side edge. 13. A sheet of wristbands as defined in claim 12, wherein each wristband strap has an edge that is exposed. 14. A she... | WRISTBAND RESOURCES, INC. | A61B005/117 A44C005/00
US20110105940 | A1 | 2011-05-05 | METHOD FOR QUANTIFICATION OF TITER OF NEUTRALIZING ANTIBODY TO NEUROTOXIN | 2008-08-20 | A method for quantification of a titer of a neutralizing antibody to a neurotoxin is disclosed. The method comprises the steps: (a) mixing a standard sample containing a fixed amount of a neurotoxin and a test sample containing a neutralizing antibody to said neurotoxin; (b) administering the mixture obtained in step (a) into the muscle of a non-human mammal; (c) applying electric stimulus to said non-human mammal; (d) measuring a compound muscle action potential (CMAP) due to contraction of the muscle of said mammal by application of electric stimulus with an electromyograph; and (e) analyzing CMAP amplitude data obtained in step (d) for an extent of decrease in amplitude by a non-neutralized neurotoxin to thereby quantify a titer of the neutralizing antibody contained in the test sample. | 1. A method for quantification of a titer of a neutralizing antibody to a neurotoxin comprising the following steps: (a) mixing a standard sample containing a fixed amount of a neurotoxin and a test sample containing a neutralizing antibody to said neurotoxin;(b) administering the mixture obtained in step (a) into the muscle of a non-human mammal;(c) applying electric stimulus to said non-human mammal;(d) measuring a compound muscle action potential (CMAP) due to contraction of the muscle of said mammal by application of electric stimulus with an electromyograph; and(e) analyzing CMAP amplitude data obtained in step (d) for an extent of decrease in amplitude by the neurotoxin not neutralized to thereby quantify a titer of the neutralizing antibody contained in the test sample. 2. The method for quantification of claim 1 wherein the neurotoxin is selected from a neurotoxin produced by bacterium of Clostridium, a neurotoxin produced by fish and shellfish or a snake toxin. 3. The method for quantification of claim 2 wherein the bacterium of Clostridium is Clostridium botulinum. 4. The method for quantification of claim 3 wherein said Clostridium botulinum is selected from type A, B, C, D, E, F or G of Clostridium botulinum. | JURIDICAL FDN THE CHEMO-SERO-THERAPEUTIC RES. INS. JP AS REP BY DIR-GEN NATL INST OF INFECTIOUS DIS | A61B005/0488
US07938782 | B2 | 2011-05-10 | Prediction of disordered breathing | 2008-05-12 | An approach for predicting disordered breathing involves detecting one or more conditions associated with disordered breathing. The detected conditions are compared to disordered breathing prediction criteria. A prediction of disordered breathing is performed based on the comparison of the detected conditions to the prediction criteria. At least one of comparing the detected conditions to the prediction criteria and predicting disordered breathing is performed at least in part implantably. | 1. A method of operating a medical device to predict disordered breathing in a patient, comprising: measuring a transthoracic impedance of the patient;detecting a pulmonary congestion condition of the patient based on the measured transthoracic impedance;comparing the detected pulmonary congestion condition to a prediction criterion; andpredicting disordered breathing based on the comparison, wherein each of measuring, detecting, comparing and predicting is performed by the medical device. 2. The method of claim 1, wherein the detecting includes detecting a plurality of conditions associated with disordered breathing, of which the pulmonary congestion condition is one. 3. The method of claim 1, wherein the detecting includes detecting a plurality of conditions predisposing the patient to disordered breathing, of which the pulmonary congestion condition is one. 4. The method of claim 1, wherein the detected pulmonary congestion condition is one of a plurality of detected conditions, wherein the prediction criterion is included in one or more sets of prediction criteria, and wherein the comparing includes comparing the plurality of detected conditions to the one or more sets of prediction criteria. 5. The method of claim 1, wherein the predicting is performed at least in part implantably. 6. The method of claim 1, wherein each of measuring, detecting, comparing, and predicting is performed at least in part implantably. 7. The method of claim 1, wherein the detected pulmonary congestion condition is one of a plurality of detected conditions, and wherein the comparing includes (a) computing an estimated probability that disordered breathing will occur based on the detected conditions, and (b) comparing the estimated probability to a threshold probability associated with an onset of disordered breathing. 8. The method of claim 1, wherein the prediction criterion is included in one or more sets of prediction criteria, the method further comprising: adjusting a particular set of prediction criteria based on the detected pulmonary congestion condition. 9. The method of claim 1, wherein the predicting includes predicting the disordered breathing will occur within a particular time interval. 10. A medical device operable to predict disordered breathing in a patient, comprising: a sensor configured to sense a transthoracic impedance of the patient;a detector system coupled to the sensor and configured to detect conditions associated with disordered breathing, the detected conditions including a pulmonary congestion condition of the patient; anda prediction engine coupled to the detector system and configured to predict disordered breathing based on the detected conditions. 11. The device of claim 10, wherein the prediction engine includes an implantable component. 12. The device of claim 10, wherein the sensor, the detector system, and the prediction engine each include an implantable component. 13. The device of claim 10, wherein the prediction... | Cardiac Pacemakers, Inc. | A61B005/08
US07963921 | B1 | 2011-06-21 | Method and measuring device for determining blood pressure | 1998-12-23 | The invention is directed to a method and a device for determining the blood pressure, in which a pressure sensor is applied to an individual's limb to detect the blood pressure prevailing in said limb. According to the present invention, the orientation of the limb is detected by means of an orientation sensing unit, and the detected blood pressure is corrected in an evaluating unit in response to the limb's detected orientation. | 1. A method of measuring blood pressure using a wrist sphygmomanometer including a body and including a cuff attached to the body and wrapped around the wrist, said body being attached to the cuff in a manner to allow the body to be located on a thumb side of an arm when the wrist sphygmomanometer is fitted on the wrist, comprising: fitting the wrist sphygmomanometer on a subject's wrist;placing the wrist with the wrist sphygmomanometer fitted thereon on the subject's chest; andreading a value detected by the wrist sphygmomanometer with the wrist placed on the chest, wherein a positioning system provides a first signal that indicates that the sphygmomanometer is at an appropriate measuring level and a second signal that guides the subject to the appropriate measuring level. 2. The method of measuring blood pressure of claim 1, wherein the first signal is an acoustic signal. 3. The method of measuring blood pressure of claim 1, wherein the first signal is a visual signal. 4. The method of measuring blood pressure of claim 3, wherein the first signal includes a red light indicating an inappropriate measuring level and a green light indicating an appropriate measuring level. 5. The method of measuring blood pressure of claim 1, wherein the first signal is provided at the start of, or during blood pressure measurement. 6. The method of measuring blood pressure of claim 1, wherein the second signal includes displaying arrows that guide a user to the appropriate measuring level. 7. The method of measuring blood pressure of claim 1, wherein the second signal includes a flashing light that guides a user to the appropriate measuring level. 8. A method of measuring blood pressure of a user using a wrist sphygmomanometer including a body, and a cuff attached to the body, comprising: fitting the wrist sphygmomanometer on the user's wrist including locating the body on a thumb side of the user's arm;placing the wrist on the user's chest; andreading a value detected by said wrist sphygmomanometer with the wrist placed on the chest; andprompting the user to adopt an appropriate measuring position including displaying a visual signal on the display that guides the user to an appropriate measuring position. 9. The method of claim 8, wherein the visual signal includes arrows that guide the user to an appropriate measuring level. | Kaz USA, Inc. | A61B005/00
US08269636 | B2 | 2012-09-18 | Method and apparatus for monitoring ingestion of medications using an implantable medical device | 2011-01-24 | An implantable medical device, such as a pacemaker or implantable cardioverter defibrillator (ICD), is configured to automatically detect ingestion of medications to verify that prescribed medications are taken in a timely manner and at the correct dosage. Briefly, individual pills are provided with miniature radio frequency identification (RFID) devices capable of transmitting RFID tag signals, which identify the medication contained within the pill and its dosage. The implanted device is equipped with an RFID transceiver for receiving tag signals from a pill as it is being ingested. The implanted system decodes the tag to identify the medication and its dosage, then accesses an onboard database to verify that the medication being ingested was in fact prescribed to the patient and to verify that the correct dosage was taken. Warning signals are generated if the wrong medication or the wrong dosage was taken. Therapy may also be automatically adjusted. Non-RF-based ID devices are also described, which instead transmit ID data via biphasic current pulses. | 1. In a system comprising an implantable medical device and an external device in communication with the implantable medical device, a method comprising: sensing, with the implantable medical device, a signal transmitted by a pill being ingested by a patient;detecting ingestion of the pill based on the sensed signal;transmitting data from the implantable medical device to the external device; andprocessing the data at the external device to determine medication compliance by the patient by comparing the pill against a database listing medications prescribed to the patient, and generating a warning signal if the pill is not listed in the database. 2. The method of claim 1 wherein the pill is equipped with radio frequency identification (RFID) device and wherein sensing the signal transmitted by the pill comprises sensing an RFID signal transmitted by the RFID device of the pill. 3. The method of claim 2 wherein the RFID device is an active ID device and wherein sensing the signal transmitted by the pill comprises sensing active ID signals transmitted by the ID device. 4. The method of claim 2 wherein the pill is equipped with a passive RFID device and wherein sensing signals comprises: transmitting a power signal to the passive RFID device of the pill; andreceiving a responsive signal returned by the passive RFID device of the pill. 5. The method of claim 1 wherein generating a warning signal is performed by transmitting a warning signal to the external device. 6. The method of claim 1 further comprising automatically recording, within the implantable medical device, data indicative of the time the pill was ingested. 7. In a system comprising an implantable medical device and an external device in communication with the implantable medical device, a method comprising: sensing, with the implantable medical device, a signal transmitted by a pill being ingested by a patient;detecting ingestion of the pill based on the sensed signal;transmitting data from the implantable medical device to the external device;processing the data at the external device to determine medication compliance by the patient;determining a dosage of the pill ingested and determining whether the dosage is appropriate for the patient; andgenerating a warning signal if the dosage being ingested is not the appropriate dosage for the patient. 8. The method of claim 7 and further comprising determining the dosage taken by the patient over a period of time, comparing the dosage over the period of time to a schedule, and determining whether the patient has taken a correct dosage over the period of time. 9. A system for use in monitoring the ingestion of pills by a patient, the system comprising: an implantable medical device adapted to sense a signal from a pill being ingested by a patient, and to transmit data corresponding to the pill; andan external device in communication with the implantable medical device and adapted to receive the data from the implantable medical device... | Pacesetter, Inc. | G08B023/00
US08768430 | B2 | 2014-07-01 | MRT interference field detection medical implant with prone posture position sensor | 2010-12-18 | Allows the specificity of an automatic MRT detection to be increased in a simple manner. This is achieved using an automatically calibrating position sensor, so that the user does not have to perform additional calibration of this sensor. Incorrect sensor calibrations are thus eliminated as well. | 1. An implantable medical device comprising a cardiac pacemaker that also comprises: a magnetic resonance imaging interference fields detection unit, comprising at least one sensor, or at least one indicator, or both configured to detect electromagnetic interference fields typcial of magnetic resonance imaging or magnetic fields typical of magnetic resonance imaging;at least one control unit coupled to the magnetic resonance imaging interference fields detection unit comprising at least one diagnostic unit orat least one treatment unit,or both said at least one diagnostic unit and said at least one treatment unit;a position sensor coupled to the at least one control unit;wherein the magnetic resonance imaging interference fields detection unit is configured to detect interference from magnetic resonance imaging interference fields only when, in addition to the at least one sensor, or the at least one indicator, or both the at least one sensor and the at least one indicator that are configured to detect electromagnetic interference fields typcial of magnetic resonance imaging, the position sensor indicates a prone posture of a patient associated with said implantable medical device;a calibration unit coupled to said position sensor and configured to automatically detect the prone posture from said position sensor;wherein the calibration unit is connected to said position sensor and to one or more of a heart analysis detection unit,a respiratory rate detection unit, andan implant clock;wherein the calibration unit is further configured to detect one or more parameter values from one or more of said heart analysis detection unit, said respiratory rate detection unit and said implant clock; and,wherein the calibration unit is further configured to store and update data from said position sensor when said one or more detected parameter values suggest night sleep. 2. The implantable medical device according to claim 1, wherein the position sensor is self-calibrating. 3. The implantable medical device according to claim 2, wherein the calibration of the position sensor is carried out after evaluation of a time of day from said implant clock. 4. The implantable medical device according to claim 1, wherein the identification of the prone posture include evaluation of diurnal phases with minimum heart rate that are evaluated from said heart analysis detection unit. 5. The implantable medical device according to claim 1, wherein the at least one control unit is configured to identify night sleep by an evaluation of a heart rate variability and a circadian rhythm thereof from said heart analysis detection unit. 6. The implantable medical device according to claim 1, wherein the at least one control unit is configured to identify night sleep by an evaluation of a respiratory rate from said respiratory rate detection unit. 7. The implantable medical device according to claim 1, wherein the at least one control unit is configured to identify night sleep b... | Biotronik CRM Patent AG | A61B005/103 A61N001/37 A61B005/00 A61B005/11 A61B005/06 G01R033/28
US08750973 | B2 | 2014-06-10 | Method and system for detecting P-waves in the surface ECG signal | 2010-11-29 | The invention provides an ECG system and apparatus for detecting P-waves even in the patients with arrhythmia. The system is capable of obtaining sufficient data from eight leads and to display an ECG construct with marked or emphasized P-waves. | 1. A system for noninvasive measuring ECG signals produced by the heart of a human subject, and for detecting P-wave in said signals, comprising i) measuring means consisting of at least 9 electrodes, capable of detecting a potential difference at a plurality of points on the skin of said subject during a predetermined time interval (lead signal), providing at least 8 lead signals;ii) electronic circuitry for amplifying said lead signals, reducing noise, converting the lead signals to data and transferring the data to storage means;iii) storage means consisting of an electronic memory for storing all data obtained in measuring said lead signals, for storing software used in processing said signals, and for storing all data resulting from said processing;iv) processing means for calculating linear combinations of said at least eight lead signals and for segmenting said linear combination into time segments one of which is a predetermined marked time segment, followed by applying SUMER technique (Separation Using Maximum Energy Ratio) to calculate a cost function of the energy ratio between said predetermined segment and other segments within said time interval, to maximize said energy ratio and to provide an ECG construct with emphasized P-wave; andv) presenting means for graphical representation of said lead signals and said ECG constructs in a printed form or on display. 2. The system according to claim 1, wherein said measuring means consist of 9 electrodes, capable of providing 8 lead signals. 3. The system according to claim 1, wherein said measuring means comprise a software for performing the steps of i) representing said at least 8 lead signals as eight signal-time functions in said predetermined time interval;ii) calculating a linear combination of said at least eight signal-time functions;iii) segmenting said linear combination into time segments, one of which is a predetermined time segment, and calculating a cost function of the energy ratio between said predetermined segment and all other segments; andiv) repeating steps ii) and iii) in order to maximize said cost function. 4. The system according to claim 1, wherein said subject suffers from arrhythmia. 5. The system according to claim 1, wherein said processing means include a software for calculating linear combinations of eight leads, which combination converges to a signal that has the maximum energy ratio between a predetermined marked time segment and all other segments. 6. The system according to claim 1, wherein said predetermined segment is marked manually by an experienced person. 7. The system according to claim 1, wherein said predetermined segment is marked automatically by said processing means. 8. The system according to claim 1, wherein said measuring processing means separate atrial activity from ventricular activity, and emphasize the atrial activity in said construct signal and in said lead signals. 9. The system according to claim 1, comprising i) measurin... | Ben-Gurion University of the Negev Research and Development Authority | A61B005/04
US08473041 | B2 | 2013-06-25 | Cardiography system and method using automated recognition of hemodynamic parameters and waveform attributes | 2011-02-02 | A cardiography system and method using automated recognition of hemodynamic parameters and waveform attributes is provided. The cardiography system and method includes at least one sensor, a knowledge base and a processing device. The at least one sensor provides a waveform signal and a hemodynamic parameter input. The knowledge base includes data corresponding to various disease states. The processing device receives the waveform signal and hemodynamic parameter input from the sensor, identifies waveform attributes on the waveform signal, measures the waveform attributes, accesses the knowledge base, cross-references the waveform attributes and the hemodynamic parameters with data in the knowledge base, and outputs a suggested likelihood of a particular disease state. The knowledge base optionally includes goal-directed therapies associated with particular disease states for providing suggested goal-directed therapies based on the cross-referencing of the waveform attributes and the hemodynamic parameters with the knowledge base. | 1. A system for automated recognition of hemodynamic parameters and waveform attributes comprising: an impedance cardiography device including at least one sensor configured to be externally placed on a patient and adapted to noninvasively provide at least one impedance waveform signal;a knowledge base including known data corresponding to various disease states;a processing device in operable communication with said at least one sensor and said knowledge base, said processing device configured to receive said at least one impedance waveform signal and calculate a hemodynamic parameter of the patient therefrom, identify waveform attributes on said waveform signal, measure said waveform attributes, measure said hemodynamic parameter, and cross-reference said waveform attributes and said hemodynamic parameter with said knowledge base data, and output a suggested likelihood of a particular disease state based on said cross-reference. 2. The cardiography system of claim 1 further comprising a display device configured to display at least said output. 3. The cardiography system of claim 2 wherein said display device is selected from the group consisting of computer monitor, flat-screen display, projector, printing device, and audible device. 4. The cardiography system of claim 2 wherein said display device further comprises user input devices configured to communicate with said display device and said processing device. 5. The cardiography system of claim 1 wherein said knowledge base known data further comprises goal directed therapies correlated with particular disease states. 6. The cardiography system of claim 5 wherein said processing device is configured to output at least one suggested goal directed therapy based on the suggested likelihood of a particular disease state. 7. The cardiography system of claim 1 wherein said impedance waveform signal is selected from the group consisting of an ICG signal, an ECG signal and a PCG signal. 8. The cardiography system of claim 1 wherein said hemodynamic parameter is selected from the group consisting of thoracic fluid content, heart rate, pre-ejection period, left ventricular ejection time, systolic time ratio, isovolumic relaxation time, stroke volume, stroke volume index, cardiac output, cardiac index, blood pressure, Heather Index, rate pressure product, ejection fraction, end diastolic volume, pulmonary artery occlusion pressure, central venous pressure and systemic vascular resistance. 9. The cardiography system of claim 1 wherein said knowledge base known data comprises a plurality of waveforms, waveform attributes and hemodynamic parameters. 10. The cardiography system of claim 9 wherein said plurality of waveforms, waveform attributes and hemodynamic parameters are associated with systolic heart failure. 11. The cardiography system of claim 9 wherein said plurality of waveforms, waveform attributes and hemodynamic parameters are associated with diastolic heart failure. 12. The cardiog... | Vasamed, Inc. | A61B005/04
US20110152631 | A1 | 2011-06-23 | MEDICAL DIAGNOSTIC APPARATUS AND METHOD OF OPERATING THE SAME | 2010-12-17 | The present disclosure relates to a medical diagnostic apparatus and a method of operating the same. The medical diagnosis apparatus performs measurement based on a preset diagnosis system environment in response to a measurement start instruction input from a user, and only if the apparatus receives an instruction for changing the diagnosis system environment, the apparatus performs measurement based on a changed diagnosis system environment in response to the instruction. In the apparatus and method, a workflow is arranged to perform measurement accurately, efficiently and safely, thereby reducing time for measurement or diagnosis and decreasing the occurrence of user error. | 1. A method of operating a medical diagnostic apparatus, wherein the medical diagnostic apparatus performs a first measurement based on a preset diagnosis system environment in response to a measurement start instruction input from a user, and wherein only if the apparatus receives an instruction for changing the diagnosis system environment, the apparatus performs a second measurement based on a changed diagnosis system environment in response to the instruction. 2. The method of claim 1, comprising: 1) receiving the measurement start instruction;2) confirming whether a change of the preset diagnosis system environment is needed; and3) performing the first measurement,wherein the second measurement is performed by reflecting the change in diagnosis system environment only upon confirming that the change of the preset diagnosis system environment is needed. 3. The method of claim 2, wherein the diagnosis system environment comprises information about at least one of a probe, a diagnosis application, an operation mode, and at least one parameter. 4. The method of claim 3, wherein the step 2) of confirming whether the change of the preset diagnosis system environment is needed comprises: a) confirming whether a change of a diagnosis application is needed; andb) confirming whether a change of an operation mode is needed,the step b) being performed after changing the diagnosis application only upon confirming that the change of the diagnosis application is needed. 5. The method of claim 3, wherein the step 2) of confirming whether the change of the preset diagnosis system environment is needed comprises: a) confirming whether a change of an operation mode is needed; andb) confirming whether a change of a diagnosis application is needed,the step b) being performed after changing the operation mode only upon confirming that the change of the operation mode is needed. 6. The method of claim 4, wherein the step 2) of confirming whether the change of the preset diagnosis system environment is needed further comprises: c) confirming whether a change of a probe is needed, before the step a), after the step a), or after the step b),the probe being changed to proceed to the next step only upon confirming that the change of the probe is needed. 7. The method of claim 4, wherein the step 2) of confirming whether the change of the preset diagnosis system environment is needed further comprises optionally: d) confirming whether a change of a parameter is needed, before the step a), after the step a), or after the step b),the parameter being changed to proceed to the next step only upon confirming that the change of the parameter is needed. 8. The method of claim 4, wherein the step 2) of confirming whether the change of the preset diagnosis system environment is needed further comprises: c) confirming whether a change of a probe is needed; andd) confirming whether a change of a parameter is needed,wherein the steps c) and d) are performed at any selective ... | MADISON CO., LTD. | A61B005/00
US20110144452 | A1 | 2011-06-16 | APPARATUS AND METHOD FOR DETERMINING EMOTIONAL QUOTIENT ACCORDING TO EMOTION VARIATION | 2010-10-20 | The present invention relates to an apparatus and a method for determining emotional quotient according to emotional variation. The method for determining emotional quotient according to emotional variation includes sensing bio signals and environmental signals from a plurality of sensors; defining an emotional variation threshold value for each signal sensed at the sensing; extracting the emotional signal by analyzing the bio signals and inferring the emotional quotient for each bio signal based on the extracted emotional signal and the environmental signals; determining the emotional variation threshold values and the emotional quotient based on the inference results of the emotional quotient; and evaluating the determined emotional quotient. With the present invention, it can provide the emotional care service capable of informing and communicating the true emotional state of the person to user. | 1. A method for determining emotional quotient according to emotional variation, comprising: defining an emotional variation threshold value based on bio signals and environmental signals sensed from a plurality of sensors;extracting emotional signal by analyzing the bio signals and inferring the emotional quotient for each bio signal based on the extracted emotional signal and the environmental signals;determining the emotional variation threshold values and the emotional quotient based on the inference results of the emotional quotient;evaluating the determined emotional quotient; andre-determining the emotional variation threshold value and the emotional quotient of the determining according to the evaluation results of the evaluating. 2. The method for determining emotional quotient of claim 1, wherein the determining the emotional variation threshold values and the emotional quotient includes performing multimodal emotional inference by using the emotional quotient for each bio signal as the multi input, the emotional variation threshold value and the emotional quotient being determined according to the multimodal emotional inference results. 3. The method for determining emotional quotient of claim 1, wherein the evaluating includes determining a way for evaluating the emotional quotient for evaluating the emotional quotient among self diagnosis and emotional care service feedback. 4. The method for determining emotional quotient of claim 3, wherein the evaluating further includes querying whether the user agrees on the recognized emotional state when self diagnosis is used. 5. The method for determining emotional quotient of claim 4, wherein the re-determining includes: if agreed on the recognized emotional state by the response to the query, fixing the emotional variation threshold values in an emotional threshold management table, andif not, adjusting the level of the information sensed from the bio signals and the environmental signals to correct the emotional variation threshold value. 6. The method for determining emotional quotient of claim 3, wherein the evaluating further includes: analyzing service consuming patterns according to the user's emotional state to select the emotional care service meeting the emotional quotient when the emotional care service feedback is used;monitoring the emotional variation while the selected emotional care service is provided; andanalyzing the emotional variation in the current emotional quotient according to the monitored results. 7. The method for determining emotional quotient of claim 6, wherein the re-determining further includes: when the current emotion of the user is not varied, adjusting the level of the emotional quotient if it is determined that it is negative emotion by searching the emotional quotient, andif not, fixing the emotional variation threshold values in the emotional threshold management table. 8. The method for determining emotional quotient of claim 6, wherein the re... | - | A61B005/00
US20110105893 | A1 | 2011-05-05 | TISSUE TRACKING ASSEMBLY AND METHOD | 2009-11-02 | A tissue tracking assembly and method is disclosed herewith. The tissue tracking assembly comprises: a flexible probe with a proximal end and distal end, the proximal end incorporates an electric connector and the distal end incorporates at least one trackable sensor. A delivery system is provided in association with the flexible probe to accurately embed the flexible probe within a tissue. The delivery system is configured to be attached to the flexible probe during insertion of probe within the tissue and detaches from the flexible probe after embedding the probe within the tissue. | 1. A tissue tracking assembly comprising: a flexible probe with a proximal end and a distal end, the proximal end incorporates an electric connector and the distal end incorporates a sensor assembly; and a delivery system associated with the flexible probe, adapted for embedding the flexible probe within a tissue;wherein the delivery system is configured to be attached to the flexible probe during insertion of the flexible probe within the tissue and to be detached from the flexible probe after embedding the flexible probe within the tissue. 2. The assembly as claimed in claim 1, wherein the delivery system includes a needle assembly, a reinforcing tube and a clamping mechanism. 3. The assembly as claimed in claim 1, wherein the delivery system includes an attachment for performing a secondary function, the attachment including ablator, depth marker and anchoring mechanisms. 4. A tissue tracking assembly comprising: a flexible probe with a proximal end and a distal end, the proximal end incorporates an electric connector and the distal end incorporates at least one trackable sensor;a delivery system associated with the flexible probe, adapted for embedding the flexible probe within tissue, comprising:a needle assembly having an open first end and an open second end with a lumen extending there between, the first end comprises a pointed tip and the second end comprises a detachable clamping mechanism configured to be attached to the flexible probe while inserting the flexible probe within the tissue and to be detached from the flexible probe after embedding the flexible probe within the tissue; anda reinforcing tube associated with the needle assembly for supporting the flexible probe during insertion of the flexible probe within the tissue. 5. The assembly as claimed in claim 1, wherein the flexible probe includes at least one trackable sensor embedded in a flexible enclosure for tracking tissue motion. 6. The assembly as claimed in claim 1, wherein the flexible probe has surface friction sufficient to prevent relative motion of the flexible probe with respect to the tissue, and the flexible probe has strength sufficient to allow flexible probe insertion and removal. 7. The assembly as claimed in claim 1, wherein the trackable sensor includes a locking mechanism for locking the trackable sensor to the reinforcing tube. 8. The assembly as claimed in claim 1, wherein the open lumen is configured to accommodate the flexible probe while embedding the probe within the tissue. 9. The assembly as claimed in claim 1, wherein the reinforcing tube is configured to axially support the flexible probe while embedding the probe within the tissue. 10. The assembly as claimed in claim 1, wherein reinforcing tube is associated with needle wall. 11. The assembly as claimed in claim 1, wherein reinforcing tube is associated with the clamping mechanism. 12. The assembly as claimed in claim 1, wherein the clamping mechanism is configured to clamp and seal ... | GENERAL ELECTRIC COMPANY | A61B005/05
US20110112380 | A1 | 2011-05-12 | Method and System for Optimal Estimation in Medical Diagnosis | 2010-11-12 | A system and method for estimating and updating the current state of disease of a patient, called the disease impression. The disease impression of the patient, for a disease with a set of possible disease conditions, is updated based upon the current disease impression, the outcome of a test performed with respect to the disease in the patient, and the conditional probabilities of obtaining the test outcome given each of the disease conditions for the disease. The conditional probabilities with respect to all possible outcomes of the test and conditions of the disease may be stored in a likelihood matrix, which comprises a medical knowledge base. Because tests are error prone, the estimate of the patient's state of disease evolves according to a hidden Markov model, which allows the clinical impression of the patient to be computed in real-time. By applying the method to many diseases and many tests, large-scale medical diagnosis is achieved and is computationally tractable because as a result of the process, such diagnosis becomes stochastic filtering problem that is trivially parallel. In addition, the expected diagnostic utility of any test can be predicted. The same knowledge base is used to perform both inference and prediction. | 1. A method of inferring the impression of the state of a disease in a patient, wherein said disease has a set of possible disease conditions, said method comprising: (a) identifying a current disease impression of the patient;(b) obtaining the outcome of a test performed with respect to said disease in said patient; and(c) updating said current disease impression to a new disease impression based upon the conditional probabilities of obtaining said test outcome given each of said disease conditions. 2. The method of claim 1, wherein said test has a set of possible outcomes, and further comprising providing a likelihood matrix comprising the conditional probability distribution of obtaining each test outcome given each of said disease conditions, wherein the conditional probabilities in step (c) are selected from said likelihood matrix. 3. The method of claim 2, further comprising (d) providing a plurality of tests, each of said plurality of tests having a set of possible outcomes;(e) providing for each of said plurality of tests a likelihood matrix comprising the conditional probability distribution of each of said test outcomes given each of said disease conditions; and(f) repeating steps (a)-(c) with respect to at least one of said plurality of tests. 4. The method of claim 3, wherein steps (a)-(c) are repeated until at least one updated disease impression is within a predetermined threshold with respect to one of said disease conditions. 5. The method of claim 3, further comprising (g) defining for each of a plurality of diseases a set of possible disease conditions;(h) providing a likelihood matrix for each of said plurality of tests with the conditional probability distribution of each of said test outcomes given each of said disease conditions for each of said plurality of diseases; and(i) performing step (f) for each of said diseases, wherein each of the current disease impressions for each of said diseases is updated in parallel to a new disease impression based upon the conditional probabilities of obtaining said test outcome given each of said disease conditions selected from the corresponding likelihood matrix, such that the clinical impression of said patient is provided. 6. The method of claim 5, wherein steps (a)-(c) are repeated until at least one updated disease impression is within a predetermined threshold with respect to one of said disease conditions for at least one of said diseases. 7. The method of claim 1, wherein said test may randomly produce erroneous outcomes, whereby said random error causes the disease impression to update stochastically according to a hidden Markov model. 8. The method of claim 2, wherein said disease conditions are mutually exclusive and exhaustive. 9. The method of claim 3, wherein each of said tests is conditionally independent of one another. 10. The method of claim 3, wherein at least two of said tests are conditionally dependent upon one another. 11. The method of claim 5, wherein s... | eTenum, LLC | A61B005/00
US20110160555 | A1 | 2011-06-30 | Universal Models for Predicting Glucose Concentration in Humans | 2009-07-31 | An embodiment of the invention provides a system for predicting future glucose levels in an individual including a glucose measuring device for generating glucose signals representing glucose levels obtained from the individual at fixed time intervals and an analyzer. The analyzer uses a glucose prediction function that is portable between individuals irrespective of health of the individuals. The glucose prediction function includes model coefficients that are invariant between the individuals. The glucose prediction function outputs the future glucose levels by weighing the previous glucose signals obtained from the individual by the model coefficients. | 1-15. (canceled) 16. A system for predicting at least one future glucose level of an individual, said system including: a glucose measuring device, the glucose measuring device generates a series of glucose signals representing glucose levels obtained from the individual at fixed time intervals; andan analyzer having a glucose prediction function that is portable between individuals irrespective of health of individuals, said glucose prediction function including a plurality of model coefficients that are invariant between individuals, said glucose prediction function outputs the at least one future glucose level by weighing the current and a plurality of previous series of glucose signals obtained from the individual by said model coefficients, said glucose prediction function outputs a series of future glucose levels by omitting the oldest predicted or actual glucose level used in the last iteration of said glucose prediction function, multiplying a most recent predicted future glucose level by a first model coefficient, and multiplying a next most recent predicted or actual glucose level by a next model coefficient. 17-27. (canceled) 28. A method, including: receiving a time horizon as an input or retrieving the time horizon from memory;receiving series of glucose signals from a glucose measuring device, the series of glucose signals representing glucose levels obtained from an individual at fixed time intervals;predicting at least one future glucose level of the individual by weighing the series of glucose signals by a plurality of model coefficients of a glucose prediction function that is portable between individuals irrespective of health of individuals, said plurality of model coefficients are invariant between individuals, said weighing of the series of glucose signals by said plurality of model coefficients of said glucose prediction function includes omitting a least recent predicted or actual glucose level from said glucose prediction function, multiplying a most recent predicted future glucose level by a first model coefficient, and multiplying a next most recent predicted or actual glucose level by a next model coefficient, and said predicting being performed with a processor having code to perform calculations of said glucose prediction function; andrepeating said predicting for the number of required samples to reach the time horizon with each new prediction being one sampling time period later. 29. The method according to claim 28, wherein the health of the individual includes a diabetes type of the individual. 30. The method according to claim 28, wherein the health of the individual includes an age of the individual. 31. The method according to claim 30, wherein the health of the individual includes whether the individual is hospitalized. 32. The method according to claim 28, wherein said plurality of model coefficients are invariant between individuals irrespective of a type of said glucose measuring device utilized to ... | - | A61B005/145
US20110105865 | A1 | 2011-05-05 | DIFFUSE REFLECTANCE SPECTROSCOPY DEVICE FOR QUANTIFYING TISSUE ABSORPTION AND SCATTERING | 2009-04-24 | A diffuse reflectance spectroscopy system for quantifying electromagnetic absorption and scattering in a tissue is provided. Also provided are optical probes and methods for imaging a tissue mass. In some embodiments, the methods include the steps of contacting a tissue mass with an optical probe, wherein the optical probe includes at least one entity for emitting light that interacts with a tissue mass and then is remitted to a collecting entity, for collecting the light that has interacted with the tissue mass, wherein the collecting entity comprises a detector comprising one or more photodiodes; measuring turbid spectral data of the tissue mass using the optical probe; converting the turbid spectral data to at least one of absorption and scattering spectral data via a Monte Carlo algorithm or a diffusion algorithm; and quantifying tissue compositions and scatterer size in a tissue mass using the at least one of absorption and scattering spectral data. | 1. A diffuse reflectance spectroscopy system for quantifying light absorption and scattering in a tissue mass, the system comprising: an optical probe comprising at least one entity for emitting light that interacts with a tissue mass and then is remitted into a collecting entity, wherein the collecting entity comprises a detector comprising one or more photodiodes; anda processing unit for converting collected light, via a Monte Carlo algorithm or a diffusion algorithm into absorption and scattering data. 2. The diffuse reflectance spectroscopy system of claim 1, wherein the entity for emitting light is present at a fixed distance external to a photodiode. 3. The diffuse reflectance spectroscopy system of claim 1, wherein the entity for emitting light comprises one or more illumination fibers, each illumination fiber being present at a fixed distance external to a photodiode, optionally adjacent to a photodiode. 4. The diffuse reflectance spectroscopy system of claim 1, wherein the entity for emitting light comprises one or more illumination fibers, each illumination fiber being present within a photodiode. 5. (canceled) 6. The diffuse reflectance spectroscopy system of claim 4, wherein the photodiode comprises an aperture, and the illumination fiber is disposed within the aperture, optionally wherein spacing is present to vary the distance between the center of the aperture and/or fiber and an edge of the photodiode. 7. The diffuse reflectance spectroscopy system of claim 1, further comprising a light source coupled to the entity for emitting light, wherein the light source optionally comprises a lamp or a plurality of light-emitting diodes (LEDs). 8-9. (canceled) 10. The diffuse reflectance spectroscopy system of claim 1, wherein the entity for emitting light comprises direct illumination via a lamp or a plurality of light-emitting diodes (LEDs). 11-12. (canceled) 13. The diffuse reflectance spectroscopy system of claim 10, wherein the entity for emitting light and collecting entities are encased in a housing, where the entity for emitting light is at a proximal end of the housing and the one or more photodiodes are at a distal end of the housing, the one or more photodiodes each comprising an aperture, whereby the entity for emitting light provides backlit illumination through each aperture into one or more photodiodes. 14. The diffuse reflectance spectroscopy system of claim 13, wherein the housing comprises one or more reflective interior surfaces. 15. The diffuse reflectance spectroscopy system of claim 1, wherein the one or more photodiodes comprises an array of photodiodes. 16-17. (canceled) 18. An optical probe comprising at least one entity for emitting light into a tissue mass and at least one collecting entity for collecting light that has interacted with a tissue mass, wherein the collecting entity comprises one or more photodiodes. 19. The optical probe of claim 18, wherein the entity for emitting light is present at... | DUKE UNIVERSITY | A61B005/1455
US20110112383 | A1 | 2011-05-12 | Devices, systems and methods for treatment of eating disorders | 2011-01-04 | An ingestible gastric device comprising a formed body that is adapted to expand upon exposure to a hydration medium and then degrade after a first residence time in the stomach cavity, allowing the device to be passed by the patient's normal digestive process, and monitoring means for monitoring the gastric device. In some embodiments, the monitoring means comprises a biocompatible RFID tag. | 1. An ingestible gastric device, comprising: a formed body, said body being adapted to expand upon exposure to a hydration medium; andmonitoring means for monitoring the gastric device. 2. The gastric device of claim 1, wherein said body includes an outer layer adapted to receive a core material therein. 3. The gastric device of claim 1, wherein said monitoring means comprises a passive RFID tag. 4. The gastric device of claim 3, wherein said RFID tag is encapsulated in a material selected from the group consisting of poly(urethanes), poly(siloxanes), poly(methyl methacrylate), poly(vinyl alcohol) for hydrophilicity and strength, poly(ethylene), poly(vinyl pyrrolidone, poly(2-hydroxy ethyl methacrylate), poly(N-vinyl pyrrolidone), poly(methyl methacrylate), poly(vinyl alcohol), poly(acrylic acid), polyacrylamide, poly(ethylene-co-vinyl acetate), poly(ethylene glycol), poly(methacrylic acid), polylactides (PLA), polyglycolides (PGA), poly(lactide-co-glycolides) (PLGA), polyanhydrides, polyorthoesters, and mixtures and combinations thereof. 5. A food grade ingestible gastric device, comprising: a flexible pouch formed from a food grade biodegradable film, said pouch having a first volume that allows normal ingestion of said gastric device into a subject's stomach, said pouch enclosing a sealed cavity having a core material disposed therein,said biodegradable film comprising a first gel-forming material selected from the group consisting of acacia (gum Arabic), agar, agarose, alginate gels, alginic acid, alginate gums, amylopectin, arabinogalactan, carob bean gum, carrageenan, chondroitin sulfate, eucheuma, fucoidan, furcellaran, gelatin, gellan, guar gum, gum ghatti, gum karaya, hypnea, laminaran, locust bean gum, natural gum, pectins, starches, polypeptides, polyamino acids, tragacanth, xanthan, psyllium, maltodextrin, and mixtures and combinations thereof,said core material comprising a hydratable material that is adapted to effectuate expansion of said sealed cavity upon hydration of said hydratable material by a gastric fluid, whereby upon a first hydration of said degradable film and, thereby, said hydratable material, by a first gastric fluid, said gastric device expands to a second volume greater than said first volume, said second volume being sufficient to preclude passage of said gastric device into said subject's duodenum while still allowing food to pass therethrough,said gastric device being adapted to commence said first hydration after contact with said first gastric fluid for a period of time in the range of approximately 5-30 min,said gastric device being adapted to sufficiently degrade in said subject's stomach after a first period of time to allow passage of said gastric device into said subject's duodenum;said gastric device further including monitoring means for non-invasively monitoring the gastric device while disposed in a subject's gastrointestinal tract. 6. The gastric device of claim 5, wherein said monitoring means ... | Plensat, Inc | A61B005/07
US08372014 | B2 | 2013-02-12 | Post-operative bacteria test strip spool and method | 2011-02-18 | A method for testing body fluids is disclosed. The method includes providing a media-dispensing apparatus, placing at least the distal opening of the elongated tubular member adjacent body fluid (e.g., near a patient's abdomen) or adjacent a medical device such that a portion of the media strip contacts body fluid or the medical device, and advancing the media strip such that the portion of the media strip that contacted the body fluid or the medical device advances at least through the second lumen. Additionally, the step of cutting the media strip with a cutting instrument is also disclosed. | 1. A method of testing body fluids, comprising the steps of: providing a media-dispensing apparatus including a housing having a lower portion and an upper portion, an elongated tubular member extending distally from the housing and having a distal opening adjacent a distal end thereof, a first lumen disposed within the elongated tubular member extending from the lower portion of the housing to the distal opening, a second lumen disposed within the elongated tubular member extending from the upper portion of the housing to the distal opening, and a media strip extending from the lower portion of the housing, through the first lumen, out the distal opening of the elongated tubular member, through the second lumen and into the upper portion of the housing;placing at least the distal opening of the elongated tubular member adjacent a surgical site such that a portion of the media strip contacts the surgical site; andadvancing the media strip such that the portion of the media strip that contacted the surgical site advances at least through the second lumen. 2. The method of testing body fluids according to claim 1, wherein the media-dispensing apparatus further includes a cover hingedly coupled to the housing and a cutting instrument disposed on a portion of at least one of the housing and the cover, the cover being moveable between a first position where the cover is in proximity with the upper portion of the housing and at least a second position where the cover is spaced apart from the upper portion of the housing, such that moving the cover from the second position towards the first position cuts the media strip and wherein the method further includes the step of cutting the media strip. 3. The method of testing body fluids according to claim 1, wherein at least the distal opening of the elongated tubular member remains adjacent the surgical site as the media strip is advanced. 4. The method of testing body fluids according to claim 1, wherein at least a portion of the media-dispensing apparatus is inserted near a patient's abdomen. 5. The method of testing body fluids according to claim 1, wherein the media-dispensing apparatus comprising a seal disposed within at least one of the first lumen and the second lumen, the seal being configured to allow the media strip to pass therethrough and being configured to limit the passage of body fluids during advancement of the media strip. 6. The method of claim 5, wherein the seal is disposed adjacent a distal end of the first lumen. 7. The method of testing body fluids according to claim 1, wherein one lumen of the media-dispensing apparatus extends farther distally than the other lumen. 8. The method of testing body fluids according to claim 1, wherein the media strip is made of a material which is able to at least detect or identify at least one abnormality or at least one infection. 9. The method of testing body fluids according to claim 1, wherein at least a portion of the media strip is a c... | Covidien LP | A61B005/00
US08401631 | B2 | 2013-03-19 | Method and system for monitoring physiological conditions | 2010-10-18 | A method for providing an indication of a state of awareness for a patient, includes the steps of arranging data of an EEG and EMG power spectrogram to provide power versus frequency in a log-log arrangement; calculating a first best-fit line for a lower frequency region of the EEG power spectrogram; calculating at least a second best-fit line for a higher frequency region of the EEG power spectrogram. The display of these lines is augmented by displaying a template that identifies different regions on the display that help confirm the state of the patient. Secondly, the time domain EEG signals can be filtered and displayed such that different frequency bands can be simultaneously displayed or a single frequency band can be displayed according to different time scales. | 1. A method for providing an indication of a state of awareness for a patient, comprising the steps of: displaying data of an EEG power spectrum to provide a power versus frequency graph;determining, using a processor, a frequency value that separates a lower frequency region from a higher frequency region of the EEG power spectrum;generating, using the processor, a template, wherein the template has a plurality of contiguous angular sections originating at the frequency value; andproviding, using the processor, an indicator of the state of awareness of the patient based on a position of at least a portion of the higher frequency region coinciding with one of the plurality of contiguous angular sections. 2. The method of claim 1, wherein the plurality of contiguous angular sections has three angular sections. 3. The method of claim 1, wherein data of the EEG power spectrum is displayed in a log-log arrangement. 4. The method of claim 3, further comprising: calculating a first best-fit line for a lower frequency region of the EEG power spectrum; andcalculating at least a second best-fit line for a higher frequency region of the EEG power spectrum. 5. The method of claim 4, wherein the frequency value is the intersection point of the first and second best-fit lines. 6. The method of claim 5, further comprising: providing the indicator of the state of awareness of the patient based on a position of the second best-fit line coinciding with one of the plurality of contiguous angular sections. 7. The method of claim 4, further comprising: displaying the first and second best-fit lines. 8. The method of claim 4, wherein the plurality of contiguous angular sections includes: a first angular section which represents a region having a first slope steeper than about −40 dB/decade;a second angular section which represents a region having a second slope more shallow than about −30 dB/decade; anda third angular section which represents a region between the first and second angular sections. 9. The method of claim 8, wherein each of the plurality of contiguous angular sections has a respective associated color. 10. The method of claim 9, further comprising: determining in which particular one of the plurality of contiguous angular sections the second best-fit line is located; andwherein providing the indicator includes displaying the respective color associated with the particular one of the plurality of contiguous angular sections. 11. The method of claim 4, further comprising: determining the intersection point of the first best-fit line and the second best-fit lines. 12. The method of claim 1, wherein data of the EEG power spectrum is displayed in a log-linear arrangement. 13. The method of claim 1, wherein the plurality of contiguous angular sections includes: a first angular section which represents a region having a first slope steeper than about −40 dB/decade;a second angular section which represents a region having a second slope more shallo... | - | A61B005/04
US08631799 | B2 | 2014-01-21 | Respiratory therapy system including a nasal cannula assembly | 2009-01-23 | A nasal cannula for supplying a respiratory gas to a patient and a method of treating a patient with sleep disorder. The nasal cannula comprises a pair of spaced apart supply lines which each have a head at one end thereof with a discharge opening located therein. The opposite end of each supply line is connectable to a high flow respiratory gas source. Each head is sized to be snugly received and retained within one of the nasal cavities of the patient while forming a sufficient leakage passage, between a portion of inwardly facing nasal cavity skin of a patient and a portion of an exterior surface of the head, to facilitate exhausting of any excess respiratory gas supplied to the patient through the leakage passage and also facilitate inhalation of any room air required in excess of the respiratory gas to be supplied to the patient. | 1. A nasal cannula for supplying a respiratory gas to a patient, the nasal cannula comprising: a pair of supply lines which each have a head at one end thereof with a discharge opening therein for discharging a respiratory gas, and the opposite end of each of the pair of supply lines being connectable to a respiratory gas source, the respiratory gas being supplied at a flow rate of at least about 120 liters per minute and at a pressure of between 4 and 20 cm of H2O;a central bridge member, having a sufficiently short axial length, configured to span substantially no more than a width of a philtrum of a patient and being located adjacent the heads for integrally connecting to the supply lines to one another; andeach head being sized to be snugly received and retained within one of the nasal cavities of the patient while forming a sufficient leakage passage, between a portion of inwardly facing nasal cavity skin of a patient and a portion of an exterior surface of the head, to facilitate exhausting of excess respiratory gas supplied to the patient through the leakage passage;wherein the central bridge member defines an internal flow passage therein which provides an interconnecting gas flow path between which generally operates to equalize the flow of gas from the pair of supply lines to the pair of heads and provides alternate flow in the event that one of the heads becomes partially or completely obstructed or restricted during use; andthe nasal cannula forming an open system such that when the respiratory gas is introduced into the nostrils of the patient's nose, a portion of the respiratory gas, along with a portion of the exhaled gases from the patient, being allowed to leak out through the leakage passage via a nostril/head interface. 2. The nasal cannula according to claim 1, wherein each gas supply line, once attached to a patient in use, upon exiting the nasal cavities of the patient generally follows along a cheek of the patient and passes below a chin and near a mid-line of a face of the patient;a pair of adjustable straps are attached to the gas supply line and around a head of the patient to facilitate retention of the nasal cannula to the head of the patient;a first adjustable strap is attached to a first gas supply line and is adapted to pass from the first gas supply line, above a first ear of the patient, around a head of the patient, under a second ear of the patient and attach again to a second gas supply line; anda second adjustable strap is attached to the second gas supply line and is adapted to pass from the second gas supply line, above the second ear of the patient, around a head of the patient, under the first ear of the patient and attach again to the second gas supply line. 3. The nasal cannula according to claim 1, wherein an element is located within the gas supply line, along a gas flow path, and the element is coupled to a controller which controls flow of power to the element for heating the gas flowing through th... | Salter Labs | A61M016/00 A61M015/08
US08585617 | B2 | 2013-11-19 | Diagnosis and prediction of obstructive sleep apnea | 2009-12-21 | A tongue location monitoring system, including, one or more position circuits that respond to transmissions from a transceiver, a transceiver that transmits to the one or more position circuits, a control circuit coupled to the transceiver: wherein the control circuit determines the location of a person's tongue based on the responses of the position circuits; and wherein either the transceiver or the position circuits are implanted in the muscle of the person's tongue or placed on the tongue, and the latter is placed outside the person's head. | 1. A tongue location monitoring system, comprising: at least one position circuit configured to respond to wireless transmissions;a transceiver configured to wirelessly transmit to the at least one position circuit;a control circuit coupled to the transceiver;wherein said control circuit determines the location of a person's tongue based on the responses of the at least one position circuit;and wherein the at least one position circuit is configured for placement in or on the person's tongue, and the transceiver is configured for location outside the person's head. 2. A system according to claim 1, wherein said at least one position circuit is configured for implantation beneath the skin of the subject and on a muscle of the tongue. 3. A system according to claim 1, wherein said at least one position circuit is embedded in an implantable stimulator that is configured to be implanted in the tongue muscle and said implantable stimulator is adapted to stimulate the tongue muscle. 4. A system according to claim 1, wherein said at least one position circuit is embedded in a biodegradable tape that is configured to be adhesively attached to the tongue. 5. A system according to claim 1, wherein said control circuit determines the location of the tongue based on the travel time of the signal to the at least one position circuit and back. 6. A system according to claim 1, wherein said control circuit determines the location of the tongue based on the strength of the signal returning to the transceiver. 7. A system according to claim 1, wherein an obstructive sleep apnea event is detected by monitoring the location of the tongue. 8. A system according to claim 1, wherein an obstructive sleep apnea event is detected by monitoring the angle of arrival of a response signal from the at least one position circuit. 9. A system according to claim 1, wherein the at least one position circuit is a plurality of position circuits, and the plurality of position circuits are queried sequentially. 10. A system according to claim 1, wherein the at least one position circuit is a plurality of position circuits, and the plurality of position circuits are queried in parallel. 11. A system according to claim 1, wherein the at least one position circuit is queried continuously. 12. A system according to claim 1, wherein the at least one position circuit is queried periodically. 13. A system according to claim 1, wherein the transceiver is provided power wirelessly and the transceiver queries the at least one position circuit as long as it is provided power wirelessly. 14. A method of monitoring the location of the tongue, comprising: positioning a transceiver outside a head of a person;transmitting signals from the transceiver to at least one position circuit positioned in or on a tongue of the person;receiving a response from the at least one position circuit;calculating the location of the at least one position circuit relative to the transceiver from the respo... | Nyxoah SA | A61B005/103 A61B005/08
US08792977 | B2 | 2014-07-29 | Quantitative motor function evaluation system | 2008-02-26 | The system of the present invention includes (a) means for displaying image information including a target image and a cursor image for tracking the target image; (b) means used when the subject moves the cursor image; (c) means for detecting the state of tracking the target image by the cursor image; (d) means for detecting the muscle active state of the subject using the means (b); (e) means for analyzing the tracking state detected by the means (c) and the muscle active state detected by the means (d); and (f) means for evaluating the motor function of the subject by using results of analysis obtained by the means (e) as indexes. | 1. A motor function evaluation system for evaluating the motor function of a wrist motion of a subject, which comprises: a display configured for displaying image information including a target image and a cursor image for tracking the target image;a device configured for use by the subject for moving the cursor image by using wrist motion;a detector configured for detecting the state of tracking of the target image by the cursor image;a sensor configured for detecting myoelectric signals of extensor carpi radialis brevis muscle and extensor carpi radialis longus muscle (ECR), extensor carpi ulnaris muscle (ECU), flexor carpi ulnaris muscle (FCU) and flexor carpi radialis muscle (FCR) as the muscle active state of muscles related to the wrist motion of the subject using the device;an analyzer configured for analyzing the tracking state detected by the detector and the muscle active state detected by the sensor; andan evaluator configured for evaluating the motor function of the wrist motion of the subject by using results of analysis obtained by the analyzer as indexes. 2. The system according to claim 1, wherein the display comprises a display screen configured for displaying the image information. 3. The system according to claim 1, wherein the target image is at least one selected from the group consisting of: (i) an image which moves along a predetermined locus or moves in any direction;(ii) at least two images which are fixed at a predetermined interval;(iii) a line-like image having a predetermined length and width, which is constituted by a straight line and/or a curved line; and(iv) an image only consisting of a starting point and an end point. 4. The system according to claim 3, wherein the predetermined locus comprises at least one selected from the group consisting of a straight line, a curved line, a circle and a polygon. 5. The system according to claim 3, wherein each of the target images (i) and (ii) comprises at least one shape selected from the group consisting of a circle, an ellipse, a polygon and a star shape. 6. The system according to claim 3, wherein in the target image (ii), one target image is centered and two or more target images are positioned on a concentric circle of the centered target image. 7. The system according to claim 1, wherein the device is provided separately from the display. 8. The system according to claim 1, wherein the device comprises a movable part operated in any direction by the subject and an output part for transmitting, to the display, motion information of the movable part as information for moving the cursor image. 9. The system according to claim 8, wherein the device further comprises a sensor part configured for detecting a predetermined parameter regarding motion information of the movable part. 10. The system according to claim 9, wherein the predetermined parameter is at least one selected from the group consisting of a position of a wrist joint of a subject involved in operati... | Tokyo Metropolitan Institute of Medical Science Tokyo Metropolitan Government | A61B005/0488 A61B005/00 A61B005/11
US08380282 | B2 | 2013-02-19 | Method for acquiring electromagnetic signals and contrast product therefor | 2010-07-06 | The invention concerns a system capable of generating a magnetic indication B0 comprising gradients (Gx, Gy, Gz) in certain directions, transmitting ratio frequency wave pulse sequences (RF) perpendicular to B0 in a range of adjustable frequencies, and detecting electromagnetic signals received from a body part (4). The method includes injecting a contrast product in said body part, capable of being temporarily fixed in an observed zone (1), and comprising an element capable of causing chemical displacement of a resonance frequency of water hydrogen protons; exciting said body part, using a radio wave pulse sequence; in a range of frequencies adjusted on the basis of the magnetic induction B0 and the chemical displacement for some of said waves; detecting the electromagnetic signals received in said body part, substantially corresponding to the magnetic resonance signals of the protons of the observed zone having undergone the chemical displacement. | 1. A method for acquiring electromagnetic signals received from at least one part of a body placed in a system comprising means for generating a magnetic induction B0, said magnetic induction comprising gradients in certain directions in space, means for transmitting radio frequency wave pulse sequences perpendicular to the magnetic induction B0 in a range of adjustable frequencies, and means for detecting electromagnetic signals received from said body part, the method comprising the following steps: a) injecting, into said body part, an amount of contrast product capable of being temporarily fixed in or of passing through an observed zone of said body part, said contrast product comprising at least one element capable of causing a chemical shift of a resonance frequency of water hydrogen protons of said body part;b) determining a new resonance frequency, shifted with respect to the Larmor frequency (v0) for the water hydrogen protons of said body part nearby said contrast product;c) exciting said body part by means of a radio frequency wave pulse sequence in a range of frequencies adjusted according to the magnetic induction B0 and to the new resonance frequency determined at step b) resulting from the chemical shift of the resonance frequency of water hydrogen protons for at least some of said radio frequency waves;d) detecting, coherently with the excitation of step c), electromagnetic signals received from said body part, said signals corresponding substantially to magnetic resonance signals of the water hydrogen protons of the observed zone of said body part having undergone the chemical shift of the resonance frequency. 2. The method as claimed in claim 1, in which the element capable of causing a chemical shift of a resonance frequency of water hydrogen protons and included in the contrast product comprises a lanthanide. 3. The method as claimed in claim 2, in which the lanthanide is chosen from at least one of dysprosium, praseodymium and europium. 4. The method as claimed in claim 1, in which the contrast product also comprises a cage that incorporates the element capable of causing a chemical shift of a resonance frequency of water hydrogen protons. 5. The method as claimed in claim 1, also comprising a step consisting in forming an image from the electromagnetic signals received from said body part that are detected, according to a spatial coding dependent on the gradients of said magnetic induction. 6. The method as claimed in claim 1, in which said observed zone comprises a group of blood vessels. 7. The method as claimed in any one of claims 1 to 5, in which the contrast product is injected with a targeting molecule capable of being fixed to at least one target that is part of the observed zone. 8. The method as claimed in claim 7, in which the target is a group of cells expressing a gene of said body part. 9. The method as claimed in claim 1, in which the observed zone-comprises a tumor zone of said body part and in which ... | Centre National de la Recherche and Scientifique—CNRS Universite Victor Segalen Bordeaux 2 | A61B005/05
US07956894 | B2 | 2011-06-07 | Apparatus and method for computerized multi-media medical and pharmaceutical data organization and transmission | 2003-02-19 | An apparatus for multi-media data organization and transmission is provided. The apparatus has a computer having a microprocessor, a memory storage, a display for providing information to a user, and an input device. An image-recording device is electrically-coupled to the computer for capturing images for storage in the memory storage of the computer. A database, which has a structure defined in the memory storage, receives and stores a plurality of information relating to an event. A program, being executable by the computer, provides a graphical user interface on the display. The program has an imaging module with document and image capture filing and scanning functions. The graphical user interface receives an input from the input device and from the image-recording device. In a further aspect of the invention, the program has a communications module for transmission of the plurality of information relating to the event to a remote location. | 1. An apparatus for multi-media data organization and transmission of information relating to the provision of service on-site, the apparatus comprising: an on-site computer having a microprocessor, a memory storage, a display for providing information to an on-site service provider, and an input device for receiving information and commands from an on-site service provider;a remote computer having a microprocessor and a memory storage;a display for providing information to a remote service provider, and an input device for receiving information and commands from a remote service provider;an imaging device operatively coupled to said memory storage of said on-site computer for capturing images for storage in said memory storage of said on-site computer;a database defined in said memory storage of said on-site computer for storing a plurality of information relating to an event, said database including provision for annotations of said images;a database defined in said memory storage of said remote computer for storing a plurality of information relating to an event, said database including provision for annotations of said images;a program executable by said on-site computer, said program providing a graphical user interface on said display of said on-site computer, said program having an imaging module with document and image capture filing and scanning functions, said graphical user interface accepting an input from said input device of said on-site computer to control said program;a data transmission connection between said on-site computer and said remote computer for transferring data between the memory storage of said on-site computer and the memory storage of said remote computer; anda data synchronizer for making the data contents of fields of filed information available to both the on-site service provider and the remote service-provider in, respectively, said database defined in said memory storage of said on-site computer and said database defined in said memory storage of said remote computer the same, the data synchronizer further comprising:a sync system determining whether a sync file is present before the data contents of fields of information are transmitted. 2. The apparatus of claim 1 wherein the sync file comprises patient profile data. 3. The apparatus of claim 1 wherein the sync file comprises a patient folder. 4. The apparatus of claim 1 wherein the on-site computer comprises an excerpt transfer system transmitting an excerpt of a data record to a third party. 5. The apparatus of claim 1 further comprising a record encapsulation system encapsulating the data contents of the fields of filed information. 6. The apparatus of claim 4 further comprising a detail encapsulation system encapsulating the excerpt of the data record prior to transmission. 7. The apparatus of claim 1 wherein the remote computer further comprises a record notification system generating notification data after receipt of the fields of filed informa... | William Rex Akers | H04N005/225 H04N007/00 H04N011/00 G06Q010/00 G06Q050/00 A61B005/00
US07935057 | B2 | 2011-05-03 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2009-01-14 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. A system for processing data signals obtained from a glucose sensor, the system comprising: a receiving module that receives a signal from a glucose sensor over a period of time; anda processor module comprising programming that detects and measures a duration of detected transient non-glucose related signal artifacts and replaces at least some of the signal artifacts with estimated glucose values in response to the duration exceeding a predetermined amount of time. 2. The system of claim 1, further comprising a glucose sensor. 3. The system of claim 1, wherein the programming configured to replace at least some of the signal artifacts is configured to output the estimated glucose values in response to detection of transient non-glucose related signal artifacts, wherein the estimated glucose values comprise filtered, calibrated glucose values. 4. The system of claim 3, wherein the processor module further comprises programming configured to output unfiltered, calibrated glucose values when transient non-glucose related signal artifacts are not detected. | DexCom, Inc. | A61B005/00 A61B005/05
US07945083 | B2 | 2011-05-17 | Method for supporting diagnostic workflow from a medical imaging apparatus | 2006-05-25 | A method for supporting diagnostic workflow from a medical imaging apparatus. A set of at least two images are obtained from a patient and displayed according to a user-specified image display layout selected from a plurality of image display layouts. One or more markers are associated with a region of interest in the displayed images. A list of regions of interest is generated, each having an entry for each associated marker. A classification is assigned to each entry in the list of regions of interest according to health risk. | 1. A method for supporting diagnostic workflow from a medical imaging apparatus, the method comprising the steps of: a) obtaining a set of at least two images of a patient, wherein each of the at least two images is of the same anatomical feature of the patient but taken at different time periods;b) displaying, on a display, the set of images according to a user-specified image display layout selected from a plurality of image display layouts, and displaying each of the images of said set of images by electronically scanning the displayed images with a magnified view to provide a scanned view by the method comprising the steps of: (i) defining a magnification outline window that is smaller than the diagnostic image;(ii) positioning the magnification outline window over a portion of the diagnostic image to define a magnified image portion therein;(iii) displaying, on the display, the magnified image portion in a magnification display window; and(iv) changing the magnified image portion within the magnification display window with a repeated process of: changing the relative position of the magnification outline window according to a scan pattern, thereby redefining the magnified image portion and forming a redefined magnified image portion therein; and displaying the redefined magnified image portion in the magnification display window;c) associating one or more markers with each of the at least two images, each marker identifying a region of interest in the displayed image of each of the at least two images;d) generating a list of regions of interest having an entry for each marker associated in step c); ande) assigning a classification to each entry in the list according to possible health risk. 2. The method of claim 1 wherein the step of associating one or more markers includes automatically detecting the region of interest identified by each marker according to tissue characteristics. 3. The method of claim 1 wherein a marker operation is available to a user and wherein the step of associating one or more markers comprises the marker operation of responding to a user instruction for marker assignment. 4. The method of claim 1 wherein a number of utilities including an image processing algorithm are provided that assist in region of interest list review and wherein the step of assigning a classification includes the step of applying an image processing algorithm to detect tissue characteristics. 5. The method of claim 1 wherein the regions of interest identified by the markers in the associating step are presented as a list of regions of interest for review by a user. 6. The method of claim 1 wherein the generating step and the assigning a classification step further includes the step of generating a report that lists the list of regions of interest and the assigned classifications. 7. The method of claim 1 wherein the step of obtaining the set of at least two images is accomplished by obtaining mammography images. 8. The method of clai... | Carestream Health, Inc. | G06K009/00 A61B005/05
US07945306 | B2 | 2011-05-17 | Method for generating an image exposure of the heart requiring a preparation | 2006-09-06 | A method for generation of an image exposure of the heart of an examination subject with an imaging medical examination apparatus in particular a magnetic resonance apparatus, the image exposure requiring preparation includes the steps of: preparation of the heart in a heart position that is relevant for the image acquisition and determination of the associated heart position, determination of at least one current heart position in the further course of the heart cycle, comparison of the determined current heart position with the heart position relevant for the image acquisition, and given correlation of the current heart position and the heart position relevant for the image acquisition, starting the image acquisition through a control device of the imaging medical examination apparatus. | 1. A method for generating an image exposure of a heart, said image exposure being acquired with a medical imaging apparatus using an image exposure protocol that gives said image exposure a diagnostically-sufficient image quality, said method comprising the steps of: acquiring a first image comprising a depiction of the heart in a heart position, within cardiac cycle, that is relevant for said image exposure, by operating said medical imaging apparatus with a different imaging protocol from said image exposure imaging protocol;acquiring a plurality of second images, each at a different time within the cardiac cycle, also by operating said medical imaging apparatus with said different imaging protocol, each comprising a depiction of the heart in a different current position at a different time within the cardiac cycle;in a control unit of the medical imaging apparatus, automatically determining when substantial depiction correlation exists between the current position depicted in one of said second images and the position of the heart depicted in said first image; andwhen said substantial depiction correlation is determined to exist, triggering, from said control unit, operation of said medical imaging apparatus to acquire said image exposure with said image exposure protocol. 2. The method as claimed in claim 1 comprising starting said image exposure acquisition completely automatically through said control unit. 3. The method as claimed in claim 1 comprising starting said image exposure acquisition semi-automatically through said control unit, dependent on an input manually entered into said control unit. 4. The method as claimed in claim 1 comprising automatically, without manual intervention, determining at least one of said heart position relevant for the image exposure acquisition and the current heart position. 5. The method as claimed in claim 1 comprising determining at least one of the heart position relevant for the image exposure acquisition and the current heart position based on a parameterization entered into said control unit. 6. The method as claimed in claim 1 comprising the first image acquisition as a first slice image, and determining said current heart position as second slice images image, and determining said substantial correlation of the current heart position and the heart position relevant for the image exposure acquisition by comparing said first slice image and said second slice images in said control unit image. 7. The method as claimed in claim 1 comprising determining said substantial correlation of the current heart position and the heart position relevant for the image acquisition by executing a computer program in said control unit. 8. The method as claimed in claim 1 comprising determining said substantial correlation by superimposition of said first image with each of said second image. 9. The method as claimed in claim 1 comprising determining said substantial correlation by detection of image featur... | Siemens Aktiengesellschaft | A61B005/05
US07955269 | B2 | 2011-06-07 | Method and apparatus for pulmonary artery pressure signal isolation | 2009-06-22 | An implantable system senses a pulmonary artery pressure (PAP) signal using an implantable sensor placed in the pulmonary artery and isolates a plurality of signals from the PAP signal for diagnostic and/or therapeutic use. Each signal is extracted from the PAP signal using its known frequency characteristics and/or timing relationship with one or more detectable events. | 1. A method, comprising: receiving a pulmonary artery pressure (PAP) signal from an implantable PAP sensor through a wireless communication link;adjusting the PAP signal using a signal indicative of posture; andisolating a plurality of signals from the adjusted PAP signal using a signal processor. 2. The method of claim 1, wherein the plurality of signals comprises one or more cardiac signals and one or more respiratory signals. 3. The method of claim 2, wherein isolating the plurality of signals from the PAP signal comprises: isolating a cardiac signal using an adaptive filter having a pass-band; andadjusting the pass-band dynamically using a physiological signal or parameter. 4. The method of claim 2, wherein isolating the plurality of signals from the PAP signal comprises: isolating a respiratory signal using an adaptive filter having a pass-band; andadjusting the pass-band dynamically using a physiological signal or parameter. 5. The method of claim 1, comprising: receiving the PAP signal using an implantable medical device; andisolating the plurality of signals from the PAP signal using the implantable medical device. 6. The method of claim 1, wherein receiving the PAP signal comprises receiving an ultrasonic signal through an ultrasonic telemetry link and demodulating the ultrasonic signal. 7. The method of claim 1, wherein receiving the PAP signal comprises receiving an electromagnetic signal through a far-field radio-frequency telemetry link and demodulating the electromagnetic signal. 8. The method of claim 1, wherein receiving the PAP signal comprises receiving a magnetic signal through an inductive telemetry link and demodulating the magnetic signal. 9. The method of claim 1, further comprising adjusting the PAP signal using an atmospheric pressure before isolating the plurality of signals. 10. The method of claim 1, wherein isolating the plurality of signals comprises: filtering the PAP signal using at least one adaptive filter having a pass-band; anddynamically adjusting the pass-band using a physiological signal or parameter. 11. The method of claim 10, wherein dynamically adjusting the pass-band comprises: detecting a heart rate; andadjusting the pass-band dynamically as a function of the heart rate. 12. The method of claim 1, wherein isolating the plurality of signals comprises isolating the one or more cardiac signals by sampling the PAP signal at a predetermined type event in each cycle of a cyclic physiological signal. 13. The method of claim 1, wherein the plurality of signals comprises one or more of an atmospheric pressure signal, a posture signal, a weather signal, an altitude signal, a Valsalva signal, and a Mueller signal. 14. The method of claim 1, further comprising delivering one or more therapies using an implantable medical device. 15. The method of claim 14, further comprising controlling the delivery of the one or more therapies using one or more signals selected from the plurality of signals isolated... | Cardiac Pacemakers, Inc. | A61B005/02
US07959630 | B2 | 2011-06-14 | Medical probes for creating and diagnosing circumferential lesions within or around the ostium of a vessel | 2008-09-11 | The present inventions provide assemblies, probes, and methods for creating circumferential lesions in tissue, e.g., the tissue within or around the ostium of a vessel. An ablation probe with an ablative structure can be placed in contact within or around the ostium of the vessel. A diagnostic probe can be introduced through a lumen within the ablation probe and inserted into the vessel. The energy can be provided to the ablative structure to create a circumferential lesion within or around the ostium of the vessel, and the diagnostic structure can be used to diagnose the tissue to determine whether the circumferential lesion can be properly created. | 1. A probe for ablating tissue, comprising: an elongate probe body having a main member disposed in a first plane, and a distal member configured for forming a loop-shaped structure, wherein the probe body is pre-shaped in a manner that disposes the loop-shaped structure in a second plane in the range of 30 to 60 degrees with the first plane; andone or more ablative elements mounted to the loop-shaped structure, wherein the one or more ablative elements are arranged to create a circumferential lesion. 2. The probe of claim 1, wherein the probe body is an intravascular catheter body. 3. The probe of claim 1, wherein the loop-shaped structure is configured to be disposed around the ostium of a pulmonary vein, and the one or more ablative elements are configured to circumferentially contact tissue around the ostium of the pulmonary vein. 4. The probe of claim 1, further comprising a lumen extending through the probe body, the lumen configured to slidably receive an inner elongate probe body, the lumen having an exit port out which the inner probe body can extend through the loop-shaped structure. 5. The probe of claim 4, wherein the exit port is proximal to the loop-shaped ablative structure. 6. The probe of claim 1, further comprising a pull-wire connected to the distal member, wherein the pull-wire is configured for being pulled to form the distal member into the loop-shaped ablative structure. 7. The probe of claim 1, wherein the second plane is out-of-plane with the first plane by 45 degrees. 8. The probe of claim 1, wherein the probe body further includes a resilient core wire that pre-shapes the probe body in a manner that disposes the loop-shaped structure in the second plane. 9. A probe assembly for ablating tissue, comprising: an outer probe including an elongate probe body having a main member disposed in a first plane, and a distal member configured for forming a loop-shaped ablative structure, wherein the probe body is pre-shaped in a manner that disposes the loop-shaped structure in a second plane in the range of 30 to 60 degrees with the first plane, the outer probe further including a lumen having an exit port, and one or more ablative elements mounted to the loop-shaped ablative structure; andan inner probe configured to be slidably disposed within the lumen of the outer probe, the inner probe including an elongate probe body having a distal diagnostic structure configured to extend out the exit port and through the loop-shaped structure, and one or more diagnostic elements mounted to the distal diagnostic structure. 10. The probe assembly of claim 9, wherein the outer probe body is an intravascular catheter body. 11. The probe assembly of claim 9, wherein the loop-shaped ablative structure is configured to be disposed around the ostium of a pulmonary vein, and the one or more ablative elements are configured to circumferentially contact tissue around the ostium of the pulmonary vein. 12. The probe assembly of claim 9, whe... | Boston Scientific Scimed, Inc. | A61B018/14 A61B005/042
US07963917 | B2 | 2011-06-21 | System and method for continuous non-invasive glucose monitoring | 2005-12-05 | A system and method for continuous non-invasive glucose monitoring is disclosed. According to one embodiment of the present invention, the method includes the steps of (1) contacting a remote device to an area of biological membrane having a permeability level, the remote device comprising a sensor and a transmitter; (2) extracting the at least one analyte through and out of the area of biological membrane and into the sensor; (3) generating an electrical signal representative of a level of the at least one analyte; (4) transmitting the electrical signal to a base device; (5) processing the electrical signal to determine the level of the at least one analyte; and (6) displaying the level of the at least one analyte in real time. The system includes a remote device that includes a sensor that generates an electrical signal representative of the concentration of the at least one analyte; and a transmitter that transmits the electrical signal. The system further includes a base device that includes a receiver that receives the electrical signal; a processor that processes the electrical signal; and a display that displays the processed signal in real time. | 1. A method for real time remote monitoring and display of a level of at least one analyte in a body fluid of a subject, comprising: contacting a remote device to an area of biological membrane having a permeability level, the remote device comprising a sensor and a transmitter;extracting the analyte through and out of the area of biological membrane and into the sensor;generating an electrical signal representative of a level of the analyte;transmitting the electrical signal to a base device;processing the electrical signal to determine the level of the analyte and the rate of change of the level of analyte, wherein the one or more future analyte level(s) are calculated based on the determined level of analyte and an average of the determined rate of change of the level of analyte;displaying the level of the at least one analyte in real time; andtriggering an alarm if (1) one or more determinations of the level of analyte is/are below a hypoglycemic preset limit or above a hyperglycemic preset limit and (2) the determined rate of change of the level of analyte predicts that one or more future analyte level(s) will be below the hypoglycemic preset limit or above the hyperglycemic preset limit within a period of time. 2. The method of claim 1, wherein the at least one analyte is glucose. 3. The method of claim 1, further comprising the step of: increasing the permeability level of the area of biological membrane. 4. The method of claim 1, wherein the step of contacting a remote device to an area of biological membrane having a permeability level comprises: affixing the remote device to the area of biological membrane with an adhesive. 5. The method of claim 1, wherein the step of transmitting the signal to a base device comprises: converting the electrical signal representative of the level of the at least one analyte to a digital signal; andtransmitting the digitized signal, an identification number of the remote device, and a time stamp to the base device. 6. The method of claim 1, wherein the step of transmitting the signal to a monitoring device is performed periodically. 7. The method of claim 1, wherein step of the transmitting the electrical signal to a base device comprises: transmitting the signal to the base device by at least one of a wireless application protocol link, a general packet radio service link, a Bluetooth radio link, an IEEE 802.11-based radio frequency link, a RS-232 serial connection, an IEEE-1394 (Firewire) connection, a fibre channel connection, an infrared (IrDA) port, a small Computer Systems Interface (SCSI) connection, and a Universal Serial Bus (USB) connection. 8. The method of claim 1, further comprising the step of: periodically determining an amount of a drug to be injected in response to the analyte level; andautomatically providing the determined amount of the drug to the subject. 9. The method of claim 8, wherein the drug is insulin. 10. The method of claim 1, wherein the analyte level is displayed ... | Echo Therapeutics, Inc. | A61B005/00 A61B005/05
US07938788 | B2 | 2011-05-10 | Foot typing method | 2007-01-04 | A method for classifying foot types by measuring the position of the Rearfoot Supinatory End Range of Motion, the Rearfoot Pronatory End Range of Motion, the Forefoot Supinatory End Range of Motion and the Forefoot Pronatory End Range of Motion, and classifying the foot based on the results of these measurements is disclosed. Once a human foot is classified and typed according to the disclosed method, a treatment plan can be customized for the specific foot type. The treatment can include the use of pre-orthotic pads and orthotic inserts which are customized according to a patient's specific foot type. | 1. A method for classifying foot types based on positional measurements of a human foot comprising the steps of: with a patient in a supine position, a. measuring a Rearfoot Supinatory End Range of Motion position in the foot and determining whether the position is one of inverted or everted;b. measuring a Rearfoot Pronatory End Range of Motion position in the foot with reference to a bisection of the lower one third of a patient's leg after applying an eversion force upon the Calcaneous until it can no longer move, and determining whether the position is one of inverted, perpendicular or everted;c. measuring a Forefoot Supinatory End Range of Motion position in the foot with reference to the foot's fifth metatarsal after applying a plantarflectory force downward upon the foot's first metatarsal from above until it can no longer move and determining whether the position is one of dorsiflexed or plantarflexed;d. measuring a Forefoot Pronatory End Range of Motion position in the foot with reference to the fifth metatarsal after applying a dorsiflectory force upward upon the first metatarsal from below until it can no longer move and determining whether the position is one of dorsiflexed, vertical or plantarflexed; ande. classifying the foot into a distinct foot type selected from a group of sixteen different foot types, each of said sixteen foot types being determined based on the position of the foot according to results of steps a) through d). 2. The method according to claim 1 wherein a Rearfoot in a human foot is classified based on the positional measurements of steps a) and b). 3. The method according to claim 1 wherein a Forefoot in a human foot is classified based on the positional measurements of steps c) and d). 4. A method of treating podiatric conditions in a human patient comprising the steps of: with a patient in a supine position, a. measuring a Rearfoot Supinatory End Range of Motion position in a human foot andb. measuring a Rearfoot Pronatory End Range of Motion in the foot with reference to a bisection of the lower one third of a patient's leg after applying an eversion force upon the Calcaneous until it can no longer move, and determining whether the position is one of inverted, perpendicular or everted;c. measuring a Forefoot Supinatory End Range of Motion position in the foot with reference to the foot's fifth metatarsal after applying a plantarflectory force downward upon the foot's first metatarsal from above until it can no longer move and determining whether the position is one of dorsiflexed or plantarflexed;d. measuring a Forefoot Pronatory End Range of Motion position in the foot with reference to the fifth metatarsal after applying a dorsiflector force upward upon the first metatarsal from below until it can no longer move; and determining whether the position is one of dorsiflexed, vertical or plantarflexed;e. classifying the foot into a distinct foot type selected from a group of sixteen different foot types, each... | - | A61B005/103
US07949383 | B2 | 2011-05-24 | Imaging apparatus with selectable moveable stage | 2005-05-27 | An imaging box assembly is provided for capturing an image of a sample. The imaging box assembly includes a body having an interior cavity for receiving the sample, and having a front portion defining an opening into the cavity. The body further includes a view port enabling viewing of the sample contained in the interior cavity. A door is mounted to the body that is movable between an opened condition, enabling access to the interior cavity through the cavity opening, and a closed condition, positioning a door rear portion substantially adjacent the body front portion to prevent access through the cavity opening. The box assembly further includes a moveable stage disposed in the cavity interior that supports the sample. The moveable stage is adapted to selectively position the sample at a selected one of a plurality of positions relative to the view hole. | 1. A light-tight imaging system for capturing an image of a sample, the imaging system comprising: an imaging box including: a body including an interior cavity for receiving the sample, and having a front portion defining an opening into said cavity, said body further including a view hole enabling viewing of the sample; anda door having a rear portion, said door movable between an opened condition, enabling access to the interior cavity through the cavity opening, and a closed condition, positioning said rear portion substantially adjacent the body front portion to prevent access through the cavity opening;a moveable stage disposed in the cavity interior that supports the sample, said moveable stage being adapted to selectively position the sample at a selected one of a plurality of vertical positions relative to the view hole, wherein each of the plurality of vertical positions has substantially the same horizontal position in the interior cavity and each is substantially linearly closer to or substantially linearly further away from said view hole with respect to every other vertical position; anda light shield supported by said moveable stage, said light shield preventing light from said sample to translate horizontally to a portion of the moveable stage that does not support said sample, wherein said light shield includes a substantially vertical wall upstanding from a surface of the moveable stage. 2. The imaging system according to claim 1, further including: a temperature control element configured to alter the temperature of one of the sample and the interior cavity. 3. The imaging system according to claim 2, wherein the temperature control element includes a thermal sheet in thermal communication with said moveable stage. 4. The imaging system according to claim 1, further including: a gas manifold device in communication with said interior cavity, and configured to supply a gas to the sample, wherein said gas manifold device is coupled to said moveable stage. 5. The imaging system according to claim 1, further including: a positioning sensor, wherein said positioning sensor cooperates with the moveable stage to determine its position relative to said view port. 6. The imaging system according to claim 5, wherein said positioning sensor includes a laser positioned in said interior cavity to intercept said moveable stage at the selected one of a plurality of positions relative to the view hole at said first vertical position. 7. The imaging system according to claim 1, further including: a drive assembly coupled to the moveable stage to selectively move said moveable stage to the selected one of a plurality of positions. 8. The imaging system according to claim 1, further including: a partition located in the cavity interior, the partition forming a compartment within the body separate from the interior cavity. 9. An imaging system for capturing an image of a sample, the imaging system comprising: an imaging box having: a body ... | Xenogen Corporation | A61B005/05
US07944208 | B2 | 2011-05-17 | Active image area positioning with adjustable patient bench permitting chair-shaped patient transport into and out of an open MRI apparatus system | 2008-11-24 | A Magnetic Resonance Imaging System consisting of a magnet (10) with a U-shaped frame (15), whose pole faces define an open magnetic imaging area [R] for a patient, and a magnetic field generator (17) that is controlled to generate magnetic fields inside the magnetic imaging area; a transport system (30) associated with the magnet has a support structure (20) that defines a movement path through the magnetic imaging area. The transport system (30) also includes a bench (40) to support the patient. The transport system (30) can slide the bench (40) along the support structure so that the patient can be introduced into and extracted from the magnetic imaging area. The bench (40) may further include a non-reclinable support part (41) coupled to reclinable support parts (42 and 43) at opposing ends to allow the position of patient to be rotatably changed on the bench (40). | 1. A magnetic resonant imaging system comprising: a main magnet having opposing poles, positioned along a horizontal axis, that define an open magnetic resonance imaging area configured to detect a magnetic field generated by the main magnet; anda transport system comprising: a bench comprising at least two parts configured to support a patient, that is capable of moving the patient into and out of the open magnetic resonance imaging area with the at least two parts of the bench being capable of forming and maintaining the form of a chair shaped seat, while the patient is being moved into and out of the open magnetic resonance imaging area of the main magnet,wherein said at least two parts are pivotally attached to one another at, at least one rotation point around the horizontal axis, anda support structure configured to support the bench, wherein the support structure comprises: a track passing between said opposing poles of the main magnet, that is also at a fixed position relative to the main magnet, such that the bench moves along a movement path on the fixed position of the track into and out of the open magnetic resonance imaging area, andwherein at least one part, of the at least two parts of the bench which are configured to support a patient is rotatable at said at least one rotation point while within the open magnetic resonance imaging area in order to position the patient in any desired slanted position after the patient is, or has been moved by the bench and positioned within the open magnetic resonance imaging area which allows any part of the anatomy of the patient, on the bench, to be properly positioned at the center of the open magnetic resonance imaging area. 2. The magnetic resonant imaging system of claim 1, wherein the bench comprises: a non-reclinable support part andat least one reclinable support part rotatably coupled to the non-reclinable support part at the at least one rotation point. 3. The magnetic resonant imaging system of claim 2, wherein two reclinable support parts are rotatably coupled to the opposing ends of the non-reclinable support part at corresponding rotation points. 4. The magnetic resonant imaging system of claim 1, wherein said track comprises a transmission belt loop that is made from a non-metallic material and is configured to move the bench along the movement path. 5. The magnetic resonant imaging system of claim 4, further comprising at least one pulley to engage with the transmission belt loop at a drive point. 6. The magnetic resonant imaging system of claim 1, wherein said transport system comprises a pantograph mechanism, wherein a rotational movement at a first part of the bench corresponds to a rotational movement at a second part of the bench. 7. The magnetic resonant imaging system of claim 6, wherein said pantograph mechanism comprises a plurality of substantially parallel telescopic supports correspondingly coupled to the first part and to the second part of the bench. 8. The m... | Paramed Medical Systems, Inc. | G01V003/00 A61B005/05
US07970450 | B2 | 2011-06-28 | Wireless medical monitoring apparatus and system | 2002-05-03 | An apparatus for monitoring an electrical signal from a patient's body includes a disposable electrode patch having a thin flexible housing with an adhesive exterior, a power source, a printed circuit board, a plurality of electrodes, a converter for converting a detected electrical signal from the patient's body to a digital signal, a processor for processing the digital signal, and a transmitter connected for transmitting the processed digital signal as a wireless signal. A monitoring unit communicating with the electrode patch includes a power source, a transceiver, a global positioning receiver, a processor, at least one communication port for external communications, and a display. A system of the invention includes a plurality of patients having medical monitors wirelessly communicating biometric information to a central processor for archiving and accessing. | 1. An apparatus for wireless monitoring of at least one electrical signal from a patient's skin, said apparatus comprising in combination a disposable electrode patch and a monitoring unit: the disposable electrode patch having a relatively thin and flexible sealed housing having an adhesive surface effective for releasably adhering said electrode patch to the patient's skin, said housing having a single continuous surface with a pair of end portions and a center portion disposed between the end portions, each of the end portions having a width that is greater than a width of the center portion, a power source comprising a disposable battery positioned in the housing for providing power, the housing further containing a printed circuit board operatively connected to the power source for distributing power, a plurality of electrodes operably connected with said circuit board and positioned only in the end portions of said housing and are spaced apart so as to detect the electrical signal when the patch is properly adhered to the patient's skin, one of the electrodes being a ground electrode for grounding the electrode patch to the patient, a converter connected with said plurality of electrodes for receiving a plurality of detected electrical signals and for converting the plurality of detected electrical signals to a single digital signal, a processor having a clock and connected with said converter for processing the digital signal responsive to time, a transceiver for transmitting and receiving wireless signals, the transceiver connected with said processor for transmitting a processed digital signal via an antenna positioned in the housing as a wireless signal by short range radio frequency to the monitoring unit; andthe monitoring unit being reusable and portable by the patient, the monitoring unit comprising a transceiver effective in wireless communication with said electrode patch for receiving the wireless signal transmitted by the electrode patch and then transmitting the processed digital signal via a cellular telephone system to a computer, the monitoring unit further comprising at least one communication port for external communications and the monitoring unit is remotely programmable from a base station processor communicating therewith through at least one communication port. 2. The apparatus of claim 1, wherein said monitoring unit further comprises an activation switch for manually activating said apparatus. 3. The apparatus of claim 1, wherein said monitoring unit further comprises a plurality of light sources connected to function as visual indicators of apparatus status. 4. The apparatus of claim 1, wherein at least one communication port comprises an RS-232 port. 5. The apparatus of claim 1, wherein at least one communication port is responsive to infrared light. 6. The apparatus of claim 1, wherein said monitoring unit further comprises a power source that is rechargeable. 7. The apparatus of claim 6, wherein said monit... | Halthion Medical Technologies, Inc. | A61B005/0402
US09498152 | B2 | 2016-11-22 | Method and system for expediting the rescue of victims experiencing sudden cardiac arrest (SCA) when used in conjunction with an automated external defibrillator (AED) | 2010-10-25 | A method and system for expediting the rescue of victims experiencing sudden cardiac arrest (SCA) when used in conjunction with an automated external defibrillator (AED). The system involves an apparatus that calls for help (i.e. 911), detects SCA and cardiac arrhythmias, locates an AED, guides the rescuer through CPR and connection of the AED. A second apparatus is in the enclosure that contains the AED, sends location information to the other apparatus, triggers audible and visual indicators, and provides information on the victim's location. | 1. A system comprising a sudden cardiac arrest (SCA) detection apparatus and a location apparatus, the location apparatus being separate from and located with an automatic external defibrillator (AED), the SCA detection apparatus comprising: at least one battery;one or more electrodes for attachment to a patient;a wire and connector for connection to an AED;a status indication system;a voice messaging system;at least one controller;controls, components and circuitry to allow wireless communications; andcontrols, components and circuitry to allow Global Positioning System (GPS);wherein the SCA detection apparatus is configured to: call for help when an event occurs, wherein the call for help is made in accordance with a list of prioritized numbers, with the SCA detection apparatus calling a secondary number if there is no answer at a primary number, and further wherein the call for help is performed as a background task so as not to affect further operation of the SCA detection apparatus;analyze a victim and detect whether the victim is in SCA;if SCA is detected, guide a rescuer through performing cardiopulmonary resuscitation (CPR);locate a nearby AED by communicating with the location apparatus located with the nearby AED;indicate the location of the victim to the nearby AED;guide the rescuer to the location of the nearby AED;guide the rescuer on how to connect an AED to the SCA detection apparatus;wherein the location apparatus is configured to: provide the SCA detection apparatus with the location of the nearby AED;notify the nearby AED that the event has occurred;provide the rescuer with the location of the victim. 2. A system according to claim 1 wherein the SCA detection apparatus has controls and circuitry to measure and analyze the patient's ECG signal. 3. A system according to claim 2 wherein the SCA detection apparatus has controls and circuitry to measure and analyze the patient's transthoracic impedance signal. 4. A system according to claim 3 wherein the SCA detection apparatus has controls and circuitry to measure and analyze one or more the patient's: temperature;blood oxygen level;blood pressure;pulse;heart rate. 5. A system according to claim 2 wherein the SCA detection apparatus analyzes the patient's ECG to determine if the patient's heart rhythm is shockable. 6. A system according to claim 5 wherein the SCA detection apparatus determines if the patient's heart rhythm is shockable and communicates the results to another wireless device. 7. A system according to claim 5 wherein the SCA detection apparatus determines if the patient's heart rhythm is shockable and uses the voice messaging system to notify the rescuer of the results. 8. A system according to claim 5 wherein the SCA detection apparatus measures and stores the patient data in internal memory. 9. A system according to claim 8 wherein the SCA detection apparatus upon request transmits patient data stored in internal memory to another wireless device during non-... | Scion Medical Limited | A61B005/0205 A61B005/11 A61B005/0408 G08B025/08 G08B025/12 G09B023/28 A61B005/053 A61N001/39 A61B005/00
US09649089 | B2 | 2017-05-16 | Portable ultrasound scanner and docking system | 2009-11-17 | An apparatus includes a plurality of portable scanners and docks. The scanners include first and second configurations to provide different functionality. The docks include at least one of floor, a wall, and a table dock. The scanners are selectively mateable to the various docks. | 1. A system comprising: a first dock including a first common portion and a base portion, wherein the first common portion of the first dock includes a support member with a first end region, a second end region, and a middle region there between, a keyboard mounted to only the first end region, and a portable ultrasound scanner support mounted to only the second end region, wherein the portable ultrasound support comprises at least one mechanical connector, and the base portion of the first dock is configured to support the first common portion of the first dock relative to a first surface, and wherein the first common portion of the first dock further includes one or more electrical connectors;a second dock including a first common portion and a base portion, wherein the first common portion of the second dock includes a support arm with a first end region, a second end region, and a middle region there between, a keyboard mounted to only the first end region, and a portable ultrasound scanner support mounted to only the second end region, wherein the portable ultrasound scanner support comprises at least one mechanical connector, the base portion of the second dock is configured to support the first common portion of the second dock relative to a second surface and wherein the first common portion of the first dock and the first common portion of the second dock are the same and the base portion of the first dock and the base portion of the second dock are different, and wherein the first common portion of the second dock further includes one or more electrical connectors; anda first portable ultrasound scanner including an ultrasound transducer interface, a display on a first side of the scanner, at least one complementary mechanical connector and electrical connectors on a second side of the first scanner which is opposite the first side and which is configured to attach the first portable ultrasound scanner to the at least one mechanical connector and electrical connectors, and a dock interface that selectively connects the first scanner to the first and second docks for operation in conjunction therewith. 2. The apparatus of claim 1 including a second portable ultrasound scanner including an ultrasound transducer interface, a display on a first side of the scanner, at the first side, and which is configured to attach the second portable ultrasound scanner to the at least one mechanical connector, and a dock interface that selectively connects the second scanner to the first and second docks for operation in conjunction therewith, wherein the first scanner performs a first set of functions and the second scanner performs a second set of functions, and the first set is a subset of the second. 3. The apparatus of claim 2 wherein the first set includes a digital information and communication in medicine function and a communication and color flow mode function and the second set includes a three dimensional image display function and a brach... | B-K Medical Aps | A61B008/00 A61B005/00
US09492226 | B2 | 2016-11-15 | Graphical user interface for real-time RF lesion depth display | 2010-12-08 | A system for displaying characteristics of target tissue during an ablation procedure is provided that includes an electronic control unit (ECU) configured to receive data regarding electrical properties of the target tissue for a time period. The ECU is also configured to determine a value responsive to the data and indicative of at least one of a predicted depth of a lesion in the target tissue, a predicted temperature of the target tissue, and a likelihood of steam pop of the target tissue for the time period. The system further includes a display device operatively connected to the ECU. The display device is configured to receive the value and display a visual representation indicative of at least one of a predicted depth of a lesion in the target tissue, a predicted temperature of the target tissue, and a likelihood of steam pop of the target tissue for the time period. | 1. A graphical user interface system for displaying characteristics of target tissue during an ablation procedure, the system comprising: an electronic control unit (ECU) configured to: receive as an input data a complex impedance measured between a first electrode and the target tissue for a time period during the ablation procedure, wherein the time period is a time during which ablation energy is applied to the target tissue,receive at least one characteristic of the target tissue,receive a signal indicative of a contact force measured between a second electrode and the target tissue for the time period,determine a predicted depth value of a lesion in the target tissue based on each of the following: i) the complex impedance, ii) the time period, and iii) the signal indicative of the contact force,associate the predicted depth value with the at least one characteristic of the target tissue, andoutput the predicted depth value; anda display device operatively connected to the ECU, wherein the display device is configured to receive the predicted depth value and display a visual representation of the predicted depth value of the lesion in the target tissue for the time period. 2. The graphical user interface system of claim 1, wherein the visual representation is displayed during the time period. 3. The graphical user interface system of claim 1, wherein the at least one characteristic of the target tissue comprises a thickness of the target tissue. 4. The graphical user interface system of claim 3, wherein the display device is further configured to display a visual representation indicative of the thickness of the target tissue. 5. The graphical user interface system of claim 4, wherein the visual representation of the predicted depth value of the lesion in the target tissue for the time period comprises an expandable graph or meter. 6. The graphical user interface of claim 5, wherein the expandable graph or meter is at least partially overlaid over the visual representation indicative of the thickness of the target tissue. 7. The graphical user interface system of claim 1, wherein the data regarding the complex impedance is generated by measuring electrical properties of the target tissue with a sensor on an ablation catheter. 8. The graphical user interface system of claim 1, wherein the system is configured to store a first color associated with a first range of predicted depth values of the lesion in the target tissue for the time period. 9. The graphical user interface system of claim 8, wherein the system is configured to store a second color associated with a second range of predicted depth values of the lesion in the target tissue for the time period. 10. The graphical user interface system of claim 9, wherein the system is configured to store a third color associated with a third range of predicted depth values of the lesion in the target tissue for the time period. 11. The graphical user interface system of claim 10, wherei... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B018/04 A61B018/14 A61B018/12 A61B017/00 A61B018/00
US20110152661 | A1 | 2011-06-23 | Method and apparatus for identifying the viability of ischemic myocardium of a patient | 2010-09-22 | A method for identifying the viability of ischemic myocardium of a patient. The method includes the steps of measuring in real-time the ischemic myocardium of a beating or non-beating heart. There is the step of determining in real-time whether the ischemic myocardium of the beating or non-beating heart is stunned or is nonviable. An apparatus for identifying the viability of ischemic myocardium of a patient's heart. The apparatus includes an electrode array having at least four electrodes for electrical communication with the heart which produces an array signal. The apparatus includes a processor portion in communication with the array which receives the array signal and determines in real-time whether the ischemic myocardium of the heart is stunned or is nonviable. An apparatus for analyzing living tissue. The apparatus includes an electrode array having at least four electrodes for electrical communication with the tissue which produces an array signal. The apparatus includes an admittance magnitude and phase detection circuit in communication with the array which receives the array signal from the array and produces a detection circuit signal corresponding to phase angle and magnitude of the array and signal. The apparatus includes a processor in communication with the detection circuit which receives the detection circuit signal and determines permittivity and conductivity of the tissue in real-time. A method for analyzing living tissue. | 1. A method for identifying the viability of ischemic myocardium of a patient comprising the steps of: measuring in real-time the ischemic myocardium of a beating or non-beating heart; anddetermining in real-time whether the ischemic myocardium of the beating or non-beating heart is stunned or is nonviable. 2. A method as described in claim 1 wherein the measuring step includes the step of measuring the ischemic myocardium's electrical admittivity of the beating or non-beating heart. 3. An apparatus for identifying the viability of ischemic myocardium of a patient's heart comprising: an electrode array having at least four electrodes for electrical communication with the heart which produces an array signal; anda processor portion in communication with the array which receives the array signal and determines in real-time whether the ischemic myocardium of the heart is stunned or is nonviable. 4. An apparatus as described in claim 3 wherein the heart is nonviable if its relative permittivity is greater than about 19,000. 5. An apparatus for analyzing living tissue comprising: an electrode array having at least four electrodes for electrical communication with the tissue which produces an array signal;an admittance magnitude and phase detection circuit in communication with the array which receives the array signal from the array and produces a detection circuit signal corresponding to phase angle and magnitude of the array and signal; anda processor in communication with the detection circuit which receives the detection circuit signal and determines permittivity and conductivity of the tissue in real-time. 6. An apparatus as described in claim 5 wherein the admittance magnitude detection circuit is an analog phase angle and magnitude detection circuit. 7. An apparatus as described in claim 6 wherein the admittance magnitude and phase angle detection circuit and the processor together are portable. 8. A method for analyzing living tissue comprising the steps of: producing an array signal from an electrode array having at least four electrodes in electrical communication with the tissue;receiving the array signal from the array at an admittance magnitude and phase angle detection circuit in communication with the array;producing a detection circuit signal corresponding to phase angle and magnitude of the array signal with the phase angle and magnitude detection circuit; anddetermining permittivity and conductivity of the tissue in real-time with a processor in communication with the detection circuit which receives the detection circuit signal. | Board of Regents, The University of Texas System | A61B005/0408 A61B005/0402
US20110112389 | A1 | 2011-05-12 | Analyte Monitoring Device and Methods of Use | 2011-01-13 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for processing sensor data from a continuous analyte sensor, comprising: receiving sensor data from a continuous analyte sensor for a first time period, wherein the sensor data comprises at least one sensor data point;processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period; andproviding output, wherein providing output comprises displaying the target analyte value or range of target analyte values on a user interface. 3. The method of claim 2, wherein the target analyte value or range of target analyte values is based on an individual physiological pattern. 4. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining whether the target analyte value or range of target analyte values is based on a clinical acceptability. 5. The method of claim 2, wherein the continuous analyte sensor is a continuous glucose sensor. 6. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on at least one parameter that is customizable by the user. 7. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on an amount and/or timing of a medicament delivery and/or caloric consumption. 8. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on a host's physiological pattern. 9. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on the sensor data. 10. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on an analysis of the sensor data using a clinical, physiological, and/or statistical analysis. 11. The method of claim 10, wherein determining the target analyte value or range of target analyte values based at least in part on an analysis of the sensor data using a clinical, physiological, and/or statistical analysis includes determining the t... | Abbott Diabetes Care Inc. | A61B005/1473
US20110160547 | A1 | 2011-06-30 | WIRELESS TRANSMISSION DEVICE FOR PHYSIOLOGICAL INFORMATION | 2009-12-29 | A wireless transmission device for physiological information is carried personally by a user and is attached on skin to measure the user's physiological condition by touching the skin. Related physiological signals are continuously and automatically transmitted to other monitoring device through an automatic wireless transmission unit in the wireless transmission device. Therefore, the user's physical condition can be monitored at any time and the related physiological information can be provided to a medical facility or other monitoring facility for understanding and tracking. | 1. A wireless transmission device for physiological information being directly worn on a user and touching skin, comprising a measuring unit to detect a physiological condition of the user; an interior of the wireless transmission device being provided with an automatic wireless transmission unit which continuously and automatically transmit related physiological signals that have been measured, the automatic wireless transmission unit transmitting these physiological signals to a wireless transmitter which is provided with a receiving unit and a transmission unit, the receiving unit analyzing the received physiological signals, whereas the analyzed physiological signals being transmitted wirelessly to at least one monitoring device through the transmission unit. 2. The wireless transmission device for physiological information according to claim 1, wherein the receiving unit is a Bluetooth or a signal connection line. 3. The wireless transmission device for physiological information according to claim 1, wherein the monitoring device is a cell phone or a server. | - | A61B005/00 H04B007/00
US20110137208 | A1 | 2011-06-09 | DEVICE AND METHOD FOR AUTOMATICALLY SAMPLING AND MEASURING BLOOD ANALYTES | 2009-07-23 | The present invention includes a device and method used for automated sampling and measurement of blood analytes, such as blood glucose, from a subject patient. The blood sampling and measurement may be performed automatically by the device on a patient without requiring user intervention or supervision. The sampling and measuring device comprises a sensor unit operably attached to a replaceable cartridge. The replaceable cartridge provides a disposable body that contains a number of lancets and test strips used for collecting and analyzing a blood sample. The sensor unit contains an actuator and microcontroller that fires the lancet and collects blood sample data from the test strips. The blood sampling and measurement process may be initiated by an external controller or may be autonomously initiated in a specific interval. A further embodiment involves integration and use of the sampling and measurement device within an automated monitoring and treatment system. | 1. An automated device for sampling and measuring blood analytes, comprising: a sensor unit structured to be positioned on a patient body, the sensor unit including electronic circuitry, lancet firing means, and variable pressure control means; anda replaceable cartridge having a plurality of consumable products disposed therein, the consumable products including one or more lancets and one or more test strips for measuring a blood analyte of the patient;wherein the replaceable cartridge and the sensor unit are structured to temporarily mate with one another via an attachment means such that the replaceable cartridge is removable from the sensor unit;wherein the electronic circuitry enables automated blood extraction and analysis of a blood analyte from the patient body iteratively over time without need for manual intervention; andwherein the automated blood extraction and analysis is performed through electronically controlled use of the variable pressure control means, the lancet firing means, the one or more lancets, and the one or more blood test strips. 2. The automated device of claim 1, wherein the sensor unit comprises an upper portion and a lower portion operably connected thereto. 3. The automated device of claim 2, wherein the upper portion and the lower portion of the sensor unit is configured to be positioned around a finger of the patient, to obtain a blood analyte measurement from a measurement site at a tip of the finger of the patient. 4. The automated device of claim 1, wherein the sensor unit is adjustable during positioning of the sensor unit upon the patient body. 5. The automated device of claim 1, wherein the lancet firing means provided by the sensor unit comprises an actuator, a tension spring, and an actuator carriage enclosed within an actuator casing. 6. The automated device of claim 1, wherein the lancet firing means includes an actuator, and wherein depth of lancet actuation is adjustable for the one or more lancets individually or in combination. 7. The automated device of claim 1, wherein the automated device automatically applies anesthetic or analgesic solution to skin proximate to a measurement site of the patient prior to firing of the one or more lancets. 8. The automated device of claim 1, wherein the sensor unit further comprises an electronic contact used to read a test strip provided by the replaceable cartridge during automated blood analysis. 9. The automated device of claim 1, wherein the replaceable cartridge provides a plurality of reaction areas, each reaction area containing a lancet and a test strip deployed for obtaining a single blood sample. 10. The automated device of claim 9, further comprising a barrier structure comprising a combination of an absorbent material and a material non-permeable to blood arranged around each reaction area, and wherein the barrier structure separates the consumable products of the replaceable cartridge from the sensor unit. 11. The automated device of cl... | ADMETSYS CORPORATION | A61B005/151
US20110137157 | A1 | 2011-06-09 | IMAGE PROCESSING APPARATUS AND IMAGE PROCESSING METHOD | 2010-11-08 | This invention concerns the acquisition of diagnosis information data effective for diagnosing eye disease independently of the eye state. An image processing apparatus for processing a tomogram of an eye includes a unit configured to determine an eye feature based on the tomogram and thus determine the eye state, a unit configured to detect, from the tomogram, a detection target to be used to calculate diagnosis information data quantitatively representing the determined eye state, and a unit configured to calculate the diagnosis information data using position information of the detection target. In accordance with the eye state, the detection unit changes the detection target or an algorithm to be used to detect the detection target. | 1. An image processing apparatus for processing a tomogram of an eye, comprising: a determination unit configured to determine a state of a disease in the eye based on information of the tomogram; anda detection unit configured to change, in accordance with the state of the disease in the eye determined by said determination unit, one of a detection target to be used to calculate diagnosis information data quantitatively representing the state of the disease and an algorithm to be used to detect the detection target. 2. The apparatus according to claim 1, wherein the detection target includes a predetermined layer of the tomogram, andif a shape of the predetermined layer has changed, or the tomogram includes a predetermined tissue, said detection unit changes a detection parameter to be used to detect the predetermined layer included in the detection target, and then redetects the predetermined layer. 3. The apparatus according to claim 2, wherein presence/absence of the change of the shape of the predetermined layer includes presence/absence of distortion of retinal pigment epithelium layer of the eye, and presence/absence of the predetermined tissue includes one of presence/absence of a white spot and presence/absence of a cyst. 4. The apparatus according to claim 3, wherein said determination unit determinesa first state upon determining that the distortion of the retinal pigment epithelium layer of the eye does not exist, and neither the white spot nor the cyst exists,a second state upon determining that the distortion of the retinal pigment epithelium layer of the eye exists, or not the cyst but the white spot exists, anda third state upon determining the cyst exists, andsaid detection unit detectsan inner limiting membrane, a nerve fiber layer boundary, and a retinal pigment epithelium layer boundary as the detection target when said determination unit has determined the first state,the inner limiting membrane, the retinal pigment epithelium layer boundary, and a retinal pigment epithelium layer boundary assuming that the retinal pigment epithelium layer has no distortion as the detection target when said determination unit has determined the second state, andthe inner limiting membrane and the retinal pigment epithelium layer boundary as the detection target when said determination unit has determined the third state. 5. The apparatus according to claim 4, wherein when said determination unit has determined that the distortion of the retinal pigment epithelium layer of the eye exists, said detection unit changes a detection parameter to be used to detect the retinal pigment epithelium layer boundary in a region with the distortion, and then redetects the retinal pigment epithelium layer boundary,when said determination unit has determined that the white spot exists, said detection unit changes a detection parameter to be used to detect the retinal pigment epithelium layer boundary located at a deep position in a depth direction relativ... | CANON KABUSHIKI KAISHA | A61B005/05
US20110160614 | A1 | 2011-06-30 | BLOOD TEST APPARATUS | 2011-03-09 | A blood sampling cartridge is attachable to and removable from an attaching part of a blood test apparatus that includes a housing. The attaching part is provided in one side of the housing. The blood sampling cartridge includes a holder, a blood sensor that is attachable to one end of the holder, a lancet that is movable inside the holder, a blood collection needle attached to one end of the lancet, and a guide provided on the lancet. The guide of the lancet is configured to latch with the holder after puncturing with the blood collection needle, such that the blood collection needle is prevented from projecting beyond the blood sensor. | 1. A blood sampling cartridge that is attachable to and removable from an attaching part of a blood test apparatus having a housing, the attaching part being provided in one side of the housing, the blood sampling cartridge comprising: a holder;a blood sensor that is attachable to one end of the holder;a lancet that is movable inside the holder;a blood collection needle attached to one end of the lancet; anda guide provided on the lancet,wherein the guide of the lancet is configured to latch with the holder after puncturing with the blood collection needle, such that the blood collection needle is prevented from projecting beyond the blood sensor. 2. The blood sampling cartridge according to claim 1, wherein: the holder has a hole,the guide is provided on a side surface of the lancet, andthe guide is configured to latch with the hole of the holder. 3. The blood sampling cartridge according to claim 1, wherein: the holder has a cylindrical shape, and has a plurality of holes in an outer surface of the holder,the lancet has a rod shape,the guide of the lancet comprises a plurality of guides provided on an outer surface of the lancet,the plurality of guides are configured to latch with the plurality of holes, respectively. 4. The blood sampling cartridge according to claim 3, wherein the plurality of guides face each other. 5. The blood sampling cartridge according to claim 1, wherein: the holder has a convex part,the guide is provided on a side surface of the lancet, andthe guide is configured to latch with the convex part of the holder. 6. The blood sampling cartridge according to claim 1, wherein: the blood collection needle is accommodated inside the holder above the blood sensor, andthe blood collection needle is configured to vertically pass through a surface of the blood sensor upon puncturing. 7. The blood sampling cartridge according to claim 1, wherein: a front part of the holder of the blood sampling cartridge, with respect to an insertion direction of the attaching part, is thinner than a rear part of the holder. 8. The blood sampling cartridge according to claim 1, wherein: the blood collection needle and the lancet of the blood sampling cartridge are integrated. 9. The blood sampling cartridge according to claim 1, wherein: the blood sampling cartridge is dividable into the integrated blood collection needle and lancet, and the holder to which the blood sensor is attached. 10. The blood sampling cartridge according to claim 1, further comprising a second holder that covers a surface of the blood sensor, the surface facing outwardly from the holder, wherein the second holder includes a circular concave part of 4 to 15 mm in diameter on the surface facing outwardly from the holder. 11. The blood sampling cartridge according to claim 1, further comprising a protective cap which covers the blood collection needle. 12. The blood sampling cartridge according to claim 10, further comprising a protective cap which covers the second ... | PANASONIC CORPORATION | A61B005/153 A61B005/151
US20110130666 | A1 | 2011-06-02 | ENHANCED REPORTING OF PATHOLOGICAL EPISODES | 2010-11-01 | An apparatus comprises a sensor circuit configured to produce a time-varying physiologic sensor signal of a subject and a pathology detection circuit communicatively coupled to the sensor. The pathology detection circuit is configured to detect a first pathological episode using the sensed physiologic sensor signal, deem that the first pathological episode has ended, detect at least one second pathological episode using the sensed physiologic sensor signal, and indicate the first and second pathological episodes as one pathological episode if the first and second episode are detected within a specified time interval. | 1. An apparatus comprising: a sensor circuit configured to produce a time-varying physiologic sensor signal of a subject; anda pathology detection circuit communicatively coupled to the sensor and configured to: detect a first pathological episode using the physiologic sensor signal;deem that the first pathological episode has ended;detect at least one second pathological episode using the physiologic sensor signal; andindicate the first and second pathological episodes as one pathological episode if the first and second episodes are detected within a specified time interval. 2. The apparatus of claim 1: wherein the sensor circuit includes an implantable cardiac signal sensing circuit configured to produce a sensed cardiac signal representative of cardiac activity of a subject; andwherein the pathology detection circuit is configured to: detect a first episode of arrhythmia using the sensed cardiac signal;deem that the first episode of arrhythmia has ended;detect at least one second episode of arrhythmia using the sensed cardiac signal; andindicate the first and second episodes of arrhythmia as one episode of arrhythmia if the first and second episodes are detected within the specified time interval. 3. The apparatus of claim 2, wherein the pathology detection circuit is configured to adjust at least one arrhythmia detection parameter when one or more arrhythmia episodes are detected within the specified time interval of a previous arrhythmia episode. 4. The apparatus of claim 3, wherein, when multiple tachyarrhythmia episodes are detected within the specified time interval, the pathology detection circuit is configured to, at least one of: reduce an atrial tachyarrhythmia response entry count;increase an atrial tachyarrhythmia response exit count; orreduce a tachyarrhythmia detection rate zone threshold. 5. The apparatus of claim 2, wherein the pathology detection circuit is configured to: determine whether the first and second episodes are one episode of arrhythmia by applying an additional detection criterion to the first and second arrhythmia episodes when the second episode is detected within the specified time interval; andindicate the first and second episodes of arrhythmia as one episode of arrhythmia according to the additional detection criterion. 6. The apparatus of claim 5, wherein the pathology detection circuit is configured to obtain, as the additional detection criteria, at least one of: similarity of heart rate among detected episodes;an assessment of heart rate stability; ora morphology analysis of the detected arrhythmia. 7. The apparatus of claim 5, including: a second implantable sensor circuit configured to produce a second sensor signal representative of hemodynamic function of the heart, andwherein the pathology detection circuit is configured to determine whether the first and second episodes are one episode of arrhythmia according to the second sensor signal. 8. The apparatus of claim 1, wherein the sensor circuit... | - | A61B005/0205 A61B005/00 A61B005/0402 A61B005/0464
US20110160565 | A1 | 2011-06-30 | DETECTING PROXIMITY TO MRI SCANNER | 2010-11-30 | A plurality of separate indications of a scanner field can be received using a corresponding plurality of scanner field sensors of an implantable medical device (IMD). In an example, the IMD can be switched from a first therapy mode to a second therapy mode using one or more of the plurality of scanner field sensors, and from the second therapy mode back to the first therapy mode using each of the plurality of scanner field sensors. In certain examples, shock therapy can be terminated or inhibited using a detected proximity of the IMD to a magnetic resonance imaging (MRI) scanner, or antitachycardia pacing (ATP) can be terminated or inhibited using a detected active scan of the MRI scanner. | 1. A system comprising: an implantable medical device (IMD) configured to be implanted into a subject, the IMD including a first therapy mode and a second therapy mode, the second therapy mode different than the first therapy mode;a plurality of scanner field sensors configured to receive separate indications of a scanner field; andwherein the IMD is configured to receive each of the separate indications of the scanner field, to switch from the first therapy mode to the second therapy mode using one or more of the plurality of scanner field sensors, and to switch from the second therapy mode to the first therapy mode using each of the plurality of scanner field sensors. 2. The system of claim 1, wherein the scanner field includes a static magnetic or time-varying magnetic or electric field associated with a nuclear magnetic resonance (NMR) device. 3. The system of claim 1, wherein the IMD is configured to switch from the first therapy mode to the second therapy mode if any one or more of the plurality of scanner field sensors receives an indication of a scanner field. 4. The system of claim 3, wherein the IMD is configured to switch from the second therapy mode to the first therapy mode only if each of the plurality of scanner field sensors receives an indication of no scanner field. 5. The system of claim 1, wherein the first therapy mode includes normal operation of the IMD, and wherein the second therapy mode includes at least one of a terminated or inhibited shock therapy or a terminated or inhibited antitachycardia pacing (ATP). 6. The system of claim 1, wherein the plurality of scanner field sensors includes at least one static magnetic or electric field sensor and at least one active magnetic or electric field sensor. 7. The system of claim 6, wherein the plurality of scanner field sensors includes: a first scanner field sensor configured to detect an indication of a magnetic field, the first scanner field sensor including an inductor having a ferromagnetic core and a core saturation detector configured to detect saturation of the inductor and the ferromagnetic core in a strong magnetic field;a second scanner field sensor configured to detect an indication of a magnetic field, the second scanner field sensor including a Hall effect sensor; anda third scanner field sensor configured to detect an indication of a magnetic or electric field, the third scanner field sensor including an electromagnetic interference (EMI) sensor configured to detect EMI from an active magnetic resonance imaging (MRI) scan. 8. The system of claim 1, wherein at least one of the plurality of scanner field sensors are configured to detect a proximity to an MRI scanner; and wherein the IMD is configured to terminate or inhibit shock therapy using the detected proximity to the MRI scanner. 9. The system of claim 1, wherein at least one of the plurality of scanner field sensors are configured to detect an active scan of the MRI scanner; and wherein the IMD is co... | - | A61N001/08 A61N001/37 A61B005/055
US20110137148 | A1 | 2011-06-09 | Method and device comprising fused ultrasound and magnetic resonance imaging | 2010-11-15 | A device and method for concurrently obtaining ultrasound and MRI image data are disclosed. A probe member comprises ultrasound imaging means at a first predetermined position on the probe member and an MRI coil at a second predetermined position on the probe member. An MRI imaging system having an MRI pulse sequence comprising an MRI image data acquisition period and an MRI pulse sequence time gap generates an MRI image data set. An ultrasound image data set is acquired only during the MRI pulse sequence time gap in which the operation of the ultrasound imaging means does not interfere with the MRI imaging operation. | 1. A method for imaging a region comprising: providing a probe member comprising ultrasound imaging means,positioning the probe member a predetermined distance from the region to be imaged,positioning the probe member and region within an MRI system,imaging the region with the MRI system to define an MRI image data set utilizing a predetermined MRI pulse sequence,the MRI pulse sequence comprising a TR period comprising an MRI image data acquisition period followed by an MRI pulse sequence time gap,imaging the region with the ultrasound transceiver within the MRI pulse sequence time gap to define an ultrasound image data set. 2. The method of claim 1 further comprising the step of co-registering the MRI data set and ultrasound image data set to provide a fused image data set. 3. The method of claim 1 wherein the MRI system further comprises a receiving MRI coil disposed on the probe member. 4. The method of claim 1 wherein the MRI pulse sequence comprises a spin echo MRI pulse sequence. 5. The method of claim 3 wherein the MRI system further comprises an endorectal MRI coil disposed on the probe member. 6. The method of claim 3 wherein the MRI pulse sequence comprises a spin echo MRI pulse sequence. 7. An apparatus for imaging a region comprising: An ultrasound system comprising a probe member comprising an ultrasound transceiver,an MRI system utilizing a predetermined MRI pulse sequence,the MRI pulse sequence comprising a TR period comprising an MRI image data acquisition period followed by an MRI pulse sequence time gap to define an MRI image data set,wherein the ultrasound system is configured to image the region within the MRI pulse sequence time gap to define an ultrasound image data set. 8. The apparatus of claim 7 further comprising co-registration means for co-registering the MRI data set and the ultrasound image data set to provide a fused image data set. 9. The apparatus of claim 7 wherein the MRI system further comprises a receiving MRI coil disposed on the probe member. 10. The apparatus of claim 9 wherein the MRI pulse sequence comprises a spin echo MRI pulse sequence. 11. A probe member comprising ultrasound imaging means at a first predetermined position on the probe member and an MRI coil at a second predetermined position on the probe member. | Irvine Sensors Corporation | A61B005/055
US20110137211 | A1 | 2011-06-09 | OPTICAL DEVICE AND TRAINING METHOD FOR CORRECTING POSTURE | 2010-12-02 | A device and a method whereby a computer user seated at a first predetermined position can be provided with a continuous visual indication of the user's spinal posture with respect to the computer monitor. A device is provided which shows either a first indication or a second indication, as a function of the angle at which it is viewed by the user. The first indication is indicative of a viewing angle with a most preferred spinal posture, and the second indication is indicative of a viewing angle with a less preferred spinal posture. The device is placed in a second predetermined position on the computer monitor, so that the user will see either the first or the second indication, depending on whether the user is currently seated with the most preferred or the less preferred spinal posture. Alternatively, an indicator or an appearance exhibited by the device may depend on the angle at which it is viewed by the viewer, rather than the user's spinal posture per se. | 1. A method of providing a user with a visual indication of a viewing angle defined between the user's eyes and a target being viewed, comprising the steps of: providing a device whose appearance changes as a function of the angle at which it is viewed by a user; andinstructing the user to place the device on the viewing target. 2. The method of claim 1, wherein said device shows a first indicator when viewed at a first angle and a second indicator when viewed at a second angle. 3. The method of claim 1, wherein said device includes a lenticular lens. 4. The method of claim 2, wherein said first indicator indicates that the viewer is viewing the target at a most preferred viewing angle, and said second indicator indicates that the viewer is viewing the target at a less preferred viewing angle. 5. The method of claim 4, wherein said first and second indicators include different respective colors. 6. The method of claim 5, wherein said first and second indicators include different respective designs. 7. The method of claim 4, wherein said first and second indicators include different respective designs. 8. The method of claim 4, wherein said device includes a lenticular lens. 9. The method of claim 4, wherein said device shows at least one additional indicator that indicates that the target is being viewed at a least preferred viewing angle. 10. The method of claim 4, wherein said first and second indicators correspond respectively to a higher position and a lower position of the user's eyes. 11. A method of viewing a target with a visual indication of a viewing angle defined between the viewer's eyes and the target being viewed, comprising the steps of: obtaining a device whose appearance changes as a function of the angle at which it is viewed by the viewer; andplacing the device on the viewing target. 12. A method of providing a computer user seated at a first predetermined position with a continuous visual indication of the user's spinal posture with respect to the computer monitor, comprising the steps of: providing a device having either a first appearance or a second appearance as a function of the angle at which it is viewed by the user, the first appearance being indicative of a viewing angle with a most preferred spinal posture and the second appearance being indicative of a viewing angle with a less preferred spinal posture; andinstructing the user to place the device on the computer monitor in a second predetermined position so as to show the first or the second appearance to the user in response to the user's spinal posture. 13. A method for a computer user seated at a first predetermined position to view a computer monitor with a continuous visual indication of the user's spinal posture with respect to the computer monitor, comprising the steps of: obtaining a device having either a first appearance or a second appearance as a function of the angle at which it is viewed by the user, the first appearance being indicative of... | - | A61B005/103
US20110124974 | A1 | 2011-05-26 | MOBILE MEDICAL COMMUNICATIONS SYSTEM | 2009-11-24 | A mobile medical communications system, having a processor, at least one video capture device, an audio emitting device, an audio receiving device, and a communications device, the system being communicatively coupled to a remote terminal via the communications device. In one embodiment the system may include a patient scanning device, a vital sign measuring and recording device and a global positioning unit. | 1. A mobile medical communications system, comprising: a processor;at least one video capture device;an audio emitting device;an audio receiving device; anda communications device, that communicatively couples the system to a remote terminal. 2. The mobile medical communications system of claim 1, further comprising: a patient scanning device;a vital sign measuring and recording device; anda global positioning unit. 3. The mobile medical communications system of claim 1, wherein said remote terminal comprises: a processor;at least one video display screen;an audio emitting device;an audio receiving device; anda communications device. 4. The mobile medical communications system of claim 1, wherein said mobile medical communications system is integrated into an emergency vehicle. 5. The mobile medical communications system of claim 1, wherein said mobile medical communications system is portable. 6. A method of providing medical treatment to a patient, comprising: establishing a communications connection between a mobile medical communications system and a remote terminal;focusing at least one video capture device on an area of the patient's body in need of treatment;displaying video data captured by said at least one video capture device on at least one video display screen of said remote terminal such that said video data is displayed in real time; andestablishing audio communication between said mobile medical communications system and said remote terminal such that said audio communication is carried out in real time. 7. The method of claim 6, further comprising: analyzing the patient using a patient scanning device;communicating the obtained data to a remote terminal; anddisplaying said data on at least one video display screen of said remote terminal. 8. The method of claim 6, further comprising: measuring the patient's vital signs using a vital sign measuring and recording device;communicating the obtained data to a remote terminal; anddisplaying said data on a vital sign display of a remote terminal. 9. The method of claim 6, further comprising: determining the location of an emergency vehicle using a global positioning unit;communicating the obtained data to a remote terminal; anddisplaying the location of the emergency vehicle on a mapping display of the remote terminal. | - | A61B005/00 H04N007/18
US20110151310 | A1 | 2011-06-23 | HIGH ENERGY DENSITY BATTERY FOR USE IN IMPLANTABLE MEDICAL DEVICES AND METHODS OF MANUFACTURE | 2009-08-07 | A high energy density battery is provided that improves energy density through efficient placement of the inter-plate connections within the battery enclosure. The placement of the current carrying leads in the high energy density battery allows for a greater volume of active material to be placed within the battery enclosure. This placement design can also be used to reduce the size of existing power sources. Methods for constructing high energy density batteries and methods for increasing the volumetric energy density of an implantable battery are also provided. The resulting high energy density battery can be used to power electronics associated with a variety of devices such as medical devices. | 1. A high energy density battery comprising: a case comprising an outer surface, an inner surface and a case opening;a header assembly inserted in the case opening, the header assembly comprising: an electrical feed-through, the electrical feed-through comprising a terminal pin,a glass-to-metal seal, anda surrounding sidewall extending through the case opening to the outer surface and the inner surface of the case, the case opening being sized to receive the header assembly, with the header assembly surrounding sidewall contacting the case opening;an electrode stack, the electrode stack comprising: an anode layer, the anode layer comprising electrically conductive, chemically active material, an upper surface, and a lower surface,an anode lead, the anode lead comprising electrically conductive material, an anode lead origin and an anode lead end, wherein:the anode lead origin is connected to the anode layer, andthe anode lead end extends into the case and attaches to the terminal pin:an insulative separator layer, the insulative separator layer comprising an upper surface and a lower surface;a cathode layer, the cathode layer comprising electrically conductive, chemically active material, an upper surface and a lower surface;a cathode lead, the cathode lead comprising electrically conductive material, a cathode lead origin and a cathode lead end, wherein:the cathode lead origin is connected to the cathode layer, andthe cathode lead end extends into the case and attaches to the positive terminal;a current collecting lead, the current collecting lead disposed between a cathode tab and the terminal pin, wherein the current collecting lead:is electrically connected across the cathode layer of the electrode stack, andis insulated with an insulative material; andan electrolyte solution, the electrolyte solution disposed within the case and contacting the electrode stack. 2. The high energy density battery of claim 1 that is a primary battery. 3. The high energy density battery of claim 1 that is a secondary battery. 4. The high energy density battery of claim 1 wherein the battery case comprises a material selected from the group consisting of nickel, stainless steel, aluminum, titanium, glass and ceramic. 5. The high energy density battery of claim 1 wherein the battery ease is a deep-drawn battery case. 6. The high energy density battery of claim 1 wherein the battery case is a multi-part or clam-shell case. 7. The high energy density battery of claim 1 wherein the battery case is a liner or insulating bag. 8. The high energy density battery of claim 1 comprising a plurality of electrode stacks, wherein the anode layer of one or more electrode stacks of the plurality is connected in series or in parallel to at least one other anode layer of an electrode stack of the plurality, and the cathode layer of one or more electrode stacks of the plurality is connected in series or in parallel to at least one other cathode layer of an electrode stack of... | Balan Biomedical, Inc. | H01M006/42 H01M002/30 H01M006/00 A61B005/00 A61B001/00
US08795957 | B2 | 2014-08-05 | Blood collection container | 2011-02-23 | The present invention provides a blood collection container capable of allowing blood to coagulate in a short time and preventing blood from coagulating with bubbles contained therein and thus preventing formation of bubbly clots when used to collect blood therein. A blood collection container 1 stores a blood coagulation promoting agent 4 for promoting blood coagulation and an antifoaming agent 5, which is a polyoxyalkylene or a derivative thereof. The amount of the antifoaming agent 5 is 2.0×10−3 to 0.2 mg per mL of blood to be collected in the blood collection container 1. | 1. A blood collection container comprising: a tubular container having an opening at one end and a bottom at the other end opposite to the one end anda blood coagulation promoting agent (A) being a hydrolase capable of hydrolyzing the bond between arginine and any amino acid residue and/or the bond between lysine and any amino acid residue in peptide chains; andan anti foaming agent,said antifoaming agent being a polyoxyalkylene or a derivative thereof,wherein the amount of said antifoaming agent is 2.0×10−3 to 0.2 mg per mL of blood to be collected in the blood collection container, and a blood coagulation promoting agent-containing composition comprising said blood coagulation promoting agent (A) and an antifoaming agent-containing composition comprising said antifoaming agent are housed in said tubular container so that a lowermost portion of said blood coagulation promoting agent-containing composition with respect to said bottom of the tubular container is located nearer to the opening end as compared to a lowermost portion of said antifoaming agent-containing composition with respect to said bottom of the tubular container. 2. The blood collection container according to claim 1, further comprising a water-soluble binder. 3. The blood collection container according to claim 2, wherein said water-soluble binder is at least one selected from the group consisting of polyvinyl alcohol, polyvinyl pyrrolidone, acrylic acid copolymers, and polyoxyalkylene block copolymers. 4. The blood collection container according to claim 1, further comprising a tubular container having an opening at one end and a plug sealingly attached to said opening. 5. The blood collection container according to claim 4, wherein said antifoaming agent is deposited substantially on an entire inner surface of said tubular container. 6. The blood collection container according to claim 1, wherein said hydrolase is serine protease. 7. The blood collection container according to claim 6, wherein said serine protease is applied to the inner surface of the blood collection container by spray coating. 8. The blood collection container according to claim 1, further comprising a blood coagulation promoting agent (B) which comprises an adsorbent inorganic substance. 9. The blood collection container according to claim 8, wherein said blood coagulation promoting agent (B) is deposited substantially on the entire inner surface of said tubular container. 10. The blood collection container according to claim 1, which is a vacuum blood collection tube, the interior of which is at a reduced pressure. 11. The blood collection container according to claim 8, wherein said antifoaming agent-containing composition further comprises said blood coagulation promoting agent (B). 12. The blood collection container according to claim 2, which is a vacuum blood collection tube, the interior of which is at a reduced pressure. 13. The blood collection container according to claim 3, which is ... | Sekisui Medical Co., Ltd. | A01N001/02
US08543191 | B2 | 2013-09-24 | Method to acquire measurement data of a breathing examination subject by magnetic resonance technology, and associated computer program | 2010-11-24 | A method for the acquisition of measurement data of a breathing examination subject by magnetic resonance includes the following steps: (a) detect the physiological breathing signal of the examination subject with a breathing signal detection unit, (b) evaluate the detected breathing signal in an evaluation unit, (c) based on the evaluated breathing signal, calculate in a computer at least one parameter affecting the type of acquisition of measurement data by means of magnetic resonance, (d) detect a current physiological breathing signal with the breathing signal detection unit, (e) compare the last detected breathing signals with at least one trigger condition, (f) initiate the acquisition of measurement data using the calculated parameter from step (c) upon satisfaction of the trigger conditions from step (e), (g) repeat the steps (d) through (f) until all desired measurement data have been acquired, and (h) store and/or process the acquired measurement data in a memory and/or processing unit. After the evaluation of the detected breathing signal, at least one parameter of a following acquisition of measurement data is thus determined automatically without an input by an operator of the MR apparatus in use being required. | 1. A method for acquiring magnetic resonance measurement data from a breathing examination subject, comprising the steps of: (a) detecting a physiological breathing signal of an examination subject using a breathing signal detection unit, to obtain a detected breathing signal;(b) in a computerized evaluation unit, automatically evaluating the detected breathing signal to obtain an evaluation result;(c) in a computerized processor, automatically calculating, based on said evaluation result, at least one parameter for operating a magnetic resonance data acquisition unit that has an effect on acquisition of magnetic resonance data from the examination subject by the magnetic resonance data acquisition unit;(d) detecting a current physiological breathing signal from the examination subject with said breathing signal detection unit;(e) in a computerized comparison unit, comparing the last physiological breathing signals that were detected with said breathing signal detection unit with at least one trigger condition, to obtain a comparison result;(f) from said comparison unit, automatically initiating acquisition of said magnetic resonance data from the examination subject by said magnetic resonance data acquisition unit using said at least one parameter calculated in step (c), upon said comparison result indicating satisfaction of said at least one trigger condition by said last physiological breathing signals; and(g) repeating steps (d) through (f) until all desired magnetic resonance data have been acquired. 2. A method as claimed in claim 1 wherein step (c) comprises calculating at least one parameter that affects or determines a duration of acquisition of said magnetic resonance data after triggering said acquisition of said magnetic resonance measurement data. 3. A method as claimed in claim 2 wherein step (c) comprises calculating said at least one parameter depends on a percentile proportion of a duration of a cycle of the detected breathing signal, and statistically determining said duration of said cycle as said evaluation result in step (b). 4. A method as claimed in claim 3 wherein step (c) comprises calculating said at least one parameter as the maximum number of slices of the examination subject from which said magnetic resonance data are acquired after initiating said acquisition of said magnetic resonance measurement data in step (f). 5. A method as claimed in claim 4 comprising calculating said maximum number of slices comprises comparing a duration required for acquisition of magnetic resonance data for each slice of the examination subject with said statistical duration of said breathing cycle. 6. A method as claimed in claim 2 comprising calculating said at least one parameter as a number of echoes of an echo train used by said magnetic resonance data acquisition unit to acquire said magnetic resonance measurement data. 7. A method as claimed in claim 1 comprising calculating said at least one parameter dependent on a predeter... | Siemens Aktiengesellschaft | A61B005/05
US20110118563 | A1 | 2011-05-19 | Portable multi-Parameter physiological monitoring and recording device | 2010-01-08 | A portable multi-parameter physiological monitoring and recording device includes multiple sensors, multiple detection channel units, a multiplexer, a clock generator, a controller, an analog switch, a recording unit and a storage unit. The sensors are attached on the body of a person for sensing physiological parameter signals. The detection channel units amplify the received physiological parameter signals. The multiplexer captures the amplified physiological parameter signals. The clock generator generates clock signals in the audio-frequency range. The controller controls the multiplexer to capture the amplified physiological parameter signals based on the clock signals generated by the clock generator. The analog switch outputs the amplified physiological parameter signals as an audio signal. The recording unit records the audio signals as an audio file. The storage unit stores the audio file. | 1. A portable multi-parameter physiological monitoring and recording device, which is used for measuring multiple physiological signals of a person. These measured results can be outputted as an audio signal and saved as an audio file, comprising: Several sensors attached on specified locations on the body skin of a person for sensing the corresponding physiological signals respectively;multiple-detection channel units respectively connected to the corresponding sensors to amplify the received physiological signals;a multiplexer connected to the detection channel units for capturing the amplified physiological parameter signals;a clock generator for generating clock signal in the audio-frequency range;a controller connected to the multiplexer and the clock generator for controlling the multiplexer to capture the amplified physiological signals based on the clock signals generated by the clock generator;an analog switch connected to the detection channel units and the multiplexer for outputting the amplified physiological signals as an audio signal;a recording unit connected to the analog switch and the controller for recording the audio signal as an audio file under control of the controller; anda storage unit connected to the recording unit for storing the audio file. 2. The portable multi-parameter physiological monitoring and recording device as claimed in claim 1, further comprising a connection interface connecting to the analog switch for outputting the audio signals. 3. The portable multi-parameter physiological monitoring and recording device as claimed in claim 2, wherein the connection interface is connected to an external storage device through cable or wireless transmission for storing the audio signal outputted from the connection interface as the audio file. 4. The portable multi-parameter physiological monitoring and recording device as claimed in claim 3, wherein the external storage device is provided to process the received audio signal to proceed analysis or determination instantly. 5. The portable multi-parameter physiological monitoring and recording device as claimed in claim 3, wherein the external storage device is a computer device, a recording device or a personal digital assistant. 6. The portable multi-parameter physiological monitoring and recording device as claimed in claim 3, wherein the external storage device is connected to the connection interface to transfer the audio signals through audio cables. 7. The portable multi-parameter physiological monitoring and recording device as claimed in claim 3, wherein the external storage device is connected to the connection interface to transfer the audio signals through Bluetooth wireless transmission. 8. The portable multi-parameter physiological monitoring and recording device as claimed in claim 1, further comprising a display panel connected to the controller for displaying the amplified physiological signals captured by the multiplexer. 9. The portable multi... | Tatung Company | A61B005/00
US08764669 | B2 | 2014-07-01 | Blood pressure information display device, blood pressure information display system, blood pressure information display method, and recording medium recorded with blood pressure information display program | 2011-01-31 | A blood pressure information display device for calculating a blood pressure value based on a pulse wave amplitude through the oscillometric method and the like, and displaying the calculated blood pressure value as blood pressure information performs a process of changing a magnitude of a specified pulse wave amplitude of a plurality of extracted pulse wave amplitudes according to the operation by a user (medical staff) when a predetermined instruction is input. A predetermined algorithm is applied to the plurality of pulse wave amplitudes reflecting the change to calculate a reference blood pressure value, and the calculated reference blood pressure value is displayed as the blood pressure information. | 1. A blood pressure information display device for calculating a blood pressure value, and displaying the calculated blood pressure value as blood pressure information, the blood pressure information display device comprising: a display unit;an operation unit for accepting an operation by a user;a cuff to be wrapped around a predetermined measurement site of a person to be measured;an adjustment unit for adjusting an inner pressure of the cuff;a pressure detection unit for detecting a cuff pressure signal representing the pressure in the cuff;an acquiring unit for acquiring the cuff pressure signal in time-series from the pressure detection unit when a drive control of the adjustment unit is being carried out;an extracting unit for extracting a pulse wave amplitude for every beat based on the cuff pressure signal;a display processing unit for displaying pulse wave amplitude information representing a plurality of extracted pulse wave amplitudes on the display unit;a change processing unit for performing a process of changing a magnitude of a specified pulse wave amplitude of the plurality of extracted pulse wave amplitudes according to an instruction from the operation unit when a predetermined instruction is input from the operation unit; anda calculating unit for calculating a reference blood pressure value by applying a predetermined algorithm to a plurality of pulse wave amplitudes reflecting the change on the specified pulse wave amplitude; whereinthe display processing unit displays the reference blood pressure value on the display unit as the blood pressure information. 2. The blood pressure information display device according to claim 1, further comprising a storage unit for storing the plurality of extracted pulse wave amplitudes in association with at least one of a cuff pressure or a time, and for storing an amplitude value subsequent to the specified pulse wave amplitude being changed in association with a cuff pressure or a time at which the specified pulse wave amplitude is detected. 3. The blood pressure information display device according to claim 2, wherein the display processing unit displays the plurality of extracted pulse wave amplitudes along a cuff pressure axis or a time axis as the pulse wave amplitude information. 4. The blood pressure information display device according to claim 3, wherein the change processing unit changes a position of the specified pulse wave amplitude on the cuff pressure axis or the time axis according to the instruction from the operation unit. 5. The blood pressure information display device according to claim 3, wherein the change processing unit adds a pulse wave amplitude having a magnitude corresponding to an instruction to a specified position on the cuff pressure axis or the time axis according to the instruction from the operation unit. 6. The blood pressure information display device according to claim 1, wherein the change processing unit accepts the specification of change for o... | Omron Healthcare Co., Ltd. | A61B005/02
US08394037 | B2 | 2013-03-12 | Systems and devices for evaluating the integrity of a uterine cavity | 2009-11-11 | Methods, systems and devices for evaluating the integrity of a uterine cavity. A method comprises introducing transcervically a probe into a patient's uterine cavity, providing a flow of a fluid (e.g., CO2) through the probe into the uterine cavity and monitoring the rate of the flow to characterize the uterine cavity as perforated or non-perforated based on a change in the flow rate. If the flow rate drops to zero or close to zero, this indicates that the uterine cavity is intact and not perforated. If the flow rate does not drop to zero or close to zero, this indicates that a fluid flow is leaking through a perforation in the uterine cavity into the uterine cavity or escaping around an occlusion balloon that occludes the cervical canal. | 1. A system for characterizing a patient uterus, comprising: a probe having a distal end;an expandable member at the distal end of the probe, the expandable member having an expansible-collapsible frame which supports a fluid-tight wall and being adapted to occupy a uterine cavity when expanded;a source of a pressurized flow of a fluid;an integrity test lumen located in the probe, connected to the source, and configured to deliver the fluid from the source to a region surrounding an exterior of the fluid-tight wall of the expandable member when the expansible-collapsible frame is expanded in the uterine cavity of a patient;a flow sensor for measuring a flow rate of the fluid from the lumen into the region surrounding the fluid-tight wall of the expandable member when expanded in the uterine cavity; anda controller operatively coupled to the flow sensor and to the source of the pressurized fluid flow, said controller being configured to (1) initiate a flow of fluid from the fluid source through the lumen into the uterus to the region surrounding the exterior of the fluid-tight wall of the expandable member after the expandable-collapsible frame has been expanded to cause the expandable member to occupy the uterine cavity and before neutral gas is delivered to an interior of the expandable member to enable ablation, and (2) measure the flow rate of the flow of fluid into the uterus region surrounding the exterior of the fluid-tight wall of the expandable member using the flow sensor, and (3) determine if the measured flow rate of the fluid flow into the region drops below a predetermined minimum threshold rate within a predetermined period of lime in the range from 1 second to 60 seconds, wherein the controller generates a signal indicating that the uterus is perforated if the flow rate fails to fall below said predetermined minimum threshold rate within said predetermined time period. 2. The system of claim 1, wherein the signal is at least one of visual, aural and tactile. 3. The system of claim 1, wherein the flow sensor is a mass flow sensor. 4. The system of claim 1, wherein the controller generates a signal indicating that the uterus is not perforated if the flow rate drops below said predetermined minimum threshold rate within said predetermined period of time. 5. The system of claim 4, wherein the signal is at least one of visual, aural and tactile. 6. The system of claim 1, wherein the expandable member comprises a thin-wall dielectric which defines an energy-delivery surface for ablating uterine cavity tissue. 7. The system of claim 6, wherein the energy-delivery surface is configured to deliver RF energy to the tissue via capacitive coupling. 8. The system of claim 6, wherein the energy-delivery surface includes at least one electrode coupled to an RF generator. 9. The system of claim 6, wherein the controller is configured to disable activation of the energy-delivery surface if the flow rate fails to fall below said predetermine... | Minerva Surgical, Inc. | A61B005/103 A61B005/117
US08945090 | B2 | 2015-02-03 | Implantable radiopaque marking | 2010-11-23 | Specified information can be provided to a user, to a machine, or to an automated process using a specified configuration of a plurality of radiopaque rings about at least a portion of an implantable lead. | 1. A system comprising: an implantable lead including: a plurality of radiopaque rings arranged in a specified configuration about a portion of the implantable lead; andwherein the specified configuration of the plurality of radiopaque rings is configured to provide specified information to a user, to a machine, or to an automated process; andwherein the specified configuration is selected from a first specified configuration of the plurality of radiopaque rings that is configured to provide an indication that the implantable lead is MR compatible to a first magnetic field strength, and a second specified configuration that is configured to provide an indication that the implantable lead is MR compatible to a second magnetic field strength, the second specified configuration different than the first specified configuration and the second magnetic field strength different than the first magnetic field strength. 2. The system of claim 1, wherein each of the plurality of radiopaque rings have different inner diameters. 3. The system of claim 1, wherein the implantable lead includes a rear seal configured to be at least partially inserted into a lead port of an implantable medical device (IMD) header; and wherein the plurality of radiopaque rings are configured to be injection molded into the rear seal. 4. The system of claim 1, wherein the specified configuration of the plurality of radiopaque rings includes the number of radiopaque rings, and wherein different numbers of radiopaque rings are configured to provide different information about the MR compatibility of the implantable lead. 5. The system of claim 1, wherein the specified configuration of the plurality of radiopaque rings includes a radiopaque ring shape. 6. The system of claim 5, wherein different radiopaque ring shapes are configured to provide different information about the implantable lead itself or about an implantable medical device coupled to the implantable lead. 7. The system of claim 1, wherein the first specified configuration includes a first number of radiopaque rings and the second specified configuration includes a second number of radiopaque rings, the second number different than the first number. 8. A system comprising: an implantable lead including: a rear seal configured to be at least inserted into a lead port of an implantable medical device (IMD) header; anda plurality of radiopaque rings each having different inner diameters arranged in a specified configuration about a portion of the implantable lead, wherein the plurality of radiopaque rings are configured to be injection molded into the rear seal; andwherein the specified configuration of the plurality of radiopaque rings provides an indication of a level of magnetic resonance (MR) compatibility of the implantable lead;wherein a first specified configuration is configured to provide an indication that the implantable lead is MR Conditional to a first magnetic field strength and a second specified config... | Cardiac Pacemakers, Inc. | A61M025/098 A61N001/05 A61B019/00 A61B005/04 A61N001/08
US08386042 | B2 | 2013-02-26 | Omnidirectional accelerometer device and medical device incorporating same | 2009-11-03 | A portable medical device is provided with an internal accelerometer device. The medical device includes a circuit board, the accelerometer device, and a response module coupled to the accelerometer device. The accelerometer device is mechanically and electrically coupled to the circuit board, and it includes a plurality of mass-supporting arms for a plurality of electrically distinct sensor electrodes, piezoelectric material for the mass-supporting arm, and a proof mass supported by the mass-supporting arms. Each of the mass-supporting arms has one of the sensor electrodes located thereon. Acceleration of the proof mass causes deflection of the piezoelectric material, which generates respective sensor signals at one or more of the sensor electrodes. The response module is configured to initiate an acceleration-dependent operation of the portable medical device in response to generated sensor signals present at the sensor electrodes. | 1. An omnidirectional accelerometer device comprising: a piezoelectric sensor element comprising an electrically conductive support substrate, a layer of piezoelectric material overlying the support substrate, and a plurality of electrically conductive sensor electrodes overlying the piezoelectric material, the piezoelectric sensor element further comprising a mass-supporting platform and a plurality of mass-supporting arms, each of the sensor electrodes being located on a corresponding one of the mass-supporting arms;a proof mass coupled to the mass-supporting platform by a connecting rod that holds the proof mass above a surface of the piezoelectric sensor element at a specified height;an electromechanical mounting arrangement configured to mechanically and electrically couple the piezoelectric sensor element to a circuit board such that each of the sensor electrodes is electrically coupled to a respective contact pad of the circuit board;an electrically conductive offset block that mechanically and electrically couples the support substrate of the piezoelectric sensor element to a ground contact pad of the circuit board; anda signal processing module coupled to the sensor electrodes, the signal processing module being configured to process sensor signals generated in response to deflection of the piezoelectric sensor element caused by the proof mass, the signal processing module comprising a first signal processing sub-module configured to process the sensor signals for human activity monitoring, and a second signal processing sub-module configured to process the sensor signals for impact detection; wherein:the first signal processing sub-module and the second signal processing sub-module are coupled in parallel to the sensor electrodes to enable concurrent processing of the sensor signals;the first signal processing sub-module comprises a sensor signal input element, an amplifier-rectifier coupled to an output of the sensor signal input element, a high pass filter coupled to an output of the amplifier-rectifier, a voltage limiter coupled to an output of the high pass filter and an integrator coupled to an output of the voltage limiter; andthe second signal processing sub-module comprises the sensor signal input element, the amplifier-rectifier, and a peak and hold circuit coupled to an output of the amplifier-rectifier. 2. The omnidirectional accelerometer device of claim 1, wherein each of the sensor electrodes represents a separate and distinct electrical sensing node. 3. The omnidirectional accelerometer device of claim 1, the electromechanical mounting arrangement comprising: a plurality of electrically conductive mounting bases; anda plurality of electrically conductive mounting tabs, wherein each of the mass-supporting arms is mechanically and electrically coupled to the circuit board with one of the mounting bases and one of the mounting tabs, the mounting bases are electrically coupled to the support substrate, and each of the mount... | Medtronic Minimed, Inc. | A61N001/00
US07950163 | B2 | 2011-05-31 | Foot type and plantar pressure measuring device | 2009-12-30 | A foot type and plantar pressure measuring device is provided. The device includes a color pad and a foggy film on the color pad. The device further includes a mesh pad on the foggy film and/or a resilient pad under the color pad, the resilient pad having protrusions. When a patient stands still on the device for a few seconds and then steps off, the patient's foot type and plantar pressure is shown by color pattern displayed on surface areas of the foggy film. The color pattern corresponds to the plantar pressure exerted on the device, and tones of the color pattern represent the plantar pressure exerted on the surface areas. | 1. A foot type and plantar pressure measuring device, comprising: a soft, flexible color pad made of a sticky material; anda semi-transparent film on the color pad;wherein surface areas of the semi-transparent film display colors in response to a pressure exerted on the semi-transparent film and the color pad, andwherein tones of the colors are associated with the pressure exerted on the surface areas. 2. The device of claim 1, further comprising a meshed pad having a plurality of holes, the meshed pad being placed on the semi-transparent film. 3. The device of claim 2, wherein at least one of the holes are of a circular shape, a circle-like shape, or a polygonal shape. 4. The device of claim 1, wherein the semi-transparent film comprises a foggy film. 5. The device of claim 1, wherein the color pad is of a dark color. 6. The device of claim 5, wherein the dark color comprises any color from gray to black. 7. The device of claim 1, further comprising a resilient pad formed under the color pad. 8. The device of claim 7, wherein the resilient pad includes a plurality of protrusions for exerting a reactive force on the color pad. 9. A device for measuring plantar pressure, comprising: a resilient pad including a plurality of protrusions;a color pad on the resilient pad; anda semi-transparent film on the color pad, wherein surface areas of the semi-transparent film display colors in response to a pressure exerted on the semi-transparent film and the color pad. 10. The device of claim 9, wherein the protrusions contact the color pad. 11. The device of claim 9, wherein tones of the colors are associated with the pressure exerted on the surface areas. 12. The device of claim 9, further comprising a meshed pad on the semi-transparent film, the meshed pad having a plurality of holes. 13. A device for measuring plantar pressure, comprising: a resilient pad including a plurality of protrusions;a color pad on the resilient pad, the color pad being made of a sticky material;a light transmitting film on the color pad; anda meshed pad on the light transmitting film, the meshed pad having a plurality of holes. 14. The device of claim 13, wherein at least one of the holes are of a circular shape, a circle-like shape, or a polygonal shape. 15. The device of claim 13, wherein the protrusions contact the color pad. | - | A61B005/103 G01B021/20
US07959574 | B2 | 2011-06-14 | System and method for analyzing a patient status for respiratory insufficiency for use in automated patient care | 2007-03-26 | A system and method for analyzing a patient status for respiratory insufficiency for use in automated patient care. A set of one or more physiological measures relating to patient information recorded on a substantially continuous basis or derived therefrom are received. The physiological measures set is stored into a patient care record for an individual patient into a database. One or more of the physiological measures in the physiological measures set are analyzed relative to one or more other physiological measures to determine a pathophysiology indicating an absence, an onset, a progression, a regression, and a status quo of respiratory insufficiency. | 1. A system for analyzing a patient status for respiratory insufficiency for use in automated care, comprising a server system receiving a set of one or more physiological measures relating to patient information recorded on a substantially continuous basis or derived therefrom, the server system further receiving one or more baseline physiological measures relating to patient information recorded during an initial time period or derived therefrom;a database module storing the physiological measures into a patient care record for an individual patient into a database, the database module further storing the baseline physiological measures into the patient care record for the individual patient into the database;an analyzer analyzing one or more of the physiological measures in the physiological measures set relative to one or more other physiological measures to determine a patient status indicating at least one of a change and a status quo of respiratory insufficiency, the analyzer analyzing the physiological measures in the physiological measures set relative to the baseline physiological measures to determine a further patient status indicating at least one of a change and a status quo of respiratory insufficiency;wherein the further patient status is selected to indicate a progression that is trending into a state of likely respiratory insufficiency. 2. A system according to claim 1, further comprising: the analyzer analyzing the physiological measures in the physiological measures set relative to the other physiological measures to determine a further patient status indicating at least one of a change and a status quo of diseases other than respiratory insufficiency. 3. A system according to claim 1, further comprising: the server system receiving one or more quality of life measures relating to patient information recorded on a substantially continuous basis or derived therefrom;the database module storing the quality of life measures into the patient care record for the individual patient into the database; andthe analyzer analyzing the quality of life measures in the physiological measures set relative to the other quality of life measures to determine a further patient status indicating at least one of a change and a status quo of respiratory insufficiency. 4. A system according to claim 1, wherein the further patient status to indicate the progression is inverted as the patient status to indicate a regression that is trending from a state of likely respiratory insufficiency. 5. A system according to claim 1, wherein the patient status to indicate the progression into a state of likely respiratory insufficiency is selected from a group comprising: an absolute limit of arterial oxygen less than 85 mm Hg; andan absolute limit of arterial carbon dioxide falling below 25 mm Hg in an absence of marked exercise as indicated by hyperventilation; andan absolute limit of arterial carbon dioxide greater than 50 mm Hg as indicated by hypoventil... | Cardiac Pacemakers, Inc. | A61B005/08 A61B005/02
US07970460 | B2 | 2011-06-28 | Method and apparatus for suppressing power frequency common mode interference | 2010-02-23 | An apparatus for suppressing power frequency common mode interference in a bioelectrical signal measuring system includes a driving circuit configured to amplify and change the phase of a common mode interference signal to produce a first driving signal, wherein the common mode interference signal is extracted from a plurality of first electrodes attached to a patient; phase compensating and processing circuitry electrically connected to the driving circuit for receiving the first driving signal, the phase compensating and processing circuitry configured to produce a second driving signal by phase-compensating the first driving signal based on a characteristic value of power frequency interference in a bioelectrical signal acquired through the plurality of first electrodes; and a switch to receive the first driving signal and the second driving signal, the switch configured to selectively switch between providing the first driving signal and providing the second driving signal to a second electrode attached to the patient. | 1. An apparatus for suppressing power frequency common mode interference in a bioelectrical signal measuring system, comprising: a driving circuit configured to amplify and change the phase of a common mode interference signal to produce a first driving signal, wherein the common mode interference signal is extracted from a plurality of first electrodes attached to a patient;phase compensating and processing circuitry electrically connected to the driving circuit for receiving the first driving signal, the phase compensating and processing circuitry configured to produce a second driving signal by phase-compensating the first driving signal based on a characteristic value of power frequency interference in a bioelectrical signal acquired through the plurality of first electrodes; anda switch to receive the first driving signal and the second driving signal, the switch configured to selectively switch between providing the first driving signal and providing the second driving signal to a second electrode attached to the patient. 2. The apparatus of claim 1, wherein the phase compensating and processing circuitry comprises: at least one analog-to-digital (A/D) converter to convert the bioelectrical signal and the first driving signal from the driving circuit into respective digital signals;a microprocessor to receive the digital signals and to determine a phase compensation amount of the first driving signal according to the characteristic value of the power frequency interference in the bioelectrical signal; anda digital to analog (D/A) converter for receiving a signal from the microprocessor and converting the signal into an analog signal. 3. The apparatus of claim 1, further comprising: a low-pass filter to filter the phase-compensated second driving signal. 4. A method for suppressing power frequency common mode interference in a bioelectrical signal measuring system, comprising: receiving a common mode interference signal extracted from a plurality of first electrodes attached to a patient;amplifying and changing the phase of the common mode interference signal to produce a first driving signal;producing a second driving signal by phase-compensating the first driving signal based on a characteristic value of power frequency interference in a bioelectrical signal acquired through the plurality of first electrodes; andselectively switching between providing the first driving signal and providing the second driving signal to a second electrode attached to the patient. 5. The method of claim 4, wherein phase-compensating the first driving signal comprises: initializing a system state value of a measuring system;extracting a first characteristic value of the power frequency interference in a current system state of the measuring system;adding 1 to the system state value and extracting a second characteristic value;if the second characteristic value has decreased from the first characteristic value, determining that the power frequency interfere... | Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | A61B005/04
US07961921 | B2 | 2011-06-14 | System and method for medical diagnosis and tracking using three-dimensional subtraction in a picture archiving communication system | 2005-09-27 | Certain embodiments of the present invention provide for a system for medical diagnosis and tracking. In an embodiment, a medical diagnosis and tracking system includes one or more medical imaging devices adapted to create one or more current images. The medical diagnosis and tracking system also includes a historical images archive adapted to storing the one or more current images and/or one or more historical images. The system also includes a three-dimensional subtraction application adapted to create one or more three-dimensional subtracted images by performing three-dimensional subtraction on the one or more current images and the one or more historical images. The medical diagnosis and tracking system also includes a Picture Archiving Communication System (PACS) workstation adapted to display the one or more three-dimensional subtracted images. | 1. A system for medical diagnosis and tracking, comprising: at least one medical imaging device adapted to create at least one current image from a current examination;a historical images archive adapted to storing at least one of: said at least one current image, andat least one historical image from a previous examination;a rules engine adapted to apply at least one rule to select at least one historical image for three-dimensional subtraction with said at least one current image;a three-dimensional subtraction application adapted to create at least one three-dimensional subtracted image by performing three-dimensional subtraction on said at least one current image and said at least one historical image, wherein said at least one current image and said at least one historical image do not contain contrast; anda Picture Archiving Communication System (PACS) workstation adapted to display said at least one three-dimensional subtracted image. 2. The system of claim 1, wherein said at least one rule is at least one of: a predefined rule, anda rule created dynamically by a user of said system for medical diagnosis and tracking. 3. The system of claim 1, wherein said at least one rule is defined by at least one of: software,a user, anda system administrator. 4. The system of claim 1, wherein said system is used for tumor diagnosis and tracking. 5. The system of claim 1, wherein said historical images archive, said rules engine, and said three-dimensional subtraction application are integrated on the Picture Archiving Communication System (PACS) workstation. 6. The system of claim 1, wherein said three-dimensional subtraction application uses image registration techniques to align said at least one current image and said at least one historical image before creating said at least one three-dimensional subtracted image. 7. A method for medical diagnosis and tracking, the method comprising: at least one processing device for performing at least: receiving at least one current medical image from a current examination;applying at least one rule for selecting at least one historical image from a previous examination for three-dimensional subtraction with said at least one current image from said current examination;generating at least one current three-dimensional volume image using said at least one current medical image;generating at least one historical three-dimensional volume image using said at least one historical image;registering said at least one current three-dimensional volume image and said at least one historical three-dimensional volume image using image registration techniques;constructing at least one three-dimensional subtracted image from said registered at least one current three-dimensional volume image and said at least one historical three-dimensional volume image, wherein said at least one current three-dimensional volume image and said at least one historical three-dimensional volume image do not contain contrast; anddisplayi... | General Electric Company | G06K009/00 A61B005/05 A61B006/00
US08326411 | B2 | 2012-12-04 | Trunk width measuring unit and visceral fat measuring device | 2011-01-19 | A trunk width measuring unit including a first contact portion attached movably in the vertical direction relative to a unit main body, the first contact portion to be brought into contact with an upper surface of a trunk of a subject, second and third contact portions attached movably in the horizontal direction relative to the unit main body, a vertical width detection encoder for detecting a lowering distance of the first contact portion from an original position on the upper side of the trunk of the user to a contact point with the upper surface of the trunk, and a horizontal width detection encoder attached movably in the horizontal direction relative to the unit main body by the same distance as a moving distance of the second contact portion, the horizontal width detection encoder detecting a moving distance of the third contact portion relative to the second contact portion. | 1. A trunk width measuring unit, comprising: a unit main body;a first contact portion attached movably in the vertical direction relative to the unit main body, the first contact portion being configured to be brought into contact with an upper surface of a trunk of a user in a supine position, the first contact portion being made of a material extending in a horizontal direction;a second contact portion attached movably in the horizontal direction relative to the unit main body, the second contact portion being configured to be brought into contact with one side surface of the trunk of the user in the supine position;a third contact portion attached movably in the horizontal direction relative to the unit main body, the third contact portion being configured to be brought into contact with an other side surface of the trunk of the user in the supine position;a vertical width detection encoder for detecting a lowering distance of the first contact portion from an original position on the upper side of the trunk of the user to a contact point with the upper surface of the trunk; anda horizontal width detection encoder attached movably relative to the unit main body in conjunction with a moving distance of one contact portion among the second contact portion and the third contact portion, the horizontal width detection encoder being configured to detect a moving distance of the other contact portion relative to the one contact portion. 2. The trunk width measuring unit according to claim 1, wherein the unit main body has a conjunction mechanism for moving the horizontal width detection encoder in the horizontal direction in conjunction with movement of the one contact portion. 3. The trunk width measuring unit according to claim 1, wherein the horizontal width detection encoder is movable by the same distance as the moving distance of the one contact portion among the second contact portion and the third contact portion. 4. The trunk width measuring unit according to claim 2, wherein the conjunction mechanism is provided with a first rack-and-pinion mechanism for converting linear motion of the one contact portion into rotary motion, a rotation shaft driven and rotated by the first rack-and-pinion mechanism, and a second rack-and-pinion mechanism for converting the rotary motion of the rotation shaft into linear motion of a casing that houses the horizontal width detection encoder. 5. The trunk width measuring unit according to claim 2, wherein the conjunction mechanism is provided with a wire or a belt for coupling the one contact portion and a casing that houses the horizontal width detection encoder. 6. The trunk width measuring unit according to claim 1, wherein the horizontal width detection encoder is attached integrally with the one contact portion. 7. The trunk width measuring unit according to claim 1, comprising: a trunk width calculating unit for: determining vertical width of the trunk of the user by calculating height from a plan... | Omron Healthcare Co., Ltd. | A61B005/05 A61B001/00 A61B005/117 B43L007/10 B43L013/00 G01B001/00 G01B005/00 G01B003/08
US07973540 | B2 | 2011-07-05 | Coupling loop and method of positioning coupling loop | 2010-04-23 | A coupling loop or antenna is provided that can be used with a system that determines the resonant frequency of a sensor by adjusting the phase and frequency of an energizing signal until the frequency of the energizing signal matches the resonant frequency of the sensor. In one embodiment orientation features are provided for positioning the coupling loop relative to the sensor to maximize the coupling between the sensor and the coupling loop. | 1. A method of positioning a coupling loop relative to a sensor, comprising: positioning the coupling loop relative to the sensor so that a first area of a housing for the coupling loop is in a predetermined position with respect to the sensor; andadjusting an angle of the coupling loop relative to the sensor until a predefined pattern is formed by markings on the sensor and markings on the housing,wherein a medical imaging technology is used for positioning the coupling loop and adjusting the angle of the coupling loop. 2. The method of claim 1, wherein the sensor includes a flat side and an inductor parallel to the flat side and wherein the predefined pattern is formed when a plane formed by the coupling loop is approximately parallel to the flat side of the sensor. 3. The method of claim 1, wherein the predefined pattern is selected from the group consisting of: a bulls eye, a cross hair, a cross, a logo, an image, and an alphanumeric string. 4. The method of claim 1, wherein the first area of the housing is approximately at a center of a receive loop within the housing. 5. A coupling loop, comprising: a sensor coupling loop for receiving a signal from a sensor; anda housing surrounding the sensor coupling loop and having a first area that is approximately at the center of the sensor coupling loop, wherein the first area includes an orientation mark that is visible using a medical imaging technology and is used to position the coupling loop relative to the sensor. 6. The coupling loop of claim 5, wherein the orientation mark is formed from support ribbing inside the housing. 7. The coupling loop of claim 5, wherein the first area is co-planar with a plane extending through the sensor coupling loop. 8. The coupling loop of claim 5, wherein the orientation mark is selected from the group consisting of: a circle, a bulls eye, a cross hair, a cross, a logo, an image and an alphanumeric string. 9. The coupling loop of claim 5, wherein the orientation mark comprises a plurality of marks that indicate a length or a distance. | CarioMEMS, Inc. | G01R027/28
US07941209 | B2 | 2011-05-10 | Signal analysis method | 2005-05-06 | Improvement in the reliability of segmentation of a signal, such as an ECG signal, is achieved through the use of duration constraints. The signal is analysed using a hidden Markov model. The duration constraints specify minimum allowed durations for specific states of the model. The duration constraints can be incorporated either in the model itself or in a Viterbi algorithm used to compute the most probable state sequence given a conventional model. The derivation of a confidence measure from the model can be used to assess the quality and robustness of the segmentation and to identify any signals for which the segmentation is unreliable, for example due to the presence of noise or abnormality in the signal. | 1. A computer-implemented method for analysing a signal which has been segmented using a probabilistic segmentation model, the method comprising: calculating a confidence measure for each of a plurality of segmented signal features;plotting each calculated confidence measure against a length of its respective signal feature;applying density modelling techniques to the plot to determine a suitable region associated with high confidence measures;determining whether a confidence measure for a specific signal feature falls outside this region; andwherein the signal is a bio-medical signal selected from the group consisting of an electrocardiogram signal, an electroencephalogram signal, an electromyogram signal, a cardiac electrophysiology signal, and an endocardial electrogram. 2. A method according to claim 1, wherein each of the plurality of segmented signal features comprise a portion of a longer signal. 3. A method according to claim 1, wherein the confidence measure is one selected from the group consisting of: the log likelihood of the signal feature given the model; and the joint log likelihood of the signal feature and the most probably state sequence given the model. 4. A method according to claim 1, further comprising displaying the plot of each calculated confidence measure against the length of its respective signal feature. 5. A method according to claim 1, wherein the signal is an electrocardiogram signal and information on a QT interval is obtained from the probabilistic segmentation model. 6. A non-transitory computer-readable medium storing a computer program capable of being executed on a computer system and capable when so executed of causing the computer system to perform a method according to claim 1. | OBS Medical Limited | A61B005/044 A61B005/0476 A61B005/0488
US07935065 | B2 | 2011-05-03 | Oral device | 2006-04-26 | An oral device for use with a person in which at least one non-toxic gas pulse is delivered to a predetermined location in the mouth via a device conduit. The oral device may be provided as a kit with at least one device for measuring the subject's responses and representing them as feedback to the subject/clinician. A method of creating a gas bolus pulse train, delivering it to a predetermined mouth area, and monitoring the subject's responses to it, is also shown. The oral device and method may be used as a diagnostic tool, or a therapeutic tool, in swallowing or speech rehabilitation of children and adults who have swallowing, speech, salivary, and/or oral sensorimotor impairments. | 1. A method of delivering a gas in a subject's mouth comprising: inserting at least one conduit into the subject's mouth, said at least one conduit comprising at least one exit port;positioning said at least one conduit in the subject's mouth such that said at least one exit port is directed at a predetermined location in the subject's mouth, and such that said at least one conduit is not disposed between the subject's tongue and a roof of the mouth, wherein said tongue is free to touch said roof of the mouth when said at least one conduit is positioned in the subject's mouth;delivering a gas to said predetermined location in the subject's mouth; andstimulating said predetermined location in the subject's mouth with said gas and thereby inducing swallowing by the subject with said at least one conduit remaining in the subject's mouth during said swallowing. 2. A method as claimed in claim 1 wherein said delivering said gas comprises delivering said gas at a predetermined frequency between 0.5 and 30 Hz. 3. A method as claimed in claim 2 wherein said delivering said gas comprises delivering a gas pulse train having a duration of between 2 and 20 seconds. 4. A method as claimed in claim 3 further comprising delivering said gas pulse train at a pressure between 3 cm water and 15 cm water. 5. A method as claimed in claim 1 wherein the predetermined location is the peri-tonsillar area. 6. A method as claimed in claim 1 wherein the predetermined location is the tongue. 7. A method as claimed in claim 1 wherein the predetermined location is the roof of the mouth. 8. A method as claimed in claim 1 wherein said at least one conduit is carried by a dental splint. 9. A method as claimed in claim 1 wherein said at least one conduit comprises a plurality of holes formed therein. 10. A method as claimed in claim 1 wherein said gas is air. 11. A method as claimed in claim 1 further comprising closing the subject's jaws with said conduit positioned in the subject's mouth such that a spacing between the subject's jaws is no more than a just noticeable difference for jaw position. 12. A method as claimed in claim 11 wherein said closing the subject's jaw such that said spacing therebetween is no more than said just noticeable difference for jaw position comprises substantially closing said subject's upper and lower teeth such that a space between said upper and lower teeth is less than 2 mm. 13. A method as claimed in claim 12 wherein said space between said upper and lower teeth is between 1 and 2 mm. 14. A method of delivering a gas in a subject's mouth comprising: inserting first and second conduits on opposite sides of the mouth, said first and second conduits each comprising at least one exit port;positioning said first and second conduits in the subject's mouth such that said exit ports are directed at predetermined locations in the subject's mouth;delivering a gas to said predetermined locations in the subject's mouth; andstimulating said prede... | The University of Western Ontario | A61B005/103 A61B005/117
US07945315 | B2 | 2011-05-17 | System for analysis of electrograms | 2007-06-29 | A system for use in analysis of electrograms comprising: an input signal generator (302); an input electrode (304) for applying an input signal to a driving region of a heart (316); an output electrode (306a-c) for receiving an output signal at a driven region of the heart; a processing system (300) operable to receive signals indicative of said recorded value from the output electrode for analysing conduction paths through the heart, wherein the signal generator is operable to generate an input signal comprising a plurality of pulses, being spaced from each other by a pacing interval; and the processing system being arranged to identify signal delay between the input signal and the output signal on the basis of the signal received by the output electrode in relation to the plurality of pulses, and to identify a rate of variation in signal delay over a range of values of pacing interval. | 1. A system for use in analysis of electrograms, the system comprising: a signal generator for generating an input signal;an input electrode for applying an input signal to a driving region of a heart organ;an output electrode for receiving an output signal at a driven region of the heart organ, so as to record a value thereof;a processing system operable to receive signals indicative of said recorded value from the output electrode for analysing conduction paths through the heart organ,wherein the signal generator is operable to generate an input signal comprising a plurality of pulses, at least some of the pulses from said plurality being spaced from each other by a pacing interval;and wherein the processing system is arranged to identify signal delay between the input signal and the output signal on the basis of the signal received by the output electrode in relation to the plurality of pulses,wherein the processing system is arranged to identify a rate of variation in signal delay over a range of values of pacing interval. 2. A system according to claim 1, wherein said processing system is arranged to identify first and second rates of variation in signal delay within the range of values of pacing interval. 3. A system according to claim 2, wherein said processing system is further arranged to compare the first and second rates of variation in signal delay so as to generate a parameter having a value indicative of a difference in said first and second rates of variation in signal delay. 4. A system according to claim 3, wherein said processing system is further arranged to compare said parameter with at least one known parameter value, so as to identify one of a plurality of physiological cardiac conditions. 5. A system according to claim 3, wherein said processing system is arranged to use the output signal so as to construct a graphical representation of output potential against the pacing interval, and said first and second rates of variation in signal delay relate to said graphical representation. 6. A system according to claim 5, wherein said first and second rates of signal delay relate to the gradient of a first portion of the graphical representation, and the gradient of a second portion of the graphical representation, respectively. 7. A system according to claim 6, wherein said parameter is derived from a relative angle between the gradient of the first portion of the graphical representation, and the gradient of the second portion of the graphical representation. 8. A system according to claim 7, wherein if said parameter value is lower than 150° a first physiological cardiac condition is identified and, if said parameter value is higher than approximately 150°, a second physiological cardiac condition is identified. 9. A system according to claim 7, wherein if said parameter value is between approximately 115° and 135° a first physiological cardiac condition is identified and, if said parameter value is between approximate... | Medilec Limited | A61B005/04
US20110152249 | A1 | 2011-06-23 | EVALUATING CENTRAL NERVOUS SYSTEM | 2005-12-23 | Evaluating hemodynamic parameters in the cortex can be used as an objective test for CNS activity, e.g., pain. | 1. A method for measuring a central nervous system (CNS) activity in a subject, the method comprising: identifying a subject as a subject who experiences pain or analgesia; anddetermining a level of activation in a cortical region of the subject by detecting a hemodynamic parameter in the subject, wherein the level of activation is indicative of the CNS activity in the subject. 2. The method of claim 1, wherein the pain is acute pain or chronic pain. 3. The method of claim 2, wherein the chronic pain is chronic nociceptive pain or neuropathic pain. 4. The method of claim 1, wherein the subject is anesthetized. 5. The method of claim 1, wherein detecting a hemodynamic parameter comprises performing diffuse optical tomography (DOT) and producing an image of the cortical region from which the level of activation is determined. 6. The method of claim 1, wherein the level of activation is determined relative to a reference level of activation. 7. The method of claim 6, wherein the reference level of activation is based on an assessment of the subject at different time. 8. The method of claim 1, further comprising administering an anesthetic or pain reductant if the level of activation is indicative of pain. 9. The method of claim 1, further comprising reducing the concentration an anesthetic in a subject if the level of activation is indicative of analgesia. 10. The method of claim 1, further comprising administering an anesthetic if the level of activation is indicative of pain. 11. The method of claim 1, wherein the SI cortical region is evaluated. 12. A method for identifying a therapeutic intervention for a neurological condition, the method comprising: (a) administering a test intervention to a mammalian subject, wherein the subject has a neurological condition or is subjected to a stimulus capable of inducing CNS activity indicative of the neurological condition;(b) performing diffuse optical tomography (DOT) on the subject; and(c) determining a level of activation of a cortical region, wherein the level of activation is indicative of the effectiveness of the test intervention as a therapeutic intervention. 13. The method of claim 12, wherein the neurological condition comprises a psychiatric disorder. 14. The method of claim 13, wherein the psychiatric disorder comprises depression or anxiety. 15. The method of claim 12, wherein the neurological condition comprises pain. 16. The method of claim 12, wherein the stimulus is a noxious stimulus. 17. The method of claim 16, wherein the noxious stimulus is an acute painful stimulus. 18. The method of claim 12, wherein the stimulus is a non-noxious stimulus. 19. The method of claim 12, wherein the level of activation is determined relative to a control. 20. The method of claim 12, wherein the level of activation is determined relative to a control image of the cortical region of a control subject, wherein the control image is obtained during the presence of the stimulus, and the tes... | THE GENERAL HOSPITAL CORPORATION THE BRIGHAM AND WOMEN'S HOSPITAL, INC. | A61K031/55 A61P025/00 A61P025/22 A61P025/24 A61P025/18 A61P029/00 A61B005/05
US20110118556 | A1 | 2011-05-19 | Systems and Methods for Diagnostics, Control and Treatment of Neurological Functions and Disorders by Exposure to Electromagnetic Waves | 2010-11-12 | Methods, systems and devices are provided in which electromagnetic waves (EMWs) are applied to a neural target. The neuronal function of at least a portion of the neural target is monitored, and the effect of EMW application on neural properties, neural function, and/or neural disorders is assessed. Embodiments include methods of diagnosis, treatment, drug delivery, cognitive enhancement, and EMW stimulation systems. | 1. A method of diagnosing a neural disorder, comprising: applying electromagnetic waves (EMWs) at a frequency in the range of 3 to 300 GHz to a neural target of a subject so as to alter a neural property of the target,comparing the neural property of the target after EMW application to the same neural property of a similarly treated control target, anddiagnosing a neural disorder in the subject when a difference in the neural properties is observed. 2. The method of claim 1, further comprising discontinuing the EMW application in the event that the EMWs are applied in a manner that exceeds a predetermined upper EMW application limit. 3. The method of claim 1, wherein the applying of the EMWs comprises applying several EMW amplitudes or applying in a duty cycled manner. 4. The method of claim 1, wherein the neural target is selected from the group consisting of cerebral cortex, deep brain, brainstem, spinal cord, optic tract, optic nerve, retina, cranial nerves, spinal roots, peripheral nerves, and nerve endings in the skin, and a combination thereof. 5. A method of treating a neural disorder, comprising: exposing a neural target of a subject in need of such treatment to electromagnetic waves (EMWs) at a frequency in the range of 3 to 300 GHz in an amount sufficient to alter a neural property of the neural target, thereby conferring a therapeutic benefit on the subject, wherein the neural target is identified with a neural disorder. 6. The method of claim 5, further comprising discontinuing the EMW exposure in the event that the EMWs are applied in a manner that exceeds a predetermined upper EMW exposure limit. 7. The method of claim 5, wherein the exposing of the neural target comprises exposing the neural target to several EMW amplitudes or to EMWs in a duty cycled manner. 8. The method of claim 5, further comprising monitoring neuronal function of the subject. 9. The method of claim 5, wherein the neural target is selected from the group consisting of cerebral cortex, deep brain, brainstem, spinal cord, optic tract, optic nerve, retina, cranial nerves, spinal roots, peripheral nerves, and nerve endings in the skin, and a combination thereof. 10. A method of delivering a drug, comprising: exposing a neural target in a subject's central nervous system or optic system to electromagnetic waves (EMWs) at a frequency in the range of 3 to 300 GHz in an amount sufficient to alter the subject's blood-brain or blood-retinal barrier, respectively, andadministering a drug to the subject before, during, or after the EMW exposure, or any combination thereof, such that the drug crosses the blood-brain or blood-retinal barrier. 11. A method of cognitive enhancement, comprising exposing a neural target in a subject's central nervous system to electromagnetic waves (EMWs) at a frequency in the range of 3 to 300 GHz in an amount sufficient to alter a neural property of the neural target, thereby increasing cognitive ability in the subject. 12. An elect... | - | A61B005/00 A61N005/00 A61M037/00
US07949390 | B1 | 2011-05-24 | Time domain monitoring of myocardial electrical stability | 2006-02-14 | Embodiments of the present invention relate to implantable systems, and methods for use therewith, for monitoring myocardial electrical stability. A patient's heart is paced for a period of time using a patterned pacing sequence that repeats every N beats, and an electrical signal is obtained that is representative of a plurality of consecutive beats of the patient's heart while it is being paced using the patterned pacing sequence that repeats every N beats. Myocardial electrical stability is then analyzed using time domain techniques that are tailored to the patterned pacing sequence used to pace the patient's heart. In other embodiments, the patient's heart need not be paced. This abstract is not intended to be a complete description of, or limit the scope of, the invention. | 1. In an implantable system, a method for time domain monitoring of myocardial electrical stability, comprising: (a) pacing a patient's heart for a period of time using a patterned pacing sequence that repeats every N beats, where N is an integer that is at least 3;(b) obtaining an electrical signal that is representative of a plurality of consecutive beats of the patient's heart while the patient's heart is being paced using the patterned pacing sequence that repeats every N beats;(c) dividing the plurality of consecutive beats into a plurality of sets of N consecutive beats;(d) for each set of N consecutive beats, of the plurality of sets of N consecutive beats, determining one or more pairwise combination of consecutive pairs of beats within the set of N consecutive beats;(e) cumulative averaging or cumulative summing corresponding pairwise combinations determined for the plurality of sets of N consecutive beats to thereby produce a plurality of cumulative values; and(f) monitoring myocardial electrical stability based on the cumulative values;wherein steps (b), (c), (d), (e) and (f) are all performed in the time domain without any transformation into the frequency domain. 2. The method of claim 1, wherein step (f) comprises determining, based on the cumulative values, whether electrical alternans are present. 3. The method of claim 2, wherein: step (e) comprises cumulative averaging corresponding pairwise differences determined for the plurality of sets of N consecutive beats to thereby produce the plurality of cumulative values; andstep (f) comprises determining that electrical alternans are present when the cumulative values exceed a specified threshold. 4. The method of claim 2, wherein: step (e) comprises cumulative summing corresponding pairwise differences determined for the plurality of sets of N consecutive beats to thereby produce the plurality of cumulative values; andstep (f) comprises determining that electrical alternans are present when a slope of the cumulative values exceeds a specified slope threshold. 5. The method of claim 1, wherein step (f) comprises tracking changes in myocardial electrical stability as steps (a) through (e) are repeated over time. 6. The method of claim 1, further comprising bounding the pairwise combinations determined at step (d) prior to the cumulative averaging or cumulative summing performed at step (e). 7. The method of claim 1, further comprising: detecting disruptive beats, if they exist, based on the cumulative values determined at step (e); andif one or more disruptive beats are detected, compensating for the disruptive beats, and performing steps (c), (d) and (e) again before monitoring myocardial electrical stability at step (f). 8. The method of claim 7, wherein: step (e) comprises cumulative averaging corresponding pairwise differences determined for the plurality of sets of N consecutive beats to thereby produce the plurality of cumulative values; andthe detecting disruptive beats ... | Pacesetter, Inc. | A61B005/02
US07941208 | B2 | 2011-05-10 | Therapy delivery for identified tachyarrhythmia episode types | 2007-03-07 | Methods and systems for identifying tachyarrhythmia episode types and delivering therapy to mitigate the identified tachyarrhythmia episode types are described. Electrogram signals of cardiac activity are sensed and stored by an implantable cardiac device. Tachyarrhythmia episodes are detected and tachyarrhythmia episode types are identified based on characteristics of the electrogram signals. In preparation for performing ablation, a tachyarrhythmia episode is induced. The features of the induced tachyarrhythmia episode are compared to characteristics of the identified episode types. A similarity between the induced tachyarrhythmia episode and at least one of the episode types identified from the stored electrogram signals is indicated to facilitate performing the ablation. | 1. A method of operating a cardiac system that includes an implantable cardiac device that is implanted in a patient, the implantable cardiac device including implantable electrodes adapted to be positioned at respective specific sites within the patient, the method comprising: for a first period of time: sensing and storing electrogram signals of cardiac electrical activity from the patient in the implantable cardiac device, the sensing being carried out using the implantable electrodes;detecting first intrinsic tachyarrhythmia episodes using the electrogram signals; andidentifying first tachyarrhythmia episode morphology types based on morphological characteristics of the electrogram signals for the detected first intrinsic tachyarrhythmia episodes; andafter the first period of time, in preparation for performing ablation: inducing a tachyarrhythmia episode;sensing electrogram signals from the induced tachyarrhythmia episode using the implantable electrodes;using the sensed electrogram signals from the induced tachyarrhythmia episode to compare morphological characteristics of the induced tachyarrhythmia episode to the morphological characteristics of the detected first intrinsic tachyarrhythmia episodes; andindicating a similarity between the induced tachyarrhythmia episode and at least one of the first tachyarrhythmia episode morphology types identified from the stored electrogram signals to facilitate performing the ablation;wherein the method further comprises: for each identified first tachyarrhythmia episode morphology type, determining how many of the detected first intrinsic tachyarrhythmia episodes exhibited that identified episode morphology type; andfor each identified first tachyarrhythmia episode morphology type, displaying a number of how many first intrinsic tachyarrhythmia episodes were detected for that identified episode morphology type. 2. The method of claim 1, further comprising ranking the episode types based on the number of detected first intrinsic tachyarrhythmia episodes determined for each identified episode type, a rank of an episode type corresponding to a tachyarrhythmia burden of the episode type. 3. The method of claim 1, further comprising performing the ablation in such a way as to prevent future occurrences of one or more of the identified episode types. 4. The method of claim 1, further comprising: for a second period of time after performing the ablation: sensing and storing second electrogram signals of cardiac electrical activity from the patient in the implantable cardiac device;detecting second intrinsic tachyarrhythmia episodes using the second electrogram signals; andidentifying second tachyarrhythmia episode morphology types based on characteristics of the second electrogram signals; andcomparing the first tachyarrhythmia episode morphology types identified from the first intrinsic tachyarrhythmia episodes stored before ablation to the second tachyarrhythmia episode morphology types identified from... | Cardiac Pacemakers, Inc. | A61B005/04
US07965851 | B2 | 2011-06-21 | Modular wireless auditory test instrument with intelligent transducers | 2003-03-21 | A method and apparatus for performing various auditory tests utilizing a hand-held, portable, wireless testing device is provided. Within the device is a diagnostic subsystem used to implement and analyze selected tests. Attached to the device, either directly or via a flexible cable, are one or more probes. A processor, storage means, input means and display means are included, thus allowing the device to process and store instructions as well as process, store and display data. A wireless networking subsystem is included, enabling the device to communicate with other, similarly enabled, systems and devices within the device's communication range. Due to the inclusion of the wireless networking subsystem, the device can be configured to allow the user to transmit and/or print data, connect to a network, obtain device configuration updates, and send and receive patient and office updates. | 1. An auditory test instrument comprising: a connector configured to connect to at least one test probe;a power supply for supplying power to the test instrument, wherein said power supply is portable;a wireless networking subsystem configured to wirelessly transmit first signals via a short distance wireless network to a peripheral electronic device or system;a diagnostic subsystem coupled to said at least one test probe, said diagnostic subsystem configured to implement at least one auditory diagnostic test;input means configured to accept commands from a user;a display configured to display results from said at least one auditory diagnostic test; andat least one processor coupled to said diagnostic subsystem and to said input means;wherein the auditory test instrument further comprises a housing containing the diagnostic subsystem, the peripheral electronic device or system is external to the housing, and the wireless network subsystem is configured to transmit the first signals to the peripheral electronic device or system that is external to the housing. 2. The auditory test instrument of claim 1, wherein said at least one probe is detachably coupled to said auditory test instrument. 3. The auditory test instrument of claim 1, wherein said at least one probe is an intelligent probe. 4. The auditory test instrument of claim 1, wherein said power supply is rechargeable. 5. The auditory test instrument of claim 1, wherein said power supply is rechargeable via a contactless recharging system. 6. The auditory test instrument of claim 1, wherein said wireless networking subsystem is further adapted to receive second signals via said short distance wireless network from said peripheral electronic device. 7. The auditory test instrument of claim 1, wherein said wireless networking subsystem is a Bluetooth enabled wireless networking subsystem and wherein said peripheral electronic device or system is a Bluetooth enabled electronic device or system. 8. The auditory test instrument of claim 1, wherein said peripheral electronic device or system is a LAN system. 9. The auditory test instrument of claim 1, wherein said peripheral electronic device or system is a device selected from the group of devices consisting of computers, personal digital assistants, printers, facsimile devices, and cellular telephones. 10. The auditory test instrument of claim 1, wherein said diagnostic subsystem is modular and detachable. 11. The auditory test instrument of claim 1, wherein said diagnostic subsystem is configured to perform at least one middle-ear test. 12. The auditory test instrument of claim 1, wherein said diagnostic subsystem is configured to perform at least one test selected from the group of tests consisting of tympanometry, tympanometry sweeps, acoustic reflex, acoustic reflex threshold determination, reflex decay measurements, pure tone thresholds, Eustachian tube function, acoustic reflex latency, Gelle's test, outer ear acoustic impedance ... | GN ReSound A/S | H04R025/00 A61B005/00
US08261428 | B2 | 2012-09-11 | Method for assembling a 3-dimensional microelectrode structure | 2009-11-25 | The present invention discloses a method for assembling a 3D microelectrode structure. Firstly, 2D microelectrode arrays are stacked to form a 3D microelectrode array via an auxiliary tool. Then, the 3D microelectrode array is assembled to a carrier chip to form a 3D microelectrode structure. The present invention uses an identical auxiliary tool to assemble various types of 2D microelectrode arrays having different shapes of probes to the same carrier chip. Therefore, the method of the present invention increases the design flexibility of probes. The present invention also discloses a 3D microelectrode structure, which is fabricated according to the method of the present invention and used to perform 3D measurement of biological tissues. | 1. A method for assembling a 3-dimensional microelectrode structure, comprising the steps of: Step a: fabricating 2-dimensional microelectrode arrays, wherein each of said 2-dimensional microelectrode arrays includes a base, a plurality of probes, a plurality of alignment members and a plurality of positioning members, and wherein said probes, said alignment members and said positioning members are connected to said base, and wherein each of said probes has at least one electrode, and wherein said at least one electrode is electrically connected to at least one of the plurality of the alignment members via a corresponding wire;Step b: stacking said 2-dimensional microelectrode arrays to form a 3-dimensional microelectrode array, wherein said alignment members of said 2-dimensional microelectrode arrays are neatly arranged at predetermined positions; andStep c: assembling said 3-dimensional microelectrode array to a carrier chip to form the 3-dimensional microelectrode structure, wherein said alignment members are electrically connected to a plurality of electric-connection pads of said carrier chip correspondingly. 2. The method for assembling a 3-dimensional microelectrode structure according to claim 1, wherein each of the plurality of the alignment members has an exposed electric-connection member; said at least one electrode is connected to corresponding said electric-connection member via the corresponding wire; in said Step c, said electric-connection members are respectively electrically connected to corresponding said electric-connection pads of said carrier chip. 3. The method for assembling a 3-dimensional microelectrode structure according to claim 2, wherein said carrier chip has a plurality of slots; in said Step c, said positioning members are accommodated in said slots, whereby said alignment members are aligned to contact said electric-connection pads, and said electric-connection members are electrically connected to said electric-connection pads. 4. The method for assembling a 3-dimensional microelectrode structure according to claim 1, wherein said 2-dimensional microelectrode arrays are made of a material selected from a group consisting of a glass, PDMS (Polydimethylsiloxane), PI (polyimide), and a thick SU-8 photoresist. 5. The method for assembling a 3-dimensional microelectrode structure according to claim 1, wherein said at least one electrode is made of a material selected from a group consisting of carbon nanotubes, iridium oxide, platinum, gold, titanium, platinum black, and an electric-conduction polymer PEDOT (polyethylenedioxythiophene). 6. The method for assembling a 3-dimensional microelectrode structure according to claim 1, wherein said alignment members are a silicon structure. 7. The method for assembling a 3-dimensional microelectrode structure according to claim 1, wherein said alignment members of different 2-dimensional microelectrode arrays are alternately arranged. 8. The method for assembling a 3... | National Tsing Hua University | G01R003/00
US08167811 | B2 | 2012-05-01 | Apparatus and method for outputting heart sounds | 2011-01-11 | An apparatus for outputting heart sounds includes an implantable system and an external system. The implantable system includes a sensor for generating sensed signals representing detected heart sounds, an interface circuit and a control circuit for receiving the sensed signals, generating data representing the heart sounds therefrom, and transmitting the data to the external system via the interface circuit. The external system includes an interface circuit for communicating with the implantable system, and a control circuit for receiving the data representing the heart sounds and for generating control signals that cause an output device to generate outputs representing the sounds. The implantable system may also include a sensor(s) for detecting cardiac electrical signals. In this case, outputs representing the cardiac electrical signals are also output. | 1. An external system for communicating with an implanted system, comprising: an interface circuit configured to communicate with the implanted system;an output device configured to generate outputs representative of heart sounds detected by the implantable system in response to control signals, the outputs including visual display of the heart sounds; anda control circuit, coupled to the interface circuit and the output device, to receive data representative of the heart sounds from the implanted system and generate the control signals, the control circuit configured to apply the data representative of the heart sounds to a first processing path configured for machine detection of the heart sounds using the data representative of the heart sounds and a second processing path configured for the visual display of the heart sounds. 2. The external system of claim 1, wherein the interface circuit is configured to communicate with the implanted system using optical signals. 3. The external system of claim 1, wherein the output device comprises an audio device configured to generate audio outputs representative of the heart sounds. 4. The external system of claim 1, wherein the first processing path comprises a first band-pass filter, a rectifier, a low-pass filter, and a first ensemble averager, and the second processing path comprises a second band-pass filter and a second ensemble averager. 5. The external system of claim 4, wherein the control circuit is configured to trigger the first ensemble averager and the second ensemble averager using detection of systole. 6. The external system of claim 1, wherein the output device is configured to generate outputs representative of the heart sounds and one or more cardiac electrical signals in response to the control signals, and wherein the control circuit is configured to receive the one or more cardiac electrical signals and generate the control signals causing the output device to generate the outputs representative of the heart sounds and the one or more cardiac electrical signals. 7. The external system of claim 6, wherein the output device is configured to generate outputs representative of the heart sounds and a surface electrocardiograph (ECG) signal in response to the control signals, and wherein the control circuit is configured to receive the surface ECG signal and generate the control signals causing the output device to generate the outputs representative of the heart sounds and the surface ECG signal. 8. The external system of claim 7, wherein the control circuit is configured to generate timing comparison control signals causing the output device to output timing comparison information indicative of relative timing between the heart sounds and surface ECG events. 9. The external system of claim 8, wherein the control circuit is configured to generate the timing comparison control signals causing the output device to display calipers for indicating time intervals between the heart s... | Cardiac Pacemakers, Inc. | A61B005/02
US07967762 | B2 | 2011-06-28 | Ultra miniature pressure sensor | 2007-01-04 | A method of measuring blood pressure and velocity proximally and distally of a stenosis in a vessel carrying blood includes the steps of providing a guide wire having both a pressure sensor and a velocity sensor disposed on a distal region of the guide wire, introducing the guide wire into the vessel, advancing the guide wire to position the pressure sensor and the velocity sensor proximally and distally of the stenosis, and measuring the blood pressure and velocity proximally and distally of the stenosis with the pressure sensor and the velocity sensor. | 1. An apparatus for measurement of a plurality of flow-related characteristics comprising: a guidewire having an outside diameter of 0.018″ or less;a first sensor disposed in a distal region of the guidewire for measuring a first characteristic of flow;a second sensor disposed in a distal region of the guidewire for measuring a second characteristic of flow, wherein the first sensor and second sensor are configured for measuring the first characteristic and the second characteristic simultaneously;a sensor housing containing at least one of the first and second sensors, the sensor housing having an external wall, a lumen, and an opening extending through the external wall to the lumen;a first coil disposed proximally from the sensor housing; anda second coil disposed distally from the sensor housing. 2. The apparatus of claim 1, wherein the first sensor is a pressure sensor. 3. The apparatus of claim 1, wherein the second sensor is a velocity transducer. 4. The apparatus of claim 1, further comprising a sensor housing having an external wall, a lumen, and a hole in the external wall extending from an external surface of the sensor housing to the lumen and being open to ambient fluid. 5. The apparatus of claim 4, wherein the sensor housing is disposed near a distal extremity of the guidewire. 6. The apparatus of claim 1, wherein the first sensor comprises a solid state pressure sensor. 7. The apparatus of claim 6, wherein the solid state pressure sensor has a pressure sensitive region located at or near one end thereof. 8. The apparatus of claim 6, wherein the solid state pressure sensor apparatus is mounted within the sensor housing such that the pressure sensitive region projects into the lumen of the sensor housing without contacting the external wall of the sensor housing and such that a portion of the pressure sensitive region is disposed opposite the opening and is in fluid communication with ambient fluid via the opening. 9. The apparatus of claim 6, wherein the solid state pressure sensor is mounted within the sensor housing using an elastic medium. 10. The apparatus of claim 9, wherein the elastic medium is an epoxy. 11. The apparatus of claim 7, wherein the pressure sensitive region includes a pressure sensitive diaphragm. 12. The apparatus of claim 1, wherein the solid state pressure sensor comprises a diaphragm structure supported by a base plate. 13. The apparatus of claim 1, further comprising a core wire extending from a proximal extremity of the guidewire, through the sensor housing, and to the a distal extremity of the guidewire. 14. The apparatus of claim 13, wherein the core wire comprises a tapered distal extremity. 15. The apparatus of claim 4, wherein the lumen is filled with viscous fluid. 16. The apparatus of claim 15, wherein the viscous fluid comprises an oil. 17. The apparatus of claim 15, wherein the opening has a diameter configured to hold the viscous fluid within the lumen and prevent bleeding of the ... | Volcano Corporation | A61B005/02
US07967760 | B2 | 2011-06-28 | Respiration detecting means | 2006-04-24 | The present application is directed to a detector to detect or generate measuring signals that are indicative of the respiration of a person, wherein said detector can be used advantageously in view of hygiene and stands out due to a simple and robust configuration. According to a first aspect, the detector detects a signal that is indicative of the activity of the respiratory muscles of a person to be examined, the detector comprising a band which, in the application position, is passed around a torso region which widens and narrows when the person to be examined breathes, as well as a structure that is included in the band and loaded in accordance with the narrowing or widening of the torso region. The structure is configured such that it causes a change in the volume of a measuring space device depending on a tensile force. It is thus advantageously possible to detect or measure the forces acting on the band in a non-electric manner and, on the basis of this detection or measurement, draw conclusions relating to the widening or narrowing of the torso. | 1. A detecting means for detecting a signal that is indicative of the activity of the respiratory muscles of a person to be examined, said detecting means comprising: a band means for widening and narrowing in response to the respiration of the person to be examined, the band means being passable around a torso region of the person to be examined; anda structure which is included in the band means and is loaded by said band means in accordance with narrowing and widening of the torso region caused by the respiration of the person to be examined,wherein said structure is configured such that it causes a change in the volume of a measuring space means for receiving the load,wherein the change in the volume is caused by squeezing or widening an elastically deformable chamber at a tube portion thereof,wherein said tube portion is squeezed or widened by the structure in accordance with tensile force applied to the band means caused by the respiration of the person to be examined, andwherein walls surrounding the tube portion have a geometry such that deformation conditions are obtained which, in turn, cause the squeezing or widening of the tube portion depending on the tensile force applied to the band means. 2. The detecting means according to claim 1, wherein the structure is configured as a buckle. 3. The detecting means according to claim 2, wherein the buckle is located adjacent to the tube portion. 4. The detecting means according to claim 2, wherein the buckle directly causes the squeezing or widening of the tube portion in accordance with the tensile force applied to the band means. 5. The detecting means according to claim 1, further comprising first and second straps adapted to fit over the person's shoulders. 6. The detecting means according claim 2, wherein the buckle is divided into two segments, the two segments surrounding the tube portion for squeezing or widening it depending on tensile force applied to the band means. 7. A measuring arrangement comprising the detecting means according to claim 1, the detecting means being configured to measure tensile force signals, said measuring arrangement further comprising: a measuring tube and a measuring device which is coupled to the detecting means via the measuring tube for recording the tensile force signals measured by the detecting means. 8. The detecting means according to claim 1, wherein the detecting means forms part of a multi-channel measuring system. 9. The detecting means according to claim 1, wherein a change in the volume depending on tensile force applied to the band means is passed on to a measuring device via a measuring tube arrangement. 10. The detecting means according to claim 1, wherein the measuring device comprises a signal recording device. 11. The measuring arrangement according to claim 7, wherein the measuring device is configured such that its signal recording capacity extends over a pressure signal range which also comprises the dynamic pressure signal... | MAP Medizin-Technologie GmbH | A61B005/08
US08187199 | B2 | 2012-05-29 | Myocardial contractile reserve measured during activities of daily living | 2010-10-13 | A contractile reserve indicator, corresponding to a predicted value of a maximum change in myocardial contractility that can be achieved by a subject, can be determined using a detected indication of cardiac contractility across various different physical activity levels. | 1. A system comprising: a physical activity sensor, configured to detect a physical activity signal from a subject;a cardiac contractility sensor, configured to detect an indication of cardiac contractility from the subject;a processor, coupled to the cardiac contractility sensor and the physical activity sensor, the processor configured to determine a contractile reserve indicator using a characteristic of a function of the indication of cardiac contractility across various different activity levels, wherein the contractile reserve indicator corresponds to a predicted value of a maximum change in contractility that can be achieved by the subject; anda memory, coupled to the processor and configured to receive and store the contractile reserve indicator to be provided to a process or a user. 2. The system of claim 1, wherein the processor is configured to use the physical activity signal to define time periods corresponding to different steady-state activity levels, and wherein the processor is configured to determine the contractile reserve indicator using a characteristic of a function of the indication of cardiac contractility across the different steady-state activity levels. 3. The system of claim 1, wherein the cardiac contractility sensor is configured to detect the indication of cardiac contractility using at least one of: a pulse pressure, a rate of change of blood pressure during a cardiac ejection period, a rate of change of intraventricular blood pressure during a cardiac ejection period, a S1 heart sound magnitude, a pre-ejection period, or a change in cardiac stroke. 4. The system of claim 1, wherein the processor is configured to classify the subject as one of (1) a therapy responder or (2) a therapy non-responder using the contractile reserve indicator. 5. The system of claim 1, comprising: a dyssynchrony sensor configured to detect an indication of cardiac dyssynchrony from the subject, the dyssynchrony sensor coupled to the processor;wherein the processor is configured to: determine a dyssynchrony indicator using the indication of cardiac dyssynchrony; anddetermine a diagnostic indicator using both the dyssynchrony indicator and the contractile reserve indicator; andwherein the memory is coupled to receive and store the diagnostic indicator to be provided to a user or process. 6. The system of claim 5, wherein the dyssynchrony sensor is configured to detect an indication of cardiac dyssynchrony using at least one of an electrocardiogram, an intracardiac impedance, an indication of heart wall motion, or a magnitude of a heart sound. 7. The system of claim 1, comprising a therapy circuit, coupled to the processor, the therapy circuit configured to provide a therapy to the subject; and wherein the processor is configured to control the therapy provided by the therapy circuit using information about the contractile reserve indicator. 8. The system of claim 7, wherein the therapy circuit includes at least one of a cardiac resy... | Cardiac Pacemakers, Inc. | A61B005/02
US07942825 | B2 | 2011-05-17 | Method and device for monitoring thermal stress | 2008-06-09 | A method and device for monitoring thermal stress in a user is described. The device is designed to include a material having specific thermodynamic properties and physical dimensions defined as a function of those thermodynamic properties. A system for thermal stress monitoring including a thermal stress monitoring device configured within a garment is also described. | 1. A method for continuous non-invasive monitoring of thermal stress of a user, the method comprising the steps of: a) providing a device comprising, at least one first temperature sensor configured to be placed in thermal contact with skin of the user,at least one second temperature sensor, andan insulating layer positioned between the first and second temperature sensor,where the insulating layer comprises an outer periphery and a thermal conductance (Lband),where the thermal conductance (Lband) of the insulating layer satisfies the following equation1+2⁢Lband2Lcore2+2⁢LbandLcore≤10where Lcore is between 36 and 140 W/m2K, andwhere the temperature sensors are located at least a distance (w) from a peripheral edge of the insulating layer, where the distance (w) satisfies the following equationw=0.0436⁢Dcore⁢72090+1101⁢Qevap+118⁢Qevap⁢LbandLbandwhere Qevap is 250 W/m2 or less, andwhere Dcore is between 0.002 and 0.005 metersb) placing said device in thermal contact with the skin of the user;c) determining heat flux from the skin through the device at known time intervals; andd) determining core temperature and change in core temperature as a function of heat flux. 2. The method of claim 1, further comprising the subsequent step of comparing the determined core temperature to a threshold temperature range. 3. The method of claim 2, further comprising the subsequent step of providing an alert when the determined core temperature is outside the threshold temperature range. 4. The method of claim 1, further comprising the subsequent step of comparing the determined change in core temperature to a threshold temperature change range. 5. The method of claim 4, further comprising the subsequent step of providing an alert when the determined change in core temperature is outside the threshold temperature range. 6. The method of claim 5, where the alert is provided to the user. 7. The method of claim 6, where the alert is selected from a group consisting of audible alert, visual display, vibration, and a combination thereof. 8. The method of claim 1, where the step of placing said device in contact with the skin of the user comprises placing said device in contact with the skin of the user proximate to a temporal artery of the user. 9. The method of claim 1, where the step of placing said device in contact with the skin of the user comprises affixing said device to the skin of the user. 10. The method of claim 1, where the thermal conductance (Lband) of the insulating layer is between 230 and 910 W/m2 K, and where the distance (w) is between 3.5 cm and 7 cm. 11. A device for continuous non-invasive monitoring thermal stress of a user, the device comprising: at least one first temperature sensor configured to be placed in thermal contact with skin of the user,at least one second temperature sensor, andan insulating layer positioned between the first and second temperature sensor,where the insulating layer comprises an outer periphery and a thermal ... | Kimberly-Clark Worldwide Inc. | A61B005/00 G01K017/00 G01K003/00 G01K001/02 G01K005/70 G01K005/00 G01K009/00
US07957794 | B2 | 2011-06-07 | System and method for transducing, sensing, or affecting vaginal or body conditions, and/or stimulating perineal musculature and nerves using 2-way wireless communications | 2009-08-17 | A system for sensing vaginal conditions or for affecting vaginal or body conditions in a mammal. The system includes a portable probe unit and a separate controller unit, both of which include two-way communication devices to both transmit signals to, and receive signals from, each other, wirelessly and in real time. The substantially smooth and sealed outer surface of the cylindrical body of the probe unit has a rounded end and is dimensioned to permit comfortable and repeated insertion into, removal from, and containment entirely within a mammal's vagina. The probe body has a device substantially flush with its outer surface and configured to deliver electrical pulses, receive electrical pulses, sense or sample vaginal conditions and/or deliver medication. Signals to the probe unit comprise control and programming signals acting as a feedback loop configured to start, stop and/or alter activity of the substantially flush device of the probe unit, either automatically or manually, for sensing vaginal conditions or for affecting vaginal or body conditions. | 1. A system for sensing vaginal conditions or for affecting vaginal or body conditions in a mammal, comprising: a portable probe unit, said probe unit comprising a substantially cylindrical body having a substantially smooth and substantially sealed outer surface with a rounded end and so dimensioned as to permit comfortable and repeated insertion into, removal from, and containment entirely within a mammal's vagina; means substantially flush with the outer surface of the body of the probe and configured to accomplish at least one of delivering electrical pulses, receiving electrical pulses, sensing or sampling vaginal conditions, and delivering medication; a programmable microprocessor; a memory; a battery; and two-way communication means with antenna and adapted to both transmit signals to a controller unit and receive signals from said controller unit wirelessly and in real time, even when said probe unit is contained entirely within a mammal's vagina; andthe controller unit comprising two-way communication means adapted to both receive signals from said probe unit and transmit signals to said probe unit wirelessly and in real time, even when said probe unit is contained entirely within a mammal's vagina, wherein said signals to said probe unit comprise control and programming signals acting as a feedback loop configured to start, stop, and/or alter the activity of said substantially flush means of said probe unit, either automatically or manually, for sensing vaginal conditions or for affecting vaginal or body conditions. 2. A system according to claim 1, wherein said substantially flush means of said probe unit comprises means for sampling cervical fluid or other changes in the vaginal environment. 3. A system according to claim 1, wherein said substantially flush means of said probe unit comprises means for sensing at least one of temperature, pH, secretion viscosity, vaginal pathogens and atypical cervical cells. 4. A system according to claim 1, wherein said substantially flush means of said probe unit includes at least one of a conductive band and a sensor transducer. 5. A system according to claim 1, wherein said probe unit is adapted to be programmed to start, stop and/or alter at least one of sensing vaginal conditions, sampling vaginal conditions and delivering medication after a predetermined period of time. 6. A system according to claim 1, wherein said probe unit is less than one inch in diameter and less than four inches long. 7. A system according to claim 1, wherein said probe unit is provided with means for facilitating removal from a vagina. 8. A system according to claim 1, wherein said system is adapted to permit a human user to operate said controller unit while said probe unit is in the human user's vagina. 9. A system according to claim 1, wherein said two-way communication means of at least one of said probe unit and said controller unit includes means for transmitting signals to and/or receiving signals from an... | Athena Feminine Technology | A61B005/04 A61B005/05 A61B005/00 A61B010/00 A61B005/103 A61B005/117
US08660626 | B2 | 2014-02-25 | System and method for mitigating interference in pulse oximetry | 2011-02-04 | A pulse oximetry sensor adapted to emit light from an emitter proximate to a patient's tissue and detect a portion of the emitted light on a detector proximate to the tissue. The pulse oximetry system is adapted to acquire position data for the emitter and for the detector with one or more position indicators. The pulse oximetry system is adapted to process the position data to obtain a change in position of the emitter and the detector; and to process pulse oximetry measurements using the change in position to obtain a motion-corrected pulse oximetry data. | 1. A pulse oximetry system comprising: a sensor comprising: a sensor body;an emitter disposed on the sensor body;a detector disposed on the sensor body; andone or more position indicators that generate a change in voltage or resistance in response to changes in position of the emitter and the detector; anda monitor comprising: a processor configured to:receive a signal from the detector and a signal from the one or more position indicators;determine a corrected path length based on the signal from the one or more position indicators; anddetermine a physiological parameter based at least in part upon the signal from the detector and the corrected path length. 2. The pulse oximetry system of claim 1, wherein the one or more position indicators comprise a first position indicator disposed laterally relative to the emitter and a second position indicator disposed laterally relative to the detector, wherein the first position indicator and the second position indicator are each separated from the respective emitter and detector by known, fixed distances. 3. The pulse oximetry system of claim 1, wherein the one or more position indicators comprise a first contact point and a second contact point of a linear resistor of a linear potentiometer device. 4. The pulse oximetry system of claim 3, wherein a digit, when placed within the sensor body, acts as a variable linear resistor between the first contact point and the second contact point. 5. The pulse oximetry system of claim 1, wherein the one or more position indicators comprise a first contact point and a second contact point of a potentiometer device that measures bending or flex of a substrate based on changes in resistance between the first contact point and the second contact point. 6. The pulse oximetry system of claim 1, wherein the one or more position indicators comprise one or more Hall Effect devices that generate a voltage in response to movement of a magnet relative to the Hall Effect devices. 7. A pulse oximetry monitor, comprising: one or more processors configured to: derive relative positions of an emitter and a detector of a pulse oximetry sensor using non-optical data generated by one or more position indicators disposed on a sensor body of the pulse oximetry sensor;determine a path length of light from the emitter and the detector based on the relative positions; anddetermine a physiological parameter based on a physiological data signal from the sensor and a relationship between light absorbance or light scatter probabilities as a function of the path length. 8. The pulse oximetry monitor of claim 7, wherein the one or more processors are configured to calculate correction factors or parameters for the physiological data signal acquired using the pulse oximetry sensor based on the derived relative positions. 9. The pulse oximetry monitor of claim 7, wherein the one or more processors are configured to derive at least a distance between the emitter and the detector based on ... | Covidien LP | A61B005/1455
US09179827 | B2 | 2015-11-10 | Systems and methods for determining the position and orientation of medical devices inserted into a patient | 2010-12-02 | A medical device system includes an elongated body with a distal end that is configured and arranged for insertion into a patient. A housing is disposed in the distal end of the body. A rotatable magnet is disposed in the housing. At least one magnetic field winding is configured and arranged to generate a magnetic field at the location of the magnet. The magnetic field causes rotation of the magnet at a target frequency. An array of magnetic field sensors is disposed external to the patient. The magnetic field sensors are configured and arranged to sense the location and orientation of the magnet in relation to the array of magnetic field sensors. | 1. A medical device system comprising: an imaging pill configured and arranged for ingestion by a patient;a housing disposed within the pill;a rotatable magnet disposed in the housing and configured to rotate freely relative to the housing;at least one magnetic field winding, the magnetic field winding configured and arranged to generate a magnetic field at the location of the magnet, the magnetic field causing rotation of the magnet at a target frequency;an array of magnetic field sensors disposed external to the patient, the magnetic field sensors configured and arranged to sense the location and orientation of the magnet in relation to the array of magnetic field sensors;electronics for capturing at least one image;a power generation system, the power generation system comprising a plurality of layers of wire disposed over the housing;an electronic circuit configured and arranged for storing power generated in the layers of wire;one or more input leads coupling the plurality of layers of wire to the electronic circuit; andone or more output leads coupling the electronic circuit to the electronics for capturing at least one image. 2. The medical device system of claim 1, wherein the electronic circuit further comprises a rectifier for converting AC power to DC power. 3. The medical device system of claim 1, wherein the electronic circuit further comprises a filter. 4. The medical device system of claim 1, wherein the magnet is spherical. 5. The medical device system of claim 1, wherein the housing is spherical. | Boston Scientific Scimed, Inc. | A61B005/00 A61B001/04 A61B005/06 A61B005/07 A61B019/00
US08942676 | B2 | 2015-01-27 | Systems and devices for emergency tracking and health monitoring | 2009-11-06 | The present disclosure provides for systems, devices, and methods which address needs of seniors, disabled persons, or any other similar users, at a stage in life when they can still live independently while benefiting from monitored care. Embodiments described herein enable a higher level of connectivity without requiring the member to be technologically savvy, and provides a simple mechanism for personalized emergency support. Embodiments utilize at least one personal communication device which may allow for remote interactions with medical and concierge personnel in a monitoring center or other emergency contacts, in order to fulfill requests for everyday needs, ranging from providing directions to a locations, to health alerts based on health and status measurements collected electronically and remotely. | 1. A monitoring and communication system comprising: a companion device communicatively coupled to a cellular network, said companion device configured to establish two-way voice and data communication with a monitoring center via said cellular network, wherein said monitoring center includes a help desk application in communication with a data services application and a telephony application, said companion device further configured to allow for communication between a user and at least one other designated source via said cellular network;a user device, said user device configured to be disposed proximate to said user and configured to communicate with said companion device, said user device further configured to actively monitor at least one user condition and to report said condition to said monitoring center through said companion device, wherein said companion device initiates two-way voice communication with the monitoring center when said user condition fall outside a predetermined range, and wherein said condition is reported to the help desk application by the data services application, and the two-way voice communication is established with the help desk application using the telephony application; andan emergency button on the monitoring and communication system, the emergency button operable by the user to initiate a two-way voice and data communication with a monitoring center via said cellular network, wherein said companion device is configured to transmit location information for at least one of said companion device and said user device automatically when an emergency contact is initiated between said companion device and said monitoring center. 2. The system of claim 1 further comprising at least one fall detection device. 3. The system of claim 2 wherein said companion device is configured to contact said monitoring center upon receiving a signal indicating that said fall detection device has registered a user emergency. 4. The system of claim 1 further comprising at least one health sensor selected from the group consisting of: blood pressure sensor, blood-oxygen sensor, blood-sugar level sensor, and pulse sensor. 5. The system of claim 1 wherein said companion device and said user device are configured to receive configuration instructions from said monitoring center. 6. The system of claim 1 wherein said companion device is configured to receive GPS location information and to forward said GPS location information to said monitoring center. 7. The system of claim 1 wherein said user device connects to said companion device using a Bluetooth-compliant communication device. 8. The system of claim 1 wherein said user device connects to said companion device using a ZigBee-compliant communication device. 9. The system of claim 1 wherein said companion device is configured to communicate with said monitoring center using GPRS protocol. 10. The system of claim 1 wherein said companion device is configured to communicate ... | ActiveCare, Inc. | H04B007/00 A61B005/00 G08B025/00 A61B005/11 G06F019/00 A61B005/021 A61B005/024 A61B005/145
US08838227 | B2 | 2014-09-16 | Portable automatic brain state assessment apparatus | 2011-02-01 | A method and apparatus for performing rapid brain assessment may provide emergency triage to head trauma patients by analyzing a combination of spontaneous and evoked brain potentials. The spontaneous and evoked potentials are analyzed, and the results classified, to present a real-time assessment of a patient's brain, diagnosing any potential abnormalities therein. | 1. An apparatus for determining the neurological state of a subject, comprising: a set of electrodes configured to be placed on the subject to collect brain electrical signals;a processor;a memory operatively coupled to the processor, wherein the memory stores one or more operating instructions; andan input-output interface having at least one channel operatively coupled to the processor, whereinthe input-output interface is configured to receive brain electrical signals through the set of electrodes; andthe processor is configured to utilize the one or more operating instructions to perform one or more operations on the signals received through the input-output interface, wherein the one or more operations comprise:de-noising the received signals;extracting features from the de-noised signals;classifying the extracted features; anddetermining a neurological state of the subject based on the classification, wherein determining the neurological state comprises the steps of:determining if the neurological state is normal or abnormal; anddetermining if an abnormal neurological state is psychiatric or “functional” in nature, organic in nature, or an emergency or “Alert” condition. 2. The apparatus of claim 1, further comprising: a display device, wherein the display device is operatively connected to the processor and the processor is further configured to display results of the one or more operations on the display device. 3. The apparatus of claim 1, further comprising: a bi-directional communications channel operatively connected to the input-output interface and to the processor to communicate the received signals. 4. The apparatus of claim 1, further comprising: a low voltage preamplifier having at least one channel operatively connected to the input-output interface. 5. The apparatus of claim 4, wherein the low voltage preamplifier has a high noise tolerance. 6. The apparatus of claim 1, wherein the input-output interface includes an electroencephalograph testing interface and an auditory evoked potentials testing interface; and wherein the processor is configured with the one or more stored operating instructions for auditory tests selected from the group comprising electroencephalograph test procedures, auditory evoked potentials test procedures, and combinations thereof. 7. The apparatus of claim 1, wherein the processor has an internal memory, and the processor is configured to utilize the internal memory to execute one or more operating instructions for processing the received signals. 8. The apparatus of claim 1, wherein the extracted features are classified using discriminant analysis. 9. The apparatus of claim 1, wherein the instructions for denoising the signals include instructions to filter the received signals. 10. The apparatus of claim 9, wherein the filtering comprises: using a dual-tree complex wavelet transform of the received signals. 11. The apparatus of claim 1, wherein the set of electrodes comprises: a plurality ... | BrainScope Company, Inc. | A61B005/04 A61B005/048 A61B005/0484 A61B005/00
US09451897 | B2 | 2016-09-27 | Body adherent patch with electronics for physiologic monitoring | 2010-12-02 | In one configuration, an adherent device to adhere to a skin of a subject includes a stretchable base layer having an upper side and a lower side and an adhesive coating on the lower side to adhere the base layer to the skin of the subject. The base layer has at least two openings extending therethrough, each of the at least two openings having a size. The adherent device also includes a stretchable covering layer positioned above and adhered to the base layer with an adhesive to define at least two pockets. The adherent device also includes at least two gels, each gel having a size larger than the size of openings to retain the gel substantially within the pocket, and a circuit carrier supported with the stretchable base layer to measure at least one physiologic signal of the subject. Other configurations and methods are also claimed. | 1. An adherent device to adhere to a skin of a subject, comprising: a stretchable base layer having an upper side and a lower side and an adhesive coating on the lower side to adhere the base layer to the skin of the subject, the base layer having at least two openings extending therethrough, each of the at least two openings having a size; a stretchable covering layer positioned above and adhered to the base layer with an adhesive to define at least two pockets, wherein the stretchable covering layer is thin, flexible, and configured to stretch with the skin of the subject; at least two gels, wherein each gel is positioned within one of the corresponding pockets, each gel having a size larger than the size of the at least two openings to retain said gel substantially within said corresponding pocket; a flexible circuit that includes at least two electrodes in contact with the at least two gels, the flexible circuit including a first portion located on the upper side of the stretchable base layer and a second portion that extends away from the first portion and through an opening in the stretchable covering layer; a circuit carrier positioned above the stretchable covering layer and supported with the stretchable base layer to measure at least one physiologic signal of the subject, wherein the circuit carrier is connected to the at least two electrodes via the second portion extending through an opening in the stretchable covering layer; and a compliant connection that includes a plurality of bridging loops formed between the upper side of the stretchable base layer and the circuit carrier that permits at least some movement of the circuit carrier in a plane parallel to the stretchable base layer. 2. An adherent device to monitor a subject having a skin, comprising: a stretchable base layer having an upper side and a lower side and an adhesive coating disposed on the lower side to adhere the base layer to the skin of the subject, the base layer having at least two openings extending therethrough, each opening having a size; a flexible circuit having at least two electrodes disposed thereon, each electrode positioned with a respective one of the at least two openings to couple to the skin of the subject the flexible circuit including a first portion located adjacent to the upper side of the stretchable base layer and a second portion that extends away from the stretchable base layer through an opening in a stretchable covering layer; at least two gels positioned with the at least two openings in the base layer, each gel having a size larger than the size of said each opening; the stretchable covering layer positioned above the at least two gels and adhered to the base layer, such that each gel is constrained substantially within a corresponding pocket disposed between the base layer and the covering layer, wherein the stretchable covering layer is thin, flexible, and configured to stretch with the skin of the subject; a circuit carrier positioned... | Medtronic Monitoring, Inc. | A61B005/04 A61B005/0408 A61B005/0432 A61B005/00
US20110118570 | A1 | 2011-05-19 | OPTIC SENSOR DEVICE WITH SERS | 2009-03-31 | The present invention relates to a system for determining a concentration of a substance of interest in a body fluid, more specifically to systems comprising sensor devices for optically detecting compounds such as glucose. The system of the present invention is, e.g., suitable for continuous determination of a concentration of a substance of interest which is present in a body fluid, e.g. a concentration of glucose in blood or interstitial fluid. | 1.-21. (canceled) 22. A system for determining a concentration of a substance of interest in a body fluid, the system comprising: a sensor device adapted to be positioned in direct contact with a body fluid to be analysed, said sensor device defining an analysis volume, at least partly delimited towards the body fluid by a semi-permeable membrane allowing substances of interest from the body fluid to enter the analysis volume,a light source adapted to transfer primary light to the sensor device,wherein the sensor device comprises at least one substantially flat area constituting a first area and at least one substantially flat area constituting a second area, at least one area comprising an optically responsive element,wherein groups of areas are provided, each group of areas forming a common level and the common levels having a predetermined spacing from each other to provide the said difference in distance over which light interacts with compounds and tissue and wherein the first group of areas is formed at a base level of the device and the second group of areas is formed by projections from the said base level to a top level. 23. The system according to claim 22, wherein at least one of the first or second substantially flat areas, or at least one member of the first and second group of areas, comprises an optically responsive element comprising any number and combination of a SERS, a photo detector and a reflective surface. 24. The system according to claim 22, further comprising first light guiding means arranged to guide primary light from the light source to the analysis volume. 25. The system according to claim 24, further comprising second light guiding means arranged to guide secondary light away from the analysis volume. 26. The system according to claim 22, wherein the sensor device forms part of a probe head adapted to be positioned invasively, a semi-permeable membrane being arranged in a wall part of the probe head. 27. The system according to claim 26, wherein at least one of the first and second light guiding means is positioned with an end face in contact with the internal of the probe head. 28. A device according to claim 22, wherein at least a part of at least one of said areas is covered by a permeable membrane forming an analysis volume. 29. A device according to claim 22, wherein the sensor device is positioned/implanted under the skin of an animal. 30. The device according to claim 29 wherein the animal is a human, the sensor device being positioned/implanted under the skin of a human being. 31. The device according to claim 29, wherein the light source and possibly first and/or second light guiding means are external to the animal/human body, all transmitted and/or reflected light penetrating the skin of the animal/human body. 32. The system according to claim 22, further comprising at least one laser for emitting primary light, said laser(s) being connected to the first light guiding means. 33. The system ac... | P & V Consulting GmbH & Co. KG | A61B005/1455
US20110125040 | A1 | 2011-05-26 | WIRELESS ECG MONITORING SYSTEM | 2009-03-04 | An ECG monitoring system for ambulatory patients includes a reusable battery-powered ECG monitor with an electrode attached to a patient for receiving ECG signals. A processor analyzes the received ECG signals for predefined arrhythmia. If an arrhythmia is detected, a wireless transceiver in the ECG monitor transmits the event information and an ECG strip to a cellphone handset. The cellphone handset automatically relays the event information and ECG strip to a monitoring center for further diagnosis and necessary intervention. The ECG monitor sends notifications to the cellphone handset whenever the status of the monitor changes. The cellphone handset forwards status notifications to which a patient response has not been received to the monitoring center, as well as status notifications generated by the cellphone handset. | 1. (canceled) 2. (canceled) 3. A wireless cardiac monitoring system comprising: a battery powered ECG monitor adapted to receive cardiac signals from an electrode attached to a patient, the ECG monitor including a transceiver for wirelessly transmitting ECG signal data, the ECG monitor further including a status notifier, coupled to the transceiver, for transmitting a status notification when a status of the ECG monitor changes; anda cellphone handset including cellphone electronics operable for communication over a cellular network and a transceiver for wirelessly receiving ECG signal data or status notifications from the ECG monitor and a controller operable to cause the cellphone electronics to send ECG signal data received from the ECG monitor to a monitoring center over the cellular network,wherein the controller is further operable to respond to a status notification from the ECG monitor by at least one of a) bringing a status notification to the attention of the patient with the cellphone handset, or b) sending the status notification to the monitoring center over the cellular network, wherein the controller is further operable to delay sending a given status notification to the monitoring center to allow time for receipt of a rectifying status notification. 4. The wireless cardiac monitoring system of claim 3, wherein a rectifying status notification comprises a status notification providing information that an undesired status notified by a previous notification has been corrected. 5. The wireless cardiac monitoring system of claim 4, wherein the undesired status notification comprises at least one of a notification that an electrode has become loose or that the ECG monitor is not attached to a patient. 6. The wireless cardiac monitoring system of claim 3, wherein the controller is further operable to immediately send every status notification, except those previously determined to be delayed, to the monitoring center. 7. The wireless cardiac monitoring system of claim 6, wherein the controller is further operable to send a delayed status notification to the monitoring center if a rectifying status notification has not been received in a predetermined amount of time. 8-11. (canceled) 12. A wireless cardiac monitoring system comprising: a battery powered ECG monitor adapted to receive cardiac signals from an electrode attached to a patient, the ECG monitor including a transceiver for wirelessly transmitting ECG signal data, the ECG monitor further including a status notifier, coupled to the transceiver, for transmitting a status notification when a status of the ECG monitor changes; anda cellphone handset including cellphone electronics operable for communication over a cellular network and a transceiver for wirelessly receiving ECG signal data or status notifications from the ECG monitor and a controller operable to cause the cellphone electronics to send ECG signal data received from the ECG monitor to a monitoring center over t... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/0402
US20110144528 | A1 | 2011-06-16 | System for circadian rhythm monitor with synchrony and activity planning | 2009-12-11 | A personal health system which includes a suitable Core Body Temperature (CBT) monitor that can be worn for all or part of a 24 hour day and collect continuous CBT data. The CBT data is collected and compared to determine circadian desynchrony. A conveniently carried or worn processor/display unit, in communication with the CBT monitor, algorithmically determines activity types and activity timing based on the collected CBT data to improve synchrony. The activities and when to perform them are displayed to the user. | 1. A Personal Health Care system, comprising; a wearable Core Body Temperature (CBT) monitor, including at least one sensor, power supply and communications link, and;a processor and display provision including a programmable controller, data storage, a display and at least one communications link to the CBT monitor,wherein;CBT data is collected and compared to at least one of normative or previously stored CBT data, and deviations from predetermined desirable circadian CBT synchrony and oscillation amplitude are detected and analyzed algorithmically to determine activity type and timing to restore circadian behavior, whereby advice to perform the activities at the determined time is displayed. 2. The system of claim 1 wherein the CBT monitor includes a sensor mounted in the ear. 3. The system of claim 2 wherein the sensor includes at least one of a thermopile, a thermistor, or a multiple sensor arrangement of thermopiles and thermistors to provide improved signal to noise. 4. The system of claim 1 wherein the communication link between the CBT monitor and the processor and display is wireless, including standard wireless protocols such as Bluetooth and Zigbee. 5. The system of claim 1 wherein the processor and display is a standard personal appliance including smartphones and PDAs, executing a program on a phone processor for the circadian data collection, algorithms, and display functions, and the monitor communicates with the appliance directly. 6. The system of claim 1 wherein the processor and display are web-based, and the monitor communicates by accessing webpages over an internet gateway. 7. The system of claim 1 wherein the processor and display is a purpose-built device and at least one of contains the complete user interface, analysis algorithms and display, or serves as a gateway to a PC or the internet for at least part of the user interface, analysis algorithms and display functions. 8. A computerized method for analyzing and displaying circadian synchrony information for an individual, comprising; acquiring continuous or semi-continuous CBT data from an individual in normal daily activity situations,comparing at least one of; real time CBT, CBT over an interval, and CBT transition intervals and amplitude, to normative CBT circadian data,displaying the acquired data to allow visual comparison to normative data; and,providing suggested activities to reduce the desynchrony between the individual's actual CBT data and normative. 9. The method of claim 8 further comprising scheduling the suggested activities and alerting the individual when the activities are due. 10. The method of claim 8 further comprising selecting a source for normative CBT data. | - | A61B005/01
US08251911 | B2 | 2012-08-28 | Monitoring physiological condition and detecting abnormalities | 2010-11-30 | A system and method for monitoring an individual's physiological condition and detecting abnormalities therein, comprising concurrently receiving an electrocardiograph signal and a ballistocardiograph signal. The electrocardiograph and ballistocardiograph signals are conditioned to minimize background extraneous noise after which, each signal is concurrently processed and analyzed to detect repeating cyclical patterns and further characterized to identify individual components of the repeating cycles. At least one individual component in one signal is selected as a reference marker or a selected component in the other signal. The two signals are then synchronized, outputs produced therefrom and stored in a database. | 1. A method for determining one or more parameters characteristic of an individual's cardiac function, said method comprising: acquiring electrocardiograph and ballistocardiograph signals from said individual;converting said electrocardiograph and ballistocardiograph signals into electrocardiograph data and ballistocardiograph data;receiving said electrocardiograph data and ballistocardiograph data at a computer, said computer including a microprocessor;synchronizing, by said microprocessor, said electrocardiograph data with said ballistocardiograph data to provide at least one synchronized electrocardiograph-ballistocardiograph waveform;identifying, by said microprocessor, one or more components of said synchronized electrocardiograph-ballistocardiograph waveform, said components indicative of said one or more parameters characteristic of an individual's cardiac function. 2. The method according to claim 1, wherein identifying one or more components of the synchronized electrocardiograph-ballistocardiograph waveform comprises detecting and marking one or more of the h, i, j, a, a1, g, H, I, J, K, L, M and N peaks on said synchronized electrocardiograph-ballistocardiograph waveform. 3. The method according to claim 1, wherein said one or more parameters are one or more of: time period between any two of the h, i, j, a, a1, g, H, I, J, K, L, M and N peaks, time duration of the H-I wave, H-I slope, I-J slope, J-K slope, K-L slope, L-M slope, M-N slope, amplitude of the H peak, amplitude of the I peak, amplitude of the J peak, amplitude of the K peak, amplitude of the L peak, amplitude of the M peak, amplitude of the N peak, amplitude of the j peak, amplitude of the a1 peak and j-a slope. 4. The method according to claim 1, wherein said one or more parameters are one or more of: stroke volume, cardiac output, ending diastolic volume, ending systolic volume, ventricular ejection fraction, cardiac output index, pre-ejection period, cardiac performance index and isovolumetric contraction time. 5. The method according to claim 1, wherein converting said electrocardiograph and ballistocardiograph signals into electrocardiograph data and ballistocardiograph data comprises converting, by an analog-digital converter, analog electrocardiograph and ballistocardiograph signals into digital electrocardiograph and ballistocardiograph data. 6. The method according to claim 1, wherein said ballistocardiograph signals are acquired from an accelerometer. 7. The method according to claim 6, wherein said accelerometer is positioned on the sternum of said individual. 8. The method according to claim 1, wherein synchronizing said electrocardiograph data with said ballistocardiograph data includes aligning an R peak of at least one electrocardiograph waveform corresponding to said electrocardiograph data with an H peak of at least one ballistocardiograph waveform corresponding to said ballistocardiograph data. 9. The method according to claim 1, further comprisin... | Heart Force Medical Inc. | A61B005/02
US07944551 | B2 | 2011-05-17 | Systems and methods for a wavelet transform viewer | 2008-10-10 | Techniques for the display of a signal with a wavelet transform of that signal in a wavelet transform viewer are disclosed, according to embodiments. According to embodiments, the wavelet transform viewer can display a plot of physiological signals such as a photoplethysmograph (PPG) signal. A portion of the plot of the signal can be selected. A wavelet transform the selected portion of the signal can be calculated and a wavelet plot of the tranformed signal can be displayed simultaneously with that signal. A plot of the selected portion of the signal can also be simultaneously displayed with both the plot of the signal and the wavelet plot. | 1. A method for displaying a signal in a wavelet transform viewer, the method comprising: receiving a signal in a processor;receiving a user selection of a portion of the signal corresponding to an area of interest via control input;transforming the selected portion of the signal using a wavelet transform using the processor to generate a transformed signal;generating a wavelet plot based on the transformed signal using the processor; anddisplaying the wavelet plot simultaneously with a plot of the selected portion of the signal on a display. 2. The method of claim 1 wherein the signal comprises a photoplethysmographic (PPG) signal. 3. The method of claim 1 wherein the wavelet plot comprises a continuous wavelet scalogram of the transformed signal. 4. The method of claim 1 wherein the wavelet plot comprises a plot of a phase of the transformed signal, a real part of the transformed signal, an imaginary part of the transformed signal, and/or a cross-correlation between the transformed signal and a second signal, and/or combinations thereof. 5. The method of claim 1 wherein the selected portion of the signal is adjustable. 6. The method of claim 1 further comprising displaying a plot of the received signal simultaneously with the plot of the selected portion of the signal and the wavelet plot. 7. The method of claim 1 further comprising: receiving a selection of a second portion of the signal;transforming the selected second portion of the signal using a wavelet transform to generate a further transformed signal;regenerating the wavelet plot based on the further transformed signal; andreplacing the displayed wavelet plot with the regenerated wavelet plot. 8. The method of claim 1 further comprising displaying at least one other signal plot with the wavelet plot and the plot of the selected portion of the signal. 9. The method of claim 1 further comprising displaying a wavelet plot of at least one other signal. 10. The method of claim 1 further comprising displaying the wavelet plot with the plot of the selected portion of the signal within a polysomnograph (PSG) display. 11. A wavelet transform viewer comprising: a processor capable of: receiving a signal;receiving a user selection of a portion of the signal corresponding to an area of interest;transforming the selected portion of the signal using a wavelet transform to generate a transformed signal;generating a wavelet plot based on the transformed signal; anda display capable of displaying the wavelet plot simultaneously with a plot of the selected portion of the signal. 12. The wavelet transform viewer of claim 11 wherein the signal comprises a photoplethysmographic (PPG) signal. 13. The wavelet transform viewer of claim 11 wherein the wavelet plot comprises a continuous wavelet scalogram of the transformed signal. 14. The wavelet transform viewer of claim 11 wherein the wavelet plot comprises a plot of a phase of the transformed signal, a real part of the transformed signal, an ima... | Nellcor Puritan Bennett Ireland | A61B005/1455
US08666672 | B2 | 2014-03-04 | System and method for interpreting a user's psychological state from sensed biometric information and communicating that state to a social networking site | 2009-11-21 | The present invention is a mood sensing and communicating system and method that includes a biometric sensor configured to sense biometric information from a person. A mood interpretive system is supported on a computing device or on a remote server. The mood interpretive system is in communication with the biometric sensor and is configured to assign a psychological or emotional mood state with the sensed biometric information. A communication system is supported on the computing device and is configured to communicate the assigned psychological mood state across an Internet to a web-site for positing on a web page. The method includes sensing biometric information from a user with a biometric sensor, associating a psychological mood state with the sensed biometric information with a mood interpretive module, and communicating the assigned psychological mood state to a web-site supported on a server across an Internet. | 1. A method comprising: selecting a primary biometric parameter and a secondary biometric parameter, the primary and secondary biometric parameters being two different parameters selected from the group consisting of body temperature, skin temperature, heart rate, blood oxygen levels, and skin moisture;creating a first linear correlation between biometric information of the primary biometric parameter and an ordered range of primary mood states;creating a second linear correlation between biometric information of the secondary biometric parameter and an ordered range of secondary mood states;receiving primary biometric information for the primary biometric parameter from a first sensor;interpreting the primary biometric information using the first linear correlation, thereby determining a primary mood state of a person;receiving secondary biometric information for the secondary biometric parameter from a second sensor;interpreting the secondary biometric information using the second linear correlation, thereby determining a secondary mood state of the person;using a processor to determine a specific psychological mood state of the person, which comprises the primary mood state and secondary mood state;communicating the specific psychological mood state to a server. 2. The method of claim 1, further comprising: selecting a tertiary biometric parameter, the tertiary biometric parameter being a different parameter from the primary biometric parameter and the secondary biometric parameter, and further being selected from the group consisting of body temperature, skin temperature, heart rate, blood oxygen levels, and skin moisture;creating a third linear correlation between biometric information of the tertiary biometric parameter and an ordered range of tertiary mood states;receiving tertiary biometric information for the tertiary biometric parameter from a third sensor; andinterpreting the tertiary biometric information using the third linear correlation, thereby determining a tertiary mood state of the person;and wherein the specific psychological mood state of the person further comprises the tertiary mood state. 3. The method of claim 1, further comprising periodically repeating the steps of sensing primary biometric information, interpreting the primary biometric information, sensing secondary biometric information, interpreting the secondary biometric information, determining a specific psychological mood state, and communicating the specific psychological mood state. 4. The method of claim 1, wherein the ordered range of primary mood states or the ordered range of secondary mood states comprise one or more of happy, sad, relaxed, angry, stressed, tired, excited, or combinations thereof. 5. The method of claim 1, further comprising allowing manual adjustment of the first linear correlation or the second linear correlation. 6. The method of claim 1, further comprising allowing manual adjustment of the specific psychological mood state. 7. ... | Radial Comm Research L.L.C. | A61B005/16 G06F019/00 G06F003/01 A61B005/01 A61B005/024 A61B005/145 A61B005/1455 A61B005/053
US08290574 | B2 | 2012-10-16 | ECG monitoring system with configurable alarm limits | 2009-03-04 | An ECG monitoring system for ambulatory patients includes a small multi-electrode patch that adhesively attaches to the chest of a patient. A reusable battery-powered ECG monitor clips onto the patch and receives patient electrical signals from the electrodes of the patch. A processor continuously processes received ECG signals and analyzes the received ECG signals for pre-defined arrhythmia. If an arrhythmia is detected, a wireless transceiver in the ECG monitor transmits the event information and an ECG strip to a monitoring center. At the commencement of a monitoring procedure a message is sent to the monitoring center and configuration information for arrhythmia detection is down-loaded and installed in the monitor. The configuration file is determined by a screen of selectable standard and custom arrhythmias and alarm limits at the monitoring center. | 1. A method for configuring a cardiac monitoring system for detection of a predetermined arrhythmia of an ambulatory patient, the cardiac monitoring system including a battery powered ECG monitor worn by the patient and having a patient ECG signal processor, arrhythmia detector, and a wireless transceiver for sending status messages and receiving arrhythmia detector configuration information, the cardiac monitoring system further including a cellphone handset having cellphone electronics operable for communication with a monitoring center over a cellular network, a transceiver for wirelessly communicating with the ECG monitor, and a controller operable to cause the cellphone handset to forward status messages received from the ECG monitor to the monitoring center and to forward to the ECG monitor configuration information received from the monitoring center, the method comprising: determining, prior to cardiac monitoring of a patient, a parameter of a predetermined arrhythmia which is to be detected and storing the parameter in a configuration file at the monitoring center;attaching the ECG monitor to the patient to begin ECG monitoring;sending a status message from the ECG monitor to the monitoring center in response to the commencement of ECG monitoring;responding to the status message by sending the configuration file to the ECG monitor; andusing the configuration file to configure the arrhythmia detector. 2. The method of claim 1 wherein determining further comprises determining one or more alarm limits for one or more arrhythmias which are to be detected by the arrhythmia detector. 3. The method of claim 2, wherein determining further comprises: displaying on a computer display a plurality of selectable arrhythmias with selectable alarm limits. 4. The method of claim 3, wherein the alarm limits comprise one of more of beats per minute (BPM), time, or priority. 5. The method of claim 3, wherein displaying further comprises displaying a custom alarm which may be selected and configured by an operator. 6. The method of claim 1, further comprising: checking for an updated configuration file each time a message is sent to the monitoring center during ECG monitoring. 7. The method of claim 1, wherein using further comprises: detecting an arrhythmia when an arrhythmia identified in the configuration file meets or exceeds an alarm limit for the arrhythmia identified in the configuration file; andsending an event notification message to the monitoring center when an arrhythmia is detected. 8. The method of claim 7, wherein sending an event notification message further comprises: sending a concurrent ECG strip to the monitoring center. 9. The method of claim 1, further comprising: after an ECG monitoring procedure has commenced, displaying on a computer display a plurality of selected arrhythmias with selected alarm limits of the configuration file for the procedure;changing one or more of a selected arrhythmia or a selected alarm limit of the... | Koninklijke Philips Electronics N.V. | A61B005/04
US07962192 | B2 | 2011-06-14 | Systems and methods for aligning a tool with a desired location or object | 2006-04-28 | Systems and methods for aligning a tool positioned on a moveable arm with a location or an object of interest on a body surface. Some methods comprise using one or more cameras and identifying position offsets of the location or the object of interest from the one or more cameras. Some methods comprise using at least two cameras and rectifying the images from the first and the second cameras. A variety of different tools and assemblies may be located on the moveable arm for performing various diagnostic and therapeutic procedures that take advantage of the ability of the moveable arm to rapidly and precisely position the respective tool or assembly at desired locations on a skin. | 1. A method for aligning a tool positioned on a moveable arm of an automated system with a location or an object of interest on a skin, comprising: determining a fixed relationship between the tool positioned on the moveable arm and at least two cameras positioned on the movable arm, wherein the tool is capable of manipulating tissue;selecting a location or an object of interest on a skin;identifying position offsets of the location or the object of interest from the at least two cameras;automatically moving the moveable arm to align the tool with the location or the object of interest on the skin based, at least in pail, on the identified offsets and on maintaining while aligning the tool the fixed relationship between the at least two cameras and the tool. 2. The method of claim 1, wherein identifying position offsets comprises determining whether the at least two cameras are aligned with the location or object of interest within acceptable tolerances, andmoving the moveable arm, if the at least two cameras are not aligned with the location or object of interest within acceptable tolerances, until the at least two cameras are aligned with the location or object of interest within acceptable tolerances. 3. The method of claim 1, further comprising determining whether the tool is aligned with the location or the object of interest within acceptable tolerances, and repeating the acts of identifying position offsets and moving the movable arm, if the tool is not aligned with the location or object of interest within acceptable tolerances, until the tool is aligned with the location or object of interest within acceptable tolerances. 4. The method of claim 1, further comprising continually calculating and adjusting position and orientation of the moveable arm as the tool is moved into alignment with the location or the object of interest. 5. The method of claim 1, wherein the object of interest has certain rotational orientation in 3D, and wherein the method comprises identifying 3D orientation offsets of the object of interest. 6. The method of claim 5, wherein the object of interest is a follicular unit and the tool is a hair removal tool, and wherein aligning the tool comprises aligning an elongate axis of the hair removal tool with an elongate axis of the follicular unit. 7. The method of claim 1, wherein the location or the object of interest is selected from the group comprising an injection site, a tattoo placement or removal site, a tissue removal site, a follicular unit, a wrinkle line, a wart, and a mole. 8. The method of claim 6, further comprising operating the hair removal tool to remove the follicular unit. 9. The method of claim 1, wherein the tool comprises an elongated axis and each of the at least two cameras comprises a respective axis, and wherein the tool and the at least two cameras are positioned such that their respective axes are parallel to each other. 10. The method of claim 1, wherein the tool comprises an elonga... | Restoration Robotics, Inc. | A61B005/103
US07967749 | B2 | 2011-06-28 | Monitoring system and method using rules | 2003-07-23 | A method of monitoring a subject, such as a patient, using rules is provided. The method includes receiving data from the subject, and applying the data to a rule-based algorithm. The method also preferably includes generating a response based on the application of the rule-based algorithm. Some examples of responses include generating an alarm, generating a suggested reason for an identification of an abnormal condition, and generating a suggested response to an abnormal condition that has been identified. Preferably, more than one set of variables is applied to a particular rule-based algorithm. A monitor using the method may be configured to have multiple rule-based algorithms selected and applied simultaneously. Each rule-based algorithm can generate a unique response, or a combination can be used to generate a smaller number of responses. Also, a monitor using this method may be configured to allow rule-based algorithms to be exchanged and/or altered. | 1. A patient physiologic monitoring assembly comprising: a plurality of sensors that generate a real-time physiologic data stream, said real-time physiologic data stream including a plurality of physiologic variables;a first logic rule set for diagnosing a physiological condition of a patient, the first logic rule set including a plurality of logic rules for interpreting the plurality of physiologic variables of the real-time physiologic data stream;a second logic rule set for diagnosing the physiological condition of the patient, the second logic rule set including a plurality of logic rules for interpreting the physiologic variables of the real-time physiologic data stream; anda controller that receives said real-time physiologic data stream as the data stream is generated, said controller including a logic that:simultaneously cross references said plurality of physiologic variables of the real-time physiologic data stream with the first logic rule set and second logic rule set as the controller receives the real-time physiologic data system;generates at least a first diagnostic interpretation of said plurality of physiologic variables utilizing said first logic rule set and a second diagnostic interpretation of said plurality of physiologic variables utilizing said second logic rule set; anddisplays said first and second diagnostic interpretations on a display element;wherein the first diagnostic interpretation and the second diagnostic interpretations are different diagnostic interpretations of the same physiological condition, using the same real-time physiologic data stream. 2. A patient physiologic monitoring assembly as described in claim 1, wherein said logic is further adapted to select said first logic rule set and said second logic rule set from a rules database, said rules database including a plurality of logic rule sets. 3. A patient physiologic monitoring assembly as described in claim 2, wherein said logic is further adapted to modify one of said plurality of logic rules within said first logic rule set. 4. A patient physiologic monitoring assembly as described in claim 3, wherein said logic edits one of said plurality of logic rules. 5. A patient physiologic monitoring assembly as described in claim 3, wherein said logic deletes one of said plurality of logic rules. 6. A patient physiologic monitoring assembly as described in claim 3, wherein said logic adds a new logic rule to said first logic rule set. 7. A patient physiologic monitoring assembly as described in claim 2, wherein said logic is further adapted to add a new logic rule set to said rules database. 8. A patient physiologic monitoring assembly as described in claim 1, further comprising a plurality of networked medical facilities in communication with said controller such that said first logic rule set may be received from any of said plurality of networked medical facilities. 9. A method for providing diagnostic aid to a clinician monitoring the medical cond... | GE Medical Systems Information Technologies, Inc. | A61B005/00
US07939093 | B2 | 2011-05-10 | Method and apparatus for representing myocardial tissues in different states of damage | 2007-08-06 | The invention relates to a method for differentially representing myocardial tissue in different states of damage, comprising the following steps: administering a myocardium-suitable contrast agent to a patient under examination; entering at least one patient-specific parameter affecting the speed of uptake by and elimination from the myocardium of said contrast agent; calculating a point in time after administration of the contrast agent at which a difference between a contrast agent content in necrotic myocardial tissue and a contrast agent content in non-necrotic myocardial tissue attains a maximum value, on the basis of the at least one patient-specific parameter, and carrying out, at the point in time calculated, a late-phase CT scan for accentuation of necrotic myocardial tissue compared to non-necrotic myocardial tissue. The invention likewise relates to apparatus, in particular for carrying out the method. | 1. A method for representing a myocardial tissue of a patient, comprising: administering a contrast agent to the patient;inputting a patient-specific parameter that relates to a speed of perfusing the contrast agent by a myocardial tissue and eliminating the contrast agent from the myocardial tissue;calculating a time at which a difference between a content of the contrast agent in a necrotic myocardial tissue and in a non-necrotic myocardial tissue attains a maximum value based on the patient-specific parameter;performing a late-phase CT scan for detecting the necrotic myocardial tissue compared to the normally perfused myocardial tissue at the calculated time; andrepresenting the necrotic myocardial tissue and the normally perfused myocardial tissue for medically examining the patient,wherein the maximum value of the difference in the content of the contrast agent is calculated by:Sm,n,o⁡(t)=m·ⅇ(t+tconst)-n·(t+tconst)owherein: S(t) is a time-dependent approximation function for a response of a contrast agent signal intensity versus time,m, n and o are constants derived from the patient-specific parameters,tconst is a time after which the contrast agent has been flushed out of the normally and the less perfused myocardial tissues, and t is the time with which the contrast agent signal intensity varies and is zero at tconst. 2. The method as claimed in claim 1, further comprising performing an early-phase CT scan for detecting the normally perfused myocardial tissue compared to a less perfused myocardial tissue and the necrotic myocardial tissue immediately after administering the contrast agent. 3. The method as claimed in claim 2, wherein a subtraction data record is calculated from the early-phase CT scan and the late-phase CT scan for detecting the less perfused myocardial tissue. 4. The method as claimed in claim 2, wherein the early-phase CT scan and the late-phase CT scan are displayed on a display device. 5. The method as claimed in claim 1, wherein the patient-specific parameter is selected from the group consisting of: pulse, blood pressure, respiration, ventricular volume, cardiac output, amount and type of the contrast agent, body size, and weight. 6. The method as claimed in claim 1, wherein a contour of areas of the necrotic myocardial tissue is extracted from the late-phase CT scan data. 7. The method as claimed in claim 1, wherein t is varied until S(t) is in a typical grayscale range in an image of the necrotic myocardial tissue. 8. The method as claimed in claim 1, wherein S(t) is approximately in a middle of a grayscale range of an image of the myocardial tissue. 9. An apparatus for determining a time for performing a late-phase CT scan of a myocardial tissue of a patient, comprising: an inputting device that inputs a patient-specific parameter related to a speed of perfusing a contrast agent by a myocardial tissue and eliminating the contrast agent from the myocardial tissue;a computing unit that calculates a time at w... | Siemens Aktiengesellschaft | A61B005/05
US07970459 | B2 | 2011-06-28 | Sleepiness level detection method and apparatus | 2009-05-29 | A sleepiness level detection method of determining a sleepiness level of a subject, includes: detecting a heart beat signal indicating a signal of heart beats from the subject and storing the heart beat signal in a storage device; detecting amplitude peaks of the heart beat signal and detecting an interval between the amplitude peaks as a heart beat interval; and computing a spectral density corresponding to fluctuation in the heart beat interval detected and determining the sleepiness level based on a maximum frequency at which the spectral density is maximum. | 1. A sleepiness level detection method of determining a sleepiness level of a subject, comprising: detecting a heart beat signal indicating a signal of heart beats from the subject and storing the heart beat signal in a storage device;detecting amplitude peaks of the heart beat signal and detecting an interval between the amplitude peaks as a heart beat interval; andcomputing a spectral density corresponding to fluctuation in the heart beat interval detected and determining the sleepiness level based on a maximum frequency at which the spectral density is maximum. 2. The sleepiness level detection method according to claim 1, wherein the determining includes: classifying maximum frequencies at which the spectral density is maximum into a plurality of clusters;determining a representative value of the maximum frequencies included in each cluster; anddetermining the sleepiness level based on change in the representative value and the spectral density corresponding to a frequency of the representative value. 3. The sleepiness level detection method according to claim 2, wherein the determining of the sleepiness level includes determining that the sleepiness level has decreased if the maximum frequency decreases and/or a magnitude of the spectral density corresponding to the maximum frequency increases. 4. The sleepiness level detection method according to claim 3, wherein the determining of the representative value includes generating a histogram of the maximum frequencies stored in each cluster and determining the representative value based on the histogram. 5. The sleepiness level detection method according to claim 1, further comprising generating a graph indicating fluctuation in the maximum frequency, and outputting the graph generated. 6. A non-transitory computer readable storage medium containing instructions that, when executed by a computer, cause the computer to perform, for determining a sleepiness level of a subject: detecting a heart beat signal indicating a signal of heart beats from the subject and storing the heart beat signal in a storage device;detecting amplitude peaks of the heart beat signal and detecting an interval between the amplitude peaks as a heart beat interval; andcomputing a spectral density corresponding to fluctuation in the heart beat interval detected and determining the sleepiness level based on a maximum frequency at which the spectral density is maximum. 7. The non-transitory computer readable storage medium according to claim 6, wherein the determining includes: classifying maximum frequencies at which the spectral density is maximum into a plurality of clusters;determining a representative value of the maximum frequencies included in each cluster; anddetermining the sleepiness level based on change in the representative value and the spectral density corresponding to a frequency of the representative value. 8. The non-transitory computer readable storage medium according to claim 7, wherein the determi... | Fujitsu Limited | A61B005/04
US08444621 | B2 | 2013-05-21 | Sampler cap | 2011-02-24 | A sampler cap includes a sampler connector, an analyzer connector and a liquid impermeable closure member, which may be a membrane or a filter. The sampler cap may be used to transfer a test sample to an analyzer without removing the sampler cap from the sampler. | 1. A sampler cap for capping an end of a sampler and for venting the sampler, the sampler cap comprising: a hollow sampler cap body including a sampler connector for coupling to the sampler;venting means disposed inside of the hollow sampler cap body and configured to vent the sampler with a sample; anda closure member closing the hollow body, the closure member having a weakening zone through which an analyzer collects the sample from the sampler via an analyzer probe, wherein the weakening zone allows penetration of the analyzer probe through the closure member at a penetration spot and provides air passage adjacent to the penetration spot. 2. The sampler cap according to claim 1, wherein the weakening zone comprises a longitudinal slit. 3. The sampler cap according to claim 2, wherein the analyzer probe has a diameter, and wherein the longitudinal slit is longer than the diameter of the analyzer probe. 4. The sampler cap according to claim 2, wherein the closure member has a thickness and the longitudinal slit passes entirely through the closure member thickness. 5. The sampler cap according to claim 2, wherein the closure member has a thickness and the longitudinal slit passes partly through the closure member thickness. 6. The sampler cap according to claim 1, wherein the closure member comprises a flexible membrane. 7. The sampler cap according to claim 1, wherein the venting means comprises a venting conduit for venting air from the sampler to the surroundings, the venting conduit comprising a liquid impermeable filter positioned so that air expelled from the sampler passes through the filter to the surroundings and liquid expelled from the sampler through the venting conduit is trapped by the filter. 8. The sampler cap according to claim 7, wherein the closure member comprises a flexible membrane, and wherein the flexible membrane flexes between a first position and a second position, wherein, when the flexible membrane is in the first position, the venting conduit is closed, thereby preventing passage of fluid from the sampler connector and along the venting conduit; andwherein, when the flexible membrane is in the second position, the venting conduit is open, thereby permitting passage of fluid from the sampler connector and along the venting conduit. 9. The sampler cap according to claim 7, wherein the venting conduit comprises at least one venting chamber for accommodating liquid expelled from the sampler between the sampler connector and the liquid impermeable filter. 10. The sampler cap according to claim 1, wherein the sampler cap comprises at least one collection chamber for collecting liquid expelled from the sampler upon penetration of the analyzer probe. 11. The sampler cap according to claim 1, further comprising an analyzer connector for engaging the sampler cap with the analyzer for sample collection. 12. The sampler cap according to claim 11, wherein the analyzer connector comprises air passage means for equilibr... | Radiometer Medical ApS | A61B019/00
US08965485 | B2 | 2015-02-24 | Integrated surgical cutting system | 2008-11-03 | Integrated surgical systems and methods for using same are provided that can receive scanned images and produce a three dimensional model from the images. Based at least in pan on the three dimensional model, a processor generates code defining an optimized tool path, which is sent to a surgical machining system that can machine the desired portion of the patient. In one exemplary aspect, the integrated system operates in a clearly defined and pre-programmed manner with no necessity for in-situ sensory or guidance feedback. | 1. An integrated surgical system, comprising: a scanning device configured to scan a selected portion of a patient's body and to generate a plurality of two-dimensional images of the selected portion;a surgical machining system comprising a machine tool and means for selectively moving the machine tool relative to a plurality of axes; anda processor operatively connected to the scanning device and the surgical machining system, the processor configured to: receive two-dimensional images from the scanning device;generate a three-dimensional image of the selected portion of the patient's body from the received two-dimensional images,display the three-dimensional image,receive input from the user, the input indicating a location of the selected portion to be machined during surgery and at least one machining property,generate an optimized tool path based at least in part on the input and at least in part on geometry of the three-dimensional image of the selected portion of the patient's body, andtransmit the optimized tool path to the surgical machining system to effect movement of the machine tool in accordance with the optimized tool path in an automated manner; wherein the processor is configured to update the optimized tool path based on detected differences between actual properties of the selected portion of the patient's body and expected properties of the three-dimensional image of the selected portion. 2. The integrated surgical system of claim 1, wherein the step of displaying further comprises displaying at least one of the two-dimensional images to a user. 3. The integrated surgical system of claim 1, wherein the surgical machining system comprises a variable high speed machine. 4. The integrated surgical system of claim 1, wherein the surgical machining system comprises an automated computer numerical controlled surgical machining system. 5. The integrated surgical system of claim 1, wherein the machining property comprises at least one of a shape, a size, and an orientation. 6. The integrated surgical system of claim 5, wherein the machining property is selected from the group consisting of a machining axis, a machine hole radius, and a machine hole depth. 7. The integrated surgical system of claim 1, wherein the scanning device comprises a computed tomography (CT) device. 8. The integrated surgical system of claim 1, wherein the scanning device comprises a magnetic resonance imaging (MRI) device. | University of Utah Research Foundation | A61B005/05 G05B019/402 A61B019/00 A61B017/32
US08934959 | B2 | 2015-01-13 | Physiological pharmacokinetic analysis for combined molecular MRI and dynamic PET imaging | 2009-06-22 | Magnetic resonance (MR) imaging typically has excellent spatial resolution, but relatively poor temporal resolution. In contrast, positron emission tomography (PET) typically has excellent temporal resolution, but poor spatial resolution relative to MR. Resultantly, it is advantageous to use combined PET-MR imaging sequences to create hybrid or enhanced images that reap the benefits of both modalities. A contrast agent (80) that includes both a PET tracer (82) and MR contrast enhancement (86) can be used in such a combined modality setting. The contrast agent (80) also includes a targeting system (84) that allows the contrast agent (80) to pool in a region of interest. | 1. A diagnostic imaging apparatus for contrast enhanced imaging of a patient, the apparatus comprising: a magnetic resonance imaging apparatus including: a main magnetic coil assembly for generating a main magnetic field in an imaging region of the apparatus,a gradient coil assembly for overlaying gradient magnetic fields on the main magnetic field, andan RF coil assembly for at least transmitting RF signals into the imaging region and receiving magnetic resonance signals from the imaging region;a nuclear PET imaging apparatus including: a detector array configured to detect nuclear decay events concurrently with the receiving of the magnetic resonance signals by the RF coil assembly, andat least one reconstruction processor for producing an image representation from the nuclear decay events detected by the detector array that gather in the imaging region and from the detected magnetic resonance;a common PET and MRI agent configured for injection into a patient disposed in the imaging region, the common agent including: a carrier including a fluorine-19 component which is imaged by the magnetic resonance imaging apparatus and a radiopharmaceutical which includes fluorine-18 which decays to generate the nuclear decay events detected by the PET imaging apparatus, the fluorine-18 being imaged concurrently with the imaging of the fluorine-19 component by the magnetic resonance imaging apparatus,a targeting system that takes advantage of biological processes to cause the common agent to collect in a target region, andan MRI contrast agent that is imaged by the magnetic resonance imaging apparatus. 2. The diagnostic imaging apparatus as set forth in claim 1, wherein the common agent is able to pass through a capillary endothelium, but unable to pass into a cell. 3. The diagnostic imaging apparatus as set forth in claim 1, wherein the targeting system includes at least one of a monoclonal antibody that attaches to a certain type of cell wall and a fibrin-specific paramagnetic nanoparticle, and the common agent is able to pass through a capillary endothelium, but unable to pass into a cell. 4. A method of diagnostic imaging comprising: injecting a subject with a common PET/MRI agent including: a first component including fluorine-19 that is imaged by a magnetic resonance imaging apparatus,a second component that includes a radiopharmaceutical including fluorine-18 that is imaged by a PET imaging apparatus concurrently with the imaging of the first component by the magnetic resonance imaging apparatus,a targeting system that takes advantage of biological processes to cause the common agent to collect in a target region, andan MRI contrast agent that is imaged by the magnetic resonance imaging apparatus;wherein the common agent includes a carrier including the first and second components;exciting magnetic resonance in selected dipoles of the subject and the first component of the common agent;with the magnetic imaging apparatus, detecting the magnetic r... | Koninklijke Philips N.V. | A61B005/055 G01R033/56 A61K049/18 A61K051/12 G01R033/48
US08473046 | B2 | 2013-06-25 | Diffusion-based magnetic resonance methods for characterizing bone structure | 2011-01-06 | A method of in vitro or in vivo nuclear magnetic resonance and/or magnetic resonance imaging, to determine bone properties by measuring the effects of molecular diffusion inside the bone specimen to derive parameters that are related to the structure of the trabecular bones. The method is a non-invasive probe that provides topological information on trabecular bone without requiring a full high-resolution image of its structure, and is compatible with clinical use. | 1. A method for obtaining a map of a structural property of a bone, the method comprising: obtaining a plurality of diffusion weighted images of a sample of the bone, wherein each diffusion weighted image is obtained using a diffusion time period of a different duration;determining statistical information based on the plurality of diffusion weighted images;extracting at least one bone property from the statistical information so as to provide an indication of bone structure for a plurality of voxels; anddetermining a map of the bone property for at least a portion of the sample of the bone using the plurality of voxels. 2. The method as claimed in claim 1, wherein obtaining a plurality of diffusion weighted images includes measuring diffusion using a pulsed-field gradient pulse sequence followed by an imaging sequence. 3. The method as claimed in claim 1, wherein obtaining a plurality of diffusion weighted images includes measuring diffusion using a decay from diffusion in an internal field pulse sequence followed by an imaging sequence. 4. A method for obtaining a structural property at a specific voxel in a bone, the method comprising: obtaining a plurality of diffusion weighted signals from a sample of the bone using a 3D localization sequence, wherein each diffusion weighted signal is obtained using a diffusion time period of a different duration;determining statistical information based on the plurality of diffusion weighted signals; andextracting at least one bone property from the statistical information so as to provide an indication of bone structure for the voxel. 5. The method as claimed in claim 4, wherein obtaining a plurality of diffusion weighted signals includes measuring diffusion using a pulsed-field gradient pulse sequence. 6. The method as claimed in claim 4, wherein obtaining a plurality of diffusion weighted signals includes measuring diffusion using a decay from diffusion in an internal field pulse sequence. | Schlumberger Technology Corporation | A61B005/04
US08579812 | B2 | 2013-11-12 | System and methods for management of disease over time | 2009-12-15 | A system for delivering health information is provided. The system comprises at least one apparatus configured to receive information related to a condition of a patient. The analysis system further comprises a reception system configured to receive data related to the patient's condition and a processor configured to assess a condition of a patient and to identify a data set including information related to the patient's condition. The system also comprises a communication system configured to communicate the data set to a user. | 1. A method for communicating health information by a communication system comprising an analysis system, implemented by one or more processors, the method comprising: collecting patient information associated with a patient, by one or more reception devices, wherein the patient information comprises physiological information related to a health condition of a patient;processing, by the analysis system, the physiological information to assess the health condition associated with the patient;determining, by the analysis system, a time status of the patient's condition;assigning, by the one or more processors, a major data block to the patient based on the health condition associated with the patient, wherein the major data block includes; information that is identified as relevant to all patients who have the health condition or are at risk for the health condition, andminor data blocks, which are subsets of the major data block and comprise information that is identified as potentially relevant to a subset of the patients who have the health condition or are at risk for the health condition;selecting, by the one or more processors, a plurality of minor data blocks from the major data block based on the time status of the health condition and an intended recipient of the minor data block, wherein each of the plurality of minor data blocks is intended for a different recipient; andcommunicating, by the communication system, each minor data block as a data set to the intended recipient of each minor data block. 2. The method of claim 1, wherein collecting patient information includes performing at least one of either a diagnostic test and entry of data into a data interface. 3. The method of claim 2, wherein performing the diagnostic test includes collecting data related to brain electrical activity. 4. The method of claim 1, further including transmitting the data set to a location distant from the analysis system. 5. The method of claim 1, wherein [selecting at least one of the plurality of minor data blocks includes comparing the health condition of the patient to health conditions of other patients represented by data stored in a database and selecting the minor data block based on the comparison] assignment of the major data block includes classifying the patient condition relative to the condition of other patients represented by data stored in a database. 6. The method of claim 5, wherein the assignment is based on at least one of a stage of disease progression, a life expectancy of the patient, a comorbid condition, a treatment received by the patient, and a genetic profile of the patient's condition. 7. The method of claim 1, wherein the data set communicated to one of the intended recipients includes at least one therapeutic recommendation. 8. The method of claim 1, wherein the data set communicated to one of the intended recipients includes a prediction of a length of time remaining in a current disease stage. 9. The method of cla... | Brainscope Company, Inc. | A61B005/00
US20110160619 | A1 | 2011-06-30 | METHOD AND APPARATUS FOR A SCENTED ALARM CLOCK BASED ON SLEEP STATE | 2009-12-31 | A method and apparatus for affecting the dreams of the subject or individual is presented. The apparatus first predicts or detects the desired sleep state and presents a scent to the subject. The apparatus then awakens the subject after a given period of time using external stimulus. The desired sleep state may be the REM sleep state, and the period of time may be one minute. Upon awakening, the subject will recall dreams that have been affected by the emitted scent. If the scent is pleasant, then the subject will recall having pleasant dreams. | 1. An alarm clock comprising: a sleep state detector coupled to a subject to determine when said subject is in a given sleep state;a scent applied to said subject in said given sleep state;a processor coupled to said sleep state detector to determine when to apply at least one stimulus to awaken said subject. 2. The apparatus of claim 1, whereby said given sleep state is a dream-heavy REM (Rapid Eye Movement) state. 3. The apparatus of claim 1, whereby said stimulus is a way to awaken said subject selected from the group consisting of light, sound, touch and movement. 4. The apparatus of claim 1, whereby said processor calculates a time period. 5. The apparatus of claim 4, whereby said time period is one minute. 6. The apparatus of claim 3, whereby a second stimulus is selected from said group. 7. The apparatus of claim 1, whereby said scent creates undesirable dreams. 8. The apparatus of claim 1, whereby said scent creates pleasant dreams. 9. The apparatus of claim 1, whereby said subject is awoken in a preferred state. 10. A method of waking a subject comprising the steps of: applying a scent to said subject sleeping in a given sleep state;waiting a first time period; andwaking said subject. 11. The method of claim 10, further comprising the steps of: a sleep state detector coupled to said subject to determine when said subject is in said given sleep state. 12. The method of claim 11, wherein said given sleep state is a dream-heavy state. 13. The method of claim 11, further comprising the steps of: a processor coupled to said sleep state detector to calculate when to apply at least one stimulus. 14. The method of claim 13, wherein said waking occurs by the application of said at least one stimulus to awaken said subject. 15. The method of claim 14, wherein said stimulus is a way to awaken said subject selected from the group consisting of light, sound, touch and movement. 16. The method of claim 10, wherein said subject is placed into a preferred state. 17. The method of claim 10, wherein said scent creates pleasant dreams. 18. The method of claim 10, wherein said scent creates undesirable dreams. 19. An alarm clock comprising: a sleep state detector coupled to a subject;a scent being applied to said subject in a sleep state; anda processor coupled to said sleep state detector to determine when to apply at least one stimulus to awaken said subject. 20. The apparatus of claim 19, whereby said sleep state is the dream-heavy REM (Rapid Eye Movement) state. | LCtank LLC | A61B005/11 A61M021/00 G04B047/00
US20110105889 | A1 | 2011-05-05 | NANOPARTICLE LABELING AND SYSTEM USING NANOPARTICLE LABELING | 2009-02-20 | A nanoparticle labeling which is simultaneously usable in the combination of X-ray imaging with optical imaging and the combination of X-ray imaging and optical imaging with magnetic resonance imaging characterized by comprising core/shell type semiconductor nanoparticles having an average shell thickness of 0.1 nm or more but not more than 10.0 nm together with an X-ray sensitive material for the former combination, and core/shell type semiconductor nanoparticles having an average shell thickness of 0.1 nm or more but not more than 10.0 nm together with an X-ray sensitive material and magnetic particles for the latter combination. | 1-16. (canceled) 17. A nanoparticle labeling agent, comprising: core/shell semiconductor nanoparticles each having a core and a shell which covers the core and has an average thickness of 0.1 nm or more and 10.0 nm or less; andan X-ray sensitive material. 18. The nanoparticle labeling agent of claim 17, wherein the X-ray sensitive material contains at least one X-ray sensitive material selected from scandium (Sc), titanium (Ti), vanadium (V), chromium (Cr), iron (Fe), cobalt (Co), nickel (Ni), copper (Cu), zinc (Zn), gallium (Ga), selenium (Se), bromine (Br), rubidium (Rb), strontium (Sr), yttrium (Y), zirconium (Zr), niobium (Nb), molybdenum (Mo), ruthenium (Ru), rhodium (Rh), palladium (Pd), silver (Ag), cadmium (Cd), iridium (Ir), tellurium (Te), and iodine (I), caesium (Cs), barium (Ba), lanthanum (La), cerium (Ce), praseodymium (Pr), neodymium (Nd), promethium (Pm), samarium (Sm), europium (Eu), gadolinium (Gd), terbium (Tb), dysprosium (Dy), holmium (Ho), erbium (Er), thulium (Tm), ytterbium (Yb), lutetium (Lu), hafnium (Hf), a tantalum (Ta), tungsten (W), osmium (Os), platinum (Pt), and gold (Au). 19. The nanoparticle labeling agent of claim 17, further comprising: a substance to combine the core/shell type semiconductor nanoparticles and the X-ray sensitive material. 20. The nanoparticle labeling agent of claim 19, wherein the substance to combine is SiO2. 21. The nanoparticle labeling agent of claim 20, wherein the SiO2 is contained in the shell of the core/shell type semiconductor nanoparticle. 22. The nanoparticle labeling agent of claim 17, further comprising: magnetic particles. 23. The nanoparticle labeling agent of claim 22, wherein the magnetic particles are a superparamagnetic substance, a paramagnetic substance, or a ferromagnetic substance. 24. The nanoparticle labeling agent of claim 23, wherein the superparamagnetic substance, the paramagnetic substance, or the ferromagnetic substance is a metal oxide. 25. The nanoparticle labeling agent of claim 24, wherein the metal oxide is selected from a group consisting of an oxide of cobalt, an oxide of nickel, an oxide of manganese, and an oxide of iron. 26. The nanoparticle labeling agent of claim 25, wherein the oxide of iron is Fe3O4 or γ-Fe2O3. 27. The nanoparticle labeling agent of claim 23, wherein the paramagnetic substance includes a chelated gadolinium complex as a mother substance and one or more kinds of paramagnetic ions contained in a chelating agent. 28. The nanoparticle labeling agent of claim 27, wherein the paramagnetic ions include one or more kinds of manganese (Mn), praseodymium (Pr), neodymium (Nd), samarium (Sm), europium (Eu), terbium (Tb), dysprosium (Dy), holmium (Ho), erbium (Er), thulium (Tm), ytterbium (Yb), and lutetium (Lu). 29. A system for obtaining an in vivo image with high resolution, comprising: a first modality capable of detecting fluorescence of the core/shell type semiconductor particles in the nanoparticle labeling agent of claim 1... | KONICA MINOLTA MEDICAL & GRAPHIC, INC. | A61B005/055 B32B001/00 A61B006/00
US07937133 | B2 | 2011-05-03 | Method for determining size, pathology, and volume of embolic material | 2007-06-13 | Methods for determining the size, pathology, and volume of embolic debris captured in an embolic protection filtering device. The methods may include providing an embolic protection filtering device and scanning the filtering device with a computed tomography scanner. The methods may also include analyzing digital images produced during the scanning step. | 1. A method for quantifying embolic debris, comprising the steps of: advancing an embolic protection filtering device through a body lumen to a position near an area of interest, the device including an elongate shaft and a filter coupled to the shaft;capturing embolic material in the filter;scanning the body lumen with a computed tomography scanner to produce a plurality of digital images; andanalyzing the digital images to calculate the volume of the embolic material captured in the filter. 2. The method of claim 1, wherein the step of scanning the body lumen with a computed tomography scanner to produce a plurality of digital images includes scanning the filter within the body lumen. 3. The method of claim 1, wherein the step of scanning the body lumen with a computed tomography scanner to produce a plurality of digital images includes scanning the filter with a cone beam x-ray computed tomography scanner. 4. The method of claim 3, wherein the digital images produced by the cone beam x-ray computed tomography scanner include a number of pixels. 5. The method of claim 4, wherein the pixels define a visual representation of the filter, space within the filter, or the embolic material disposed within the space. 6. The method of claim 5, further comprising the step of identifying groups of pixels that represent an individual particle of embolic material and determining the size of the individual particle. 7. The method of claim 6, further comprising the step of assessing the pathology of the individual particle. 8. The method of claim 7, wherein the step of assessing the pathology of the individual particle includes determining if the individual particle is a calcified plaque. 9. The method of claim 8, further comprising the step of breaking up the calcified plaque. 10. The method of claim 5, wherein the step of analyzing the digital images to calculate the volume of the embolic material captured in the filter includes using a computer to count the number of pixel representing embolic material on each individual digital image. 11. The method of claim 10, wherein each pixel correlates to a known volume and wherein the number of pixels counted on the digital images is multiplied by the known volume to calculate the volume of the embolic debris. 12. The method of claim 1, further comprising the step of removing the filter from the body lumen. 13. The method of claim 12, wherein the filter has a filtering capacity, and wherein the step of removing the filter from the body lumen occurs when the volume of the embolic debris captured in the filter approaches the filtering capacity. 14. A method for analyzing embolic debris captured in an embolic protection filtering device, comprising the steps of: advancing an embolic protection filtering device through a body lumen to a position near an area of interest, the device including an elongate shaft and a filter coupled to the shaft;capturing embolic material in the filter;scanning the body lumen... | Boston Scientific Scimed, Inc. | A61B005/05
US07945303 | B2 | 2011-05-17 | Noninvasive pulmonary performance measurement method and system | 2005-08-03 | A system and method for measuring a pulmonary performance value of a subject. In one embodiment, a system includes one or more imaging devices, each configured to capture one or more images of a subject inside an interior chamber, such as a plethysmograph chamber. An exemplary system also includes an imaging processor in communication with the one or more imaging devices for estimating a change in volume of the subject from information about the one or more images. | 1. An imaging system for a plethysmograph, the system comprising: one or more imaging devices, wherein each of said one or more imaging devices is operatively configured to capture one or more images of a subject inside the plethysmograph as the subject breathes one or more gases contained within the plethysmograph and operable to provide an output signal including information regarding said one or more images, said one or more imaging devices including at least two imaging devices positioned orthogonal to each other with respect to the subject;a contrast assistance element positioned so as to be opposite at least one of said one or more imaging devices; andan imaging processor in communication with said one or more imaging devices and operatively configured to estimate a change in volume of said subject from said information. 2. A system according to claim 1, wherein said imaging processor is a computer. 3. A system according to claim 1, wherein said imaging processor comprises computed tomography technology for estimating said change in volume. 4. A plethysmograph for a non-invasive pulmonary measurement of a subject, the plethysmograph comprising: a structure having an interior chamber for receiving the subject and for containing one or more gases; one or more imaging devices, wherein each of said one or more imaging devices captures one or more images of the subject, when positioned in the structure, as the subject breathes said one or more gases and provides an output signal including information regarding said one or more images, said one or more imaging devices including at least two imaging devices positioned orthogonal to each other with respect to the subject; andan imaging processor in communication with said one or more imaging devices so as to receive said output signal, wherein said imaging processor estimates a change in volume of the subject using said information included in said output signal, wherein said structure includes an inside surface defining a portion of said interior chamber, said plethysmograph further including a contrast assistance element positioned on said inside surface so as to be opposite said one or more imaging devices. 5. A plethysmograph according to claim 4, wherein said one or more imaging devices is inside said interior chamber. 6. A plethysmograph according to claim 4, wherein said one or more imaging devices is outside said interior chamber. 7. A plethysmograph according to claim 4, further comprising a temperature regulation element for maintaining a temperature of said one or more gases at or about the temperature of a pulmonary airway of the subject. 8. A plethysmograph according to claim 4, further comprising a moisture regulation element for maintaining a predetermined level of moisture of said one or more gases. 9. A plethysmograph according to claim 4, wherein said contrast assistance element comprises an electroluminescent panel. 10. A plethysmograph according to claim 4, wherein said... | The University of Vermont and State Agricultural College | A61B005/05
US07938784 | B2 | 2011-05-10 | Compact diagnostic neurological tool | 2007-02-02 | A compact diagnostic neurological tool comprising a two-point discriminator; a brush; and a sharp-dull instrument; wherein the two-point discriminator, the brush, and the sharp-dull instrument are all housed within a case; wherein the two-point discriminator, the brush, and the sharp-dull instrument are all retractable; wherein the two-point discriminator comprises two tines, and wherein the distance between the tines can be increased by extending the two-point discriminator and decreased by retracting the two-point discriminator. The case is preferably rectangular in shape and roughly five inches long, one and one-quarter inches wide, and one-half inch thick. | 1. A compact diagnostic neurological tool comprising: (a) a variable two-point discriminator;(b) a brush; and(c) a sharp-dull instrument having a fine with a pointed end and a fine with a blunt end, wherein the tines are at a constant distance from one another;wherein the variable two-point discriminator, the brush, and the sharp-dull instrument are all housed within a single case;wherein the variable two-point discriminator, the brush, and the sharp-dull instrument may be retracted so that they are fully enclosed by the case or extended so that they extend outside of the case;wherein the variable two-point discriminator comprises two tines with distal ends, and wherein the distance between the distal ends of the fines can be increased by the user by pushing a thumb grip upward to extend the two-point discriminator outward from the case and decreased by the user by moving the thumb grip downward to retract the two-point discriminator into the case, and wherein a scale situated adjacent to the thumb grip corresponds to the distance between the distal ends of the tines; andwherein neither the variable two-point discriminator, the brush, nor the sharp-dull instrument is removable from the case, the two-point discriminator, the brush, the sharp-dull instrument, and a Whartenberg wheel are all fully enclosed within the case when not in use, and the variable two-point discriminator, the brush, and the sharp-dull instrument are used without being removed from the case. 2. The compact diagnostic neurological tool of claim 1, wherein the case is rectangular in shape. 3. The compact diagnostic neurological tool of claim 1, further comprising thumb grips that enable the brush and the sharp-dull instrument to be extended outside of the case or retracted into the case. 4. The compact diagnostic tool of claim 3, wherein the tool comprises a front side, wherein the thumb grips are located on the front side of the case, and wherein the case is concave in shape such that the front side of the case is curved inward to facilitate use of the thumb grips. 5. The compact diagnostic neurological tool of claim 1, further comprising a hinge, wherein the tool comprises a top end, wherein the hinge is located on the top end of the tool, wherein the brush and the sharp-dull instrument are adjacent one another, wherein the two-point discriminator is located adjacent to a swing arm when the swing arm is in a closed position, wherein the swing arm can be pivoted upward on the hinge until the swing arm is parallel to and directly above the two-point discriminator, and wherein the swing arm comprises the Wartenberg wheel. 6. The compact diagnostic tool of claim 5, further comprising a Buck-type hammer, wherein the swing arm comprises a top end, and wherein the Buck-type hammer is located on the outside of the case on the top end of the swing arm when the swing arm is in a closed position. 7. The compact diagnostic neurological tool of claim 6, further comprising a Taylor-t... | J & R Enterprises, Inc. | A61B019/00 A61B005/00 A61B005/103 A61B005/117
US07963923 | B2 | 2011-06-21 | Blood pressure meter | 2005-10-04 | A blood-pressure meter having storing means for storing a blood pressure measurement result, and time measurement means for measuring a time is disclosed. The blood-pressure meter includes: input means for entering behavior history; means for storing the blood pressure measurement result into the storing means while associating the blood pressure measurement result with a time when blood pressure measurement is performed; and means for storing the entered behavior history into the storing means while associating the behavior history with a time corresponding to the behavior history. | 1. An ear blood-pressure meter comprising: storing means for storing a blood pressure measurement result;time measurement means for measuring a time;input means for receiving behavior history; means for storing the blood pressure measurement result into the storing means while associating the blood pressure measurement result with a time when blood pressure measurement is performed;means for storing the behavior history into the storing means while associating the behavior history with a time corresponding to the behavior history;a pair of cuffs adapted to apply pressure directly onto opposite surfaces of a tragus or a periphery of the tragus in a temporal region;a pair of opposed arms linked together at one end of each arm by a spindle;a first air tube connected to one of the cuffs, the first air tube extending through a first end of a first arm of the pair of opposed arms and exiting through a second end of the first arm;a second air tube connected to another of the cuffs, the second air tube extending through a first end of a second arm of the pair of opposed arms and exiting through a second end of the second arm;a distance variable mechanism arranged to adjust a distance between ends of the arms remote from the spindle by changing an angle between the spindle and one of the arms, the cuffs being attached to their respective arms at the remote ends thereof;a detector configured to detect a pulse wave at the tragus or tragus periphery, the detector including a light-emitting element and a light receiving element, and both the light-emitting element and the light-receiving element are located on a same expanding surface of one of the cuffs, so as to be pressed in use against the tragus or tragus periphery; anda controller configured to control the cuffs to increase or decrease pressure applied onto the tragus or tragus periphery,wherein the ear blood pressure meter is adapted to measure blood pressure by pinching the tragus or tragus periphery between the cuffs to compress blood vessels. 2. The ear blood-pressure meter as claimed in claim 1, further comprising: display means for displaying candidates for being entered as the behavior history. 3. The ear blood-pressure meter as claimed in claim 1, wherein the ear-type blood-pressure meter performs blood pressure measurement at predetermined time intervals by measuring the predetermined time intervals by the time measurement means. 4. The ear blood-pressure meter as claimed in claim 1, further comprising: search means for searching blood pressure measurement results and behavior history stored in the storing means based on a condition received from the input means to extract data satisfying the condition. 5. The ear blood-pressure meter as claimed in claim 4, wherein the search means extracts information including a blood pressure value corresponding to behavior information received from the input means. 6. The ear blood-pressure meter as claimed in claim 4, wherein the search means extracts... | Nippon Telegraph and Telephone Corporation | A61B005/02
US07953470 | B2 | 2011-05-31 | Method, device and navigation aid for navigation during medical interventions | 2001-10-23 | The invention relates to one or several navigation aids whose position can be used to maintain the recorded state of at least one partial structure of a non-bone structure. The navigation aid(s) is/are inserted into the non-bone structure. | 1. A device for navigation during a medical intervention on a non-osseous structure, comprising: (a) a plurality of navigation aids for navigation during a medical intervention on a non-osseous structure, each navigation aid comprising one marker defining a respective local system of coordinates for determining a position of at least one substructure of the non-osseous structure based on position of the navigation aids in the non-osseous structure, wherein each navigation aid has at least five degrees of freedom to be sensed and wherein each respective local system of coordinates of each navigation aid is locally varying; and(b) means for sensing the spatial position of each of the plurality of navigation aids in at least five degrees of freedom of each navigation aid. 2. The device as claimed in claim 1, comprising a plurality of navigation aids and a plurality of sensors for sensing the position of the navigation aids. 3. The device as claimed in claim 1, comprising means for localizing surgical instruments. 4. The device as claimed in claim 1, comprising a sensing field, which covers the effective radius of necessary instruments and of the navigation aids. 5. The device as claimed in claim 1, comprising means for the representation of pre- or intraoperative image data in relation to a necessary instrument. 6. The device as claimed in claim 1, comprising a plurality of navigation aids and means for generating a field, with which a sensor of each navigation aid can interact. 7. The device as claimed in claim 6, wherein the field-generating means generate at least one gradient field. 8. The device as claimed in claim 7, wherein the field-generating means generate at least two gradient fields. 9. A method for navigation during a medical intervention on a non-osseous structure comprising the steps of: (a) registering a non-osseous structure with preoperative data of the non-osseous structure;(b) fixing a plurality of navigation aids on or in the non-osseous structure in selected positions, each navigation aid comprising one marker defining a respective local system of coordinates, wherein each navigation aid has at least five degrees of freedom to be sensed;(c) defining the respective local system of coordinates via each navigation aid, wherein each respective local system of coordinates of each navigation aid is locally varying; and(d) determining a substructure position of at least one substructure of the non-osseous structure based on the varying positions of the navigation aids and based on the respective varying local coordinate systems. 10. The method as claimed in claim 9, wherein the non-osseous structure is registered before introduction of the navigation aids, so that the navigation aids can be attached or introduced in a navigated manner. 11. The method as claimed in claim 9, wherein a preoperative plan for the operation is prepared on the basis of the preoperative data and this plan for the operation is checked and, if appropr... | Deutsches Krebsforschungszentrum Stiftung des öffentlichen Rechts | A61B005/05
US07970463 | B2 | 2011-06-28 | Anti-tachyarrhythmia system with unified atrial tachyarrhythmia rate threshold | 2008-11-04 | A cardiac rhythm management (CRM) system includes an implantable cardioverter defibrillator (ICD) and an external system. The ICD includes a plurality of functional modules performing tachyarrhythmia classification and therapy control functions using atrial tachyarrhythmia rate thresholds that are set to a unified value. In one embodiment, the CRM system allows a user to activate and deactivate each of the functional modules and program the unified value using the external system. | 1. A cardiac rhythm management (CRM) system, comprising: an implantable cardioverter defibrillator (ICD) including: a cardiac sensing circuit configured to sense one or more cardiac signals including an atrial electrogram;a rate detector configured to detect an atrial rate from the atrial electrogram;a plurality of functional modules each configured to be activated and deactivated in response to a selection command, the functional modules each performing a tachyarrhythmia classification or therapy control function using the atrial rate and an atrial tachyarrhythmia rate threshold; anda threshold circuit configured to receive a unified threshold value and set the atrial tachyarrhythmia rate threshold for each activated functional module of the plurality of functional modules to the unified threshold value. 2. The system of claim 1, wherein the plurality of functional modules comprises an atrial fibrillation (AF) detector configured to compare the atrial rate to an AF rate threshold and declare a detection of AF in response to the atrial rate exceeding the AF rate threshold, wherein the threshold circuit is configured to set the AF rate threshold to the unified threshold value. 3. The system of claim 1, wherein the plurality of functional modules comprises an atrial tachyarrhythmia response (ATR) controller configured to adjust delivery of pacing pulses in response to the atrial rate exceeding an ATR rate, wherein the threshold circuit is configured to set the ATR rate to the unified threshold value. 4. The system of claim 3, wherein the ATR controller is configured to switch a pacing mode in response to the atrial rate exceeding the ATR rate. 5. The system of claim 4, wherein the ATR controller is configured to switch the pacing mode from DDD mode to VVI, DDI, or VDI mode, or from VDD mode to VDI mode, in response to the atrial rate exceeding the ATR rate. 6. The system of claim 1, wherein the plurality of functional modules comprises an atrial flutter response (AFR) controller configured to trigger an AFR window in response to a detection of an intrinsic atrial depolarization, the AFR window used for controlling detection of intrinsic atrial depolarizations that trigger delivery of ventricular pacing pulses during an atrial tracking pacing mode, wherein the threshold circuit is configured to set the AFR window to an atrial tachyarrhythmia interval value corresponding to the unified threshold value. 7. The system of claim 6, wherein the AFR controller is configured to retrigger the AFR window in response to a detection of the intrinsic atrial depolarization during the AFR window. 8. The system of claim 1, wherein the plurality of functional modules comprises an atrial tachyarrhythmia therapy controller configured to cause delivery of an anti-tachyarrhythmia therapy in response to the atrial rate exceeding an atrial tachyarrhythmia therapy rate, wherein the threshold circuit is configured to set the atrial tachyarrhythmia therapy rate to the... | Cardiac Pacemakers, Inc. | A61N001/00 A61B005/04
US20110137155 | A1 | 2011-06-09 | DELIVERY DEVICE FOR LOCALIZED DELIVERY OF A THERAPEUTIC AGENT | 2010-12-08 | Therapeutic agent delivery devices and methods for delivering a therapeutic agent to a target location as well as methods for determining the location of a lesion on a vessel wall are disclosed. Various embodiments disclose an expandable member comprising a drug delivery matrix for selectively delivering a therapeutic agent to a lesion on a vessel wall. The drug delivery matrix may comprise one or more sensors and an electroactive polymer for releasing the therapeutic agent. Other embodiments disclose an expandable member comprising a plurality of radially-expanding flexible walls forming a plurality of channels for selectively delivering therapeutic agent to a target area adjacent one or more of the channels. Detecting a lesion may comprise using a plurality of Hall effect sensors disposed on a distal end of a catheter. | 1. A therapeutic agent delivery device comprising: an elongate member having a distal end;an expandable member disposed on the distal end of the elongate member; anda drug delivery matrix disposed on at least a portion of the expandable member, the drug delivery matrix comprising: one or more drug delivery areas, each drug delivery area comprising an electroactive polymer;one or more sensors adapted to detect a condition of a target location on a vessel wall; andone or more conductive elements for transmitting at least one signal from the one or more sensors and for transmitting at least one signal to the electroactive polymer of the one or more drug delivery areas;wherein when a first sensor of the one or more sensors detects the condition of the target location, at least one signal is transmitted from the first sensor, and, based on the detection of the condition of the target location, at least one signal is transmitted to a first drug delivery area of the one or more drug delivery areas, thereby causing the therapeutic agent to be delivered from the electroactive polymer of the first drug delivery area to the target location. 2. The therapeutic agent delivery device according to claim 1, wherein the first drug delivery area is adapted to communicate with a second drug delivery area of the one or more drug delivery areas; and wherein when the signal is transmitted to the first drug delivery area, the first drug delivery area communicates the signal to the second drug delivery area, thereby causing the therapeutic agent to be delivered from an electroactive polymer of the second drug delivery area to the target location. 3. The therapeutic agent delivery device according to claim 1, wherein the one or more sensors comprise at least one of a mechanical, environmental and biochemical sensor. 4. The therapeutic agent delivery device according to claim 1, wherein the signals are electrical signals. 5. The therapeutic agent delivery device according to claim 1, wherein the one or more conductive elements comprise at least one of metal and polymer wiring. 6. A method of delivering a therapeutic agent to a target location, the method comprising: (a) using a therapeutic agent delivery device comprising: (i) an elongate member having a distal end;(ii) an expandable member disposed on the distal end of the elongate member; and(iii) a drug delivery matrix disposed on at least a portion of the expandable member, the drug delivery matrix comprising: (1) one or more drug delivery areas, each drug delivery area comprising an electroactive polymer;(2) one or more sensors adapted to detect a condition of the target location on a vessel wall; and(3) one or more conductive elements for transmitting at least one signal from the sensors and for transmitting at least one signal to the electroactive polymer of the one or more drug delivery areas;(b) positioning the device in the vessel;(c) detecting the condition of the target location; and(d) transmitting at le... | BOSTON SCIENTIFIC SCIMED, INC. | A61M025/10 A61B005/05
US07941205 | B2 | 2011-05-10 | System and method for separating cardiac signals | 2002-07-03 | EKG sensors ((150) are placed on a patient (140) to receive electrocardiogram (EKG) recording signals, which are typically combinations of original signals from different sources, such as pacemaker signals, QRS complex signals, and irregular oscillatory signals that suggest an arrhythmia condition. A computing module (120) uses independent component analysis to separate the recorded EKG signals. The separated signals are displayed to help physicians to analyze heart conditions and to identify probably locations of abnormal heart conditions. At least a portion of the separated signals can be further displayed in a chaos phase space portrait to help detect abnormality in heart conditions. | 1. A medical system for processing electrocardiogram (EKG) signals, the medical system comprising: a receiving module configured to receive a plurality J of recorded EKG signals Xj from a plurality of EKG sensors, one of the recorded EKG signals per one of the EKG sensors;a computing module configured to separate the received signals using independent component analysis to produce a plurality I of statistically independent separated signals Yi;a display module configured to display the separated signals;a database storing a plurality of EKG signal triggers and one or more corresponding diagnoses, wherein at least one of the EKG signal triggers is based on history of the patient during an observation period; anda matching module configured to match the separated signals with one or more of the stored EKG signal triggers, the one or more of the stored EKG signal triggers comprising the at least one of the EKG signal triggers based on history of the patient during the observation period. 2. The medical system of claim 1, wherein the display module is further configured to display at least a portion of the separated signals in a chaos phase space portrait. 3. The medical system of claim 1, wherein the separated signals include three components of QRS complex, and wherein the display module is further configured to display at least the three QRS complex components in a chaos phase space portrait. 4. The medical system of claim 1, wherein the computing module is configured to separate the recorded signals by performing steps comprising multiplying a vector of the recorded signals by a matrix Wij such that Yi=Wij*Xj. 5. The medical system of claim 1, wherein the computing module is configured to separate the recorded signals using a neural-network implemented method, the neural-network implemented method comprising: selecting a plurality I of bias weights Wi0 and a plurality I*J of scaling weights Wij;adjusting the bias weights Wi0 and the scaling weights Wij to minimize information redundancy among separated signals; andproducing the statistically independent separated signals Yi such that Yi=Wij*Xj+Wi0. 6. The medical system of claim 1, wherein, when the matching module matches one of the EKG signal triggers to a particular diagnosis, the computing module causes the particular diagnosis to be displayed. 7. A computer-implemented method of processing electrocardiogram (EKG) recording signals, the method comprising: receiving a first plurality of EKG recording signals from EKG sensors placed on a patient, one of the EKG recording signals per one of the EKG sensors;separating the first plurality of EKG recording signals using independent component analysis to produce a second plurality of statistically independent separated signals;storing a plurality of EKG signal triggers and one or more diagnoses, each diagnosis of the one or more diagnoses corresponding to at least one EKG signal trigger of the plurality of EKG signal triggers, wherein the at l... | Sigmed, Inc. | A61B005/0452
US07966061 | B2 | 2011-06-21 | Processing and analyzing physiological signals to detect a health condition | 2007-08-29 | Described herein is a method of developing a fuzzy logic system to detect a non-normal health condition. In particular, signal processing and transformation of electrocardiogram (EKG) signals for sleep disorder breathing are provided. The method includes: recording EKG measurements during sleep; clipping the measurements into clips of a consistent length; calculating heart rate and obtaining an evenly sampled discrete time series data clip; performing Short-Time Discrete Fourier Transform on each data clip generating STDFT respective matrices; encoding each STDFT matrix into a grey-level image; calculating Grey-Level Co-occurrence Matrices; extracting textural features; performing statistical analysis on the features to formulate rules; and employing the rules in a Fuzzy Logic system. The system and method described herein yields an accuracy of 75.88%, or better, in detection of sleep apnea. | 1. A method of developing a sleep disorder detection system, the method comprising: obtaining EKG recordings from known normal patients and known sleep disorder patients;clipping the EKG measurements into specific time intervals, forming clips;calculating an R-R interval discrete time series for each clip, and evenly resampling the R-R interval discrete time series using cubic spline interpolation, creating data clips;performing Short-Time Discrete Fourier Transform (STDFT) on each data clip, which generates STDFT complex matrices;encoding each STDFT matrix into at least one grey-level image and calculating at least one Grey-Level Co-occurrence Matrix (GLCM) for each at least one grey-level image;calculating textural features for each GLCM;performing statistical analysis on each textural feature between GLCMs from known normal patients and GLCMs from known sleep disorder patients;formulating fuzzy logic rules based on the statistical analysis; andemploying fuzzy logic to implement the rules and training the sleep disorder detection system using the textural features from known normal and known sleep disorder patients as inputs. 2. The method according to claim 1, wherein at least two different grey-scale encoding methods are used to form at least two images. 3. The method according to claim 1, wherein STDFT matrices are encoded into color images. 4. The method according to claim 1, wherein at least two GLCMs are calculated for each image. 5. The method according to claim 1, further comprising: segmenting encoded images into at least four equal squares or rectangles. 6. The method according to claim 5, further comprising: calculating at least two GLCMs for each segment. 7. The method according to claim 6, wherein the at least two GLCMs for each segment, each have a different orientation (Θ) parameter. 8. The method according to claim 6, wherein the at least two GLCMs for each segment, each have a different distance parameter. 9. The method according to claim 1, further comprising: using four different grey-scale encoding methods are used to form respective four images for each STDFT matrix. 10. The method according to claim 9, further comprising: segmenting each image into at least four equal segments;calculating at least two different GLCMs for each segment, for at least four segments. 11. A method of screening a patient for sleep apneic episodes, the method comprising: recording nocturnal EKG signals from the patient; clipping the EKG measurement into specific time intervals, forming clips;calculating an R-R interval discrete time series for each clip, and evenly resampling the R-R interval discrete time series using cubic spline interpolation, creating data clips;performing Short-Time Discrete Fourier Transform (STDFT) on each data clip, which generates STDFT complex matrices;encoding each STDFT matrix into at least one grey-level image and calculating at least one Grey-Level Co-occurrence Matrix (GLCM) for each at least one grey-leve... | Board of Regents, The University of Texas System | A61B005/0456
US07970449 | B2 | 2011-06-28 | Method and apparatus for background current arrangements for a biosensor | 2010-04-12 | Embodiments of the present invention provide for raising a background current setting for a biosensor above the actual (measured) background current present (i.e., overestimating the background current), particularly in the hypoglycemic range, to improve sensor accuracy and decrease the chance of glucose value overestimation by the sensor. | 1. A method for forming an adjusted analyte sensor value, comprising: determining by a computing device for an analyte sensor an initial analyte sensor value using an initial background current value;adjusting by the computing device the initial background current value by a factor based on the determined initial analyte sensor value to form an adjusted background current value; andforming by the computing device an adjusted analyte sensor value using the adjusted background current value. 2. The method of claim 1, wherein said factor is a multiplication factor applied to the initial background current value. 3. The method of claim 2, wherein said multiplication factor is less than 3. 4. The method of claim 1, wherein said factor comprises an increase or decrease of the initial background current value to an adjusted background current value selected from a predefined list of acceptable adjusted background current values. 5. The method of claim 4, wherein said predefined list of acceptable adjusted background current values comprises 0, 0.8, 1.5, 2.2, and 3.0 nA. 6. The method of claim 1, wherein the adjusted background current value is greater than the initial background current value. 7. The method of claim 1, wherein said based on the initial analyte sensor value comprises determining by the computing device whether the initial analyte sensor value falls within a predetermined range. 8. The method of claim 7, wherein said predetermined range is 70 mg/dl and below. 9. The method of claim 8, wherein said adjusted background current value is approximately 1.3-1.7 nA. 10. The method of claim 7, wherein said predetermined range is above 70 mg/dl. 11. The method of claim 10, wherein said adjusted background current value is approximately 0.6-1.0 nA. 12. The method of claim 1, further comprising determining by the computing device the initial background current value by providing by a computing device an arbitrary background current level within a range of approximately 0.5 to 2.0 nA. 13. The method of claim 1, wherein said adjusted analyte sensor value is displayed on an associated display device. 14. The method of claim 1, wherein said adjusted background current value is further adjusted in real time as the adjusted analyte sensor value changes. 15. The method of claim 1, wherein said using the initial background current value comprises subtracting by the computing device the initial background current value from a first measured analyte sensor value to form said initial analyte sensor value, and said using the adjusted background current value comprises subtracting by the computing device the adjusted background current value from a second measured analyte sensor value to form said adjusted analyte sensor value. 16. The method of claim 1, wherein said factor is defined as a percentage of the initial background current value. 17. The method of claim 1, further comprising determining a rate at which an analyte level has changed duri... | Isense Corporation | A61B005/04
US07957780 | B2 | 2011-06-07 | Physiological parameter confidence measure | 2006-03-01 | Confidence in a physiological parameter is measured from physiological data responsive to the intensity of multiple wavelengths of optical radiation after tissue attenuation. The physiological parameter is estimated based upon the physiological data. Reference data clusters are stored according to known values of the physiological parameter. At least one of the data clusters is selected according to the estimated physiological parameter. The confidence measure is determined from a comparison of the selected data clusters and the physiological data. | 1. A method of determining a measure of confidence in a physiological parameter, the physiological parameter determined by transmitting multiple wavelengths of optical radiation into a tissue site and detecting the optical radiation after tissue attenuation, the method comprising: deriving physiological data responsive to the intensity of multiple wavelengths of optical radiation transmitted into a tissue site and detected after tissue attenuation;estimating a physiological parameter based upon the physiological data;providing a physiological data reference;obtaining at least one data cluster from the physiological data reference wherein the at least one data cluster is a probability distribution of normalized pleth ratios; anddetermining a measure of confidence in the estimated physiological parameter based upon the at least one data cluster and the derived physiological data. 2. The method according to claim 1 wherein the providing step comprises: predetermining the physiological data for known values of the physiological parameter across a sample population;clustering the data according to the physiological parameter values; andstoring the data clusters so as to be retrievable according to the physiological parameter values. 3. The method according to claim 2 wherein the obtaining step comprises selecting the at least one data cluster according to the estimated physiological parameter. 4. The method according to claim 3 wherein the selecting step comprises: determining at least one data cluster having a corresponding physiological parameter value closest to the estimated physiological parameter; andreading the determined at least one data cluster from the memory. 5. The method according to claim 4 wherein the physiological parameter is at least one of SpO2, MetHb and HbCO. 6. A physiological parameter confidence measurement method comprising: deriving physiological data responsive to the intensity of multiple wavelengths of optical radiation transmitted into a tissue site and detected after tissue attenuation;estimating a physiological parameter based upon the physiological data;providing a physiological data reference having a plurality of data clusters each corresponding to a particular value of the physiological parameter,;comparing at least one of the data clusters to the physiological data;indicating confidence in the estimated physiological parameter based upon the comparison; andassociating a probability function with each of the data clusters. 7. The physiological parameter confidence measurement method according to claim 6 wherein the comparing step comprises determining a probability that the derived physiological data corresponds to the estimated physiological parameter. 8. The physiological parameter confidence measurement method according to claim 7 wherein the indicating step comprises generating at least one of a visual indication and an audible indication corresponding to the determined probability. 9. The physiological... | MASIMO Laboratories, Inc. | A61B005/1455
US07959577 | B2 | 2011-06-14 | Method, system, and apparatus for neural localization | 2008-03-31 | Described herein are devices, systems and methods for determining if a nerve is nearby a device or a region of a device. In general, a device for determining if a nerve is nearby a device includes an elongate body having an outer surface with one or more bipole pairs arranged on the outer surface. Bipole pairs may also be referred to as tight bipoles. The bipole pairs may be arranged as a bipole network, and may include a cathode and an anode that are spaced relatively close together to form a limited broadcast field. In general, the broadcast filed is a controlled or “tight” broadcast field that extends from the bipole pair(s). Methods of using these devices and system are also described. | 1. A method of detecting if a nerve is above or below a given region of a device in tissue, the method comprising: positioning a guidewire within the tissue of a patient;positioning a device within the tissue using the guidewire, wherein the device comprises a flexible elongate body having an outer surface including an upper surface that is opposite a lower surface and plurality of anodes and cathodes on the outer surface;determining a threshold amount of energy required to stimulate a response in the patient by applying increasing levels of energy to form a substantially continuous broadcast field between the plurality of anodes and cathodes on the upper surface such that the broadcast field is formed by adjacent bipole pairs of anodes and cathodes which share alternating anodes, to determine the first stimulation level at which the nerve responds;determining a threshold amount of energy required to stimulate a response in the patient from a plurality of anodes and cathodes on the lower surface of the device by applying energy to form a substantially continuous broadcast field from the lower surface; anddetermining if the nerve is closer to the upper or lower surface regions by comparing the threshold amounts from the upper surface and the lower surface. 2. The method of claim 1, wherein the step of determining a threshold amount of energy required to stimulate a response in the patient comprises determining the threshold current required to stimulate an EMG response in the patient. 3. The method of claim 1, wherein the step of positioning the device comprises positioning a device that is substantially flat, having a first major side on the upper surface and a second major side on the lower surface. 4. The method of claim 1, wherein the step of positioning the device comprises moving the device over the guidewire. 5. The method of claim 1, wherein the step of positioning the guidewire comprises passing the guidewire from outside of the patient, through a neural foramen, and out of the patient, so that both ends of the guidewire extend from the patient. 6. The method of claim 1, wherein the step of determining a threshold amount of energy required to stimulate a response in the patient comprises incrementally increasing the current applied to the plurality of anodes and cathodes to form the substantially continuous broadcast field. 7. A method of detecting if a nerve is above or below a given region of a device in a tissue, the method comprising: positioning a device within the tissue so that an upper surface of the device faces the dorsal region of the patient and a lower surface of the device faces ventral region of the patient, wherein the device comprises a flexible elongate body having a first plurality of anodes and cathodes on the upper surface and a second plurality of anodes and cathodes on the lower surface;determining a threshold amount of energy required to stimulate a response in the patient from the upper surface by applying i... | Baxano, Inc. | A61B005/05
US20110118588 | A1 | 2011-05-19 | Combination MRI and Radiotherapy Systems and Methods of Use | 2009-03-12 | A combination MRI and radiotherapy system comprising: a) an MRI system for imaging a patient receiving radiotherapy, comprising a magnetic field source suitable for generating a magnetic field of strength and uniformity useable for imaging, capable of being ramped up to said magnetic field in less than 10 minutes, and ramped down from said magnetic field in less than 10 minutes; b) a radiation source configured for applying radiotherapy; and c) a controller which ramps the magnetic field source down to less than 20% of said magnetic field strength when the radiation source is to be used for radiotherapy, and ramps the magnetic field source up to said magnetic field strength when the MRI system is to be used for imaging. | 1. A combination MRI and radiotherapy system comprising: a) an MRI system for imaging a patient receiving radiotherapy, comprising a magnetic field source suitable for generating a magnetic field of strength and uniformity useable for imaging, capable of being ramped up to said magnetic field in less than 10 minutes, and ramped down from said magnetic field in less than 10 minutes;b) a radiation source configured for applying radiotherapy; andc) a controller which ramps the magnetic field source down to less than 20% of said magnetic field strength when the radiation source is to be used for radiotherapy, and ramps the magnetic field source up to said magnetic field strength when the MRI system is to be used for imaging. 2. A system according to claim 1, wherein the magnetic field source comprises a non-superconducting coil. 3. A system according to any of the preceding claims, wherein the MRI system comprises a prepolarized MRI system, and the magnetic field source comprises a high field source for polarization and a low field source for readout. 4. A system according to claim 3, wherein the high field source comprises a superconducting coil capable of ramping up to a high magnetic field used for polarization, and ramping down from the high magnetic field, each in less than 10 minutes. 5. A system according to claim 3 or claim 4, wherein the low field source comprises a superconducting coil. 6. A system according to claim 5, wherein the superconducting coil of the low field source is capable of ramping up to a low magnetic field used for readout, and ramping down from the low magnetic field, each in less than 10 minutes. 7. A system according to any of the preceding claims, wherein the controller controls the MRI system not to acquire images when the radiation source is being used for radiotherapy. 8. A system according to any of the preceding claims, also comprising a real-time tracker which tracks changes in position of a radiotherapy target in the patient. 9. A system according to claim 8, wherein the tracker comprises an RF tracker. 10. A system according to claim 8 or claim 9, wherein the tracker includes a radioactive marker. 11. A system according to any of claims 8-10, wherein the tracker comprises an image-based tracking system. 12. A system according to any of claims 8-11, wherein the tracker comprises an implanted leadless marker. 13. A system according to any of the preceding claims, wherein the radiation source comprises a linac. 14. A system according to any of claims 1-13, wherein the radiation source comprises a radioactive source. 15. A system according to any of the preceding claims, wherein the magnetic field source of the MRI system comprises an open magnet. 16. A system according to any of the preceding claims, wherein the magnetic field source is sufficiently well shielded magnetically such that the magnetic field used for imaging is less than 100 gauss throughout any volume where the radiation source is loca... | - | A61B005/055 A61N005/00
US20110130674 | A1 | 2011-06-02 | Method and System for Regional Assessment of Lung Physiology | 2006-11-23 | The invention provides a system and method for regional assessment of lung physiology. The system includes a plurality of sound transducers configured to be fixed on a surface of the individual over the thorax. A processor is configured to receive signals generated by the transducers and to determine from the signals a value of a parameter in each of one or more regions of the lungs. The method of the invention includes obtaining signals indicative of pressure waves at locations over the thorax; and determining from the signals a value of a parameter in each of the regions of the lungs. | 1-27. (canceled) 28. A system for regional assessment in two or more regions of an individual's lungs comprising: (a) a plurality of N transducers, each transducer configured to be fixed on a surface of the individual over the thorax, the ith transducer being fixed at a location xi and generating a signal P(xi,t) indicative of pressure waves at the location xi; for i=1 to N; and(b) a processor configured to: i) receive the signals P(xi,t),ii) determine a value of a first parameter at a plurality of locations xi of transducers, in a calculation involving one or more of the signals P(xi,t),iii) determine a value of the first parameter at a plurality of locations x by interpolation of the values determined in step (ii) andiv) for each of the two or more regions determine a value of a second parameter in a calculation involving the values of the first parameter determined in step (iii) at a plurality of locations x in the region. 29. The system according to claim 28 wherein the first or second parameter is a total average acoustic energy of the region over a time interval from a first time t1 to a second time t2. 30. The system according to claim 28 wherein the first or second parameter is a total acoustic energy of the region over a time interval from a first time t1 to a second time t2 divided by a total average acoustic energy of the lungs over the same time interval. 31. The system according to claim 30 further comprising a two-dimensional display device. 32. The system according to claim 30 wherein the processor is further configured to display an image of the lungs divided into the regions and the regional assessments of the regions. 33. The system according to claim 32 wherein the image is obtained in a calculation involving the signals P(xi,t). 34. The system according to claim 34 wherein the image is obtained in a calculation involving average acoustic energies {tilde over (P)}(xi,t1,t2) obtained at locations x over the lungs over a time interval from a first time t1 to a second time t2. 35. The system according to claim 29 wherein the average acoustic energy {tilde over (P)} over a time interval from t1 to t2 is determined at a location xi of a transducer using the algebraic expression:P~(xi,t1,t2)=∫t1t2P2(xi,t)t. 36. The system according to claim 35 wherein an average acoustic energy is determined at least one location x by interpolation of the determined {tilde over (P)}(xi,t1,t2) using the algebraic expression:P~(x,t1,t2)=∑i=1NP~(xi,t1,t2)g(x,xi,σ)(2)where g(x,xi,σ) is a kernel satisfying∇2g=∂g∂σ(3)∑i=1Ng(x,xi,σ)isapproximatelyequalto1.(4) 37. The system according to claim 36 wherein g(x,viσ) is the kernelg(x,x1,σ)=Exp-((x1-xi1σ)22σ)·Exp-((x2-xi2σ)22σ).(5) 38. The system according to claim 28 wherein the processor is configured to perform the regional assessment of the lungs over a plurality of time intervals. 39. A method for regional assessment in two or more regions of an individual's lungs ... | - | A61B005/08
US09173582 | B2 | 2015-11-03 | Adaptive performance trainer | 2010-04-23 | Techniques for accelerating training through optimization of the psychophysiological state of the trainee are provided. These techniques include an adaptive performance training system configured to acquire, analyzed, display, and translate data that reflects the psychophysiological state of the user, including the electrical activity of the brain (EEG), the heart (EKG), the musculature (EMG), respiration and other parameters that characterize the state of the user in real-time. The system includes a plurality of feedback mechanisms for providing visual, auditory, and/or tactile feedback based on the current psychophysiological state of the user and for facilitating moving the user toward a goal psychophysiological state for performing a particular task and for optimizing performance of that task. | 1. A system for accelerating learning of a task using real-time data acquisition and analysis of real-time psychophysiological signals obtained from a user, the system comprising: one or more hardware processors;a monitoring module configured to, when executed by the one or more hardware processors, continuously monitor psychophysiological signal data collected from a plurality of sensors attached to a user during a pre-shot period of the task to be learned;an expertise profiler module configured to analyze the psychophysiological signal data to generate a current psychophysiological state of the user by computing one or more power spectral density indices from electroencephalographic activity of the user and converting the one or more power spectral density indices into one or more probabilities that the user's level of mental effort and engagement belongs to one of a plurality of classes using a discriminant function,compare the current psychophysiological state of the user to one or more goal states associated with expert performance of the task to be learned, andgenerate a set of control signals based on difference between the current psychophysiological state of the user to one or more goal states; anda real-time feedback generator configured to, when executed by the one or more hardware processors, during the pre-shot period, receive the control signals from the expertise profiler module, andgenerate at least one multimodality feedback signal to provide feedback to guide the user toward the one or more goal states, wherein the multimodality feedback signal is based on the control signals received from the expertise profiler module, wherein the feedback comprises a tactile indication to the user of the user's current psychophysiological state at a pace that replicates the user's heartbeat, and wherein the tactile indication of the user's current psychophysiological state decreases as the user's current psychophysiological state approaches an ideal shooting state for the task to be learned. 2. The system of claim 1 wherein the monitoring module is configured to continue to monitor the psychophysiological signal data after the feedback has been delivered to the user; and wherein the system is configured to determine whether the feedback delivered to the user has helped the user to reach the one or more goal states. 3. The system of claim 1 wherein the monitoring module is further configured to monitor additional parameters of task performance, including monitoring user interaction with one more interfaces for capturing user interactions with the system. 4. The system of claim 1 wherein the system further comprises one or more feedback devices, and wherein the multimodality feedback signal activates the one or more feedback devices. 5. The system of claim 4 wherein the one or more feedback devices includes a video display. 6. The system of claim 4 wherein the one or more feedback devices includes an audio output device. 7. The system of c... | ADVANCED BRAIN MONITORING, INC. | A61B005/00 A61B005/04 A61B005/0402 A61B005/024 A61B005/021 A61B005/0476 A61B005/0488 A61B005/08 A61B005/11 A61B005/16
US20110160618 | A1 | 2011-06-30 | APPARATUS FOR THE DIAGNOSIS AND TREATMENT OF BRUXISM | 2009-08-06 | The present invention is an apparatus for detecting and treating bruxism. The apparatus of the invention comprises a sensing unit, a calculation unit and a stimulation unit. The sensing unit is in charge of measuring the interocclusal distance and reporting this distance to the calculation unit. After receiving the information from the sensing unit, the calculation unit archives and processes the information to determine if the sensing event is a bruxism event. When a sensing event is confirmed as a bruxism event, the calculation unit selects an adequate stimulation and sends a request to the stimulation unit, which applies the selected stimulation. The apparatus of the present invention may be suitable for treating both diurnal and nocturnal bruxism and the intrabuccal components may be removed from the mouth when bruxism events are not expected to occur, i.e. during the day for an individual who suffers only from nocturnal bruxism. | 1. An apparatus for detecting and treating bruxism comprising three dedicated units, the elements of which are located into one or more components, wherein said units are: a) a sensing unit comprised of at least one element adapted to make measurements of distance:b) a calculation unit adapted to determine distances from said measurements, to detect the possibility of the occurrence of a bruxism event from the value of said distances, to confirm that said possible event is an actual bruxism event that requires treatment, to make logic decisions, and to deliver a signal that a stimulus should be applied to end said confirmed event; andc) a stimulation unit adapted to receive said signal from said calculation unit and to create and apply said stimulus to end said confirmed bruxism event;said apparatus characterized in that at least one of said elements of said sensing unit are located in a component that is attached to a first tooth located in one jaw of a patient and the distance measurements made by said sensing unit are indicative of the interocclusal distance between said first tooth and the opposing tooth located in said patient's other jaw. 2. An apparatus according to claim 1, wherein the sensing unit comprises at least one signal emitting element and at least one detector element configured in one of the following ways: a) one signal emitting element and one detector element are included in one component attached to a tooth of a patient;b) one signal emitting element and one detector element are included in one component attached to a tooth of a patient and a passive element is attached to the opposing tooth in the other jaw,c) one of said signal emitting elements is located in a first component, which is attached to a first tooth in one jaw of a patient and one of said detector elements is located in a second component, which is attached to the opposing tooth in the other jaw. 3. (canceled) 4. (canceled) 5. An apparatus according to claim 1, wherein the calculation unit comprises a microprocessor, a dedicated software package, memory, two way communication means and an electronic circuit, wherein the elements of said calculation unit are included in at least one component. 6. (canceled) 7. An apparatus according to claim 5, wherein the elements of the calculation unit are included in one component, which comprises at least one element of the sensing unit. 8. (canceled) 9. (canceled) 10. An apparatus according to claim 1, wherein the calculation unit is adapted to detect a possible bruxism event if an interocclusal distance measured by the sensing unit is less than a predetermined threshold vale. 11. An apparatus according to claim 1, wherein the calculation unit is adapted to confirm that a possible bruxism event is a real bruxism event on the basis of one or more of the following: a) a measured interocclusal distance and the duration of said possible bruxism event;b) a measured interocclusal distance, the duration of said possib... | WISEMED LTD. | A61B005/103
US08641621 | B2 | 2014-02-04 | Systems, methods, apparatuses, and computer-readable media for image management in image-guided medical procedures | 2011-01-26 | Presented herein are methods, systems, devices, and computer-readable media for image management in image-guided medical procedures. Some embodiments herein allow a physician to use multiple instruments for a surgery and simultaneously provide image-guidance data for those instruments. Various embodiments disclosed herein provide information to physicians about procedures they are performing, the devices (such as ablation needles, ultrasound transducers or probes, scalpels, cauterizers, etc.) they are using during the procedure, the relative emplacements or poses of these devices, prediction information for those devices, and other information. Some embodiments provide useful information about 3D data sets and allow the operator to control the presentation of regions of interest. Additionally, some embodiments provide for quick calibration of surgical instruments or attachments for surgical instruments. | 1. A method for image management in image-guided medical procedures, comprising: determining, with one or more computing devices, pose information for a set of 3D visualizable medical data;determining, with the one or more computing devices, real-time pose information for a medical device;selecting an imaging plane from the set of 3D visualizable medical data based at least on the real-time pose information for the medical device and the pose information for the set of 3D visualizable medical data;determining a perspective view of the selected imaging plane in a virtual 3D space based at least on real-time relative pose information of the medical device with respect to a position of a user;determining a selected portion of the set of 3D visualizable medical data, wherein the selected portion of the set of 3D visualizable medical data corresponds to the set of 3D visualizable medical data in the selected imaging plane;displaying at a first focus setting, on the one or more displays, a perspective view of the selected portion of the set of 3D visualizable medical data in the virtual 3D space based at least on the determined perspective view of the selected imaging plane; anddisplaying at one or more second focus settings, on the one or more displays, a perspective view of the set of 3D visualizable medical data that is in at least a portion of one or more imaging planes that are parallel to the selected imaging plane, wherein the one or more second focus settings vary in proportion to a distance between the one or more imaging planes and the selecting imaging plane. 2. The method of claim 1, wherein the set of 3D visualizable medical data comprises preoperative imaging data. 3. The method of claim 1, wherein the set of 3D visualizable medical data comprises annotations of imaging data. 4. The method of claim 1, wherein the set of 3D visualizable medical data is a first set of 3D visualizable medical data and the selected portion of the set of 3D visualizable medical data is a first selected portion of the set of 3D visualizable medical data, the method further comprising: determining pose information for a second set of 3D visualizable medical data relative to the real-time pose information for the medical device;determining a second selected portion of the second set of 3D visualizable medical data, based at least in part on the real-time pose information for the medical device; anddisplaying, on the one or more displays, a perspective view of the second selected portion of the second set of 3D visualizable medical data in the virtual 3D space. 5. The method of claim 1, wherein the imaging plane is a first imaging plane, the medical device is a first medical device, and the selected portion of the 3D visualizable data is a first selected portion of the 3D visualizable data, the method further comprising: determining real-time pose information for a second medical device;selecting a second imaging plane from the set of 3D visualizable medical d... | InnerOptic Technology, Inc. | A61B008/00
US08696725 | B2 | 2014-04-15 | Devices and methods for controlling patient temperature | 2011-02-07 | Relatively non-invasive devices and methods for heating or cooling a patient's body are disclosed. Devices and methods for treating ischemic conditions by inducing therapeutic hypothermia are disclosed. Devices and methods for inducing therapeutic hypothermia through esophageal cooling are disclosed. Devices and methods for operative temperature management are disclosed. | 1. A system for controlling core body temperature of a subject and inducing systemic hypothermia, comprising: a heat transfer tube configured for insertion within the esophagus of the subject, wherein the tube is configured to contact the epithelial lining of the esophagus;an external heat exchanger containing a heat transfer fluid;a pump for flowing the heat transfer fluid through a circuit within the heat transfer tube;a heat transfer element in contact with the external heat exchanger;a hollow tube having a distal end configured to extend into a stomach of the subject; anda pressure transducer for generating a signal representative of abdominal or esophageal pressure, wherein the signal is transmitted to a microprocessor. 2. The system of claim 1, wherein the microprocessor calculates an estimated transpulmonary pressure based on the esophageal pressure. 3. The system of claim 1, wherein the microprocessor responds to the signal from the pressure transducer with a proportional integrated differential response to control mechanical ventilation. 4. The system of claim 1, wherein the heat transfer tube comprises a heat transfer region capable of contacting substantially all of the esophageal epithelium. 5. The system of claim 1, wherein the system is capable of cooling at a rate of about 1.2° C./hr to about 1.8° C./hr. 6. The system of claim 1, wherein the system is capable of cooling a mass at a rate of about 350 kJ/hr to about 530 kJ/hr. 7. The system of claim 1, wherein the system is capable of cooling a mass at a rate of about 430 kJ/hr. 8. The system of claim 1, wherein the system includes a heat transfer region with a surface area of at least about 100 cm2. 9. The system of claim 8, wherein the heat transfer region has a surface area of about 140 cm2. 10. The system of claim 1, further comprising an ancillary tube, wherein the ancillary tube is insertable within the subject's nasopharynx or auricular canal, and wherein the ancillary tube includes a heat transfer medium flow channel. | Advanced Cooling Therapy, LLC | A61F007/00 A61F007/12 A61B005/00
US20110144484 | A1 | 2011-06-16 | METHOD AND DEVICE FOR MONITORING A MEDICAL PROCEDURE SUCH AS A VASCULAR SURGERY OPERATION | 2010-12-10 | In a monitoring method and device for monitoring a medical procedure, in particular a vascular surgery operation, at least one monitoring measurement in a monitoring region is acquired by a medical imaging device, and the at least one monitoring measurement takes place essentially simultaneously with the medical procedure. | 1. A method for monitoring a vascular surgery operation wherein vascular surgery is performed at a vascular surgery site in the body of a patient, comprising the steps of: with a medical imaging device, monitoring a region of the body of the patient that does not encompass said vascular surgery site, which is at risk of damage due to said vascular surgery, by obtaining an image of said region; andobtaining said image of said region in a monitoring measurement that takes place simultaneously with said vascular surgery. 2. A monitoring method as claimed in claim 1 comprising making multiple monitoring measurements, each including an image of said region, continuously in succession during said vascular surgery. 3. A monitoring method as claimed in claim 1 comprising obtaining at least one reference variable and making a comparison between information contained in said image and said at least one reference variable to detect an impermissible deviation of said information from said reference variable. 4. A monitoring method as claimed in claim 3 comprising making said comparison at a time selected from the group consisting of before making said monitoring measurement and before beginning said vascular surgery. 5. A monitoring method as claimed in claim 3 comprising making said comparison substantially in parallel with said monitoring measurement or said vascular surgery. 6. A monitoring method as claimed in claim 3 comprising obtaining at least one monitoring variable, as said information, from said monitoring measurement. 7. A monitoring method as claimed in claim 3 comprising emitting a warning signal if said deviation exceeds a predetermined tolerance range with respect to said reference variable. 8. A monitoring method as claimed in claim 1 comprising obtaining said image by an imaging modality selected from the group consisting of perfusion imaging, angiography imaging, diffusion imaging, susceptibility-weighted imaging and T2*-weighted imaging. 9. A monitoring method as claimed in claim 1 comprising administering a contrast agent to the patient before said monitoring measurement, and generating said image in said monitoring measurement using said contrast agent. 10. A device for monitoring a vascular surgery operation wherein vascular surgery is performed at a vascular surgery site in the body of a patient, comprising: a medical imaging device that monitors a region of the body of the patient that does not encompass said vascular surgery site, which is at risk of damage due to said vascular surgery, by obtaining an image of said region;said medical imaging device being configured to obtain said image of said region in a monitoring measurement that takes place simultaneously with said vascular surgery; anda computer evaluation device configured to evaluate said monitoring measurement simultaneously with said vascular surgery. 11. A monitoring device as claimed in claim 10 comprising wherein said medical imaging device is configured to m... | - | A61B005/05
US20110118623 | A1 | 2011-05-19 | ELECTRONIC THERMOMETER | 2011-01-18 | The invention provides an electronic thermometer in which a contact state with a human body can be confirmed by a simple, easy-to-assemble configuration. The electronic thermometer includes a hollow outer case that includes a probe unit with a temperature measuring unit, a temperature sensor, an inner case, a control circuit, and a pair of electrodes. The electrodes are positioned inside the probe unit by mounting the inner case on the outer case. A determination unit is also provided and the determination unit measures an electrostatic capacitance between the pair of electrodes and determines whether the probe unit is in proper contact with the measured region of the user based on a change of the measured electrostatic capacitance. | 1. An electronic thermometer comprising: a hollow outer case that includes a probe unit, the probe unit including a temperature measuring unit that abuts on a measured region of a user at a leading end thereof, a temperature sensor being disposed in the temperature measuring unit in order to detect a temperature;an inner case that is mounted on a hollow center of the outer case while a electronic circuit board is attached to the inner case, a control circuit that processes data detected with the temperature sensor being formed in the electronic circuit board; anda pair of electrodes that is fixed to the inner case, the electrodes that are not exposed to an outside of the probe unit being positioned inside the probe unit by mounting the inner case on the outer case,wherein a determination unit is provided in the control circuit, the determination unit measuring an electrostatic capacitance between the pair of electrodes and determining whether the probe unit is in proper contact with the measured region of the user based on a change of the measured electrostatic capacitance. 2. The electronic thermometer according to claim 1, wherein the pair of electrodes is disposed in a longitudinal direction of the probe unit while separated from each other with an interval. 3. The electronic thermometer according to claim 1, wherein the pair of electrodes is fixed to the inner case by fitting a recess or a projection, provided in the pair of electrodes, and a projection or a recess, provided in the inner case, in each other. 4. The electronic thermometer according to claim 3, wherein the pair of electrodes and/or the inner case includes the plurality of recesses or projections. 5. The electronic thermometer according to claim 1, wherein screw fitting units, which are being able to be fitted in each other, are provided in the pair of electrodes and the inner case. 6. The electronic thermometer according to claim 1, wherein an electrode fixing unit in the inner case includes an elastic portion, and the pair of electrodes is positioned by pressing the electrodes against an inner wall surface of the probe unit such that the electrodes are attached firmly to the inner wall surface of the probe unit. 7. The electronic thermometer according to claim 2, wherein the pair of electrodes is fixed to the inner case by fitting a recess or a projection, provided in the pair of electrodes, and a projection or a recess, provided in the inner case, in each other. 8. The electronic thermometer according to claim 2, wherein screw fitting units, which are being able to be fitted in each other, are provided in the pair of electrodes and the inner case. 9. The electronic thermometer according to claim 2, wherein an electrode fixing unit in the inner case includes an elastic portion, and the pair of electrodes is positioned by pressing the electrodes against an inner wall surface of the probe unit such that the electrodes are attached firmly to the inner wall surface of the ... | OMRON HEALTHCARE CO., LTD. | A61B005/01
US07966051 | B2 | 2011-06-21 | Fluorescent agent concentration measuring apparatus, dose control apparatus, administration system, fluorescent agent concentration measuring method, and dose control method | 2006-01-09 | A fluorescent agent accumulation concentration measuring apparatus is configured to include a single-wavelength LED to radiate excitation light to a test bottle loaded in the inside, a barrier filter to transmit only fluorescence from the test bottle, a photoreceptor to receive the fluorescence through the barrier filter and output an electric signal, a detection processing circuit to conduct signal processing of the electric signal from the photoreceptor and detect the fluorescence intensity, and an operation circuit 26 to compare the detection result from the detection processing circuit with an analytical pattern stored in a pattern storage portion and calculate the peak time of the accumulation concentration in a tissue of a sample in the test bottle. | 1. A fluorescent agent concentration measuring method comprising: a fluorescence detection step of detecting fluorescence from a sample containing a body fluid taken from a living body administered with a fluorescent drug;a fluorescent agent concentration calculation step of calculating the concentration of the drug in the living body, based on a detection signal from the fluorescence detection step; anda peak time estimation step of estimating the peak time after administration of the drug into the living body, when the concentration of the drug in the living body tissue is the highest, based on the elapsed time after administration of the drug to the living body and the calculated concentration of the drug. 2. A dose control method comprising: a fluorescence detection step of detecting fluorescence from a sample containing a body fluid taken from a living body administered with a fluorescent drug;a fluorescent agent concentration calculation step of calculating the concentration of the drug in the living body, based on a detection signal from the fluorescence detection step;a peak time estimation step of estimating the time after administration of the drug into the living body, when the concentration of the drug in the living body tissue is the highest, based on the elapsed time after administration of the drug to the living body and the calculated concentration of the drug;an administration start judgment step of judging whether a predetermined time has elapsed or not and judging the start of administration of the drug when the predetermined time has elapsed; a dose setting step of setting at least one of the dose and the administration time of the drug, based on the predetermined dose information of the drug after the judgment of the start of administration in the administration start judgment step; andan output step of performing at least one of functions of allowing the dose to be displayed and allowing the drug to be administered automatically, based on the dose of the drug set in the dose setting step. | Olympus Corporation | A61B005/00 A61M031/00
US07951078 | B2 | 2011-05-31 | Method and apparatus for determining familial risk of disease | 2006-02-01 | Personal and family health history information can be used to assess familial risk of disease. For example, information can be collected about the disease history of a person and the person's first- and second-degree relatives and then analyzed to determine the familial risk of common diseases such as coronary heart disease, stroke, type 2 diabetes, and colorectal, breast, and ovarian cancer. Assessed familial risk of disease can then be used by researchers to better estimate the contribution of personal history and family history to the etiology and natural history of a disease of interest, and by consumers and health professionals to determine recommendations for disease management, prevention and screening that are personalized and targeted to the familial risk. Other embodiments are also described and claimed. | 1. A computerized method for determining the extent to which personal and family health histories play a role in developing a disease of interest in a particular subject, the method comprising: accessing, by a computer system, the particular subject's stored, personal and family health histories including a history of one of the group consisting of a) the disease of interest, b) indicator diseases, and c) indicator traits; anddetermining, by the computer system, the subject's assigned level of familial risk of the disease of interest by recognizing familial patterns of disease using a stored array of elements, wherein each element of the array gives a predetermined level of familial risk for the disease of interest corresponding to a different combination of a personal disease history scenario and a familial disease history scenario, wherein the predetermined level of risk was set based on having recognized familial patterns of disease,wherein said determining the subject's assigned level of familial risk of the disease of interest further comprises,indexing into the stored array of elements using a plurality of items taken from the subject's accessed personal and family health histories, to find one or more elements that match said plurality of items,wherein the items taken from the subject's accessed personal and family health histories include (a) whether any first- or second-degree relatives of the subject have developed the disease of interest, (b) whether any first- or second-degree relatives have developed an indicator disease, other than the disease of interest, associated with an increased risk of developing the disease of interest, and (c) for any first- or second-degree relative that developed the disease of interest or indicator disease, information comprising (i) an age of onset at which the disease of interest or indicator disease developed, (ii) a number of first-degree relatives of the subject that developed the disease of interest or indicator disease, and (iii) a number of second-degree relatives of the subject that developed the disease of interest or indicator disease. 2. The method of claim 1 further comprising: compiling, by the computer system, clarifying statements that clarify or further explain the assigned level of familial risk of the disease of interest, wherein the compiled clarifying statements correspond to the same element of the array that was indexed and gave the assigned level of familial risk. 3. An article of manufacture comprising: a non-transitory machine-readable medium having instructions stored therein which program a computer system to access a particular subject's stored personal and family histories, recognize familial patterns of disease to determine an assigned level of familial risk of disease of interest for the particular subject based on having accessed the subject's stored personal and family histories, and present the assigned level of risk to a user, wherein the programmed computer system i... | - | A61B005/00
US07957784 | B2 | 2011-06-07 | Body surface mapping system | 2006-12-29 | A body-surface mapping system is disclosed that uses a plurality of electrodes to map at least a portion of a human torso without having to adjust the positions of the electrodes. The body-surface mapping system energizes groupings or regions of electrodes, then compares and adjusts the current driven through each grouping or region of electrodes to produce near-uniform fields. The electrodes of the body-surface mapping system may be interconnected by wires capable of sensing interelectrode distances, such that the system can reconstruct a detailed model of a patient's torso surface. The body-surface mapping system may also use a catheter in addition to the body surface electrodes to compute both endocardial and epicardial voltage distributions. | 1. A system for body-surface mapping of electrical potentials of at least a portion of a human body, the system comprising: a) a flexible garment adapted to fit at least a portion of the human body, said flexible garment comprising a plurality of interconnected electrodes wherein at least a portion of the plurality of electrodes are arranged in a plurality of regions of electrodes;b) an electronic device having a software program adapted to measure drive currents for at least two regions of electrodes and to measure linearities of at least two electrical fields created by the at least two regions of electrodes, such that the program may compare the at least two regions of electrodes and identify at least one region that creates an electrical field that is more linear than an electrical field generated by at least one other region; andc) a memory coupled to said electronic device for storing information regarding at least the identified at least one region of electrodes, said stored information comprising the drive currents for the at least one region. 2. The system of claim 1 wherein said software program is further adapted to measure homogeneities of the at least two electrical fields created by the at least two regions of electrodes, such that the program may identify at least one region that creates an electrical field that is more homogeneous than an electric field generated by at least one other region. 3. The system of claim 1 wherein said software program is further adapted to control drive currents for the at least two regions of electrodes. 4. The system of claim 1 wherein the plurality of interconnected electrodes are interconnected by piezoelectric wires such that relative distances between electrodes within the plurality of interconnected electrodes may be determined using information about stress forces that are applied to the plurality of piezoelectric wires. 5. The system of claim 1 further comprising a localization system to determine relative distances between at least two regions of electrodes. 6. The system of claim 5 wherein the localization system comprises a device that can determine position information for each of the plurality of interconnected electrodes, said position information comprising interelectrode spacings, such that three dimensional positions of each of the plurality of interconnected electrodes may be calculated. 7. The system of claim 1 wherein the garment comprises: a plurality of electrodes arranged in a plurality of rows, wherein the plurality of rows may be placed about a circumference of the human body; anda plurality of spacers designed to locate at least one of said plurality of rows at fixed distances from adjacent rows;wherein the lengths of said plurality of spacers is known to the system. 8. The system of claim 7 wherein each of said plurality of rows is adjustable such that a circumference measurement for each row of electrodes of the plurality of rows may be determined by the system. 9. ... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/04
US08454538 | B2 | 2013-06-04 | Filtering method for surface modeling | 2011-01-28 | A system for mapping a tissue surface includes a probe for mapping a tissue surface, a localization system to measure a location data point indicative of the probe's location, a memory in which to store the location data point, a servo mechanism to move the probe along at least a portion of the tissue surface, a controller to move the probe to a plurality of locations and to record in the memory a plurality of location data points, and a contact-sensing processor to analyze the plurality of location data points and to identify a subset thereof on the tissue surface. A modeling processor generates a model of the tissue surface using the subset of location data points. The contact-sensing processor utilizes probe velocity, or a rate of change in the distance moved by the probe, to determine contact between the probe and the tissue surface. | 1. A method of monitoring a degree of contact between a probe and a tissue surface within a patient's body, comprising: moving the probe within the patient's body;measuring a location of the probe a plurality of times as it moves; andgenerating a signal indicative of a change in a degree of contact between the probe and the tissue surface based upon a distance moved by the probe over time,wherein the step of generating a signal indicative of a change in a degree of contact between the probe and the tissue surface based upon a distance moved by the probe over time comprises: converting the plurality of measured locations of the probe into a first output signal;filtering the plurality of measured locations of the probe using a low pass filter to create a second output signal;comparing the first output signal with the second output signal; andgenerating a signal indicative of a change in a degree of contact between the probe and the tissue surface if the first output signal and the second output signal diverge by more than an expected amount. 2. The method according to claim 1, wherein the step of generating a signal indicative of a change in a degree of contact between the probe and the tissue surface based upon a distance moved by the probe over time comprises: determining a rate of change of the location of the probe with respect to time; andgenerating a signal indicative of a change in a degree of contact between the probe and the tissue surface based upon the rate of change of the location of the probe with respect to time. 3. The method according to claim 2, wherein the step of generating a signal indicative of a change in a degree of contact between the probe and the tissue surface based upon the rate of change of the location of the probe with respect to time comprises generating a signal indicative of an increase in the degree of contact between the probe and the tissue surface when the rate of change of the location of the probe with respect to time is less than an expected rate of change. 4. The method according to claim 2, wherein the step of generating a signal indicative of a change in a degree of contact between the probe and the tissue surface based upon the rate of change of the location of the probe with respect to time comprises generating a signal indicative of a decrease in the degree of contact between the probe and the tissue surface when the rate of change of the location of the probe with respect to time is greater than an expected rate of change. 5. The method according to claim 2, wherein the step of determining a rate of change of the location of the probe with respect to time comprises measuring a velocity of the probe. 6. The method according to claim 2, wherein the step of determining a rate of change of the location of the probe with respect to time comprises computing a velocity of the probe from successive measured locations of the probe. 7. The method according to claim 1, wherein the step of generating a sig... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/103
US08676304 | B2 | 2014-03-18 | Ischemia monitoring system for patients having periods of left bundle branch block | 2010-06-01 | A device for detecting cardiac ischemia is disclosed. The device includes a processor that is configured to distinguish between two different heart beats types such as left bundle branch block beats and normal sinus beats. The processor applies different ischemia tests to the two different beat types, and generates alert when it detects ischemia. | 1. A cardiac monitoring device configured to analyze electrical signals from the heart to detect ischemic events, including: at least one sensor designed to collect the electrical signal from the heart;electrical circuitry coupled to the sensor and configured to amplify and digitize the electrical signal to produce sensed data;memory storage coupled to the electrical circuitry to store the sensed data;a processor configured to: analyze the sensed data in the memory storage;based on the sensed data, distinguish between sinus beats, left bundle branch block beats and other types of beats, and measure at least one feature of beats associated with sinus beats and left bundle branch block beats during a predetermined time interval, respectively; anddetect the presence or absence of an ischemic event by applying a first ischemia detection criterion to the sinus beat measured feature and a second ischemia detection criterion to the left bundle branch block measured feature dependent upon a number of sinus beats and left bundle branch block beats during said predetermined time interval;wherein (1) when said number of left bundle branch block beats are less than a first threshold value, then apply said first ischemia detection criterion to said sinus beat measured feature, (2) when said number of left bundle branch block beats are greater than a second threshold value, then apply said second ischemia detection criterion to said left bundle branch block measured feature, and (3) when said number of left bundle branch block beats and said sinus beats are greater than a third threshold value, respectively then apply said first and second ischemic detection criteria to both said left bundle branch block measured feature and said sinus beat measured feature. 2. The device of claim 1, wherein the processor is configured to distinguish other types of beats from both sinus and left bundle branch block beats by computing the value of a heart signal parameter associated with an early portion of a QRS complex and comparing the value of the heart signal parameter to an expected value that applies to both sinus and left bundle branch beats. 3. The device of claim 2 wherein the heart signal parameter is a function of one of: Q wave amplitude, Q wave slope, R wave amplitude, R wave slope, or duration between the Q and R waves. 4. The device of claim 1, wherein the processor is configured to distinguish between sinus and left bundle branch block beats by computing the value of a heart signal parameter associated with a late portion of a QRS complex and comparing the value of the heart signal parameter to an expected value that applies to sinus beats exclusively or left bundle branch beats exclusively. 5. The device of claim 4 wherein the heart signal parameter is a function of one of: S wave slope, ratio of S wave slope to R wave slope, duration between the R and S waves, or duration of the S wave. 6. The device of claim 5 wherein the duration of the S wave is a fun... | Angel Medical Systems, Inc. | A61B005/02 A61B005/04 A61N001/00
US08583210 | B2 | 2013-11-12 | Apparatus and method for communication between MRI system and implantable medical devices | 2010-12-18 | A device and a method for handling and withstanding electromagnetic fields, specifically such fields as occur in magnetic resonance tomography examinations (i.e., MRT or MRI). This refers in particular to an IMD that can transmit data and/or parameters to an MRT device. | 1. An implantable medical device comprising: a memory configured to hold information;at least one antenna and/or at least one telemetry antenna;a communications device connectable to the at least one antenna and/or the at least one telemetry antenna, wherein said communications device is configured to communicate with a magnetic resonance imaging (MRI) device;at least one sensor or indicator configured to detect electromagnetic interference fields;a control unit coupled with said memory, said at least one antenna and/or at least one telemetry antenna, said communications device and said at least one sensor or indicator configured to detect electromagnetic interference fields;wherein said at least one sensor or indicator is further configured to send a first signal to said control unit;wherein the control unit is configured to receive said first signal from said at least one sensor or indicator configured to detect electromagnetic interference fields whenever an electromagnetic interference field is detected,receive a second signal transmitted from the MRI device through the communications device, andcontrol the communications device to emit a third signal through the at least one antenna and/or the at least one telemetry antenna, wherein said third signal is interpretable by the MRI device, wherein the control unit initiates the transmission of the third signal either directly after said control unit receives the first and/or second signal, or after said control unit receives the first and/or the second signal using a time delay;wherein said transmission is effected through a targeted generation of an image interference with said third signal; andwherein the communications device is configured to provide communication between the magnetic resonance imaging device and the implantable medical device. 2. The implantable medical device according to claim 1, wherein the memory contains data and/or parameters configured to allow a risk-free MRI examination or an MRI examination with reduced risks, compared to an MRI examination performed without said communications device, wherein said data and/or parameters are transmitted to the magnetic resonance imaging device from the communications device with the third signal via the at least one antenna and/or the at least one telemetry antenna whenever an electromagnetic interference field has been detected and/or the communications device has detected a connection to the magnetic resonance imaging device. 3. The implantable medical device according to claim 1, wherein said at least one sensor or indicator configured to detect electromagnetic interference fields is further configured to transmit measured values to the control unit, and the control unit is configured to perform an adjustment of information and/or parameters in accordance with the measured values before the measured values are transmitted from the communications device to the magnetic resonance imaging device, wherein the adjustment is based o... | BIOTRONIK CRM Patent AG | A61B005/05
US08608678 | B2 | 2013-12-17 | Orthopedic technical device | 2009-05-19 | The invention relates to an orthopedic technical device for the lower extremities, with at least one receptacle (12, 16) for an extremity (14, 18), a sensor (30) and an electric evaluating unit (34), which is connected with the sensor (30) for processing data of a sensor (30). According to the invention, it is proposed that the electric evaluating unit (34) be set up in order to carry out a method with the following steps: (i) switching into a learning mode, (ii) determining at least one characteristic value of the system (R, S, N, aw,threshold), (iii) storing the characteristic value of the system (R, S, N, aw,threshold), (iv) switching into an operating method and (v), in the operating mode, influencing, optimizing or selecting a control algorithm of the orthopedic device by means of the characteristic value of the system (R, S, N, aw,threshold). | 1. An orthopedic technical device for the lower extremities, comprising: (a) a thigh receptacle for attachment to a thigh of a person,(b) an angle sensor to register an inclination with respect to the vertical and to emit an inclination signal,(c) an actuation device to release a switching command, and(d) an electrical evaluation unit, connected to the sensor, for processing data from the sensor, and to the actuation device,(e) wherein the electrical evaluation unit is configured to carry out a method with the following steps: (i) upon receiving the switching command from the actuation device, switching into a learning mode,(ii) establishing at least one fitting-error characteristic value that encodes a fitting error of the sensor,(iii) storing the fitting-error characteristic value,(iv) switching into an operating mode, and(v) correcting inclination sensor data from the angle sensor with the aid of the at least one fitting-error characteristic value in the operating mode. 2. The orthopedic technical device as claimed in claim 1, wherein the electrical evaluation unit is configured to establish the at least one fitting-error characteristic value only if it is in the learning mode. 3. The orthopedic technical device as claimed in claim 1, further comprising: (a) a shank receptacle for attachment to a shank, and(b) a locking and/or damping mechanism for releasably locking and/or damping a relative movement between thigh receptacle and shank receptacle,(c) wherein the orthopedic technical device is an orthosis, and(d) wherein the electrical evaluation unit has a digital storage medium and is designed to actuate the locking and/or damping mechanism as a function of the inclination signal in the operating mode for activating and for releasing the locking and/or damping mechanism and to carry out a method with the following steps: (i) switching into the learning mode,(ii) establishing the fitting-error characteristic value that encodes a positional deviation between a correct position of the angle sensor and an actual position of the angle sensor,(iii) storing the fitting-error characteristic value in the digital storage medium, and(iv) activating and releasing the locking and/or damping mechanism with the aid of the fitting-error characteristic value in the operating mode. 4. The orthopedic technical device as claimed in claim 3, wherein the electrical evaluation unit is configured to change the fitting-error characteristic value only when it is in the learning mode. 5. The orthopedic technical device as claimed in claim 3, wherein the electrical evaluation unit is configured to activate the locking and/or damping mechanism and to establish the fitting-error characteristic value while the locking and/or damping mechanism locks or are damped the thigh receptacle and the shank receptacle relative to one another. 6. The orthopedic technical device as claimed in claim 3, wherein a system characteristic value is a fitting-error characteristic value th... | Otto Bock Healthcare GmbH | A61F005/00
US08369915 | B2 | 2013-02-05 | Integrated miniaturized fiber optic probe | 2009-11-06 | A fiber optic probe having one or more photodetectors bound thereto is provided. By directly integrating thin, flexible photodetectors with an optical fiber, the probes provide a compact structure that increases throughput and decreases cost, making it practical for a clinical use. In some embodiments, the fiber optic probes are small enough for insertion into the shaft of a needle, such as a biopsy needle. | 1. An apparatus comprising: (a) light source;(b) an optical fiber having a distal end, a proximal end and a side-wall, the proximal end in optical communication with the light source, wherein the optical fiber is configured to emit light from its distal end, at least one location along its side-wall, or both;(c) at least one photodetector having a device layer bonded to and conforming to the curvature of the curved outer surface of the side-wall of the optical fiber, wherein the at least one photodetector is positioned such that light from the optical fiber will not directly impinge upon the at least one photodetector; and(d) an electrically conductive lead in electrical communication with at least one photodetector. 2. An apparatus comprising: (a) an optical fiber;(b) at least one photodetector having a device layer bonded to and conforming to the curvature of the curved outer surface of the side-wall of the optical fiber;(c) an electrically conductive lead in electrical communication with at least one photodetector; and(d) a layer of electrically conductive material disposed between the optical fiber and at least one photodetector, wherein the layer of electrically conductive material provides a back electrode for the at least one photodetector. 3. The apparatus of claim 1, further comprising a layer of dielectric material coating a portion of the photodetector. 4. The apparatus of claim 1, wherein the at least one photodetector is located no more than 10 mm from the distal end of the optical fiber. 5. The apparatus of claim 1, wherein the optical fiber is configured to emit light from its distal end and from at least one additional location along its side-wall, the apparatus further comprising at least one additional photodetector having a device layer bonded to and conforming to the curvature of the curved outer surface of the side-wall of the optical fiber in the vicinity of the at least one additional light-emitting location. 6. The apparatus of claim 1, comprising a plurality of photodetectors disposed around the circumference of the optical fiber, wherein each of the plurality of photodetectors is bonded to and conforms to the curvature of the curved outer surface of the side-wall of the optical fiber. 7. The apparatus of claim 6, wherein the optical fiber is configured to emit light at a plurality of locations around its circumference, and further wherein at least one of the plurality of photodetectors is located in the vicinity of each of the plurality of light-emitting locations. 8. The apparatus of claim 1, further comprising a biopsy needle containing the optical fiber. 9. A method for using the apparatus of claim 1, the method comprising: (a) placing the optical fiber in close proximity to a sample; and(b) emitting light from the optical fiber;wherein return light from the sample is detected by the at least one photodetector via absorption of said return light by the photodetector. 10. The method of claim 9, wherein the sam... | Wisconsin Alumni Research Foundation Duke University | A61B005/1459
US20110105874 | A1 | 2011-05-05 | METHOD FOR MEASURING BODY PARAMETERS | 2008-08-12 | A system comprises a capacitive sensing unit for capturing an electrophysiological signal from a body part. The capacitive sensing unit includes a first electrode plate that forms a capacitor with the body part. Motion of the electrode plate with respect to the body part may be detected by a motion sensitive unit mechanically coupled to the capacitive sensing unit. The motion sensitive unit detects motion by self-mixing interferometry. Then, a processing unit rejects the electrophysiological signal if a great displacement of the electrode plate is detected. | 1. A system comprising: a capacitive sensing unit comprising an electrode plate forming in combination with a body part a capacitor for capturing an electrophysiological signal from the body part;a motion sensitive unit mechanically coupled to the capacitive sensing unit for detecting by self-mixing interferometry a motion of the electrode plate with respect to the body part; and,a processing unit for altering the electrophysiological signal on the basis of the detected motion of the electrode plate. 2. The system of claim 1, wherein the processing unit rejects the electrophysiological signal when the detected motion lies within a predefined range. 3. The system of claim 1, wherein: the motion sensitive unit comprises: a light source for illuminating the body part;a cavity at the light source where light scattered by the body part interferes with light transmitted by the light source causing an interference signal representative of the motion of the electrode plate with respect to the body part. 4. A method for measuring body parameters comprising: detecting an electrophysiological signal from a capacitor formed from an electrode plate in combination with a body part;optically detecting a motion of the electrode plate with respect to the body part by optical means using self-mixing interferometry; andaltering the electrophysiological signal on the basis of the detected motion of the electrode plate. | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/05
US07957785 | B2 | 2011-06-07 | Waterproof bioelectrode | 2006-08-04 | A waterproof bioelectrode includes an electrode pad (101) to be mounted on a living body and a lead wire (110) to be connected to the electrode pad. The electrode pad includes a waterproof base (106) having an adhesive contacting surface and a hole substantially at its center, a waterproof seal member (105) which fixes the lead wire to a lower surface of the contacting surface of the base while covering the hole such that a detection electrode provided to a distal end of the lead wire is exposed from the hole of the waterproof base, and a conductive gel (103) arranged on the contacting surface to come into contact with the detection electrode. | 1. A disposable waterproof bioelectrode comprising a plurality of electrode pads to be attached on a living body anda plurality of waterproof lead wires respectively connected to said plurality of electrode pads with one-to-one relation, said plurality of waterproof lead wires being formed integrally,wherein each of said plurality of electrode pads comprises: a waterproof base member having an adhesive contacting surface and a hole substantially at a center thereof,a waterproof seal member having a notch substantially corresponding to a sectional shape of one of said plurality of waterproof lead wires in order to prevent water from entering from said one of said plurality of waterproof lead wires, wherein said waterproof seal member fixes said one of said plurality of waterproof lead wires to a back surface of said waterproof base member, wherein said back surface is opposite to said adhesive contacting surface, while covering the hole such that a detection electrode provided to a distal end of said lead wire is exposed through the hole of said waterproof base member, anda conductive gel arranged on said adhesive contacting surface of said waterproof base member to come into contact with said detection electrode, andwherein said plurality of waterproof lead wires comprises: a base film which is common to said plurality of waterproof lead wires and an electrode pattern is provided on one of an upper surface and a lower surface of the base film, andan upper foam material and a lower foam material each made of a water non-permeable flexible insulating material respectively provided on the upper surface and the lower surface of the base film. 2. The disposable waterproof bioelectrode according to claim 1, wherein ends of said plurality of lead wires which are not to be connected to said electrode pads form a connector. 3. The disposable waterproof bioelectrode according to claim 1, further comprising a first separator to protect said adhesive contacting surfaces of said plurality of electrode pads when said bioelectrode is not in use, wherein said first separator is common to said plurality of electrode pads. 4. The disposable waterproof bioelectrode according to claim 3, wherein said waterproof base member is made of a flexible film like material, and the waterproof bioelectrode further comprises a second separator provided to a part of said adhesive contacting surface of said waterproof base member,wherein said second separator is provided on at least a part of an outer border of said adhesive contacting surface. 5. The disposable waterproof bioelectrode according to claim 1, wherein at least said waterproof base member of said waterproof base member and said waterproof seal member is formed of a moisture-permeable waterproof film having a thickness of 20 μm to 70 μm. 6. The disposable waterproof bioelectrode according to claim 1, wherein said seal member fixes said lead wire to said waterproof base member from a position inside an outer circu... | Fukuda Denshi Co., Ltd. | A61B005/0408
US07957787 | B2 | 2011-06-07 | Method of examining dynamic cardiac electromagnetic activity and detection of cardiac functions using results thereof | 2008-07-25 | A method of examining cardiac electromagnetic activity over a heart for diagnosing the cardiac functions of the heart is disclosed. The method includes collecting a plurality of sets of spatially distributed, time-varying magnetic field signals of the heart of a subject, wherein the magnetic field signals exhibit features of at least a wave, identifying a time corresponding to a local maximum intensity of the magnetic field signals of the wave at each measurement position and plotting a temporal evolution of the local maximum intensity of the magnetic field signals during a time interval of the wave. | 1. A method of examining cardiac electromagnetic activity of a heart of a subject for diagnosing cardiac functions, the method comprising: performing a non-contact magnetic measurement over a thorax region of the subject at a plurality of measurement positions to obtain a plurality of sets of spatially distributed, time-varying magnetic field signals of the heart of the subject, wherein each set of the spatially distributed, time-varying magnetic field signals at each of the measurement positions exhibits features of at least a wave;offsetting each set of the time-varying magnetic field signals;identifying a time (tmax,chN) corresponding to a local positive or negative maximum intensity of the each set of the offset time-varying magnetic field signals of the wave at the each of the measurement positions; andplotting a temporal evolution of the local maximum intensity of the offset time-varying magnetic field signals during a time interval of the wave to obtain a propagation behavior of the wave, wherein the temporal evolution of the local maximum intensity of the offset time-varying magnetic field signals during the time interval of the wave is characterized by a time-dependent area ratio of a positive wave area to a negative wave area. 2. The method of claim 1, wherein the each set of the time-varying magnetic field signals at the each of the measurement positions exhibits the features of at least a P-wave, a Q-wave, a R-wave, a S-wave, a T-wave or a combination thereof corresponding to that of an electrocardiography. 3. The method of claim 1, wherein the wave is a T-wave. 4. The method of claim 1, wherein the each set of the time-varying magnetic field signals is offset by zeroing an interval of the magnetic field signals before a P-wave. 5. The method of claim 1, wherein the time-varying magnetic field signals comprise magnetic field components perpendicular (Bz) to a surface of the subject. 6. The method of claim 1, wherein the time-dependent area ratio starts from zero, then increases with time, and finally reaches a fixed value for a normal cardiac activity. 7. The method of claim 1, wherein an abnormal cardiac activity shows the time-dependent area ratio peaks at least once at a certain time before the time-dependent area ratio reaches a fixed value. 8. The method of claim 1, wherein the propagation behavior of the wave is applicable in diagnosing coronary artery diseases. 9. The method of claim 1, wherein the propagation behavior of the wave is useful applicable in localizing ischemic regions of the heart. | - | A61B005/05
US07941202 | B2 | 2011-05-10 | Modular multichannel microelectrode array and methods of making same | 2006-10-10 | Some embodiments of the invention comprise a customizable multichannel microelectrode array with a modular planar microfabricated electrode array attached to a carrier and a high density of recording and/or stimulation electrode sites disposed thereon. Novel methods of making and using same are also disclosed. | 1. A neuroprobe system comprising: (a) a carrier, and (b) at least one planar microelectrode array attached to the carrier, wherein the planar microelectrode array includes a substrate formed separately from the carrier and further includes: a plurality of interconnects disposed on the substrate between layers of dielectric material,a plurality of electrode sites, at least some of which are in contact with a corresponding interconnect through a corresponding aperture in a dielectrical layer that is configured to permit metallization beyond the perimeter of the aperture,a plurality of bond pads, at least some of which are in contact with a respective interconnect, andat least one lateral projection with an electrode site disposed thereon, the lateral projection configured to attach radially to the carrier,wherein at least one of the electrode sites is a recording site or a stimulation site. 2. The neuroprobe system of claim 1, wherein at least one of the electrode sites is a stimulation site. 3. The neuroprobe system of claim 1, wherein at least some of the interconnects comprise multilevel metal. 4. The neuroprobe system of claim 1, wherein at least one of the bond pads is configured to interface with an integrated circuit. 5. The neuroprobe system of claim 1, wherein the substrate comprises an internal channel configured to deliver a fluid therethrough to a target. 6. The neuroprobe system of claim 1, wherein the carrier comprises a polymer tube with a lumen. 7. The neuroprobe system of claim 6, further comprising a wire microelectrode for recording and/or stimulation which is disposed in the lumen. 8. The neuroprobe system of claim 6, further comprising a stiffener or a stylet insertable into the lumen. 9. The neuroprobe system of claim 8, wherein the stiffener or the stylet comprises a wire microelectrode. 10. The neuroprobe system of claim 1, wherein the carrier comprises a metal wire. 11. The neuroprobe system of claim 10, wherein the wire comprises a microelectrode for recording and/or stimulation. 12. The neuroprobe system of claim 1, wherein at least one of the electrode sites comprises a composite electrode site formed by grouping two or more smaller sites. 13. The neuroprobe system of claim 1, wherein at least one of the electrode sites is larger than its corresponding aperture. 14. The neuroprobe system of claim 1, wherein the lateral projection is configured to wrap around the carrier to form a ring electrode. 15. The neuroprobe system of claim 14, wherein at least one of the electrode sites comprises a composite electrode site formed by grouping two or more smaller sites. 16. The neuroprobe system of claim 15, wherein the composite electrode site is a strip along the microelectrode array. 17. The neuroprobe system of claim 15, wherein the substrate comprises an internal channel configured to deliver a fluid therethrough to a target. 18. The neuroprobe system of claim 17, wherein the carrier comprises a metal wire. 1... | NeuroNexus Technologies | A61B005/04 A61N001/05
US07945302 | B2 | 2011-05-17 | Wound mapping system | 2003-12-05 | A tissue mapping system comprising a two dimensional array of test electrodes 10 for application to the surface of tissue under investigation and circuit means 50-66 for measuring an electrical characteristic of the tissue underlying each test electrode. In one embodiment the electrical characteristics is the impedance of the tissue underlying each test electrode. | 1. A system for monitoring surface changes in a skin wound over time, the system comprising: a wound dressing including a two-dimensional array of test electrodes for application to the surface of the wound,a circuit electrically coupled to the two-dimensional array of test electrodes, the circuit for measuring an electrical characteristic of tissue immediately underlying each test electrode, said circuit comprising: a switching device for selecting successive test electrodes, andan analyser for passing an electrical alternating current between each selected test electrode and at least one further electrode applied to the skin at a location away from the wound, said analyser further calculating the electrical characteristic of the tissue immediately under each currently selected test electrode as a function of the voltage difference between the currently selected test electrode and at least one reference electrode, adjacent to the currently selected test electrode, the voltage difference being measured while passing the electrical alternating current between each currently selected test electrode and the at least one further electrode; anda display device electrically coupled to the analyser, the display device for presenting a visual map generated using analyser circuitry of said analyser, the map indicating the size and shape of the wound based upon the measured electrical characteristic. 2. A system as claimed in claim 1, wherein the array of test electrodes is arranged on a flexible backing of insulating material. 3. A system as claimed in claim 2, wherein the further electrode is disposed on the flexible backing of insulating material. 4. A system as claimed in claim 2, wherein the array of electrodes is a rectangular array. 5. A system as claimed in claim 2, wherein each test electrode is covered with a conductive gel, the resistance between adjacent test electrodes being high relative to the resistance via the gel between each test electrode and the underlying tissue. 6. A system as claimed in claim 5, wherein the gel is a hydrogel. 7. A system as claimed in claim 2, wherein leads for the test electrodes are also disposed on the flexible backing of insulating material and covered with an insulating material. 8. A system as claimed in claim 1, wherein the two-dimensional array comprises at least 25 test electrodes. 9. A system as claimed in claim 1, wherein the electrical characteristic is an impedance of the tissue immediately underlying each test electrode. 10. A system as claimed in claim 1, wherein for each test electrode a measurement is made at a plurality of different frequencies. 11. A system as claimed in claim 1, wherein each measurement is made at a frequency of from 1 milliHz to 100 kHz. 12. A system as claimed in claim 1, wherein said switching device for selecting successive test electrodes also selects, for the currently selected test electrode, at least one test electrode, adjacent to the currently selected test e... | University of Ulster | A61B005/04
US07937137 | B2 | 2011-05-03 | Methods of identifying and/or assessing cardiac arrhythmias using an implantable medical device | 2006-01-30 | Apparatus using one or more modes of statistical analysis with one or more monitored parameters of a patient's heart to identify and/or assess arrhythmias. Through use of the one or more modes of statistical analysis, a medical professional can be aided during evaluation of patient data for diagnosis of the patient. At least one of the monitored parameters may include one or more values used representatively for storage intervals of a selected length. As such, for each storage interval, a value may be determined for the one monitored parameter occurring at an upper percentile and a lower percentile. In addition, a median value may be determined for the one monitored parameter for each storage interval. Over a plurality of the storage intervals, these determined values can be used in one or more modes of statistical analysis to better identify and assess the arrhythmias. | 1. A non-transitory computer-readable medium programmed with instructions for performing a method of identifying and/or assessing patient arrhythmias, the medium comprising instructions for causing a programmable processor to: obtain sets of high, median, and low percentile measures of a heart parameter of a patient collected over a plurality of successive time intervals, each set of high, median, and low percentile measures corresponding to a respective time interval of the plurality of successive time intervals, the sets gathered from a memory of a medical device implanted within the patient;for each set, calculate a ratio corresponding to the magnitude of the gaps between the percentile measures; andusing each ratio, for the time interval associated with each ratio, determine at least one of an indication of an arrhythmia and at least a partial duration of the arrhythmia. 2. The computer readable medium of claim 1, wherein the indication of the arrhythmia for the time interval associated with each ratio involves a determination that such ratio is below a predetermined ratio threshold. 3. The computer readable medium of claim 2, further including instructions to increment a counter for any storage interval in which such ratio is below the predetermined ratio threshold, whereby the counter is used to indicate the number of storage intervals defining the corresponding arrhythmia; andreset the counter for any storage interval in which such ratio is above the predetermined ratio threshold. 4. The computer readable medium of claim 1, further including instructions to: aggregate each ratio with a dynamic base line ratio for the patient; determine whether the aggregate exceeds a high rate burden threshold; and provide an indication of the presence of an arrhythmia for the time intervals associated with the ratios in which the aggregate exceeds the high rate burden threshold. 5. The computer readable medium of claim 4, further including instructions to increment a counter for any storage interval in which the aggregate exceeds the predetermined high rate burden threshold, whereby the counter is used to indicate the number of storage intervals defining the corresponding arrhythmia; andreset the counter for any storage interval in which the aggregate is equal to or below the predetermined high rate burden threshold. 6. The computer readable medium of claim 4, wherein the aggregating step involves the use of a cumulative sum. | Medtronic, Inc. | A61B005/0452
US20110124978 | A1 | 2011-05-26 | HEALTH AND FITNESS SYSTEM | 2009-11-24 | The present invention provides a health and fitness system, comprising a means for providing access to the health and fitness application, a means for retrieving data from a health and fitness data source, a means for selectively integrating health and fitness data among multiple health and fitness applications, sensors, and devices, a means for logging health and fitness data, a means for purchasing a health and fitness item from a catalog of items, a means for sharing health and fitness data with a remote computing device, and a means for messaging to and from a smart phone. | 1. A health and fitness system, comprising: a means for providing access to the health and fitness application;a means for retrieving data from a health and fitness data source;a means for selectively integrating health and fitness data among multiple health and fitness applications, sensors, and devices;a means for logging health and fitness data;a means for purchasing a health and fitness item from a catalog of items;a means for sharing health and fitness data with a remote computing device; anda means for messaging to and from a smart phone. 2. The health and fitness system of claim 1, wherein the means for sharing health and fitness data includes interacting with a personal trainer website, the medical website, or the practice management software. 3. The health and fitness system of claim 1, wherein the health and fitness data source is a health and fitness device. 4. The health and fitness system of claim 3, wherein the device is a heart rate monitor, a blood pressure monitor, a glucose meter, a cycling sensor, a cycling computer, a rowing sensor, a rowing computer, a treadmill sensor, a treadmill computer, an elliptical trainer sensor, an elliptical trainer computer, a global-positioning system device, a running sensor, a pedometer, a scale, or a fitness watch. 5. The health and fitness system of claim 3, wherein the device is a mobile device configured to interact with a health monitoring device. 6. The health and fitness system of claim 1, wherein the health and fitness data source is a website. 7. The health and fitness system of claim 1, wherein the health and fitness item includes health monitoring equipment and health and fitness content. 8. The health and fitness system of claim 1, wherein the means for logging health and fitness data provides a personal health diary, a weight scale log, caloric intake, a blood pressure log, or glucose log. 9. The health and fitness system of claim 1, wherein the remote computing device is on the Internet. 10. The health and fitness system of claim 1, wherein the message is a positive psychological message, or a suggestion relating to health or fitness. 11. The health and fitness system of claim 1, further comprising a graphical user interface that interfaces with at least one of the means. 12. A computer program product comprising a computer useable medium having computer program code embodied therein configured to cause a computing device to: provide access a health and fitness application;retrieve data from a health and fitness data source;selectively integrate health and fitness data among multiple health and fitness applications, sensors, and devices;log health and fitness data;purchase a health and fitness item from a catalog of items;communicate health and fitness data with a remote computing device; andsend messages to and receive messages from a smart phone. 13. The computer program product of claim 12, wherein the means for sharing health and fitness data includes interacting wi... | - | A61B005/00 G06Q030/00 G06Q050/00 G06F003/01 H04B001/38 G06F003/048
US20110112382 | A1 | 2011-05-12 | SYSTEMS AND METHODS FOR COMBINED PHYSIOLOGICAL SENSORS | 2010-11-12 | Systems and methods are provided for monitoring the physiological state of a subject. One or more physiological parameters of a subject may be determined from a photoplethysmograph (PPG) signal or signals obtained using at least one PPG sensor. In some embodiments, an electrical physiological signal (EPS) sensor may be located in or near a PPG sensor. A sensor configuration including both PPG sensors and EPS sensors may be advantageously used to detect a PPG signal or signals in combination with one or more EPS signal or signals. To reduce potential interference between an EPS sensor and a PPG sensor, fiber-optic input and output lines may be used to transmit optical signals from light generating circuitry and light detecting circuitry. In some embodiments, the generating and detecting circuitry may be located remotely from one another and may further be located remotely from the EPS sensor, PPG sensor, or both. | 1. A physiological sensor device comprising: a first sensor configured to receive a first electrical signal of a subject; anda second sensor configured to transmit to or receive from the subject a first optical signal, the second sensor comprising: an optical aperture configured to transmit or receive the first optical signal, wherein the second sensor is capable of being coupled with circuitry for generating the first optical signal or converting the first optical signal into a second electrical signal and wherein the circuitry is disposed remote from the first sensor to reduce electrical interference between the circuitry and the first sensor. 2. The device of claim 1, wherein the optical aperture is disposed near the first sensor. 3. The device of claim 1, wherein the first sensor is an EEG sensor, and the first electrical signal is an EEG signal. 4. The device of claim 1, wherein the second sensor is a PPG sensor. 5. The device of claim 1, wherein the optical aperture and the circuitry are coupled by at least one fiber optic line. 6. The device of claim 1, wherein the first optical signal is transmitted by the second sensor into the subject, and wherein the second sensor is configured to receive a second optical signal from the subject. 7. The device of claim 6, wherein the first optical signal is generated by the circuitry, and wherein the second optical signal is converted into the second electrical signal by the circuitry. 8. The device of claim 6, wherein the first optical signal includes light of at least one wavelength. 9. The device of claim 6, wherein the first optical signal includes light of a plurality of wavelengths. 10. The device of claim 1, wherein the circuitry is included as part of the second sensor. 11. The device of claim 1, wherein the circuitry is included in a monitor remote from the second sensor. 12. A method for receiving physiological signals, comprising: receiving a first electrical signal of a subject with a first sensor; andtransmitting a first optical signal into the subject with a second sensor, the second sensor comprising: an optical aperture configured to transmit the first optical signal, wherein the second sensor is capable of being coupled with circuitry for generating the first optical signal and wherein the circuitry is disposed remote from the first sensor to reduce electrical interference between the circuitry and the first sensor. 13. The method of claim 12, wherein the optical aperture is disposed near the first sensor. 14. The method of claim 12, wherein the first sensor is an EEG sensor, and the first electrical signal is an EEG signal. 15. The method of claim 12, wherein the second sensor is a PPG sensor. 16. The method of claim 12, wherein the optical aperture and the circuitry are coupled by at least one fiber optic line. 17. The method of claim 12, wherein the first optical signal includes light of at least one wavelength. 18. The method of claim 12, wherein the first optical ... | Nellcor Puritan Bennett LLC | A61B005/00
US20110130646 | A1 | 2011-06-02 | PATIENT SUPPORT DEVICE, SUCH AS A PATIENT BED, TABLE OR CHAIR FOR USE WITH MAGNETIC RESONANCE IMAGING APPARATUSES | 2010-11-01 | A patient support device for use with a Magnetic Resonance Imaging apparatus includes a patient support surface and at least one receptacle for housing and/or removably coupling a coil adapted to receive signals from anatomic regions upon excitation thereof by the Magnetic Resonance Imaging apparatus, the at least one receptacle being arranged in an area of the patient support surface corresponding to a position of an anatomic region to be imaged. The receptacle is adapted for removable fastening of removably connectable bases of the receiver coil, and the base of each coil and its respective housing or coupling receptacle include complementary mechanisms for electrical connection to at least one dedicated electrical line such that the electrical connection is automatically effected when the base of each coil is fitted and/or mechanically fastened in a respective housing or coupling receptacle. | 1. A patient support device for use with Magnetic Resonance Imaging apparatuses, the device comprising: a patient support surface having such a size as to accommodate at least a part of the patient body;at least one receptacle for housing and/or removably coupling a coil adapted to receive signals from anatomic regions upon excitation thereof by the Magnetic Resonance Imaging apparatus, said at least one receptacle being arranged in an area of the patient support surface corresponding to a position of an anatomic region to be imaged, wherein the receptacle for housing and/or removably coupling the coil is formed by a niche or an enclosure with an open upper side extending flush with the patient support surface and wherein the coil is completely embedded inside the receptacle without protruding above the patient support surface;wherein the receptacle is adapted for removable fastening of removably connectable bases of the receiver coil;wherein the base of each coil and its respective housing or coupling receptacle include complementary means for electrical connection of the coil to at least one dedicated electrical line which transmits a signal of the receiver coil outside the devicewherein said means for electrical connection of the coil are constructed in such a manner that electrical connection is automatically effected when the base of each coil is fitted and/or mechanically fastened in a respective housing or coupling receptacle. 2. A device as claimed in claim 1, wherein the receptacle for housing/coupling the receiver coil is accessible from the patient support surface side. 3. A device as claimed in claim 1, wherein the housing/coupling receptacle is adapted for removable fastening of removably connectable bases of the receiver coil. 4. A device according to claim 1, wherein each receptacle provides hidden housings for said coils that are accessible from outside the device and beneath the patient support surface, whereby the coils are hiddenly received beneath the patient support surface. 5. A device as claimed in claim 1, wherein the base of the coil and its respective housing or coupling receptacle include complementary mechanical means for removable connection of the base and the receptacle. 6. A device as claimed in claim 1, wherein the electrical line for transmitting the signal of the receiver coil connects the receiver coil to an input of at least one controller of the Magnetic Resonance Imaging apparatus, which is designed to process the electric signal from the receiver coil. 7. A device as claimed in claim 1, wherein the base of each coil has an electrical output connector that corresponds with an electrical input connector of the housing and coupling receptacle, and wherein the electrical output connectors and input connectors of the housing and coupling receptacle are complementary and configured for mutual engagement and disengagement, each housing and/or coupling receptacle of each coil having said input connector conne... | ESAOTE S.P.A. | A61B005/055
US09215999 | B2 | 2015-12-22 | EEG monitor of fetal health | 2008-03-28 | A method for predicting health compromise in a fetus comprises acquiring one or more EEG signals and a fetal heart rate (FHR) signal from the surface of the head of a fetus; and determining the spectral edge frequency (SEF) of the one or more EEG signals. A repetitive temporal correlation between the FHR signal and the SEF of the one or more EEG signals is indicative of fetal health compromise. A system for detecting the correlation between FHR signal and the SEF of the one or more EEG signals and predicting fetal health compromise is also described. | 1. A method for the early prediction of fetal health compromise due to fetal hypoxia/asphyxia, comprising: (a) acquiring and identifying patterns in a Fetal Heart Rate (FHR) signal, wherein identifying patterns in the FHR signal comprises determining minima of the FHR signal and first time intervals between said minima;(b) acquiring one or more electroencephalogram (EEG) signals from at least one sensor placed on the surface of the head of a fetus and identifying patterns in the one or more EEG signals by performing an analysis based on a spectral edge frequency (SEF) of the one or more EEG signals, wherein identifying patterns in the one or more EEG signals comprises determining maxima of the SEF of the one or more EEG signals and second time intervals between said maxima;(c) determining a repetitive temporal correlation between the patterns in the FHR signal and the patterns in the one or more EEG signal by determining a correlation value between the first time intervals and the second time intervals; and(d) predicting the fetal health compromise based on the repetitive temporal correlation. 2. The method of claim 1, wherein the minima of the FHR signal comprise nadirs of FHR deceleration and the maxima of the SEF of the one or more EEG signals comprise peaks of the SEF. 3. The method of claim 1, wherein the correlation value is indicative of fetal health compromise when the correlation value is above 0.4. 4. The method of claim 1, wherein the correlation value is indicative of fetal health compromise when the correlation value is above 0.5. 5. The method of claim 1, wherein the SEF of the one or more EEG signals is determined in a frequency domain below 40 Hz. 6. The method of claim 1, wherein the SEF of the one or more EEG signals is determined in a frequency domain below 30 Hz. 7. The method of claim 1, wherein performing the analysis based on SEF of the one or more EEG signals comprises determining the frequency below which about 90% of the EEG activity occurs. 8. The method of claim 1, wherein performing the analysis based on SEF of the one or more EEG signals comprises determining the frequency below which about 95% of the EEG activity occurs. 9. The method of claim 1, wherein the one or more EEG signals are acquired from one or more electrodes placed on the fetus' head over the parasagittal region/fronto-parietal-occipital cortex. 10. The method of claim 1, wherein the fetal health compromise is fetal acidosis. 11. The method of claim 1, wherein the fetal health compromise is a cerebral fetal compromise. 12. The method of claim 1, further comprising applying to the fetus a therapy against development of acidosis after the FHR minima becomes temporally correlated with the maxima of the SEF. 13. The method of claim 12, wherein the therapy is expediting the delivery of the fetus. 14. The method of claim 1, wherein the FHR signal is derived from an electro-cardiogram (ECG) signal. 15. The method of claim 1, wherein the one or m... | The London Health Sciences Centre Research Inc. | A61B005/048 A61B005/00 A61B005/024 B63B029/06 B63C011/02 F16M011/24 F16M013/02 A61B005/0448
US07955277 | B2 | 2011-06-07 | Method and apparatus for physiological testing | 2008-12-31 | A device and method to test the level of physical effort of an individual is disclosed. The device has a moveable member which may be a shaft, a sensor for measuring force applied to the shaft; and a resistance means to vary the resistance to movement of the moveable member such as an electric brake. The relationship between the varied resistive force applied to the moveable member and the measured force may be determined to establish the level of effort of the user. Typically the resistance will be rapidly varied at a rate of five to ten times per second which is too quick for a person being tested and providing a sub-maximal effort, to sense a change in resistance and compensate for that change. | 1. A method of measuring a person's joint performance and simultaneously measuring sincerity of effort comprising: providing a movable member to be moved by a single movement of one or more joints of the person; rapidly varying the resistance to movement of the member at least twice during the single movement wherein the rate of change of resistance is at least twice per second such that the person cannot react repeatedly and consistently to the change in resistance due to physiological constraints;measuring the force applied to the moveable member by the person during at least the single movement,recording the resistance applied to the movable member;inputting the resistance applied and the measured force applied into a means for calculating the correlation between the measured force applied to the movable member during the single movement and the resistance to movement of the movable member; andoutputting a coefficient of correlation indicative of the correlation between the measured force applied to the movable member during the single movement and the resistance to movement of the movable member, wherein the coefficient of correlation provides an indication of the person's sincerity of effort in moving the movable member. 2. The method of claim 1, wherein the method is carried out over a single movement of the person's joint. 3. A method of sincerity of effort testing of a person's joint performance by measuring the torque applied to a shaft by one of the person's joints comprising: applying a varying resistance to turning of the shaft, the resistance to turning of the shaft being rapidly varied at least twice during a single movement wherein the rate of change of resistance is at least twice per second such that the person cannot react repeatedly and consistently to the change in resistance due to physiological constraints;measuring the torque applied to the shaft by the person during at least the single-movement; recording the resistance applied to the shaft; and inputting the resistance applied and the measured torque applied into a means for calculating the correlation between the measured torque applied to the shaft during the single movement and the resistance to movement of the shaft; andoutputting a coefficient of correlation indicative of the correlation between the measured torque applied to the shaft during the single movement and the resistance to movement of the shaft, wherein the coefficient of correlation provides an indication of the person's sincerity of effort in moving the shaft. 4. The method of claim 2, wherein the method is carried out over a single movement of the person's joint. 5. A method of measuring a person's joint performance and simultaneously measuring level of effort comprising: providing a movable member having a variable resistance to movement, wherein the person applies a force to the moveable member to move the moveable member;rapidly varying the resistance to movement of the moveable member at least t... | Tek Solutions PTY LTD | A61B005/103 A61B005/117
US07951062 | B2 | 2011-05-31 | Blood pump system and method of operation | 2008-06-05 | A blood pump system includes a blood pump having a motor with a rotor and a stator. The stator has a plurality of stator windings situated therein. A motor controller is coupled to the motor, and a processor has inputs coupled to the motor controller for receiving a time continuous signal from the pump. The processor is programmed to transform the time continuous signal to the frequency domain, and control the pump and detect excess suction in response to the time continuous signal in the frequency domain. | 1. A method of controlling a blood pump, comprising: sampling a time continuous signal from the blood pump to obtain a first number of data points;transforming a second number of data points relating to the sampled time continuous signal to the frequency domain, wherein the second number is greater than the first number;analyzing the sampled time continuous signal in the frequency domain;controlling the blood pump in response to the analysis of the sampled time continuous signal in the frequency domain; anddetecting excess suction in response to analysis of distortion in the sampled time continuous signal in the frequency domain. 2. The method of claim 1, further comprising determining parametric data in response to the analysis of the sampled time continuous signal in the frequency domain. 3. The method of claim 2, wherein the parametric data includes heart rate. 4. The method of claim 2, wherein the parametric data includes respiratory rate. 5. The method of claim 2, wherein the parametric data includes pump flow rate. 6. The method of claim 1, wherein sampling the time continuous signal from the blood pump includes sampling between approximately 100 and 200 data points of the time continuous signal. 7. The method of claim 1, further comprising zero padding the first number of data points to obtain the second number of data points. 8. The method of claim 1, further comprising validating the sampled time continuous signal in response to the analysis of the sampled time continuous signal in the frequency domain. 9. The method of claim 8, wherein validating the sampled time continuous signal includes evaluating the signal to noise ratio. 10. The method of claim 8, wherein validating the sampled time continuous signal includes evaluating the signal to noise plus distortion ratio. 11. The method of claim 1, wherein the time continuous signal is selected from the group consisting of a pump flow rate, a pump speed, and a pump current. 12. The method of claim 1, wherein the time continuous signal is selected from the group consisting of a pump inlet pressure, a pump outlet pressure, and a differential pressure across the pump. 13. A blood pump system, comprising: a blood pump comprising a rotor and a stator, the stator including a plurality of stator windings;a controller operatively coupled to the pump;a processor operatively coupled to the controller and adapted to receive a first number of data points relating to a time continuous signal from the pump; andwherein the processor is programmed to transform a second number of data points relating to the time continuous signal to the frequency domain, wherein the second number is greater than the first number, and to detect excess suction in response to distortion in the transformed time continuous signal. 14. The blood pump system of claim 13, wherein the controller applies current to the stator windings in a sequence to create a rotating field, and wherein the time continuous signal includ... | Micromed Technology, Inc. | A61B005/00
US07970462 | B2 | 2011-06-28 | Implantable medical devices evaluating thorax impedance | 2007-05-29 | Implantable medical device with an impedance determination unit with constant current/voltage source having current feed terminals connected to electrodes for intracorporal placement which generates measuring current pulses having constant current/voltage, for causing a current through a body via intracorporally placed electrodes, a measuring unit for measuring voltage/current strength of voltage/current fed through body, an impedance value determination unit connected to the current/voltage source and adapted to determine an impedance value for each measuring current pulse, and an impedance measuring control and evaluation unit connected to the impedance determination unit which controls the unit and evaluates a sequence of consecutive impedance values, the impedance determination unit further adapted to determine at least intrathoracic and intracardiac impedance values for same period of time, the intrathoracic values sampled with a lower sampling rate than the intracardiac values. | 1. An implantable medical device comprising: an impedance determination unit comprising a constant current source or a constant voltage source having current feed terminals that are connected or can be connected to electrodes for intracorporal placement and that is configured to generate measuring current pulses having a constant current strength or a constant voltage, respectively, to cause a current to be fed through a body via intracorporally placed electrodes;a measuring unit configured to measure a voltage corresponding to a current fed through a body by said constant current source or a current strength of a current fed through a body by said constant voltage source, respectively;an impedance value determination unit that is connected to said constant current source or said constant voltage source and to the measuring unit wherein said impedance value determination unit is configured to determine an impedance value for each measuring current pulse; and,an impedance measuring control and evaluation unit, that is connected to said impedance determination unit and that is configured to control said impedance determination unit and to evaluate a sequence of consecutive impedance values determined by said impedance determination unit; andsaid impedance determination unit further configured to determine at least intrathoracic impedance values and intracardiac impedance values concurrently, wherein the intrathoracic impedance values are sampled with a lower sampling rate than the intracardiac impedance values,determine respiration rate values from said intrathoracic impedance values,determine heart rate values from said intracardiac impedance values,evaluate said sequence of consecutive impedance values based on said respiration rate values and said heart rate values. 2. The implantable medical device according to claim 1, wherein the impedance determination unit is further configured to determine at least two kinds of intrathoracic impedance values comprising: a first kind corresponding to an impedance between an IMD case and at least one electrode to be placed in one ventricle of a heart; and,a second kind corresponding to an impedance between an IMD case at least one intracardiac electrode. 3. The implantable medical device according to claim 1, wherein the impedance determination unit is further configured to determine at least two kinds of intrathoracic impedance values comprising: a first kind corresponding to an impedance between an IMD case and an electrode to be placed in a right ventricle of a heart, anda second kind corresponding to an impedance between an IMD case and an electrode to be placed in a left ventricle of a heart. 4. The implantable medical device according to claim 2 wherein the impedance determination unit is further configured to determine said intrathoracic impedance values wherein: said constant current source is connected to the IMD case and a right ventricular or a left ventricular ring electrode, respectively, suc... | Biotronik CRM Patent AG | A61B005/02 A61B005/08 A61B005/05 A61N001/00
US07945313 | B2 | 2011-05-17 | Method and device for monitoring heart rhythm in a vehicle | 2007-05-10 | An heart rhythm monitoring device for a vehicle, which determines whether a driver has an arrhythmia includes a vehicle state determining portion that determines whether the vehicle is stopped; an electrode arranged on a steering wheel in a position where the driver grips the steering wheel; an electrocardiogram waveform obtaining portion that obtains a first electrocardiogram waveform from the electrode; and a signal processing and calculating portion that determines whether the heart rhythm of the driver is erratic based on the first electrocardiogram waveform. When the vehicle is in motion, the signal processing and calculating portion determines whether the heart rhythm of the driver is erratic based on the waveform component that is strong with respect to noise in the first electrocardiogram waveform. | 1. A heart rhythm monitoring device for a vehicle, which determines whether a driver has an arrhythmia based on an electrocardiogram waveform having a waveform component that is strong with respect to noise and a waveform component that is not strong with respect to noise, comprising: a vehicle state determining portion that determines whether the vehicle is stopped;an electrode that is arranged on a steering wheel in a position where the driver grips the steering wheel;an electrocardiogram waveform obtaining portion that obtains a first electrocardiogram waveform from the electrode; anda signal processing and calculating portion that determines whether the heart rhythm of the driver is erratic based on the first electrocardiogram waveform,wherein when the vehicle state determining portion determines that the vehicle is not stopped, the signal processing and calculating portion determines whether the heart rhythm of the driver is erratic based on the waveform component that is strong with respect to noise in the first electrocardiogram waveform, andwherein when the vehicle state determining portion determines that the vehicle is stopped, a second electrocardiogram waveform is obtained and the signal processing and calculating portion determines whether the heart rhythm of the driver is erratic based on a waveform component other than the waveform component that is strong with respect to noise in the second electrocardiogram waveform. 2. The heart rhythm monitoring device according to claim 1, wherein the waveform component that is strong with respect to noise is an R-wave component of the first electrocardiogram waveform. 3. The heart rhythm monitoring device according to claim 1, further comprising: a vehicle stopping assist portion that assists the driver in stopping the vehicle,wherein the vehicle stopping assist portion assists the driver in stopping the vehicle when the vehicle state determining portion determines that the vehicle is not stopped and the signal processing and calculating portion determines that the heart rhythm of the driver is erratic. 4. The heart rhythm monitoring device according to claim 1, further comprising: a pulse wave sensor that obtains a pulse wave waveform from the electrode,wherein when the vehicle state determining portion determines that the vehicle is stopped, the determination of whether the heart rhythm of the driver is erratic is also based on the pulse wave waveform. 5. The heart rhythm monitoring device according to claim 1, further comprising: a template storing portion that stores at least one template having a reference electrocardiogram waveform;an electrocardiogram waveform comparing portion that determines whether the first electrocardiogram waveform resembles the reference electrocardiogram waveform of each of the at least one template stored in the template storing portion; anda template updating portion,wherein when the electrocardiogram waveform comparing portion determines that the first e... | Toyota Jidosha Kabushiki Kaisha Nippon Medical School Foundation Denso Corporation | A61B005/0408
US07946984 | B2 | 2011-05-24 | Transcutaneous analyte sensor | 2005-03-10 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A sensor system comprising: an analyte sensor configured for placement through a skin of a host;a mounting unit configured for placement on the skin of the host, wherein the mounting unit exhibits articulation against the skin of the host, such that, in use, the mounting unit conforms to movement of the host when placed adjacent to the skin of the host relative to the sensor, and wherein the sensor is configured to extend through the mounting unit and into the host; andan adhesive pad configured to adhere the mounting unit to the skin of the host. 2. The sensor system of claim 1, wherein the mounting unit comprises a base configured for fastening to the skin of the host, wherein the base provides the articulation against the skin of the host. 3. The sensor system of claim 1, wherein the mounting unit comprises a base portion and a sealing portion, wherein the sealing portion comprises a flexible material. 4. The sensor system of claim 3, wherein the flexible material comprises a watertight sealing member. 5. The sensor system of claim 4, wherein the sealing member comprises an elastomeric material. 6. The sensor system of claim 4, wherein the sealing member comprises silicone. 7. The sensor system of claim 4, wherein the sealing member is configured to seal at least one of a needle or the sensor at an opening. 8. The sensor system of claim 1, wherein the articulation is provided by a flexible plastic material. 9. The sensor system of claim 1, wherein the sensor is flexible. 10. A sensor system comprising: a glucose sensor configured for transcutaneous placement through a skin of a host; andan articulating mounting unit configured to provide electrical connection between the sensor and sensor electronics, wherein the articulating mounting unit comprises a body configured for placement adjacent to the skin of the host, wherein the sensor is configured to extend through an area of the body of the mounting unit and into the host, and wherein the body of the articulating mounting unit is configured to articulate with movement of the skin of the host when placed on the skin of the host. 11. The system of claim 10, further comprising an adhesive material configured to adhere the mounting unit to the skin of the host. 12. The sensor system of claim 10, wherein the body of the mounting unit comprises a base configured for fastening to the skin of the host, wherein the base is at least partially formed from a flexible material. 13. The sensor system of claim 10, further comprising a covering configured to be placed over at least a portion of the mounting unit. 14. The sensor system of claim 10, wherein the body of the mounting unit comprises a flexible plastic. 15. The sensor system of claim 10, wherein the sensor is flexible. | DexCom, Inc. | A61B005/00 A61B005/05
US09101278 | B2 | 2015-08-11 | RR interval based beat rejection for a cardiac monitor | 2011-02-11 | Disclosed is a system for the detection of cardiac events (a guardian system) that includes an implanted device called a cardiosaver, a physician's programmer and an external alarm system. The system is designed to provide early detection of cardiac events such as acute myocardial infarction or exercise induced myocardial ischemia caused by an increased heart rate or exertion. The system can also alert the patient with a less urgent alarm if a heart arrhythmia is detected. Using one or more detection algorithms, the cardiosaver can detect a change in the patient's electrogram that is indicative of a cardiac event within five minutes after it occurs and then automatically warn the patient that the event is occurring. To provide this warning, the guardian system includes an internal alarm sub-system (internal alarm means) within the cardiosaver and/or an external alarm system (external alarm means). If the guardian system is put into a pacemaker, the algorithm can utilize a different analysis of the electrogram depending on whether or not the pacemaker is actually pacing the heart. | 1. A system for detecting a cardiac event in a human patient the system including: (a) at least two sensors configured for obtaining an electrical signal from the patient's heart;(b) analog-to-digital converter circuitry for digitizing the signal to produce a plurality of data segments having a time duration that is between a range of 3 to 30 seconds; and,(c) a processor configured for: (1) processing at least one data segment to identify activity related to beats from the patient's heart in said data segment, each of said beats being a sub-segment of said at least one data segment and each of said beats having a measurable R-R interval from the preceding beat;(2) calculating an average R-R interval of each of the plurality of said sub-segments by averaging the R-R intervals of the beats in the at least one data segment;(3) identifying short beats within at least one of said sub-segment of said at least one data segment, a short beat being defined as a beat that has an R-R interval less than a predetermined value lower than the average R-R interval of said at least one data segment;(4) determining a plurality of values of a selected heart signal parameter for said at least one data segment by processing only the sub-segment beats that are not short beats;(5) storing the plurality of values;(6) determining whether a cardiac event has occurred by comparing a current value of the selected heart signal parameter in a current data segment with a selected heart signal value that is based on at least one of the plurality of selected heart signal parameter values from a set of previous data segments. 2. The system of claim 1 in which the sensors are electrodes adapted for implantation into a body of the patient. 3. The system of claim 1 further comprising a housing adapted for implantation into a body, and wherein the housing contains the analog-to-digital converter, the means for processing data segments, the means for calculating the average R-R interval, the means to identify short beats, and the means to detect the cardiac event. 4. A system for detecting a cardiac event in a human patient the system including: at least two electrodes adapted to be implanted in the patient for obtaining an electrical signal from the patient's heart, the electrical signal being an electrogram;a cardiosaver adapted to be implanted including: (a) analog-to-digital converter circuitry for digitizing the electrogram to produce a plurality of electrogram segments having a time duration that is between a range of 3 to 30 seconds;(b) a processor configured for:(1) processing at least one electrogram segment to identify beats of the electrogram in said electrogram segment, each of said beats being a sub-segment of said at least one electrogram segment having a measurable cardiac feature;(2) calculating an average of the cardiac feature within each sub-segment of said at least one electrogram segment by averaging the feature values of the beats in the at least one electrogr... | Angel Medical Systems, Inc. | A61B005/0452 A61B005/0468 A61B005/00
US08882759 | B2 | 2014-11-11 | Microwave ablation system with dielectric temperature probe | 2009-12-18 | An electromagnetic surgical ablation system having a generator adapted to selectively provide surgical ablation energy to an ablation probe, and methods of operating same, are disclosed. The system includes a controller operatively coupled to the generator, and a tissue sensor probe operatively coupled to the controller. The tissue sensor provides a tissue temperature and at least one tissue dielectric property to the controller. During an electromagnetic tissue ablation procedure, the controller monitors tissue temperature and the at least one tissue dielectric property to determine tissue status, and to activate and deactivate the generator in accordance therewith. | 1. An ablation system, comprising: a generator adapted to selectively provide ablation energy to an ablation probe;an ablation probe operably coupled to the generator and adapted to receive the ablation energy from the generator and to deliver the ablation energy to tissue;a controller operatively coupled to the generator, including: a processor;a memory operatively coupled to the processor;a database operatively coupled to the processor and configured to store a tissue status lookup table having a first dimension representative of tissue temperature, a second dimension representative of tissue permittivity, and a third dimension representative of tissue loss factor;a tissue sensor probe operatively coupled to the controller and having a distal end and a proximal end, including: a dielectric sensor disposed at a distal end of the tissue sensor probe and configured to transmit a test signal to tissue, receive at least a portion of a reflected test signal from tissue, and provide a dielectric sensor signal corresponding to a measured tissue permittivity and loss factor;a temperature sensor disposed at a distal end of the tissue sensor probe and configured to provide a temperature sensor signal corresponding to a measured temperature of tissue;wherein a dielectric sensor circuit is operatively coupled to the processor and is adapted to receive the dielectric sensor signal from the tissue sensor probe; and a temperature sensor circuit is operatively coupled to the processor and is adapted to receive the temperature sensor signal from the tissue sensor probe; andwherein the processor is configured to utilize each of the measured tissue temperature, the measured tissue permittivity, and the measured tissue loss factor as indices into the tissue status lookup table to identify a tissue status. 2. The ablation system in accordance with claim 1, wherein the dielectric sensor is selected from the group consisting of a time domain reflectometry dielectric sensor, a broadband vector reflectometry dielectric sensor, and a scattering parameter analysis dielectric sensor. 3. The ablation system in accordance with claim 1, wherein the temperature sensor is selected from the group consisting of a fluoroptic temperature sensor, a thermocouple, a thermistor, and an infrared thermometer. 4. The ablation system in accordance with claim 1, wherein the tissue sensor probe includes an absorbent sleeve disposed on an outer surface of the tissue sensor probe. 5. The ablation system in accordance with claim 1, wherein the tissue sensor probe includes a handle disposed at a proximal end of the tissue sensor probe. 6. The ablation system in accordance with claim 5, wherein the handle includes a grip-enhancing feature. 7. The ablation system in accordance with claim 1, further comprising an actuator operatively coupled to the controller and configured to selectively activate the generator. 8. The ablation system in accordance with claim 7, wherein the actuator is selec... | Covidien LP | A61B018/12 A61B018/18 A61B017/00
US08585568 | B2 | 2013-11-19 | Systems and methods for neuro-EEG synchronization therapy | 2010-11-09 | Described are methods, devices, and systems for a novel, inexpensive, easy to use therapy for a number of disorders. Described are methods and devices to treat disorders that involves no medication. Methods and devices described herein use alternating magnetic fields to gently “tune” the brain and affect mood, focus, and cognition of subjects. | 1. A method of treating a subject, comprising: (a) determining the intrinsic frequency (f) of the subject within a specified EEG band by: (1) obtaining EEG data of the subject's brain;(2) removing any DC component in the signal;(3) performing a Fast Fourier Transformation the EEG data to determine a FFT curve X(f); and(4) achieving a fitted Gaussian curve, A(f), of a local maximum of the FFT curve X(f) by: (A) using the equationA⁡(f)=Ge-(f-s)22⁢⁢d2, wherein G is gain, d is standard deviation, and s is the mean frequency based on the specified EEG band, and(B) determining a goodness of fit measure by optimizing using the following equationH⁡(X,A)=∑f=s-1⁢⁢Hzs+1⁢⁢Hz⁢X⁡(f)-A⁡(f), by: 1) estimating a first mean frequency, a first standard deviation, and first gain for a first optimizing loop,2) shifting the gain, G, up or down slightly from the first gain, determining a new gain resulting in a better fit than that of the first gain,3) shifting the standard deviation, d, up or down slightly from the first standard deviation, determining a new standard deviation resulting in a better fit than that of the first standard deviation,4) shifting the mean frequency, s, up or down slightly from the first mean frequency, determining a new mean frequency resulting in a better fit than that of the first mean frequency, and5) repeating steps 2), 3), and 4), in which the first gain, the first standard deviation, and the first mean frequency are replaced with a second gain, a second standard deviation, and a second mean frequency, respectively, until the three parameters are optimized, thereby generating an optimized mean frequency that is the intrinsic frequency;(b) comparing the intrinsic frequency from step (a) to an average intrinsic frequency of a healthy population database or comparing a Q factor to an average Q factor of a healthy population database, wherein comparing the Q factor comprises determining the Q factor of the intrinsic frequency within the specified band of the subject using at least one of the intrinsic frequency and the fitted Gaussian curve;(c) if the intrinsic frequency from step (a) is higher than the average intrinsic frequency of the healthy population database, shifting down the intrinsic frequency within the specified EEG band of the subject by applying a magnetic field close to a head of the subject, wherein said magnetic field has a frequency lower than the intrinsic frequency of the subject, or if the Q factor from step (b) is higher than the average Q factor of a healthy population database, shifting down the Q factor of the subject by applying a magnetic field close to a head of the subject, wherein said magnetic field comprises a plurality of frequencies around the intrinsic frequency of the subject; and(d) if the intrinsic frequency from step (a) is lower than the average intrinsic frequency of the healthy population database, shifting up the intrinsic frequency within the specified EEG band of the subject by applying a magnet... | NeoSync, Inc. | A61B018/18 A61B005/0476
US08372153 | B2 | 2013-02-12 | Method for detecting body parameters | 2011-01-27 | A method for detecting biometric parameters includes the steps of performing a bone graft procedure on a knee, providing a biometric sensor at the knee, and measuring a biometric parameter at the knee with the sensor. The sensor is capable of measuring parameters in an adjacent surrounding including pressure, tension, shear, relative position, and vascular flow. Data relating to the biometric parameter is transmitted to an external source and the data are analyzed to evaluate a biometric condition of the knee. A set of the sensors can be placed on the knee. | 1. A method for detecting biometric parameters, which comprises: performing a surgical procedure on at least one knee;implanting at least one biometric transceiver at the at least one knee;transmitting a first energy wave from the transceiver into a procedure area including at least one of the knee and an area adjacent the knee;quantitatively assessing the behavior of the energy wave with the transceiver;after transmitting the first energy wave: transmitting a second energy wave from the transceiver into the procedure area; andquantitatively assessing the behavior of the second energy wave;wherein at least one of the first and second energy waves are pulsed during transmission in a vibratory manner to stimulate the procedure area in accordance with at least one detected parameter of the procedure area; anddetermining a current status of the at least one parameter of the procedure area selected from the group consisting of pressure, tension, shear, relative position, bone density, fluid viscosity, temperature, strain, angular deformity, vibration, venous flow, lymphatic flow, load, torque, distance, tilt, shape, elasticity, motion, bearing wear, subsidence, bone integration, change in viscosity, particulate matter, kinematics, stability, and vascular flow with the transceiver based upon a comparison between the assessed behavior of the first and second energy waves. 2. The method according to claim 1, which further comprises: transmitting data relating to the at least one biometric parameter to an external source; andanalyzing the data to evaluate a biometric condition of the at least one knee. 3. The method according to claim 2, which further comprises, based upon the evaluation of the biometric condition of the at least one knee, changing a currently ongoing interoperative procedure at the procedure area. 4. The method according to claim 2, which further comprises, based upon the evaluation of the biometric condition of the at least one knee, chronically monitoring the anatomic condition relating to the procedure area. 5. The method according to claim 1, which further comprises providing a set of the transceivers on the at least one knee. 6. The method according to claim 1, which further comprises providing energy from outside the procedure area to the transceiver to power the transceiver and, thereby, create the energy wave and quantitatively assess the behavior of the energy wave. 7. The method according to claim 6, which further comprises providing the energy through at least one of an electromagnetic couple, a magnetic couple, a capacitive couple, an inductive couple, a sonic couple, an ultrasonic couple, a fiber optic couple, an optical couple, and an infrared couple. 8. The method according to claim 1, which further comprises: providing a set of the biometric transceivers at the procedure area;transmitting an energy wave from the transceivers into the procedure area;quantitatively assessing the behavior of the energy wave with at le... | - | A61F002/38
US08579824 | B2 | 2013-11-12 | Method and apparatus to monitor change in inspiratory effort using intrathoracic blood pressure waveform morphology | 2009-10-30 | An implantable medical device and associated method detect obstructed inspiration by monitoring an blood pressure signal. A respiration signal is monitored and a phase of respiratory inspiration is detected from the respiration signal. A trend in the pressure signal is measured during the inspiration phase. Obstructed inspiration for the inspiration phase is detected in response to the measured the trend. | 1. A method for detecting obstructed inspiration for use in an implantable medical device system, comprising: receiving by a processor of an implantable medical device a blood pressure signal from an implantable pressure sensor;receiving by the processor a respiration signal;detecting an onset and an end of a phase of respiratory inspiration from the respiration signal;detecting the pressure signal occurring only between the onset and the end of the detected inspiration phase;measuring a trend in the pressure signal detected during the inspiration phase;detecting an obstructed breath cycle for the inspiration phase in response to the trend;updating a counter in response to detecting the obstructed breath cycle; andenabling the processor to detect an obstructive sleep apnea episode in response to the counter reaching a detection threshold number of obstructed breath cycles. 2. The method of claim 1 wherein measuring the trend comprises measuring a decreasing diastolic pressure. 3. The method of claim 1 wherein measuring the trend comprises comparing a pressure measured from the pressure signal during the inspiration phase and a previous pressure measurement obtained during a previous inspiration phase. 4. The method of claim 1 wherein measuring the trend comprises measuring an inspiratory slope between a first pressure measurement and a second pressure measurement, the first and second pressure measurements occurring on two different cardiac cycles during the same inspiration phase. 5. The method of claim 4 wherein detecting the obstructed inspiration further comprises comparing the inspiratory slope to a previously measured inspiratory slope. 6. The method of claim 1 wherein measuring the trend in the pressure signal comprises: measuring a first blood pressure variable during the inspiration phase;detecting an expiration phase from the respiration signal;measuring a second blood pressure variable during the expiration phase; andcomputing a difference between the first blood pressure variable and the second blood pressure variable. 7. The method of claim 1 further comprising: setting an obstructive sleep apnea threshold;counting a number of consecutive inspiration phases detected as obstructed inspiration; anddetecting an episode of obstructive sleep apnea in response to the number exceeding the threshold. 8. The method of claim 7 further comprising storing one of a frequency and a duration of detected obstructive sleep apnea episodes. 9. The method of claim 1 further comprising adjusting a therapy in response to detecting the obstructed inspiration, the therapy comprising one of a sleep apnea therapy and a cardiac therapy. 10. The method of claim 1 further comprising: when an inspiration phase is not detected, measuring a blood pressure on each cardiac cycle of the blood pressure measurement;determining a beat-to-beat change in the blood pressure measurement;detecting a lack of inspiratory effort in response to the beat-to-beat change in... | Medtronic, Inc. | A61B005/02
US20110152723 | A1 | 2011-06-23 | BODY INDEX APPARATUS | 2010-12-16 | A body index apparatus (100) has a measurer for measuring, at plural measurement points, the distance indicating the width of abdomen (11) of a human subject (10) who lies on a reference plane (S) in the supine position. The distances from these measurement points to the reference plane (S) are different from one another. The measurement points are included in a plane (G) intersecting the reference plane (S) and are positioned so that the maximum width of the abdomen (11) in the plane (G) is between a measurement point that is the closest to the reference plane (S) and a measurement point that is the most distant from the reference plane (S). The body index apparatus (100) estimates an obesity index by calculation using the value of the ratio of the first distance to the second distance that is a distance at a predetermined measurement point, the first distance indicating the maximum width of plural distances measured at the plural measurement points. | 1. A body index apparatus comprising: a measurer for measuring, at plural measuring points, a distance indicating an abdominal width of a human subject lying on a reference plane in the supine position, the plural measurement points having different distances to the reference plane and being included in a plane intersecting the reference plane; andan estimator for estimating an index relating to obesity, by calculation using a variable relating to a difference between a first distance and a second distance, the first distance corresponding to the maximum width of plural distances at the plural measurement points measured by the measurer and the second distance corresponding to a distance of the width of a predetermined measurement point of the plural measurement points,wherein the plural measurement points are determined in such a way that the maximum value of the abdominal widths in the plane intersecting the reference plane is located between a measurement point closest to the reference plane and a measurement point that is the most distant from the reference plane. 2. A body index apparatus according to claim 1, wherein the estimator executes calculation using the first distance and the variable. 3. A body index apparatus according to claim 2, further comprising a bioelectrical impedance measurer for measuring a bioelectrical impedance of the abdomen by bringing plural electrodes into contact with the abdomen, wherein the estimator executes calculation using the bioelectrical impedance measured by the bioelectrical impedance measurer, the first distance, and the variable. 4. A body index apparatus according to claim 1, wherein the variable is the ratio between the first distance and the second distance. 5. A body index apparatus according to claim 1, wherein the variable is the difference between the first distance and the second distance. 6. A body index apparatus according to claim 1, wherein the predetermined measurement point is a measurement point that is the most distant from the reference plane among the plural measurement points. 7. A body index apparatus according to claim 6, wherein the estimator is capable of executing a first calculation in which the index is estimated when the predetermined measurement point is a measurement point that is the most distant from the reference plane and executing a second calculation in which the index is estimated when the predetermined measurement point is another measurement point, andwherein the estimator executes the second calculation in a case in which the distance of a measurement point that is measured by the measurer and is the most distant from the reference plane indicates zero, and executes the first calculation in a case in which the distance of a measurement point that is measured by the measurer and is the most distant from the reference plane does not indicate zero. 8. A body index apparatus according to claim 1, wherein at least one measurement point of the plural measurement... | Tanita Corporation | A61B005/103
US20110125036 | A1 | 2011-05-26 | THERAPEUTIC SYSTEM, THERAPEUTIC DEVICE, AND CONTROL METHOD | 2007-10-03 | To provide a therapeutic system which is based on the KAATSU training and which is suitable for treating metabolic syndrome.;A therapeutic system is comprised of a tight fitting device 100, a body segment 200, a measuring segment 300, and a control segment 400. The tight fitting device 100 is wrapped around a predetermined range on the limb. The tight fitting device 100 has an air-tight gas bag and can change a compression pressure to be applied to the limb by means of introducing and removing air into and from the gas bag. The body segment 200 controls the flow of the air into and from the gas bag. The measuring segment 300 is attached to the limb around which the tight fitting device 100 is wrapped and measures the amplitude of pulse wave. The control segment 400 determines an air pressure within the gas bag at the time point at which the amplitude of the pulse wave reaches its maximum, during preprocessing that is performed prior to the treatment, and uses it as the pressure within the gas bag for the treatment. | 1. A therapeutic system comprising: a tight fitting device including a belt having the length that is enough to be wrapped around a predetermined range of muscles of one of the limbs; fastening means for fastening said belt with said belt being wrapped around said predetermined range of muscles; and a gas bag provided in or on said belt, said gas bag being adapted to apply a predetermined compression pressure to said predetermined range of muscles by means of compressing said predetermined range of muscles when said gas bag is filled with gas while said belt that has been wrapped around said predetermined range of muscles is fastened by said fastening means;pressure setting means which is capable of setting a gas pressure within said gas bag to a predetermined pressure;control means for controlling said pressure setting means in order to change said compression pressure; andpulse wave measuring means for measuring a predetermined parameter that varies according to the variation of amplitude of arterial pulse wave that varies depending on said compression pressure, at a position near said predetermined range of muscles or a position closer to the distal end of the limb than there, to generate a pulse wave data associated with the parameter;said control means being adapted tomake said pressure setting means execute two processing, i.e., preprocessing and normal processing;control said pressure setting means during said preprocessing so that said pressure setting means changes the gas pressure within said gas bag, and determine a maximum pulse wave pressure that is a gas pressure within said gas bag at the time point at which the amplitude of the pulse wave has reached its maximum, by means of receiving two or more said pulse wave data from said pulse wave measuring means while the pressure within said gas bag is changing; andcontrol said pressure setting means during said normal processing so that said pressure setting means sets the gas pressure within said gas bag to said maximum pulse wave pressure. 2. The therapeutic system as claimed in claim 1, wherein said control means is adapted to control, during said preprocessing, said pressure setting means in such a manner that said pressure setting means once raises the pressure within said gas bag until it exceeds a pressure that is expected to be higher than the maximum pulse wave pressure and thereafter reduces the pressure within said gas bag. 3. The therapeutic system as claimed in claim 2, wherein said control means is adapted to continuously receive said pulse wave data from said pulse wave measuring means while the pressure within said gas bag is decreasing during said preprocessing, and adapted to determine the maximum pulse wave pressure at the time point at which the pulse wave has reached its maximum, from at least one preceding pulse wave data, when said pulse wave data indicates that the amplitude of said pulse wave becomes smaller than before. 4. The therapeutic system as claimed in... | Toshiaki Nakajima SATO SPORTS PLAZA CO., LTD. | A61B005/024
US20110105941 | A1 | 2011-05-05 | ELECTRODE CONNECTION MONITORING | 2008-09-22 | The present invention relates to an apparatus and a method for measuring muscular activity, in particular in relation to bruxism, and providing an electrical stimulation in response to the measured muscular activity through electrodes applied to the skin of an individual, wherein the quality of the connectivity of the electrodes to the skin is monitored and the electrical stimulation signal is changed based on said quality measurement. | 1-35. (canceled) 36. Apparatus comprising: an electrode assembly comprising three electrodes: a first, a second and a reference electrode, mounted on a common substrate in a fixed spatial relationship one to another wherein said reference electrode lies approximately equidistant from said first and second electrodes, said electrode assembly adapted for being applied to the skin of an individual, wherein the apparatus is operative to receive and monitor electrical signals and provide electrical stimulation to said individual, said apparatus further comprising:inputs for receiving signals from said electrodes,means for measuring EMG signals from said electrode(s) when the electrodes are applied to the skin,a source of a monitoring signal having a frequency and outputs for use in applying said monitoring signal alternately to said first and second electrodes,means for determining the amplitude of the frequency of the monitoring signal at the input of said reference electrode,means for providing an electrical stimulation feedback signal to the reference electrode upon the detection of EMG signals indicative of a muscular activity relating to bruxism of the individual,wherein said apparatus is configured so that said monitoring signal is applied at a voltage sufficient to produce a preset current between the first electrode and the reference electrode and between the second electrode and the reference electrode,wherein the amplitude of the frequency of the monitoring signal at said reference electrode is measured to determine the quality of the connection via the skin between the first electrode and the reference electrode and the quality of the connection via the skin between the second electrode and the reference electrode, andwherein the feedback signal is modulated if the quality of the electrical connection between the electrodes is changed. 37. Apparatus as claimed in claim 36, comprising a Fourier transform analyser for extracting frequency content information from said received signals. 38. Apparatus as claimed in claim 36, further comprising means for measuring the amplitude within the received signals of at least one selected frequency other than that of said monitoring signal. 39. Apparatus as claimed in claim 37, wherein said apparatus is for measuring EMG signals from said electrodes when the electrodes are applied to the skin and the amplitude of the said frequency of the monitoring signal in signals received at said inputs as determined by said Fourier transform analyser is measured for monitoring the quality of the electrical connection between the electrodes and the skin. 40. Apparatus as claimed in claim 36, wherein said EMG signals indicative of the muscular activity are identified by measuring therein the amplitude of one or more selected frequencies different from the frequency of the monitoring signal. 41. Apparatus as claimed in claim 36, wherein the feedback signal is reduced when the quality of the electrical connection ... | Medotech A/S | A61B005/0488
USRE042378 | E1 | 2011-05-17 | Implantable pressure sensors and methods for making and using them | 2006-07-20 | An implant includes a pressure sensor, a controller for acquiring pressure data from the sensor, and an acoustic transducer for converting energy between electrical energy and acoustic energy. A capacitor is coupled to the acoustic transducer for storing electrical energy converted by the transducer and/or for providing electrical energy to operate the implant. The acoustic transducer may operate alternatively or simultaneously as an energy exchanger or an acoustic transmitter. During use, the implant is implanted within a patient's body, and an external transducer transmits a first acoustic signal into the patient's body, to energize the capacitor. The implant then obtains pressure data, and transmits a second acoustic signal to the external transducer, the second acoustic signal including the pressure data. | 1. A surgicalAn implant, comprising: a sensor for measuring intra-body diagnostic data;a controller configured for generatingthat generates an electrical communication signal containing the measured diagnostic data;one or more acoustic transducers;circuitry for collectively configuring the one or more acoustic transducers to convert acoustic energy received from a location external to the implant into electrical energy used to support operation of the implant, and to convert the electrical communication signal receivedgenerated by the controller into an acousticalacoustic communication signal for transmissiontransmitted to a location external to the implant; andan energy storage device configured for storing the electrical energy converted from acoustic energy by the one or more transducers, wherein the energy storage device comprises a first relatively fast-charging capacitor and a second relatively slow-charging capacitor, the first and second capacitors being coupled to the one or more acoustic transducers such that the first capacitor is charged first and the second capacitor is charged only upon substantiallysubstantial charging of the first capacitor. 2. The implant of claim 1, wherein the one or more acoustic transducers are configured by the circuitry in a full-duplex mode, such that the one or more acoustic transducers can simultaneously convert the acoustic energy into electrical energy and convert the electrical communication signal into the acoustical communicationacoustic communication comunication signal. 3. The implant of claim 2, wherein the one or more transducers comprise at least one receive only transducer for converting the acoustic energy into electrical energy, and at least one transmit only transducer for converting the electrical communication signal into the acousticalacoustic communication signal. 4. The implant of claim 2, wherein the one or more transducers comprises at least one transducer, each of which is configured by the circuitry for converting the acoustic energy into electrical energy, and at least one transducer configured for converting the electrical communication signal into the acoustic communication signal. 5. The implant of claim 1, wherein the one or more acoustic transducers are configured by the circuitry in a half-duplex mode, such that the one or more acoustic transducers can alternately convert the acoustic energy into electrical energy, and convert the electrical communication signal into the acoustic communication signal. 6. The implant of claim 1, wherein the one or more transducers are collectively configured by the circuitry for converting an acoustic communication signal transmitted from a location external to the implant to another electrical communication signal, the controller configured for detecting the other electrical communication signal. 7. The implant of claim 6, wherein the controller is configured for extracting one or more commands from the other electrical communication s... | Remon Medical Technologies, Ltd. | A61B005/04
US08611981 | B2 | 2013-12-17 | MRI-CEST diagnostic technique based on non-punctual analysis | 2009-08-18 | An embodiment in the MRI-CEST field is proposed for analyzing a body-part, which includes a CEST agent providing a magnetization transfer with a bulk substrate of the body-part. A corresponding diagnostic system includes input means for providing an input map including a plurality of input elements each one for a corresponding location of the body-part; each input element is indicative of a spectrum of a magnetic response of the location, which spectrum includes the magnetic response at an agent frequency of resonance of the contrast agent (with the agent frequency that is at an agent offset of frequency from a bulk frequency of resonance of the bulk substance), and at a reference frequency at the opposite of the agent offset from the bulk frequency. The system further includes calculation means for calculating an agent value and a reference value for each one of a set of selected locations; the agent value is calculated according to the magnetic responses of the selected location in a non-punctual agent range of frequencies including the agent frequency, and the reference value is calculated according to the magnetic responses of the selected location in a non-punctual reference range of frequencies including the reference frequency (with the reference range that is symmetric to the agent range with respect to the bulk frequency). Comparison means is then provided for calculating a parametric value for each selected location; the parametric value is calculated according to a comparison between the agent value and the reference value of the selected location. | 1. An MRI-CEST diagnostic system, comprising: an input module configured to provide an input map including a plurality of input elements each one for a corresponding selected location of a body-part, each input element being indicative of a spectrum of a magnetic response of the corresponding selected location including the magnetic response at an agent frequency of resonance of a CEST agent providing a magnetization transfer with a bulk substrate of the body-part, the agent frequency being at an agent offset of frequency from a bulk frequency of resonance of the bulk substrate, and at a reference frequency at the opposite of the agent offset from the bulk frequency,a calculation module configured to calculate an agent value and a reference value for each one of a set of the selected locations, the agent value being calculated according to multiple magnetic responses of the selected location in a non-punctual agent range of frequencies centered around the agent frequency, and the reference value being calculated according to multiple magnetic responses of the selected location in a non-punctual reference range of frequencies centered around the reference frequency, the reference range being symmetric to the agent range with respect to the bulk frequency, anda comparator module configured to calculate a parametric value for each one of the set of selected locations, the parametric value being calculated according to a comparison between the agent value and the reference value of the selected location. 2. The system according to claim 1, wherein the agent range and the reference range extend around the bulk frequency from a first agent limit to a second agent limit being different from the agent frequency, and from a first reference limit to a second reference limit being different from the reference frequency, respectively. 3. The system according to claim 2, wherein the first agent limit and the first reference limit consist of the bulk frequency. 4. The system according to claim 1, wherein the calculation module includes an integrator module configured to calculate the agent value according to an integral of the magnetic responses of the selected location in the agent range, and to calculate the reference value according to an integral of the magnetic responses of the selected location in the reference range. 5. The system according to claim 1, wherein the calculation module includes a complement module configured to calculate the agent value according to a complement of an integral of the magnetic responses of the selected location in the agent range, and to calculate the reference value according to a complement of an integral of the magnetic responses of the selected location in the reference range. 6. The system according to claim 1, wherein the calculation module includes a module configured to calculate the parametric value according to an intermediate value, the intermediate value being based on a first ratio between the reference va... | Bracco Imaging S.p.A. | A61B005/05 G01V003/00
US08750450 | B2 | 2014-06-10 | Method for producing a dental 3D X-ray image, and X-ray device therefor | 2009-07-29 | In a method for the creation of a dental 3D radiograph of at least a subregion, forming a volume, of an object using an X-ray apparatus, the volume is recorded as a 3D radiograph from a plurality of projected images during a revolution around the object. At least a part of the object is displayed in a graphical representation before the radiograph of the volume is created, the position of the graphical representation relative to the current position of the apparatus and the patient being known. The volume to be radiographed, governed by the positioning of the object relatively to the apparatus and by the choice of the settings data and/or control data, is placed in the graphical representation in an at least approximately correct position, and when there is a change in position and/or size of the volume to be radiographed within the graphical representation, the settings data and/or control data for the creation of the 3D radiograph are defined. | 1. A method for creating a dental three-dimensional (3D) radiograph of at least one subregion, forming a volume, of an object using an X-ray apparatus, the method comprising: positioning the object relative to the apparatus;creating a 3D radiograph of the volume from a plurality of projected X-ray images during an at least partial revolution of a system of the X-ray apparatus including an X-ray emitter and an X-ray detector around at least a subregion of the object;prior to the creating of the radiograph of the volume, displaying at least a part of the object in a graphical representation, a position of the graphical representation relative to a current position of the apparatus and a current position of a patient being predetermined, the volume being governed by the positioning of the object relative to the apparatus and by at least one of settings data and control data, and the volume being shown on the object in the graphical representation and having a position and size that are selectively alterable; andautomatically adjusting at least one of the settings data and the control data for creating the 3D radiograph, in response to at least one of the position and size of the volume shown on the object within the graphical representation being altered,wherein the settings data and the control data define at least one of a position of an axis of rotation of at least part of the X-ray apparatus, a path of movement of at least part of the X-ray apparatus, a position of at least part of the X-ray apparatus, and a position of the object. 2. The method as defined in claim 1, wherein the graphical representation includes at least one of at least part of a surface and a contour of object. 3. The method as defined in claim 2, further comprising: calculating a three-dimensional contour from a plurality of optical images; anddisplaying the volume within the three-dimensional contour. 4. The method as defined in claim 2, further comprising: creating a plurality of images of the object from a plurality of directions; anddisplaying the plurality of images, wherein at least part of the volume is displayed in each of the images. 5. The method as defined in claim 2, further comprising: displaying a stored 3D image in addition to at least one of the surface and the contour of the object, anddisplaying the volume in a representation of at least one of the surface and the contour of the object and in the 3D image. 6. The method as defined in claim 5, further comprising at least one of: automatically aligning a position of the representation of at least one of the surface and the contour of the object and a position of the existing 3D image; andproviding an input unit that controls alignment of the position of the representation of at least one of the surface and the contour of the object and the position of the existing 3D image. 7. The method as defined in claim 1, wherein the graphical representation is an existing 3D image or part of the existing 3D image. ... | Sirona Dental Systems GmbH | A61B006/14 A61B005/07
US20110105869 | A1 | 2011-05-05 | Sensor for Internal Monitoring of Tissue O2 and/or pH/CO2 In Vivo | 2010-11-09 | Provided is a durable tissue pH/pCO2 and/or tissue oxygen sensitive probe of sufficient strength to withstand direct tissue pressures in vivo, the probe comprising one or more sensor chambers within a biocompatible, gas-permeable membrane containing together in a single chamber, or in separate chambers, respectively, a pH sensitive fluorophor from which pCO2 level(s) are calculated when the fluorophor is excited and the resulting fluorescence is measured and/or an oxygen sensitive phosphor solution producing oxygen quenchable phosphorescence when excited. Further provided is a tissue pH/pCO2 and/or tissue oxygen detection and measurement system comprising the probe, and methods for use of the probe and the system to directly, rapidly and accurately measure tissue pH/pCO2 and/or tissue oxygen levels in a patient without reliance on blood vessels or fluid protection of the probe. | 1. An tissue-insertable, in vivo system for real-time measurement of tissue pCO2 of an animal or human patient, the system comprising: a fiber optic sensor chamber forming a probe having the wall-strength to withstand external tissue pressure;a highly fluorescent, aqueous, buffered, pH-sensitive fluorophor, sealed within the fiber optic sensor chamber;an excitation light source for activating the fluorescence of the fluorophor; andan instrument for measuring and reporting fluorescence from the activated fluorophor in the fiber optic sensor chamber in place within the tissue of the patient, from which level(s) of pCO2 are calculated. 2. The system of claim 1, wherein the fluorophor comprises a pH sensitive porphyrin-based dye. 3. The system of claim 1, wherein the pH buffer is bicarbonate buffer, sealed within the fiber optic sensor chamber. 4. The system of claim 1, further comprising optical fibers to 1) transport the excitation light to the fluorophor, and 2) to transport the fluorescence from the fluorophor following excitation to the measuring and reporting instrument. 5. The system of claim 3, further comprising wireless connections to 1) transport the excitation light to the fluorophor, and 2) to transport the fluorescence from the fluorophor following excitation to the measuring and reporting instrument. 6. The system of claim 3, further comprising a combiner for coupling the excitation light source to the optical fiber of the fiber optic sensor chamber. 7. The system of claim 6, further comprising an amplifier to amplify the fluorescence signals. 8. The system of claim 1, further comprising a central processor for calculating and reporting fluorescence measurements. 9. The method of claim 1, further comprising a temperature sensor circuit for measuring temperature at the site of insertion. 10. A tissue-insertable probe device containing therein a fluororphor analyte within a sealed sensor chamber, wherein the fluorophor operably responds to pH levels in the surrounding tissue, thereby providing calculated pCO2 levels in the tissue. 11. The probe of claim 10, further comprising one or more aligned optic fibers, each having two opposing ends, and each of which is operably connected and sealed at the distal end to the probe, thereby forming an operably-linked light guide for collecting emitted fluorescence from the fluorophor at less than the numerical aperture of the light guide; wherein at the proximal end, at least one first fiber is externally, operably-connected to the light source to transmit excitation light to the fluorophor within the sealed sensor chamber, and wherein at least one second fiber is externally connected to the detecting device to collect and transmit emitted fluorescence from the fluorophor to the detector device. 12. A method for making real-time, in vivo measurement of tissue pCO2 in the animal or human patient, the method comprising: inserting into the tissue of a patient the probe containing the sealed ... | The Trustees of the University of Pennsylvania | A61B005/1455
US07953469 | B2 | 2011-05-31 | System and method for real-time localization for gated MR imaging | 2008-02-01 | A system and method for MR imaging is disclosed that includes displaying a series of MR images of an ROI in real-time and localizing a slice within the ROI. The method also includes using a parallel imaging technique to acquire gated MR data from the localized slice and reconstructing a prescribed fixed number of gated MR images of the localized slice. | 1. An MRI apparatus to acquire images comprising: a MRI system having a plurality of gradient coils positioned about a bore of a magnet to impress a polarizing magnetic field and an RF transceiver system and an RF switch controlled by a pulse module to transmit RF signals to an RF coil assembly to acquire MR images; anda computer programmed to: acquire RF coil sensitivity data for a region-of-interest (ROI);acquire un-gated MR data from the ROI;reconstruct a series of real-time MR images from the acquired un-gated MR data;identify at least one slice from at least one image of the series of real-time MR images to obtain a slice identified for localization within the ROI;acquire triggered MR data from the at least one slice to partially fill k-space for the slice identified for localization;determine MR data for unacquired k-space from the RF coil sensitivity data and the triggered MR data to continue to fill k-space;perform a CINE interpolation of the triggered MR data; andreconstruct high-resolution images of the slice identified for localization. 2. The MRI apparatus of claim 1 wherein the computer is further programmed to acquire all the triggered MR data for the given slice within a single R-R interval. 3. The MRI apparatus of claim 2 wherein the computer is further programmed to acquire the triggered MR data in an R-R interval immediately following an R-R interval used for steady-state inducement. 4. The MRI apparatus of claim 1 wherein the computer is further programmed to perform an ASSET acquisition to acquire the triggered MR data and determine the MR data for the unacquired k-space. 5. The MRI apparatus of claim 1 wherein the computer is further configured to perform self-encoded parallel imaging to acquire the triggered MR data. 6. The MRI apparatus of claim 1 wherein the computer is further programmed to acquire triggered MR data for all the slices identified for localization within a single breath-hold, wherein the triggered MR data is triggered by a cardiac cycle of an imaging subject. 7. The MRI apparatus of claim 1 wherein the computer is further caused to retrospectively interpolate the triggered MR data and reconstruct a prescribed fixed number of images of the slice identified for localization. 8. The MRI apparatus of claim 1 wherein the computer is further caused to automatically return to the acquisition of the un-gated MR data following the acquisition of triggered MR data. 9. The MRI apparatus of claim 1 wherein the computer is further caused to acquire the triggered MR data with a center frequency, transmit/receive gains, and shim values used to acquire the un-gated MR data. 10. The MRI apparatus of claim 1 wherein the computer is programmed to automatically identify the at least one slice to cover a whole heart of a medical patient. 11. A method of MR imaging moving objects comprising: acquiring un-gated MR data for reconstruction into a series of real-time MR images of an ROI having motion therein;displaying the... | General Electric Company | A61B005/05
US08202227 | B2 | 2012-06-19 | Pressure sensing adapter for gastric band system injector | 2009-12-14 | The body of a pressure sensing component includes a pressure sensor, a first opening, and a first pair of electrical contacts. A cap of the pressure sensing component comprises a second opening, a wiper seal, and a second pair of electrical contacts. The cap is rotatable relative to the body. When the cap is in a first rotational position, the first and second openings are in fluid communication with each other; and the first pair of contacts are not coupled with the second pair of contacts. When the cap is in a second rotational position, the seal isolates the first opening from the second opening; and the first pair of contacts are coupled with the second pair of contacts. The pressure sensing component may be coupled with a syringe and needle to externally adjust and monitor the pressure of fluid in an implanted gastric band system. | 1. The apparatus, comprising: (a) a body having a hollow interior portion, wherein the body comprises: (i) a first opening in communication with the hollow interior portion,(ii) a pressure sensing device in communication with the hollow interior portion, and(iii) a first port in communication with the hollow interior portion, wherein the first port is configured to couple with a selected one of a syringe or a needle;(b) a cap coupled with the body, wherein the cap comprises: (i) a second opening,(ii) a seal, wherein the seal is configured to selectively seal the first opening relative to the second opening when the cap is at a first position relative to the body, wherein the seal is further configured to selectively permit fluid communication between the first opening and the second opening when the cap is at a second position relative to the body, and(iii) a second port in communication with the second opening, wherein the second port is configured to couple with a selected one of a syringe or a needle; and(c) a communication device coupled with the cap, wherein the communication device is configured to selectively couple with the pressure sensing device when the cap is at the first position, wherein the communication device is further configured to transmit pressure data to an external device when the communication device is coupled with the pressure sensing device. 2. The apparatus of claim 1, wherein the communication device comprises a cable. 3. The apparatus of claim 1, wherein the body comprises a first electrical contact in communication with the pressure sensing device, wherein the cap comprises a second electrical contact in communication with the communication device, wherein the first and second electrical contacts are configured such that the first and second electrical contacts are coupled together when the cap is at the first position, wherein the first and second electrical contacts are configured such that the first and second electrical contacts are not coupled together when the cap is at the second position. | Ethicon Endo-Surgery, Inc. | A61B005/00 A61F002/00 A61F002/02
US07937275 | B2 | 2011-05-03 | Identifying clinical trial candidates | 2008-05-08 | In general in one aspect, one or more clinical study candidates are identified by consolidating health care information for a consumer, applying the consumer's health care information to a rules engine, identifying a consumer as eligible to participate in one or more clinical studies, and presenting an opportunity to participate in the study to the consumer. | 1. A computer-implemented method of identifying one or more clinical study candidates, the method comprises: receiving consolidated health care information for a consumer of medical services that has an established real-time communication channel between the consumer of medical services and a provider of medical services;retrieving by one or more computers attributes defining a suitable candidate for a clinical study;causing the one or more computers to compare the attributes defining the suitable candidate for the clinical study to the consolidated health information for the consumer;determining by the one or more computers that the consumer's consolidated health information includes at least one of the attributes defining the suitable candidate for the clinical study;identifying the consumer as eligible to participate in the clinical study;retrieving by the one or more computers one or more qualifications defining a suitable provider of medical services to consult with the consumer of medical services about the clinical study;determining by the one or more computers that the provider of medical services is associated with at least one of the one or more qualifications defining the suitable provider of medical services;identifying by the one or more computers the provider of medical services as qualified to consult with the consumer of medical services about the clinical trial;generating a graphical user interface that when rendered on a display device, renders a visual representation of a field that provides an opportunity for the identified consumer to participate in the clinical study; andduring the established real-time communication channel between the consumer and the provider of medical services: sending by the one or more computers the graphical user interface to at least one of the consumer and the provider of medical services. 2. The method of claim 1 further comprising: adding an event associated with the opportunity for the consumer to participate in the clinical study to an agenda associated with the consumer. 3. The method of claim 1 furthering comprising: receiving an acceptance of the opportunity from the consumer;enrolling the consumer in the study; andadding tasks related to the study to an agenda associated with the consumer. 4. The method of claim 1 further comprising detecting changes in a state of the consumer's health care information. 5. The method of claim 4 wherein the consumer's health care information is applied to a rules engine upon detecting at least one change in the state of the consumer's health care information. 6. The method of claim 4 wherein the rules engine defines one or more rules. 7. The method of claim 6 wherein one or more rules included in the rules engine are based on criterion specified by the study. 8. The method of claim 4 wherein the consumer's health care information is input into the rules engine at a predefined time. 9. The method of claim 1 wherein the consumer is identified as eligib... | American Well Corporation | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07970470 | B2 | 2011-06-28 | Diagnosis and/or therapy using blood chemistry/expired gas parameter analysis | 2004-08-30 | Methods and systems for diagnosing disorders, including, for example, disordered breathing, involve sensing one or more of a blood chemistry parameter and/or an expired gas parameter, such as expired respiratory gas concentration, blood gas concentration, and blood pH. Diagnosis of the disorder may be performed by a medical device, such as a respiratory therapy device or a cardiac therapy device, based on implantably detected blood gas/pH concentration/level or externally detected expired respiratory gas concentration. Cardiac and respiratory therapies for addressing the disorder may be adjusted based on the detected parameters. | 1. A method of providing disordered breathing therapy for a patient whose respiration cycle includes inspiration and expiration intervals, comprising: providing a patient-external respiratory therapy and a patient-internal cardiac therapy to a patient, the respiratory therapy being adapted to treat disordered breathing;measuring an expired respiratory gas concentration of an expired respiratory gas sample acquired at the end of an expiration interval of the patient as determined by a timing element, the measured expired respiratory gas concentration representative of blood gas concentration;determining one or more parameters influenced by disordered breathing, the one or more parameters comprising at least the expired respiratory gas concentration;providing one or more parameter thresholds comprising at least an expired respiratory gas threshold; andadjusting the respiratory therapy and the cardiac therapy based on the one or more parameters and the one or more parameter thresholds. 2. The method of claim 1, the one or more parameters further comprising blood pH measured by a sensor. 3. The method of claim 1, wherein determining the one or more parameters comprises determining the one or more parameters using one or more implantable sensors. 4. The method of claim 1, wherein determining the one or more parameters comprises determining the one or more parameters using one or more patient-external sensors. 5. The method of claim 1, wherein determining the one or more parameters comprises determining the one or more parameters using one or more implantable sensors and one or more patient-external sensors. 6. The method of claim 1, wherein adjusting the respiratory therapy comprises adjusting a positive airway pressure therapy. 7. The method of claim 1, wherein adjusting the cardiac therapy comprises adjusting a cardiac pacing therapy. 8. The method of claim 1, wherein determining the one or more parameters comprises determining one or both of expired oxygen concentration or carbon dioxide gas concentration. 9. The method of claim 1, wherein determining the one or more parameters comprises determining one or more of blood oxygen concentration, blood carbon dioxide concentration, or blood pH. 10. The method of claim 1, further comprising monitoring for a change in the one or more parameters. 11. The method of claim 10, wherein monitoring for the change in the one or more parameters comprises detecting a change in the one or more parameters indicative of hypoxemia. 12. The method of claim 10, wherein monitoring for the change in the one or more parameters comprises detecting a change in the one or more parameters indicative of apnea. 13. The method of claim 10, wherein monitoring for the change in the one or more parameters comprises detecting a change in the one or more parameters indicative of hypercapnea or hypocapnea. 14. The method of claim 1, wherein adjusting the respiratory therapy and the cardiac therapy comprises using the therap... | Cardiac Pacemakers, Inc. | A61N001/00 A61B005/00 A61B005/05
US07967756 | B2 | 2011-06-28 | Respiratory therapy control based on cardiac cycle | 2004-09-15 | Methods and systems involve adjusting respiratory therapy based on cardiac cycle phase. A parameter indicative of cardiac cycle is sensed and the respiratory therapy is adjusted based on cardiac cycle phase. Modulation of respiratory therapy pressure reinforces the pumping action of the heart and results in increased cardiac output with decreased expenditure of myocardial energy output. | 1. A method for delivering airway pressure to a patient, comprising: providing a patient-implantable cardiac rhythm monitoring system adapted to electrically couple to the patient's heart and to communicate with a patient-external respiratory therapy device;determining cardiac cycle phase using the cardiac rhythm monitoring system, including determining if the cardiac cycle phase is systole or diastole based on one or more of (a) implantably sensing a parameter indicative of the cardiac cycle phase, and (b) determining delivery of a cardiac pacing pulse; andcontrolling the airway pressure delivered by the patient-external respiratory therapy device in synchrony with the cardiac cycle phase, wherein controlling the airway pressure is performed at least in part implantably. 2. The method of claim 1, wherein sensing the parameter comprises sensing oxygen saturation. 3. The method of claim 1, wherein sensing the parameter comprises sensing blood pressure. 4. The method of claim 1, wherein sensing the parameter comprises sensing cardiac electrical activity. 5. The method of claim 1, wherein sensing the parameter comprises sensing a parameter associated with cardiac motion. 6. The method of claim 5, wherein sensing the parameter associated with cardiac motion comprises sensing transthoracic impedance. 7. The method of claim 5, wherein sensing the parameter associated with cardiac motion comprises sensing heart sounds. 8. The method of claim 1, wherein implantably controlling the airway pressure comprises matching intrathoracic pressure with cardiac cycle phase. 9. The method of claim 1, wherein implantably controlling the airway pressure comprises: delivering a first airway pressure during a first cardiac cycle phase; anddelivering a second airway pressure during a second cardiac cycle phase. 10. The method of claim 1, wherein implantably controlling the airway pressure comprises: modulating airway pressure above a baseline pressure during cardiac systole; andmodulating airway pressure below a baseline pressure during cardiac diastole. 11. The method of claim 1, wherein: determining the cardiac cycle phase comprises anticipating the cardiac cycle phase; andimplantably controlling the airway pressure comprises controlling the airway pressure based on the anticipated cardiac cycle phase. 12. The method of claim 11, wherein anticipating the cardiac cycle comprises anticipating the cardiac cycle based on recent cardiac cycle history. 13. The method of claim 1, wherein implantably controlling the airway pressure comprises implantably controlling airway pressure to improve at least one of ventricular preload or ventricular afterload. 14. The method of claim 1, wherein implantably controlling the airway pressure comprises implantably controlling the airway pressure to enhance cardiac pumping action. 15. A therapy control system, comprising: a patient-implantable pulse generator configured to deliver cardiac pacing pulses to a patient's heart bas... | Cardiac Pacemakers, Inc. | A61N001/365 A61B005/08
US07945305 | B2 | 2011-05-17 | Adaptive acquisition and reconstruction of dynamic MR images | 2005-09-01 | A method for acquiring MR data from a beating heart during subject respiration includes a prescan phase in which a respiratory compensation table and a k-space sampling schedule are produced. The k-space sampling table is produced using a spatio-temporal model of the beating heart and time sequential sampling theory. During the subsequent scan an imaging pulse sequence which is prospectively compensated for respiratory motion is used to acquire k-space data from the subject. The imaging pulse sequence is repeated to play out the phase encodings in the order listed in the k-space sampling schedule. | 1. A method for acquiring and producing images of a time-varying subject with a magnetic resonance imaging (MRI) system, the steps comprising: a) performing a prescan of the time-varying subject that includes: a)i) estimating a dynamic model for the time-varying subject;a)ii computing a dynamic sampling schedule using the estimated dynamic model, the sampling schedule including information indicative of locations in k-space, and an order and timing in which the locations in k-space are to be sampled; anda)iii selecting a pulse sequence for directing the MRI system to acquire k-space image data from the time-varying subject;b) performing a scan of the time-varying subject to acquire k-space image data with the MRI system that includes: b)i acquiring information indicative of the motion of the subject caused by respiration;b)ii altering during the scan one of the information in the dynamic sampling schedule indicative of the locations in k-space, the order in which the locations in k-space are sampled, and the timing in which the locations in k-space are sampled, or the selected pulse sequence to compensate for respiratory motion of the subject; andb)iii acquiring k-space image data with the MRI system using the sampling schedule and the selected pulse sequence; andc) reconstructing an image using k-space image data acquired in step b). 2. The method as recited in claim 1 wherein the dynamic model in step a)i includes information about the dual-k-t support of the time-varying subject. 3. The method as recited in claim 1 in which during the pre-scan, step a)i includes acquisition of data with the MRI system. 4. The method as recited in claim 3 which further includes altering the data acquisition during the pre-scan to compensate for respiratory motion of the subject. 5. The method as recited in claim 1 in which the dynamic sampling schedule calculated in step a) includes information that enables additional reference data to be acquired with the MRI system during step b). 6. The method as recited in claim 5 in which step b) includes acquiring reference data with the MRI system that includes subject motion information. 7. The method as recited in claim 1 in which step a)i) includes producing a respiratory motion calibration rule that relates motion of the subject to respiratory phase, step b)i) includes estimating the respiratory phase, and step b)iii) employs the respiratory phase and the respiratory motion calibration rule to determine the alteration. 8. The method as recited in claim 7, in which the respiratory motion calibration rule depends on the state of the time-varying subject. 9. The method as recited in claim 1, in which the selected pulse sequence directs the MRI system to sample k-space image data along parallel lines in k-space. 10. The method as recited in claim 1, in which the selected pulse sequence directs the MRI system to sample k-space image data along radial lines in k-space. 11. The method as recited in claim 1 in whic... | The Board Of Trustees Of The University of Illinois | A61B005/05
US07967750 | B2 | 2011-06-28 | Lie detection via electrogastrography | 2007-04-13 | The present invention provides a device and methods of lie detection utilizing electrogastrography to monitor changes in the gastric electric rhythm of a subject. Also provided are methods of discriminating between the stress of deception induced in a subject during a lie detection examination and the stress of the examination itself. | 1. A method for evaluating the truthfulness of a subject's response to questions posed during administration of a polygraph examination thereto, comprising: connecting the subject to a polygraph instrument comprising: channels for measurement of heart rate variability via electrocardiography and gastric electric rhythm via gastroelectrography; anda processing unit operatively connected thereto; said processing unit configured to record and compare the measurements;posing a plurality of control questions to the subject;posing a plurality of relevant questions to the subject;measuring heart rate variability and gastric electric rhythm in the subject after posing the controlling questions to obtain a baseline for each of said heart rate variability and gastric electric rhythm;measuring heart rate variability and gastric electric rhythm in the subject after posing the relevant questions; andcomparing the heart rate variability and gastric electric rhythm after posing said relevant questions with the heart rate variability baseline and the gastric electric rhythm baseline, wherein an occurrence of bradygastria during posing of said relevant questions compared to the corresponding baseline is indicative that the response to the relevant questions is truthful even if heart rate variability increases compared to its corresponding baseline during posing of said relevant questions. 2. The method of claim 1, further comprising: measuring independently an additional one or more involuntary physiological responses in the subject after posing said controlling and said relevant questions. 3. The method of claim 2, wherein said involuntary physiological response is respiration rate, blood pressure or galvanic skin response or a combination thereof. 4. The method of claim 1, wherein said occurrence of bradygastria with said increase in heart rate variability from baseline is indicative of stress in the subject induced by posing said relevant questions. 5. The method of claim 1, wherein an occurrence of tachygastria with said increase in heart rate variability from baseline is indicative of stress in the subject induced by deception in responding to said relevant questions. 6. A method of improving the accuracy of lie detection during a polygraph examination of a subject, comprising: connecting the subject to a polygraph instrument comprising: channels for monitoring heart rate variability via electrocardiography and gastric electric via a gastroelectrography; anda processing unit operatively connected thereto said processing unit configured to record and compare heart rate variability and gastric electric rhythm:monitoring heart rate variability during administration of the polygraph examination; andmonitoring concurrently a gastric electric rhythm of the individual, wherein a bradygastric change in gastric electric rhythm is indicative of a truthful response from the subject even if an increase in heart rate variability sufficient to indicate a false respon... | The Board of Regents of the University of Texas System | A61B005/00
US08781564 | B2 | 2014-07-15 | Method and apparatus in connection with exercise | 2008-06-02 | An apparatus in connection with a fitness exercise of a person, including means for processing multidimensional electrocardiographic data of the person, the multidimensional electrocardiographic data comprising at least two spatially separately measured electrocardiographic signal components, means for forming a vectorcardiographic measure on the basis of the multidimensional electrocardiographic data, and means for applying the vectorcardiographic measure in determination of a fitness exercise related parameter. | 1. An apparatus in connection with a fitness exercise of a person, the apparatus comprising: a processing device to process multidimensional electrocardiographic data of the person, the multidimensional electrocardiographic data comprising at least two spatially separately measured electrocardiographic signal components obtained by using two couplings of electrodes and a total of at least three electrodes, each of the couplings of electrodes including two electrodes;a forming device to form a vectorcardiographic measure on the basis of the multidimensional electrocardiographic data;an applicator to apply the vectorcardiographic measure and a heart rate parameter to determine a fitness-exercise-related parameter; andan instruction module to provide exercise instructions from the fitness-exercise-related parameter, the exercise instructions to be given with a user interface to the person for carrying out the fitness exercise, wherein the instruction module applies a lookup table or a calculation algorithm to determine the exercise instructions from a variety of exercise instructions. 2. The apparatus of claim 1, wherein the apparatus further comprises a transformer to transform at least a portion of the multidimensional electrocardiographic data into an orthogonal representation of the data having three mutually different directions that are not all in the same plane. 3. The apparatus of claim 1, wherein the apparatus comprises: a monitoring device to monitor a heart rate of the person; anda determiner to determine the fitness-exercise-related parameter on the basis of the heart rate of the person and the vectorcardiographic measure. 4. The apparatus of claim 1, wherein the apparatus further comprises: an identifier to identify a multidimensional R wave and a multidimensional T wave from the multidimensional electrocardiographic data; anda forming device to measure, as the vectorcardiographic measure, a relationship between the multidimensional R wave and the multidimensional T wave. 5. The apparatus of claim 1, wherein the apparatus comprises: a forming device to form a vector representation of a multidimensional R wave in the electrocardiographic data;a forming device to form a vector representation of a multidimensional T wave in the electrocardiographic data; andan estimator to estimate an angle measure of the angle between the vector representation of the multidimensional R wave and the vector representation of the multidimensional T wave. 6. The apparatus of claim 1, wherein the apparatus further comprises: a monitoring device to monitor if the vectorcardiographic measure exceeds a predetermined threshold value; andan indicator to indicate if the vectorcardiographic measure exceeds a predetermined threshold value. 7. The apparatus of claim 1, wherein the applicator applies the vectorcardiographic measure to determine the fitness-exercise-related parameter during the fitness exercise and/or after the fitness exercise. 8. The apparatus o... | Polar Electro Oy | A61B005/0402
US08948853 | B2 | 2015-02-03 | Cardiac mapping with catheter shape information | 2011-03-10 | A non-contact cardiac mapping method is disclosed that includes: (i) inserting a catheter into a heart cavity having an endocardium surface, the catheter including multiple, spatially distributed electrodes; (ii) measuring signals at the catheter electrodes in response to electrical activity in the heart cavity with the catheter spaced from the endocardium surface; and (iii) determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface. Related systems and computer programs are also disclosed. | 1. A method comprising: processing information relating to characteristics of catheter shape;inserting a catheter into a heart cavity having an endocardium surface, the catheter comprising multiple, spatially distributed electrodes;moving the catheter to each of multiple, different positions in the heart cavity;for each of the different catheter positions, measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;subsequent to processing the information relating to the characteristics of the catheter shape, determining physiological information at multiple locations of the endocardium surface based at least in part on the measured signals and the processed catheter shape information;using the determined physiological information to guide ablation of one or more selected regions of the heart;repeating the measurement of catheter electrode signals and the determination of the physiological information after the ablation treatment, anddisplaying a difference map comprising information about how the determined physiological information changed in response to the ablation treatment. 2. The method of claim 1, wherein the information relating to characteristics of the catheter shape is processed prior to inserting the catheter into the heart cavity. 3. The method of claim 1, further comprising storing the processed information relating to characteristics of the catheter shape. 4. The method of claim 1, wherein determining the physiological information at multiple locations of the endocardium surface based at least in part on the measured signals and the processed catheter shape information comprises accessing the stored processed catheter shape information. 5. The method of claim 1, wherein the catheter is configured to assume multiple different shapes. 6. The method of claim 1, further comprising, subsequent to inserting the catheter into the heart cavity, deploying the multiple, spatially distributed electrodes into a specified electrode arrangement. 7. The method of claim 1, wherein processing information relating to the characteristics of the catheter shape comprises processing information relating to the relative arrangement of the multiple, spatially distributed electrodes. 8. The method of claim 1, wherein processing information relating to the characteristics of the catheter shape comprises determining locations of the electrodes based on a location of the catheter. 9. The method of claim 8, wherein determining physiological information at multiple locations of the endocardium surface based at least in part on the measured signals and the processed catheter shape information comprises determining physiological information based at least in part on the determined locations of the electrodes. 10. The method of claim 1, wherein processing the information relating to characteristics of the catheter shape comprises processing information relating to the distribution of the electrodes on the catheter.... | Rhythmia Medical, Inc. | A61B005/02 A61B005/042
US08740750 | B2 | 2014-06-03 | Device and method for controlling and/or regulating a training and/or rehabilitation unit | 2008-07-24 | In at least one embodiment, a device for controlling and/or regulating a training and/or rehabilitation unit comprises a training and/or rehabilitation unit, a sensor unit having sensors for determining oxygen concentration and for determining carbon dioxide concentration, a control unit for the sensors, a microcontroller in the control unit for controlling the heating power of heating elements of the sensors, depending on breath flow volume, for maintaining constant sensor temperatures, and a resistance and/or braking arrangement of the training and/or rehabilitation unit that can be controlled and/or regulated based on a breathing gas composition determined by the sensor unit. The device can be used to increase endurance, preferably by way of altitude training. | 1. A device for controlling and/or regulating a training and/or rehabilitation unit, wherein the device comprises: a training and/or rehabilitation unit;a sensor unit with a heatable electrochemical solid-state electrolyte sensor, which in order to determine the oxygen concentration, contains yttrium-contaminated zirconium oxide as an electrolyte, two electrodes, a support element, and a heating element, and in order to determine the carbon dioxide concentration, contains a further heatable electrochemical solid-state electrolyte sensor, which comprises a sodium ion conductor as an electrolyte, two electrodes, and a support element and a heating element;a control unit for the sensors;a microcontroller in the control unit configured to control a heating power of the heating elements of the sensors and maintain constant sensor temperatures, wherein the heating power is controlled dependent on a flow volume of breathing of a person; anda controllable and/or adjustable resistance and/or brake arrangement of the training and/or rehabilitation unit that depends on a breathing gas composition determined by the sensor unit. 2. The device according to claim 1, wherein the sensor unit is arranged directly in a component through which inspired and expired air of the person flows. 3. The device according to claim 1, wherein the sensor unit is arranged in a breathing mask. 4. The device according to claim 1, further comprising an integrated ear clip for measuring oxygen concentration in the person's blood. 5. The device according to claim 1, further comprising an integrated ear clip for measuring the person's pulse. 6. The device according to claim 1, further comprising a data storage for storing performance data with further medical parameters of the person. 7. The device according to claim 1, wherein data measured by the sensor unit is recorded with the aid of a connected personal digital assistant. 8. The device according to claim 1, wherein the training and/or rehabilitation unit is an ergometer, fitness machine, cross trainer, rowing ergometer machine, rowing machine, treadmill, walking machine, spin bike, or bicycle. 9. The device according to claim 1, wherein the resistance and/or brake arrangement of the training and/or rehabilitation unit contains a pneumatic, hydraulic, mechanical or electro-magnetic brake, an eddy-current brake, or band brake. 10. The device according to claim 1, further comprising a two- and/or three-dimensional visual display, at least one acoustic output and/or recording device, and an output unit configured to generate wind, temperatures, and/or odors. 11. The device according to claim 1, further comprising an output unit configured to stimulate the sense of touch. 12. The device according to claim 1, wherein components of the training and/or rehabilitation unit, including the controllable and/or adjustable resistance and/or brake arrangement, the sensor unit, and the control unit for the sensors are connected with e... | - | A63B021/00
US08583225 | B2 | 2013-11-12 | System and method for detecting skin penetration | 2007-08-24 | A system and method for detecting skin penetration. The system comprises an invasive component for penetrating the skin; a dummy electrode for making contact with the surface of the skin; at least one penetrating electrode disposed in the invasive component; and a Wheatstone bridge circuit; wherein a resistance across the dummy electrode and the penetrating electrode constitutes one of the resistive legs of the Wheatstone bridge circuit and skin penetration of the invasive component is detected based on a differential output voltage from the Wheatstone bridge circuit. | 1. A system for detecting skin penetration, the system comprising: an invasive component for penetrating the skin;a dummy electrode for making contact with the surface of the skin;at least one penetrating electrode disposed in the invasive component;a Wheatstone bridge circuit;wherein a resistance across the dummy electrode and the penetrating electrode constitutes one of the resistive legs of the Wheatstone bridge circuit, anda first pair of reference electrodes for making contact with the surface of the skin, wherein a skin resistance across the first pair of reference electrodes constitutes the mirroring resistive leg, with respect to ground, of the Wheatstone bridge circuit;wherein the system is configured to detect a depth of skin penetration of the invasive component based on a differential output voltage from the Wheatstone bridge circuit. 2. The system as claimed in claim 1, wherein one of the reference electrodes is the dummy electrode. 3. The system as claimed in claim 1, further comprising second and third pairs of reference electrodes, each for making contact with the surface of the skin, wherein respective skin resistances across the second and third pairs of reference electrodes constitute the remaining resistive legs of the Wheatstone bridge circuit respectively. 4. The system as claimed in claim 1, wherein a plurality of penetrating electrodes are disposed in the invasive component. 5. The system as claimed in claim 4, wherein resistances across the dummy electrode and the respective penetrating electrodes are multiplexed across one of the resistive legs of the Wheatstone bridge circuit connected to the respective penetrating electrodes, and the depth of skin penetration of the invasive component is detected based on differential output voltages from the Wheatstone bridge circuit induced by the respective resistances across the dummy electrode and the respective penetrating electrodes. 6. The system as claimed in claim 1, wherein the invasive component comprises a microneedle disposed in a lancing device, and the dummy electrode for making contact with the surface of the skin is disposed on a skin-contact face of the lancing device. 7. The system as claimed in claim 1, wherein the invasive component comprises a microneedle disposed in a lancing device, and the dummy electrode for making contact with the surface of the skin and the first pair of reference electrodes are disposed on a skin-contact face of the lancing device. 8. The system as claimed in claim 3, wherein the invasive component comprises a microneedle disposed in a lancing device, and the dummy electrode for making contact with the surface of the skin and the first, second, and third pairs of reference electrodes are disposed on a skin-contact face of the lancing device. 9. The system as claimed in claim 1, wherein the invasive component comprises a hollow or solid microneedle. 10. The system as claimed in claim 1, wherein the invasive component comprises a co... | Agency for Science, Technology and Research | A61B005/053
US08483814 | B2 | 2013-07-09 | SCG point pick process for use with multiaxis accelerometers | 2010-04-29 | A seismocardiograph using multiple accelerometer sensors to identify cardiac valve opening and closing times. A methodology for selecting event times is also disclosed. | 1. A device for monitoring and displaying cardiac function of a patient comprising: a.) at least one accelerometer sensor having multi-axis output, placed on the patient's body to collect vector accelerometer data along two axis selected from the set of xyphoid to apex of the heart axis, and xyphoid to carotid artery axis; said sensor generating signals in response to compression waves originating the heart along at least the two axes;b.) a signal processing system comprising multiple seisomocardiographic SCG channels coupled to said sensor for generating a SCG waveform along each axis, forming a carotid SCG channel waveform along one axis, and a xyphoid channel waveform along one axis, and a apex SCG channel waveform along one axis, and for recording and storing a segment of said SCG waveforms for analysis;c.) a module for extracting the time of mechanical events in the heart from said SCG waveform segment and for presenting the time of each such events to a system user on a display;d.) wherein said set of mechanical events is selected from the group comprising; mitral valve closing event MC, mitral valve opening event MO, aortic valve closing event AC, atrial valve opening AO;said device further comprising;e.) a set of electrocardiographic ECG electrodes placed on said patient's body for generating signals in response to electrical depolarization of originating the heart;f.) signal processing system coupled to said ECG electrodes for generating an ECG waveform, and for recording and storing a segment of said ECG waveform for analysis;g.) said atrial valve opening event AO module for carrying out a selection process in one SCG channel includes; forming an average carotid acceleration wavelet from said carotid SCG channel waveform;identify the R wave peak in the ECG channel;identify the largest positive peak in said SCG channel waveform within 220 msec of the R peak and mark this point as AO carotid, if no positive peak is found, start search again at R wave peak and search ahead for the largest negative peak, when found, search back for the first negative knee (concave down), and mark this point as AO carotid. 2. The device of claim 1 further comprising; an AO module for carrying out a selection process in one SCG channel includes; forming an average xyphoid acceleration wavelet from said xyphoid SCG channel waveform;starting at AO carotid time, search back 60 msec for largest positive peak, and mark this point as AO xyphoid. 3. The device of claim 1 further comprising; an AO module for carrying out a selection process in one SCG channel includes; forming an average apex acceleration wavelet from said apex SCG channel waveform;starting at AO carotid time, search back 60 msec for largest negative peak, and mark this point as AO apex;and if¦AO xyphoid−AO apex¦&lt;=10 msec; then select AO apex as aortic valve opening AO;and if (AO xyphoid−AO apex)>=0 and¦AO xyphoid−AO apex¦>10 msec; select AO xyphoid as pulmonic valve opening PO, and select AO ap... | - | A61B005/04
US08509881 | B2 | 2013-08-13 | True ECG measurement during cardio pulmonary resuscitation by adaptive piecewise stitching algorithm | 2009-11-03 | A method and apparatus utilizing a piecewise stitching adaptive algorithm (PSAA) to filter signal artifacts, such as those induced by cardiopulmonary resuscitation (CPR) from sensed signals in real-time. PSAA is a method of estimating artifact component present in a first signal that is highly correlated with a second signal. The PSAA may utilize autocorrelation and cross-correlation calculations to determine signal sample windows in the first and second signals. The PSAA may estimate a signal artifact in a primary signal segment based on the determined correlations between the primary signal and an artifact signal. The PSAA may remove the estimated signal artifact from the primary signal. In the absence of an artifact signal, PSAA is able to estimate artifacts in the first signal utilizing filters. The PSAA may be implemented in Automated External Defibrillators, Monitor Defibrillators or other devices capable sensing highly correlated signals such as, for example, ECG and CPR signals. | 1. An apparatus for filtering artifacts from ECG signals in real-time comprising: means for sensing an ECG signal, the ECG signal representing a physical impulse of cardiac tissue;means for sensing an artifact signal, the artifact signal representing a physiological function;means for removing the artifact signal from the ECG signal using a piecewise stitching adaptive algorithm that provides self-corrective, self-deconvolved, time-varying, self-adaptive signal estimation including analysis of overlapped adjacent segments that form a plurality of sample windows of the ECG signal and are stitched together with the artifact signal removed in response to the piecewise stitching adaptive algorithm. 2. An apparatus for filtering artifacts from ECG signals in real-time comprising: means for sensing an ECG signal, the ECG signal representing a physical impulse of cardiac tissue;means for sensing an artifact signal, the artifact signal representing a physiological function;means for removing the artifact signal from the ECG signal using a piecewise stitching adaptive algorithm that provides time-varying, self-adaptive signal estimation;wherein the piecewise stitching adaptive algorithm comprises:means for selecting signal sample windows from the ECG signal and the artifact signal;means for generating a primary ECG signal segment from the ECG signal and a primary artifact signal segment from the artifact signal based on the selected sample windows;means for determining a relationship between the primary ECG signal segment and primary artifact signal segment;means for estimating a ECG signal artifact in the primary ECG signal segment based on the determined relationship;means for removing the estimated ECG signal artifact from the primary signal segment of the ECG signal. 3. An apparatus having a processor adapted to filter CPR compression artifacts from ECG signals in real-time comprising: an ECG sensor for sensing a ECG signal representing a physical impulse of cardiac tissue;an artifact sensor for sensing an artifact signal representing a physiological function;a piecewise stitching adaptive algorithm processor coupled to the ECG sensor and the artifact sensor programmed to calculate a ECG signal artifact created by the artifact signal and to remove the ECG signal artifact from the ECG signal utilizing self-corrective, self-deconvolved, time-varying, self-adaptive signal estimation including analysis of overlapped adjacent segments that form a plurality of sample windows of the ECG signal and are stitched together with the artifact signal removed in response to the piecewise stitching adaptive algorithm. 4. An apparatus having a processor adapted to filter CPR compression artifacts from ECG signals in real-time comprising: an ECG sensor for sensing a ECG signal representing a physical impulse of cardiac tissue;an artifact sensor for sensing an artifact signal representing a physiological function;a piecewise stitching adaptive algorithm processor cou... | Cardiac Science Corporation | A61B005/04
US08636794 | B2 | 2014-01-28 | Grafts and stent grafts having a radiopaque marker | 2006-11-09 | A graft device comprising a layer of synthetic non-metallic material having a first surface and a second surface spaced apart from the first surface. The device further, includes a radiopaque marker at least partially embedded in the layer. | 1. A graft device comprising: an extruded layer of synthetic non-metallic material having a first surface and a second surface spaced apart from the first surface; anda co-extruded first radiopaque marker and a coextruded second radiopaque marker each at least partially embedded in the layer, the first radiopaque marker being formed of a first radiopaque material and the second radiopaque marker being formed of a second radiopaque material different from the first radiopaque material. 2. The graft device according to claim 1, wherein the first radiopaque marker defines a portion of the first surface. 3. The graft device according to claim 1, wherein the first radiopaque marker defines a portion of the second surface. 4. The graft device according to claim 1, wherein the first radiopaque marker is disposed between the first and second surfaces. 5. The graft device according to claim 1, wherein the synthetic non-metallic material comprises a material selected from a group consisting essentially of Dacron, polyester, PTFE, ePTFE, polyurethane, polyurethaneurea, siloxane, and combinations thereof. 6. The graft device according to claim 1, wherein the first radiopaque marker is formed from a paste having about 20% to about 60% tantalum powder. 7. The graft device according to claim 1, wherein the first radiopaque marker is formed from a paste having about 20% to about 40% Barium Sulfate. 8. The graft device according to claim 7, wherein the Barium Sulfate is mixed with a biocompatible coloring agent. 9. The graft device according to claim 8 wherein the coloring agent is green, blue or black colored. 10. The graft device according to claim 9 wherein the coloring agent is cobalt blue. 11. The graft device according to claim 1, wherein the graft device has a length extending axially from end to end, and the first radiopaque marker comprises a first continuous line that extends along a majority of the length of the graft device. 12. The graft device according to claim 11, wherein the second radiopaque marker comprises a second continuous line parallel to the first continuous line, the second continuous line extending along the majority of the length of the graft device. 13. The graft device according to claim 12, wherein one of the first continuous line is formed from a paste including Barium Sulfate and the second continuous line is formed from a paste including tantalum powder. 14. The graft device according to claim 12, wherein the layer is tubular in shape, and wherein the first continuous line and the second continuous line are formed within a same circumferential third of the layer. 15. A device comprising: a stent frame;an extruded synthetic non-metallic material that surrounds a portion of the stent frame, the synthetic non-metallic material having first and second surfaces; anda first co-extruded radiopaque strip and a second coextruded radiopaque strip embedded in the nonmetallic material, the first radiopaque strip and the second ... | C. R. Bard, Inc. | A61F002/06
US20110112428 | A1 | 2011-05-12 | Body composition measuring apparatus using a bioelectric impedance analysis associated with a neural network algorithm | 2010-01-21 | A body composition measuring apparatus using a bioelectric impedance analysis and a neural network algorithm for obtaining two or more anthropometry variables from testees and then inputting the anthropometry variables into the internal processing unit that has a built-in back propagation-artificial neural network that has one input layer, 1-10 hidden layers each having 1-15 hidden neurons and one output layer having one output neuron. By means of the aforesaid artificial neural network, the invention accurately predict the fat free mass of the testee so as to further obtain the amount of body fat, showing higher accuracy than conventional linear regression equation (LRE). | 1. A body composition measuring apparatus using a bioelectric impedance analysis associated with a neural network algorithm, comprising: an apparatus body, said apparatus body comprising at least one anthropometry variables acquiring means for obtaining anthropometry variables of a testee, including at least two of the age, body height, body weight and bioelectrical impedance values of the testee; and a processing unit mounted inside said apparatus body and connected to said at least one anthropometry variables acquiring means, said processing unit having built therein at least one back propagation-artificial neural network (BP-ANN), each said back propagation-artificial neural network comprising: an input layer, said input layer comprising a plurality of input neurons adapted for receiving said anthropometry variables from said at least one anthropometry variables acquiring means; 1-10 hidden layers, each said hidden layer comprising 1-15 hidden neurons and a transfer function corresponding to each said hidden neuron, each said transfer function being a Log-Sigmoid or Hyperbolic Tangent Sigmoid; and an output layer, said output layer comprising an output neuron and a linear transfer function for outputting fat free mass (FFM), said fat free mass (FFM) being processable by said processing unit to obtain the body fat of the testee. 2. The body composition measuring apparatus using a bioelectric impedance analysis associated with a neural network algorithm as claimed in claim 1, wherein the number of said hidden layers is 1-5 layers. 3. The body composition measuring apparatus using a bioelectric impedance analysis associated with a neural network algorithm as claimed in claim 2, wherein the number of the hidden neuron of each said hidden layer is 1-10. 4. The body composition measuring apparatus using a bioelectric impedance analysis associated with a neural network algorithm as claimed in claim 1, wherein the number of the hidden layers is 2-3 layers. 5. The body composition measuring apparatus using a bioelectric impedance analysis associated with a neural network algorithm as claimed in claim 4, wherein the number of the hidden neuron of each said hidden layer is 1-10. 6. The body composition measuring apparatus using a bioelectric impedance analysis associated with a neural network algorithm as claimed in claim 1, wherein said at least one back propagation-artificial neural network (BP-ANN) uses the age, body height, body weight and bioelectrical impedance values of a number of persons as the network input for training, and the fat free mass of said persons as a network output for training. 7. The body composition measuring apparatus using a bioelectric impedance analysis associated with a neural network algorithm as claimed in claim 1, wherein said at least one back propagation-artificial neural network (BP-ANN) uses Levenberg-Marquardt algorithm to run training. 8. The body composition measuring apparatus using a bioelectric impedance... | CHARDER ELECTRONIC CO., LTD. | A61B005/053
US20110137195 | A1 | 2011-06-09 | Method and Apparatus for Monitoring Respiratory Activity | 2007-02-28 | A method of monitoring respiratory activity includes obtaining a respiration sound signal. The respiration sound signal is processed to identify a frequency component corresponding to respiration, and an amplitude related parameter of the component is measured. If the measured parameter is in an indeterminate region then a temporal domain related parameter of the identified frequency component is measured. An apparatus for monitoring respiratory activity is also disclosed. | 1. A method of monitoring respiratory activity comprising obtaining a respiration sound signal, processing the respiration sound signal to identify a frequency component corresponding to respiration, measuring an amplitude related parameter of the identified frequency components and, if the parameter is in an indeterminate region, measuring a temporal domain related parameter of the identified frequency component. 2. A method as claimed in claim 1 in which, if the amplitude related parameter falls below a lower limit of the indeterminate region then a cessation of respiration state is identified and if the amplitude related parameter exceeds an upper limit of the indeterminate region, then a respiration continuation state is identified. 3. A method as claimed in claim 1 in which the temporal domain related parameter comprises at least one of duration and integral of the identified frequency component. 4. A method as claimed in claim 1 in which, if the temporal domain related parameter is measured to be within a respiration range then a respiration continuation state is identified and if the parameter is outside the respiration range then a cessation of respiration state is identified. 5. A method as claimed in claim 3 in which, if cessation of breathing is detected then an apnoea alarm is triggered. 6. A method as claimed in claim 1 in which the amplitude related parameter comprises signal power. 7. A method as claimed in claim 1 in which the respiration sound signal is processed to identify a frequency component comprising a frequency pass band. 8. A method as claimed in claim 7 in which the frequency pass band is in the region of approximately 500 to 900 hertz. 9. A method as claimed in claim 1 in which the temporal domain related parameter is measured for a signal segment. 10. A method as claimed in claim 9 in which the signal segment is defined between adjacent maxima or adjacent minima. 11. A method as claimed in claim 10 in which the adjacent maxima or adjacent minima are detected by a peak detection algorithm. 12. A method as claimed in claim 1 in which ambient noise is detected and filtered out of the respiration sound signal. 13. A method as claimed in claim 12 in which the ambient noise is filtered out by signal subtraction. 14. An apparatus for monitoring respiratory activity comprising a respiration sound monitor and a processor arranged to identify a frequency component corresponding to respiration from the respiration sound signal, measure an amplitude related parameter of the identified frequency component and, if the parameter is in an indeterminate region, measure a temporal-domain related parameter of the identified frequency component. 15. An apparatus as claimed in claim 14 in which the monitor comprises at least one microphone. 16. An apparatus as claimed in claim 14 in which the processor includes a band pass filter for identifying the frequency components corresponding to respiration. 17. An apparatus as cla... | IMPERIAL INNOVATIONS LIMITED | A61B005/08
US20110152650 | A1 | 2011-06-23 | ADAPTIVE PUMP CONTROL DURING NON-INVASIVE BLOOD PRESSURE MEASUREMENT | 2009-12-21 | A method of operating a non-invasive blood pressure (NIBP) monitor having a blood pressure cuff. During operation of the NIBP monitor, the blood pressure cuff is initially inflated at a rapid inflation rate. Once the blood pressure cuff reaches a first pressure, the inflation rate of the blood pressure cuff is reduced from the rapid inflation rate to a measurement inflation rate. The blood pressure cuff continues to inflate at the measurement inflation rate while the NIBP monitor receives signals from the patient. Based upon the signals received from the patient, the controller of the NIBP monitor calculates an initial inflation pressure. The blood pressure cuff is inflated to the calculated initial inflation pressure and inflation is terminated. In this manner, signals received from the patient during inflation are used to calculate the initial inflation pressure to reduce the amount of time required to make a blood pressure measurement. | 1. A method of monitoring blood pressure in a patient, the method comprising the steps of: positioning a blood pressure cuff on the patient, the blood pressure cuff being selectively inflatable to restrict blood flow past the blood pressure cuff;inflating the blood pressure cuff at a rapid inflation rate to a first pressure;continuing to inflate the blood pressure cuff above the first pressure at a measurement inflation rate, wherein the measurement inflation rate is less than the rapid inflation rate;determining an initial inflation pressure for the patient during the inflation of the blood pressure cuff at the measurement inflation rate;terminating the inflation of the blood pressure cuff at the initial inflation pressure;decreasing the pressure in the blood pressure cuff from the initial inflation pressure while monitoring oscillation pulses from the blood pressure cuff; andcalculating the blood pressure for the patient based on the monitored oscillation pulses. 2. The method of claim 1 wherein the measurement inflation rate is sufficient to allow a determination of an estimated systolic pressure before the cuff pressure reaches the initial inflation pressure. 3. The method of claim 1 wherein the first pressure is below a predicted systolic pressure for the patient. 4. The method of claim 1 further comprising the steps of: positioning an inflation valve between a source of pressurized air and the blood pressure cuff; andoperating the inflation valve to selectively inflate the blood pressure cuff at both the rapid inflation rate and the measurement inflation rate. 5. The method of claim 1 further comprising the steps of: positioning a sensor of a pulse monitor on the patient;monitoring for the presence of pulse signals from the pulse monitor during inflation of the blood pressure cuff at the measurement inflation rate; anddetermining the initial inflation pressure based upon the pulse signals from the pulse monitor. 6. The method of claim 5 wherein the pulse monitor is an SpO2 monitor and the sensor is positioned distal to the blood pressure cuff. 7. The method of claim 1 further comprising the steps of: monitoring for the presence of oscillation pulses from the blood pressure cuff during inflation of the blood pressure cuff at the measurement inflation rate; anddetermining the initial inflation pressure based on the oscillation pulses detected during the inflation of the blood pressure cuff at the measurement inflation rate. 8. The method of claim 7 wherein the first pressure is below a predicted systolic pressure such that the blood pressure cuff is inflated at the measurement inflation rate from the first pressure to the initial inflation pressure. 9. The method of claim 7 further comprising the steps of: determining an estimated systolic pressure based on the oscillometric pulses detected during the inflation of the blood pressure cuff at the measurement inflation rate; anddetermining the initial inflation pressure based upon the es... | GENERAL ELECTRIC COMPANY | A61B005/1455 A61B005/021
US20110102177 | A1 | 2011-05-05 | Diver Monitoring and Communication System | 2008-08-15 | The present invention relates to a system, method and software for monitoring a diver comprising: a topside monitoring system 13 on a vessel comprising a receiver for receiving data from at least one diver 12, the data being transmitted using an acoustic modem 207 and specifying information relating to the diver 12, wherein the information relating to the diver comprises the depth of diver, a computer for processing the data 101, and a user interface 111 for conveying the information and the elapsed time of dive to a dive supervisor so that the dive supervisor can manage the dive. | 1. A method of monitoring a diver comprising the steps of: receiving data from at least one diver, the data being transmitted using an acoustic modem and specifying at least one parameter relating to the diver,conveying the parameter and elapsed time of dive to a dive supervisor so that the dive supervisor can manage the dive, the parameter and elapsed time of dive forming diver information,wherein the parameter relating to the diver comprises the depth of the diver. 2. A method according to claim 1 further comprising the step of determining a decompression schedule for the diver and conveying the decompression schedule. 3. A method according to claim 2 wherein the step of determining a decompression schedule comprises receiving input specifying one or more decompression schedules or dive tables, and determining a decompression schedule based on the received input. 4. A method according to claim 2 further comprising the steps of: determining one or more alternative decompression schedules based on the conveyed diver information, the step being triggered by received or calculated diver information or by input from a dive supervisor, andconveying the alternative decompression schedules. 5. (canceled) 6. (canceled) 7. (canceled) 8. A method according to claim 4 wherein the step of determining one or more alternative decompression schedules comprises selecting one or more dive tables from a set of dive tables, each dive table specifying one or more decompression schedules, wherein the selection is based on some or all of the diver information. 9. A method according to claim 8 wherein selecting one or more dive tables based on some or all of the diver information comprises the step of determining which of the set of dive tables specify decompression schedules that can still be adhered to by the diver based on diver information. 10. A method according to claim 1 further comprising one or more alarms for alerting a dive supervisor of actions that may be required, wherein the alarms are triggered by the relationship of: diver information to the decompression schedule, and/ordiver information to specified values and/or states. 11. (canceled) 12. (canceled) 13. A method according to claim 1 further comprising the step of recording dive information related to the dive, wherein dive information comprises one or more of any dive related parameter, data, information or input received or generated prior to or throughout the dive. 14. A system for monitoring a diver comprising: a topside monitoring system on a vessel comprising:a receiver for receiving data from at least one diver, the data being transmitted using an acoustic modem and specifying information relating to the diver, wherein the information relating to the diver comprises the depth of diver,a computer for processing the data, anda user interface for conveying the information and the elapsed time of dive to a dive supervisor so that the dive supervisor can manage the dive. 15. A system... | Prink Limited | G08B021/00 G01N033/00 A61B005/08 A61B005/00
US07959569 | B2 | 2011-06-14 | System and methods for processing analyte sensor data | 2010-03-23 | Systems and methods for processing sensor analyte data, including initiating calibration, updating calibration, evaluating clinical acceptability of reference and sensor analyte data, and evaluating the quality of sensor calibration. During initial calibration, the analyte sensor data is evaluated over a period of time to determine stability of the sensor. The sensor may be calibrated using a calibration set of one or more matched sensor and reference analyte data pairs. The calibration may be updated after evaluating the calibration set for best calibration based on inclusion criteria with newly received reference analyte data. Fail-safe mechanisms are provided based on clinical acceptability of reference and analyte data and quality of sensor calibration. Algorithms provide for optimized prospective and retrospective analysis of estimated blood analyte data from an analyte sensor. | 1. A method for processing sensor data from a continuous glucose sensor, the method comprising: receiving sensor data from a continuous glucose sensor, wherein the sensor data comprises one or more sensor data points;receiving reference data from a reference glucose monitor that is independent from the continuous glucose sensor, wherein the reference data comprises one or more reference data points; anddetermining, using an electronic system, a stability of the continuous glucose sensor at least in part by evaluating a correlation between sensor and reference data points. 2. The method of claim 1, wherein determining a stability of the continuous glucose sensor further comprises waiting a predetermined time period of from about ten to about thirty minutes after implantation of the continuous glucose sensor in a host before determining a stability associated with the continuous glucose sensor. 3. The method of claim 1, wherein determining a stability of the continuous sensor further comprises evaluating at least one of oxygen and amplitude. 4. The method of claim 1, wherein receiving reference data from a reference glucose monitor comprises receiving a wired internal communication. 5. The method of claim 1, wherein determining a stability of the continuous glucose sensor comprises evaluating sensor data to check at least for an indication of a change in at least one of an amplitude or a variability of a sensitivity associated with the continuous glucose sensor. 6. The method of claim 1, further comprising predicting one or more glucose values at one or more future points in time. 7. The method of claim 1, further comprising forming one or more matched data pairs, wherein the one or more matched data pairs are formed by matching at least one reference data point to at least one substantially time corresponding sensor data point. 8. The method of claim 1, further comprising providing output to a user interface or to an external device responsive to a level of stability of the continuous glucose sensor. 9. The method of claim 1, wherein providing output comprises providing at least one of an audible output, visual output, a tactile output, or combinations thereof. 10. The method of claim 9, wherein providing output further comprises indicating at least one of a numeric estimated glucose value, a directional trend of glucose concentration, a graphical representation of an estimated glucose value, or combinations thereof. 11. The method of claim 8, wherein providing output comprises alerting a user of a present and/or upcoming hypoglycemic and/or hyperglycemic event. 12. The method of claim 8, wherein providing output comprises sending the sensor data to an insulin pump. 13. The method of claim 8, wherein the output is provided only after a predetermined level of stability has been determined. 14. The method of claim 8, wherein the output comprises sensor data. 15. The method of claim 8, wherein the output comprises an indication of insta... | DexCom, Inc. | A61B005/00
US20110144455 | A1 | 2011-06-16 | SYSTEMS AND METHODS FOR MONITORING A SUBJECT AT REST | 2011-02-25 | Disclosed herein are methods and devices for monitoring a subject at rest. One such device comprises a sensing unit having a fluid-filled bladder configured to be placed under a substrate on which the subject lays and a sensor in fluid communication with the bladder. The sensor is configured to sense pressure variations within the bladder generated by a heart beat, respiration and body weight of the subject and to generate signals indicative of the pressure variations. A processor is configured to receive the signals and to determine and generate output indicative of the subject's heart beat and respiration and presence on the substrate. An external device is configured to display one or more of the output. | 1. A device for monitoring a subject at rest comprising: a sensing unit having a fluid-filled bladder configured to be placed under a substrate on which the subject lays and a sensor in fluid communication with the bladder, wherein the sensor is configured to sense pressure variations within the bladder generated by one or more of a heart beat, respiration and body weight of the subject and to generate signals indicative of the pressure variations;a processor configured to receive the signals and to determine and generate output indicative of one or more of the subject's heart beat and respiration and presence on the substrate; andan external device configured to display one or more of the output. 2. The device of claim 1, wherein the subject is a patient in a medical facility, the substrate is a mattress and the external device is operable to be monitored by personnel of the medical facility, the external device configured to display output for a plurality of patients each associated with a sensing unit. 3. The device of claim 2, wherein the external device comprises an alarm for alerting the personnel of output that is out of a predetermined range. 4. The device of claim 2, wherein the output generated by the processor further includes one or more of length of time on mattress, frequency of movement of patient on mattress and degree of movement by patient on the mattress, wherein frequency of movement is determined from number of changes in static pressure and degree of movement is determined from degree of change of static pressure and change of strength of one or more of heart beat and respiration. 5. The device of claim 4, wherein the external device is configured to generate an alarm based on the output, the alarm configured to generate an action by personnel to prevent pressure ulcers. 6. The device of claim 1, wherein the subject is a participant in a prison release program and the external device is operable to be monitored by personnel of the program, the external device configured to display output for a plurality of participants each associated with a sensing unit, the output including heart rate, respiration, and presence or absence in bed for each of the participants, wherein the output can be further analyzed to determine drug or alcohol abuse and compliance with curfews. 7. The device of claim 2 further comprising a database configured to store output over a period of time, wherein the stored output can be further analyzed to determine staffing requirements based on an amount of patient activity. 8. The device of claim 1 further comprising monitoring system software and a user interface, wherein the monitoring system software is configured to display the output in an arrangement based on user input through the user interface. 9. The device of claim 2, wherein the external device comprises a live view screen including a real-time status for each of the plurality of patients. 10. The device of claim 4, wherein the processor... | BAM LABS, INC. | A61B005/00
US20110137199 | A1 | 2011-06-09 | VISCERAL FAT MEASURING DEVICE | 2011-02-10 | A visceral fat measuring device capable of simply and safely measuring a visceral fat amount is provided. In the visceral fat measuring device for calculating the visceral fat amount based on trunk measurement information, impedance information of the entire trunk, and impedance information of a surface layer of the trunk obtained by measuring a potential difference in the body axis direction of the trunk on the dorsal side of the trunk, a belt has a hollow pressed member pressed onto the dorsal side of the trunk, the pressed member having a pressed surface provided with electrodes, a wiring member including a circuit substrate connected to the electrodes for measuring the potential difference is accommodated inside the pressed member, the pressed member has flexibility in the vertical direction to the body axis direction so as to be curved along a surface shape of the dorsal side of the trunk, and when the pressed member is curved, an opposite surface to the pressed surface is extended relatively to the pressed surface and deflected. | 1-9. (canceled) 10. A visceral fat measuring device for calculating a visceral fat amount based on trunk measurement information serving as a basis for calculating a trunk sectional area in a section on an abdominal part of a trunk vertical to a body axis of the trunk, impedance information of the entire trunk obtained by applying an electric current from hands and legs to the trunk and measuring a potential difference in part of a surface of the trunk, and impedance information of a surface layer of the trunk obtained by winding a belt having a plurality of electrodes around the trunk so as to apply the electric current through the vicinity of the surface layer of the trunk and measure a potential difference in part of the surface of the trunk, wherein the belt has a hollow pressed member pressed onto the trunk, the pressed member having a pressed surface provided with a plurality of the electrodes, a wiring member including a circuit substrate connected to a plurality of the electrodes for measuring the potential difference is accommodated inside the pressed member, due to an opposite surface to the pressed surface provided with a plurality of the electrodes having a stretching portion having stretchability in the belt longitudinal direction and flexibility in the vertical direction to the body axis direction, and a non-stretching portion having neither stretchability nor flexibility, the pressed member has flexibility in the vertical direction to the body axis direction so as to be curved along a surface shape of the trunk, when the pressed member is curved, the opposite surface to the pressed surface provided with a plurality of the electrodes is extended relatively to the pressed surface and deflected so as not to narrow inner space where the wiring member is accommodated, a wiring portion having pliability of the wiring member is arranged in the inner space where the stretching portion is positioned, and a wiring portion having no pliability of the wiring member is arranged in the inner space where the non-stretching portion is positioned. 11. The visceral fat measuring device according to claim 10, wherein the stretching portion is formed by a concave-convexo surface portion in which a concave part and a convex part extending in the short-side direction are provided alternately and continuously in the longitudinal direction, the non-stretching portion is formed by a flat surface portion, and due to extension of the opposite surface to the pressed surface by extension of the concave-convexo surface portion when the pressed member is curved, the pressed member is curved without narrowing the inner space where the wiring member is accommodated. | OMRON HEALTHCARE CO., LTD. | A61B005/053
US08068904 | B2 | 2011-11-29 | Devices and methods for monitoring physiological information relating to sleep with an implantable device | 2010-12-17 | Described here are implantable devices and methods for monitoring physiological information relating to sleep. The implantable devices are generally designed to include at least one sensor for sensing physiological information, a processor for processing the physiological information using low computational power to detect a sleep stage, and a battery. The detected sleep stage information may then be used to indicate sleep quality, identify or monitor a medical condition, or guide treatment thereof. | 1. A method for monitoring physiological information relating to sleep comprising: sensing physiological information;detecting a plurality of half waves in the sensed physiological information based on a local maximum amplitude and a subsequent local minimum amplitude, or based on a local minimum amplitude and a subsequent local maximum amplitude;determining a duration and an amplitude of each half wave of the plurality of detected half waves;determining, for each half wave, whether the duration is within a range of a plurality of different duration ranges, and whether the amplitude of said half wave is within a range of a plurality of different amplitude ranges, each different duration range having a preset minimum duration threshold and a preset maximum duration threshold, each different amplitude range having a preset minimum amplitude threshold and a preset maximum amplitude threshold;for each duration range of the plurality of different duration ranges, counting a number of occurrences of half waves within said each duration range for each amplitude range of the plurality of different amplitude ranges; anddetermining a detection of an occurrence of at least one sleep state or stage if the number of occurrences of half waves within specified amplitude ranges and specified duration ranges correspond to the at least one sleep state or stage. 2. The method of claim 1 further comprising storing at least a portion of the physiological information or the half wave physiological information. 3. The method of claim 1 further comprising transmitting at least a portion of the physiological information or the half wave physiological information to a remote location. 4. The method of claim 1 further comprising: determining whether detected occurrences of the at least one sleep state or stage indicate a physiological condition; andtreating the physiological condition. 5. The method of claim 4, wherein the physiological condition is selected from the group consisting of a neurological condition, a psychological condition, a cardiac condition, a respiratory condition, and a sleep condition. 6. The method of claim 4, wherein treating the physiological condition comprises providing stimulation to a patient. 7. The method of claim 6, wherein providing stimulation comprises selecting stimulation based on which at least one sleep state or stage is detected. 8. The method of claim 7, wherein selecting stimulation comprises selecting frequency, pulse-width, amplitude, or duration of stimulation, or a stimulation montage. 9. The method of claim 7, wherein selecting stimulation comprises initiating or terminating stimulation based on which at least one sleep state or stage is detected. 10. The method of claim 4, wherein treating the physiological condition comprises delivering a drug. 11. The method of claim 10, wherein delivering the drug delivery comprises initiating or terminating drug delivery based on which at least one sleep state or stage is detecte... | NeuroPace, Inc. | A61B005/048
US07949385 | B2 | 2011-05-24 | System and method for reconstruction of the human ear canal from optical coherence tomography scans | 2006-04-17 | A method for reconstructing an ear canal from optical coherence tomography (OCT) scan data of an ear comprises extracting frame numbers and line numbers of interference intensities corresponding to one or more markers on an OCT scan guide, receiving reference frame numbers and lines numbers for one or more markers, determining a starting position and direction for the OCT ear scan from the ear scan marker frame and line numbers and the reference marker frame and line numbers, for each scan line, finding a pixel number of a maximum interference intensity value, and determining an offset distance of said pixel from said scan guide, and reconstructing a surface of the ear canal from the distance offset data. | 1. A method for reconstructing an ear canal, comprising the steps of: providing a set of optical coherence tomography (OCT) scan data of an ear comprising frames and scan lines of pixels that form a 3D data set whose values are interference intensities;extracting frame numbers and line numbers of interference intensities corresponding to one or more markers on an OCT scan guide;receiving reference frame numbers and lines numbers for said one or more markers;determining a starting position and direction for said OCT ear scan from the ear scan marker frame and line numbers and said reference marker frame and line numbers;for each scan line, finding a pixel number of a maximum interference intensity value, and determining an offset distance of said pixel from said scan guide; andreconstructing said surface of said ear canal from said distance offset data. 2. The method of claim 1, wherein an offset distance d of a pixel j from the guide is computed from d(j)=d0+j/ds,wherein d0 is an offset from the center of the guide and ds is a sampling rate. 3. The method of claim 1, wherein said ear canal surface can be reconstructed from {right arrow over (p)}(i,j)={right arrow over (p)}0+{right arrow over (c)}(t)+d(j)(sin(α){right arrow over (u)}+cos(α){right arrow over (v)}),wherein {right arrow over (p)}(i, j) represents a scan line i perpendicular to guide, j a pixel along the scan line, {right arrow over (p)}0 is a starting position in space of the scan, {right arrow over (c)}(t) a curved path of the guide, t is a position on the guide corresponding to the scan line i, and α is a scan line angle, and {right arrow over (u)} and {right arrow over (v)} are two unit vectors perpendicular to each other defining a 2D coordinate frame in a plane perpendicular to the curve {right arrow over (c)}(t). 4. The method of claim 3, whereint=i·vf,wherein v is a scan speed of a scanning probe contained within said guide, and f is a scan line sampling frequency, andα=i·ωf,wherein ω is the angular speed of rotation the probe. 5. The method of claim 3, wherein said unit vectors {right arrow over (u)} and {right arrow over (v)} are defined byu->=t->×w->t->×w->,⁢v->=t->×u->,wherein {right arrow over (w)} is a vector that is not perpendicular to the curve {right arrow over (c)}(t) at any time, and {right arrow over (t)} is a tangent to the curve defined ast->=ⅆⅆt⁢c->⁡(t). 6. The method of claim 1, wherein said OCT scan data is obtained by providing an OCT scanning apparatus including, a guide containing a rotatable probe, sliding and rotating said probe in said guide, emitting near infrared light from said probe at predetermined intervals, measuring interference of reflected light with a reference signal, and saving said interference data in a computer readable storage medium. 7. The method of claim 1, comprising, for each scan line, selecting those pixels with an interference intensity value above a pre-determined threshold and determining an offset distance of said sel... | Siemens Medical Solutions USA, Inc. | A61B005/05
US08348862 | B2 | 2013-01-08 | Muscular energy state analysis system and method for swing motion and computer program product thereof | 2009-12-09 | A muscular energy state analysis system and method for a swing motion and a computer program product thereof are provided. The system includes: a swing implement, for a user to perform the swing motion, and including an acceleration sensor for sensing acceleration of the swing implement when the swing implement is swung, so as to generate a swing speed data; a plurality of signal detection modules, for detecting electromyographic (EMG) signals generated by a muscles of the user; a database, for storing a muscular energy sample value,; a muscular energy analysis module, for analyzing the EMG signals and the swing speed data so as to obtain a plurality of muscular performance values; and a comparison module, for comparing the swing speed data and the muscular performance values with the at least one muscular energy sample value in the database, so as to generate a comparison result data. | 1. A muscular energy state analysis system for a swing motion, comprising: a swing implement, for a user to perform the swing motion, and comprising an acceleration sensor for sensing acceleration of the swing implement when the swing implement is swung, so as to generate a swing speed data;a plurality of signal detection modules, for detecting a plurality of electromyographic (EMG) signals generated by a plurality of muscles of the user;a database, for storing a muscular energy sample value, wherein the muscular energy sample value comprises a swing speed sample value and a plurality of corresponding muscular performance sample values;a muscular energy analysis module, for analyzing the EMG signals and the swing speed data so as to obtain a plurality of muscular performance values, wherein the muscular energy analysis module comprises a signal analysis module and a muscular energy determination module, the signal analysis module performs a time-domain analysis on the EMG signals so as to obtain a plurality of muscle exertion intensity values, and the muscular energy determination module calculates the muscular performance values by using the muscle exertion intensity values, wherein the signal analysis module further performs a frequency-domain analysis on the EMG signals so as to obtain a fatigue index value and forms a two-dimensional coordinate system by using one of the muscle exertion intensity values and the fatigue index value; anda comparison module, for comparing the swing speed data and the plurality of muscular performance values with the muscular energy sample value in the database, so as to generate a comparison result data, wherein the comparison module analyzes a target quadrant where a point defined by the muscle exertion intensity value and the fatigue index value is, the comparison module determines whether the user is in the a fatigue state according to the quadrant where the point is, and when the comparison module determines that the user is in the fatigue state, the comparison module generates a rest suggestion information. 2. The muscular energy state analysis system for a swing motion according to claim 1, wherein the comparison module compares the swing speed data and the plurality of muscular performance values with the swing speed sample value and the plurality of corresponding muscular performance sample values of each of the muscular energy sample values in the database respectively, so as to generate the comparison result data. 3. The muscular energy state analysis system for a swing motion according to claim 1, wherein the muscular performance values and the muscular performance sample values are respectively corresponding to one of a plurality of muscular properties; and the muscular energy analysis module compares the muscular performance value and the muscular performance sample value of the same muscular property, so as to generate the comparison result data. 4. The muscular energy state analysis system for ... | Institute for Information Industry | A61B005/103 A61B005/117 A61B005/04 A61B005/22 A61B001/24 A63B057/00 A63B069/00 A63B069/38 A63B053/00 A63B021/00 A63B021/23 A63B023/16 G01L003/24 G01L005/00
US20110160609 | A1 | 2011-06-30 | Method and system for monitoring pressure in a body cavity | 2010-04-30 | An implantable pressure sensor system having a sensor assembly configured and adapted to measure pressure in a volume, the sensor assembly including at least a first MEMS pressure sensor, an application-specific integrated circuit (ASIC) having memory means, temperature compensation system, drift compensation system, a sensor catheter, and power supply means for powering the sensor assembly, the first MEMS pressure sensor having a pressure sensing element that is responsive to exposed pressure, the pressure sensing element being adapted to generate a pressure sensor signal representative of the exposed pressure, the temperature compensation system being adapted to correct for temperature induced variations in the pressure sensor signal, the drift compensation system being adapted to correct for pressure and temperature induced pressure sensor signal drift. | 1. An implantable pressure sensor system, comprising: a sensor assembly configured and adapted to measure pressure in a volume, said sensor assembly including at least a first MEMS pressure sensor, an application-specific integrated circuit (ASIC) having memory means, a temperature compensation system, a drift compensation system, a sensor catheter, and power supply means for powering said sensor assembly,said first MEMS pressure sensor having a first pressure sensing element that is responsive to first exposed pressure, said first pressure sensing element being adapted to generate a first pressure sensor signal representative of said first exposed pressure,said temperature compensation system being adapted to correct for at least one temperature induced variation in said first pressure sensor signal,said drift compensation system being adapted to correct for pressure induced first pressure sensing element signal drift. 2. A catheter for a pressure sensor, comprising: a housing having a pressure transmitting fluid disposed therein, said housing having first region on a first end, a flexible tip region on a second end and a membrane region adapted to couple pressure transmitted to said pressure transmitting fluid to the pressure sensor, said housing first end having means for sealingly engaging the pressure sensor. 3. The catheter of claim 2, wherein said flexible tip region has a first rigidity and said housing first region has a second rigidity, said first rigidity being less than at least 50% of said second rigidity. 4. The catheter of claim 2, wherein said pressure transmitting fluid comprises a biocompatible gel. 5. The catheter of claim 2, wherein said pressure transmitting fluid comprises a biocompatible liquid. 6. The catheter of claim 5, wherein said pressure transmitting fluid comprises saline. 7. The catheter of claim 2, wherein said housing comprises silicone. 8. The catheter of claim 2, wherein said catheter includes a biocompatible sleeve, said sleeve being disposed proximate said catheter housing outer surface. 9. The catheter of claim 8, wherein said catheter sleeve comprises a bioabsorbable polymeric material. 10. The catheter of claim 9, wherein said catheter sleeve comprises polyglycolic acid (PGA). | - | A61B005/00
US07945311 | B2 | 2011-05-17 | Retroreflective marker-tracking systems | 2007-02-09 | A marker-tracking system includes an object, a marker illuminating device, a marker sensing device, and a computing device. The object includes a first retroreflective marker having a shape that is substantially defined by two spherical caps of different radii that are disposed substantially concentric in relation to one another. The marker illuminating device substantially illuminates the first retroreflective marker, the marker sensing device detects the illuminated first retroreflective marker and generates first data indicative of the location of the illuminated first retroreflective marker in space, and the computing device processes the first data generated by the marker sensing device to determine a position and/or orientation of the object in space. | 1. A system comprising: an object to be tracked including a first retroreflective marker having a substantially spherical shape comprising a first spherical cap having a first diameter and a second spherical cap having a second diameter greater than the first diameter, the first spherical cap and the second spherical cap being disposed substantially concentric in relation to one another, a split line defining a plane that intersects the caps, wherein the second spherical cap terminates at the split line,a portion of a surface of an outer edge of at least one of the first and second spherical caps having a reflective material, thereby providing retroreflection from the marker;a marker illuminating device configured to illuminate the first retroreflective marker;a marker sensing device configured to detect the illuminated first retroreflective marker and for generating first data indicative of the location of the illuminated first retroreflective marker in space; anda computing device configured to process the first data generated by the marker sensing device to determine a position of the object in space. 2. The system of claim 1, wherein the shape of the first retroreflective marker is further defined by one or more flanges. 3. The system of claim 1, wherein the first retroreflective marker has a uniform refractive index. 4. The system of claim 1, wherein the first spherical cap comprises a sphere and the second spherical cap comprises a spherical cavity, the first spherical cap and the second spherical cap separated by a medium. 5. The system of claim 4, wherein: the sphere is defined by a radius R1; andthe spherical cavity is defined by an inner radius R2, an outer radius R3, a height h, and a base radius a. 6. The system of claim 4, wherein the outer radius R3 of the spherical cavity is equal to the height h of the spherical cavity. 7. The system of claim 4, wherein the outer radius R3 of the spherical cavity is greater than the height h of the spherical cavity. 8. The system of claim 4, wherein the outer radius R3 of the spherical cavity is less than the height h of the spherical cavity. 9. The system of claim 4, wherein the inner radius R2 of the spherical cavity is equal to the radius R1 of the sphere. 10. The system of claim 4, wherein the inner radius R2 of the spherical cavity is greater than the radius R1 of the sphere. 11. The system of claim 10, wherein the reflective material is applied directly on a medium-side surface of the spherical cavity. 12. The system of claim 10, wherein the reflective material is applied directly on a non-medium-side surface of the spherical cavity. 13. The system of claim 4, wherein the reflective material is applied directly on a surface of the sphere. 14. The system of claim 4, wherein the reflective material is applied directly on a non-medium-side surface of the spherical cavity. 15. The system of claim 1, wherein: the first spherical cap is further defined by a radius R4, a height h1, an... | Northern Digital Inc. | A61B005/05
US07938776 | B2 | 2011-05-10 | Examination device for blood detection | 2008-03-18 | The present invention describes examination devices, such as examination gloves, for detecting the presence of blood, and in particular hidden blood, in biological samples. Examination devices according to the present invention include, in certain embodiments, developer and indicator reagents in a single encapsulation, which offers numerous clinical and other benefits. Kits including devices according to the present invention in sterile packaging, and methods of making and using devices according to the present invention, are also described. | 1. An examination device for detecting the presence of blood in a sample, comprising: an examination glove having at least first, second, and third finger receiving members, each of said first, second, and third finger receiving members having a first open end and a second closed end and being spaced apart from and continuous with each other of said first, second, and third finger receiving members,said first finger receiving member including a sample collection area disposed upon said first finger receiving member and generally located proximal said second closed end of said first finger receiving member for receiving a sample to be tested for the presence of blood,said second finger receiving member including at least one reagent region including a developer reagent capable of reacting with a blood component disposed upon said second finger receiving member and generally located proximal to said second closed end of said second finger receiving member, andsaid third finger receiving member including at least one reagent region including an indicator reagent disposed upon said third finger receiving member and generally located proximal to said second closed end of said third finger receiving member,wherein, when a sample containing blood is received at said sample collection area of said first finger receiving member and caused to contact said developer reagent disposed upon said second finger receiving member and said indicator reagent disposed upon said third finger receiving member, a visible indication that said sample contains blood results at said second closed end of said third finger receiving member, and when a sample not containing blood is received at said sample collection area of said first finger receiving member and caused to contact said developer reagent disposed upon said second finger receiving member and said indicator reagent disposed upon said third finger receiving member, a visible indication that said sample contains blood result does not result at said second closed end of said third finger receiving member. 2. An examination device for detecting the presence of blood in a sample according to claim 1, wherein said examination device is an examination glove. 3. An examination device for detecting the presence of blood in a sample according to claim 1, wherein said first and second finger receiving members are continuous with and located adjacent or non-adjacent one to the other. 4. An examination device for detecting the presence of blood in a sample according to claim 1, further comprising a sample collection element generally located at said sample collection area, wherein said sample collection element is integral with said first finger receiving member or is distinct from said first finger receiving member. 5. An examination device for detecting the presence of blood in a sample according to claim 4, wherein, when said sample collection element is distinct from said first finger receiving member, said sample col... | Mederistic Solutions, Inc. | A61B005/00
US07959857 | B2 | 2011-06-14 | Radiation sterilization of medical devices | 2007-06-01 | Disclosed herein is a method of radiation sterilizing a plurality of medical devices. | 1. A method of radiation sterilizing a plurality of stent-catheter assemblies, the method comprising: Positioning a plurality of stent-catheter assemblies on a fixture, each of the stent catheter assemblies being arranged in a planar configuration and disposed in corresponding planar packaged supported on the fixture, wherein the packages are stacked horizontally on the fixture, and wherein the stents of the assemblies are positioned at a position in each package such that the stents are exposed to the same or substantially the same radiation; andExposing the packages to an incoming radiation beam, the radiation beam being at an acute angle to the planes of the planar configuration of the assemblies, wherein the packages are arranged such that a front end of the stack faces the radiation beam and a back end of the stack faces away from the radiation beam, wherein a radiation shield is positioned at the front end of the stack to reduce radiation exposure to selected assemblies adjacent the front end from the incoming radiation. 2. The method according to claim 1, wherein the radiation beam comprises an e-beam radiation beam. 3. The method according to claim 1, wherein the planar packages in the stack are staggered along an edge of the stack facing the radiation beam. 4. The method according to claim 1, wherein radiation shield is positioned and the properties of the radiation shield are selected such that variation in radiation exposure to the selected assemblies from incoming radiation and the other assemblies in the stack is reduced, substantially eliminated, or eliminated. 5. The method according to claim 1, wherein the stents of the assemblies are positioned at a position in each package such that the stents are exposed to the same or substantially the same radiation. 6. The method according to claim 1, wherein each of the stent catheter assemblies are arranged in a coiled configuration. 7. The method according to claim 1, wherein the incoming radiation is 35-45 KGy and the stent in each of the assemblies is positioned at 36°-108° in the coiled configuration. 8. The method according to claim 1, wherein the acute angle is 35°-45°. 9. The method according to claim 1, wherein exposing the packages to a radiation beam comprises movement of the fixture with respect to the incoming radiation beam. 10. The method according to claim 9, wherein the assemblies are sufficiently sterilized after one movement of the fixture with respect to the incoming radiation beam. 11. The method according to claim 1, wherein the assemblies are cooled prior to, during or after positioning the assemblies on the fixture. 12. A method of radiation sterilizing a plurality of stent-catheter assemblies, the method comprising: Positioning a plurality of stent-catheter assemblies on a fixture, each of the stent catheter assemblies being arranged in a planar configuration and disposed in corresponding planar packages supported on the fixture, wherein the packaged are ... | Abbott Cardiovascular Systems Inc. | A61L002/00 B01J019/08 G01T001/185 G01N021/00 H01J027/00 A61B006/00 A61B005/00 A61M029/00
US07963924 | B2 | 2011-06-21 | Heart simulator | 2006-10-25 | A development system for a cardiac pacemaker control system, the development system comprising: a cardiac pacemaker prototype including an adaptive control system; and a heart simulator in communication with the cardiac pacemaker prototype comprising: a means for receiving pacing stimulations output from the cardiac pacemaker prototype; and a hemodynamic sensor model responsive to the received pacing stimulations, the hemodynamic sensor model being operable to output a signal simulating a hemodynamic sensor in a plurality of heart states, the hemodynamic sensor model being further operative to output the signal simulating the hemodynamic sensor in the plurality of heart states in a plurality of heart conditions. | 1. A heart simulator exhibiting an output simulated hemodynamic sensor whose temporal pattern reflects a heart characteristic which is optimal when pacing stimulations are optimal, the heart simulator comprising: an input device arranged to provide at least one input parameter;a bidirectional intra-cardiac electrogram path;an electrical heart model in communication with said bidirectional intra-cardiac electrogram path, said electrical heart model arranged to output the simulated intra-cardiac electrogram via the bidirectional intra-cardiac electrogram path responsive to said provided at least one input parameter and the pacing stimulations received via said bidirectional intra-cardiac electrogram path; anda hemodynamic sensor model in communication with said electrical heart model and comprising an atrioventricular (AV) timer, an optimal timer and a QT timer, said hemodynamic sensor model arranged to output a simulated hemodynamic sensor signal, said hemodynamic sensor model arranged to load said AV timer with an optimal AV delay responsive to an atrial event of said electrical heart model, to load said optimal timer with the optimal AV delay plus a QT interval and to load said QT timer with the QT interval, each of the optimal AV delay and the QT interval responsive to the provided at least one input parameter, said hemodynamic sensor model exhibiting a plurality of distinct simulated heart states, and wherein said output simulated hemodynamic sensor signal exhibits a particular predetermined change over time in each of the distinct simulated heart states, and wherein transition between each of the distinct simulated heart states is responsive to one of said AV timer, said optimal timer and said QT timer in relation to events output by said electrical heart model. 2. A heart simulator according to claim 1, wherein said plurality of simulated heart states are selected from the group consisting of passive filling, active filling, ejection, regurgitation and late passive filling. 3. A heart simulator according to claim 1, wherein said at least one input parameter is selected from the group consisting of a heart rate value, a heart contractility value, a short AV delay and a long AV delay. 4. A heart simulator according to claim 1, wherein said hemodynamic sensor model simulates one of an impedance sensor, a pressure sensor, a cardiac wall accelerometer, a QT interval sensor and a non-invasive hemodynamic sensor. 5. A heart simulator according to claim 1, wherein said at least one input parameter comprises a cardiac cycle time. 6. A heart simulator according to claim 1, wherein said hemodynamic sensor model is further arranged to add a selectable amount of random noise to said output simulated hemodynamic sensor signal. 7. A heart simulator according to claim 1, wherein said plurality of distinct simulated heart states comprises an active filling state, an ejection state, a regurgitation state, a late passive state, and a passive filling stat... | Sorin CRM SAS | A61B005/02
US07957782 | B2 | 2011-06-07 | Binding strap used in connection with SpO2 sensor | 2006-09-08 | A binding strap used in connection with a SpO2 sensor comprises two clamping portions each of which is provided with an access opening for passing the light emitting portion or light receiving portion of the SpO2 sensor and a holding recess for holding corresponding light emitting portion or light receiving portion, wherein said access openings are designed to allow corresponding light emitting portion or light receiving portion to pass only when said access openings are undergoing a certain amount of deformation. The binding strap of the present invention is advantageous in that the light emitting portion and the light receiving portion of the SpO2 sensor might be tightly secured to the binding strap, and the lost of the binding strap when not use is ensured to be hard. | 1. A binding strap used in connection with a SpO2 sensor, which comprises: a main body consisting of a first clamping portion, a second clamping portion and a connection portion connecting said first clamping portion and said second clamping portion, wherein the first clamping portion defines an inner side that faces toward the second clamping portion and an outer side that faces away from the second clamping portion, and wherein the second clamping portion defines an inner side that faces toward the first clamping portion and an outer side that faces away from the second clamping portion;a wrapping body integrally extended from said first clamping portion of said main body;a first access opening extending through the outer side and the inner side of the first clamping portion and configured to allow a light emitting portion of said SpO2 sensor to selectively pass therethrough;a first holding recess concavely formed in the inner side of said first clamping portion, wherein the first holding recess is configured to hold said light emitting portion after the light emitting portion has passed through said first access opening;a second access opening extending through the outer side and the inner side of the second clamping portion and configured to allow a light receiving portion of said SpO2 sensor to selectively pass therethrough; anda second holding recess concavely formed in the inner side of said second clamping portion, wherein the second holding recess is configured to hold said light receiving portion after the light receiving portion has passed through said second access opening;wherein each of said first and second access openings is configured to allow said light emitting portion or said light receiving portion, respectively, to pass therethrough only when said access opening is undergoing deformation. 2. The binding strap according to claim 1, wherein said binding strap is made of elastically deformable material. 3. The binding strap according to claim 2, wherein said elastically deformable material is silica gel. 4. The binding strap according to claim 1, wherein each of said light emitting portion and said light receiving portion is a plate member in an elliptical shape, said first access opening having an elliptical profile with a major axis slightly shorter than a minor axis of said light emitting portion, said second access opening having an elliptical profile with a major axis slightly shorter than a minor axis of said light receiving portion. 5. The binding strap according to claim 1, wherein each of said light emitting portion and said light receiving portion is a plate member in an elliptical shape, said first access opening having a circular profile with a diameter slightly shorter than a minor axis of said light emitting portion, said second access opening having a circular profile with a diameter slightly shorter than a minor axis of said light receiving portion. 6. The binding strap according to claim 1, wherein each of... | Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | A61B005/1455
US08366640 | B2 | 2013-02-05 | Devices and methods for cervix measurement | 2010-11-11 | A device for measuring a length of a cervix includes an elongate measurement member, a hollow member, a flange, a handle, and a locking mechanism. The elongate measurement member extends along a longitudinal axis and includes a measurement scale thereon. The hollow member is coaxial with and disposed over the elongate measurement member. The flange is offset from the longitudinal axis and attached to a distal end of the hollow member. The handle is attached to a proximal end of the measurement member. The locking mechanism is configured, when locked, to fix the hollow member relative to the measurement member and, when unlocked to allow the hollow member to slide along the measurement member and rotate about the longitudinal axis so as to place the flange in a desired position without moving the measurement scale. | 1. A device for measuring a length of a cervix, comprising: an elongate measurement member extending along a longitudinal axis and including a measurement scale thereon; a hollow member coaxial with and disposed over the elongate measurement member; a flange offset from the longitudinal axis and attached to a distal end of the hollow member; a handle attached to a proximal end of the measurement member; and a locking mechanism on the handle configured, when locked, to fix the hollow member relative to the measurement member and, when unlocked, to allow the hollow member to slide along the measurement member and rotate about the longitudinal axis so as to place the flange in a desired position without moving the measurement scale, wherein a proximal end of the hollow member is slideable into the handle. 2. The device of claim 1, wherein the flange has an opening through which the measurement member can advance distally. 3. The device of claim 1, wherein the flange has a flat surface perpendicular to the longitudinal axis. 4. The device of claim 1, wherein the locking mechanism includes a button, the button including a through-hole configured such that the hollow member can slide therethrough and a lock channel configured such that the hollow member cannot slide therethrough. 5. The device of claim 4, wherein the button further comprises at least one lock ramp between the through-hole and the lock channel. 6. The device of claim 1, wherein the measurement scale is a millimeter scale. 7. The device of claim 1, wherein the measurement scale extends from 0 mm to 50 mm. 8. The device of claim 1, wherein the hollow member is transparent. 9. The device of claim 1, wherein the measurement scale includes an opaque background. 10. The device of claim 1, further comprising an indicator line on the hollow member. 11. The device of claim 10, wherein the indicator line is a color other than black. 12. A method for measuring a length of a cervix, comprising: holding a handle of a device, the device further comprising an elongate measurement member having a measurement scale thereon, a hollow member coaxial with and disposed over the elongate measurement member, and a flange attached to a distal end of the hollow member;rotating the hollow member about the elongate measurement member so as to place the flange at a desired orientation without rotating the measurement scale;advancing the device distally within a vagina until the flange contacts a cervix at an external uterine opening;advancing the measurement member distally within the vagina until a distal end of the measurement member contacts a cervical uterine junction at a formix vaginae, wherein advancing the measurement member distally comprises sliding the hollow member into the handle;locking the measurement member relative to the hollow member by locking a locking mechanism on the handle; andobserving a position of the hollow member with respect to the measurement member to determine a length o... | CerviLenz Inc. | A61B005/103 A61B005/107
US08731650 | B2 | 2014-05-20 | Single Trial detection in encephalography | 2010-10-05 | An EEG cap (8) having 64 or 128 electrodes (10) is placed on the head of the subject (11) who is viewing CRT monitor (14). The signals on each channel are amplified by amplifier (17) and sent to an analog-to-digital converter (20). PC (23) captures and records the amplified signals and the signals are processed by signal processing PC (26) performing linear signal processing. The resulting signal is sent back to a feedback/display PC (29) having monitor (14). | 1. A method for interpreting brain activity monitored by an array of brain activity sensors applied to the head of a subject, wherein each of the sensors are adapted to generate a response signal upon prompting of the subject by an external stimulus, comprising: receiving two or more response signals from the array of brain activity sensors, each response signal corresponding to a different brain activity sensors;determining two or more spatial weighting coefficients, each corresponding to one of the two or more response signals;weighting each of the two or more received response signals in accordance with a corresponding one of the determined weighting coefficients; andgenerating an output from the weighted two or more response signals, the output having a correlation to the brain activity of the subject in response to the external stimulus, wherein the determining of two or more weighting coefficients comprises determining two or more weighting coefficients such that the weighted responses are maximally discriminating between first and second predetermined times corresponding to two stimuli where a maximum difference in brain activity is expected. 2. The method of claim 1, wherein the determining two or more spatial weighting coefficients comprises applying a linear classifier to each of the two or more response signals. 3. The method of claim 1, wherein the generating the output signal comprises linearly integrating the two or more weighted response signals. 4. The method of claim 3, wherein the linear integration comprises summation. 5. The method of claim 1, wherein the correlation comprises a relatively high correlation. 6. The method of claim 1, wherein the external stimulus comprises a visual stimulus. 7. The method of claim 1 further comprising: identifying an expected scalp locality for discriminating components of the output signal relative to the external stimulus;determining a sensor projection for each of the plurality of brain activity sensors;using the determined sensor projections to identify the scalp locality for the actual discriminating components for the external stimulus; andevaluating the validity of the output signal by determining whether the identified locality for the actual discriminating components is substantially the same as the expected locality. 8. A system for interpreting brain activity monitored by an array of brain activity sensors applied to the head of a subject, wherein each of the sensors are adapted to generate a response signal upon prompting of the subject by an external stimulus, comprising: an input adapted to receive two or more response signals generated by at least a portion of the brain activity sensors, each response signal corresponding to a different brain activity sensors;a data processor, coupled to the input, programmed to: receive the two or more response signals from the input;determine two or more spatial weighting coefficients, each corresponding to one of the two or more respons... | The Trustees of Columbia University in the City of New York | A61B005/04
US08369926 | B2 | 2013-02-05 | Assessing the condition of a joint and devising treatment | 2011-01-31 | Methods are disclosed for assessing the condition of a cartilage in a joint, particularly in a human knee. The methods include converting an image such as an MRI to a three dimensional map of the cartilage. The cartilage map can be correlated to a movement pattern of the joint to assess the effect of movement on cartilage wear. Changes in the thickness of cartilage over time can be determined so that therapies can be provided. Information on thickness of cartilage and curvature of cartilage or subchondral bone can be used to plan therapy. Information on movement pattern can be used to plan therapy. Planning therapy may include generating one or more physical models. | 1. A method of creating a physical model used in the treatment of an articular surface of a diseased or damaged joint of a patient, the method comprising: electronically receiving and processing image data of at least a portion of the joint of the patient;creating a three dimensional representation of at least a portion of a subchondral bone surface of the diseased or damaged joint from the image data;transmitting the three dimensional representation to a site for analysis; andgenerating a physical model that includes a surface portion that is derived from the three dimensional representation;wherein the physical model is configured for use in the treatment of the joint. 2. The method of claim 1, wherein said image data includes both normal and diseased cartilage tissue. 3. The method of claim 1, wherein said articular surface includes bone. 4. The method of claim 1, wherein said articular surface includes cartilage. 5. The method of claim 1, further comprising estimating a surface of a region of cartilage based at least in part on the representation of the subchondral bone surface. 6. The method of claim 1, further comprising estimating the dimensions of a region of cartilage based at least in part on the representation of the subchondral bone surface. 7. The method of claim 1, further comprising estimating a thickness of a region of cartilage based at least in part on the representation of the subchondral bone surface. 8. The method of claim 1, further comprising estimating a curvature of a region of cartilage based at least in part on the representation of the subchondral bone surface. 9. The method of claim 1, further comprising estimating a volume of a region of cartilage based at least in part on the representation of the subchondral bone surface. 10. The method of claim 1, further comprising estimating a depth of a region of cartilage based at least in part on the representation of the subchondral bone surface. 11. The method of claim 1, further comprising estimating an area of a region of cartilage based at least in part on the representation of the subchondral bone surface. 12. The method of claim 1, further comprising estimating a width of a region of cartilage based at least in part on the representation of the subchondral bone surface. 13. The method of claim 1, wherein at least a portion of said physical model has a thickness that is similar to normal cartilage. 14. The method of claim 1, wherein the surface portion substantially matches the three dimensional representation of at least a portion of the subchondral bone surface of the joint. 15. The method of claim 1, wherein the portion of the subchondral bone surface of the joint encompasses substantially all of a femoral condyle of the joint of the patient. 16. The method of claim 1, wherein the portion of the subchondral bone surface of the joint encompasses substantially all of a femoral notch of the joint of the patient. 17. The method of claim 1, wherein the por... | The Board of Trustees of the Leland Stanford Junior University | A61B005/00
US08652126 | B2 | 2014-02-18 | Method and computer program for authenticating a physiological sensor, a sensor system, a patient monitor, and a physiological sensor | 2009-11-24 | A mechanism for authenticating a physiological sensor is disclosed. When a sensor is connected to a monitor, the monitor examines whether a first sensor-specific usage identifier of the connected sensor is consistent with a second sensor-specific usage identifier thereof. The first and second sensor-specific usage identifiers are indicative of the cumulative usage of the connected sensor, but may nevertheless be unequal. The first sensor-specific usage identifier is maintained in the sensor and the second sensor-specific usage identifier in a host memory external to the sensor. The use of the connected sensor is allowed when the examining indicates that the said two identifiers of the connected sensor are consistent, and rejected when the identifiers are inconsistent. The two identifiers are further updated in response to the use of the connected sensor in the patient monitor, thereby to keep the said identifiers consistent and updated for a subsequent use attempt of the sensor. | 1. A method for authenticating a physiological sensor, the method comprising: examining, in response to a sensor being connected to a patient monitor, whether a first sensor-specific usage count of the connected sensor is greater than or equal to a second sensor-specific usage count of the connected sensor, wherein the first and second sensor-specific usage counts are indicative of cumulative usage of the connected sensor and wherein the first sensor-specific usage count is stored in the connected sensor and the second sensor-specific usage count is stored in a host memory external to the connected sensor;allowing use of the connected sensor in the patient monitor when the examining indicates that the first sensor-specific usage count is greater than or equal to the second sensor-specific usage count for the connected sensor;rejecting the use of the connected sensor in the patient monitor when the examining indicates that the first sensor-specific usage count is less than the second sensor-specific usage count of the connected sensor; andupdating the first and second usage counts for the connected sensor in response to the use of the connected sensor in the patient monitor, thereby to keep the first and second sensor-specific usage counts of the connected sensor consistent and updated for a subsequent use attempt of the connected sensor. 2. The method according to claim 1, wherein the updating includes updating the first usage count stored in the connected sensor, in which the first usage count is stored in a generic memory of one of EEPROM and EPROM memory types. 3. The method according to claim 1, further comprising storing a fixed sensor-specific count in a plurality of sensors, wherein the storing is performed prior to use of the plurality of sensors and wherein the connected sensor is any of the plurality of sensors. 4. The method according to claim 3, further comprising encrypting the fixed sensor-specific count and the first sensor-specific usage count prior to the storing; andencrypting the first sensor-specific usage identifier prior to the updating. 5. The method according to claim 1, wherein the examining includes retrieving the second sensor-specific usage through a network. 6. The method according to claim 1, wherein the examining includes relieving the second sensor-specific usage count from the host memory, wherein the host memory is in the patient monitor. 7. The method according to claim 6, wherein the updating includes incrementing the first usage count and the second usage count, wherein the first usage count is indicative of total number of times the connected sensor is used and the second usage count is indicative of total number of times the sensor is used in the patient monitor. 8. A computer program product for authenticating a physiological sensor connected to a patient monitor, the computer program product comprising: a first program product portion configured retrieve, upon connection of a sensor to a patient mon... | General Electric Company | A61B018/18
US09095349 | B2 | 2015-08-04 | Systems and methods for determining the likelihood of endocardial barotrauma in tissue during ablation | 2010-12-10 | A method and system for determining a likelihood of barotrauma occurring in tissue during formation of a lesion therein is provided. The system includes an electronic control unit (ECU). The ECU is configured to acquire at least one value of at least one component of a complex impedance between an electrode and the tissue. The ECU is configured to calculate an index responsive to the at least one value of the at least one complex impedance component. The index is indicative of a likelihood of barotrauma occurring in the tissue. The method comprises acquiring at least one value of at least one component of a complex impedance between an electrode and the tissue. The method comprises calculating an index responsive to the at least one value of the at least one complex impedance component. The calculated index is indicative of a likelihood of barotrauma occurring in the tissue. | 1. A system for determining a likelihood of barotrauma occurring in a tissue during formation of a lesion in said tissue as a result of an ablation procedure being performed thereon, comprising: an electronic control unit (ECU) configured to: acquire at least one value for at least one component of a complex impedance between an electrode and said tissue; andcalculate an index responsive to said at least one value for at least one component of said complex impedance, said index indicative of said likelihood of barotrauma occurring in said tissue;wherein said ECU is further configured to assess stability of said electrode based on predetermined stability criteria, and said calculated index is responsive to a determination by said ECU that said electrode meets said predetermined stability criteria. 2. The system of claim 1, wherein said electronic control unit is further configured to: acquire a value for power applied to said tissue during the formation of said lesion in said tissue; and said calculated index is responsive to said at least one value for at least one component of said complex impedance further using said value for power. 3. The system of claim 1, wherein said electronic control unit is configured to acquire values for first and second components of said complex impedance, said value for said first component of said complex impedance comprising said at least one value for at least one component of said complex impedance, and said calculated index is responsive to said at least one value for at least one component of said complex impedance further using said value for said second component of said complex impedance. 4. The system of claim 3, wherein said first and second complex impedance components collectively comprise two members of the group including the following members: a resistance between said electrode and said tissue, a reactance between said electrode and said tissue, a phase angle between said electrode and said tissue, and an impedance magnitude between said electrode and said tissue. 5. The system of claim 1, wherein said electrode comprises an ablation electrode on a radio-frequency ablation catheter. 6. The system of claim 1, wherein said electronic control unit is further configured to output said index to a display device. 7. The system of claim 1, wherein said electronic control unit is further configured to evaluate said index to determine said likelihood of barotrauma occurring in said tissue. 8. The system of claim 7, wherein said electronic control unit is further configured to: generate a signal representative of an indicator corresponding to said determined likelihood of barotrauma occurring in said tissue; and output said signal to a display device. 9. The system of claim 1, wherein said at least one component of said complex impedance is reactance, and said electronic control unit is configured to calculate one of a change in said reactance between a first reactance value and a second reactance va... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/053 A61B018/14 A61B017/00 A61B018/00 A61B019/00 A61B005/03
US08797714 | B2 | 2014-08-05 | Cables for patient monitoring and related systems with integrated front end | 2009-12-22 | Patient monitoring systems can include a system for transmitting information from a patient parameter sensor to a patient monitor. The system can include an analog-to-digital converter close to the patient parameter sensor and can transmit digital signals through a cable to the patient monitor. | 1. A patient monitoring system comprising: a patient monitor configured to display information regarding a patient, wherein the patient monitor comprises a first port;a cable comprising a proximal end and a distal end, and further comprising a cable body that extends between the proximal and distal ends, wherein the proximal end of the cable comprises a first connector configured to electrically couple with the first port of the patient monitor;an electrical interface configured to receive analog signals from a patient parameter sensor; anda first circuit including a front-end processing module configured to condition the analog signals from the patient parameter sensor and convert the analog signals into digital signals, wherein the first circuit is external to the patient monitor, wherein the first circuit is closer to the distal end of the cable than it is to the proximal end of the cable when the system is in an assembled state and the cable is fully extended to limit a distance over which the analog signal is transmitted over the cable, and wherein the first circuit is at a distance of no less than about 3.0 feet from the first connector and no more than about 3.0 feet from the distal end when the system is in an assembled state and the cable is fully extended. 2. The system of claim 1, wherein the cable comprises the first circuit. 3. The system of claim 1, wherein the first circuit is physically separate from the cable, and wherein the distal end of the cable comprises a second connector configured to couple with the first circuit. 4. The system of claim 3, further comprising a second circuit configured to convert analog signals received from a patient parameter sensor into digital signals, and wherein the second connector is configured to selectively couple with either the first circuit or the second circuit. 5. The system of claim 1, wherein the first circuit is further from the patient monitor than is the distal end of the cable when the system is in an assembled state and the cable is fully extended. 6. The system of claim 1, wherein the first circuit is at a distance of no more than about 2.0 feet from the distal end when the system is in an assembled state and the cable is fully extended. 7. The system of claim 1, wherein the patient monitor is configured to process digital information received via the first port in any of two or more operational modes, wherein the first circuit is configured to communicate an operational state of the first circuit to the patient monitor, and wherein the patient monitor is configured to automatically process said digital information received via the first port in an operational mode that corresponds with the communicated operational state of the first circuit. 8. The system of claim 7, wherein the first circuit is configured to transition among two or more operational states, and wherein each of said operational states corresponds with an operational mode of the patient monitor. 9. The system... | Mindray DS USA, Inc. | A61B005/00 A61B005/03 A61B005/0215
US08862204 | B2 | 2014-10-14 | Reducing mechanical stress on conductors and connection points in a position determinable interventional medical device | 2010-12-30 | A medical device configured for diagnosis or treatment of tissues within a body is provided. The device includes an elongate, deformable member configured to be received within a lumen in the body and having proximal and distal ends. A position sensor is disposed at the distal end. In one embodiment, a conductor is wound about the member. The conductor is connected to the position sensor and has a first winding pitch over a first portion of the deformable member and a second winding pitch, different from the first winding pitch, over a second portion of the deformable member. In another embodiment, the member defines a neutral longitudinal axis extending between the proximal and distal ends. A conductor extending between the proximal and distal ends is connected to the position sensor at a connection node on the neutral axis. | 1. A medical device configured for diagnosis or treatment of a tissue within a body, comprising: an elongate, deformable member configured to be received within a lumen in said body, having a proximal end and a distal end;a position sensor disposed at said distal end of said deformable member; and,a conductor wound about said deformable member, said conductor connected to said position sensor at a connection node, said conductor having a first winding pitch over a first portion of said deformable member and a second winding pitch over a second portion of said deformable member, said second winding pitch different from said first winding pitch. 2. The medical device of claim 1 wherein said second portion of said deformable member is distal to said first portion of said deformable member and said second winding pitch is smaller than said first winding pitch. 3. The medical device of claim 2 wherein the difference in said first and second winding pitches causes a change in a mechanical property of said deformable member. 4. The medical device of claim 3 wherein said change comprises an increase in said mechanical property and said mechanical property comprises flexibility. 5. The medical device of claim 1 wherein said position sensor comprises a coil. 6. The medical device of claim 5 wherein said coil is wound about said distal end of said deformable member. 7. The medical device of claim 1 further comprising a connector disposed at said connection node and connected to said position sensor and said conductor. 8. The medical device of claim 1 wherein said connection node is located on a neutral longitudinal axis extending between said proximal and distal ends of said member. 9. The medical device of claim 1 further comprising a compression coil disposed within said first portion of said deformable member. 10. The medical device of claim 1 further comprising: a plurality of pull wires disposed within said deformable member; and,a handle having a first actuator coupled to at least one pull wire of said plurality of pull wires. 11. The medical device of claim 1 wherein said first portion of said deformable member is made from a material having a first durometer and said second portion of said deformable member is made from a material having a second durometer, said second durometer less than said first durometer. 12. The medical device of claim 1 wherein said conductor is configured to transmit signals to or from said position sensor. 13. A medical device configured for diagnosis or treatment of a tissue within a body, comprising: an elongate, deformable member configured to be received within a lumen in said body, having a proximal end and a distal end, said deformable member defining a neutral longitudinal axis extending between said proximal and distal ends;a position sensor disposed at said distal end of said deformable member; and,a conductor extending between said proximal and distal ends of said deformable members, said conductor con... | MediGuide Ltd. | A61B005/05 A61B017/3207 A61B018/14 A61N007/02 A61M029/02 A61B019/00 A61B017/00 A61B018/24
US08417314 | B2 | 2013-04-09 | Ruthenium purple biosensor | 2008-01-03 | The invention relates to a biosensor comprising an electrically conductive substrate, with a first layer comprising Ruthenium Purple formed on the substrate, a second layer comprising polyaniline or a derivative thereof comprising one or more non-polar substituents formed on the first layer, and a third layer comprising one or more enzymes trapped within a matrix formed on the second layer. The biosensor is for use in the detection of analytes such as purines and derivatives thereof, particularly in the detection of hypoxanthine. | 1. A biosensor comprising: (i) an electrically conductive substrate;(ii) a first layer on the substrate, the first layer comprising Ruthenium Purple;(iii) a second layer on the first layer, the second layer comprising polyaniline or a derivative thereof comprising one or more non-polar substituents; and(iv) a third layer on the second layer, the third layer comprising one or more enzymes preferably trapped within a matrix or adsorbed onto the second layer. 2. A biosensor according to claim 1, wherein the non-polar substituents on the polyaniline derivative comprise one or more alkyl chains and/or ether groups. 3. A biosensor according to claim 1, wherein the non-polar substituents on the polyaniline derivative are selected from the group consisting of C1-C12 alkyl chains ether groups containing C1-C12 alkyl chains, and combinations thereof. 4. A biosensor according to claim 1, wherein the layer comprising the matrix is a sol-gel layer. 5. A biosensor according to claim 4, wherein the sol-gel layer comprises one or more silicon-based compounds, alumina, a ceramic oxide, a colloidal metal hydroxide. 6. A biosensor according to claim 1, wherein the matrix layer comprises a material selected from the group consisting of hydrogel, polyvinyl, polyacrylamide, polypyrrole, and sugar-derivatised polypyrroles. 7. A biosensor according to claim 1, wherein the matrix layer comprises one or more compounds selected from the group consisting of a silicate, a silane having the general formula: and (R9O)3—Si—CH2—CH2—Si—(OR9)3 where:R1=straight chain, branched chain, cyclic, non-cyclic, saturated or non-saturated, substituted or non-substituted alkyl; substituted or non-substituted aryl; —NR5; and —COR6;R2, R3 and R4 are independently selected from; straight chain, branched chain, cyclic, non-cyclic, saturated or non-saturated alkyl; —COR6; —O-alkyl; and —O—COR6; —R7R8; R7N(R6)2, and R7NHR6R8;R5=branched or non-branched cyclic or non-cyclic, saturated or non-saturated alkyl; or benzyl;R6=C1 to C3 alkyl;R7=C1 to C6 alkyl;R8=Epoxy, —NH2 or —SH;R9=Straight or branched C1 to C6 alkyl. 8. A biosensor according to claim 7, wherein R1 is methyl, ethyl, propyl, —NCHCH2CH3, —NC(CH3) CH2CH3, —NC(CH3)CH2CH(CH3)2, COCH3, orand R2, R3, and R4 are each independently methyl, ethyl, propyl, O-methyl, O-ethyl, O-propyl, —CHCH2, benzyl, or —OCOCH3. 9. The biosensor of claim 7, wherein R1 is an alkyl group having 1-6 carbon atoms. 10. The biosensor of claim 7, wherein R2, R3, and R4 are each aminoalkyl. 11. The biosensor of claim 7, wherein R2, R3, and R4 are each an alkyl group having 1-6 carbon atoms. 12. The biosensor of claim 7, wherein R5 contains 1-6 carbon atoms. 13. The biosensor of claim 7, wherein R7 is an alkyl group containing 1-3 carbon atoms. 14. The biosensor of claim 7, wherein R8 is 15. The biosensor of claim 7, wherein R9 is an alkyl group containing 1-3 carbon atoms. 16. A biosensor according to claim 1, further comprising a silane coupling agent in... | University of Warwick | G01N027/327
US08634903 | B2 | 2014-01-21 | Measuring T-Wave alternans | 2009-10-30 | An implantable medical device (IMD), such as an implantable pacemaker, cardioverter, or diagnostic device, generates an EGM signal, e.g., a far field EGM signal, samples the EGM signal to obtain a single T-wave amplitude value for each T-wave over a plurality of beats, and stores the T-wave amplitude values in memory. The IMD creates a time series of the T-wave amplitude values stored in memory, calculates the power spectral density for the times series, and selects a power spectral density of a particular frequency, e.g., 0.5 cycles per beat, as the TWA value. The IMD may periodically determine TWA values for the patient and store the values in memory. The TWA values may be presented to medical personnel, e.g., as a trend. The IMD may deliver or modify therapy, or provide an alert, based on the TWA values. | 1. A method comprising: generating an electrocardiogram (ECG) signal of a patient via one or more medical leads attached to the patient;sampling the ECG signal at a plurality of sample times over each of a plurality of T-waves for each of a plurality of beats to obtain a sample point from each of the T-waves for each of the sample times;determining a plurality of ECG-based T-wave alternans (TWA) values, wherein each of the ECG-based TWA values is determined from the sample points from each of the T-waves associated with a respective sample time during each of the T-waves;determining one of the plurality of sample times associated with a largest one of the ECG-based TWA values;selecting an interval based on the sample time associated with the largest one of the plurality of ECG-based TWA values;generating an electrogram (EGM) signal of the patient using an implantable medical device;storing one or more samples of the EGM signal per heartbeat of the patient for a plurality of heartbeats, wherein for each heartbeat a stored sample is an amplitude of a T-wave that occurs after a fiducial point of the EGM signal by an amount of time equal to the selected interval; anddetermining an EGM-based TWA value based on the T-wave amplitudes of the EGM signal. 2. The method of claim 1, wherein storing and determining comprises storing and determining by the implantable medical device. 3. The method of claim 1, wherein the fiducial point is a QRS complex. 4. The method of claim 1, further comprising: selecting a plurality of intervals based on sample times associated with the largest ECG-based TWA values for different heart rates;determining a heart rate of the patient based on the EGM signal; andselecting one of the plurality of intervals based on the determined heart rate. 5. The method of claim 1, further comprising: determining a heart rate of the patient based on the EGM signal; andautomatically adjusting the interval based on the heart rate. 6. The method of claim 1, further comprising generating an alert if the EGM-based TWA value exceeds a threshold. 7. The method of claim 1, further comprising: repeating, over multiple time periods, the generating, storing, and determining to determine a plurality of EGM-based TWA values; anddisplaying the EGM-based TWA values as a trend. 8. The method of claim 1, wherein determining the EGM-based TWA value comprises: creating a time series of the T-wave amplitude values of the EGM signal;computing a power spectrum of the time series; andselecting the power spectrum value at a particular frequency as the EGM-based TWA value. 9. The method of claim 1, wherein determining the EGM-based TWA value comprises: computing a difference between samples of consecutive T-waves of the EGM signal; anddetermining the EGM-based TWA value based on the difference values. 10. The method of claim 1, wherein storing one or more samples of the EGM signal comprises storing fifty samples or less of the EGM signal per heart beat of the... | Medtronic, Inc. | A61B005/0468
US20110152643 | A1 | 2011-06-23 | "Band-aid"-type potassium ion (K+) biosensor | 2010-10-12 | Potassium ion (K+) is important in regulating normal cell function in the human body, specifically the heartbeat and the muscle function. It is important to be able to monitor potassium ion concentrations in human fluids. This invention describes a novel concept for a potassium ion biosensor that accurately, rapidly, and efficiently monitors the presence and records the concentration of potassium ions with high specificity, not only in serum and urine, but also in the sweat or even eye fluid. This specific biosensor design utilizes a nanomanufacturing technique, i.e. electrospinning, to produce advanced nano-bio-composites that specifically trace even minute quantities of potassium ions through the use of selective bio-receptors (ionophores) attached to high surface area nanofibers. Electroactive polymers are then employed as transducers to produce an electronic (rather than ionic) output that changes instantly with the change in K+ concentration. Such biosensors may be manufactured in a skin patch configuration. | 1. This invention claims a device for the monitoring of potassium ion concentration in sweat and/or eye fluid, and/or urine, and/or serum. 2. The device of claim 1 whereas the device is a skin patch. 3. The device of claim 1 whereas the potassium ion sensor is a bio-active nanofibrous mat. 4. The device of claim 1 whereas the biosensor is produced by the electrospinning processing of polyvinylpyrrolidone powder dissolved in ethanol, and (ionophore) valinomycin solutions together. 5. The device of claim 1 whereas an ion-to-electron transducer is also used. 6. The device of claim 1 whereas the ion-to-electron transducer may be electro-active polymer such as polyaniline (PANI). 7. The device of claim 1 whereas PANI may be processed by electrospinning. 8. The device of claim 1 whereas this device can be a resistive chemodetector. 9. The device of claim 1 whereas the sensor output is transmitted wirelessly. 10. The device of claim 1 whereas this device is a medical diagnostics tool. | - | A61B005/00
US20110144514 | A1 | 2011-06-16 | Tracheal Tube with Pressure Monitoring Lumen and Method for Using the Same | 2009-12-16 | According to various embodiments, methods and systems for determining pressure in the lungs may employ tracheal pressure measurements. The tracheal pressure measurements may be obtained through a pressure monitoring lumen associated with a tracheal tube. Such systems may include a purging or flushing mechanism to keep the pressure monitoring lumen free of any obstructions. The flushing mechanism may utilize respiratory gases diverted from the airway stream and regulated to flush the lumen at relatively low pressures. The resulting pressure measurements may be used to determine a more accurate estimate of lung pressure, which in turn may be used to control a ventilator and provide breathing assistance to a patient. | 1. A system for trachea pressure measurement comprising: a tracheal tube comprising a conduit configured to be inserted into the trachea of a subject to deliver respiratory gases to the subject;a pressure monitoring lumen associated with the tracheal tube and in fluid communication with a pressure transducer;a conduit configured to divert a portion of the respiratory gases to the pressure monitoring lumen;a fluid reservoir in fluid communication with the conduit and configured to allow respiratory gases to accumulate to a first pressure within the reservoir; anda pressure regulator configured to allow respiratory gases from the fluid reservoir to flow into the pressure monitoring lumen. 2. The system of claim 1, wherein the tracheal tube comprises an information element comprising a memory circuit storing calibration data for the system. 3. The system of claim 1, comprising a processor coupled to the pressure transducer, wherein the processor comprises encoded instructions for determining pressure based on measurements from the pressure transducer. 4. The system of claim 3, wherein the encoded instructions for determining pressure comprise instructions for accounting for flow from the fluid reservoir. 5. The system of claim 4, wherein the instructions for accounting for flow from the fluid reservoir comprise calibrating to a baseline pressure. 6. The system of claim 3, wherein the pressure regulator is coupled to the processor and the processor is configured to control the flow of respiratory gases from the fluid reservoir. 7. The system of claim 1, comprising one or more valves configured to facilitate one-way flow of respiratory gases into the fluid reservoir. 8. The system of claim 1, comprising one or more valves configured to facilitate one-way flow of respiratory gases from the fluid reservoir into the pressure monitoring lumen. 9. A method for determining trachea pressure in a patient under mechanical ventilation comprising: communicating a trachea pressure present in a patient's trachea with a pressure monitoring lumen associated with a tracheal tube;measuring the trachea pressure with a pressure transducer in fluid communication with the pressure monitoring lumen; anddiverting a portion of a patient's respiratory gases from a point on a respiratory circuit upstream of the patient's airway such that a generally continuous pressure of the respiratory gases is applied to the pressure monitoring lumen. 10. The method of claim 9, wherein the continuous pressure is less than about 5 cm H2O. 11. The method of claim 9, wherein the continuous pressure is between about 1 cm H2O and 5 cm H2O. 12. The method of claim 9, wherein the continuous pressure is between about 2 cm H2O and 3 cm H2O. 13. The method of claim 9, wherein the continuous pressure is greater than a lowest pressure after exhalation and less than about 10 cm H2O. 14. The method of claim 9, comprising determining a baseline pressure in the pressure monitoring lumen using ... | NELLCOR PURITAN BENNETT LLC | A61B005/08 A61M016/04
US20110144539 | A1 | 2011-06-16 | Concussion Warning Apparatus | 2009-12-11 | The present invention is related to detecting possible concussions due to blunt force blows or shocks to the head. In the present invention, areas of the athlete's head are covered with a blunt force blow indicating sheet prior to the possible concussion causing activity and the sheet is examined for indications of blunt force blows that may have caused a concussion. The athlete with indications of possible concussion is then examined to determine if the athlete suffered a concussion. The blunt force blow indicating sheet may be used to line a helmet or headgear used by athletes or the sheet may be shaped like a helmet for use in activities where headgear is not worn. Objects that may be subject to blunt force blows, during for instance shipping, is covered with the sheet to detect if the object experienced a blunt force blow that may have damaged the object. | 1. A concussion warning apparatus comprising: an indicator sheet with indicators where an indicator changes when struck with a blunt force blow impact that is sufficient to cause a concussion;areas of an athlete's head are covered by the indicator sheet prior to an activity that may cause a concussion;the indicator sheet is examined for a changed indicator as a warning of a blunt force blow that may have caused a concussion. 2. The indicator sheet of claim 1 is used to line the inside of a helmet or headgear used by the athlete. 3. The indicator sheet of claim 1 provides a first indicator that changes at first blunt force blow impact and a second indicator that changes at a second blunt force blow impact different from the first blunt force blow impact. 4. The indicator sheet of claim 1 provides indicators that change at different blunt force blow impact and the areas of the athlete's head that are subject to different blunt force blow impacts that may cause a concussion are covered with indicators with change force impact appropriate for the area. 5. The indicator sheet of claim 1 provides indicators that change after two blunt force blows that are of lower impact where a single blow of lower impact would not change the indicator but the indicator would change with a single blow of higher impact. 6. The indicator sheet of claim 1 where a changed indicator provides a color change indication that spreads across an area of the indicator sheet. 7. The indicator sheet of claim 1 provides a first indicator of a first size and a second indicator of a second size different from the first size. 8. The indicator sheet of claim 1 provides indicators of different sizes and the areas of the athlete's head that require higher resolution indication of the impact area are covered with smaller size indicators than areas with lower resolution indication with larger size indicators. 9. The indicator sheet of claim 1 is shaped to conform to an athlete's head and worn by the athlete without a helmet or headgear. 10. The indicator sheet of claim 1 where the indicators are integrated with the cushion layer of a helmet and the cushion layer is examined for a blunt force blow indication as a warning of a possible concussion. 11. A blunt force blow indicating sheet for warning of concussions where areas of the head are covered by the indicating sheet prior to an activity that may cause concussions and the sheet is examined for a blunt force blow indication as a warning of a possible concussion. 12. The indicating sheet of claim 11 is used to line a helmet or headgear. 13. The indicating sheet of claim 11 where the sheet provides a visual indicator that spreads across the area where a blunt force has been experienced. 14. The indicating sheet of claim 11 where the sheet provides an indication of two blunt force blows of lower intensity to an area of the sheet where a single blow of lower intensity would not have been sufficient for an indication but a single ... | - | A61B005/103
US20110118590 | A1 | 2011-05-19 | System For Continuous Cardiac Imaging And Mapping | 2010-07-19 | A system improves precision and reliability of intra-cardiac catheter position tracking and monitoring. An interventional system for internal anatomical examination includes a catheterization device for internal anatomical insertion. The catheterization device includes, at least one magnetic field sensor for generating an electrical signal in response to rotational movement of the at least one sensor about an axis through the catheterization device within a magnetic field applied externally to patient anatomy and a signal interface for buffering the electrical signal for further processing. A signal processor processes the buffered electrical signal to derive a signal indicative of angle of rotation of the catheterization device relative to a reference. The angle of rotation is about an axis through the catheterization device. A reproduction device presents a user with data indicating the angle of rotation of the catheterization device. | 1. An interventional system for internal anatomical examination, comprising: a catheterization device for internal anatomical insertion including, at least one magnetic field sensor for generating an electrical signal in response to rotational movement of said at least one magnetic field sensor about an axis through said catheterization device within a magnetic field applied externally to patient anatomy anda signal interface for buffering said electrical signal for further processing;a signal processor for processing the buffered electrical signal to derive a signal indicative of angle of rotation of said catheterization device relative to a reference, said angle of rotation being about an axis through a catheter; anda reproduction device for presenting a user with data indicating said angle of rotation of said catheterization device. 2. A system according to claim 1, wherein said at least one magnetic field sensor comprises a plurality of sensors in substantially mutually orthogonal orientation for generating a corresponding plurality of electrical signals,said signal interface buffers said electrical signals for further processing andsaid signal processor processes the buffered electrical signals to derive a signal indicative of angle of rotation of said catheterization device. 3. A system according to claim 2, wherein said plurality of sensors are located in a tip of said catheterization device and said signal indicative of angle of rotation of said catheterization device indicates angle of rotation of said tip of said catheterization device. 4. A system according to claim 2, wherein said plurality of sensors comprises three sensors. 5. A system according to claim 2, wherein said signal processor processes the buffered electrical signals to derive a signal indicative of angle of rotation of said plurality of sensors. 6. A system according to claim 2, wherein said reference comprises a rotational angle of said catheterization device substantially at initial entry of said catheterization device into patient anatomy. 7. A system according to claim 2, wherein said reference is determined by said magnetic field applied externally to patient anatomy. 8. A system according to claim 2, wherein said reference comprises a rotational angle of another portion of said catheterization device. 9. An interventional system for internal anatomical examination, comprising: a catheterization device for internal anatomical insertion including, a plurality of sensors in substantially mutually orthogonal orientation for generating electrical signals in response to rotational movement of said plurality of sensors about an axis through said catheterization device within a magnetic field applied externally to patient anatomy anda signal interface for buffering said electrical signals for further processing;a signal processor for processing the buffered electrical signals to derive a signal indicative of angle of rotation of said catheterization device relative... | SIEMENS MEDICAL SOLUTIONS USA, INC. | A61B005/05
US07953468 | B2 | 2011-05-31 | Magnetic resonance imaging using technique of positioning multi-slabs to be imaged | 2006-01-12 | A magnetic resonance imaging system is provided for obtaining MR images by scanning a region of the object previously located on an object's positioning image. The system comprises a displaying unit, inputting unit, approximating unit, and locating unit. The displaying unit displays a plurality of tomographic images of the object as the positioning image, each of the tomographic images including an indication of a target of interest thereon. The inputting unit enables information about a running state of the target in a direction along the target to be supplied toward each of the tomographic images. The approximating unit calculates three-dimensionally an approximated curve indicating the running state of the target in the direction on the basis of the supplied information about the running state. The locating unit locates the region substantially perpendicular to the approximated curve. | 1. A magnetic resonance imaging system in which an object to be imaged is subjected to scanning to obtain an MR image, the scanning being applied to a three-dimensional region of the object previously located on previously obtained tomographic positioning images of the object, the system comprising: a displaying unit configured to display two tomographic images of the object as pre-scan positioning images, each of the two tomographic images including a target of interest thereon, the two tomographic images intersecting each other;an inputting unit configured to enable an operator to specify a plurality of points along the target on each of the two tomographic images;an approximating unit configured to calculate an approximated three-dimensional curve passing through a plurality of virtual points determined by the plurality of operator-specified points on each of the two tomographic images; anda locating unit configured to locate a region to be imaged by using the calculated approximated curve. 2. The magnetic resonance imaging system of claim 1, wherein the region is one or more slabs each consisting of one or more slices, and the region is substantially perpendicular to the approximated curve. 3. The magnetic resonance imaging system of claim 2, wherein the two tomographic images are substantially perpendicular to each other. 4. The magnetic resonance imaging system of claim 2, wherein: the approximating unit is configured to include a first element configured to project points specified on one tomographic image onto the other tomographic image,a second element configured to allow the operator to move the projected points on the other tomographic image,a third element configured to calculate said approximated three-dimensional curve as passing through the operator-specified and adjusted points on each of the two tomographic images; andthe locating unit includes a first element configured to allow the operator to select a desired position along the approximated curve, anda second element configured to locate a said slab on the approximated curve so that the slab is substantially perpendicular to the approximated curve. 5. The magnetic resonance imaging system of claim 4, wherein the two tomographic images are a sagittal image and a coronal image of the target of the object, the target being a vertebral column of the object. 6. The magnetic resonance imaging system of claim 5, wherein the locating unit includes a third element configured to allow the operator to specify the number of slices composing each of the slabs, wherein the second element of the locating unit is configured to set the slab including slices of which number is specified by the third element of the locating unit. 7. The magnetic resonance imaging system of claim 6, wherein the third element of the locating unit is configured to determine which slice of each slab is positioned at the operator-specified point. 8. The magnetic resonance imaging system of claim 1, further co... | Kabushiki Kaisha Toshiba | A61B005/055
US07962194 | B2 | 2011-06-14 | Method and system for propelling and controlling displacement of a microrobot in a blood vessel | 2003-04-15 | A magnetic resonance imager (MRI) controllably propels at least one magnetic field responsive body within a subject. The MRI comprises a bore magnet producing an intense magnetic field sufficient to magnetically saturate the magnetic field responsive body. A magnetic field gradient generator is configured to apply variable magnetic gradients to at least a portion of the subject containing the magnetically-saturated body. A controller instructs the magnetic field gradient generator to acquire image data of the magnetically-saturated body within the subject and to apply a calculated magnetic gradient that will propel the magnetically-saturated body towards a target location in the subject. A tracking unit of the MRI provides position feedback information of the magnetically-saturated body within the subject. | 1. An apparatus for tracking, propelling and controlling the displacement of an untethered body including magnetic field responsive material in a blood vessel of a subject, the apparatus comprising: a magnetic resonance imager (MRI) including: a bore magnet configured to produce an intense magnetic field sufficient to magnetically saturate the magnetic field responsive material of the body, said bore magnet having a hollow interior sized to receive the subject therein;a multi-dimensional magnetic field generator configured to generate a variable magnetic field gradient having a direction and magnitude; anda tracking unit configured to acquire an MRI image of the position of the body in the blood vessel and to provide position feedback information;a trajectory calculator configured to calculate a trajectory of the body in the blood vessel based at least upon the position feedback information of the body determined by the tracking unit of the MRI, a configuration of the blood vessel, and a target location to be reached by the body; anda gradient calculator configured to calculate the direction and magnitude of the gradient of the multi-dimensional magnetic field required to be applied to the magnetically-saturated body in order to propel and displace the body along the trajectory calculated by the trajectory calculator;wherein a) the generator of the MRI is configured to apply the gradient calculated by the gradient calculator to the body, and b) both the generator and the tracking unit are part of the same MRI. 2. An apparatus as defined in claim 1, wherein the body includes a shell that comprises the magnetic field responsive material. 3. An apparatus as defined in claim 1, wherein the magnetic field responsive material is a ferromagnetic material. 4. An apparatus as defined in claim 1, wherein the gradient calculator is configured to calculate the direction and magnitude of the gradient of the multi-dimensional magnetic field using: the equation {right arrow over (τ)}={right arrow over (m)}×{right arrow over (B)}=Vferro{right arrow over (M)}×{right arrow over (B)}; andthe equation {right arrow over (F)}magnetic=Vferro{right arrow over (M)}·∇{right arrow over (B)}, wherein:{right arrow over (τ)}=torque induced on the body;{right arrow over (m)}=magnetic moment of the magnetic field responsive material;{right arrow over (B)}=multi-dimensional magnetic field;Vferro=volume of the magnetic field responsive material;{right arrow over (M)}=magnetization of the magnetic field responsive material;{right arrow over (F)}magnetic=force induced on the body;∇{right arrow over (B)}=gradient of the multi-dimensional magnetic field. 5. A method for tracking, propelling and controlling the displacement of an untethered body including magnetic field responsive material in a blood vessel of a subject, the method comprising: providing a MRI (Magnetic Resonance Imager) including: a bore magnet configured to produce an intense magnetic field sufficient to magnetic... | Polyvalor, Limited Partnership Val-Chum, Limited Partnership | A61B005/05
US08015978 | B2 | 2011-09-13 | Locating an implanted object based on external antenna loading | 2010-12-22 | In general, the invention is directed to techniques for locating an implanted object using an external antenna. The implanted object may be, for example, an internal antenna that facilitates recharging of and/or communication with an implantable medical device. An external device coupled to the external antenna drives the antenna with a plurality of waveforms. Asymmetry in the loading profile of the external antenna when it is driven by the plurality of waveforms allows the external device or another device to determine the location of the implanted object relative to the external antenna. The external device or other device may provide information to a user based on the determined location of the implanted object relative to the external antenna, such as information to help a user position the external antenna with respect to an internal antenna in embodiments in which the implanted object is an internal antenna. | 1. A system comprising: an external antenna;an external device coupled to the external antenna that generates a first waveform, drives the external antenna with the first waveform while the external antenna is moved through a three-dimensional region proximate to an implantable object a first time, periodically measures values of an electrical parameter associated with the external antenna while the external antenna is moved through the three-dimensional region the first time and driven with the first waveform, generates a second waveform that has at least one different waveform parameter than the first waveform, drives the external antenna with the second waveform while the external antenna is moved through the three-dimensional region a second time, and periodically measures values of the electrical parameter while the external antenna is moved through the three-dimensional region the second time and driven with the second waveform; anda processor that determines a first position/electrical parameter relationship based on the electrical parameter values measured while the external antenna was driven with the first waveform and positions within the three-dimensional region at which the external antenna was located when the electrical parameter values were measured, determines a second position/electrical parameter relationship based on the electrical parameter values measured while the external antenna was driven with the second waveform and positions within the three-dimensional region at which the external antenna was located when the electrical parameter values were measured, and determines first and second equations based on the first and second position/electrical parameter relationships,wherein the first equation defines a first position variable as a function of the electrical parameter and the second equation defines a second position variable as a function of the electrical parameter. 2. The system of claim 1, wherein the first position variable comprises a radial distance, and the second position variable comprises a depth. 3. The system of claim 1, wherein the external device generates a third waveform that has at least one different waveform parameter than the first and second waveforms, drives the external antenna with the third waveform while the external antenna is moved through the three-dimensional region a third time, and periodically measures values of the electrical parameter while the external antenna is moved through the three-dimensional region the third time and driven with the third waveform,wherein the processor determines a third position/electrical parameter relationship based on the electrical parameter values measured while the external antenna is driven with the third waveform and positions within the three-dimensional region at which the external antenna was located when the electrical parameter values were measured, and determines the first and second equations based on the first, second and third position/elec... | Medtronic, Inc. | A61B019/00 A61B005/00
US07942824 | B1 | 2011-05-17 | Integrated sleep diagnostic and therapeutic system and method | 2007-07-19 | The present invention relates to an integrated sleep diagnosis and treatment device, and more particularly to an integrated apnea diagnosis and treatment device. The present invention additionally relates to methods of sleep diagnosis and treatment. The sleep disorder treatment system of the present invention can use a diagnosis device to perform various forms of analysis to determine or diagnose a subject's sleeping disorder or symptoms of a subject's sleep disorder, and using this analysis or diagnosis can with or in some embodiments without human intervention treat the subject either physically or chemically to improve the sleeping disorder or the symptoms of the sleeping disorder. The diagnostic part of the system can use many different types of sensors and methods for diagnosing the severity of the symptoms of or the sleep disorder itself. The treatment part of the system can use a device to physically or chemically treat the subject's symptoms or sleep disorder itself. | 1. A sleep-related breathing disorder treatment system comprising: a portable data acquisition system comprising patient interface box further comprising at least one sensor having a signal, the sensor for determining a physiological characteristic of a subject, and at least one electronic component for receiving the signal and retransmitting the signal from the at least one sensor or transmitting a signal based at least in part on the signal from the at least one sensor to a separate PAP or CPAP device; andthe PAP or CPAP device for treating the subject's sleep related breathing disorder, the PAP or CPAP device comprising an electrical component for receiving the retransmitted or transmitted signal and a nasal cannula or mask for delivering a gas for treating the subject's sleep related breathing disorderwherein the PAP or CPAP device is titrated or adjusted based on at least in part the retransmitted or transmitted signal and the titration or adjustment is also based at least in part on training data representative of physiological signals that are specific to that subject, derived from data collected from two or more days of use of the sleep-related breathing disorder treatment system. 2. The system of claim 1, wherein the signal is transmitted or retransmitted wirelessly. 3. The system of claim 1, wherein the signal is transmitted or retransmitted via USB. 4. The system of claim 1, wherein the data acquisition system further includes at least one additional sensor, the additional sensor for collecting a video signal. 5. The system of claim 1, wherein the data acquisition system comprises at least three sensors having signals, the sensors for determining at least respiratory effort, pulse oximetry, and airflow of a subject; and at least one electronic component for receiving the signals and retransmitting the signals or transmitting a signal based at least in part on the signal from the at least one of the at least three sensors. 6. The system in claim 1, wherein the system further comprises a remote monitor adapted to assess the quality of signals received from the data acquisition system to ensure receipt of sufficient and adequate data to perform the titration correctly, and wherein the system is programmed to ignore sensor signals if the signals are assessed as being corrupted or if the signals are assessed as coming from a sensor that has been disconnected. 7. A sleep-related breathing disorder treatment system comprising: a portable data acquisition system comprising a patient interface box further comprising at least one sensor having an electrical signal, the sensor for determining a physiological characteristic of a subject; at least one electronic component for receiving the electrical signal from the at least one sensor, generating a processed signal based at least in part on the signal from the at least one sensor, and transmitting the processed signal; and a component for wireless radio frequency transmission;a separate PAP ... | Cleveland Medical Devices Inc. | A61B005/08 A61B005/02 A62B007/04
US07955268 | B2 | 2011-06-07 | Multiple sensor deployment | 2007-07-20 | A method of measuring pressure within the human body including implanting a pressure sensing assembly having a flexible structure and first and second sensor elements having self-contained power supplies coupled to the flexible structure. Periodic data collection events are performed to collect data from the sensor elements. A data collection event includes a request for data from a remote communication device, a transfer of sensor data to the remote communication device and a processing of the sensor data. The invention also includes a sensor assembly for implantation into a human body. The sensor assembly includes a first sensor having a self-contained power supply, a sensing element and an integral communication device capable of communicating with a remote communication device. The sensor assembly also includes a second sensor and a flexible structure to which the first and second sensor are attached. | 1. A method of measuring blood pressure within a human body, comprising: implanting a pressure sensing assembly into a vessel near the heart in the human body, the sensor assembly including a flexible structure capable of expanding when implanted within the vessel, a first sensor member having a self contained electric power supply coupled to the flexible structure, and a second sensor member having a self contained electric power supply coupled to the flexible structure;wirelessly transmitting a sensor data request message from a pulse generator implanted within the body to the first and second sensor members, the sensor data request message configured to prompt each of the first and second sensor members to transition from a power save state to an initialization state for a predetermined period of time for processing incoming communication data;comparing an address in the sensor data request message against a unique address of the first and second sensor members and, upon receiving a matching address, selectively activating one of the first or second sensor members from the initialization state to an activation state in which a pressure sensor is activated for sensing pressure data; andperforming periodic data collection events, including: collecting a sensor data packet wirelessly transmitted to the pulse generator from the activated sensor member, the sensor data packet including a power supply value representing the charge level of the self-contained electric power supply of the activated sensor member;comparing the power supply value of the activated sensor member against a threshold value;wherein, if the activated sensor member has an adequate charge remaining, waiting for a predetermined period of time and then collecting another sensor data packet wirelessly transmitted to the pulse generator from the activated sensor member;wherein, if the activated sensor member does not have an adequate charge remaining, waiting for a predetermined period of time and then collecting a sensor data packet wirelessly transmitted to the pulse generator from the other sensor member, the sensor data packet from the other sensor member including a power supply value representing the charge level of the self-contained electric power supply of the other sensor member; andperiodically performing the steps of comparing the power supply value of an activated sensor member against a threshold value to determine a charge level and alternating the collecting of sensor data packets from either the first or second sensor member based at least in part on the charge level. 2. The method of claim 1, wherein the step of performing a data collection event includes collecting a plurality of sensor data packets from one of the first sensor member and the second sensor member. 3. The method of claim 1, wherein the step of performing a data collection event includes collecting a plurality of sensor data packets from each of the first sensor member and the second sensor membe... | Cardiac Pacemakers, Inc. | A61B005/02
US07966052 | B2 | 2011-06-21 | Textile-based electrode | 2007-02-08 | Two garments with textile based electrodes are disclosed. First, a wrist band for use with a cardiac patient remote monitoring system includes two fabric layers with integral textile-based electrodes. The textile based electrodes include a fabric portion having stretch-recovery non-conductive yarns and an electrically conductive region having stretch-recovery electrically conductive yarn filaments. The skin contacting surface of the band includes a conductive region formed as a continuous ring or stripe. A connector links the conductive region to a lead to a device. Second, an infant garment includes textile based electrodes at the torso region and optionally at other regions in order to monitor the infant's biophysical characteristics as the garment is worn. | 1. A wrist band for a cardiac monitoring system, comprising: a first fabric portion comprising a skin contacting surface and an inner surface and that includes stretch-recovery non-conductive yarns and at least one electrically conductive region comprising stretch-recovery electrically conductive yarn filaments;a second fabric portion comprising an outer surface forming the wrist band outer surface and an inner surface, and that includes stretch-recovery non-conductive yarns, wherein the second fabric portion overlays at least a portion of the first fabric portion; anda connector for connecting the at least one electrically conductive region of the first fabric portion to the cardiac monitoring system or to a lead for such system for transmitting signals representing biophysical data. 2. The wrist band of claim 1, wherein the first fabric portion and the second fabric portion are formed from a single fabric and wherein the second fabric portion overlays the first fabric portion upon folding the single fabric. 3. The wrist band of claim 2, wherein the folded fabric is attached to form a two layer construction. 4. The wrist band of claim 3, wherein the folded fabric is attached at its selvages by stitching. 5. The wrist band of claim 1, wherein the second fabric portion includes at least one electrically conductive region comprising stretch recovery electrically conductive yarn filaments, wherein the second fabric portion overlays at least a portion of the first fabric portion and the at least one electrically conductive region of the second fabric portion contacts the at least one electrically conductive region of the first fabric portion. 6. The wrist band of claim 5, wherein the connector is selected from the group consisting of: snap, rivet, pin and clamp. 7. The wrist band of claim 6, wherein the first fabric portion and the second fabric portion are formed from a single fabric and wherein the second fabric portion overlays the first fabric portion upon folding the single fabric, and wherein the first fabric portion is attached to the second fabric portion at a position adjacent to the connector. 8. The wrist band of claim 6, where the connector comprises one or more snaps adapted to face away from the skin. 9. The wrist band of claim 6, where the connector comprises one or more snaps adapted to face toward the skin. 10. The wrist band of claim 5, wherein the conductive region of the second fabric portion is a patch bordered by non-conductive region. 11. The wrist band of claim 5, wherein at least one of: (i) the electrically conductive region of the first fabric portion; and (ii) the electrically conductive region of the second fabric portion; comprises at least a portion of an elastified electrically conductive yarn. 12. The wrist band of claim 5, wherein at least one of: (i) the electrically conductive region of the first fabric portion; and (ii) the electrically conductive region of the second fabric portion; comprises an elasti... | Textronics, Inc. | A61B005/04
US07945301 | B2 | 2011-05-17 | Methods and apparatus for the administration of CO | 2009-07-14 | The present invention relates to a method and apparatus for the administration of carbon monoxide (CO) to a patient. It allows the safe and effective use of carbon monoxide as a medicament. | 1. A method for administering carbon monoxide to a patient, comprising the steps of: a) administering exogenous carbon monoxide to the patient,b) determining the concentration of carbon monoxide in the patient's blood,c) comparing the actual concentration of carbon monoxide in the blood with a preset, desired value using a central processing unit;and optionally repeating steps b)-d)wherein the administration of carbon monoxide is performed via sequences of pulses, wherein the number and length of the pulses in each sequence and/or the number of sequences is regulated depending on the determination of the concentration of carbon monoxide in the blood,wherein the carbon monoxide is delivered to a patient as a pure gas, in a gas mixture, dissolved in a fluid, or by administering a carbon monoxide donor, and,wherein the carbon monoxide containing gas mixture is administered by way of admixing it into the breathing air of a patient. 2. The method according to claim 1 wherein the concentration of carbon monoxide in the blood is determined by a method selected from the group consisting of: measuring the concentration of carboxyhemoglobin (HbCO) in the blood,measuring the concentration of oxyhemoglobin (HbO2) in the blood,measuring the activity of enzymes in the blood; andmeasuring the CO content of the air expired by the patient. 3. The method according to claim 2 wherein the carboxyhemoglobin (HbCO) is determined by non-invasive measurement, or is determined from a blood sample. 4. The method according to claim 2 wherein the measurement of the oxyhemoglobin (HbO2) in the blood is carried out by oxymetry. 5. The method according to claim 2 wherein a measurement of the composition of the expired air is carried out by spectroscopic or electrochemical methods. 6. The method according to claim 1 wherein the carbon monoxide is delivered to the patient through inhalation, insufflation, intravenously, or rectally. 7. The method according to claim 6 wherein the carbon monoxide is delivered for inhalation in pulses, wherein the pulses may be inspiration and/or expiration triggered. 8. The method according to claim 1 wherein the concentration of carbon monoxide in the blood is controlled by the oxygen content of the breathing gas. 9. The method according to claim 1 wherein the patient is spontaneously breathing or artificially breathing. 10. The method according to claim 1 wherein the concentration of carbon monoxide in the blood is determined by at least two separate methods of measurement. 11. An apparatus for administering carbon monoxide to a patient, the apparatus comprising: a delivering unit,a carbon monoxide source,a dosing unit for administering carbon monoxide to the patient,sensor means for regulating the dosing unit depending on the feedback from the sensor unit, and,a filter unit through which the air expired by the patient is passed in order to remove excess carbon monoxide from the expired gas, wherein the filter is a physical or a chemic... | Ikaria, Inc. | A61B005/05 A61M031/00 F16K031/02
US07938775 | B2 | 2011-05-10 | Device, system, and method for in-vivo analysis | 2005-06-28 | A device, system and method for in-vivo analysis. An autonomous in-vivo device may include a magnet to detain at least a portion of a sample collected from a body lumen; a sensor to sense a property of the detained sample portion; and a transmitter to transmit data of the sensed property. | 1. An autonomous in-vivo device comprising: a magnetized reagent to be mixed, within said autonomous in-vivo device, with a sample collected from a body lumen;a detector reagent to be mixed, within said autonomous in-vivo device, with said sample; whereby particles within a first portion of said sample, which contains a substance to be detected by said detector reagent, bind to said magnetized reagent and to said detector reagent, and particles within a second portion of said sample, which does not contain the substance to be detected by said detector reagent, do not bind to said magnetized reagent and to said detector reagent;a magnet to detain said first portion of said sample, said magnet not detaining said second portion of said sample;a sensor to sense a property of the detained sample portion; anda transmitter to transmit data of the sensed property. 2. The autonomous in-vivo device of claim 1, comprising an inlet to transfer said sample into said autonomous in-vivo device. 3. The autonomous in-vivo device of claim 1, comprising a duct to transfer said sample within said autonomous in-vivo device. 4. The autonomous in-vivo device of claim 1, comprising a micropump to pump in said sample into said autonomous in-vivo device. 5. The autonomous in-vivo device of claim 1, wherein the magnetized reagent includes a magnetized antibody. 6. The autonomous in-vivo device of claim 1, wherein the detector reagent includes a detector antibody. 7. The autonomous in-vivo device of claim 6, wherein said detector antibody comprises a color label. 8. The autonomous in-vivo device of claim 1, comprising an outlet to transfer out of the autonomous in-vivo device a portion of said sample not detained by said magnet. 9. The autonomous in-vivo device of claim 1, comprising an illumination source to illuminate said detained portion, and wherein said sensor comprises an imager to acquire in-vivo an image of said detained portion. 10. The autonomous in-vivo device of claim 1, comprising a first reservoir to store said magnetized reagent and a second reservoir to store said detector reagent. 11. The autonomous in-vivo device of claim 1, comprising an imager. 12. The autonomous in-vivo device of claim 1, wherein said autonomous in-vivo device comprises a swallowable capsule. 13. A system comprising: an autonomous in-vivo device including at least a magnetized reagent to be mixed, within said autonomous in-vivo device, with a sample collected from a body lumen,a detector reagent to be mixed, within said autonomous in-vivo device, with said sample; whereby particles within a first portion of said sample, which contains a substance to be detected by said detector reagent, bind to said magnetized reagent and to said detector reagent, and particles within a second portion of said sample, which does not contain the substance to be detected by said detector reagent, do not bind to said magnetized reagent and to said detector reagent;a magnet to detain said first ... | Given Imaging, Ltd. | A61B005/00
US08280479 | B2 | 2012-10-02 | Flexible 3D microprobe structure | 2009-12-14 | The present invention discloses a flexible 3D microprobe structure, which comprises at least one probe, a base and a hinge portion. The probe is connected to the base via the hinge portion. The probe forms a bend angle with respect to a normal of the base by attracting the probe through an electrostatic force to make the hinge portion bend with respect to the base, and thus to form a 3D structure having the bend angle. The probe, the base and the hinge portion are made of a flexible polymeric material to reduce the inflammation response of creatures. Further, a fixing element is used to enhance the structural strength of the flexible 3D microprobe structure. | 1. A flexible 3-dimensional microprobe structure comprises at least one probe, a base, a hinge portion, and a fixing element, wherein said probe and said hinge portion are made of a first flexible polymeric material, and wherein said probe is connected to said base via said hinge portion, and wherein said probe forms a bend angle with respect to a normal of said base, and wherein said fixing element is formed by applying glue to said hinge portion. 2. The flexible 3-dimensional microprobe structure according to claim 1, wherein said first flexible polymeric material is selected from a group consisting of polyimide, parylene (poly-para-xylylen), a thick SU-8 photoresist, polydimethylsiloxane, and benzocyclobutene. 3. The flexible 3-dimensional microprobe structure according to claim 1, wherein said base is made of a second flexible polymeric material. 4. The flexible 3-dimensional microprobe structure according to claim 3, wherein said second flexible polymeric material is selected from a group consisting of polyimide, parylene (poly-para-xylylen), a thick SU-8 photoresist, polydimethylsiloxane, and benzocyclobutene. 5. The flexible 3-dimensional microprobe structure according to claim 1, wherein said probe, said base and said hinge portion are made of an identical material. 6. The flexible 3-dimensional microprobe structure according to claim 1, wherein said base has a micro channel, and said glue flows through said micro channel to apply to said hinge portion. 7. The flexible 3-dimensional microprobe structure according to claim 1, wherein said glue is polyethylene. 8. The flexible 3-dimensional microprobe structure according to claim 1, wherein said probe has a needle point and a connection member located at one side which is far away from said needle point; said connection member connects said needle point and said hinge portion. 9. The flexible 3-dimensional microprobe structure according to claim 8, wherein said connection member has an auxiliary fixing element which can increase a contact area between said glue and said flexible 3-dimensional microprobe structure. 10. The flexible 3-dimensional microprobe structure according to claim 9, wherein said auxiliary fixing element is made of a flexible polymeric material selected from a group consisting of polyimide, parylene(poly-para-xylylen), a thick SU-8 photoresist, polydimethylsiloxane), and benzocyclobutene. 11. The flexible 3-dimensional microprobe structure according to claim 1, wherein said probe has a structure-reinforcing element which is used to strengthen said probe so as to puncture a detected object by said probe. 12. The flexible 3-dimensional microprobe structure according to claim 11, wherein said structure-reinforcing element is made of a thick SU-8 photoresist. | National Tsing Hua University | A61B005/04 A61N001/00
US07953472 | B2 | 2011-05-31 | Method and apparatus for treatment of discogenic pain | 2009-09-17 | A method of treating discogenic back pain in a patient includes the steps of acquiring an image of at least a portion of the patient's spine, locating the spinal disc causing the back pain, and identifying a treatment volume encompassing at least a portion of the spinal disc. The method further includes prescribing a radiation dose to be delivered to the treatment volume, providing a radiation therapy machine, positioning the patient in the radiation therapy machine, programming the radiation therapy machine to deliver the radiation dose to the treatment volume, and delivering the radiation dose to the treatment volume. | 1. Non-transitory computer readable media with computer-executable instructions embodied thereon that when executed by a computer control a device for treating back pain, the media comprising: instructions for locating an spinal disc and defining a treatment volume encompassing at least a portion of one or more nerves transmitting impulses that result in the sensation of the back pain based on an image;instructions for creating a treatment plan for delivering a radiation dose to the treatment volume to decrease or stop the transmission of the impulses by the one or more nerves within the treatment volume; andinstructions for configuring the device to deliver the radiation dose to the treatment volume. 2. The non-transitory computer readable media of claim 1, further comprising: instructions for identifying a radiosensitive area proximate to the treatment volume. 3. The non-transitory computer readable media of claim 2, further comprising: instructions for configuring the treatment plan to avoid the radiosensitive area when delivering the radiation dose. 4. The non-transitory computer readable media of claim 3, wherein the radiosensitive area is avoided by delivering the radiation dose through one or more treatment portals extending through non-radiosensitive areas to reach the treatment volume. 5. The non-transitory computer readable media of claim 1, wherein the device for treating back pain is a radiation therapy machine. 6. The non-transitory computer readable media of claim 1, wherein the device for treating back pain comprises a medical imaging device and further comprising instructions for acquiring the image. 7. The non-transitory computer readable media of claim 6, wherein the medical imaging device is at least one of a CT scanner, a MRI scanner, a PET scanner, or a SPECT scanner. 8. The non-transitory computer readable media of claim 1, wherein the treatment volume encompasses at least a portion of the annulus fibrosus of the spinal disc. 9. The non-transitory computer readable media of claim 1, wherein the treatment volume encompasses at least a portion of the posterior wall of the spinal disc. 10. A discogenic back pain treatment device, comprising: a patient support;a radiation source;a radiation delivery mechanism coupled to the radiation source; anda treatment planning computer adapted to control the radiation delivery mechanism, the treatment planning computer programmed to deliver a radiation dose from the radiation source to a treatment volume by executing program instructions, the program instructions comprising: instructions for locating a spinal disc and defining the treatment volume to encompass at least a portion of the spinal disc;instructions for creating a back pain treatment plan for delivering the radiation to the treatment volume encompassing at least a portion of one or more nerves transmitting impulses that result in the sensation of the back pain based on an image; andinstructions for configuring the radiat... | Spinal Generations, LLC | A61B005/05 A61N005/10
US07951092 | B2 | 2011-05-31 | Guidewire loader apparatus and method | 2006-05-05 | A guidewire loader apparatus is provided for loading a guidewire into a lumen opening of a guidewire lumen of a medical device. The loader apparatus includes a first member defining a respective medical device receiving portion and a respective guidewire receiving portion in communication therewith via a respective interface portion. The loader apparatus further includes a second member configured for placement in an opposed relationship to the first member. The second member defines a respective medical device receiving portion and a respective guidewire receiving portion in communication therewith at a respective interface portion. The first member and the second member cooperatively associate with one another between a first position and a second position. In the first position, the medical device can be positioned into at least one respective medical device receiving portion. In the second position, the respective device receiving portions cooperate to align and support the medical device in a manner placing the lumen opening of the medical device substantially adjacent to the cooperating, respective interface portions. This permits aligned sliding receipt of an end of the guidewire into the cooperating, respective guidewire receiving portions, through the cooperating respective interface portions and into the lumen opening of the medical device. | 1. A guidewire loader apparatus for loading one or more guidewires into a plurality of guidewire lumens of a catheter comprising: a first member defining a respective catheter receiving portion having a first lumen diameter and a plurality of respective non-coaxial guidewire receiving portions in communication therewith at a respective interface portion having a lumen diameter substantially identical to the first lumen diameter; anda second member configured for placement in an opposed relationship to said first member, said second member defining a respective catheter receiving portion having lumen openings and a plurality of respective non-coaxial guidewire receiving portions in communication therewith at a respective interface portion, said first member and said second member cooperatively clamped together by mirror image clips on opposite sides of the first and second members, between a first position, enabling positioning of a catheter into the catheter receiving portion, and a second position, wherein the catheter receiving portion cooperates to align and support said catheter in a manner placing the lumen openings thereof substantially adjacent the cooperating, respective interface portions to permit aligned sliding receipt of respective ends of the guidewire into the cooperating, respective guidewire receiving portions, through the cooperating, respective interface portions and into the lumen openings of the catheter. 2. The guidewire loader apparatus according to claim 1, wherein each respective guidewire receiving portion is semi-conical-shaped, tapering inwardly toward the respective interface portion. 3. The guidewire loader apparatus according to claim 2, wherein each respective catheter receiving portion is semi-cylindrical-shaped. 4. The guidewire loader apparatus according to claim 1, wherein each the first member and the second member includes a respective base portion, configured for placement in an opposed relationship to one another, each respective base portion defining the respective catheter receiving portion, the guidewire receiving portion and the interface portion, and opposed, mirror image, respective manipulating tabs to enable manual operation between the first position and the second position. 5. The guidewire loader apparatus according to claim 4, wherein the respective manipulating tab of said first member extends from said respective base portion in a direction generally parallel to a longitudinal axis of the respective catheter receiving portion. 6. The guidewire loader apparatus according to claim 4, wherein, each respective manipulating tab is configured for placement in opposed relationship to one another on opposite sides of the first and second members. 7. The guidewire loader apparatus according to claim 4, wherein said manipulating tab of the second member includes a pair of opposed clip members hingeably mounted to the respective base portion on opposed sides of a longitudinal axis of the respective... | Abbott Cardiovascular Systems Inc. | A61B005/00 A61B006/00 A61M025/00 A61M025/16 A61M025/18 A61M039/00 A61M039/10
US07942833 | B2 | 2011-05-17 | Motion guide device, and its control system and control program | 2006-07-12 | A device for inducing a user's motion in a rhythm responsive to a motion rhythm after the user's motion rhythm changes suddenly. The walking assist device generates a first oscillator which attains mutual entrainment with a user's hip joint angular velocity to reflect a natural angular velocity. On the other hand, a new natural angular velocity is set based on the phase difference between the hip joint angular velocity and the first oscillator. A second oscillator which oscillates in rhythm reflecting the natural angular velocity is generated. An inducing oscillator is generated based on the second oscillator, and a torque responsive to the inducing oscillator acts on the user's body. If the magnitude of a periodic variation in the hip joint angular velocity exceeds a threshold value, a second oscillator which oscillates in rhythm reflecting the angular velocity of the first oscillator instead of the natural angular velocity is generated. | 1. A motion inducing device for inducing a user's motion according to an inducing oscillator, comprising: a motion oscillator measurement means which measures a user's motion oscillator;a first oscillator generation means which generates a first oscillator which attains mutual entrainment with the motion oscillator measured by the motion oscillator measurement means by using a first model defined by a natural angular velocity in such a way as to reflect the natural angular velocity to an angular velocity of the first oscillator;a natural angular velocity setting means which modifies the natural angular velocity based on a phase difference between the motion oscillator measured by the motion oscillator measurement means and the first oscillator generated by the first oscillator generation means;a second oscillator generation means which generates a second oscillator oscillating in rhythm reflecting the modified natural angular velocity set by the natural angular velocity setting means; andan inducing oscillator generation means which generates the inducing oscillator based on the second oscillator generated by the second oscillator generation means,wherein the second oscillator generation means generates the second oscillator which oscillates in rhythm reflecting one or both of an angular velocity of the motion oscillator measured by the motion oscillator measurement means and the angular velocity of the first oscillator generated by the first oscillator generation means, instead of the modified natural angular velocity set by the natural angular velocity setting means, when a magnitude of periodic variation in the motion oscillator measured by the motion oscillator measurement means exceeds a threshold value. 2. The motion inducing device according to claim 1, further comprising: a state measurement means which measures one or both of user's motion and physiological states; anda threshold value setting means which sets the threshold value according to one or both of the user's motion and physiological states determined by the state measurement means. 3. The motion inducing device according to claim 1, further comprising: a state measurement means which measures one or both of user's motion and physiological states; anda weight setting means which sets a weight of the angular velocity of the motion oscillator and a weight of the angular velocity of the first oscillator according to one or both of the user's motion and physiological states measured by the state measurement means,wherein the second oscillator generation means generates a second oscillator which oscillates in rhythm reflecting a weighted mean angular velocity, which is set by the weight setting means, between the angular velocity of the motion oscillator and the angular velocity of the first oscillator. 4. The motion inducing device according to claim 1, wherein the inducing oscillator generation means generates an inducing oscillator including the second oscillator generated by t... | Honda Motor Co., Ltd. | H03L007/00 G01N027/00 A61F005/00 A61B005/103 A61B005/117 A61H003/00 A61H001/00 A61H001/02 A63B024/00 A61N001/00
US07955275 | B2 | 2011-06-07 | Laparoscopic instrument and method for distance measurements of body parts | 2007-08-28 | A laparoscopic instrument includes an elongated barrel portion which mountably extends between operable parts for engagement with internal structures and organs of a subject patient, which are disposed at a first end of the barrel portion, and an actuating mechanism disposed at a second end of the barrel portion which remains outside the patient and which allows actuation of the instrument by the surgeon, for example by hand manipulation. The operable parts include two jaw parts which collectively define a clamp in which an opening size between the two jaw parts is selectively alterable by operation of the actuating mechanism. The operable parts further include a distance displacement sensor to effect distance measurement when the body part of the patient is received between the jaw parts and the instrument moved there along. | 1. A laparoscopic instrument for distance measurement of an elongated dimension of an internal body part of a patient, comprising: an elongated barrel portion;first and second jaw parts being carried at a first end of said barrel portion and extending from said first end codirectionally with a jaw axis, at least one of said jaw parts being movable with respect to said barrel portion to effect an increase and decrease of a jaw space between said first and second jaw parts;an actuating mechanism carried at a position along the barrel portion spaced apart from said first end for controlling movement of said at least one of said jaw parts by movement thereof; anda distance displacement sensor being disposed at said first end of said barrel portion, said distance displacement sensor being configured so as to be operable for sensing a distance of displacement of the first and second jaw parts along at least a length portion of the internal body part when a portion of the body part being measured is at least partially received in said jaw space and the instrument moved so as to effect travel of the first and second jaw parts in a direction crosswise to said jaw axis along said at least a length portion of the internal body part. 2. A laparoscopic instrument according to claim 1, further comprising an actuating coupling extending between said actuating mechanism and said at least one of said jaw parts, said movement of said actuating mechanism being converted into axial displacement of said actuating coupling, said axial displacement being converted into controlled movement of said at least one of said jaw parts effecting the increase and decrease of said jaw space. 3. A laparoscopic instrument according to claim 2, wherein said actuating mechanism includes a hand grip and an actuating trigger pivotably movable relative to said hand grip for converting pivotable movement thereof into said axial displacement of the actuating coupling. 4. A laparoscopic instrument according to claim 2, wherein said actuating mechanism includes a rotatable knob linked with said actuating coupling in a manner in which rotation thereof is converted into said axial displacement of the actuating coupling. 5. A laparoscopic instrument according to claim 1, wherein said distance displacement sensor includes a rotatable part which is rollably contactable with the body part and positioned such that said rotatable part is made to rotate in response to movement of the instrument along the body part when the body part is at least partially received in said jaw space. 6. A laparoscopic instrument according to claim 5, wherein said rotatable part is rotatably mounted in an orientation facing interior of a clamp opening bounded by inner confronting surfaces of the first and second jaw parts and the first end of the barrel portion. 7. A laparoscopic instrument according to claim 6, wherein said rotatable part is located between two ends of the first jaw, arranged to face a side of t... | - | A61B005/103 A61B005/117
US09146289 | B2 | 2015-09-29 | Targeted thermal treatment of human tissue through respiratory cycles using ARMA modeling | 2009-12-23 | The present application discloses a technique for targeting therapeutic thermal energy to human tissue that is subject to displacement during a respiratory cycle using ARMA modeling. It discloses using an ARMA treatment of MRI tracking data of salient features of the tissue of interest to predict the spacial position of the portion of the tissue to be treated and using this prediction to guide the application of the thermal energy. It also discloses that this technique is particularly useful when the tissue of interest undergoes elastic deformation in a respiratory cycle and high energy focused ultrasound (HIFU) is used to ablate diseased tissue such as a cancerous tumor. | 1. A process for an in-vivo targeted thermal treatment of biological tissue of a subject which undergoes elastic deformation during a respiratory cycle comprising: creating a model which predicts the motion of said tissue during a respiratory cycle by an elastic registration and an autoregressive moving average (ARMA) treatment of a training set of data, wherein the training set of data is created by tracking spatial positions of one or more salient features representative of different portions of said tissue with magnetic resonance imaging (MRI) over multiple respiratory cycles of one or more subjects;using said model to target heat energy to a desired portion of said tissue over multiple respiratory cycles using said model;generating MRI thermal maps of said tissue and monitoring the thermal treatment with the MRI thermal maps;tracking actual positions of said salient features over one or more respiratory cycles of the subject;determining a change in elastic properties of said tissue of the subject at one or more determined intervals of time based on a difference between the tracked positions and positions of said salient features predicted by said model;updating the model to reflect the change in elastic properties of said tissue if the difference between the tracked positions and predicted positions is less than a given threshold and continue using the updated model to target heat energy to the desired portion of said tissue; anddiscontinuing use of the model and interspersing tracking and thermal mapping of the salient features of said tissue over subsequent respiration cycles if the difference between the tracked positions and predicted positions exceeds said threshold. 2. The process of claim 1 wherein said model includes a registration of the positions of said salient features over a projected respiratory cycle. 3. The process of claim 1 wherein said model includes a factor to account for the change in the elastic properties of the portion of said tissue subjected to said thermal treatment as a result of said treatment. 4. The process of claim 1 wherein said tissue is the liver. 5. The process of claim 1 wherein said thermal energy is provided by high intensity focused ultrasound (HIFU). 6. The process of claim 1 wherein the MRI thermal maps of said tissue are generated without interruption during at least about 3 minutes of said thermal treatment. 7. The process of claim 1 wherein the MRI thermal maps of said tissue are generated without interruption over multiple respiratory cycles during said thermal treatment. 8. The process of claim 7 wherein said model is validated by interrupting said MRI thermal mapping after multiple respiratory cycles, tracking the actual positions of said salient features of said tissue being subjected to thermal treatment over one or more respiratory cycles and comparing said positions to those predicted by said model. 9. The process of claim 8 wherein said validation is only performed after a pluralit... | General Electric Company | A61B005/05 G01R033/48 A61N007/02 A61B005/055 A61B005/08 A61B005/113 A61B017/00 A61B019/00
US09138155 | B2 | 2015-09-22 | Method and a device for examination of physical magnitude in humans or animals in an object filled with liquid or gas | 2009-05-15 | The invention relates to a device for examination of a pressure (p) in an intervertebral disc, a body of vertebra or any other part of a disc in a human or animal. The device (1) comprises a hollow cannula (6) with an end portion (7) arranged to be inserted into the intervertebral disc, the body of vertebra or any other part of the disc. The device further comprises an optical fiber (8) arranged in the cannula (6) and a sensor device (10) arranged on the optical fiber (8), which in use is positioned in connection to the intervertebral disc, the body of vertebra or any other part of the disc. The device also has an injection device (11) for injection of liquid into the intervertebral disc, the body of vertebra or any other part of the disc through the cannula (6) while the sensor device (10) measures the physical magnitude (p) in the anatomical organ (2). The invention also relates to a method for examination of a pressure (p) in an intervertebral disc, a body of vertebra or any other part of a disc in humans and animals, designed in analogy to the device according to the invention. | 1. A diagnostic method for assessing the medical condition of an object filled with liquid or gas in humans and animals based on a physical parameter (p) measured in the object, the method comprising: inserting a tube-shaped, hollow cannula, and positioning an end portion of the cannula, inside the object;introducing a sensor device into the cannula;positioning the sensor device relative to the object;injecting a liquid through the cannula into the object for a predetermined time period while measuring the physical parameter (p) inside the object using the sensor device; the physical parameter (p) being chosen from the group consisting of: (a) pressure; and (b) temperature;comparing a profile of measured values of the physical parameter (p) to a profile of predetermined values of the physical parameter (p); the profile of measured values of the physical parameter (p) representing a plot of the measured values of the physical parameter (p) against time values of the predetermined time period; the profile of predetermined values representing a plot of the predetermined values of the physical parameter (p) against the time values of the predetermined time period; each predetermined value of the physical parameter (p) and associated time value corresponding to a previously defined medical condition of the object;determining the medical condition of the object based on the comparison of the profile of measured values to the profile of the predetermined values. 2. The method of claim 1 wherein the step of injecting includes continuously injecting the liquid through the cannula into the object while continuously measuring the physical parameter (p) inside the object using the sensor device. 3. The method of claim 1 wherein the object is selected from the group consisting of: (a) an anatomical organ; (b) an intervertebral disc; (c) a vertebra; and (d) a tumour. 4. The method of claim 1 wherein the liquid is selected from the group consisting of: (a) a contrast medium; (b) a bone cement; and (c) a pharmaceutical. 5. The method of claim 1 further comprising, following the step of injecting the liquid, determining one of propagation of the liquid in the object and leakage of the liquid from the object, using x-rays. 6. The method of claim 1 further comprising the step of locating the sensor device within the object using x-rays. 7. The method of claim 1 wherein the profile of the predetermined values represents a profile of values for a healthy object. 8. The method of claim 1 wherein the sensor device includes an optical sensor. 9. The method of claim 8 wherein the optical sensor is a Fabry-Pérot optical sensor. 10. The method of claim 8 wherein: the physical parameter (p) is pressure; andthe step of measuring the physical parameter (p) inside the object using the sensor device includes measuring the pressure inside the object based on optical interference. 11. The method of claim 1 wherein the step of positioning the sensor device relative to th... | FISO Technologies Inc. | A61B005/00 A61F002/44 A61B005/03 A61F002/46
US08409108 | B2 | 2013-04-02 | Multi-axial heart sounds and murmur detection for hemodynamic-condition assessment | 2010-11-05 | A vector method for monitoring a subject's hemodynamic condition including (a) utilizing at least one, external or internal, anatomy-attached, three-axis accelerometer, collecting from the subject, during a selected cardiac cycle, related, three-orthogonal-axes accelerometer signal data, (b) following such collecting, processing collected signal data to obtain associated, signal vector, magnitude and directionality information, and (c) analyzing such obtained vector information for assessment of the subject's heart hemodynamic condition. ECG and signal time-frequency data is also collected and used in certain manners and implementations of the invention. | 1. A plural-external-anatomical site vector method for monitoring, exclusively externally, a subject's hemodynamic condition utilizing, and based upon, exclusively externally collected data relating to murmur and heart sound phenomena, detected by, and collected exclusive from, plural, externally anatomically-attached, stably positioned, three-orthogonal-axis accelerometers, said method comprising attaching at least a pair of three-orthogonal-axis accelerometers in stable, spaced positions externally to a subject's anatomy,during at least one cardiac cycle of the subject, collecting signal data collaboratively from the so-attached, plural accelerometers,processing such collaboratively collected data to obtain vector magnitude and directionality information associated, respectively, with the least two accelerometers, andanalyzing such collaboratively collected and processed vector magnitude and directionality information for assessment of the subject's hemodynamic condition. 2. The method of claim 1, wherein , respectinq collaborative treatment associated with the at least two accelerometers, said processing involves examining, including frequency filtering, the frequency content of collected data (1) to distinguish heart sound content from murmur content, and (2) to develop vector magnitude and directionality information respectively related to each of these two phenomena, and said analyzing employs all of such respectively developed information. 3. The method of claim 1, wherein said analyzing involves using such obtained vector information to identify the presence of at least one of (a) a specific heart sound component, and (b) murmur. 4. The method of claim 1, wherein said analyzing involves using such obtained vector information to identify the presence of at least one of (a) a specific heart sound component and its subcomponents, and (b) murmur. 5. The method of claim 1 which further comprises performing said collecting, processing, and analyzing steps over a selected time interval involving plural cardiac cycles. 6. The method of claim 1 which further comprises obtaining from the attached, plural accelerometers signal time-frequency information, and said processing and analyzing include relating this time-frequency information to associated, signal vector, magnitude and directionality information. 7. The method of claim 1 which further comprises, during said collecting, acquiring time-related ECG data. 8. The method of claim 7, wherein said analyzing involves using such obtained vector information to identify the presence of at least one of (a) a specific heart sound component, and (b) murmur. 9. The method of claim 7, wherein said analyzing involves using such obtained vector information to identify the presence of at least one of (a) a specific heart sound component and it subcomponents, and (b) murmur. 10. A computer-based system for plural, external-anatomical-site vector monitoring, externally, of a subject's hemodynamic condi... | Inovise Medical, Inc. | A61B005/02
US08886224 | B2 | 2014-11-11 | Methods for providing enhanced telecommunication services | 2011-02-11 | A first method for facilitating personal contacts among a plurality of subscribers to a location-based person contact service provided by a service provider; and a second method for real-time monitoring of vital signs of a living being. In the first method, personal profiles are transmitted to subscribers. In the second method, information relating to a violated vital sign parameter determined from the monitoring is communicated to a local communication node. | 1. A method for facilitating personal contacts among a plurality of subscribers to a location-based person contact service provided by a service provider having a database that comprises a personal profile of each subscriber and filter criteria for accepting the personal profiles of the other subscribers, said method comprising: receiving, by the service provider, a first message from a first wireless device possessed by a first subscriber of the plurality of subscribers, wherein the first message includes data from which a region may be determined by the service provider, and wherein the message further includes a request for the first wireless device to receive those personal profiles of all subscribers whose wireless devices are determined to be within the region satisfying filter criteria specified by the first subscriber;extracting from the database, by the service provider, said those personal profiles;transmitting, by the service provider to the first wireless device after said receiving the first message, said those personal profiles;receiving, by the service provider after said transmitting, a second message from the first wireless device, wherein the second message instructs the service provider to forward the personal profile of the first subscriber to a second wireless device possessed by a second subscriber of said all subscribers whose wireless devices are determined to be within the region who satisfy the filter criteria specified by the first subscriber; andsending, by the service provider to the second wireless device after said receiving the second message, the personal profile of the first subscriber. 2. The method of claim 1, wherein the data comprises an explicit identification of the region on a geocoded digital map domain. 3. The method of claim 2, wherein the data comprises an identification of a geographic position of the first wireless device and a specified distance from the geographic position, and wherein the region is within the specified distance from the geographic position. 4. The method of claim 3, wherein the method further comprises: receiving, by the service provider prior to said transmitting, the identification of the geographic position from a global positioning system (GPS). 5. A system comprising a providing entity adapted to perform the method of claim 1, said providing entity being the service provider. 6. A method comprising: receiving, from a first wireless device, by a computing device associated with a service provider of a service including a plurality of subscribers, information from which a region may be determined and an indication of a filter criterion for accepting personal profiles of the subscribers, wherein the information indicates a geographic position at which the first wireless device is indicated to be currently located; andin response to receiving the information, sending, by the computing device associated with the service provider, to the first wireless device, one or more pers... | Amazon Technologies, Inc. | H04W024/00 G06Q030/06 H04L029/08 H04W004/02 G06Q030/02
US08491498 | B2 | 2013-07-23 | Sample collection implement | 2009-08-27 | A sample collection implement S1 includes a container 1 having an accommodation portion 10 in which a liquid 5 for suspending or diluting a sample is accommodated, a sample collection stick 2 being able to be disposed in the accommodation portion 10, a filter 3 provided inside the container 1, and a movable member 4 that can be moved in a predetermined direction inside the container 1 and has a function of pushing the liquid 5 and causing the liquid to pass through the filter 3, when moved in the predetermined direction. With such a configuration, when the liquid 5 is filtered, it is not necessary to move the filter 3 by directly pushing it and a filter with a low mechanical strength can be used as the filter 3. In addition, it is not necessary to use a centrifugal apparatus. | 1. A sample collection implement comprising: a container having an accommodation portion in which a liquid for suspending or diluting a sample is accommodated, a sample collection stick being able to be disposed in the accommodation portion, anda filter provided inside the container,the sample collection implement further comprisinga movable member that can be moved in a predetermined direction inside the container and has a function of pushing the liquid and causing the liquid to pass through the filter, when moved in the predetermined direction, whereinthe movable member is provided separately from the filter and can be moved by being pushed from outside the container, andadjacent first and second regions partitioned by a partition wall are formed inside the container,the filter is disposed in the first region, andwhen the liquid passes through the filter by the movement of the movable member, at least part of the liquid that has passed through the filter flows into the second region over the partition wall, and is retained in the second region. 2. The sample collection implement according to claim 1, wherein the movable member is slidably inserted into the second region and can move down when pushed from above to push the liquid so as to cause the liquid to pass through the filter, andwhen the liquid that has passed through the filter flows into the second region, the liquid can be retained on the movable member. 3. The sample collection implement according to claim 1, wherein the movable member is provided in a location different from the second region,the second region has a form such that a bottom portion thereof is closed by part of the container or by a fixing member separate from the container, andwhen the liquid that has passed through the filter flows into the second region, the liquid can be retained on the bottom portion. 4. A sample collection implement comprising: a container having an accommodation portion in which a liquid for suspending or diluting a sample is accommodated, a sample collection stick being able to be disposed in the accommodation portion, anda filter provided inside the container,the sample collection implement further comprisinga knob,a movable member that exists within the container that can be moved in a predetermined direction inside the container and has a function of pushing the liquid and causing the liquid to pass through the filter, when moved in the predetermined direction, further comprisinga linking means capable of linking the sample collection stick to the movable member in a state in which the sample collection stick is inserted into the container, whereinthe filter is held by the movable member,the sample collecting stick is not linked to the movable member when the sample collecting stick is outside of the container, is used to collect a sample in this condition, and is linked to the moveable member through the linking means when the sample collection stick is inserted into the accommodation p... | ARKRAY, Inc. | A61B010/00 A61B005/00 B01D035/00 B01D041/00
US08412303 | B2 | 2013-04-02 | Electrode for use with a template cap | 2010-10-29 | An electrode attached to the straps of and use with a template cap in making neurophysiological measurements. The electrode is configured to allow the user to quickly and accurately slide an electrode needle at a shallow adjustable angle into the scalp of the patient. The base of each electrode attaches to the cap at locations where measurements are to be made. The base supports a ramp that may be snapped into a clip on the base for storage and then springs resiliently from the clip when released. The electrode needle is carried by a slidable holder secured to the ramp at its slot and is slid forward following that slot when driving the needle down and through a small hole in the base of the electrode to insert it into the patient's scalp at the appropriate depth and angle. | 1. An electrode for neurophysiological measurements and for use with a template cap, said electrode comprising: (a) a base having a proximal end and a distal end, said base having a slot formed therein;(b) a ramp having a proximal end and a distal end, said distal end of said ramp being attached near said distal end of said base, said ramp having a slot running between said proximal end of said ramp and said distal end of said ramp;(c) a holder carried by said ramp in said slot of said ramp, said holder having a needle housing portion and a thumb drive portion; and(d) a needle having a proximal end and a distal end, said proximal end carried by said needle housing portion and said needle can be moved by pressing on said thumb drive portion and pushing said thumb drive portion from said proximal end to said distal end of said slot of said ramp thereby causing said distal end of said needle to pass from said base through said slot in said base. 2. The electrode of claim 1, said distal end of said ramp carrying a tab, and wherein said electrode further comprises a clip carried by said proximal end of said base, said clip having a slot dimensioned to receive said tab, wherein, when said tab is in said slot, said electrode is in a stored position, and wherein, when said tab is not in said slot of said clip and said needle is at said proximal end of said slot of said ramp, said electrode is in an insertion-ready position, and wherein, when said tab is not in said slot of said clip and said needle is at said distal end of said slot of said ramp, said electrode is in a needle-inserted position. 3. The electrode of claim 2, wherein said needle cannot be moved to said needle inserted position when said tab is in said slot of said clip. 4. The electrode of claim 1, wherein said slot of said ramp is dimensioned to limit insertion of said needle through said slot of said base. 5. The electrode of claim 1, wherein said ramp and said base are made to be at an angle, said angle being pre-determined to be the angle of insertion for said needle. 6. The electrode of claim 1, wherein said slot of said ramp has a long dimension running from said proximal end of said slot to said distal end of said slot, and wherein said holder holds said needle parallel to said long dimension. 7. The electrode of claim 1, wherein said holder holds said needle between said ramp and said base. 8. The electrode of claim 1, wherein said holder has said needle housing portion on one side of said slot of said ramp and said thumb drive portion on the other side of said slot of said ramp. 9. The electrode of claim 1, wherein said base is curved. 10. The electrode of claim 1, wherein said proximal and distal ends of said base have holes formed therein for securing said base to straps of a template cap. 11. A template cap with electrodes for neurophysiological measurements, said template cap comprising: (a) straps secured together to form a cap for use on the head of a patient to und... | Rhythmlink International, LLC | A61B005/0478
US08515548 | B2 | 2013-08-20 | Article of clothing including bio-medical units | 2011-02-17 | An article of clothing includes a clothing fabric and a plurality of bio-medical units integrated into the clothing fabric. A bio-medical includes a power harvesting module, a communication module, a processing module, a functional module, a die, and an IC package. The die supports the power harvesting module, the processing module, the communication module, and the functional module. The IC package houses the die and includes a mechanism for adhering to the clothing fabric. | 1. An integrated circuit (IC) for incorporation in a clothing fabric, the IC comprises: a power harvesting module configured to convert at least one of body heat, body motion, and electromagnetic signal, light, and radio frequency (RF) signals into a supply voltage;a communication module configured, when powered by the supply voltage, to: convert an inbound wireless signal into an inbound symbol stream;convert an outbound symbol stream into an outbound wireless signal;a processing module configured, when powered by the supply voltage, to: convert the inbound symbol stream into a bio-medical function;convert a bio-medical response into the outbound symbol stream;a functional module configured, when powered by the supply voltage, to: perform the bio-medical function; andwhen the bio-medical function is a monitoring function, generate the bio-medical response;at least one die supporting the power harvesting module, the processing module, the communication module, and the functional module; andan IC package housing the at least one die, wherein the IC package includes a clothing fabric adhering mechanism. 2. The IC of claim 1, wherein the clothing fabric adhering mechanism comprises at least of: an eyelet for facilitating sewing the IC into the clothing fabric;a hook for facilitating sewing the IC into the clothing fabric;a notch for facilitating sewing the IC into the clothing fabric; anda fabric adhesive for facilitating gluing the IC into the clothing fabric. 3. The IC of claim 1 further comprises: an encapsulant for encapsulating the IC package. 4. The IC of claim 1, wherein the bio-medical function further comprises at least one of: an image capture function;a movement capture function;a sound capture function;a topical treatment function; andan electronic stimulation function. 5. The IC of claim 1, wherein the power harvesting module comprises at least one of: an electromagnetic signal to voltage conversion module;an RF signal to voltage conversion module;a motion to voltage conversion module;a light to voltage conversion module; anda heat to voltage conversion module. 6. The IC of claim 5, wherein the electromagnetic signal to voltage conversion module comprises at least one of: an array of inductors and a rectifying circuit module to produce the supply voltage from the electromagnetic signal; andan array of Hall-effect devices and the rectifying circuit module to produce the supply voltage from the electromagnetic signal. 7. An article of clothing comprises: a clothing fabric; anda plurality of bio-medical units integrated into the clothing fabric, wherein a bio-medical unit of the plurality of bio-medical units includes: a power harvesting module configured to convert at least one of body heat, body motion, and electromagnetic light, and radio frequency (RF) signals into a supply voltage;a communication module configured, when powered by the supply voltage, to: convert an inbound wireless signal into an inbound symbol stream; andconver... | Broadcom Corporation | A61N001/00 A61N001/39
US08529470 | B2 | 2013-09-10 | Lancets for bodily fluid sampling supplied on a tape | 2011-02-15 | A supply of lancets for a multi-use lancing device are carried by a tape and sequentially brought from a storage position to an activating position by advancing the tape around a bend. The lancets are non-circular in cross-section along their longitudinal lengths, and in their activating position they extend from the tape such that their sharp tips are available for lancing tissue. The carrying tape can be arranged in reel-to-reel format in a housing having a lancing opening. The lancets can be integral with the tape and activated to move the lancet through the lancing opening by moving the tape along its tape path or by translating a dedicated service loop of the tape. The lancets can also be independent from the tape and activated through the lancing opening by a separate lancing actuator. A test media can be included on the carrying tape and the housing can contain a sensor to yield an integrated lancing and testing device. | 1. A multi-use lancing device comprising: a housing having a lancet opening and containing a supply reel, a storage reel, and a take-up reel, wherein the supply reel, the storage reel, and the take-up reel are each configured to rotate independently of the housing;an elongated flexible carrying tape configured for reel-to-reel transfer from the supply reel to the storage reel;a plurality of lancets sequentially positioned along the carrying tape and each having a lancet length and a distal portion defining a sharp tip; anda sterile removable cover over the lancets in the supply reel,wherein rotation of the take-up reel removes the cover from the lancets as the lancets are advanced from the supply reel to the storage reel. 2. The device of claim 1 wherein a substantial portion of the lancet length is high-aspect ratio in cross section. 3. The device of claim 1 wherein the lancets are composed of a hardened stainless steel or a shape memory alloy. 4. The device of claim 1 further comprising a cover application reel configured to place the cover on the lancets being advanced onto the storage reel. 5. The device of claim 1 further comprising a plurality of test media carried on the tape, at least one test media associated with each of the plurality of lancets. 6. The device of claim 1 wherein the lancets include a capillary groove for conveying a bodily fluid sample to an associated test media. 7. A multi-use lancing device comprising: a housing having a lancet opening and containing a supply reel and a storage reel;an elongated flexible carrying tape configured for reel-to-reel transfer from the supply reel to the storage reel;a plurality of lancets sequentially positioned along the carrying tape and each having a lancet length and a distal portion defining a sharp tip;a sterile removable cover over the lancets in the supply reel;a take-up reel for removing the cover from the lancets as the lancets are advanced from the supply reel to the storage reel; anda cover application reel for placing the cover on the lancets being advanced onto the storage reel. 8. A multi-use lancing device comprising: a housing having a lancet opening and containing a supply reel and a storage reel;an elongated flexible carrying tape configured for reel-to-reel transfer from the supply reel to the storage reel;a plurality of lancets sequentially positioned along the carrying tape and each having a lancet length and a distal portion defining a sharp tip;a sterile removable cover over the lancets in the supply reel; anda take-up reel for removing the cover from the lancets as the lancets are advanced from the supply reel to the storage reel, the take-up reel is rotatable independent of the housing; anda cover application reel for placing the cover over the lancets, the cover application reel is rotatable independent of the housing. 9. The device of claim 8 wherein the plurality of lancets and the carrying tape are formed from a single piece of tape stock. 10. The de... | Roche Diagnostics Operations, Inc. | A61B005/00
US20110105928 | A1 | 2011-05-05 | ECG Reconstruction For Atrial Activity Monitoring And Detection | 2009-11-05 | A system includes a mobile unit having a plurality of electrodes, numbering less than ten, that are configured to contact a patient to obtain electrical signals therefrom, and a diagnostic center disposed remotely from the mobile unit. The mobile unit and/or the diagnostic center are configured to construct a first portion of an ECG (electrocardiogram) corresponding to a first portion of a cardiac cycle of the patient by processing information based on received electrical signals using a first set of transformation parameters corresponding to the first portion of the cardiac cycle. The first portion may correspond to atrial or ventricular activity. | 1. A method comprising: using a mobile unit to contact a patient with a plurality of electrodes, numbering less than ten;obtaining an electrical signal from each electrode; andconstructing a first portion of an ECG (electrocardiogram) corresponding to a first portion of a cardiac cycle of the patient by processing information based on the electrical signals using a first set of transformation parameters corresponding to the first portion of the cardiac cycle. 2. The method of claim 1, wherein the constructed first portion of the ECG corresponds to ventricular activity. 3. The method of claim 1, wherein the constructed first portion of the ECG corresponds to atrial activity. 4. The method of claim 1, further comprising constructing a second portion of the ECG corresponding to a second portion of the cardiac cycle of the patient by processing information based on the electrical signals using a second set of transformation parameters corresponding to the second portion of the cardiac cycle. 5. The method of claim 4, wherein the constructed first portion of the ECG corresponds to ventricular activity and the constructed second portion of the ECG corresponds to atrial activity. 6. The method of claim 1, further comprising monitoring the constructed first portion of the ECG for atrial fibrillation. 7. The method of claim 1, further comprising monitoring the constructed first portion of the ECG for atrial flutter. 8. The method of claim 6, wherein said monitoring is conducted automatically. 9. The method of claim 6, further comprising over-writing and/or erasing the constructed first portion of the ECG if no atrial fibrillation is detected. 10. The method of claim 1, wherein the plurality of electrodes is less than 6 electrodes. 11. The method of claim 1, wherein said constructing is conducted at least partially by a remotely disposed device to which information based on the electrical signals is transmitted. 12. The method of claim 1, wherein said transmission is conducted in real-time. 13. The method of claim 4, further comprising combining the constructed first and second portions of the ECG into a combined ECG. 14. The method of claim 13, further comprising monitoring the combined ECG for atrial fibrillation. 15. A mobile unit comprising: a plurality of electrodes, numbering less than ten, configured to contact a patient to obtain electrical signals therefrom; anda controller configured to construct a first portion of an ECG (electrocardiogram) corresponding to a first portion of a cardiac cycle of the patient by processing information based on the electrical signals using a first set of transformation parameters corresponding to the first portion of the cardiac cycle. 16. The mobile unit of claim 15, wherein the constructed first portion of the ECG corresponds to ventricular activity. 17. The mobile unit of claim 15, wherein the constructed first portion of the ECG corresponds to atrial activity. 18. The mobile unit of claim 15, wh... | NewCardio, Inc. | A61B005/0402
US20110130673 | A1 | 2011-06-02 | Apparatus for determining respiratory muscle endurance of a person | 2004-07-12 | Apparatus (2) for determining respiratory muscle endurance of a person, which apparatus (2) comprises a mouthpiece (4) through which the person inspires, load-providing means (6) for providing a pressure against which the person inspires, and pressure control means (8) for controlling the pressure, the pressure control means (8) being such that it controls the pressure in response to a breathing pattern of the person. | 1. Apparatus for determining respiratory muscle endurance of a person, which apparatus comprises a mouthpiece through which the person inspires, load-providing means for providing a pressure against which the person inspires, and pressure control means for controlling the pressure, the pressure control means being such that it controls the pressure in response to a breathing pattern of the person. 2. Apparatus according to claim 1 in which the load-providing means is a rotary valve having an orifice which is variable in size. 3. Apparatus according to claim 2 in which the valve is a rotary valve. 4. Apparatus according to claim 3 in which the rotary valve is a servo-controlled valve. 5. Apparatus according to claim 4 in which the pressure control means comprises an electronic processor, and in which the processor receives a first input in the form of the pressure and a second input in the form of the flow. 6. Apparatus according to claim 5 in which the processor includes a display screen. 7. Apparatus according to claim 5 in which the processor includes a hard copy print-out means. 8. Apparatus according to claim 5 in which the processor is a microprocessor. 9. Apparatus according to claim 1 in which measurements are obtainable as measurements at the mouthpiece. | - | A61B005/08
US20110112396 | A1 | 2011-05-12 | SYSTEM AND METHOD FOR TARGETING CATHETER ELECTRODES | 2009-11-09 | A system and method is described for a catheter guidance system which allows an operator to use a mapping catheter to specify tissue target locations for the automatic guidance of a second therapeutic catheter. The operator places a mapping catheter at a desired location, and commands the catheter guidance system by either selecting a point on that catheter or one of the catheter electrode electrocardiograms. The operator may target the selected dynamic location, or tissue contact beyond that location on a specific side of the mapping catheter. | 1. A method for catheter guidance, comprising: introducing a mapping catheter into a patient;moving said mapping catheter within the patient to bring an electrode provided to said mapping catheter in contact with tissues at various locations to generate a desired map;identifying a target location from said map and placing an electrode of said mapping catheter proximate to said target location; andintroducing a therapeutic catheter into the patient and guiding said therapeutic catheter to a location of said electrode. 2. The method of claim 1, further comprising guiding said therapeutic catheter past said location of said electrode. 3. The method of claim 1, wherein said mapping catheter is guided using magnetic fields. 4. The method of claim 1, wherein said therapeutic catheter is guided using magnetic fields. 5. A system for catheter guidance, comprising: a mapping catheter having at least one electrode proximate to a distal end of said mapping catheter;therapeutic catheter;a first catheter guidance and control system that moves said mapping catheter to bring said at least one electrode in contact with tissues at various locations to generate a desired map; anda second catheter guidance and control system that moves a distal end portion of said therapeutic catheter to a location of said at least one electrode. 6. The system of claim 5, wherein said therapeutic catheter comprises an ablation catheter. 7. The system of claim 5, wherein said second catheter guidance and control system comprises a magnetic guidance and control system that uses a plurality of electromagnets to control a position of a magnet provided to a distal end of said therapeutic catheter. 8. The system of claim 5, wherein said first catheter guidance and control system comprises a magnetic guidance and control system that uses a plurality of electromagnets to control a position of a magnet provided to a distal end of said mapping catheter. 9. The system of claim 5, wherein said second catheter guidance and control system automatically moves a distal end portion of said therapeutic catheter to a location of said at least one electrode. | Magnetecs, Inc. | A61B005/05
US07967751 | B2 | 2011-06-28 | Ambulatory repeater for use in automated patient care | 2007-08-17 | An ambulatory repeater for use in automated patient care is presented. A local memory store includes a cryptographic key, sensitive information, and physiological measures. The cryptographic key is uniquely assigned to the implantable medical device prior to implant of the implantable medical device into a patient. The sensitive information is preencrypted under the cryptographic key and physiological measures are measured by the implantable medical device. An authentication module is in receipt of the cryptographic key. A permissions module confirms authorization of an external data processing device against the cryptographic key. A decryption module decrypts the sensitive information with the cryptographic key into decrypted information. A processor is operatively coupled to the local memory store. A communications module exchanges the decrypted information and the physiological measures with the external data processing device over a wireless interface contingent upon the authorization confirmation. An internal power supply supplies power to the foregoing components. | 1. A secure wireless ambulatory repeater system, comprising: an implantable medical device comprising a secure key repository for maintaining one or more cryptographic keys; anda secure wireless ambulatory repeater comprising: a memory for storing sensitive information preencrypted under a cryptographic key and physiological measures measured by the implantable medical device;an authentication module to securely obtain the cryptographic key from the secure key repository of the implantable medical device;a processor comprising: a decryption module to decrypt the sensitive information with the cryptographic key into decrypted information; anda communications module to exchange the decrypted information and the physiological measures with an external data processing device over a wireless interface contingent upon authorization of the external data processing device;wherein the ambulatory repeater is consists of one of a wearable or handheld form factor. 2. An ambulatory repeater according to claim 1, further comprising: a telemetry interface to the implantable medical device over which the cryptographic key, the sensitive information, and the physiological measures are securely received. 3. An ambulatory repeater according to claim 2, wherein the telemetry interface operates over at least one of inductive telemetry, radio frequency telemetry, and Bluetooth telemetry. 4. An ambulatory repeater according to claim 1, wherein the cryptographic key is maintained in a key escrow. 5. An ambulatory repeater according to claim 1, wherein the physiological measures are measured by a sensor comprised in the implantable medical device. 6. An ambulatory repeater according to claim 1, further comprising: an activation module to transition the ambulatory repeater from a standby mode; anda standby module to transition the ambulatory repeater from an active mode. 7. An ambulatory repeater according to claim 6, wherein the activation module is at least one of internally triggered by at least one of a schedule, an alert condition, and a manual override, and remotely triggered by the external data processing device. 8. An ambulatory repeater according to claim 1, wherein the physiological measures comprises one or more of physiological and behavioral data. 9. An ambulatory repeater according to claim 1, wherein the external data processing device is selected from the group comprising a server, personal computer, repeater, and programmer. 10. An ambulatory repeater system for use in automated patient care, comprising: an implantable medical device comprising a secure key repository for maintaining one or more cryptographic keys; andan ambulatory repeater comprising: a local memory store, comprising: a cryptographic key uniquely assigned to an implantable medical device; andsensitive information preencrypted under the cryptographic key and physiological measures measured by the implantable medical device;a processor operatively coupled to the local memory stor... | Cardiac Pacemakers, Inc. | A61B005/00
US07935064 | B2 | 2011-05-03 | Reducing or avoiding muscle cramps | 2008-07-31 | Muscle cramps can be reduced or avoided by monitoring the concentration of salt in perspiration to indicate the degree of salt depletion. Based on the degree of salt depletion, a sufficient amount of rehydration beverage is consumed to normalize the salt level. | 1. A method of limiting muscle cramping comprising: providing a packaging having a consumable composition comprising at least one salt and a table providing an estimate of a quantity of sweat secreted during physical activity based on a level of sweat and a weight of a user;providing with the packaging, a tester including a test portion configured to be exposed to body fluid of the user for use in obtaining a result indicative of an estimated degree of predicted muscle cramping from the user;providing a test chart associated with the packaging configured to be compared with the tester after exposure to the body fluid to provide a test result of an estimated degree of muscle cramping; andproviding dosage information regarding the consumable composition, the dosage information comprising an amount of the consumable composition to be consumed by the user based on the test result for salt replenishment and the estimate of a quantity of sweat secreted during physical activity, sufficient to substantially limit muscle cramping. 2. The method of claim 1, wherein the tester is configured to measure a concentration of components selected from the group consisting of: proteins, urea, ketones, lactate, magnesium, potassium and combinations thereof. 3. The method of claim 1, wherein the body fluid comprises perspiration. 4. The method of claim 1, wherein the consumable composition comprises a beverage. 5. The method of claim 1, wherein the consumable composition comprises a rehydration beverage concentrate and the dosage information comprises information in preparing a rehydration beverage from the rehydration beverage concentrate. | Stokely-Van Camp, Inc. | A61B005/05 A61B005/00 A61K047/00
US20110135575 | A1 | 2011-06-09 | METALLOPROTEIN MRI CONTRAST AGENTS AND RELATED METHODS | 2008-08-28 | The present invention generally relates to Magnetic Resonance Imaging (MRI) analysis of the location, presence, and/or quantity of a particular molecule or material. In some embodiments, the invention relates to compositions and methods for determination of an analyte, wherein a compositions may produce an MRI signal that can be affected by the presence of a particular composition (e.g., analyte). In a particular embodiment, the composition comprises a protein. The invention also provides related nucleic acid molecules, polypeptides, and fragments thereof. Compositions of the invention may exhibit reduced toxicity and may be readily delivered in vivo. Various embodiments of the invention may be useful as sensors, diagnostic tools, and the like. | 1. A method for determining an analyte, comprising: introducing a genetically-engineered protein molecule into a biological sample and, if an analyte is present, allowing the analyte to bind to the genetically-engineered protein molecule; andexposing the sample to magnetic resonance imaging conditions, thereby determining the presence and/or amount of the analyte in the biological sample. 2. A method as in claim 1, wherein determination of the analyte is performed in vivo. 3. A method as in claim 1, wherein determination of the analyte is performed in vitro. 4. A method as in claim 1, wherein the analyte is an organic analyte. 5. A method as in claim 1, wherein the analyte is a neurotransmitter analyte. 6. A method as in claim 1, wherein the analyte is a metal ion, peptide, purine, hormone, gastrin, fatty acid, amine, or neurotransmitter analyte. 7. A method as in claim 1, wherein the analyte is a zinc ion, acetylcholine, norepinephrine (NE), epinephrine, dopamine (DA), serotonin (5-HT), melatonin, octopamine, tyramine, glutamic acid, gamma-aminobutyric acid (GABA), aspartic acid, glycine, adenosine, vasopressin, somatostatin, neurotensin, histamine, ATP, GTP, LH, insulin, and derivatives thereof. 8. A method as in claim 1, wherein the protein is administered to a subject. 9. A method as in claim 8, wherein the subject is a human. 10. A method as in claim 1, wherein the act of introducing the genetically-engineered protein molecule into the biological sample comprises introducing a nucleic acid molecule which encodes the protein into the biological sample, either virally or non-virally mediated. 11. A method as in claim 1, wherein the genetically-engineered protein comprises a paramagnetic metal ion. 12. A method as in claim 11, wherein the paramagnetic metal ion is an ion of iron, nickel, manganese, copper, gadolinium, dysprosium, or europium. 13. A method as in claim 1, wherein the genetically-engineered protein is a cytochrome, tyrosine hydroxylase, or phenylalanine hydroxylase. 14. A method as in claim 1, wherein the genetically-engineered protein is P450 BM3. 15. A method as in claim 1, wherein the genetically-engineered protein is generated using directed evolution. 16. A method for determining of an analyte, comprising: providing a genetically-engineered protein molecule having a first relaxivity upon exposure to magnetic resonance imaging conditions;exposing the genetically-engineered protein molecule to a biological sample suspected of containing an analyte, wherein the genetically-engineered protein molecule interacts with the analyte, if present, to generate a second relaxivity of the genetically-engineered protein molecule, upon exposure to said magnetic resonance imaging conditions; anddetermining a change, or lack thereof, between the first and second relaxivities, thereby determining the presence and/or amount of the analyte in the biological sample. 17. A method as in claim 16, wherein the first and second relaxivi... | Massachusetts Institute of Technology | A61B005/055 G01N024/08 G01N033/20 G01N033/68 C12Q001/26 C07H021/00 C12N015/63 C12N005/00 C07K002/00
US07948370 | B2 | 2011-05-24 | Method and apparatus for providing data communication in data monitoring and management systems | 2009-08-14 | Method and apparatus for communicating with a sealed electronic device via the electronic device's existing data ports for programming, testing configuration or diagnosis of the electronic device such as a transmitter unit for use in a data monitoring and management system such as analyte monitoring and management system is provided. | 1. A data monitoring system, comprising: an analyte sensor for transcutaneous positioning in fluid contact with an interstitial fluid under a skin layer, the analyte sensor including an electrode; andelectronics including a data port with an electrical contact for operative coupling with the analyte sensor and configured to operate in one or more of a first operational mode or a second operational mode; wherein the electrical contact of the electronics data port is configured for signal communication with the electrode of the analyte sensor during the first operational mode; and furtherwherein the electrical contact of the electronics data port is configured for a uni-directional data communication during the second operational mode. 2. The system of claim 1 wherein the electronics includes a housing. 3. The system of claim 2 wherein the housing is substantially water tight. 4. The system of claim 1 wherein the electronics is configured to communicate data over a wireless data network. 5. The system of claim 1 wherein the electronics includes an RF transmitter. 6. The system of claim 1 wherein the electrical contact of the electronics is configured to receive one or more signals corresponding to an analyte level via the data port during the first operational mode. 7. The system of claim 1 wherein the analyte sensor includes a glucose sensor. 8. The system of claim 1 wherein the sensor is configured to monitor an analyte level under the skin layer. 9. The system of claim 1, wherein the electronics is configured to receive analyte related signals from the analyte sensor via the electrical contact of the data port during the first operational mode, and further, wherein the electronics is configured to receive an input signal from an interface unit via the electrical contact of the data port during the second operational mode. 10. The system of claim 9 wherein the received input signal from the interface unit includes a serial data input. 11. The system of claim 9 wherein the electrical contact of the data port to receive the analyte related signals during the first operational mode is the same electrical contact of the data port configured to receive the input signal from the interface unit during the second operational mode. 12. The system of claim 9 wherein the input signal received from the interface unit via the electrical contact of the electronics data port during the second operational mode includes one or more of a debugging command, a reset command, a diagnosis command, or a programming command. 13. The system of claim 1 wherein the electrode of the analyte sensor includes one of a working electrode, a counter electrode, or a guard contact. 14. The system of claim 1 wherein the first operational mode and the second operational mode do not overlap. 15. A data monitoring system, comprising: an analyte sensor for transcutaneous positioning in fluid contact with an interstitial fluid under a skin layer, the analyte sensor includin... | Abbott Diabetes Care Inc. | G08B001/08 A61B005/00
US07936909 | B2 | 2011-05-03 | Method and device for detecting chemical anomalies and/or salient features in soft tissue of an object area | 2007-03-30 | At least one embodiment of the present application relates to a method and/or a device for automatically detecting chemical anomalies and/or salient features in soft tissue of an object area. In at least one embodiment of the method, two image data records of two computed tomography pictures of the object are provided which are recorded in the context of a different spectral distribution of the X-radiation. The following steps are carried out for each voxel of at least one interesting slice whose X-ray attenuation values are characteristic of soft tissue. Firstly, two X-ray attenuation values of the voxel, or two averaged X-ray attenuation values of the environment of the voxel are used to determine a data point of the voxel in a diagram in which the X-ray attenuation values are plotted against one another in the context of the two spectral distributions of the X-radiation. A perpendicular distance of the data point from a connecting straight line that connects prescribed data points for pure fat and for pure soft tissue in the diagram is then calculated. Finally, the voxel is marked and displayed with highlighting when its distance exceeds a prescribed threshold value for the distance. In at least one embodiment, at least one of the proposed method and the associated device enable chemical anomalies to be automatically detected on the basis of computed tomography. | 1. A method comprising: automatically detecting at least one of chemical anomalies and salient features in soft tissue of an object area,wherein two computed tomography pictures of the object area are recorded in a context of a different spectral distribution of X-radiation,wherein two image data records of the object area are reconstructed from raw data of the two computed tomography pictures, the two image data records including X-ray attenuation values of voxels on the object area in the context of a respective spectral distribution of the X-radiation, andwherein, for at least a portion of voxels of at least one slice of the object area whose mean value from two X-ray attenuation values of the two image data records lies within a prescribed value range, the automatically detecting includes, determining a data point of the voxel in a diagram from at least one of the two X-ray attenuation values of the voxel and two averaged X-ray attenuation values that are obtained by averaging X-ray attenuation values of voxels inside a prescribed volume around the voxel in the respective image data record, the diagram including the two X-ray attenuation values plotted against one another in the context of two spectral distributions of the X-radiation,calculating a perpendicular distance of the data point of the voxel from a connecting straight line that connects prescribed data points for pure fat and for pure soft tissue in the diagram, andmarking and highlighting the voxel when the calculated distance exceeds a prescribed value for the distance. 2. The method as claimed in claim 1, wherein, for each voxel whose mean value from the two assigned x-ray attenuation values lies inside the prescribed value range that is characteristic of soft tissue, firstly a three-dimensional volume area with a prescribed extent around the voxel is prescribed there are selected inside the volume area all the voxels whose X-ray attenuation values fulfill a prescribed criterion that is characteristic of soft tissue, and an averaged attenuation value of the selected voxels is calculated separately for each image data record in order to determine from the two averaged X-ray attenuation values the data point of the voxel from which the perpendicular distance from the connecting straight line is subsequently calculated. 3. The method as claimed in claim 1, wherein, for each voxel whose mean value from the two assigned X-ray attenuation values lies inside the prescribed value range that is characteristic of soft tissue, firstly a three dimensional volume area with a prescribed extent around the voxel is prescribed, there are selected inside the volume area all the voxels whose X-ray attenuation values fulfill a prescribed criterion that is characteristic of soft tissue, a number of selected voxels in the volume area is determined and is compared with a prescribed threshold value for the number, and an averaged attenuation value of the selected voxels is calculated separately for ea... | Siemens Aktiengesellschaft | A61B006/00 G06K009/00 A61B005/05
US07947000 | B2 | 2011-05-24 | Cannula system for free-space navigation and method of use | 2006-06-30 | The present invention is directed to a novel shape-transferring cannula system, which provides access to tortuous and unsupported paths. The shape-transferring cannula system and method enables exploration of hollow body structures, and creates a custom-contoured access port for insertion and removal of, for example, diagnostic, surgical, or interventional instruments to and from a site within the body to which the physician does not have line-of-sight access. | 1. A cannula system comprising: a plurality of serial links;a tensioning element comprising a proximal part and a distal part;a static anchor; anda dynamic anchor;wherein a distal one of the serial links is farther from the static anchor than a proximal one of the serial links;wherein the proximal part of the tensioning element is coupled to the static anchor, and the distal part of the tensioning element is coupled to the distal one of the serial links;wherein the dynamic anchor moves towards the distal one of the serial links to compress all the serial links towards the distal one of the serial links; andwherein the movement of the dynamic anchor towards the distal one of the serial links increases friction between adjacent serial links. 2. The system of claim 1: wherein the serial links are nested in one another. 3. The cannula system of claim 1, further comprising: a recirculating guide coupled to the distal one of the serial links;wherein the tensioning element extends through the recirculating guide; andwherein a first end and a second end of the tensioning element are coupled to the static anchor. 4. The cannula system of claim 1, further comprising: a recirculating guide coupled to the static anchor;wherein the tensioning element extends through the recirculating guide; andwherein a first end and a second end of the tensioning element are coupled to the distal one of the serial links. 5. The cannula system of claim 1, further comprising: a proximal recirculating guide coupled to the static anchor; anda distal recirculating guide coupled to the distal one of the serial links;wherein the tensioning element extends through the proximal and distal recirculating guides. 6. The cannula system of claim 1: wherein each of the links comprises a central orifice such that a central lumen is formed in the plurality of serial links; andwherein the tensioning element extends outside of each link's central orifice. 7. A cannula system, in which a distal end is positioned toward a surgical site and a proximal end is remote from the distal end, comprising: a first plurality of serial links;a first static plate coupled via a first tensioning element to a distal link of the first plurality of serial links;a first dynamic plate positioned to compress the first plurality of serial links towards the distal link of the first plurality of serial links;a first actuator coupled to move the first dynamic plate towards the distal link of the first plurality of serial links;a second plurality of serial links that extends through the first plurality of serial links;a second static plate coupled via a second tensioning element to a distal link of the second plurality of serial links;a second dynamic plate positioned to compress the second plurality of serial links towards the distal link of the second plurality of serial links; anda second actuator coupled to move the second dynamic plate towards the distal link of the second plurality of serial links. 8. The sy... | Intuitive Surgical Operations, Inc. | A61B005/103
US09212988 | B2 | 2015-12-15 | Liquid concentration detector for a blood purification apparatus | 2010-12-14 | A blood purification apparatus includes a blood purification instrument for extracorporeally circulating blood of a patient and a concentration detector detecting a concentration of liquid flowing during blood purification. The concentration detector has a light emitter irradiating light onto said liquid, a light receiver receiving light from the light emitter transmitted through said liquid, and a detector detecting the received light intensity received by the light receiver. The concentration detector detects the concentration of the liquid based on the received light intensity and can be calibrated by adjusting an amount of irradiation by the light emitter so that the received light intensity has a predetermined value. | 1. A blood purification apparatus comprising: a blood purification instrument for extracorporeally circulating blood of a patient; anda concentration detector for detecting a concentration of liquid flowing during blood purification, the liquid including a dialysate and the blood of the patient, the concentration detector comprising: a light emitter configured to irradiate the liquid with light,a light receiver configured to receive the light from the light emitter transmitted through the liquid, anda detector configured to detect intensity of the light received by the light receiver, the concentration of the liquid being determined based on the intensity of the received light, anda controller programmed to perform calibration of the concentration detector before the blood purification instruction starts circulating the blood of the patient, in which the controller is programmed to supply only the dialysate to the concentration detector to adjust an amount of light emission from the light emitter so that the intensity of the received light has a predetermined value, whereinwhen the controller determines that the intensity of the received light is greater than a predetermined value, the controller controls the light emitter to reduce the amount of the light emission, andwhen the controller determines that the intensity of the received light is smaller than the predetermined value, the controller controls the light emitter to increase the amount of the light emission. 2. The blood purification apparatus of claim 1 wherein the controller further configured to: determine whether the received light intensity detected by the detector is substantially same as the predetermined value, andbased on the determination, control the light emitter so that the received light intensity has the predetermined value. 3. The blood purification apparatus of claim 1 wherein the light emitter comprises LED(s), andthe light receiver comprises a light receiving element(s) generating a voltage corresponding to the received light intensity, whereinthe controller adjusts light emitting current of the light emitter so that a voltage generated by the light receiver has a predetermined value. 4. A blood purification apparatus comprising: a blood purification instrument for extracorporeally circulating blood of a patient; anda concentration detector for detecting a concentration of liquid flowing during blood purification, the liquid including a dialysate and the blood of the patient, the concentration detector comprising: a light emitter configured to irradiate the liquid with light,a light receiver configured to receive the light from the light emitter transmitted through the liquid, anda detector configured to detect intensity of the light received by the light receiver, the concentration of the liquid being determined based on the intensity of the received light, anda controller programmed to perform calibration of the concentration detector before the blood purification ... | NIKKISO CO., LTD. | B01D011/00 B01D061/00 C02F001/44 G01N021/27 A61M001/16 G01N021/59 G01N021/05
US08652059 | B2 | 2014-02-18 | Cuff for blood pressure information measurement device and blood pressure information measurement device provided with the same | 2011-02-18 | A cuff includes a cuff main body portion and a gripping portion provided on an outer peripheral surface of the cuff main body portion. The cuff main body portion includes an air bladder, a tightening belt wrapped around an outer side of the air bladder, and a wrapping length adjustment mechanism for variably adjusting a wrapping length of the tightening belt. The wrapping length adjustment mechanism has a bias spring for pulling and biasing the tightening belt in a direction in which the wrapping length of the tightening belt is shortened, a first restriction portion for restricting extension of the wrapping length, and a second restriction portion for restricting shortening of the wrapping length. The gripping portion includes a push button for switching actions of the two restriction portions. In this configuration, the cuff may be easily attached to a measurement site regardless of the circumferential length of the site. | 1. A cuff for a blood pressure information measurement device configured to be attached to a measurement site of a living body and used for measurement of blood pressure information, comprising: a cuff main body portion formed into an annular form into which the measurement site is insertable from an axial direction, and configured to be attached to the measurement site in time of the measurement; anda gripping portion provided on an outer peripheral surface of the cuff main body portion, whereinthe cuff main body portion includes a fluid bag for compressing the living body, a tightening belt wrapped around an outer side of the fluid bag, and a wrapping length adjustment mechanism for variably adjusting a wrapping length of the tightening belt,the wrapping length adjustment mechanism has a winding roller for winding and feeding the tightening belt, a bias portion for pulling and biasing the tightening belt in a direction in which the wrapping length of the tightening belt is shortened, a first restriction portion for restricting extension of the wrapping length of the tightening belt, and a second restriction portion for restricting shortening of the wrapping length of the tightening belt, andeach of the first restriction portion and the second restriction portion is formed by a one-way clutch additionally provided in the winding roller,the cuff further comprising a switching portion for selectively switching whether the wrapping length adjustment mechanism is in a first state where restriction by the second restriction portion is released, restriction by the first restriction portion is applied, and thereby, the wrapping length of the tightening belt is variably adjusted only in the direction in which the tightening belt is pulled by the bias portion and the wrapping length thereof is shortened, or in a second state where the restriction by the first restriction portion is released, the restriction by the second restriction portion is applied, and thereby, the wrapping length of the tightening belt is variably adjustable only in a direction in which the wrapping length of the tightening belt is extended, the switching portion being provided in the gripping portion or the cuff main body portion in a vicinity of the gripping portion. 2. The cuff for the blood pressure information measurement device according to claim 1, wherein the switching portion is formed by a push button. 3. The cuff for the blood pressure information measurement device according to claim 2, wherein the wrapping length adjustment mechanism is in the second state in conjunction with a press-down state of the push button, and in the first state in conjunction with release of the press-down state of the push button. 4. The cuff for the blood pressure information measurement device according to claim 1, wherein the bias portion is formed by a spring attached to the winding roller. 5. The cuff for the blood pressure information measurement device according to claim 1, further ... | OMRON HEALTHCARE Co., Ltd. | A61B005/02
US08435493 | B2 | 2013-05-07 | Imaging agents | 2010-11-15 | The present invention provides novel amino acid compounds useful in detecting and evaluating brain and body tumors. These compounds have the advantageous properties of rapid uptake and prolonged retention in tumors and can be labeled with halogen isotopes such as fluorine-18, iodine-123, iodine-124, iodine-125, iodine-131, bromine-75, bromine-76, bromine-77, bromine-82, astatine-210, astatine-211, and other astatine isotopes. These compounds can also be labeled with technetium and rhenium isotopes using known chelation complexes. The compounds disclosed herein bind tumor tissues in vivo with high specificity and selectivity when administered to a subject. Preferred compounds show a target to non-target ratio of at least 2:1, are stable in vivo and substantially localized to target within 1 hour after administration. Preferred compounds include 1-amino-2-[18F]fluorocyclobutyl-1-carboxylic acid (2-[18F]FACBC) and 1-amino-2-[18F]fluoromethylcyclobutyl-1-carboxylic acid (2-[18F]FMACBC). The labeled amino acid compounds of the invention are useful as imaging agents in detecting and/or monitoring tumors in a subject by PET or SPECT. | 1. A method of synthesizing an amino acid analog of the formulaor salt thereof comprising the step of reacting a halogenated compound with a compound of the following formulafollowed by an acid to yield the amino acid analogR1 is tert-butoxycarbonyl(Boc) or bis(4-methoxyphenyl)methyl in the compound and hydrogen in the amino acid analog;R2 is hydrogen;X is selected from the group consisting of halogen, haloalkyl, halocycloalkyl, halocycloalkenyl, halocycloalkynyl, haloacyl, haloaryl, haloheteroaryl, haloalkenyl, haloalkynyl, where halo or halogen in X is selected from the group consisting of F, Cl, Br, I, At, F-18, Br-76, I-123, I-124;Y is selected from the group consisting of (CR5R6)n, n=1-4, N, O, S, and Se;R4, R5 and R6 are independently selected from the group consisting of H, alkyl, haloalkyl, cycloalkyl, halocycloalkyl, heteroaryl, aryl, haloaryl, haloheteroaryl, alkenyl, haloalkenyl, alkynyl, and haloalkynyl, where halo is selected from the group consisting of non-radioactive F, Cl, Br, and I; andR3 is selected from the group consisting of H, alkyl, haloalkyl, cycloalkyl, halocycloalkyl, cycloalkenyl, halocycloalkenyl, cycloalkynyl, halocycloalkynyl, acyl, haloacyl, aryl, haloaryl, heteroaryl, haloheteroaryl, alkenyl, haloalkenyl, alkynyl, and haloalkynyl. 2. The method of claim 1 wherein the amino acid analog is 1-amino-2-[18F]fluorocyclobutyl-1-carboxylic acid. 3. The method of claim 1, wherein X is F-18 or I-123. 4. The method of claim 1, wherein Y is CH2. 5. The method of claim 1, wherein R3 is tert-butyl in the compound and H in the amino acid analog. 6. The method of claim 1, wherein R4 is hydrogen. 7. The method of claim 1, wherein the halogenated compound comprises a halogen selected from F-18, Br-76, I-123, and I-124. 8. The method of claim 1, wherein the halogenated compound is a halogen salt. 9. The method of claim 8, wherein the halogen salt is K18F, K123I, or Na123I. | Emory University | A61K051/00 A61M036/14
US08423128 | B2 | 2013-04-16 | Cardiac arrest monitoring device | 2009-12-22 | The embodiments herein relate to devices, implementations, and techniques for health monitoring, specifically cardiac arrest monitoring systems. A paramedical equipment locator may receive a periodically sent message configured to identify a location information corresponding to a paramedical equipment and determine whether to select the paramedical equipment based at least in part on the location information. | 1. A cardiac arrest monitoring device, comprising: a sensor adapted to detect a cardiac rhythm;an anomaly detection circuit adapted to detect at least one anomaly based on the detected cardiac rhythm;a self-location acquisition unit for acquiring self-location information;a paramedical equipment locator adapted to: receive a periodically sent message configured to identify a location information corresponding to a paramedical equipment,determine a distance corresponding to the paramedical equipment using at least the received location information, anddetermine whether to select the paramedical equipment based at least in part on the determined distance; andan output unit configured to output the location information of a selected paramedical equipment;wherein the paramedical equipment locator and the output unit are wearable. 2. A device as recited in claim 1, wherein the anomaly detection circuit further comprises a comparison circuit adapted to compare the detected cardiac rhythm with preprogrammed data. 3. A device as recited in claim 2, wherein the preprogrammed data comprises previous cardiac rhythm data of a user of the cardiac monitoring device. 4. A device as recited in claim 1, wherein the paramedical equipment comprises an automated external defibrillator. 5. A device as recited in claim 1, further comprising an alarm unit adapted to generate an alarm response to the anomaly detection circuit detecting the anomaly of the cardiac rhythm. 6. A device as recited in claim 1, wherein the output unit further comprises a video display to plot a location associated with the location information of the selected paramedical equipment on a map. 7. A device as recited in claim 1, wherein the output unit further comprises a video display to graphically display the detected cardiac rhythm. 8. A device as recited in claim 1, further comprising; a transmitter for sending the self location information measured by the self-location acquisition unit. 9. A method as recited in claim 1, wherein the periodically sent message is sent from a server. 10. A method as recited in claim 1, wherein the periodically sent message is sent from the paramedical equipment. | Empire Technology Development LLC | A61B005/04
US08676294 | B2 | 2014-03-18 | Brain disease diagnosis system | 2009-06-19 | To perform a more accurate and detailed diagnosis of a brain disease, a diagnosis server of the brain disease diagnosis system for diagnosing a brain disease of an examined person includes: an acquiring unit for acquiring a brain image of the examined person so as to obtain an acquired image; a region setting unit for setting a plurality of regions in the acquired image; an individual index value calculating unit for calculating an individual index value based on a pixel value of the acquired image, in each of the plurality of regions; a whole index value calculating unit for calculating a whole index value by weighting the individual index value of each of the plurality of regions; a diagnosis unit for diagnosing the brain disease of the examined person based on the whole index value; and an output unit for outputting information indicating a diagnosis outcome. | 1. A brain disease diagnosis system for diagnosing a brain disease of an examined person, comprising: acquiring means for acquiring a brain image of the examined person so as to obtain an acquired image;region setting means for setting a plurality of regions in the acquired image acquired by the acquiring means;individual index value calculating means for calculating an individual index value based on a pixel value of the acquired image, in each of the plurality of regions set by the region setting means;whole index value calculating means for calculating a whole index value by weighting the individual index value of each of the plurality of regions calculated by the individual index value calculating means;diagnosis means for diagnosing the brain disease of the examined person based on the whole index value calculated by the whole index value calculating means; andoutput means for outputting information indicating a diagnosis outcome issued by the diagnosis means. 2. The brain disease diagnosis system according to claim 1, wherein the diagnosis means diagnoses the brain disease of the examined person by comparing a threshold value obtained based on an index value of sample data having the brain disease and an index value of sample data not having the brain disease and the whole index value calculated by the whole index value calculating means. 3. The brain disease diagnosis system according to claim 1, wherein the whole index value calculating means performs the weighting based on the index value of the sample data having the brain disease and the index value of the sample data not having the brain disease. 4. The brain disease diagnosis system according to claim 1, wherein the acquiring means corrects the image based on the pixel value of the acquired brain image so as to obtain the acquired image. 5. The brain disease diagnosis system according to claim 1, wherein the acquiring means also acquires information indicating the age of the examined person, andthe diagnosis means diagnoses the brain disease of the examined person according to the age of the examined person indicated by the information acquired by the acquiring means. 6. The brain disease diagnosis system according to claim 1, wherein the acquiring means anatomically standardizes the acquired brain image so as to obtain the acquired image. 7. The brain disease diagnosis system according to claim 1, further comprising imaging means for imaging the brain image of the examined person, wherein the acquiring means acquires the image imaged by the imaging means. 8. The brain disease diagnosis system according to claim 7, wherein the imaging means images a slice image of a brain of the examined person as the brain image of the examined person, andthe acquiring means produces from the slice image imaged by the imaging means a brain surface projection image obtained by projecting a brain surface of the brain so as to obtain the acquired image. | Hamamatsu Photonics K.K. | A61B005/05
US20110125206 | A1 | 2011-05-26 | SINGLE CHAMBER IMPLANTABLE MEDICAL DEVICE FOR CONFIRMING ARRHYTHMIA THROUGH RETROSPECTIVE CARDIAC SIGNALS | 2009-11-24 | An implantable medical device is provided that comprises a housing, sensors configured to be located to proximate a heart, and a sensing module to sense cardiac signals originating from the heart over a channel defined by the sensors. The cardiac signals include intrinsic R-wave events and associated intrinsic confirmation events when the heart exhibits normal sinus rhythm. The device further includes memory to store the cardiac signals sensed over a channel, and a detection module. The detection module identifies an R-wave event within the cardiac signals. The detection module captures, in the memory, a segment of the cardiac signals that precedes the R-wave event as a retrospective segment. The detection module determines whether the retrospective segment includes an intrinsic confirmation event that is associated with and occurs before the R-wave event. The detection module declares an arrhythmia based at least in part on the determination of whether the retrospective segment includes the intrinsic confirmation event is absent from the retrospective segment. | 1. An implantable device, comprising: a housing;sensors configured to be located proximate to a heart;a sensing module to sense cardiac signals originating from the heart over a channel defined by the sensors, the cardiac signals including intrinsic R-wave events and associated intrinsic confirmation events when the heart exhibits normal sinus rhythm;memory to store the cardiac signals sensed over the channel; anda detection module to identify an R-wave event within the cardiac signals, the detection module to capture, in the memory, a segment of the cardiac signals that precedes the R-wave event as a retrospective segment, the detection module to determine whether the retrospective segment includes an intrinsic confirmation event that is associated with and occurs before the R-wave event, the detection module to declare an arrhythmia based at least in part on the determination of whether the intrinsic confirmation event is absent from the retrospective segment. 2. The device of claim 1, further comprising a lead configured to be coupled to the housing and to be located proximate to a heart, the lead including at least a portion of the sensors utilized to define a primary channel and a secondary channel. 3. The device of claim 1, wherein the cardiac signals include a near field signal component and a far field signal component, the detection module identifying the R-wave event from the near field signal component, the detection module analyzing the far field signal component for the intrinsic confirmation event. 4. The device of claim 1, wherein the sensing module utilizes different sets of sensors to sense cardiac signals over primary and secondary channels. 5. The device of claim 1, wherein the sensing module includes high and low sensitivity amplifiers coupled to the sensors, the high sensitivity amplifiers separating a far field signal component, the low sensitivity amplifiers separating a near field signal component. 6. The device of claim 1, wherein the intrinsic confirmation event is one of a P-wave, a HIS signal and a PVC signal. 7. The device of claim 1, wherein the detection module declares the arrhythmia to constitute one of atrial fibrillation, flutter, and bigeminy. 8. The device of claim 1, wherein the sensing module monitors a near field (NF) channel for the R-wave event originating within a chamber of the heart proximate to a first sensor and monitors a far field (FF) channel for the intrinsic confirmation event originating within a chamber of the heart remote from the first sensor. 9. The device of claim 1, the detection module to combine a plurality of retrospective segments to form an ensemble average of retrospective segments. 10. The device of claim 9, wherein the detection module compares current and prior ensemble averages in connection with declaring the arrhythmia. 11. The device of claim 1, wherein the detection module correlates multiple retrospective segments to obtain a degree of correlation, the detection m... | PACESETTER, INC. | A61N001/39 A61B005/0452 A61B005/046
US07942831 | B2 | 2011-05-17 | Measuring method and measuring device | 2007-04-24 | The present invention is to present a measuring device which can conduct stable extraction and measurement of interstitial fluid. The measuring device comprises: a device body which comprises an extraction part for extracting body fluid from a living body and a measuring part for measuring a biological component contained in the extracted body fluid; a holding member which holds the device body and has an opening part for defining a measurement area on the living body; and a mounting member for mounting the holding member on the living body; wherein the measurement area on the living body is defined by the opening part when the holding member is mounted on the living body by the mounting member. | 1. A measuring device, comprising: a device body which comprises an extraction part for extracting body fluid from a living body and a measuring part for measuring a biological component contained in the extracted body fluid;a holding member which holds the device body and has an opening part for defining a measurement area on the living body; anda mounting member for mounting the holding member on the living body;wherein the measurement area on the living body is defined by the opening part when the holding member is mounted on the living body by the mounting member, andwherein the mounting member comprises a tightening part which is configured to tighten the living body between the tightening part and the holding part, and a connecting part which connects the tightening part and the holding part so that a distance between the tightening part and the holding part is changeable. 2. The measuring device of claim 1, wherein the opening part of the holding member is configured to receive a micropore forming device for forming a micropore on the living body to form a micropore in the measurement area via the opening part, andwherein the extraction part is configured to extract the body fluid from the micropore formed by the micropore forming device. 3. The measuring device of claim 1, wherein the extraction part is detachably attached to the device body. 4. The measuring device of claim 1, wherein the device body is held by the holding member so that the measuring device has a first state and a second state, andwherein, in the first state, the extraction part contacts the living body via the opening part, and in the second state, the extraction part is separated from the living body. 5. The measuring device of claim 4, wherein the device body is held by the holding member so that the device body is movable relative to the holding member, and the measuring device changes between the first state and the second state by movements of the device body. 6. The measuring device of claim 1, wherein the connecting part connects the tightening part to the holding part so that the tightening part is linearly movable relative to the holding part. 7. The measuring device of claim 1, wherein one end of the holding part and one end of the tightening part are connected by the connecting part, and the other end of the holding part and the other end of the tightening part are separated from each other. 8. The measuring device of claim 1, wherein the measurement area is defined by the opening part when the living body is held between the tightening part and the holding member. 9. The measuring device of claim 1, wherein the tightening part is flexible so as to become in close contact with the living body while the living body is held between the tightening part and the holding member. 10. The measuring device of claim 1, wherein a surface of the tightening part, which is facing the holding member, is formed concavely. 11. The measuring device of claim 1, where... | Sysmex Corporation | A61B005/00
US07942829 | B2 | 2011-05-17 | Biopsy planning and display apparatus | 2008-01-15 | Systems are provided for managing, planning and displaying the location of biopsy core sites across multiple same-patient visits or other stored biopsy plans to increase the confidence and accuracy of biopsy results. The systems may allow the operator to fuse multiple predefined and/or automatically generated biopsy core site plans in addition to previous biopsy sites with a current prostate image to assist in creating a final new biopsy core site plan. Further, the system may allow the operator to selectively view singular parts of the final plan to simplify the planning process. Further, the systems may assist the operator during the navigation phase of the biopsy by hiding biopsy core sites from the final plan that are not of immediate interest during the biopsy phase and thereby reducing clutter and distraction during biopsy sample acquisition. | 1. A method for use in a computer system associated with a medical imaging device for use in prostate biopsy planning and targeting, comprising: accessing, from an electronic database, a previously stored prostate first image with biopsy information, said biopsy information including biopsy core sample information for previously taken biopsy core samples, said biopsy core sample information including bio; core sample locations and pathological information including locations of any cells of interest along the lengths of said biopsy core samples;obtaining a current second image of a prostate of a patient;registering the first and second images to identify a transformation between the images;incorporating said biopsy information from said first image into said second image to generate a composite image, wherein said transformation is applied to said biopsy information to transform said biopsy information into a frame of reference of said second image; andoutputting said composite image to a display device. 2. The method of claim 1, further comprising: receiving at least one user selected biopsy site and incorporating said user selected biopsy sites into said composite image. 3. The method of claim 2, further comprising: after incorporating the at least one user selected biopsy site into the composite image, removing at least a portion of said biopsy information from said composite image to generate a biopsy guidance image. 4. The method of claim 1, further comprising accessing first and second stored prostate plans from said electronic database and incorporating said first and second stored prostate plans into said composite image. wherein at least one of said first and second stored prostate plans is selected from a group consisting of: a predefined biopsy plan identifying one or more biopsy sites;statistical information identifying location specific information for one or more areas of a prostate; andan elastographical biopsy plan identifying one or more biopsy sites. 5. The method of claim 4, wherein information associated with said stored prostate plans is displayed having a unique identifier, the unique identifier including a unique symbol. 6. The method of claim 4, further comprising: providing to a user a menu of available prostate plans; andreceiving user inputs identifying said first and second prostate plans. 7. The method of claim 4, further comprising: creating a current biopsy plan;adding at least one target biopsy site from one of said first and second stored prostate plans to said current biopsy plan; andusing said current biopsy plan to guide a biopsy procedure. 8. The method of claim 7, wherein an identifier displayed on said composite image of said at least one target biopsy site changes from an identifier associated with said one of said first and second stored prostate plans to an identifier associated with said current biopsy plan when said at least one target biopsy site is added to said current biopsy plan. 9. The met... | Eigen, Inc. | A61B010/00 A61B005/103 A61B005/117
US07957789 | B2 | 2011-06-07 | Therapy delivery system including a navigation element | 2005-12-30 | An electromagnetic receiver assembly is included in a navigation element for a therapy delivery system. If the system is modular, the navigation element may be an insertable module thereof and/or include a lumen to receive another insertable module. | 1. An electromagnetic receiver assembly for a navigation element of a therapy delivery system, the assembly comprising: a receiver coil including a proximal end, a distal end and a coiled wire, the coiled wire includes a first wire end and a second wire end, both wire ends of the coiled wire disposed in proximity to the receiver coil distal end; anda conductor including a first insulated wire extending alongside a second insulated wire, the first and second wires of the conductor each include an end, the first and second conductor wire ends extending away from one another in proximity to the receiver coil proximal end and the first and second insulated wires of the conductor are respectively wrapped about the first and second wire ends of the coiled wire in proximity to the distal end of the receiver coil and extending within the receiver coil to the proximal end of the receiver coil and outward from the proximal end of the receiver coil. 2. The assembly of claim 1, further comprising a core about which the coiled wire is wound, and wherein the conductor extends in between the coil and the core. 3. The assembly of claim 2, wherein the core includes a length that is within approximately 0.01 inch and 1 inch of the length of the coil. 4. The assembly of claim 2, wherein the core includes a channel extending along an outer surface thereof and the conductor extends within the channel. 5. The assembly of claim 2, wherein the core includes a longitudinally extending lumen and a minimum breadth of the lumen is greater than approximately 0.04 inch. 6. The assembly of claim 1, wherein an outer diameter of the receiver coil is no greater than approximately 0.08 inch. 7. The assembly of claim 1, wherein the conductor includes a strain-relief tubing extending around the first wire and the second wire and along a length of the conductor that extends away from the second end of the coil. 8. The assembly of claim 1, wherein a length of the conductor is greater than approximately 80 inches. | Medtronic, Inc. | A61B005/05
US08784654 | B2 | 2014-07-22 | Methods and systems for biological sample collection and analysis | 2008-11-14 | Filtration devices for collection and filtration of biological samples are disclosed. Devices having a filtration element oriented in a generally vertical orientation are provided, as well as filtration devices that incorporate a cooling mechanism to reduce the temperature of collected solids. Tissue collection devices, such as aspiration assemblies, tissue sampling devices, and the like incorporating filtration devices are disclosed. Methods of collecting biological samples and separating biological solids from a liquid/solids mixture are also disclosed, together with analytical techniques and protocols for analyzing biological samples. | 1. A cooled filtration assembly for filtering biological materials comprising a housing having a liquid/solids input chamber and a liquid output chamber; a liquid/solids input port communicating with the liquid/solids input chamber, a liquid output port communicating with the liquid output chamber, and a filtration member interposed between the liquid/solids input chamber and the liquid output chamber, wherein the filtration member comprises a thermoelectric cooling surface that is actively cooled during operation and filtrate comprising biological materials residing in the filtration assembly during a filtration operation is cooled to a temperature at least 20° F. cooler than the temperature of biological materials entering the liquid/solids input port. 2. The cooled filtration assembly of claim 1, wherein the filtration member is oriented in a generally vertical orientation during operation of the cooled filtration assembly. 3. The cooled filtration assembly of claim 1, additionally comprising a jacket mountable on the filtration assembly and comprising an active or passive cooling system. 4. The cooled filtration assembly of claim 1, additionally comprising an insulating jacket mountable on the filtration assembly. 5. The cooled filtration assembly of claim 1, wherein the filtration member comprises a size exclusion filtration member. 6. The cooled filtration assembly of claim 1, wherein the cooled filtration assembly comprises an additional port in proximity to the liquid/solids input chamber for introduction of a cooled fluid during the filtration operation. 7. The cooled filtration assembly of claim 1, wherein at least one of the liquid/solids input chamber and the liquid output chamber additionally incorporates an active cooling element. 8. A cooled filtration assembly for filtering biological materials comprising a liquid/solids input port communicating with a liquid/solids input chamber, a liquid output chamber and a liquid output port communicating with the liquid output chamber, and a filtration member interposed between the liquid/solids input chamber and the liquid output chamber, wherein at least one of the liquid/solids input chamber, the liquid output chamber and the filtration member is actively cooled during operation and filtrate comprising biological materials residing in the filtration assembly during a filtration operation is cooled to a temperature at least 20° F. cooler than the temperature of biological materials entering the liquid/solids input port, wherein the filtration member is oriented in a generally vertical orientation during operation of the cooled filtration assembly and extends less than the full height of the liquid/solids input chamber and liquid output chamber, leaving a gap permitting material to flow between the liquid/solids input chamber and the liquid output chamber. 9. The cooled filtration assembly of claim 8, wherein at least one of the liquid/solids input chamber, the liquid output chamber ... | Bayer Medical Care, Inc. | B01D035/18 B01D037/00 A61M001/36
US08642255 | B2 | 2014-02-04 | Materials and methods for hypothermic collection of whole blood | 2009-03-27 | The present invention relates to materials and methods for hypothermic collection of whole blood, and components thereof, which can extend the holding time of blood beyond the current useable limit. Additionally, blood can be drawn directly into a hypothermic preservation solution without the addition of standard anticoagulants. This is enabled by providing sustained cellular viability under hypothermic conditions using a nutrient matrix devoid of animal proteins and containing energy substrates, free-radical scavengers, and impermeants that is ionically balanced for storage of biologic materials at low temperatures to prevent cellular stress-induced apoptosis. | 1. A method of harvesting whole blood, the method comprising: a) providing a volume of whole blood from a donor; and,b) mixing, in a container, the volume of whole blood, or a portion thereof, with a volume of preservation solution adapted for cellular osmotic balance at hypothermic temperatures, such that the ratio of whole blood to preservation solution is at least about 1:0.1 by volume,c) maintaining the mixture of whole blood and preservation solution at hypothermic temperatures,wherein the volume of whole blood, the volume of preservation solution, and the container are free from an exogenously-introduced anticoagulant reagent, andwherein the preservation solution comprises an ingredient selected from the group consisting of: an aqueous solution of electrolytes containing potassium ions at a concentration range of from about 35 to about 45 mM, sodium ions at a concentration range of from about 80 to about 120 mM, magnesium ions at a concentration range of from about 2 to about 10 mM, chloride ions at a concentration range of from about 15 to about 20 mM, and calcium ions at a concentration range of from about 0.01 to about 0.1 mM; an impermeant anion; mannitol; a macromolecular oncotic agent; at least one simple sugar; a substrate for the regeneration of ATP; a biological pH buffer effective under physiological hypothermic conditions; and combinations thereof. 2. The method of claim 1, wherein the ratio of whole blood to preservation solution is at least about 1:0.5, 1:1, 1:2, 1:5, or 1:10. 3. The method of claim 1, wherein the mixture of whole blood and preservation solution is maintained at about 2° C. to about 18° C. 4. The method of claim 1 wherein the preservation solution comprises an aqueous solution of electrolytes containing potassium ions at a concentration range of from about 35 to about 45 mM, sodium ions at a concentration range of from about 80 to about 120 mM, magnesium ions at a concentration range of from about 2 to about 10 mM, chloride ions at a concentration range of from about 15 to about 20 mM, and calcium ions at a concentration range of from about 0.01 to about 0.1 mM; an impermeant anion; mannitol; a macromolecular oncotic agent; at least one simple sugar; a substrate for the regeneration of ATP; and a biological pH buffer effective under physiological hypothermic conditions. 5. The method of claim 4 wherein the preservation solution further comprises glutathione, a vitamin E derivative, an antioxidant, and combinations thereof. 6. The method of claim 1 wherein the whole blood is cord blood. 7. The method of claim 1, wherein the mixture of whole blood and preservation solution is maintained at about 2° C. to about 18° C. for at least 24 hours. 8. The method of claim 1, wherein the mixture of whole blood and preservation solution is maintained at about 2° C. to about 18° C. for at least 48 hours. 9. The method of claim 1, wherein the mixture of whole blood and preservation solution is maintained at about 2° C.... | Biolife Solutions, Inc. | A01N001/02
US08758275 | B2 | 2014-06-24 | Moving body posture angle processing device | 2009-07-02 | A moving body attitude angle processing apparatus capable of reducing cumulative errors in continuously measuring the attitude angle of a moving body is provided. The moving body attitude angle processing apparatus estimates the attitude angle of a moving body in accordance with outputs from an acceleration sensor, a magnetic sensor, and a gyro sensor that are mounted to the moving body, and outputs the attitude angle thus estimated. The moving body attitude angle processing apparatus includes: a gyro sensor that outputs an angular velocity vector; an acceleration sensor that outputs an acceleration vector; a magnetic sensor that outputs a magnetic field vector; a movement state determining device that determines a movement state of a moving body in accordance with an output from the acceleration sensor; and a gyro-sensor zero-point output estimating device that, when the movement state determining device determines the movement state to have been detected as a stationary state of the moving body, carries out a process of estimating a zero-point output of the gyro sensor in accordance with the output from the acceleration sensor, an output from the magnetic sensor, and an output from the gyro sensor. | 1. A moving body attitude angle processing apparatus for estimating an attitude angle of a moving body in accordance with outputs from an acceleration sensor, a magnetic sensor, and a gyro sensor that are mounted to the moving body, and for outputting the attitude angle thus estimated, the moving body attitude angle processing apparatus comprising: a gyro sensor that outputs an angular velocity vector;an acceleration sensor that outputs an acceleration vector;a magnetic sensor that outputs a magnetic field vector of an ambient magnetic field that is different from a geomagnetic field;a movement state determining device that determines a movement state of a moving body in accordance with an output from the acceleration sensor;a gyro-sensor zero-point output estimating device that, when the movement state determining device determines the movement state to have been detected as a stationary state of the moving body, carries out a process of estimating a zero-point output of the gyro sensor in accordance with the output from the acceleration sensor, an output from the magnetic sensor, and an output from the gyro sensor, the gyro-sensor zero-point output estimating device is embodied by an extended Kalman filter having, as its components, only a gravitational acceleration vector, a magnetic field vector of the ambient magnetic field, an angular velocity vector, and an offset value of zero-point output of angular velocity by the gyro sensor,a zero point of the gyro sensor being corrected by utilizing an output from the gyro-sensor zero-point output estimating device; andan attitude angle estimating device that estimates the attitude angle in accordance with the outputs from the acceleration sensor, the magnetic sensor, and the gyro sensor. 2. The moving body attitude angle processing apparatus as set forth in claim 1, wherein the attitude angle estimating device corrects an output of estimation of an attitude angle of the moving body in accordance with the output from the gyro-sensor zero-point output estimating device. 3. A moving body attitude angle processing apparatus for estimating an attitude angle of a moving body in accordance with outputs from an acceleration sensor, a magnetic sensor, and a gyro sensor that are mounted to the moving body, and for outputting the attitude angle thus estimated, the moving body attitude angle processing apparatus comprising: a gyro sensor that outputs an angular velocity vector;an acceleration sensor that outputs an acceleration vector;a magnetic sensor that outputs a magnetic field vector of an ambient magnetic field that is different from a geomagnetic field;a movement state determining device that determines a movement state of a moving body in accordance with an output from the acceleration sensor and an output from the gyro sensor;a gyro-sensor zero-point output estimating device that, when the movement state determining device determines the movement state to have been detected as a stationary state of t... | National Institute of Advanced Industrial Science and Technology | A61B005/103 A61B005/117 G01C009/00 G01C017/00 G01C019/00
US20110106201 | A1 | 2011-05-05 | IMPLANTABLE HEART FAILURE MONITOR | 2009-10-30 | An implantable medical device and associated method monitor a heart failure patient by sensing a signal responsive to oxygen availability in an extravascular volume of skeletal muscle tissue. The signal is used to compute a tissue oxygenation measurement. A change in the tissue oxygenation measurement is detected, and a time interval corresponding to the detected change in muscle tissue oxygenation is computed. The time interval is used for detecting if a heart failure condition is worsening or improving. | 1. A method for monitoring a heart failure patient, the method comprising: sensing a first signal responsive to oxygen availability in an extravascular volume of skeletal muscle tissue in the patient and computing a tissue oxygenation measurement using the first signal;detecting a change in the tissue oxygenation measurement;computing a time interval corresponding to the detected change in muscle tissue oxygenation; anddetecting if a heart failure condition is worsening in response to the computed time interval. 2. The method of claim 1 further comprising: sensing a second signal responsive to physical activity of a patient;detecting an episode of non-resting patient activity in response to the first signal; anddetecting the change in the tissue oxygenation measurement subsequent to detecting the episode of non-resting patient activity. 3. The method of claim 2, wherein detecting the change in the tissue oxygenation measurement comprises detecting a first level of tissue oxygenation and detecting a second level of tissue oxygenation during the episode, the second level being lower than the first level. 4. The method of claim 3, further comprising detecting a third level of tissue oxygenation corresponding to a plateau during the episode, the second level being intermediate to the first level and the third level, the detected change being the change between the first level and the third level. 5. The method of claim 2, wherein detecting the change in the tissue oxygenation measurement comprises detecting a first level of tissue oxygenation corresponding to a minima reached during the detected non-resting activity; and detecting a second level of tissue oxygenation greater than the first level and occurring after the first level of tissue oxygenation. 6. The method of claim 5, further comprising detecting a third level of tissue oxygenation corresponding to a plateau higher than the first level and occurring after the first level, the second level being intermediate to the first level and the third level, the detected change being the change from the first level to one of the second and the third level. 7. The method of claim 2, further comprising: determining a feature of the first signal as a first measure of activity level for the episode;detecting a next episode of non-resting activity having a next measure of activity level equal to the first measure of activity level; andcomputing a next time interval corresponding to a change in tissue oxygenation in response to detecting the next episode;wherein detecting if the heart failure condition is worsening comprises determining a trend using the time interval and the next time interval. 8. The method of claim 7 wherein the feature being one of a maximum peak activity level, a maximum sustained activity level, and an average activity level. 9. The method of claim 2, further comprising: determining a feature of the first signal as a first measure of activity level for the episode;detecting a ... | - | A61N001/365 A61B005/1459
US20110130654 | A1 | 2011-06-02 | SYSTEM FOR CHARACTERIZING A CORNEA AND OBTAINING AN OPHTHALMIC LENS | 2011-02-07 | A system for determining the shape of a cornea of an eye illuminates at least one of the interior surface, the posterior surface, and the interior region of the eye with infrared light of a wavelength that can generate fluorescent light from the portion of the cornea illuminated. The generated fluorescent light is then detected. A step of illuminating can comprise focusing the infrared light in a plurality of different planes substantially perpendicular to the optical axis of the eye. From the detected light it is possible to create a map of at least a portion of the interior surface, at least a portion of the posterior surface, and/or portion of the interior region of the cornea. Clarity of vision can be determined by generating autofluorescence from proteins in the pigment epithelial cells of the retina. | 1-23. (canceled) 24. Apparatus for generating a map of a cornea of an eye, the cornea having an anterior surface, a posterior surface, and an interior region between the anterior and posterior surfaces, the apparatus comprising: a) a laser for illuminating a selected portion of the anterior surface, the posterior surface, and the interior region, the illumination being with a light of a wavelength that can generates autofluorescent light directly from the portion of the cornea illuminated;b) focusing means for focusing the illuminating light on the selected portion of the cornea;c) a photodetector for detecting the generated autofluorescent light. andd) means for generating the map from the detected autofluorescent light, wherein the map is of a refractive property and/or shape of the portion of the cornea illuminated. 25. (canceled) 26. The apparatus of claim 24 wherein the eye has an optical axis, and the focusing means focuses the illuminating light in a plurality of different planes substantially perpendicular to the optical axis of the eye. 27-34. (canceled) 35. The apparatus of claim 24 wherein the means for generating the map comprises a computerized control unit. 36. The apparatus of claim 24 wherein the laser emits pulses having a duration of from about 50 to 100 femtoseconds and an energy level of at least 0.2 nJ. 37. The apparatus of claim 24 wherein said laser generates pulses at 50 MHz, with each pulse having a duration of less than 100 femtoseconds and an energy level of less than 1 nJ. 38. The apparatus of claim 24 wherein said laser emits a sequence of femtosecond laser pulses. 39. Apparatus for generating a map of a cornea of an eye, the apparatus comprising: a) a laser for illuminating a selected portion of the cornea, the selected portion including an anterior surface, a posterior surface, and an interior region, the illumination being with a plurality of light beams of a wavelength that generates autofluorescent light directly from the portion of the cornea illuminated;b) focusing means for focusing the illuminating light beams on the selected portion of the cornea;c) a photodetector for detecting the generated autofluorescent light; andd) a computerized control unit for determining an optical path length for a plurality of locations from the detected autofluorescent light and generating a map from the detected autofluorescent light, wherein the map is of the spatial distribution of optical path lengths for the selected portion of the cornea illuminated. 40. The apparatus of claim 39 wherein the eye has an optical axis, and the focusing means focuses the illuminating light in a plurality of different planes substantially perpendicular to the optical axis of the eye. 41. The apparatus of claim 39 wherein the means for generating the map comprises a computerized control unit. 42. The apparatus of claim 39 wherein the laser emits pulses having a duration of from about 50 to 100 femtoseconds and an energy level of at l... | - | A61B005/05
US07938781 | B2 | 2011-05-10 | Hemodynamic stability assessment based on heart sounds | 2010-06-10 | A method comprises detecting at least one episode of ventricular tachyarrhythmia in a subject using an implantable medical device (IMD), sensing at least one heart sound signal for the subject using the IMD, the heart sound signal associated with mechanical vibration of a heart of the subject; initiating, in response to and during the detected episode of tachyarrhythmia, a measurement of hemodynamic stability of the ventricular tachyarrhythmia from the heart sound signal, and deeming whether the ventricular tachyarrhythmia is stable according to the measurement of hemodynamic stability. The measurement of hemodynamic stability is determined using linear prediction. | 1. A method comprising: detecting at least one episode of ventricular tachyarrhythmia in a subject using an implantable medical device (IMD);sensing at least one heart sound signal for the subject using the IMD, the heart sound signal associated with mechanical vibration of a heart of the subject;initiating, in response to and during the detected episode of tachyarrhythmia, a measurement of hemodynamic stability of the ventricular tachyarrhythmia from the heart sound signal, wherein the measurement of hemodynamic stability is determined using linear prediction; anddeeming whether the ventricular tachyarrhythmia is stable according to the measurement of hemodynamic stability. 2. The method of claim 1, wherein obtaining a measurement of hemodynamic stability includes extracting a proxy measurement of aortic pressure from the heart sound signal. 3. The method of claim 1, wherein obtaining a measurement of hemodynamic stability includes: sampling the heart sound signal to obtain a plurality of signal samples; anddetermining an autocorrelation function using the signal samples. 4. The method of claim 1, wherein obtaining a measurement of hemodynamic stability from the heart sound signal includes: establishing a baseline measurement for the measurement of hemodynamic stability; anddeeming the ventricular tachyarrhythmia to be unstable when a measured change from the baseline measurement of hemodynamic stability exceeds a predetermined threshold value. 5. The method of claim 4, including deeming the ventricular tachyarrhythmia to be unstable when a measured change from the baseline measurement of hemodynamic stability exceeds a predetermined threshold value in X of Y consecutive cardiac cycles, wherein X and Y are integers and Y>X. 6. The method of claim 5, wherein establishing a baseline measurement includes determining a central tendency of the measurement of hemodynamic stability. 7. The method of claim 1, wherein obtaining a measurement of hemodynamic stability from the heart sound signal includes: establishing a baseline measurement for the measurement of hemodynamic stability;determining a measure of variability of ventricular time intervals; anddeeming the ventricular tachyarrhythmia to be unstable when both a measured change from the baseline measurement of hemodynamic stability exceeds a predetermined stability measurement threshold value and the measure of variability of ventricular time intervals exceeds a predetermined variability threshold value. 8. The method of claim 1, including delaying delivery of a predetermined shock stimulus according to the measurement of hemodynamic stability. 9. The method of claim 1, including: initiating a predetermined anti-tachyarrhythmia pacing (ATP) regimen upon detecting the episode of ventricular tachyarrhythmia; andextending a duration of the ATP regimen according to the measurement of hemodynamic stability. 10. The method of claim 1, including correlating the sensed heart sound signal to a sens... | Cardiac Pacemakers, Inc. | A61B005/02
US07957783 | B2 | 2011-06-07 | Electrical lead for an electronic device such as an implantable device | 2009-06-04 | A lead for an electronic device which resists the induction of a current from an electromagnetic field external to said lead includes one or more pairs of adjacent segments of electrical wire, each of the pairs including a first segment of electrical wire and a second segment of electrical wire. The lead also includes one or more shielded RF chokes, wherein each of the shielded RF chokes is provided between the first segment of electrical wire and the second segment of electrical wire of a respective one of the one or more pairs of adjacent segments. Also, an implantable device that includes a generator for generating one or more electrical pulse and a lead as described for delivering the pulses to tissue within a patient's body. A method for making the described implantable device is also provided. | 1. An apparatus that may be implanted within a patient's body which resists the induction of a current from an electromagnetic field external to said apparatus, comprising: a generator for generating one or more electrical pulses;a lead for delivering the one or more electrical pulses to tissue within the patient's body, said lead including:one or more pairs of adjacent segments of electrical wire, each of said pairs including a first segment of electrical wire and a second segment of electrical wire; andone or more shielded RF chokes, each of said shielded RF chokes being provided between the first segment of electrical wire and the second segment of electrical wire of a respective one of said one or more pairs of adjacent segments and having a first end operatively coupled to the first segment of electrical wire and a second end operatively coupled to the second segment of electrical wire of the respective one of said one or more pairs of adjacent segments, wherein said one or more pairs of adjacent segments comprises a plurality of pairs of adjacent segments of electrical wire and wherein said one or more shielded RF chokes comprises a plurality of shielded RF chokes. 2. The apparatus according to claim 1, wherein said electromagnetic field includes electromagnetic energy having a first wavelength, and wherein the first segment of electrical wire and the second segment of electrical wire in each of the plurality of pairs of adjacent segments of electrical wire each has a length of no more than about a predetermined percentage of said first wavelength. 3. The apparatus according to claim 2, wherein said predetermined percentage of said first wavelength is twenty five percent. 4. The apparatus according to claim 3, wherein said lead includes a first portion of electrical wire located within said tissue having a length equal to no more than about fifty percent of said first wavelength. 5. The apparatus according to claim 1, wherein each of the adjacent segments of electrical wire is a single conductor electrical wire. 6. The apparatus according to claim 1, wherein each of the adjacent segments of electrical wire is a multiple conductor electrical wire. 7. The apparatus according to claim 1, further comprising a layer of insulating material covering at least a portion of said lead and at least a portion of said generator. 8. The apparatus according to claim 7, wherein said layer of insulating material covers substantially all of said lead except for portions of said lead that must be exposed for proper operation of said apparatus. 9. An apparatus that may be implanted within a patient's body which resists the induction of a current from an electromagnetic field external to said apparatus, comprising: a generator for generating one or more electrical pulses;a lead for delivering the one or more electrical pulses to tissue within the patient's body, said lead including:one or more pairs of adjacent segments of electrical wire, each of said p... | Boston Scientific Neuromedulation Corporation Surgivision, Inc. | A61B005/04
US07945456 | B2 | 2011-05-17 | Documenting remote engagements | 2008-02-27 | A service provider is prompted to provide a summary of an engagement with a consumer, an audible summary provided by the provider is recorded and, and the audible summary is associated with a record of the engagement. | 1. A computer-implemented method, comprising: receiving a request from a consumer of medical services to consult with a medical provider;establishing a telephone-based communication over a telephone-based communication channel between the consumer of medical services and a provider of medical services based on the received request;establishing, by a computer, a record based on the communication;prompting the provider to furnish an audio message related to the telephone-based communication after the consumer disconnects from the telephone-based communication;receiving and recording, by the computer, the audio message from the provider after the consumer disconnects from the telephone-based communication without disconnecting the provider from the telephone-based communication channel established between the consumer and the provider; andassociating, by the computer, the audio message with the established record. 2. The method of claim 1, wherein prompting the provider comprises prompting the provider to enter responses to questions using a numeric keypad of a telephone; and the method further comprises:storing the entered numeric responses with the record of the engagement. 3. The method of claim 1, further comprising: retrieving the record of the engagement from a memory;replaying the audio message; andprompting the service provider to provide a written summary of the engagement. 4. The method of claim 1 wherein prompting the service provider comprises using an interactive voice response system to prompt the provider. 5. The method of claim 1 in which the service provider is a physician and the consumer is a patient. 6. A computer implemented method, comprising prompting a medical service provider to provide an audio message related to a telephone based engagement with a consumer of medical services;recording the audio message provided by the medical service provider;associating the audio message with a record of the engagement.retrieving the record of the engagement from a memory;replaying the audio message; andprompting the medical service provider to generate a written summary of the engagement via a computer. 7. A computer-implemented method for determining compensation for a medical service provider, the method comprising: establishing a communication between a consumer of medical services and a provider of medical services based on a request from the consumer of medical services to consult with a medical provider;establishing a record based on the established communication;prompting the service provider to furnish an electronic summary of the communication; andproviding a payment to the service provider after receipt of the electronic summary. 8. The computer-implemented method of claim 7, wherein providing payment to the service provider comprises: determining if an electronic summary of the communication has been completed; andif the electronic summary has been completed releasing payment to the service provider for the services. ... | American Well Corporation | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US08257272 | B2 | 2012-09-04 | Implantable sensor for measuring physiologic information | 2011-02-01 | An implantable sensor is provided that includes a piezopolymer sensor element including a body having a plurality of layers of a piezopolymer, and an attachment device configured to hold the piezopolymer sensor element in direct contact with at least one of a bodily fluid and bodily tissue such that the piezopolymer sensor element is configured to bend in response to motion of the at least one of bodily fluid and bodily tissue. A pair of electrodes is attached to the piezopolymer sensor element and the electrodes are configured to collect an electrical charge that is generated within the piezopolymer sensor element due to the bending of the piezopolymer sensor element. | 1. An implantable lead comprising: a first piezopolymer sensor element;a second piezopolymer sensor element; andan attachment device configured to hold the first piezopolymer sensor element and the second piezopolymer sensor element in direct contact with at least one of a bodily fluid and bodily tissue such that the piezopolymer sensor elements bend in response to motion of the at least one of bodily fluid and bodily tissue;wherein the first piezopolymer sensor element and the second piezopolymer sensor element are arranged to at least partially overlap each other and are held in approximately perpendicular orientation with respect to one another. 2. The lead of claim 1 wherein the first and second sensor elements comprise polyvinylidene fluoride (PVDF). 3. The lead of claim 1 wherein the attachment device comprises a generally cylindrical cable, and the first and second sensor elements are on an exterior surface of the cable such that the first and second sensor elements at least partially surrounds a portion of the cable. 4. The lead of claim 1 wherein the attachment device comprises a patch having a body portion configured to be attached to a surface of the heart. 5. The lead of claim 1 wherein the attachment device is configured for endocardial placement. 6. The lead of claim 1 wherein the attachment device is configured for pericardial placement. 7. The lead of claim 1 further comprising two pairs of electrodes, one pair attached to each of the first and second sensor elements, the electrodes configured to collect an electrical charge generated within the sensor elements due to the bending of the sensor elements. 8. The lead of claim 1 wherein the first and second sensor elements and the attachment device lie in a common plane. 9. An implantable lead comprising: a first piezopolymer sensor element;a second piezopolymer sensor element; andan attachment device configured to hold the first piezopolymer sensor element and the second piezopolymer sensor element in direct contact with at least one of a bodily fluid and bodily tissue such that the piezopolymer sensor elements bend in response to motion of the at least one of bodily fluid and bodily tissue;wherein the first piezopolymer sensor element and the second piezopolymer sensor element are held in approximately perpendicular orientation with respect to one another;wherein the attachment device comprises a cable defining an at least partially loop-shaped portion having an open space bounded thereby; and the first and second sensor elements are held within the open space. 10. The lead of claim 9 wherein the loop-shaped portion defines an at least partially diamond-shaped portion. 11. The lead of claim 10 wherein the diamond-shaped portion comprises one or more cable portions that lie in a common plane with the sensor element and one or more cable portions that bow outwardly from the common plane. 12. The lead of claim 9 further comprising a plurality of supports extending from the ... | Pacesetter, Inc. | A61B005/0295
US07945304 | B2 | 2011-05-17 | Ultrasound within MRI scanners for guidance of MRI pulse sequences | 2001-11-20 | A process and system acquiring both ultrasound and MRI data from a patient in an MRI scanner and using one to automatically guide or improve the other. The ultrasound data can track the motion of an organ such as a cardiac artery or the diaphragm and the resulting motion information can automatically guide MRI data acquisition to keep the imaging volume at the organ or interest, or can be used to accept or reject MRI data in the process of forming an MRI image, based on whether the organ of interest was within an acceptable range of positions. Conversely, the MRI unit can guide the ultrasound data acquisition so it is properly timed or otherwise controlled in accordance with MRI data acquisition. | 1. A process of acquiring ultrasound data and magnetic resonance imaging (MRI) data of a patient in an MRI scanner such that a magnetic resonance (MR) imaging volume moves relative to the MRI scanner, under guidance from said ultrasound data, to track movement of an organ of the patient that moves relative to the MRI scanner to a plurality of different positions relative to the MRI scanner, comprising: acquiring ultrasound data indicative of a changing position of a body organ of the patient while the organ is moving relative to the MRI scanner;acquiring MRI data from an MR imaging volume in the patient substantially concurrently with the acquisition of the ultrasound data;computer-processing the ultrasound data to generate guidance information causing said MR imaging volume to move relative to the MRI scanner to track the changing position of said body organ such that the MR imaging volume and said moving organ undergo matching movement with respect to the MRI scanner during said MRI data acquisition;andcomputer-processing the MRI data to generate images of at least said body organ with inherent correction for motion artifacts due to said movement of the organ relative to the MRI scanner. 2. A process as in claim 1 in which said acquiring of ultrasound data comprises using at least one ultrasound transducer acoustically coupled with the patient. 3. A process as in claim 2 in which said acquiring of ultrasound data comprises using a plurality of ultrasound transducer assemblies spaced from each other along the patient and acoustically coupled to the patient to transmit ultrasound energy along beam paths at an angle to each other. 4. A process as in claim 2 in which the acquiring of ultrasound data comprises forming ultrasound images of patient anatomy. 5. A process as in claim 4 in which said forming of ultrasound images comprises forming said ultrasound images at a rate exceeding 10 images per second. 6. A process as in claim 4 including using edge detection processing to identify and track motion in space of said moving patient organ at times related to said acquiring said MRI data. 7. A process as in claim 4 in which said acquiring of said MRI data includes forming an MR image. 8. A process as in claim 7 including selectively displaying said ultrasound and MR images. 9. A process as in claim 8 including displaying a fused image related to a combination of said ultrasound and MR images. 10. A process as in claim 4 comprising using said ultrasound data to alter, during said acquisition of said MRI data, at least one of a magnetic gradient amplitude, a magnetic gradient phase, and a radio frequency pulse used in said acquisition of said MRI data. | - | A61B005/055
US07941207 | B2 | 2011-05-10 | Cardiac monitoring | 2007-02-12 | Systems and techniques for monitoring cardiac activity. In one aspect, a method includes collecting information describing the variability in heart rate over a series of beats, designating variability at a lower end of physiological values as being largely irrelevant to atrial fibrillation, designating variability in a midrange of physiological values as being indicative of atrial fibrillation, designating variability in an upper range of physiological values as being negatively indicative of atrial fibrillation, and determining a relevance of the variability described in the collection to atrial fibrillation. | 1. A device, comprising: a beat detector to identify a beat-to-beat timing of cardiac activity;a ventricular beat detector to identify ventricular beats in the cardiac activity;variability determination logic to determine a variability in the beat-to-beat timing of a collection of beats;relevance determination logic to identify a relevance of the variability in the beat-to-beat timing to at least one of atrial fibrillation and atrial flutter; andan event generator to generate an event when the variability in the beat-to-beat timing is identified as relevant to the at least one of atrial fibrillation and atrial flutter in light of the variability in the beat-to-beat timing caused by ventricular beats identified by the ventricular beat detector. 2. The device of claim 1, wherein the relevance determination logic is to accommodate variability in the beat-to-beat timing caused by ventricular beats by weighting ventricular beats as being negatively indicative of the one of atrial fibrillation and atrial flutter. 3. The device of claim 1, wherein the variability determination logic is to compare times between R-waves in three successive QRS complexes to determine the variability in the beat-to-beat timing. 4. The device of claim 1, wherein: the variability determination logic is to represent the variability in the beat-to-beat timing as a factor that is lowest when a first time between beats is close to a second time between beats; andthe first time immediately proceeds the second time. 5. The device of claim 4, wherein the variability determination logic is to represent the variability in the beat-to-beat timing as a factor that increases non-linearly when the absolute difference between the first time the second time grows. 6. The device of claim 4, wherein the variability determination logic is to represent the variability in the beat-to-beat timing as a factor that increases more rapidly when the first time grows less than the second time than when the first time grows greater than the second time. 7. The device of claim 1, wherein the event generator is to generate an event by performing operations comprising: collecting data associated with the collection of beats; andtransmitting the data associated with the collection of beats to a remote receiver. 8. The device of claim 1, wherein the relevance determination logic comprises weighting logic to: weight variability at a lower end of physiological values as being substantially irrelevant to the one of atrial fibrillation and atrial flutter;weight variability in a midrange of physiological values as being positively indicative of the one of atrial fibrillation and atrial flutter; andweight variability in an upper range of physiological values as being negatively indicative of the one of atrial fibrillation and atrial flutter. 9. The device of claim 8, wherein the weighting logic is also to weight a beat identified as a ventricular beat as being negatively indicative of the one of atrial fib... | CardioNet, Inc. | A61B005/04
US07963931 | B2 | 2011-06-21 | Methods and devices of multi-functional operating system for care-taking machine | 2009-03-11 | The methods and devices of a multi-functional operating interface for a care-taking machine. The multi-functional operating interface of the care-taking machine includes: muscle stretch sensor, multiplexer amplifying wave filter, analog-to-digital signal converter, image processing unit, and muscle pattern database and control bus. Aided by the present muscle template training method and potential image data collecting and processing method, it can become an interface that offers the user of a care-taking machine instantaneous and unlimited controllability of the machine, enabling such people, either sick or handicapped, to enjoy more convenience and better life. | 1. A method for multi-functional control interface for a care-taking machine, which is used to construct an interface that provides a user with control of an equipment, comprising the following steps: detecting operation movements, for detecting movements on movable parts of a body of the user and generating corresponding electric signals by a muscle stretch sensor;amplifying and filtering electric signals by the care-taking machine and then output analog signals;converting said analog signals into digital electric signals by the care-taking machine;processing images by the care-taking machine, by arranging and converting said digital electric signal into gray level potential images, processing said potential images to highlight characteristics of said potential images, and grouping them according to the highlighted characteristics;outputting control commands by the care-taking machine, for providing and tallying a set of characteristics for tallying with those highlighted characteristics by a muscle template database, and presenting corresponding commands; andtransmitting control commands by the care-taking machine, for receiving the commands by a control bus, and transmitting said commands to related interface, signal converter, or control unit so as to control a specific equipment. 2. The method for multi-functional control interface for care-taking machine as claimed in claim 1, wherein said step of detecting operation movements is performed via a the muscle stretch sensor by detecting deformations produced by the movable parts of the body and generating corresponding electric signals. | Industrial Technology Research Institute | A61B005/103 A61B005/117
US08204786 | B2 | 2012-06-19 | Physiological and environmental monitoring systems and methods | 2011-01-06 | Systems and methods for monitoring various physiological and environmental factors, as well as systems and methods for using this information for a plurality of useful purposes, are provided. Real-time, noninvasive health and environmental monitors include a plurality of compact sensors integrated within small, low-profile devices. Physiological and environmental data is collected and wirelessly transmitted into a wireless network, where the data is stored and/or processed. This information is then used to support a variety of useful methods, such as clinical trials, marketing studies, biofeedback, entertainment, and others. | 1. A method of delivering targeted advertising to a subject, comprising: receiving physiological information from a monitoring device associated with a subject, wherein the monitoring device includes at least one physiological sensor and a motion sensor;receiving subject motion information from the motion sensor;processing the subject motion information via a processor to identify footstep information, body motion information, head motion information, jaw motion information, or swallowing information;analyzing the physiological information via the processor in context with the identified type of motion to identify a health issue associated with the subject;selecting an advertisement for delivery to the subject based upon the health issue; anddelivering the selected advertisement to the subject. 2. The method of claim 1, wherein selecting an advertisement comprises selecting an advertisement for a product or service related to the health issue. 3. The method of claim 1, wherein the monitoring device is attached to a body of the subject. 4. The method of claim 1, wherein the monitoring device is attached to an ear of the subject. 5. The method of claim 1, wherein delivering the selected advertisement comprises delivering the selected advertisement to the subject via one or more of the following delivery channels: email, postal mail, television, radio, newspaper, magazine, the internet, and outdoor advertising. 6. A system for delivering targeted advertising, comprising: a plurality of portable monitoring devices, each portable monitoring device associated with a respective subject, each portable monitoring device comprising at least one physiological sensor configured to detect and/or measure physiological information from a respective subject, a motion sensor configured to detect motion of the subject, and a processor, wherein the processor is configured to process subject motion information from the motion sensor to identify footstep information, body motion information, head motion information, jaw motion information, or swallowing information, and wherein the processor is configured to analyze physiological information from the at least one physiological sensor in context with an identified type of motion to identify a health issue associated with the subject; anda remotely located advertisement selection device that receives health issue information from the portable monitoring devices, selects advertisements based upon the received health issue information, and delivers selected advertisements to the respective subjects. 7. The system of claim 6, wherein the advertisement selection device receives health issue information from each portable monitoring device via a communication device associated with each portable monitoring device. 8. The system of claim 7, wherein the communication device is selected from the group consisting of PDAs, cell phones, laptop computers, and other portable telemetric devices. 9. The system of claim 7, whe... | Valencell, Inc. | A61B005/00
US07967758 | B2 | 2011-06-28 | Devices and methods for detecting and treating inadequate tissue perfusion | 2006-07-12 | Devices and methods for detecting inadequate tissue perfusion by measuring a parameter other than heart rate such as vascular blood pressure, intracardiac blood pressure, vascular blood flow or tissue perfusion, in addition to or as a substitute for heart rate. Such devices and methods improve the accuracy of determining when and to what degree therapy should be administered to treat inadequate tissue perfusion, such as pre-syncope, syncope, or orthostatic hypotension, particularly in the absence of abnormal cardiac function. | 1. A method for treating a patient, comprising: detecting, using a first medical device, a heart rate of the patient as an indicator of inadequate tissue perfusion;detecting, using a second medical device, at least one other indicator of inadequate tissue perfusion; anddelivering, the first medical device, a stimulus to increase heart rate as a function of both the detected heart rate and the at least one other indicator of inadequate tissue perfusion. 2. A method as in claim 1, wherein delivering the stimulus to increase heart rate comprises delivering electrical impulses to a heart of the patient. 3. A method as in claim 2, wherein delivering electrical impulses to the heart of the patient comprises providing a pacemaker. 4. A method as in claim 1, wherein detecting the at least one other indicator of inadequate tissue perfusion comprises detecting an intravascular parameter. 5. A method as in claim 4, wherein detecting the intravascular parameter comprises detecting blood pressure. 6. A method as in claim 5, wherein detecting blood pressure comprises contacting blood disposed within a blood vessel of the patient with a catheter. 7. A method for treating a patient, comprising: detecting, using a first medical device, a heart rate of the patient as an indicator of inadequate tissue perfusion;positioning a sensor of a second medical device proximate a blood vessel of the patient;measuring an intravascular parameter using the sensor, anddelivering, using the first medical device, a stimulus to increase heart rate as a function of both the detected heart rate and the measured intravascular parameter. 8. A method as in claim 7, wherein detecting the intravascular parameter comprises detecting blood pressure. 9. A method as in claim 8, wherein detecting blood pressure comprises contacting blood disposed within the blood vessel with a catheter. 10. A method for treating a patient, comprising: detecting, using a first medical device, a heart rate of the patient as an indicator of inadequate tissue perfusion;detecting, using a second medical device, a blood pressure of the patient; anddelivering, using the first medical device, a stimulus to increase heart rate as a function of both the detected heart rate and the blood pressure of the patient. 11. The method of claim 1, wherein the first medical device is a separate device from the second medical device. 12. The method of claim 1, wherein the first medical device is the same device as the second medical device. | Data Sciences International, Inc. | A61B005/02 A61N001/00
US07956603 | B2 | 2011-06-07 | Sensor inductors, sensors for monitoring movements and positioning, apparatus, systems and methods therefore | 2008-06-16 | A generally planar shaped inductor is disclosed that is particularly adaptable for use in motion or position sensors. One inductor can function as a signal input unit and another as a pick up unit in an arrangement wherein both inductors are placed in a generally parallel juxtaposition for flux flow there between. A movable armature is located between the inductors to control the amount of flux transmission between inductors. The position of the armature relative to the inductors controls the output signal generated by the pickup inductor that are adapted to be converted into indications of displacements. | 1. A sensor system for monitoring bodily functions comprising at least one sensor having attachment unit to be attached to external body skin at two separate locations for providing indications of skin deformation between the two locations, wherein the sensor comprises two generally planar coils juxtapositioned with a spacing in between wherein one of the coils functions as an excitation coil and the other as the output coil, and one of the coils is adapted to be connected to the skin, a metallic shield positioned within the coil spacing being adapted to be attached to the skin, wherein the shield and the coils are movable with respect to each other to follow deformations of the skin between the attachments of the coil and shield to provide indications of movement in the form of changes of the amount of magnetic flux transmitted between the coils. 2. A sensor system as defined in claim 1 wherein the coil and shield are adapted to be attached to the skin locations external of the chest and wherein the indications provided by the sensor are a function of expansions and contractions of the chest. 3. A sensor system as defined in claim 1 wherein the sensor is adapted to be attached to the skin locations external of the abdomen and wherein the indications provided by the sensor are a function of expansions and contractions of the abdomen. | Medility LLC | G01B007/14 A61B005/103
US08634614 | B2 | 2014-01-21 | System and method for volumetric analysis of medical images | 2009-06-02 | The invention relates to a system and a method for analysis and evaluation of at least one anatomical structure in a medical image, said medical image representing at least a part of a patient or person, said system comprising image analyzing means for calculating the volume and/or providing the shape of the at least one anatomical structure, and statistical analyzing means for providing a confidence interval of the volume and/or the shape of the analyzed anatomical structure(s) based on individual data of the patient or person, thereby quantifying the normality of said anatomical structure(s). By the present system and method a detailed and trustworthy confidence interval can be provided in connection with analysis and evaluation of a medical image of a patient, thereby illustrating the expected normality of the patient. The invention can be applied to images acquired from magnetic resonance imaging (MRI). One example of use is for diagnosing diseases and/or initial stages of diseases. | 1. A method for analyzing and evaluating of at least one anatomical structure in a medical image, said medical image representing at least a part of a patient or person, said method comprising the steps of: a. mapping and registering the medical image to a reference image thereby segmenting the medical image, wherein the reference image is a segmented medical image based on and averaged over a plurality of medical images obtained from normal persons,b. analysing the segmented medical image for calculating the volume and/or providing the shape of the said anatomical structure(s), andc. providing a confidence interval of the volume and/or the shape of the analysed anatomical structure(s) based on individual data of the patient or person by analysing a database comprising information of normal volume and/or shape of said anatomical structure(s) corresponding to said individual data thereby quantifying the normality of said anatomical structure(s). 2. The method according to claim 1 wherein said individual data comprises information about at least one of: gender, age, height, weight, genotype, ethnicity, habits, handedness, cognitive measures and/or nationality of the patient or person. 3. The method according to claim 1, wherein the medical image is normalized prior to analysis of said medical image. 4. The method according to claim 1, wherein the confidence interval of the volume and/or the shape is provided as: at least one numerical value,and/ora graphical representation in the analysed medical image and/or in the reference image. 5. The method according to claim 1, wherein the medical image comprises at least a part of the brain of the patient or person. 6. The method according to claim 1, wherein the medical mage is a magnetic resonance image, a CT image, or an ultrasound image. 7. The method according to claim 1, wherein the analysed anatomical structure(s) is any segmentable brain structure selected from the group of hippocampus, amygdala, putamen, pons, thalamus, global pallidus, brainstem or cerebellum. 8. The method according to claim 1 for diagnosing a disease, and/or diagnosing the initial stage of a disease, said diagnosis based on the normality and/or the confidence interval of the analysed volume and/or shape of the at least one anatomical structure. 9. The method according to claim 8, wherein the disease is selected from the group of depression, Alzheimer, epilepsy, post-traumatic stress disorder, schizophrenia, anxiety, or Parkinson's disease. 10. A computer program product having a non-transitory computer readable medium, said computer program product providing a system for analysis and evaluation of at least one anatomical structure in a medical image, said medical image representing at least a part of a patient or person, said computer program product comprising means for carrying out all the steps of the method according to claim 1. | Brainreader ApS | G06K009/00
US08652070 | B2 | 2014-02-18 | Systems, devices, and methods for interpreting movement | 2011-01-14 | According to some exemplary embodiments, a method or system can involve associating a plurality of biokinetographic comparison results with a first specific dysfunction from a group of specific dysfunctions, each of the biokinetographic comparison results obtained from a comparison of a biokinetographic value to a standard for a corresponding biokinetographic variable. | 1. A computer system for performing a diagnosis of an individual's specific dysfunction based on characteristics of the individual's gait, the system comprising: an input component for receiving biokinetographic values for initiation of gait movement, base of gait, cadence, gait turnaround and arm swing during gait movement for an individual from a plurality of sensors;a memory component for storing a gait characteristic database comprised of biokinetographic standard values for initiation of gait movement, base of gait, cadence, gait turnaround and arm swing during gait movement and at least two corresponding dysfunctions;a processor under programmed control for receiving the biokinetographic data for the individual's gait characteristics and comparing the biokinetographic data for the individual's gait characteristics to the gait characteristic standard database and determining one or more diagnoses for the individual's specific dysfunction, wherein the one or more diagnoses is a diagnosis of a neurological dysfunction when the biokinetographic values include (1) delayed initiation of movement with freezing or short steps and a gravitational unit difference of greater than one in heel strike amplitudes, (2) delayed initiation of movement with freezing or short steps and a Fast Fourier Transformation (FFT) power less than 0.5, (3) a highly variable cadence with a gravitational unit difference of greater than one in heel strike amplitudes, or (4) a highly variable cadence with a Fast Fourier Transformation (FFT) power less than 0.5; andan output component for presenting the one or more diagnoses for the individual's specific dysfunction. 2. The system of claim 1, wherein the diagnoses comprise one or more of a neurological dysfunction, a psychiatric condition, an endocrine dysfunction, orthopedic dysfunction and frailty. 3. The system of claim 1, wherein the initiation of gait movement is the amount of time before movement begins after a start signal. 4. The system of claim 1, wherein the base of gait is a pelvic tilt. 5. The system of claim 1, wherein cadence is one or more of the number of steps per unit of time, step rate variability, magnitude of the Fast Fourier Transformation or magnitude of gait accelerations. 6. The system of claim 1, wherein the gait turnaround is the number of steps to perform a turnaround. 7. The system of claim 1, wherein the arm swing during gait movement is one or more of arm swing amplitude or Fast Fourier power. 8. A computer-implemented method performing a diagnosis of an individual's specific dysfunction based on characteristics of an individual's gait, the method comprising: accessing at least one computerized knowledge base containing standard values for multiple characteristics of a gait cycle, the multiple characteristics comprising biokinetographic values for initiation of gait movement, base of gait, cadence, gait turnaround and arm swing during gait movement;obtaining biokinetographic values for in... | University of Virginia Patent Foundation | A61B005/103
US08913078 | B2 | 2014-12-16 | Diagnosis assisting system, computer readable recording medium having diagnosis assisting program recorded thereon, and diagnosis assisting method | 2010-09-14 | A plurality of images for observation in different observation formats and thumbnail images corresponding to each image for observation are generated from volume data obtained for the same subject on different imaging dates/times. A plurality of image windows, which can be switched between an active state and an inactive state, are arranged on a diagnosis screen, and images for observation generated from volume data obtained on specified imaging dates are displayed in each image window. In one embodiment, thumbnail images having the same observation format as an image for observation are displayed in the vicinity of the image for observation within an image window in the active state. | 1. A diagnosis assisting system, comprising: volume data storage means, for storing a plurality of sets of volume data, obtained by imaging at least one subject on different imaging dates/times, in a predetermined storage device, correlated with subject identifying data and imaging date/time data;volume data selecting means, for selecting at least two sets of volume data correlated with subject identifying data that represents a specified subject, from among the sets of volume data stored in the storage device;image generating means, for generating a plurality of images for observation in different observation formats that represent the specified subject for each set of volume data selected by the volume data selecting means, for generating a plurality of thumbnail images corresponding to each of the images for observation, and for correlating the images for observation, the thumbnail images, and the imaging date/time data of the set of volume data corresponding thereto;image window control means, for arranging a plurality of image windows, each of the plurality of image windows configured to be separately switchable between an active state and an inactive state, on a screen, and for controlling the switching from the active state and the inactive state; anddisplay control means, for selecting images for observation which are correlated with imaging date/time data that represent a specified imaging date/time from among the images for observation generated by the image generating means, for displaying the selected images for observation in the image windows, for selecting thumbnail images having the same observation format as that of the image for observation which is displayed in the active image window from among the thumbnail images generated by the image generating means, and for displaying the selected thumbnail images within the active image window,wherein active and inactive windows of the plurality of windows are shown simultaneously, and the image window control means controls the active window, andeach of the plurality of windows is correlated with a corresponding set of thumbnail images stored in memory, said set of thumbnails corresponding to an observation format for each respective image window of the plurality of windows, and only the set of thumbnails corresponding to a window among the plurality of windows set to the active state is displayed, and windows having the inactive state display a medical image according to respective observation format, without its respective corresponding set of thumbnails. 2. A diagnosis assisting system as defined in claim 1, wherein: the image generating means generates an image for observation and a thumbnail image corresponding to a specified observation format, when an operation that requests a change in the observation format of an image for observation displayed in an active image window; andthe display control means updates the display to display the generated image for observation and the ge... | FUJIFILM Corporation | G09G005/00 G06F019/00
US08355753 | B2 | 2013-01-15 | Medication injection site and data collection system | 2010-04-22 | A medication injection site is provided that includes a medication port and an identification sensor. The medication injection site, when coupled to a medication container, can wirelessly transmit data characterizing the contents of the medication container to a remote data collection system. The housing of the medication injection site has a shape and size enabling it to be held by a first hand of a user while the user administers medication from the medication container via the medication port using his or her second hand. The housing can include a disposable sub-housing and a reusable sub-housing such that the reusable sub-housing is used by many patients. In some implementations, the medication injection site can be placed on an IV drip line. Related apparatus, systems, and kits are also disclosed. | 1. A system comprising: a housing comprising a reusable sub-housing and a disposable sub-housing, the reusable sub-housing being operatively coupled to the disposable sub-housing, wherein the reusable sub-housing is intended for use by a plurality of patients and the disposable sub-housing is intended for use by a single patient;a medication port extending from an external face of the housing and configured to be fluidically and directly coupled to a fluid outlet of a manually injectable medication container;an identification sensor disposed within the housing to generate identification information indicative of contents of the medication container when the fluid outlet of the medication container is being fluidically coupled to the medication port;a transmitter disposed within the housing and in communication with the identification sensor to wirelessly transmit the information generated by the identification sensor to a remote data collection system; anda power source disposed within the housing powering the identification sensor and the transmitter. 2. A system as in claim 1, further comprising a junction element at least partially extending within the housing forming a first fluid channel and a second fluid channel, the first fluid channel extending from a first end to a second end, the second fluid channel extending from a distal end and terminating at the first fluid channel at an intersection intermediate the first end and the second end. 3. A system as in claim 1, wherein the power source is disposed within the reusable sub-housing. 4. A system as in claim 3, wherein the identification sensor is disposed within the reusable sub-housing. 5. A system as in claim 4, wherein the transmitter is disposed within the reusable sub-housing. 6. A system as in claim 1 wherein the transmitter is a transceiver. 7. A system as in claim 1, wherein the disposable sub-housing is part of a kit including a sterile pouch enveloping the disposable sub-housing. 8. A system as in claim 2, wherein the first fluid channel and the second fluid channel are disposed within the disposable sub-housing. 9. A system as in claim 2, further comprising: a self-contained fluid delivery sensor disposed within the reusable sub-housing and in communication with the transmitter to characterize fluid flow through one or more of the first fluid channel and the second fluid channel, wherein the transmitter further wirelessly transmits data characterizing fluid delivery to the remote data collection system. 10. A system as in claim 1, further comprising: a processor disposed within the reusable sub-housing that is coupled to one or more of the identification sensor, the transmitter, and the power source. 11. A system as in claim 1, wherein the power source is self-contained, coupled to a remote power source, or is a rechargeable battery that can be detachably coupled to a remote charging source. 12. A system as in claim 1, wherein the housing has a shape and size enabling... | CRISI Medical Systems, Inc. | H04W088/02
US08682432 | B2 | 2014-03-25 | Cardiac-event processor and heart treatment device | 2010-09-17 | Provided is a cardiac-event processor (1) including a storing unit (11) that stores a time history of cardiac events, including heartbeats, in association with the presence/absence of nerve stimulation; a heart rate calculator (12) that calculates a heart rate for when each heartbeat is detected, on the basis of the time history of the cardiac events stored in the storing unit (11); and a frequency data generator (13) that separately adds up frequencies of the heart rates calculated by the heart rate calculator (12) for the cases with and without the nerve stimulation. | 1. A cardiac-event processor comprising: a storing unit configured to store a time history of cardiac events, including heartbeats, in association with a presence/absence of nerve stimulation;a heart rate calculator configured to calculate a heart rate for when each heartbeat is detected, on a basis of the time history of the cardiac events stored in the storing unit; anda frequency data generator configured to separately add up frequencies of the heart rates calculated by the heart rate calculator for cases with and without the nerve stimulation,wherein the presence of nerve stimulation is given as a tag to the time history, andwherein the frequency data generator is configured to add up frequencies of the heart rates separately on the basis of existence or non-existence of the tags. 2. The cardiac-event processor according to claim 1, wherein the frequency data generator is configured to generate a histogram that shows a frequency distribution of the heart rates. 3. The cardiac-event processor according to claim 1, further comprising a display configured to separately display the frequencies of the heart rates with the nerve stimulation and the frequencies of the heart rates without the nerve stimulation, which are added up by the frequency data generator. 4. The cardiac-event processor according to claim 3, wherein the frequency data generator is configured to add the frequencies of the heart rates with the nerve stimulation to the frequencies of the heart rates without the nerve stimulation to generate a histogram showing a frequency distribution of overall heart rates, andthe display is configured to display the histogram generated by the frequency data generator in such a way that the frequencies of the heart rates are displayed with different display appearances for cases with and without the nerve stimulation. 5. The cardiac-event processor according to claim 1, wherein the frequency data generator is configured to add the frequencies of the heart rates with the nerve stimulation to the frequencies of the heart rates without the nerve stimulation, and normalizes the frequencies of the heart rates with the nerve stimulation and the frequencies of the heart rates without the nerve stimulation by dividing them by the sum of the frequencies. 6. The cardiac-event processor according to claim 1, wherein, in a case where the nerve stimulation is supplied, the storing unit is configured to store the time history of the heartbeats in association with an intensity of the nerve stimulation, andthe frequency data generator is configured to add up the frequencies of the heart rates with the nerve stimulation by sorting them according to a magnitude of the intensities of the nerve stimulation. 7. The cardiac-event processor according to claim 1, wherein, in the case where the nerve stimulation is supplied, the storing unit is configured to store the nerve stimulation in association with a heartbeat detected after a predetermined delay time from t... | Olympus Corporation | A61N001/08
US08414514 | B2 | 2013-04-09 | Wearable ambulatory traction system | 2009-11-16 | A wearable ambulatory traction system for a user. The system can include a traction device with a traction head, a traction body, and a centralizer; a wearable fabric garment; and a pouch secured to the fabric garment. The traction device can be insertable into the pouch. A customizable wearable ambulatory traction system for a user that can include a customizable traction device with a shearable traction head and body which can be cut to a custom size by the user. | 1. A wearable ambulatory traction system for a user, the system comprising: a. a traction device comprising:(i) a traction head;(ii) a traction body; and(iii) a centralizer extending from the traction head and engaging the traction body, providing an engagement between the traction head and the traction body;b. a fabric garment having a front portion and a back portion; andc. a pouch secured to the back portion of the fabric garment, wherein the traction body is insertable into the pouch;d. a left tube axially aligned with the traction head;e. a right tube axially aligned with the traction head; andf. a central tube disposed transverse to the left and right tubes; and wherein the central tube engages the centralizer. 2. The system of claim 1, wherein the traction device is engaged with a flat surface such that a pressure is applied to the traction device is configured to apply pressure to a user's back, and wherein the traction device is configured to provide a therapeutic pressure to the user's back, the user's neck, or combinations thereof. 3. The system of claim 2, wherein the pressure is applied to the traction device for a predetermined time interval. 4. The system of claim 2, wherein the pressure is a variable user-specified pressure, and wherein the user can vary the pressure using a weight of a portion of the user's body. 5. The system of claim 1, wherein the engagement between the traction head and the traction body is at a variable angle. 6. The system of claim 1, wherein the traction head and the traction body are both cylindrically shaped and have identical diameters. 7. The system of claim 1, wherein the traction head has a diameter different from a diameter of the traction body. 8. The system of claim 1, wherein the fabric garment is a member of the group consisting of: a shirt, a vest, a cardigan, a backpack, a v-neck shirt, a hoodie, a jacket, a poncho, a coat, a sweater, or a robe. 9. The system of claim 1, wherein the pouch is sewn onto the fabric garment. 10. The system of claim 1, wherein the pouch is removably attached to the fabric garment. 11. The system of claim 1, wherein the traction device or the fabric garment further comprises a message portion. 12. The system of claim 1, wherein the traction device further comprises a pressure sensor and a time sensor, each in communication with an alarm, wherein the alarm provides an alert to the user: a. when the pressure is greater than a preset pressure limit;b. when a usage time is greater than a preset time limit; orc. combinations thereof. 13. The system of claim 1, wherein the traction head and the traction body have different indention load deflections. 14. The system of claim 1, wherein the pouch contains the traction body in a removable engagement. 15. The system of claim 1, wherein the traction head further comprises a support member for engaging the centralizer. | - | A61F005/00 A61F005/02 A61F005/37 A61G015/00
US08449479 | B2 | 2013-05-28 | Squeeze activated medical puncturing device | 2005-04-29 | The medical puncturing device (10) includes a housing (12) having a proximal end (18) and a distal end (16) and a skin puncturing assembly (14) within the housing. The skin puncturing assembly includes a movable carrier (50) and a skin puncturing element (52). The skin puncturing element is adapted to puncture the skin of a patient. The carrier is movable from a retracted position wherein a distal tip of the skin puncturing element is disposed within the housing, to a puncturing position wherein the distal tip is exposed to puncture the skin of the patient. The carrier is moved from the retracted position to the puncturing position upon radial deformation of the housing. An internal guide (42) may be provided within the housing for guiding the movement of the carrier from the retracted position to the puncturing position. | 1. A medical puncturing device, comprising: a deformable housing, wherein at least a portion of the deformable housing includes a cross-section being continuous in a circumferential direction and perpendicular to a longitudinal axis of the housing, wherein a portion of the cross-section includes at least one point of deformation; anda skin puncturing element disposed within the housing, and maintained in a retracted first position within the deformable housing wherein a tip of the skin puncturing element is disposed within the deformable housing, the skin puncturing element being maintained in the first position against a biasing force prior to deformation of the deformable housing through an interference engagement between the deformable housing and the skin puncturing element,wherein, upon deformation of the housing at the at least one point of deformation, the skin puncturing element is movable from the first position to a second position causing the tip to be exposed. 2. The medical puncturing device of claim 1, further comprising: a carrier configured to move within the housing,wherein the skin puncturing element is mounted to the carrier and moves in concert with the carrier. 3. The medical puncturing device of claim 2, further comprising at least one guide within the housing for guiding the movement of the skin puncturing element from the first position to the second position. 4. The medical puncturing device of claim 3, wherein the at least one guide comprises: at least one guide element provided on one of the housing and carrier; andat least one guide track cooperating with the at least one guide element and formed in the other of the housing and carrier for guiding the movement of the carrier from the first position to the second position. 5. The medical puncturing device of claim 2, further comprising: at least one detent provided on one of the housing and carrier; andat least one engagement formed in the other of the housing and carrier cooperating with said at least one detent, wherein the at least one engagement secures the carrier in the first position until deformation of the housing. 6. The medical puncturing device of claim 2, further comprising a drive spring disposed within the housing and acting between the carrier and a proximal end of the housing, the drive spring adapted to move the carrier from the first position to the second position upon deformation of the housing. 7. The medical puncturing device of claim 6, wherein the drive spring is integral with at least one of the housing and the carrier. 8. The medical puncturing device of claim 2, further comprising a retraction spring disposed within the housing and acting between the carrier and the distal end of the housing, the retraction spring adapted to return the carrier to a position within the housing wherein the housing encompasses the skin puncturing element after the carrier reaches the second position. 9. The medical puncturing device of claim 2, wherein a... | Becton, Dickinson and Company | A61B005/00
US07968082 | B1 | 2011-06-28 | Evaluating mixtures of low molecular weight heparins by NMR | 2008-01-25 | Methods for analyzing mixtures of polysaccharides, for example heparin such as enoxaparin are described. In some instances, the mixtures are analyzed using 1D NMR and/or 2D NMR. | 1. A method of processing an enoxaparin preparation comprising: performing 2D-NMR to determine an amount of an epoxide, a reducing end 2-O sulfated uronic acid and a 1,6 anhydro mannosamine present in an enoxaparin preparation, andif the amount of the epoxide, the reducing end 2-O sulfated uronic acid and the 1,6 anhydro mannosamine is present in an amount within a preselected range, processing said preparation, wherein said processing includes one or more of selecting, accepting, processing into drug product, shipping, formulating, labeling, packaging or selling said preparation, andif the amount of the epoxide, the reducing end 2-O sulfated uronic acid or the 1,6 anhydro mannosamine is not present in an amount within the preselected range, discarding or withholding the preparation,wherein the preselected range of the epoxide corresponds to between 0.6 and 1.0 percent composition, wherein the preselected range of the reducing end 2-O sulfated uronic acid corresponds to between 0.8 and 1.4 percent composition, wherein the preselected range of the mannosamine corresponds to between 2.1 and 2.9 percent composition, andwherein the percent composition is calculated using the specific cross peak volumes for which all uronic acid residues have similar 1JCH couplings as do all glucosamine residues. 2. The method of claim 1, wherein the preparation has the reducing end 2-O sulfated uronic acid present in an amount that corresponds to between 0.9 and 1.3 percent composition, and the preparation is processed. 3. The method of claim 2, wherein the preparation has the reducing end 2-O sulfated uronic acid present in an amount that corresponds to between 0.9 and 1.2 percent composition, and the preparation is processed. 4. The method of claim 1, wherein the preparation has the epoxide present in an amount that corresponds to between 0.65 and 0.95 percent composition, and the preparation is processed. 5. The method of claim 4, wherein the preparation has the epoxide present in an amount that corresponds to between 0.7 and 0.9 percent composition, and the preparation is processed. 6. The method of claim 1, wherein the preparation has the 1,6 anhydro mannosamine present in an amount that corresponds to between 2.2 and 2.8 percent composition, and the preparation is processed. 7. The method of claim 6, wherein the preparation has the 1,6 anhydro mannosamine present in an amount that corresponds to between 2.3 and 2.7 percent composition, and the preparation is processed. 8. The method of claim 1, wherein the NMR-based determination further includes an amount of total 1,6 anhydrohexosamine and if the amount of the epoxide, the reducing end 2-O sulfated uronic acid, the 1,6 anhydro mannosamine, and total 1,6 anhydrohexosamine is present in an amount within the preselected range for the epoxide, the preselected range for the 2-O sulfated uronic acid, the preselected range for the 1,6 anhydro mannosamine, and a preselected range for the total 1,6 anhydrohexosa... | Momenta Pharmaceuticals, Inc. | A61B005/055 A61K031/727
US08290576 | B2 | 2012-10-16 | Modeling the electrical activity of the heart by a single dipole, concurrently estimating subject and measurement related conditions | 2008-11-10 | A method is provided which relates to the identification of the electrical activity in the heart in terms of a single dipole, which may be considered moving or non-moving, based on measurements from a body surface electrocardiogram (ECG), concurrently estimating subject and measurement related conditions. The invention further relates to computer program for performing said method, a data carrier containing said program as well a computer device for performing said program. A device for performing said method is also disclosed. | 1. A method for identifying the electrical activity in a heart of a mammalian body in terms of a single dipole from measurements from body surface potentials, concurrently estimating subject and measurement related conditions, comprising the steps of: a) measuring body surface potentials of a mammalian body by multiple leads via electrodes positioned on the mammalian body, thereby producing measurements of lead voltages at a given number of time points;b) providing prior values for a time course of the dipole location and of parameters reflecting the subject and measurement related conditions, wherein said parameters are the electrode positions on the mammalian body, conductivity parameters including the tissue conductivity and the decay exponent with respect to each electrode, together with said prior values for the time course of the baseline at each lead, which are collectively used as prior conditions, wherein the prior conditions are kept constant throughout an optimization process;c) providing initial values for a time course of the dipole momentum and initial values for all parameters of step 1b, said initial values being used as the starting values in the optimization process;d) providing regularizer functions for the properties of dipole momentum, dipole location, electrode positions, tissue conductivities, decay exponents and baselines, wherein said regularizer functions measure the conformance of said properties with an expectation about said properties, wherein said regularizer functions are based on said prior conditions;e) determining a lead voltage for each lead and each time point by calculating a scalar product of the estimated dipole momentum vector with a lead vector plus an estimated baseline value, wherein the lead vectors are determined by estimates of the dipole location, the positions of the electrodes, the tissue conductivities and the decay exponents;f) formulating an error function calculated using a deviation of the measured lead voltages of step 1a to the lead voltages determined in step 1e and the measurements calculated according to the regularizer functions of step 1d;g) minimizing the error function of step 1f by applying an optimization algorithm, varying the dipole momentum and optionally any values of step 1c, thereby optimizing the estimated values of said parameters of step 1b;h) repeating steps 1e to 1g until an optimization criterion is met; andi) outputting the time courses of the dipole momentum, dipole location, electrode positions, tissue conductivities, decay exponents and baselines. 2. The method of claim 1, wherein any of the conditions of step 1b are considered not to change with time. 3. The method according to claim 1, wherein the error function is:F=∑i=1NS⁢⁢∑λ=1NΛ⁢⁢wi,λ2⁡(mλ⁡(ti)-υλ⁡(ti))2+Θ⁡(q,p,p~,b,b~,e,e~,σ,σ~,κ,κ~)and wherein NS is the number of samples, NΛis the number of leads, mλ(ti) is the measurement at lead λ at time instance ti, vλ(ti) is the lead voltage determined in step 1e at le... |  | A61B005/04
US07951093 | B2 | 2011-05-31 | Composite medical device with markers | 2006-12-13 | Medical devices and methods of manufacturing medical devices. The medical devices include an elongate member having a proximal portion and a distal portion and a connector assembly disposed adjacent the elongated member to connect the proximal and distal portions, the connector assembly including a radiopaque marker. | 1. A medical device comprising: a first elongated member having a proximal portion and a distal portion;a second elongated member having a proximal portion and a distal portion;a tubular connector being disposed about and connected to the distal portion of the first elongated member and disposed about and connected to the proximal portion of the second elongated member to interconnect the first and second elongated members;a radiopaque marker disposed within the tubular connector and being disposed between the distal portion of the first elongated member and the proximal portion of the second elongated member;wherein the radiopaque marker is a wire or ribbon of radiopaque material disposed within the tubular connector;wherein the first elongate member comprises a stainless steel and the second elongate member comprises a nickel-titanium alloy, and the tubular connector comprises a material that is weld-compatible with both the stainless steel and the nickel-titanium alloy, and the connection between the first elongate member and the tubular connector includes a weld joint, and the connection between the second elongate member and the tubular connector includes a weld joint. 2. The medical device of claim 1, wherein the connection between the first elongate member and the tubular connector is spaced proximally from the radiopaque marker. 3. The medical device of claim 1, wherein the connection between the second elongate member and the tubular connector is spaced distally from the radiopaque marker. 4. The medical device of claim 1, wherein the tubular connector comprises an alloy having at least 20 percent chromium by weight and at least 50 percent nickel by weight. 5. The medical device of claim 4, wherein the alloy of the tubular connector comprises at least 2 percent iron by weight and at least 8 percent molybdenum by weight. 6. The medical device of claim 1, wherein the first elongated member is an elongated core member having a generally solid cross-section, and the second elongated member is an elongated core member having a generally solid cross-section. 7. A medical device comprising: a first elongated member having a proximal portion and a distal portion;a second elongated member having a proximal portion and a distal portion;a tubular connector being disposed about and connected to the distal portion of the first elongated member and disposed about and connected to the proximal portion of the second elongated member to interconnect the first and second elongated members;a radiopaque marker disposed within the tubular connector and being disposed between the distal portion of the first elongated member and the proximal portion of the second elongated member;wherein the first elongate member comprises a stainless steel and the second elongate member comprises a nickel-titanium alloy, and the tubular connector comprises a material that is weld-compatible with both the stainless steel and the nickel-titanium alloy, and the connection ... | Boston Scientific Scimed, Inc. | A61B005/00 A61M025/00
US07955270 | B2 | 2011-06-07 | Method and apparatus for respiratory audio-visual biofeedback for imaging and radiotherapy | 2006-10-04 | An improved method and apparatus for respiratory audio-visual biofeedback are disclosed. A guide patterned after a breathing cycle comfortable to the patient serves as a target. The target is displayed as a bar moving vertically upward during inhale and vertically downward during exhale, between fixed end ex-hale and end in-hale limits. The patient's current respiratory position is also displayed as a bar, oriented parallel to the target bar so that the difference between the current position and the target position is easy for the patient to see. | 1. A computer implemented method for respiratory audio-visual biofeedback for imaging and radiotherapy treatment, the computer performing steps comprising: displaying by the computer for a patient a continuously changing target indication of the patient's respiratory position;displaying by the computer for the patient a current indication of the patient's respiratory position; andaligning by the computer the target indication display and the current indication display so that a difference between the patient's current respiratory position and the patient's target respiratory position is readily apparent to the patientwherein the target indication, the current indication and the readily apparent difference are displayed continuously between an end-exhale position and an end-inhale position. 2. A method as in claim 1, wherein the patient's current respiratory position is indicated by a current marker along a current line from the end-exhale position to the end-inhale position and wherein the target respiratory position is indicated by a target marker along a target line from the end-exhale position to the end-inhale position, the lines for target respiratory position and current respiratory position being aligned in parallel in a first direction so that if the patient's current respiratory position is synchronized with the patient's target respiratory position the corresponding markers are aligned in a second direction perpendicular to the first direction. 3. A method as in claim 2, wherein the corresponding markers are straight line segments in said second direction. 4. A method as in claim 3, wherein each straight line segment is at one end of a bar, the bar extending along said first direction. 5. A method as in claim 4, wherein the bar measuring the patient's current respiratory position extends from the current line segment in the direction of the end-exhale position. 6. A method as in claim 5, wherein a target end-exhale limit and a target end in-hale limit are marked on the display. 7. A method as in claim 6, wherein the target end-exhale limit and the target end-inhale limit are marked by lines in said second direction. 8. A method as in claim 4, wherein the first direction is vertical and the second direction is horizontal, as the display is viewed by the patient. 9. In a computer implemented method for respiratory audio-visual biofeedback for imaging and radiotherapy treatment, said method displaying for a patient target end ex-hale and target end in-hale limits together with a current respiratory position to be kept within said limits by control of breathing by the patient, an improvement comprising: displaying by the computer a continuously changing target respiratory position of the patient, the target position varying continuously in accordance with a pattern learned from the patient; andaligning by the computer the display of the target respiratory position and the display of the current respiratory position so that a differe... | Stanford University | A61B005/08
US20110125050 | A1 | 2011-05-26 | Multiple-Electrode and Metal-Coated Probes | 2010-10-04 | Provided are probes featuring multiple electrodes, which probes have diameters in the nanometer range and may be inserted into cells or other subjects so as to monitor an electrical characteristic of the subject. The probes may also include a conductive coating on at least one probe element to improve the probes' performance. The probes may also be used to inject a fluid or other agent into the subject and simultaneously monitor changes in the subject's electrical characteristics in response to the injection. Related methods of fabricating and of using the inventive probes are also provided. | 1. A probe, comprising: An elongate insulator having a distal terminus and a lumen, the lumen having a diameter of from about 1 nm to about 1 cm;a first conductive layer on the inner surface of the lumen and in electrical communication with a first contact; anda second conductive layer surmounting at least a portion of the outer surface of the elongate insulator and in electrical contact with a second contact. 2. The probe of claim 1, wherein the second conductive layer terminates at a position remote from the distal terminus of the elongate insulator. 3. The probe of claim 1, wherein the elongate insulator has a variable cross-section. 4. The probe of claim 1, wherein the elongate insulator comprises a dielectric material. 5. The probe of claim 1, wherein the dielectric material comprises silica, silica nitride, glass, quartz, polymer, plastic, or any combination thereof. 6. The probe of claim 1, wherein the lumen of the elongate insulator has a diameter in the range of from about 10 nm to about 10000 nm. 7. The probe of claim 1, wherein the lumen of the elongate insulator has a diameter in the range of from about 100 nm to about 1000 nm. 8. The probe of claim 1, wherein the elongate insulator has a wall thickness in the range of from about 1 nm to about 10 micrometers. 9. The probe of claim 1, wherein the first conductive layer, the second conductive layer, or both, comprise carbon, metal, a conductive polymer. or any combination thereof. 10. The probe of claim 1, wherein the metal comprises gold, silver, chromium, titanium, tungsten, platinum, aluminum, nickel, or any combination thereof. 11. The probe of claim 1, wherein the first conductive layer has a thickness in the range of from about 1 nm to about 10 micrometers. 12. The probe of claim 1, further comprising a second elongate insulator disposed adjacent to the second electronically conductive layer. 13. The probe of claim 12, further comprising a third conductive layer surmounting at least at portion of the second elongate insulator, the third conductive layer in electrical communication with a third contact. 14. The probe of claim 12, wherein the second elongate insulator is coaxial with the first hollow insulator. 15. The probe of claim 12, further comprising a third elongate insulator disposed adjacent to the third conductive layer. 16. The probe of claim 1, further comprising a fourth conductive layer, the fourth conductive layer surmounting at least a portion of the third elongate insulator, and the fourth conductive layer in electrical communication with a fourth contact. 17. The electronic probe of claim 1, further comprising an injector in fluid communication with the lumen. 18. The electronic probe of claim 1, wherein at least a portion of the first conductive layer extends beyond the distal terminus of the elongate insulator. 19. The electronic probe of claim 18, wherein the portion of the first conductive layer that extends beyond the distal terminus of the e... | THE TRUSTEES OF THE UNIVERSITY OF PENNSYLVANIA | A61B005/053 C23C028/00 C25D007/00 B05D005/12 G01R027/08
US07955267 | B2 | 2011-06-07 | Apparatus for control of non-invasive parameter measurements | 2005-12-13 | Improved methods and apparatus for non-invasively assessing one or more parameters associated with fluidic systems such as the circulatory system of a living organism. In a first aspect, an improved method of continuously measuring pressure from a compressible vessel is disclosed, wherein a substantially optimal level of compression for the vessel is achieved and maintained using perturbations (e.g., modulation) of the compression level of the vessel. In one exemplary embodiment, the modulation is conducted according to a pseudo-random binary sequence (PBRS). In a second aspect, an improved apparatus for determining the blood pressure of a living subject is disclosed, the apparatus generally comprising a pressure sensor and associated processor with a computer program defining a plurality of operating states related to the sensed pressure data. Methods for pressure waveform correction and reacquisition, as well as treatment using the present invention, are also disclosed. | 1. A non-invasive and continuous blood pressure measurement apparatus, comprising: a non-invasive pressure sensor coupled to an applanation apparatus, said applanation apparatus controlled at least in part by a computing apparatus adapted to collect hemodynamic data from said sensor, said computing apparatus capable of operating in a plurality of modes, said plurality of modes comprising:a first mode, wherein blood pressure data is obtained via said sensor from the circulatory system of a living subject;a second mode, wherein at least one transient event is detected based on said blood pressure data; and eithera third mode subsequent to said detection of said event if said event is determined not to be of sufficient severity that recovery from said event in said second mode would exceed a first criterion; ora fourth mode subsequent to said detection if said event is determined to be of sufficient severity that recovery from said event in said second mode would exceed said first criterion. 2. The apparatus of claim 1, wherein said third mode further comprises operating in said third mode only if said event is determined to originate from said circulatory system of said living subject. 3. The apparatus of claim 1, wherein said fourth mode further comprises operating in said fourth mode only if said event is determined not to originate from said circulatory system of said living subject. 4. The apparatus of claim 1, wherein said first criterion comprises a time period or duration. 5. The apparatus of claim 1, wherein said event comprises a movement of at least a portion of said apparatus relative to said subject. 6. A non-invasive hemodynamic measurement system, comprising: a pressure sensor adapted to measure pressure data originating from an artery of a living subject;an applanation apparatus operatively coupled to said pressure sensor;a controller in signal communication with said sensor and said applanation apparatus, said controller capable of (i) detecting at least one transient event within hemodynamic data obtained from said sensor, and (ii) operating in a plurality of modes, said modes comprising:a first mode, wherein hemodynamic data is obtained from said sensor;a second mode, entered subsequent to said detection of said event if said event does not have a first characteristic; anda third mode, entered subsequent to said detection if said event has said first characteristic;wherein said first characteristic comprises parameters indicating that said event is not (a) physiologic in origin, and (b) of sufficient magnitude that recovery from said event in said second mode would exceed a first criterion. 7. The system of claim 6, wherein: said second mode comprises a mode adapted to selectively vary a level of compression by said applanation apparatus via a control signal; andsaid varying level of compression is applied via said pressure sensor. 8. The system of claim 6, wherein said first criterion comprises a time period or duration. ... | - | A61B005/02
US07945310 | B2 | 2011-05-17 | Surgical instrument path computation and display for endoluminal surgery | 2006-09-18 | An endoscopic surgical navigation system comprises a path correlation module that can compute the path taken by an endoscope scope or other medical instrument during an endoscopic medical procedure and various related attributes and parameters, and can compute and display a correlation between two paths. | 1. A computer-implemented method comprising: tracking a current position of a surgical instrument through a body lumen during an endoscopic procedure;generating data indicative of the current position of the surgical instrument for a plurality of positions of the surgical instrument during the endoscopic procedure;storing said data as a dataset representing a first path taken by the surgical instrument through the body lumen during the endoscopic procedure;causing an indication of the first path to be displayed to a user based on the dataset;determining a correlation between the first path and a second path through the body lumen, wherein determining the correlation includes determining a deviation of a position of the surgical instrument from the second path by projecting the first path and the second path onto each of three orthogonal planes, computing for each of the projections a first line connecting a first point of the first path and a second point of the second path and a second line connecting a third point on the first path and a fourth point on the second path, and computing an area of a shape defined by each of the projections of the first path and the second path and the corresponding first and second lines; andoutputting a summed total of the calculated areas as an indication of the correlation to a user. 2. A method as recited in claim 1, further comprising: transforming the data indicative of the positions of the surgical instrument along the first path into graphic model parameters to enable the first path to be rendered and displayed to a user. 3. A method as recited in claim 1, wherein the second path is a pre-surgery planned reference path. 4. A method as recited in claim 1 wherein the second path is a second recorded path of the surgical instrument. 5. A method as recited in claim 1, further comprising: during the endoscopic procedure, identifying a landmark along the first path and metadata associated with the landmark; andtransforming an associated position of the surgical instrument and said metadata into graphics primitives to allow rendering and display of data indicative of the associated position, orientation and the metadata. 6. A method as recited in claim 5, further comprising: during the endoscopic procedure, computing a distance between the current position of the surgical instrument and a landmark along the first path; andcausing an indication of the distance to be output to a user. 7. A method as recited in claim 1, wherein the surgical instrument is a flexible endoscope. 8. A method as recited in claim 7, wherein tracking the current position of the surgical instrument during an endoscopic procedure comprises tracking the current position of a distal tip of the flexible endoscope. 9. A system comprising: a processor; anda memory, coupled to the processor, storing instructions, which when executed by the system, cause the system to track a current position of an endoscope through a body lumen during an e... | Stryker Corporation | A61B005/05 A61B001/04
US20110125065 | A1 | 2011-05-26 | METHOD AND DEVICE FOR DETECTING ESTRUS | 2011-01-27 | The present invention provides a method and device for detecting estrus in animal by sensing along time the motion of the animal and identifying when the sensed motion is not related to eating periods of the animal. Based on the sensed motion which is not related to eating periods the estrus in the animal is identified | 1. A method for detecting estrus in a cattle animal, comprising the steps of: sensing and accumulating acceleration level of said cattle animal, over a period of time, by an acceleration sensor, wherein the acceleration level is indicated by energy level of an acceleration signal produced by the acceleration sensor;sensing, over a period of time, data indicative of eating performed by said cattle animal;attenuating the energy level of the acceleration signal as the indication of eating is stronger, the energy attenuated acceleration signal identifying neutralized motion data;extracting typical activity level of said animal based on said neutralized motion data; andidentifying abnormal behavior indicative of said estrus in said animal by comparing recently identified neutralized motion data with the extracted typical activity level. 2. A method according to claim 1, further comprising establishing a database of typical behavior, wherein said data base is continuously updated. 3. A method according to claim 1, further comprising the step of transmitting to a remote receiver at least one of a list comprising the sensed motion, the information related to eating periods, the collected neutralized motion data, the extracted activity level and the identification if the animal is in estrus. 4. A method according to claim 1, further comprising a step of indicating by indication means when the animal is identified as in estrus. 5. A method according to claim 1, further comprising a step of sending notification when the animal is identified as in estrus. 6. A method according to claim 1, wherein at least one of the steps is performed inside a case which is placed where the animal's head inclination can be sensed. 7. A method according to claim 1, wherein at least one of the steps is performed inside a case which is placed where the animal's head inclination can be sensed and wherein said case is at least one of the following: mounted on the animal's neck, implanted inside the animal or swallowed by the animal. 8. A method according to claim 1, wherein the step of sensing data indicative of eating comprises a step of measuring the inclination of the animal's head relative to the ground. 9. A method according to claim 1, wherein the step of sensing data indicative of eating comprises a step of measuring the distance of the animal's head to the ground. 10. A method according to claim 1, wherein the step of sensing data indicative of eating comprises a step of monitoring and analyzing the sounds and vibrations that occur when the animal is eating. 11. A method according to claim 1, wherein the step of sensing the acceleration level comprises a step of measuring the accelerations in up to three orthogonal directions of said animal. 12. A device for detecting estrus in a cattle animal, comprising: at least one acceleration sensor for sensing acceleration level of said cattle animal, over a period of time, wherein the acceleration level is indicated by ... | - | A61B005/11
US07966054 | B2 | 2011-06-21 | Disposable single-use external dosimeters for detecting radiation in fluoroscopy and other medical procedures/therapies | 2004-06-10 | Methods, systems, devices, and computer program products include positioning single-use radiation sensor patches that have adhesive onto the skin of a patient to evaluate the radiation dose delivered during a medical procedure or treatment session. The sensor patches are configured to be relatively unobtrusive and operate during radiation without the use of externally extending power cords or lead wires. | 1. A method for monitoring radiation exposure for a patient undergoing a medical procedure, comprising: releasably securing at least one single-use dosimeter sensor patch onto the skin of the patient, wherein the patch is self-contained;exposing the patient to radiation in a medical procedure, wherein the at least one patch comprises a radiation sensor circuit having a MOSFET that is unpowered during radiation exposure;transmitting data from the sensor patch to a dose-reader device after the exposing step to obtain data associated with a change in an operational parameter of the MOSFET in the dosimeter sensor patch; anddetermining radiation received by the patient during the exposing step based on the change in the MOSFET operational parameter. 2. A method according to claim 1, wherein an underside of the at least one patch comprises an adhesive, and wherein the releasably securing step is carried out by pressing the patch onto the skin so that it adheres thereto. 3. A method according to claim 1, wherein the transmitting data is carried out by inserting a tab extending from the at least one sensor patch into a port in the dose-reader device. 4. A method according to claim 1, further comprising at least one of calibrating and/or pre-dosing the at least one single-use dosimeter sensor patch before releasably securing the sensor patch onto the skin of the patient. 5. A method according to claim 1, wherein the at least one dosimeter patch comprises a plurality of spaced apart discrete radiation sensor circuits, each comprising a single MOSFET. 6. A method according to claim 1, further comprising storing radiation-data relating to the exposing and/or determining step on a memory device integrated with the at least one sensor patch. 7. A method according to claim 1, further comprising storing a plurality of patient specific data records in the reader, then transmitting the radiation data obtained by the reader from the at least one sensor patch to a remote computer. 8. A method according to claim 1, further comprising sterilizing the single-use dosimeter sensor patch and applying an adhesive to a selected primary surface, and packaging the single-use dosimeter patch in a sterile package suitable for human medical use. 9. A method according to claim 1, wherein the medical procedure is a fluoroscopic procedure. 10. A method according to claim 1, wherein the at least one single-use dosimeter patch has a low profile and is generally conformable to the underlying skin of the patient, and wherein the step of releasably securing comprises positioning the patch proximate a target radiation beam path configured to enter into the body about a chest portion of the patient. 11. A method according to claim 10, further comprising positioning at least one of the patches proximate to a region where radiation is not desired or in a radiation-sensitive region. 12. A method according to claim 9, wherein the at least one patch comprises a first patch that comp... | Sicel Technologies, Inc. | A61B005/05
US08498684 | B2 | 2013-07-30 | Sine saturation transform | 2011-03-08 | A transform for determining a physiological measurement is disclosed. The transform determines a basis function index from a physiological signal obtained through a physiological sensor. A basis function waveform is generated based on basis function index. The basis function waveform is then used to determine an optimized basis function waveform. The optimized basis function waveform is used to calculate a physiological measurement. | 1. A method of determining of a physiological parameter of a monitored patient attached to a noninvasive optical sensor, the method including electronically processing signals indicative of attenuation by body tissue of the patient, the method comprising: receiving one or more signals indicative of light attenuated by body tissue of pulsing blood in said monitored patient, said one or more signals usable to determine measurements for one or more physiological parameters of said monitored patient;electronically determining at least one basis function index based on said one or more signals;electronically generating at least one basis function waveform based on said at least one basis function index;electronically optmizing said at least one basis function waveform and said one or more signals to determine an optimized basis function index;electronically generating an optimized basis function waveform based on said optimized basis function index; andelectronically processing said one or more signals and said optimized basis function waveform to determine an output measurement of one or more of said one or more physiological parameters. 2. The method of claim 1, comprising generating one or more sinusoidal waves as the at least one basis function waveform. 3. The method of claim 2, comprising individually canceling the generated sinusoidal waves from said one or more signals. 4. The method of claim 1, wherein said optimizing includes determining a frequency corresponding to the sinusoidal waveform which produces an extrema. 5. The method of claim 4, wherein the extrema comprises a minima. 6. The method of claim 1, comprising sending said output measurement to a display. 7. The method of claim 6, comprising displaying said output measurement. 8. The method of claim 1, wherein said one or more physiological parameters comprises oxygen saturation. 9. The method of claim 1, wherein said at least one basis function corresponds to an estimated pulse rate. 10. The method of claim 1, wherein said at least one basis function waveform comprises two or more basis function waveforms. 11. A signal processing system for determining a physiological parameter of a monitored patient attached to a noninvasive optical sensor, electronically processing signals indicative of attenuation by body tissue of the patient, the system comprising: at least one input configured to receive one or more signals indicative of light attenuated by body tissue of pulsing blood in said monitored patient, said one or more signals usable to determine measurements for one or more physiological parameters of said monitored patient;at least one basis function indicator configured to determine at least one basis function index based on said one or more signals;at least one basis subset generator configured to generate at least one basis function waveform based on said at least one basis function index;at least one optimization calculator configured to optimize said one or more signa... | Masimo Corporation | A61B005/1455 A61B005/024
US09144383 | B2 | 2015-09-29 | Device and method for in vivo noninvasive magnetic manipulation of circulating objects in bioflows | 2010-11-12 | A device and methods for the non-invasive manipulation and detection of target objects such as cells, pathogens, microparticles, and nanoparticles in vivo using an external magnetic field are described. In one aspect, a device and method for capturing and detecting intrinsically magnetic target objects or target objects labeled with at least one magnetic particle within the area of interest using an in vivo flow cytometer are described. | 1. A device for the manipulation and detection of a magnetic target object within a moving biofluid of a living organism, comprising: an in vivo flow cytometer for detecting the magnetic target object within an area of interest; andat least one magnet for producing a magnetic field that manipulates the magnetic target object within the area of interest wherein the in vivo flow cytometer is an in vivo photoacoustic flow cytometer comprising a laser fiber for delivering at least one laser pulse to the area of interest, wherein the laser fiber comprises a focusing tip to transmit the at least one laser pulse to the area of interest; andwherein the at least one magnet comprises a magnetic material forming a channel, wherein the laser fiber passes through the channel and the focusing tip is situated at a laser pulse location that is essentially coincident to a corresponding magnetic field location. 2. The device of claim 1, wherein the manipulation of the magnetic target object is chosen from enrichment, capture, sorting, separation, concentration within a selected region of the biofluid, or any combination thereof of the target object within the area of interest. 3. The device of claim 1, wherein the at least one magnet is situated at a distance between about 0.1 mm and about 20 cm from the area of interest. 4. The device of claim 1, wherein the device detects the magnetic target object at a detection sensitivity ranging from about 1 to about 100 target objects per L of the moving biofluid. 5. The device of claim 1, wherein the at least one magnet produces a field that captures the magnetic target object within the area of interest. 6. The device of claim 1, wherein the moving biofluid is chosen from blood, lymph, cerebrospinal fluid, urine, chyme, cytosol, tears, or interstitial fluid. 7. The device of claim 1, wherein the magnetic target object is chosen from an intrinsically magnetic target object or a target object labeled with a magnetic particle. 8. The device of claim 7, wherein the magnetic particle is chosen from a magnetic nanoparticle or a magnetic microparticle having a size ranging from about 5 nm to about 10 μm. 9. The device of claim 1, wherein the magnetic target object is labeled with a contrast agent selected from an ultrasound contrast agent, a photoacoustic contrast agent, a hybrid contrast agent, a fluorescent contrast agent, a Raman contrast agent, an MRI contrast agent, a superconductivity quantum interference device contrast agent, a PET contrast agent, and a CT contrast agent. 10. The device of claim 9, wherein the hybrid contrast agent is a nanoparticle complex or a microparticle complex chosen from a gold-magnetic complex, a quantum dot-magnetic complex, a carbon nanotube-magnetic complex, a radionucleotide-magnetic complex, a surface enhanced resonance scattering-magnetic complex, or any combination thereof. 11. The device of claim 1, wherein the magnetic field has a strength ranging from about 0.1 T to about 20 ... | The Board of Trustees of the University of Arkansas | A61B008/00 A61B006/00 A61B005/00 A61K049/22 B03C001/28 B03C001/30 A61B008/08 A61B018/18 A61B018/20 G01N015/14
US08731652 | B2 | 2014-05-20 | Method and apparatus for monitoring the autonomous nervous system of a sedated patient | 2009-03-16 | A method and an apparatus for monitoring the autonomous nervous system of a sedated patient. The method comprises steps of providing a skin conductance signal measured at an area of the patient's skin through a measurement interval, calculating a characteristic of said skin conductance signal, establishing a first output signal indicating the state of pain or discomfort in the patient, and a second output signal indicating the state of awakening in the patient, based on said characteristic of said skin conductance signal. The calculating of the signal characteristic comprises calculating a value representative of a statistical dispersion, e.g., the standard deviation, of the values of the skin conductance signal through the measurement interval. | 1. A method for monitoring the autonomous nervous system of a sedated patient, comprising: measuring, by an electrode arrangement and a measurement converter, a skin conductance signal at an area of the patient's skin through a measurement interval,calculating, by a processing device, a characteristic of said skin conductance signal,establishing, by a processing device, a first output signal indicating a state of pain or discomfort in the patient, based on said characteristic of said skin conductance signal,establishing, by a processing device, a second output signal indicating a state of awakening in the patient, based on said characteristic of said skin conductance signal,wherein said step of calculating said characteristic comprisescalculating, by a processing device, a value representative of a statistical dispersion of values of said skin conductance signal through said measurement interval, the value being a standard deviation of the values of said skin conductance signal through said measurement interval,further wherein said step of establishing said first output signal indicating the state of pain or discomfort in the patient comprises comparing said standard deviation with a first limit value, the first signal being established as indicating pain or discomfort in the patient if the standard deviation exceeds the first limit value, and wherein said step of establishing said second output signal indicating the state of awakening in the patient comprises comparing said standard deviation with a second limit value, the second signal being established as indicating awakening in the patient if the standard deviation exceeds the second limit value. 2. The method according to claim 1, wherein said first limit value is in the range [0.01 μS, 0.30 μS]. 3. The method according to claim 2, wherein said first limit value is in the range [0.02 μS, 0.10 μS]. 4. The method according to claim 3, wherein said first limit value is 0.03 μS. 5. The method according to claim 1, wherein said second limit value is in the range [0.06 μS, 18.0 μS]. 6. The method according to claim 5, wherein said second limit value is in the range [0.10 μS, 3.0 μS]. 7. The method according to claim 6, wherein said second limit value is 0.5 μS. 8. The method according to one of the preceding claims, wherein said measurement interval is in the range [5 seconds, 40 seconds]. 9. The method according to claim 8, wherein said measurement interval is in the range [10 seconds, 30 seconds]. 10. The method according to claim 9. wherein said measurement interval is 15 seconds. 11. An apparatus for monitoring a sedated patient, comprising measurement equipment for providing a skin conductance signal measured at an area of the patient's skin, anda control unit, including a processing device which is configured to perform the following steps:calculating a characteristic of said skin conductance signal,establishing a first output signal indicating a state of pain or discomfort in the ... | Med Storm Innovation AS | A61B005/053
US09204835 | B2 | 2015-12-08 | Voxel-based methods for assessing subjects using positron emission tomography | 2010-11-01 | The subject matter disclosed herein relates to methods for diagnosing a neurological disorder in a subject. In certain aspects, the methods described herein involve determining one or more critical areas in the brain from PET data where two groups differ and measuring PET signal within determined critical areas in a new subject in order to assign risk or diagnosis. | 1. A computer-implemented method for diagnosing or determining risk of a neurological disorder in a subject, the method comprising: (a) determining the presence of a radioactive signal from an administered radioligand in at least one individual in a control group and at least one individual in a reference group to generate primary brain scan image voxel data of radioligand distribution in a brain of at least one individual in a control group and at least one individual in a reference group,(b) generating secondary brain scan image data for the individuals in the control and reference groups, wherein the secondary scan brain image data is generated using a different type of brain scan than the primary brain scan image data,(c) generating a probability-corrected time-activity curve data for each voxel in the primary brain scan image of the at least one individual in the control and reference group,(d) processing the probability-corrected time-activity curve data of the at least one individual in the control and reference group to generate a voxel binding outcome map data of the at least one individual in the control and reference group,(e) transforming the voxel binding outcome map data of the at least one individual in the control and reference group into a normalized space to generate a normalized voxel binding outcome map data of the at least one individual in the control and reference group(f) processing the normalized voxel binding outcome map data of the at least one individual in the control and reference group using statistical analysis to identify one or more voxels of interest (VOI) in the normalized voxel binding outcome map data to generate a VOI map data for differentiating of the at least one individual in the control and reference group, and(g) applying the VOI map data to a voxel binding outcome map data of a test subject to generate a mean cortical binding value to diagnose or determine risk of a neurological disorder in the subject wherein the voxel binding outcome map data of said test subject is generated from Positron Emission Tomography (PET) using a PET scanner. 2. The method of claim 1, wherein the transforming of step (e) comprises transforming the voxel binding outcome map data of the at least one individual in the control and reference group into a secondary scan space to generate a secondary space voxel binding outcome map data of the at least one individual in the control and reference group. 3. The method of claim 1, wherein step (g) comprises: (i) inverse transforming the VOI map data identified in step (f), into a secondary scan space of the test subject to generate a voxel of interest (VOI) mask for the test subject,(ii) multiplying the VOI mask for the subject by probabilistic brain region (BRP) map data for the subject to generate a brain region VOI mask for the subject,(iii) multiplying a secondary space voxel binding outcome map data of the subject by the brain region VOI mask for the subject to generate corti... | THE TRUSTEES OF COLUMBIA UNIVERSITY IN THE CITY OF NEW YORK | A61B005/05 A61B005/16 A61B005/055 A61B006/03 A61B006/00 G01R033/00 G01R033/10 G06T007/00 G06T007/40
US09042967 | B2 | 2015-05-26 | Device and method for wound imaging and monitoring | 2009-05-20 | A device for fluorescence-based imaging and monitoring of a target comprising: a light source emitting light for illuminating the target, the emitted light including at least one wavelength or wavelength band causing at least one biomarker associated with the target to fluoresce; and a light detector for detecting the fluorescence. | 1. A portable, handheld device for outputting a representation of a wound in tissue, the device comprising: a housing configured to receive and secure a digital image detector device therein;at least one light source configured to emit a homogeneous field of excitation light, the at least one light source being coupled to the housing and positionable to directly illuminate a target surface with the homogeneous field of excitation light, the light source being movable with respect to the housing to vary an illumination angle of the excitation light on the target surface, the target surface including at least a portion of a wound and an area around the wound;a power supply contained in the housing and configured to provide power to the at least one light source; anda digital image detector device configured to be removably secured in the housing and comprising: a wireless communication device having a sensor configured to detect signals responsive to illumination of the illuminated portion of the wound and the area around the wound, each signal indicative of at least one of endogenous fluorescence, exogenous fluorescence, absorbance, and reflectance in the illuminated portion of the wound and the area around the wound, wherein the at least one light source is positionable with respect to the wireless communication device when the digital image detector device is secured in the housing;a processor configured to receive the detected signals and to output a representation of the illuminated portion of the wound and the area around the wound based on the detected signals, wherein the processor is further configured to spatially co-register data relative to at least one of wound topography, wound anatomy, wound area, wound depth, wound volume, wound margins, and necrotic tissue, based on the detected signals, in the representation of the illuminated portion of the wound and the area around the wound; anda display for displaying the representation of the illuminated portion of the wound and the area around the wound output by the processor. 2. The portable, handheld device of claim 1, wherein the representation comprises mixed red-green-blue fluorescent representations of various bacteria and tissue components present in the illuminated portion of the wound and the area around the wound, and wherein the processor is configured to analyze the red-green-blue fluorescent representations in order to recognize and classify various components of the illuminated portion of the wound and the area around the wound, based on the received signals or the output representation, including: a location of at least one of bacteria, fungus, yeast, and other microorganisms present in the illuminated portion of the wound and the area around the wound,at least one of a population, a quantity, a distribution, a colonization, a contamination, a critical colonization, an infection, and an extent of at least one of bacteria, fungus, yeast, and other microorganisms when present ... | University Health Network | A61B005/00 G01N021/64 A61B010/00
US20110112439 | A1 | 2011-05-12 | Rotational Safety Syringe | 2009-11-06 | A rotational safety syringe includes front and rear barrels rotatable with respect to each other for moving a needle holder towards the rear barrel and compressing a spring to separate a protective cover. The needle holder is fixed by an elastic press button on the needle holder and a press opening on the rear barrel. After a through hole of the front barrel is aligned with a using position, the elastic press button can be pressed to release the needle holder, and the spring can eject the needle holder towards the through hole of the front barrel to extend the needle out from the through hole of the front barrel and pierce the needle into the using position. When the spring is released, the needle holder and needle are pulled back to hide the needle into the front and rear barrels to prevent the used needle from being touched. | 1. A rotational safety syringe, comprising: a set of a front barrel and a rear barrel, sheathed and rotated with respect to one another, and having at least one pair of oppositely engaged axial grooves formed on internal walls of the front and rear barrels respectively, and a rear end of at least one spiral groove of the front and rear barrels being coupled to the axial groove, and a press opening having a penetrating cylindrical wall and formed at a joint between the spiral groove and the axial groove of the rear barrel;a needle holder, installed in the set of front and rear barrels, and including a needle installed at a front end of the needle holder, an elastic press button installed on an external wall of the needle holder for securing and holding the press opening, and the needle being retractable into the spiral groove and the axial groove, and at least one protrusion formed on an external wall of the needle holder for extending the needle holder into the spiral groove of the front barrel; anda spring, installed into the set of front and rear barrels, and having two fixed ends fixed to the rear barrel and the needle holder respectively, such that when the spring is not compressed, the needle of the needle holder is hidden into the set of front and rear barrels. 2. The rotational safety syringe of claim 1, wherein the front barrel includes a through hole formed at a front end of the front barrel for passing the needle into and out of the through hole. 3. The rotational safety syringe of claim 1, wherein the front barrel includes a through hole formed at a front end of the front barrel for passing the needle into or out of the through hole, and the needle holder includes a protective cover covered onto the needle, and an end of the protective cover is extended out of the through hole of the front barrel, and the protective cover extended out from the front barrel has a width greater than the through hole. 4. The rotational safety syringe of claim 1, wherein the press opening is a spiral groove. 5. The rotational safety syringe of claim 1, wherein the distance between the spiral grooves in the front and rear barrels can be adjusted for changing an elasticity of the spring and a penetration force of the needle. | - | A61B005/151
US20110118557 | A1 | 2011-05-19 | Intelligent User Interface For Medical Monitors | 2010-11-18 | An intelligent learning process for a user interface of a medical monitor is disclosed. The medical monitor may record user statistics and cluster groups based on settings, configurations, and actions captured by the user statistics. The medical monitor may create classes of users based on the groups and then classify users into classes based on the user statistics. The user interface of the monitor may be adapted based on the user's class. In other embodiments, a central station may access user statistics from multiple monitors and adapt a user interface for the monitors based on the statistics. | 1. A system, comprising: a medical monitor coupled to a sensor, wherein the medical monitor is configured to display one or more physiological parameters, store user statistics of one or more users, and adapt a user interface of the monitor based on the user statistics. 2. The system of claim 1, wherein the monitor is configured to cluster commonly used settings, configurations, and/or actions of the user statistics into one or more groups. 3. The system of claim 1, wherein the monitor is configured to determine a care environment for the monitor based on the statistics. 4. The system of claim 3, wherein the monitor is configured to adapt a user interface of the monitor based on the statistics and the care environment. 5. The system of claim 1, wherein adapting the user interface comprises modifying the information displayed on a first screen of the user interface. 6. The system of claim 1, wherein adapting the user interface comprises reconfiguring the keystrokes used to access one or more elements of the user interface. 7. The system of claim 1, wherein adapting the user interface comprises modifying one or more alarms. 8. The system of claim 1, comprising creating a plurality of classes based on the user statistics. 9. The system of claim 8, comprising wherein the monitor is configured to classify the one or more users into one or more of the plurality of classes based on the user statistics. 10. The system of claim 9, wherein the monitor is configured to adapt the user interface for the one or more users based on the class of the one or more users. 11. A system, comprising: a central station; anda plurality of medical monitors coupled to the central station, and each comprising a user interface,wherein the central station is configured to access user statistics from at least one of the plurality of medical monitors and adapt the user interface of one or more of the plurality of medical monitors based on the user statistics. 12. The system of claim 11, wherein central station is configured to store the user statistics. 13. The system of claim 11, wherein each of the plurality of medical monitors is configured to store the user statistics. 14. The system of claim 11, wherein the central station is configured to push the user interface to one or more of the plurality of medical monitors. 15. The system of claim 11, wherein the central station is configured to provide instructional text to the plurality of medical monitors for display on one or more of the plurality of medical monitors. 16. The system of claim 11, wherein the central station is configured to cluster commonly used settings, configurations, and/or actions of the user statistics into one or more groups. 17. A method, comprising: storing a plurality of user statistics on a medical monitor;determining a plurality of classes based on the user statistics; andadapting a user interface of the monitor based on the classes. 18. The method of claim 17, comprising classifying... | Nellcor Purifan Bennett LLC | A61B005/00
US20110144475 | A1 | 2011-06-16 | METHOD FOR A REGULATION OF A VASCULAR INTERVENTION, AND DEVICE TO REGULATE A VASCULAR INTERVENTION | 2010-12-10 | In a method and apparatus to regulate a vessel intervention conducted with a vessel intervention unit, wherein an intervention region of the vessel intervention is at least partially imaged by means of an imaging monitoring unit, and a monitoring region to monitor the vessel intervention is selected, and the vessel intervention is at least partially automatically regulated in the intervention region dependent on at least one characteristic variable identified in the monitoring region. | 1. A method to regulate a vessel intervention in the body of a patient, comprising the steps of: implementing a vessel intervention in the body of a patient using a vessel intervention unit that interacts with a vessel in the body of the patient in an intervention region in the body of a patient;with an imaging monitoring unit, generating and displaying an image of at least a portion of said intervention region;with said imaging monitoring unit, also generating an image of a monitoring region in the patient, said image of said monitoring region comprising monitoring image contents; andat least partially automatically controlling operation of said intervention unit, to at least partially automatically regulate said vessel intervention, dependent on at least one characteristic variable of said monitoring region derived from said monitoring region contents. 2. A method as claimed in claim 1 comprising locating said monitoring region spatially separate from said intervention region. 3. A method as claimed in claim 1 comprising selecting said monitoring region as a region in the body of the patient that contains at least one vessel that is supplied by a vessel in the intervention region that is subjected to said vessel intervention. 4. A method as claimed in claim 1 comprising implementing said vessel intervention in at least two successive intervention steps. 5. A method as claimed in claim 4 comprising injecting contrast agent into the body of the patient after at least one of said intervention steps. 6. A method as claimed in claim 5 comprising deriving said at least one characteristic variable after administration of the contrast agent. 7. A method as claimed in claim 1 comprising determining said characteristic variable at least partially in parallel with said vessel intervention. 8. A method as claimed in claim 1 comprising adapting an intervention speed of said vessel intervention dependent on said characteristic variable. 9. A method as claimed in claim 1 comprising controlling said vessel intervention unit from a computerized control unit and providing said characteristic variable to said computerized control unit and, from said computerized control unit, automatically controlling said vessel intervention unit to end said vessel intervention when said characteristic variable reaches a predetermined limit value. 10. A method as claimed in claim 1 comprising implementing said vessel intervention in a plurality of intervention steps, monitoring said characteristic variable with respect to a predetermined limit value during said vessel intervention, and beginning at least one of said intervention steps as long as said characteristic value has not reached said predetermined limit value. 11. A device to regulate a vessel intervention in the body of a patient, comprising: a vessel intervention unit configured to implement a vessel intervention in the body of a patient by interacting with a vessel in the body of the patient in an interventi... | - | A61B005/055
US20110152665 | A1 | 2011-06-23 | METHOD AND APPARATUS FOR DESIGNING MRI GRADIENT PULSE WAVEFORM | 2010-12-22 | A method for designing MRI gradient pulse waveform provided by the present invention firstly defines target peripheral nerve stimulation (PNS) curve; then calculates a gradient pulse waveform by using a relation function between the gradient pulse waveform and PNS value curve based on the target PNS curve. | 1. A method for designing MRI gradient pulse waveform, said method comprising: defining a target peripheral nerve stimulation (PNS) curve;calculating a gradient pulse waveform using a relation function between the gradient pulse waveform and a PNS value curve based on the target PNS curve;Rx(t)=1rb∫0tB.x(θ)c(c+t-θ)2θwherein rb is rheobase, c is chronaxie time, rb and c are both provided by system; Rx represents PNS value, θ is an intermediate variant from 0˜t; Bx represents gradient field intensity, i.e. amplitude; {dot over (B)}x(θ) represents slew rate, t represents time. 2. A method for designing MRI gradient pulse waveform as stated in claim 1, wherein defining the target PNS curve comprises: acquiring a slew rate from a magnetic resonance system;defining target peripheral nerve stimulation (PNS) curve based on the acquired slew rate;reaching a maximal PNS value from a start point with the acquired slew rate, determining a second point of the PNS curve, the time of the second point of the PNS curve being undetermined;setting a time randomly, from the second point of the PNS curve as a start, keeping PNS value within a client-set range in the period of time, determining a third point of the PNS curve;controlling PNS to decrease from the client-set range to a negative maximal point from the third point with the acquired slew rate, determining a fourth point of the PNS curve, the time of the fourth point of the PNS curve being undetermined; andkeeping PNS value within a client-set range from the fourth point of the PNS curve to a fifth point of the PNS curve, the time of the fifth point of the PNS curve being undetermined. 3. A method for designing MRI gradient pulse waveform as stated in claim 2, wherein calculating a gradient pulse waveform comprises: setting the start point of the gradient pulse waveform to be designed;based on the acquired slew rate and the PNS value of the second point of the PNS curve, calculating the time and the amplitude of the second point of the gradient pulse waveform by using the relation function between the gradient pulse waveform and the PNS value curve, the time of the second point of the gradient pulse waveform being consistent with the time of the second point of the PNS curve, whereby determining the second point of the gradient pulse waveform;based on the time and the PNS value of the third point of the PNS curve, calculating the amplitude of the third point of the gradient pulse waveform by using the relation function between the gradient pulse waveform and the PNS value curve, the time of the third point of the gradient pulse waveform being consistent with the time of the third point of the PNS curve, whereby determining the third point of the gradient pulse waveform;based on the acquired slew rate and the PNS value of the fourth point of the PNS curve, calculating the amplitude and the time of the fourth point of the gradient pulse waveform by using the relation function between the gradient puls... | - | A61B005/055
US20110125063 | A1 | 2011-05-26 | Systems and Methods for Monitoring and Modifying Behavior | 2010-11-16 | A system for detecting non-verbal acoustic energy generated by a subject is provided. The system includes a sensor mountable on or in a body region of the subject, the sensor being capable of sensing the non-verbal acoustic energy; and a processing unit being capable of processing the non-verbal acoustic energy sensed by the sensor and deriving an activity related signature therefrom, thereby enabling identification of a specific activity associated with the non-verbal acoustic energy. | 1. A method of accumulating data relating to ingestion behavior of a subject, comprising: (a) monitoring ingestion related activity from tissue deformation of an ear of the subject over a time period;(b) generating a signature classifying said ingestion related activity of the subject over said time period; and(c) computationally logging said signature, thereby accumulating data relating to ingestion behavior of the subject. 2. The method of claim 1, further comprising associating said signature with a duration of ingestion, rate of ingestion, food mass, volume, caloric value or nutritional value of material ingested prior to, or following step (c). 3. The method of claim 1, further comprising aggregating said data relating to ingestion behavior for a plurality of subjects into a database. 4. The method of claim 1, further comprising repeating steps (a)-(c) thereby accumulating data comprising a plurality of distinct signatures. 5. The method of claim 1, wherein (a) is effected by counting a number of bites, chews and/or swallows over said time period. 6. The method of claim 5, wherein said signature reflects said number of bites chews and/or swallows over said time period. 7. The method of claim 6, wherein said signature further reflects a sequence of said bites chews and/or swallows over said time period. 8. The method of claim 5, wherein said counting is effected by a user mounted device. 9. The method of claim 8, wherein said device is capable of sensing ingestion activity related motion or acoustic energy. | - | A61B005/11
US07965833 | B2 | 2011-06-21 | Febrile convulsion alarm | 2007-01-09 | A device and method for detecting febrile convulsions, the device including a detector for detecting motion of the device and providing an output corresponding thereto, a capacitor coupled to the detector and charged by the output of the detector, a resistor coupled to the capacitor to provide discharge of the capacitor, a comparator for comparing a voltage level of the capacitor to a reference voltage and providing a signal when the voltage level of the capacitor is greater than the reference voltage, and an indicator driven by the signal from the comparator to provide an indication of febrile convulsions. | 1. A device for detecting febrile convulsions, the device comprising: a tilt movement sensor and a mono-stable pulse detector for detecting motion of the device and providing an output signal corresponding thereto;a capacitor coupled to said detector and charged by the output of said detector;a resistor coupled to said capacitor to provide discharge of said capacitor;a comparator for comparing a voltage level of said capacitor to a reference voltage and providing a signal when said voltage level of said capacitor is greater than said reference voltage;an indicator driven by said signal from said comparator to provide an indication of febrile convulsions; anda strap adapted and configured to be worn on a baby's body, said detector, capacitor, resistor, comparator and indicator being mounted on said strap. 2. The device according to claim 1, wherein said indicator is a buzzer, and said signal from said comparator powers an oscillator for actuating said buzzer. 3. The device according to claim 1, further comprising a battery monitor for indicating low battery power. 4. The device according to claim 1, wherein said strap is adapted and configured to be worn on a baby's hand. 5. The device according to claim 1, wherein said detector, capacitor, resistor, comparator and indicator are embodied in an ASIC. 6. The device according to claim 1, wherein said detector, capacitor, resistor, comparator and indicator are embodied in a Printed Circuit Board (PCB). 7. The device according to claim 1, wherein said detector, capacitor, resistor, comparator and indicator are emulated by a microprocessor. 8. A device for detecting febrile convulsions, the device comprising: a tilt movement sensor for detecting motion of the device and providing an output signal corresponding thereto;an input filter network for providing said output signal to a mono-stable pulse detector;a capacitor coupled to said detector and charged by the output of said detector;a resistor coupled to said capacitor to provide discharge of said capacitor;a signal alarm level comparator for comparing a voltage level of said capacitor to a reference voltage and providing a signal when said voltage level of said capacitor is greater than said reference voltage;a piezo buzzer driven by said signal from said comparator via an audio oscillator to provide an indication of febrile convulsions;a small coin-type Lithium battery for providing battery power;a battery level comparator to provide an indication of low battery power;an LED for indicating system operation; andan LED for indicating low battery power. 9. A method of detecting febrile convulsions, the method comprising: mounting a tilt movement sensor and a mono-stable pulse detector in a strap of a device adapted to be worn on a baby's body;detecting tilting motion of a device by means of said tilt movement sensor and said mono-stable pulse detector and providing an output signal corresponding thereto;charging a capacitor in said strap coupled... | - | H04M001/65 A61B005/11
US07970466 | B2 | 2011-06-28 | Method and apparatus for optimization and assessment of response to extra-systolic stimulation (ESS) therapy | 2003-10-07 | A method and apparatus for optimizing and assessing the response to extra-systolic stimulation (ESS) are provided. An optimization/monitoring parameter is calculated as a function of potentiation ratio, PR, and recirculation fraction, RF, derived from measurements of myocardial contractile function during and after ESS. PR may be computed as the ratio of the contractile function on post-extra-systolic beats during ESS to baseline contractile function. RF may be computed as the slope of a linear regression performed on a plot of the contractile function for a post-extra-systolic beat versus the contractile function for the previous post-extra-systolic beat after ESS is ceased. The ESI resulting in a maximum optimization/monitoring parameter, preferably computed as the product of PR and RF, is determined as the optimal ESI. The operating ESI may be automatically adjusted, and/or PR and RF data may be stored for monitoring purposes. | 1. A method of optimizing delivery of an extra-systolic stimulation (ESS) therapy for a plurality of extra-systolic intervals (ESIs), comprising: setting a test ESI for delivery of an ESS therapy;applying the ESS therapy at the test ESI;determining a potentiation ratio (PR) and a recirculation fraction (RF);storing the previously determined PR and RF with the test ESI;incrementing or decrementing the test ESI to a new value test ESI within a predetermined range of values and repeating the first four steps, wherein the new value test ESI is substituted for the test ESI;calculating a product of RF and PR for each test ESI;determining an optimal ESI that corresponds to a maximal product of RF and PR; andsetting an operating ESI to the optimal ESI. 2. A method according to claim 1, further comprising performing the ESS therapy delivery in an implantable pulse generator. 3. Previously presented) A method according to claim 1, wherein determining PR and RF comprises deriving a contractile function from an output signal from a mechanical sensor. 4. A method according to claim 3, wherein the mechanical sensor comprises an accelerometer. 5. A method according to claim 4, further comprising disposing an accelerometer in mechanical communication with a ventricle. 6. A method according to claim 1, wherein determining PR and RF comprises deriving a contractile function from an output signal from a mechanical sensor disposed in fluid communication with a ventricle. 7. A method according to claim 6, wherein the mechanical sensor comprises a fluidic pressure sensor. 8. A method according to claim 7, further comprising disposing a fluidic pressure sensor in fluid communication with a right ventricle. 9. A method of optimizing delivery of an extra-systolic stimulation (ESS) therapy for a given extra-systolic interval (ESI), comprising: determining a baseline contractile function parameter when an ESS therapy is withheld or not delivered for a sufficient time or number of cardiac cycles for any related hemodynamic enhancement to substantially cease;delivering the ESS therapy for an interval of time or for a number of cardiac cycles;calculating an ESS therapy contractile function parameter during the delivery of the ESS therapy;calculating a potentiation ratio (PR) based on the baseline contractile function parameter and the ESS therapy contractile function parameter;ceasing delivery of the ESS therapy;determining a recirculation fraction (RF) for several cardiac cycles following the ceasing delivery of the ESS therapy; andcalculating a product of PR and RF. 10. A method according to claim 9, further comprising performing the ESS therapy delivery in an implantable pulse generator. 11. A method according to claim 9, wherein determining PR and RF comprises deriving a contractile function from an output signal from a mechanical sensor. 12. A method according to claim 11, wherein the mechanical sensor comprises an accelerometer. 13. A method according to cla... | Medtronic, Inc. | A61N001/00 A61B005/04
US07955280 | B2 | 2011-06-07 | Leg alignment for surgical parameter measurement in hip replacement surgery | 2006-06-01 | A CAS system and method comprises a first reference in fixed relation with the pelvis and a registration tool. A sensor apparatus tracks the first reference and tool. A controller unit receives tracking data for the first reference and tool. A calculator tracks the pelvic frame of reference, and the tool to produce a femoral frame of reference at two sequential operative steps. A reference orientation adjustor receives tracking data for the pelvic frame of reference to orient the femoral frame of reference in a reference orientation with respect to the pelvic frame of reference, and to produce a reference adjustment value as a function of the reference orientation. A surgical parameter calculator receives tracking data from the tool to calculate surgical parameters as a function of the reference adjustment value. | 1. A method of measuring surgical parameters in computer-assisted surgery so as to guide an operator in inserting a hip joint implant in a femur, comprising: i) digitizing a frame of reference of a pelvis, the frame of reference of the pelvis being trackable in space;ii) digitizing a pre-reduction frame of reference of the femur as a function of the frame of reference of the pelvis, the pre-reduction frame of reference comprising a femoral axis;iii) obtaining a reference orientation for the femoral axis of the pre-reduction frame of reference of the femur with respect to the frame of reference of the pelvis;iv) digitizing a reduction frame of reference of the femur by virtually orienting the femoral axis in said reference orientation as a function of the frame of reference of the pelvis, after initiation of implant reduction;v) providing the surgical parameters associating the femur to the pelvis and measured as a difference between the first and second frames of reference of the femur. 2. The method according to claim 1, wherein the femoral axis is a mechanical axis, and wherein obtaining a the reference orientation comprises the reference orientation having the mechanical axis of the pre-reduction frame of reference of the femur at a predetermined angle of a vertical axis of a patient within a frontal plane of the frame of reference of the pelvis. 3. The method according to claim 2, wherein obtaining the reference orientation comprises the reference orientation having the mechanical axis where the predetermined angle is 3°. 4. The method according to claim 2, wherein digitizing the frame of reference of the pelvis comprises digitizing two points on an iliac crest and a point on one of pubic tubercles to obtain the frontal plane of the frame of reference of the pelvis. 5. The method according to claim 2, wherein obtaining the reference orientation comprises having the mechanical axis of the pre-reduction frame of reference of the femur pass through a center of rotation of the femur, and the reference orientation has the center of rotation of the femur and of an acetabulum coincident. 6. The method according to claim 1, wherein the step iv) includes updating the frame of reference of the pelvis by digitizing a center of rotation of the acetabulum subsequently to the implant reduction. 7. The method according to claim 1, wherein the surgical parameters are at least one of a limb length discrepancy and a medio-lateral offset. 8. The method according to claim 1, wherein the steps i) and ii) are performed with respect to a single trackable reference secured to the pelvis. 9. The method according to claim 1, wherein the method is performed on a bone model or a cadaver. | Orthosoft Inc. | A61B005/103 A61B005/117
US08333190 | B2 | 2012-12-18 | Nebulizer having flow meter function | 2011-01-06 | In accordance with non-limiting examples, a nebulizer includes a main body comprising an air channel section and further comprising a mixing chamber and a venturi positioned to be placed within the patient's oral cavity and configured to receive medicine and air and mix the medicine and air within the mixing chamber and receive the air flow through the venturi and cause the medicine entering the mixing chamber to be atomized by the action of air flowing through the venturi. An air flow sensor is associated with the main body and configured to measure the air flow created by the patient's one of at least inhaling and exhaling air. | 1. A method of administering a medicine to a patient using a nebulizer, comprising: nebulizing the medicine using a venturi placed within the patient's oral cavity, the nebulizer comprising a main body and at least one medicine receiver configured to receive a medicine container, a mixing chamber and a venturi positioned to be located within the oral cavity of a patient when used and extending into the mixing chamber and connected to the medicine receiver and configured to receive medicine from the medicine receiver and air and mix the medicine and air in the mixing chamber, and a diffuser element positioned at the mixing chamber and configured to receive the flow of air from the venturi and cause the medication entering the mixing chamber to be atomized by the action of the air flowing through the venturi;measuring the air flow created by the patient's one of at least inhaling and exhaling air; andprocessing the measured air flow over time to determine a respiratory function of the patient. 2. The method according to claim 1, and further comprising forming the mixing chamber as a rainfall chamber and positioning a diffuser within the rainfall chamber upon which the medicine that is nebulized impacts and configured to break up droplets of medicine expelled from the venturi nozzle into smaller sizes. 3. The method according to claim 1, and further comprising forming a plurality of medicine receivers on the main body, each shaped to match a shape uniquely associated with a different medicine receiver, and receiving medicine from a medicine container received within the medicine receiver. 4. The method according to claim 1, and further comprising supplying a curtain of air above and below the nebulized medicine and air to enhance penetration of nebulized medicine into the airway of the patient. 5. The method according to claim 1, and further comprising attaching a separate flow meter device onto the nebulizer body after nebulizing, and measuring the air flow created by the patient's one of at least inhaling and exhaling air using the separate flow meter device attached to the nebulizer. 6. The method according to claim 1, and further comprising determining the velocity of air over time during the one of at least inhaling and exhaling air. 7. The method according to claim 1, and further comprising determining the pressure of the air flow over time during one of at least inhaling and exhaling air. 8. The method according to claim 1, and further comprising adjusting the resistance to the air flow through the nebulizer to a predetermined level for respiratory exercise training and incentive spirometry use. 9. The method according to claim 1, and further comprising measuring involuntary reflex cough as induced by supplying to the throat of the patient an aerosol containing a chemo-irritant agent effective for inducing at least one involuntary cough as an involuntary reflex and measuring air flow during the period of the involuntary cough and proc... | Pneumoflex Systems, LLC | A61M015/00
US07945317 | B2 | 2011-05-17 | Bioimpedance methods and apparatus | 2004-09-10 | Methods and apparatus for providing bioimpedance analysis are provided. In certain aspects, equivalent circuit frequency response models are provided which lead to improved correlations with MRI data. The frequency response models take account of body composition, including the fat component of a body segment. Data obtained by performing bioimpedance spectroscopy (BIS) and MRI on the calves of subjects illustrates the improved correlations achieved compared to single frequency analyses at 50 kilohertz and analyses performed using the conventional Cole-Cole model. | 1. A method for analyzing bioimpedance data for a body segment of a subject, said body segment having an external skin surface, said method comprising: (a) applying alternating current at a plurality of frequencies to at least two points on the external skin surface so as to cause current to pass through the segment;(b) for each frequency, recording the voltage difference between at least two other points on the external skin surface, said recorded voltage differences comprising both magnitude and phase information; and(c) using the recorded voltage differences at the plurality of frequencies to determine at least one numerical value indicative of the muscle, fat, and/or extracellular fluid content of the segment, said numerical value being determined using an impedance model for the segment which at least comprises three parallel paths, one of which consists of a capacitor CM and a resistor RI in series which represent primarily the muscle component of the segment, one of which consists of a capacitor CF and a resistor RF in series which represent primarily the fat component of the segment, and one of which consists of a resistor;wherein the three parallel path impedance model is used at least one frequency less than or equal to 10 kilohertz and at least one frequency greater than 10 kilohertz. 2. The method of claim 1 wherein the at least one numerical value is indicative of the muscle content of the segment and is obtained from a correlation equation which transforms an RI value to a muscle content value for the segment. 3. The method of claim 2 wherein the correlation equation is obtained by: (i) performing steps (a), (b), and (c) on a plurality of calibration subjects to obtain an RI value for each of said subjects;(ii) performing magnetic resonance imaging of the segment for the plurality of calibration subjects to obtain a value for the muscle content of the segment for each of said subjects; and(iii) performing a regression analysis on the values obtained in steps (i) and (ii) to obtain the correlation equation. 4. The method of claim 3 wherein the regression analysis yields an R2 value of at least 0.5 for the correlation equation. 5. The method of claim 4 wherein the plurality of calibration subjects includes at least one calibration subject having a body mass index less than 20 and at least one calibration subject having a body mass index greater than 35. 6. The method of claim 1 wherein the at least one numerical value is indicative of the fat content of the segment and is obtained from a correlation equation which transforms an RF value to a fat content value for the segment. 7. The method of claim 6 wherein the correlation equation is obtained by: (i) performing steps (a), (b), and (c) on a plurality of calibration subjects to obtain an RF value for each of said subjects;(ii) performing magnetic resonance imaging of the segment for the plurality of calibration subjects to obtain a value for the fat content of the segment for ea... | Renal Research Institute, LLC | A61B005/05 A61B005/103 A61B005/117
US07970452 | B2 | 2011-06-28 | Open architecture imaging apparatus and coil system for magnetic resonance imaging | 2006-08-28 | Apparatus and method for using radio frequency coil systems for magnetic resonance imaging within an open architecture apparatus is provided. The MRI coil system includes a support structure with an open architecture in which secondary support structures, compression systems and plates containing RF coil systems may be introduced. These structures and RF coils can be moved relative to the patient, or removed entirely from the system. In one embodiment the system consists of a tabletop coil system, while another embodiment consists of a dedicated stretcher design. | 1. An apparatus for positioning and immobilizing an anatomy of interest of a patient and supporting the patient for MR imaging using an MRI scanner, the apparatus comprising: a patient support structure configured to position the anatomy of interest of the patient during MR imaging, the patient support structure including a contoured upper surface for supporting the patient comfortably;a first and a second compression plate;a mounting having a locking mechanism, movably engaged to one of the compression plates for receiving and holding a RF coil, the mounting movable to adjust the position of the RF coil relative to the one of the compression plates engaged thereon and the locking mechanism for locking the RF coil in a fixed position on the one of the compression plates when in a locked position;at least one attachment member for attaching the first and the second compression plates to the patient support structure, the at least one attachment member being configured to allow a user to adjust the position of the first and the second compression plates relative to each other to immobilize the anatomy of interest such that when in use the sensitive volume of the RF coil includes the anatomy of interest;a receptacle coupled to at least one of the patient support structure and the one of the compression plates, for receiving RF signals from the RF coil while the anatomy of interest remains immobilized by the first and the second compression plates;a secondary support structure housing a second RF coil, the secondary support structure when engaged with the patient support structure supports at least a portion of the weight of the patient; andan MRI connector engaged with the patient support structure and connected to the RF coil through the receptacle and connected to the second RF coil, the MRI connector for communication with a complementary connector on the MRI scanner for data acquisition by the MRI scanner of RF signals from at least one of the RF coil and the second RF coil for MRI imaging. 2. The apparatus of claim 1, wherein the MRI connector comprises at least one of an electrical, a wireless and an optical communication device for connection with the complementary connector of the MRI scanner, the complementary connector comprising at least one of a corresponding electrical, wireless or optical MRI communication device to the MRI connector. 3. The apparatus of claim 1, wherein the MRI connector comprises a mechanical alignment device for removable connection to a mating connector on the complementary connector of the MRI scanner, the mating connector comprising a corresponding mating mechanical alignment device, wherein the mechanical alignment device and the mating mechanical alignment device provide an alignment for removable connection of the MRI connector and complementary connector of the MRI scanner. 4. The apparatus of claim 3, wherein the mechanical alignment device comprises at least one of a lever, a cam, a linear actuator, a ro... | Hologic, Inc. | A61B005/05 A47B013/00
US07937139 | B2 | 2011-05-03 | Systems and methods of utilizing electrical readings in the determination of treatment | 2004-07-20 | The present invention relates to a system for determining treatment options from at least two electrical readings. The electrical readings are conductivity measurements of a particular region on the human body. The system utilizes a correlation algorithm to determine the diagnosis which can easily be correlated with appropriate treatments. The correlation algorithm may include the analysis of multiple electrical readings in determining the diagnosis. The system may also utilize a database of clinical data to further assist in determining the diagnosis. | 1. A method for determining appropriate treatments based on electrical readings, comprising: measuring at least two measured electrical readings from acupuncture points on a human body;accessing a clinical information database containing a plurality of patient listings of individual patient data, each patient listing comprising: at least two patient electrical readings from acupuncture points of an individual patient; anda known diagnosis of a patient ailment for that patient, the known diagnosis being associated with the at least two patient electrical readings;automatically using a correlation algorithm to generate a matching diagnosis of a particular ailment by searching the clinical information database for patient listings having patient electrical readings similar to the at least two measured electrical readings and selecting the associated known diagnosis of one of the patient listings having patient electrical readings similar to the measured electrical readings; anddetermining treatment information by correlating the matching diagnosis to a set of known treatment information corresponding to the matching diagnosis. 2. The method of claim 1, wherein measuring at least two measured electrical readings further includes measuring an impedance from a particular location on the human body to a second location on the human body. 3. The method of claim 1, wherein measuring at least two measured electrical readings further includes measuring a resistance from a particular location on the human body to a second location on the human body. 4. The method of claim 1, wherein each patient listing further comprises patient symptom information and wherein automatically using a correlation algorithm to generate the matching diagnosis of the particular ailment further includes: receiving input symptom information;searching the clinical information database for patient listings having patient symptom information similar to the input symptom information in addition to the similar at least two electrical readings; andgenerating diagnosis information based on the known diagnosis of the patient ailment of the associated patient listing. 5. The method of claim 4, wherein receiving input symptom information comprises at least one of: prompting a user to select symptoms; andproviding a set of predetermined questions designed to elicit exact symptom information. 6. The method of claim 1, wherein the known diagnosis of the patient ailment for each of the patient listings includes homeopathic ailments. 7. The method of claim 1, wherein the known diagnosis of the patient ailment for each of the patient listings includes medical ailments and conditions. 8. The method of claim 1, wherein the known diagnosis of the patient ailment for each of the patient listings includes chiropractic conditions. 9. The method of claim 1, wherein determining treatment information by correlating the matching diagnosis to the set of known treatment information corresponding to the... | BioMeridian International, Inc. | A61B005/05 A61B005/04 A61H039/02
US07963930 | B2 | 2011-06-21 | Methods and devices of multi-functional operating system for care-taking machine | 2009-03-11 | The methods and devices of a multi-functional operating interface for a care-taking machine. The multi-functional operating interface of the care-taking machine includes: muscle stretch sensor, multiplexer amplifying wave filter, analog-to-digital signal converter, image processing unit, and muscle pattern database and control bus. Aided by the present muscle template training method and potential image data collecting and processing method, it can become an interface that offers the user of a care-taking machine instantaneous and unlimited controllability of the machine, enabling such people, either sick or handicapped, to enjoy more convenience and better life. | 1. A method for constructing and processing potential images, used to process, analyze, and identify electric signals detected by one or more detecting devices; said construction and processing method further comprises the following steps: detecting electric signals by the one or more detecting devices;processing the electric signals by a multi-task amplifying filter and analog-to-digital converter to form digitized electric signals;receiving said digital electric signals by an image processing unit;converting said digital electric signals into grey scale pixel signals to form potential images by the image processing unit;pre-processing said potential images by highlighting characteristics of said potential images to identify iso-potential lines and iso-potential points via the image processing unit;grouping said potential images according to their characteristics, and tallying said potential images with the characteristics in a muscle template database by the image processing unit;generating movement commands, by giving movement commands control bus to conduct various controls of said care-taking machine; andobtaining a next step, by going back to step of receiving said digital electric signals by the image processing unit and starting a next control cycle. 2. The method for constructing and processing potential images as claimed in claim 1, wherein said step of converting said digital electric signals is done by normalizing irregularly or regularly shaped electric voltage signals into 256 categories and then re-arranging them in grey scaled dot matrix images. 3. The method for constructing and processing potential images as claimed in claim 1, wherein said muscle template database is generated by a muscle template training method, and said muscle template training method comprises the following steps: defining specific muscle movements and control commands, wherein a trainer communicates with a trainee to define what characteristics of muscle pattern represent what said control commands and then the trainer instructs stimulations to the trainee as to how to stimulate/produce said specific muscle movements as defined; andgenerating muscle template data, for generating said electric signals through a muscle stretch sensor by the specific muscle movements; then processing said electric signals by said image processing unit to highlight the characteristics of said electric signals and producing a set of electric signals with said characteristics; and then storing said electronic signals with said characteristics of muscle pattern, together with corresponding control commands of said specific muscle movements, are then stored in said muscle template database. | Industrial Technology Research Institute | A61B005/103 A61B005/117
US07962197 | B2 | 2011-06-14 | Radiation ablation tracking system and method | 2010-06-17 | A system and method for treating a target tissue region (e.g., malignant tissue) is provided. A hyperthermic probe is placed into contact with the target tissue region, and the target tissue region is exposed to a therapeutic x-ray radiation beam. The probe is operated to increase the temperature of the target tissue region, thereby facilitating a therapeutic effect of the radiation beam. Image data of the probe containing a fiducial datum is acquired while in contact with the target tissue region, a position of the target tissue region within a treatment coordinate system is determined based on the fiducial datum, and the radiation beam is spatially adjusted relative to the target tissue region based on the determined position of the target tissue region. | 1. A tissue treatment system, comprising: a tissue ablation subsystem comprising a hyperthermic probe configured to be placed in contact with a target tissue region and an radio frequency generator operatively coupled to the ablation probe, the probe comprising a plurality of electrodes configured to converge in a single location comprising a fiducial datum;a radiotherapy/imaging subsystem comprising a radiotherapy device configured for exposing the target tissue region to an x-ray radiation beam and an image console configured to acquire and display image data as a two-dimensional image, the image data containing the fiducial datum; anda tracking and positioning subsystem configured for tracking inadvertent movements of the target tissue region within a treatment coordinate system based on the fiducial datum, the tracking and positioning subsystem configured to spatially adjust the radiation beam relative to the target tissue region based on the determined position of the target tissue region. 2. The tissue treatment system of claim 1, wherein the hyperthermic probe comprises an elongated cannula and an inner probe slideably disposed within the cannula, the plurality of electrodes disposed at a distal end of the inner probe. 3. The tissue treatment system of claim 1, the radiotherapy/imaging subsystem comprising a moveable gantry configured to move relative to a table. 4. The tissue treatment system of claim 3, wherein the gantry comprises a C-arm with an x-ray source mounted at one of the C-arm and an image acquisition device mounted to the other end of the C-arm. 5. The tissue treatment system of claim 4, further comprising a motor/drive unit configured to move the x-ray source relative to the target tissue region. 6. The tissue treatment system of claim 3, further comprising a motor/drive unit supporting the C-arm. 7. The tissue treatment system of claim 6, wherein the tracking and positioning subsystem is operatively connected to the motor/drive unit. 8. The tissue treatment system of claim 3, wherein the table is configured for rectilinear translation. 9. The tissue treatment system of claim 1, wherein the tracking and positioning subsystem comprises a positioning processor and a controller configured for spatially adjusting the radiation beam by adjusting at least one of an absolute position and orientation of the radiation beam. 10. The tissue treatment system of claim 9, wherein the positioning processor and controller are incorporated into the image console. 11. The tissue treatment system of claim 1, wherein the x-ray radiation beam comprises a phased array. 12. The tissue treatment system of claim 1, wherein the plurality of electrodes comprise an array. 13. A method of treating a target tissue region of a subject's body, comprising: locating a subject on a table at least partially surrounded by a gantry having a radiation source and an imaging device;adjusting the gantry to place an isocenter of the radiation beam to inte... | Boston Scientific Scimed, Inc. | A61B005/05
US07954251 | B2 | 2011-06-07 | Multifunctional airway evaluator for orotracheal intubation | 2010-02-10 | A multifunctional airway evaluator for orotracheal intubation includes first and second connected elongated arms, movable relative to each other about a pivot axis, a planar plate connected to the first arm and extending upwardly therefrom, and a gravity-operated angular indicator attached to the planar plate. The plate includes a distance scale on its front face, and the distance scale includes a plurality of spaced apart grade lines aligned radially with the pivotal axis and distance parameters associated with the grade lines, each of the distance parameters indicating a distance between the distal ends of the first and second arms when the arms are in a relative position. The gravity-operated angular indicator includes an axial pin disposed perpendicular to the planar plate, a rotating disk pivotally attached to and rotatable around the axial pin by gravity, and one or more angular measurement marker provided adjacent the rotating disk. | 1. A method of evaluation of airway for orotracheal intubation comprising: (a) providing a multifunctional airway evaluator comprising:a first elongated arm having a distal end and an opposing hinge end;a second elongated arm having a distal end and an opposing hinge end connected to said hinge end of said first arm by a pivot connection; said first and second arms movable relative to each other about a pivotal axis of said pivot connection;a planar plate connected to said first arm and extending upwardly from an upper edge of said first arm, between said hinge end and a middle portion of said first arm, said plate including a distance scale on a front face of said planar plate, said distance scale including a plurality of spaced apart grade lines, each thereof aligned radially relative to said pivotal axis, and distance parameters associated with said grade lines, each of said distance parameters indicating a distance between said distal ends of said first and second arms when said arms are in a corresponding relative position; anda gravity-operated angular indicator attached to said planar plate, said gravity-operated angular indicator including an axial pin disposed perpendicular to said planar plate, a rotating disk pivotally attached to and rotatable around said axial pin by gravity, and one or more angular measurement marker affixed adjacent said rotating disk;(b) placing said distal ends of said first and second arms next to a first and a second predetermined locations of the body of a patient, respectively, and recording a measured distance between said first and second predetermined locations from said distance scale; and(c) affixing said planar plate at a third predetermined location of the body of said patient, with a vertical central axis of said rotating disk aligned with a zero-degree reference line of said gravity-operated angular indicator, and then instructing said patient to carry out a flexion movement involving tilting of said third location of the body, and recording an extent of rotation of said rotating disk from said angular measurement marker of said gravity-operated angular indicator. 2. The method of claim 1, wherein said first and second predetermined locations of the body are mandibular and maxillary central incisors, respectively, while the mouth of said patient is fully opened, and said measured distance is mouth opening distance of said patient; orsaid first and second predetermined locations of the body are the thyroid notch and bone point of the mentum of the mandible of said patient, respectively, while the neck of said patient is in full extension, and said measured distance is thyromental distance of said patient; orsaid first and second predetermined locations of the body are the left mandibular angle and the right mandibular angle of said patient, respectively, and said measured distance is mandibular angle distance of said patient. 3. The method of claim 1, wherein said third predetermined location is the... | Unique Surgical Innovations, LLC | A61B005/103
US07941324 | B1 | 2011-05-10 | Method and system for identification of a patient | 2007-04-26 | A method for obtaining healthcare that includes submitting biometric identification data to a Personal Health History (PHH) account that includes a PHH for a patient, permitting a healthcare provider to obtain the biometric identification data, wherein the healthcare provider accesses the PHH account with the biometric identification data to obtain the PHH, and receiving a treatment for a health condition of the patient, wherein the treatment is administered by the healthcare provider accessing the PHH. | 1. A method for obtaining healthcare comprising: obtaining, by a healthcare provider, biometric identification data from a non-communicative patient;generating, by the healthcare provider using a computer processor, an identification request comprising the biometric identification data;sending, via a computer network, the identification request to a Personal Health History (PHH) framework comprising a PHH account for the non-communicative patient, wherein the PHH framework selects the PHH account identified by the biometric identification data from a plurality of PHH accounts, and wherein the PHH account for the non-communicative patient comprises a common identifier;receiving, by the healthcare provider, the common identifier associated with the PHH account from the PHH framework;identifying the non-communicative patient by the healthcare provider using the common identifier; andadministering a treatment for a health condition of the non-communicative patient by the healthcare provider. 2. The method of claim 1, further comprising: submitting, by the healthcare provider, healthcare data for the non-communicative patient to the PHH account. 3. The method of claim 2, wherein the healthcare data comprises data associated with at least one selected from a group consisting of a medication, a diagnosis, the treatment, and a medical note. 4. The method of claim 1, further comprising: updating, by the healthcare provider, the PHH for the non-communicative patient based on the treatment received. 5. The method of claim 1, wherein the biometric identification data is data derived from a group consisting of a scan of a print and a retinal scan. 6. The method of claim 1, wherein the biometric identification data is data defining a deoxyribonucleic acid (DNA) sequence. 7. The method of claim 1, wherein the non-communicative patient is in an emergency room. 8. The method of claim 1, wherein the PHH framework verifies that the healthcare provider has permission to access the PHH account, wherein the healthcare provider is one of the plurality of healthcare providers. 9. A method for patient identification comprising: receiving, via a computer network by a Personal Health History (PHH) framework executing on a computer processor, an identification request from a healthcare provider, wherein the identification request comprises biometric identification data for a non-communicative patient;selecting, by the PHH framework, a PHH account identified by the biometric identification data from a plurality of PHH accounts, wherein the PHH account for the non-communicative patient comprises a common identifier;obtaining, by the PHH framework, the common identifier from the PHH account; andsending via the computer network by the PHH framework the common identifier to the healthcare provider, wherein the non-communicative patient receives treatment for a medical condition by the healthcare provider using the common identifier. 10. The method of claim 9, further co... | Intuit Inc. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US08425434 | B2 | 2013-04-23 | Method and apparatus for monitoring animal condition | 2009-07-27 | A method and system for monitoring the condition of animals provides for current height and weight measurements taken from an animal in order to automatically derive a height-to-weight ratio. The height-to-weight ratio thus obtained may be used to determine whether or not an animal's overall physical condition lies inside tolerance limits in relation to its predetermined expected development. An apparatus for implementing the method of the invention may include a control system configured to determine the height-to-weight ratio and may within a control system of an animal installation. The supply of feed to an animal may be adjusted in response to a deviation of its measured height weight ratio from a target or benchmark value. | 1. A method for automatically determining a condition of an animal, the method comprising the steps of: moving an animal to be measured adjacent a height gauge, the height gauge configured for measuring a height of the animal located on a level reference surface, the height gauge comprising a control device (10), comprise memory storing an algorithm, and a processor, and animal files relating to individual animals;automatically determining a identity of the animal using the height gauge, the height gauge further comprising a transponder receiver (14), by the transponder receiver co-operating with a transponder (20) worn by the animals;automatically determining a current height of the animal using the height gauge, the height gauge further comprising five receivers (7) positioned vertically spaced from each other at a regular interval (a) along a length direction axis (29) of the height gauge, the receivers measuring one of a signal emitted from the animal and a reflected signal reflected from at least one of a reflector panel and animal, based on signals received by the receivers and using the algorithm in the memory of the control device, making a determination of the height of the animal;automatically storing the determined current height in the control device;automatically determining a current weight of the animal and storing the determined current weight in the control device;automatically deriving, using the processor, a height-to-weight ratio of the animal from the stored current height and stored current weight;automatically storing the derived height-to-weight ratio in the control device;automatically comparing the derived height-to-weight ratio value with a predetermined expected height-to-weight ratio to obtain a result;automatically storing the result in the control device; andautomatically using the stored result to determine a condition of the animal, wherein,in the moving step, the animal is moved between the height gauge and the reflector,in the determining a current height step, the receivers measure a reflected signal reflected from the reflector panel and plural reflected signals reflected from the animal, andthe method comprises the further step of using the height gauge to emit signals, which are reflected off the reflector panel and off the animal, by plural emitters emitting the signals toward the animal and the reflector panel, the emitters positioned vertically spaced from each other at the regular interval (a) along the length direction axis (29) of the height gauge. 2. The method of claim 1, wherein the emitters emit laser radiation and the receivers measures reflected laser radiation. 3. A method for automatically determining a condition of an animal, the method comprising the steps of: moving an animal to be measured between a height gauge and a reflector panel, the height gauge configured for measuring a height of the animal located on a level reference surface, the height gauge comprising a control device (10), co... | DeLaval Holding AB | A61B005/103 A61B001/00 A01K001/10
US08374674 | B2 | 2013-02-12 | Nerve treatment system | 2011-02-01 | A method of treatment includes delivering an agent into a patient, and providing energy inside the patient so that the energy can interact with the agent to treat a nerve. A system for treatment includes a source of agent for delivery into a patient, and an energy source for providing energy inside the patient so that the energy can interact with the agent to treat a nerve. | 1. A treatment system, comprising: a therapeutic ultrasonic energy source comprising a transducer; anda processor coupled to the transducer, wherein the processor is configured to control the transducer using information regarding a blood vessel from an imaging device to apply ultrasonic energy to a nerve region next to the blood vessel;wherein the processor is configured to operate the transducer to apply the ultrasonic energy to treatment areas that are arranged approximately in a spherical or ellipsoidal configuration around a reference position associated with the blood vessel, to thereby treat nerve tissue in the nerve region on at least two sides of the blood vessel so that the nerve tissue on the at least two sides of the blood vessel receives more therapeutic effect than the blood vessel. 2. The system of claim 1, further comprising a module to track motion of the blood vessel. 3. The system of claim 1, wherein the nerve tissue extends around the blood vessel, and wherein the processor is configured to cause the transducer to create a zone of heat surrounding the blood vessel to create an annular region around the blood vessel to modulate the nerve tissue surrounding the blood vessel. 4. The system of claim 1, wherein the transducer is configured to apply power to the nerve region, the power having a value that is anywhere from 100 W/cm2 to 2500 W/cm2. 5. The system of claim 1, further comprising a patient interface configured to position the therapeutic ultrasonic energy source so that regions of bone are avoided by the ultrasonic energy from the transducer. 6. The system of claim 1, wherein the nerve region is an autonomic nerve region, and the transducer is configured to apply the ultrasonic energy at an energy level to create a reduction in a norepinephrine output of the autonomic nerve region. 7. The system of claim 1, further comprising a motion tracking module to track a movement of the blood vessel. 8. The system of claim 7, wherein the motion tracking module is configured to track speckles. 9. The system of claim 7, wherein the motion tracking module is configured to determine a three dimensional coordinate of the blood vessel. 10. The system of claim 7, wherein the motion tracking module is configured to track a doppler ultrasound signal from the blood vessel. 11. The system of claim 1, wherein the processor is configured to operate the transducer so that the transducer applies the ultrasonic energy simultaneously to the nerve tissue on the at least two sides of the blood vessel. 12. The system of claim 1, further comprising the imaging device, wherein the imaging device is configured to visualize the blood vessel in the nerve region with ultrasound. 13. The system of claim 1, further comprising the imaging device, wherein the imaging device is configured to detect a presence of the blood vessel by a Doppler signal. 14. The system of claim 1, further comprising the imaging device, wherein the imaging device comprise... | Kona Medical, Inc. | A61B005/05
US08838217 | B2 | 2014-09-16 | System and apparatus for providing diagnosis and personalized abnormalities alerts and for providing adaptive responses in clinical trials | 2009-11-10 | A personalized real-time automated cardiovascular monitoring system monitors abnormalities in a patient's cardiovascular activity data through the use of individually adjusted electrocardiogram Holter apparatus (Holter/ECG device) that provides an automatic medical diagnosis of cardiac abnormalities and generates abnormality alert signals representative of certain abnormalities in patient's cardiac activities. The signals are transmitted using a wireless network to a medical dispatcher center. A response is generated according to the abnormality detected. Individual parameters indicative of patient's cardio activities are personalized to allow for adjustments of chronic patients. A base Holter/ECG unit, includes the wireless/electric electrodes and their respective wireless/electric connections, and a Holter/ECG recording unit affixed to the base unit. Automatic real-time medical response may be provided based on automatic cardiac abnormality alert detection from the Holter/ECG data. | 1. A method for providing cardiovascular personalized patient alerts for a plurality of patients, comprising the steps of: a) maintaining physiological data and cardiac data for each of said plurality of patients, respectively, said cardiac data selected from the list consisting of ECG, EKG, blood pressure, heart rate;b) creating personalized resting alert thresholds for each of the patients, respectively, based on the physiological data and the cardiac data when each of said patients are respectively at rest;c) providing each patient with a respective wireless Holter/ECG-type device which is worn on each patient, each Holter/ECG-type device having respective buttons including a Medicine button for pressing when the patient has taken any particular medicine, a Sports button for pressing when the patient will be engaged in a sporting activity; a Meal button for pressing when the patient will have a meal and; a Sleep button for pressing when the patient will start/finish his sleep;d) providing each of said patients with respective abnormality alert thresholds, wherein each of said thresholds is respectively different between each of said patients based on each patients' respective physiological data or cardiac data wherein the alert thresholds being exceeded is detected based on said physiological data or cardiac data which is obtained;e) updating said alert thresholds provided in step d) for each of said patients, based on said respective resting alert thresholds, respective cardiac data provided by each Holter/ECG-type device as a result of monitoring each of said patients, respective abnormality alert thresholds and actuation of said respective buttons;f) generating, without human involvement, a respective diagnosis for each of said patients based on data provided by each of said Holters/ECGs, said updated alert thresholds, and whether any of said buttons have been pushed, wherein said respective diagnosis is generated by a respective one of said Holters/ECGs prior to step g); and;g) transmitting said respective diagnosis from said wireless Holter/ECG-type device, for a medical response to be provided at a location where one of said patients is located if said generated diagnosis is predetermined to be responded to by the medical response, wherein said medical response includes requesting providing of an ambulance to said location where said one of said patients is located. 2. The method of claim 1, further comprising the step of: h) storing said physiological data and said cardiac data for each of said patients, respectively, at a centralized database. 3. The method of claim 2, wherein step h) further comprises: signaling via a cellular network to the centralized database at a medical emergency center. 4. The method of claim 2, wherein the wireless device further includes means for analyzing statistical data received from the centralized database. 5. The method of claim 4, wherein the wireless device includes means for sending the analyzed ... | Makor Issues and Rights Ltd. | A61B005/00 A61B005/04
US08798719 | B2 | 2014-08-05 | Method of utilization of high dielectric constant (HDC) materials for reducing SAR and enhancing SNR in MRI | 2010-12-01 | Layers or coats of materials with high dielectric constant or permittivity with very low conductivity are inserted in between radiofrequency (RF) coil or coil's conductive elements and the sample to enhance the signal to noise ratio (SNR), improve image contrast, and reduce the specific absorption rate (SAR) of magnetic resonance imaging or magnetic resonance spectroscopy instruments. The embodiments of the present invention can be used as an auxiliary device to the standard pre-constructed RF coils or incorporated with RF coil constructions for enhancing RF coil performances in both transmission and reception. | 1. A spectroscopic method of analyzing a sample, the method comprising: identifying a region to be analyzed;applying a radiofrequency field from a radiofrequency coil through a high dielectric constant material to the region;and gathering information about the region with a computer, the information gathered from the application of the radiofrequency field to the region;wherein the radiofrequency field includes an electric field and a radiofrequency magnetic field, the radiofrequency magnetic field generated by the electric field;wherein the radiofrequency coil is positioned with respect to the high dielectric constant material, such that the radiofrequency magnetic field extends through both the region and the high dielectric constant material and increases the signal to noise ratio of all the information gathered. 2. The method of claim 1, further comprising applying the radiofrequency field during magnetic resonance imaging. 3. The method of claim 1, wherein the radiofrequency coil is a phase array coil for a spine, a head, a knee, or an extremity. 4. The method of claim 1, wherein the dielectric includes one or more of a solid, a liquid, a jell, a suspension, and a paste. 5. The method of claim 1, further comprising positioning the high dielectric constant material directly on the radiofrequency coil. 6. The method of claim 1, wherein the radiofrequency field is applied during nuclear magnetic resonance analysis. 7. The method of claim 1, wherein the high dielectric constant material results in a reduction of required radiofrequency power by at least 50%. 8. The method of claim 1, wherein the high dielectric constant material results in an increase in signal to noise ratio by between 20% and 40%. 9. The method of claim 1, wherein the high dielectric constant material is within a pad. 10. The method of claim 9, further comprising conforming the pad to body parts of a mammal. 11. The method of claim 9, further comprising positioning the pad around the neck of a human. 12. The method of claim 9, further comprising positioning the pad around extremities of a mammal. 13. The method of claim 9, wherein the pad is planar. 14. The method of claim 9, further comprising shaping the pad to conform to a human breast. 15. The method of claim 9, further comprising shaping the pad to conform to a pediatric human. 16. The method of claim 9, further comprising shaping the pad to conform to an adult human. 17. A spectroscopic method of analyzing a sample, the method comprising: identifying a region to be analyzed;applying a radiofrequency field from a radiofrequency coil through a high dielectric constant material to the region; andgathering information about the region from the application of the radiofrequency field to the region with a computer;wherein the radiofrequency field includes an electric field and a radiofrequency magnetic field, the radiofrequency magnetic field generated by the electric field;wherein the radiofrequency coil is po... | - | A61B005/00
US08615286 | B2 | 2013-12-24 | In vivo sensor for detecting bone surface | 2010-07-14 | A device and a method are described for finding the location of a bone surface in patients or living animals by using a thin probe equipped with bone contacting detection functionality for a minimally invasive procedure. | 1. Apparatus for determining the position, in three dimensional space, of a bone covered by soft tissue, said apparatus comprising: a needle having a tip for penetrating soft tissue;position-determining means for determining the position of said tip of said needle in three-dimensional space;force-sensing means for sensing the level of force encountered by said tip of said needle;discriminating means for determining the rate of change in the level of force encountered by said tip of said needle in relation to the distance travelled by said needle so as to discriminate between (i) when said tip of said needle is engaging soft tissue, and (ii) when said tip of said needle is engaging bone;recording means for recording the position of said tip of said needle in three dimensional space when said discriminating means indicate that said tip of said needle is engaging bone; andprocessing means for receiving a plurality of data points from said recording means and, using the plurality of data points from said recording means, determining the position, in three dimensional space, of a bone covered by soft tissue. 2. A method for determining the position, in three dimensional space, of a bone covered by soft tissue, said method comprising: directing a tip of a needle through the soft tissue and toward the bone while sensing the level of force encountered by said tip of said needle;determining the rate of change in the level of force encountered by said tip of said needle in relation to the distance travelled by said needle;identifying a spike in the rate of change in the level of force encountered by said tip of said needle in relation to the distance travelled by said needle, wherein the spike is reflective of said tip of said needle encountering bone;recording the position of said tip of said needle in three-dimensional space when said spike is identified; andrepeating the foregoing steps a sufficient number of times and determining the position, in three dimensional space, of a bone covered by soft tissue. 3. Apparatus for determining the position, in three dimensional space, of a first type of tissue which is covered by a second type of tissue, wherein said first type of tissue has a greater resistance to penetration than said second type of tissue, said apparatus comprising: a needle having a tip for penetrating tissue;position-determining means for determining the position of said tip of said needle in three-dimensional space;force-sensing means for sensing the level of force encountered by said tip of said needle when engaging tissue;discriminating means for determining the rate of change in the level of force encountered by said tip of said needle in relation to the distance travelled by said needle so as to discriminate between (i) when said tip of said needle is engaging said first type of tissue, and (ii) when said tip of said needle is engaging said second type of tissue;recording means for recording the position of said tip of said needle in ... | Curexo Technology Corporation | A61B005/055
US08447407 | B2 | 2013-05-21 | Method and system for detecting epileptogenesis | 2010-11-01 | Neuronal excitation and inhibition of the brain is tracked in the hippocampal CA1 network during a latent period, wherein biomarkers are observed which include a sustained increase in the firing rate of the excitatory postsynaptic field activity, paired with a subsequent decrease in the firing rate of the inhibitory postsynaptic field activity; the mean amplitude profiles of both fEPSP and fEPSP field potential activity during the latent period have characteristic shapes; both excitatory and inhibitory CA1 field activity firing rates are observed to follow a circadian rhythm that drifts during epileptogenesis; the circadian rhythms described are in-phase in controls and anti-phase during epileptogenesis; and the fEPSP rate drifts from a circadian rhythm to a greater extent than the fEPSP rate. An additional biomarker is a change in a circadian rhythm of core body temperature. Therapeutic measures can include thermal, chemical, or electrical modulation, in an open or closed loop process. | 1. A method of detecting an impending seizure in a patient, the method comprising the steps of: monitoring at least one neuronal field potential in the brain of the patient,including monitoring at least one changed field potential characteristic selected from the group consisting of:an increase in the frequency of excitatory spikes during a decrease in the frequency of inhibitory spikes,a phase drift of the frequency of excitatory spikes with respect to a circadian rhythm,a phase drift of the frequency of inhibitory spikes with respect to a circadian rhythm,a faster phase drift of the frequency of excitatory spikes than the phase drift of the frequency of inhibitory spikes,a substantially out of phase relationship between excitatory spikes and inhibitory spikes, anda substantially anti-phase relationship between excitatory spikes and inhibitory spikes;wherein any of said changed field potential characteristics is indicative of epileptogenesis; andcarrying out a therapeutic response if epileptogenesis is indicated, the response operative to reduce the likelihood of a seizure. 2. The method of claim 1, further comprising: producing a response operative to reduce the likelihood of a seizure, based on a monitored change in the at least one selected field potential characteristic, said response selected from the group consisting of:changing a temperature in the brain; releasing a chemical in the body; and administering an electrical impulse within the body. 3. The method of claim 2, wherein said produced response is changed over time, in response to further monitoring, in a closed-loop process. 4. A method of detecting epileptogenesis in a patient, comprising the steps of: monitoring an fEPSP and an fIPSP of the hippocampus of the patient over time, to detect at least one change selected from the group consisting of:an increase in the frequency of fEPSP discharges during a decrease in the frequency of fIPSP discharges,a phase drift of the frequency of fEPSP discharges with respect to a circadian rhythm,a phase drift of the frequency of fIPSP discharges with respect to a circadian rhythm,a faster phase drift of the frequency of fEPSP discharges than the phase drift of the frequency of fIPSP discharges,a substantially out of phase relationship between fEPSP discharges and fIPSP discharges, anda substantially anti-phase relationship between fEPSP discharges and fIPSP discharges;wherein each of said monitored changes is further indicative of epileptogenesis, whereupon therapeutically beneficial measures are taken on behalf of the patient selected from the group consisting of changing a temperature in the brain, releasing a chemical in the body, and administering an electrical impulse within the body. 5. The method of claim 4, wherein a further beneficial measure is selected from the group consisting of: notifying the patient;notifying a health care practitioner; andadministering a therapeutic remedy operative to reduce the likelihood of a seizure. 6.... | University of Florida Research Foundation, Inc. | A61N001/3605