## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍 | 專利權人 | IPC
US09336423 | B2 | 2016-05-10 | Analyte meter including an RFID reader | 2015-06-22 | A glucose monitoring system, includes a glucose sensor strip or package of strips. The strip includes a substrate and a glucose monitoring circuit that has electrodes and a bodily fluid application portion of selected chemical composition. An antenna is integrated with the glucose sensor strip. A RFID sensor chip is coupled with the glucose sensor strip and the antenna. The chip has a memory containing digitally-encoded data representing calibration and/or expiration date information for the strip. | 1. A method of determining a glucose level; the method comprising: positioning a device comprising an RFID reader within close proximity to an in vivo implanted or partially implanted glucose sensor comprising an RFID sensor chip;receiving via circuitry of the device one or more signals indicative of the glucose level from the RFID sensor chip;determining via the circuitry of the device the glucose level based on the one or more signals to provide a determined glucose level; anddisplaying the determined glucose level on a display of the device. 2. The method of claim 1, wherein the device comprises a port for receiving a glucose test strip. 3. The method of claim 1, wherein the RFID sensor chip is comprised by a continuous glucose sensor. 4. The method of claim 1, wherein the RFID reader is an integrated element of the device. 5. The method of claim 1, wherein the RFID reader is a detachable component of the device. 6. The method of claim 1, comprising communicating data from the RFID reader to the device through a multi-pin connection. 7. The method of claim 1, comprising communicating data from the RFID reader to the device wirelessly. 8. The method of claim 1, wherein data storage for the RFID reader and the device are integrated into a single location. 9. The method of claim 1, comprising operating the RFID sensor chip at about 13.56 MHz. 10. The method of claim 1, comprising operating the RFID sensor chip at from about 125 kHz to about 134 kHz. 11. The method of claim 1, comprising operating the RFID sensor chip at about 868 MHz. 12. The method of claim 1, comprising operating the RFID sensor chip at about 2.45 GHz. 13. The method of claim 1, wherein the RFID sensor chip is an active sensor chip. 14. The method of claim 1, wherein the RFID sensor chip is a passive sensor chip. 15. The method of claim 1, wherein the RFID sensor chip comprises a memory comprising digitally-encoded data. 16. The method of claim 15, wherein the data is glucose data. 17. The method of claim 1, wherein the display is a liquid crystal display (LCD). 18. The method of claim 1, wherein the display comprises a touchscreen. 19. The method of claim 1, wherein the method comprises displaying on the display a communication when the glucose level indicates hyperglycemia. 20. The method of claim 1, wherein the method comprises displaying on the display a communication when the glucose level indicates hypoglycemia. 21. The method of claim 1, wherein the method comprises displaying on the display a communication when the glucose level indicates impending hyperglycemia. 22. The method of claim 1, wherein the method comprises displaying on the display a communication when the glucose level indicates impending hypoglycemia. 23. The method of claim 1, comprising determining trending of glucose level based at least in part on the determined glucose level. 24. The method of claim 23, wherein the trending provides for a rate and/or acceleration of glucose leve... | Abbott Diabetes Care Inc. | G08B013/14 G06K007/10 A61B005/145 A61B005/1486 A61B005/1495 G01N033/487 G08C017/02 A61B005/00 A61B005/1473 A61B019/00 A61B005/15
US20160143594 | A1 | 2016-05-26 | MULTIDIMENSIONAL TIME SERIES ENTRAINMENT SYSTEM, METHOD AND COMPUTER READABLE MEDIUM | 2014-06-18 | Illness signatures are mathematically characterized by entrainment relationships among multiple time series representations of physiological processes. Such characteristics include time and phase lags, window lengths for optimum detection, which time series are most entrained with each other, the degree of entrainment relative to the rest of the large database, and the concordance or discordance of the time-varying changes. These optimum disease-specific characteristics can be determined, for example, from large, clinically well-annotated databases of time series representations of physiological processes during health and illness. These characteristics of the entrainment relationships among multiple time series representations of physiological processes are used to make mathematical and statistical predictive models using multivariable techniques such as, but not limited to, logistic regression, nearest-neighbor techniques, neural and Bayesian networks, principal and other component analysis, and others. These models are quantitative expressions that transform measured characteristics to the probability of an illness, or p(illness). | 1. A computer-implemented method for advance detection of sub-acute, potentially catastrophic illness in a patient from abnormal entrainment of multidimensional time series representations of physiological processes of the patient, comprising: calculating mathematical characteristics of relationships between a plurality of simultaneous time series representations of physiological processes of said patient in a plurality of domains;determining cross-measures of said time series representations in specified ranges of time lags and frequency bands as functions of at least two simultaneous time series representations of physiological processes;calculating rates of change of said cross-measures over a specified time window of predefined length;identifying a rank order of said cross-measures with respect to each other;identifying a rank order of said cross-measures with respect to their expected distributions;determining the concordance or discordance among said time series representations;determining the percentile or rank of the determined cross-measures with respect to a database of cross-measures;calculating a probability of an illness from a predefined multivariable statistical model that employs at least some observed cross-measure parameters and/or their percentile or rank; anddisplaying said probability of illness on a display device. 2. The computer-implemented method of claim 1, wherein said plurality of time domains in said calculating mathematical characteristics includes at least two of: time domain, frequency domain, wavelet domain, non-linear domain, phase domain, and information domain. 3. The computer-implemented method of claim 2, wherein mathematical characteristics of the time domain include at least one of autocorrelation, cross-correlation and covariance. 4. The computer-implemented method of claim 2, wherein mathematical characteristics of the frequency domain include at least one of frequency spectra using a Fourier transform, Lomb periodogram, cross-spectra, coherence, or transfer functions. 5. The computer-implemented method of claim 2, wherein mathematical characteristics of the wavelet domain include a cross-wavelet transform. 6. The computer-implemented method of claim 2, wherein mathematical characteristics of the non-linear domain include cross-entropy. 7. The computer-implemented method of claim 2, wherein mathematical characteristics of the phase domain include a Hilbert transform. 8. The computer-implemented method of claim 2, wherein mathematical characteristics of the information domain include at least one of Granger causality and mutual information. 9. The computer-implemented method of claim 1, wherein said specified ranges of time lags and frequency bands are determined empirically from a database of time series representations collected during periods of health and early stages of illness. 10. The computer-implemented method of claim 1, wherein said specified time window has a length determined in depen... |  | A61B005/00 G06F019/00
US20160135737 | A1 | 2016-05-19 | DETERMINING TREATMENT COMPLIANCE USING SPEECH PATTERNS CAPTURED DURING USE OF A COMMUNICATION SYSTEM | 2015-06-03 | Methods and systems for monitoring compliance of a patient with a prescribed treatment regimen are described. Compliance is determined based upon analysis of the patient's speech detected during use of a communication system such as a mobile telephone by the patient. Methods and systems as described herein can be used for monitoring patient compliance with a treatment for a brain-related disorder, for example. Identity of the patient as user of the communication system is determined through the use of e.g., biometric or authentication techniques. Speech data indicative of whether the patient has complied with the prescribed treatment regimen is transmitted from a circuitry-based system at the patient location to a monitoring location, where it may be reviewed by a medical caregiver or other party, for example. Patient speech may be analyzed at the patient location and/or subjected to analysis at the monitoring location. | 1. A communication system comprising: at least one audio sensor for sensing at least one audio signal including patient speech during use of the communication system by a patient at a patient location, the patient having a brain-related disorder and a prescribed treatment regimen for treating at least one aspect of the brain-related disorder;patient identification circuitry configured to determine a presence of the patient from at least one identity signal sensed at the patient location;speech detection circuitry in the communication system for processing the at least one audio signal to identify at least one section of the at least one audio signal containing speech of the patient, wherein the speech detection circuitry is configured to identify the at least one section of the at least one audio signal containing speech of the patient based at least in part on the determination of the presence of the patient by the patient identification circuitry;speech analysis circuitry in the communication system for analyzing the at least one section of the at least one audio signal to generate speech data including data indicative of whether the patient has complied with the prescribed treatment regimen; andat least one transmitting device at the patient location for transmitting a speech data signal containing the speech data including data indicative of whether the patient has complied with the prescribed treatment regimen from the patient location to a receiving device at a monitoring location. 2. The system of claim 1, wherein the at least one identity signal includes at least a portion of the at least one audio signal, wherein the patient identification circuitry is configured to analyze the at least one audio signal to determine the presence of the patient by identifying at least a portion of the at least one audio signal that resembles known speech of the patient, and wherein the speech detection circuitry is configured to identify the at least one section of the at least one audio signal containing speech of the patient by identifying speech in the at least one audio signal corresponding to presence of the patient detected from the at least one audio signal. 3.-23. (canceled) 24. The system of claim 1, including a data storage device. 25. (canceled) 26. The system of claim 1, wherein the speech analysis circuitry includes a speech processor for processing at least one section of the at least one audio signal to determine at least one speech pattern of the patient. 27. (canceled) 28. The system of claim 26, wherein the speech processor includes a speech analyzer for assessing the at least one speech pattern to determine at least one speech parameter indicative of whether the patient has complied with the prescribed treatment regimen, wherein the speech data includes the at least one speech parameter. 29. The system of claim 26, wherein the speech analysis circuitry includes a comparator for comparing the at least one speech pattern with at l... |  | A61B005/00 G10L017/00 G06F019/00
US20160120458 | A1 | 2016-05-05 | SKIN GAS MEASUREMENT DEVICE AND SKIN GAS MEASUREMENT METHOD | 2014-05-26 | A skin gas measurement device includes a skin gas collecting unit that includes a skin gas collecting space having an opening that is to be attached to a skin surface, a porous material that is for adsorbing and concentrating a skin gas component that is emitted from the skin surface into the skin gas collecting space and that allows the adsorbed skin gas component to be desorbed at a relatively low temperature, and a heater for heating the porous material; and a skin gas measurement unit for measuring the skin gas component that is desorbed from the heated porous material. | 1. A skin gas measurement device comprising: a skin gas collecting unit that includes a skin gas collecting space having an opening that is to be attached to a skin surface, a porous material that is for adsorbing and concentrating a skin gas component that is emitted from the skin surface into the skin gas collecting space and that allows the adsorbed skin gas component to be desorbed at a relatively low temperature, and a heater for heating the porous material; anda skin gas measurement unit for measuring the skin gas component that is desorbed from the heated porous material. 2. The skin gas measurement device according to claim 1, wherein the porous material is hydrophobic. 3. The skin gas measurement device according to claim 1, wherein a pore size of the porous material is larger than or equal to a molecular size of the skin gas component and less than or equal to three times the molecular size of the skin gas component. 4. The skin gas measurement device according to claim 1, wherein the porous material is a zeolite. 5. The skin gas measurement device according to claim 4, wherein a value of SiO2/Al2O3 [mol/mol] of the zeolite is greater than or equal to 10. 6. The skin gas measurement device according to claim 1, wherein the skin gas component is acetone or a molecule having a molecular size that is equal to that of acetone. 7. The skin gas measurement device according to claim 1, wherein the porous material is a zeolite, a pore size of the zeolite is larger than or equal to a molecular size of the skin gas component and less than or equal to three times the molecular size of the skin gas component, and a value of SiO2/Al2O3 [mol/mol] of the zeolite is greater than or equal to 100. 8. A skin gas measurement method comprising: a step of introducing, from an opening that is attached to a skin surface into a skin gas collecting space, a skin gas component that is emitted from the skin surface;a step of concentrating the introduced skin gas component by causing the introduced skin gas component to be adsorbed by a porous material;a step of desorbing the skin gas component that is adsorbed by the porous material by heating at a relatively low temperature; anda step of measuring a concentration or an amount of the skin gas component that is desorbed. | NTT DOCOMO, INC. THE UNIVERSITY OF TOKYO | A61B005/00
US20160120480 | A1 | 2016-05-05 | EXPERT SYSTEM TO FACILIATE SOURCE LOCALIZATION OF BRAIN ELECTRICAL ACTIVITY | 2016-01-08 | A system or method can facilitate source localization of brain electrical activity. In one example, an expert system can be utilized to evaluate the source localization results and, based on the acceptability of the results, adjust a data preparation phase to provide for acceptable source localization. | 1. A computer-implemented system to localize a source of electrical activity in a patient's body, the system comprising: a preprocessing component programmed to preprocess electrical data corresponding to electrical signals acquired from a patient;an inverse model to represent the source of electrical activity;a forward model to represent geometry and electrical characteristics of a spatial region of interest for the patient's body based on the inverse model;a source localization function programmed to compute a location for the source of electrical activity within the spatial region of interest based on the preprocessed electrical data, the inverse model and the forward model; andan expert system to evaluate an acceptability of the computed location for the source of electrical activity and to selectively adjust of at least one of the preprocessing, the inverse model and the forward model based on the evaluation of the acceptability, the expert system adapted to cause the source localization function to re-compute the location for the source of electrical activity based on the selective adjustment. 2. The system of claim 1, wherein expert system further comprises: an evaluator programmed to evaluate the acceptability of the computed location for the source of electrical activity based on predefined criteria; andan adjustment control programmed to perform the selective adjustment automatically based on preprogrammed rules applied to results of the evaluation of the acceptability. 3. The system of claim 2, wherein the adjustment control further comprises: a preprocessing control programmed to adjust the preprocessing automatically based on the rules; anda model selector programmed to choose at least one of the inverse model and the forward model based on the rules. 4. The system of claim 2, wherein if the results of the evaluation of the acceptability are unacceptable, the adjustment control is programmed to selectively at least one of the preprocessing component, the inverse model and the forward model in an order of increasing complexity for each subsequent location that the source localization component computes for the source of electrical activity. 5. The system of claim 1, wherein the preprocessing further comprises: a spectral analysis function to determine frequency characteristics of the electrical signals; anda filter function programmed to filter the electrical data, the expert system selectively controlling the filter based on the frequency characteristics based on the acceptability of the computed location for the source of electrical activity. 6. The system of claim 1, further comprising: a noise estimation function programmed to estimate relative noise of the of a plurality of channels that provide the electrical signals; anda channel control function programmed to remove the electrical data corresponding to noisy channels based on the relative noise, the expert system selectively controlling the channel control function to rem... |  | A61B005/00 A61B005/04 A61B005/0476
US20160124063 | A1 | 2016-05-05 | Magnetic Relaxometry to Assess Disease via Circulating Markers | 2015-09-22 | Methods and apparatuses for detecting substances circulating in a patient's blood are described. The substances, such as tumor cells, are targeted with nanoparticles labeled such that they preferentially bind with the substance of interest. As the nanoparticle-labeled substance circulates, it can pass near a magnetic relaxometry system that can detect the substance by detecting the nanoparticles bound to it. | 1. A method of determining the presence of a substance of interest in the circulatory system of a human or animal body, comprising: (a) providing probes comprising nanoparticles conjugated with a targeting agent that preferentially binds with the substance of interest in the body;(b) introducing the probes into the body;(c) providing a magnetic relaxometry measurement system configured to detect nanoparticles within a measurement region of the system;(d) positioning the magnetic relaxometry system relative to the body such that a portion of the circulatory system of the body is within the measurement region;(e) using the magnetic relaxometry system to assess the presence of nanoparticles bound to the substance and within the measurement region;(f) using the presence of nanoparticles to determine the presence of the substance in the circulatory system. 2. A method as in claim 1, wherein providing probes comprises providing superparamagnetic nanoparticles whose magnetization decays by Neel relaxation from saturated to about 0.1 of saturated in 5 seconds or less. 3. A method as in claim 2, wherein providing probes comprises providing nanoparticles having a biocompatible coating, and conjugated to an antibody that preferentially binds to one or more types of cells. 4. A method as in claim 3, wherein the one or more types of cells comprises one or more types of cancer cells. 5. A method as in claim 3, wherein the one or more types of cells comprises one or more types of cells correlated with the presence, growth or metastasis of cancer. 6. An apparatus for determining the presence of a substance of interest in the circulatory system of a human or animal body, comprising: (a) A magnetization system, configured to impose a magnetic field upon a magnetization region comprising a portion of the circulatory system of a human or animal body;(b) A magnetic measurement system, comprising one or more magnetic sensors configured to measure a magnetic field at a plurality of times in a measurement region comprising a portion of the circulatory system of a human or animal body;(c) An analysis system, configured to determine the presence of superparamagnetic nanoparticles bound to one or more substances of interest from the magnetic field measurements and for the Neel relaxation properties of the nanoparticles. 7. An apparatus as in claim 6, wherein the magnetization system comprises one or more electrically conductive coils, and a power supply, wherein the power supply provides electrical current through the coils. 8. An apparatus as in claim 6, wherein the measurement system comprises one or more superconducting quantum interference devices. 9. An apparatus as in claim 6, wherein the measurement system comprises one or more atomic magnetometers. 10. An apparatus as in claim 6, wherein the magnetization system is configured to mount with an arm of a human body. 11. An apparatus as claim 6, wherein the measurement system is configured to mount with an ar... |  | G01R033/44 G01R033/56 A61B005/00 G01R033/34 A61B005/145 A61B005/055
US10219699 | B2 | 2019-03-05 | Implantable sensors and methods of use | 2014-11-10 | Systems, methods, and apparatus are described herein for obtaining medical diagnostic measurements from implanted sensors. In various embodiments, a spinal implant-monitoring apparatus may include: a bridge defining a first hermetically-sealed interior and two or more legs; one or more strain gauges contained within the first hermetically-sealed interior of the bridge to provide a signal indicative of strain measured between the legs of the bridge; a housing defining a second hermetically-sealed interior, the housing mounted on a surface of the bridge; and control circuitry contained within the second hermetically-sealed interior. The control circuitry may be in communication with the one or more strain gauges and may be configured to convert the signal into digital data representative of the signal. Methods of using such apparatus are also disclosed. | 1. A spinal implant-monitoring apparatus, comprising: a bridge defining a first hermetically-sealed interior and including a first leg and a second leg that extend parallel to one another and laterally from a span of the bridge;one or more strain gauges contained within the first hermetically-sealed interior to provide a signal indicative of strain measured at the bridge;a first clamp and a second clamp, the clamps being configured to be rigidly secured to a rod implanted in a patient's body, wherein each of the first and second clamp defines a respective recess into which a respective first or second leg of the bridge is inserted and secured, and wherein one or both of the clamps define a platform adjacent the respective recess;a housing defining a second hermetically-sealed interior, the housing mounted on the platform defined by one or both of the clamps, adjacent the bridge; andcontrol circuitry contained within the second hermetically-sealed interior, the control circuitry in communication with the one or more strain gauges and configured to convert the signal into digital data representative of the signal. 2. The spinal implant-monitoring apparatus of claim 1, further comprising an antenna contained within the second hermetically-sealed interior and spaced from the control circuitry. 3. The spinal implant-monitoring apparatus of claim 2, wherein the antenna comprises a coil wrapped around a portion of a frame that spaces the coil from the control circuitry. 4. The spinal implant-monitoring apparatus of claim 2, wherein the control circuitry is configured to wirelessly transmit, to a reader via the antenna, the digital data representative of the signal. 5. The spinal implant-monitoring apparatus of claim 4, wherein the control circuitry is further configured to wirelessly transmit the digital data during a portion of a transmission time interval that is stochastically selected from a plurality of portions of the transmission time interval. 6. The spinal implant-monitoring apparatus of claim 5, wherein the transmission time interval is one of a plurality of transmission time intervals, each divided into a plurality of sub-transmission time intervals, and wherein the control circuitry is further configured to transmit the digital data during a stochastically selected sub-transmission time interval of each of the plurality of transmission time intervals. 7. The spinal implant-monitoring apparatus of claim 5, wherein the portion of the transmission time interval is stochastically selected based at least in part on a unique identifier assigned to the spinal implant-monitoring apparatus. 8. The spinal implant-monitoring apparatus of claim 1, wherein the housing comprises a generally planar lid mounted on a ceramic cap, wherein a portion of the ceramic cap adjacent the generally-planar lid is encompassed by a metal ring, wherein the metal ring is welded to a metal portion of the generally-planar lid. 9. The spinal implant-monitoring apparat... | Intellirod Spine Inc. | A61B005/00 A61B090/00 A61B017/70 A61F002/46 H04W084/18
US09901676 | B2 | 2018-02-27 | Bidirectional valve with improved threshold pressure accuracy | 2014-05-06 | A bidirectional valve for medical use or the like provides different threshold opening pressures in different directions using a valve seat and flapper construction providing improved characterization in operation in contrast to cross-slit valves often used for bidirectional operation. A valve disk supported in a central annular region provides movable portions engaging in valve seats at a peripheral region and a central region. | 1. A valve for bidirectional flow comprising: a housing providing a first and second port communicating with an internal housing volume;an elastomeric disk positioned between the first and second port in the internal housing volume, the elastomeric disk retained against the housing so that the housing holds the elastomeric disk at a protruding annular location on the elastomeric disk between a central region and an outer peripheral region of the elastomeric disk so that the central region and outer peripheral region of the elastomeric disk are free to move with respect to the housing with a flow fluid;the housing providing a first valve seat engaging the outer peripheral region of the elastomeric disk when the elastomeric disk is in a relaxed state and having a second valve seat engaging the central region when the elastomeric disk is in a relaxed state;the housing further providing first flow channels leading from the first port to the elastomeric disk to press the elastomeric disk against the second valve seat and to press the elastomeric disk away from the first valve seat when pressure at the first port is greater than pressure at the second port; andthe housing further providing second flow channels leading from the second port to the elastomeric disk to press the elastomeric disk against the first valve seat and to lift the elastomeric disk away from the second valve seat when pressure at the second port is greater than pressure at the first port. 2. The valve of claim 1 wherein the elastomeric disk interacts with the first and second valve seats to provide for a first threshold opening pressure at which fluid flows between the first port and the second port and a second threshold opening pressure different from the first threshold opening pressure when the fluid flows between the second port and the first port. 3. The valve of claim 2 wherein the first threshold opening pressure permits gravity flow of intravenous liquids from a standard IV bag. 4. The valve of claim 3 wherein the second opening pressure permits sampling of patient blood from a patient using a syringe or vacuum pump device. 5. The valve of claim 3 wherein the first threshold opening pressure is less than twenty millibars and the second threshold opening pressure is greater than two bars. 6. The valve of claim 2 wherein a difference between the first and second threshold opening pressures is a factor greater than one hundred. 7. The valve of claim 1 wherein the elastomeric disk is comprised of a first and second elastic material of different elasticity. 8. The valve of claim 1 wherein the elastomeric disk is comprised of a first and second elastic material of different Shore hardnesses. 9. The valve of claim 8 wherein the first elastic material of greater elasticity forms the outer peripheral region of the elastomeric disk contacting the first valve seat and the second elastic material of lesser elasticity forms the central region contacting the second valve seat. ... | ILLINOIS TOOL WORKS INC. | A61M039/24 A61M005/168 A61B005/15 A61M039/22
US09737236 | B2 | 2017-08-22 | Gas concentration apparatus having carbon foam | 2015-11-03 | A gas concentration apparatus and a method of operating the gas concentration apparatus are provided. The gas concentration apparatus includes a gas concentration module. The gas concentration module includes a base component having a chamber configured to accommodate carbon foam, a gas inlet connected to one side of the chamber, a gas outlet connected to the other side of the chamber, and a heating device disposed on at least one side of the substrate. | 1. A gas concentration apparatus, comprising: a gas concentration module comprising: a base component including a chamber that accommodates carbon foam having linear carbon materials,a gas inlet connected to the chamber, anda gas outlet connected to the chamber;a heating device disposed on at least one side of the base component;a source gas injection line connected to the gas inlet of the gas concentration module; anda droplet generation module connected to the source gas injection line. 2. The gas concentration apparatus of claim 1, wherein the linear carbon materials are irregularly arranged, and the carbon foam further comprises a calcinated carbon layer coating outer surfaces of the linear carbon materials and a plurality of pores disposed between the linear carbon materials coated with the calcinated carbon layer. 3. The gas concentration apparatus of claim 1, wherein the heating device comprises a heater pattern. 4. The gas concentration apparatus of claim 1, wherein the heating device comprises a peltier device. 5. The gas concentration apparatus of claim 1, wherein the base component is a plate-type substrate, the chamber is formed in a trench shape in an upper surface of the substrate, and the gas inlet and the gas outlet are formed in a top plate covering the substrate. 6. The gas concentration apparatus of claim 1, wherein the source gas injection line is configured to supply a source gas to the gas inlet, wherein the source gas comprises a volatile organic compound. 7. The gas concentration apparatus of claim 1, wherein the droplet generation module comprises: a liquid storage unit configured to store a liquid;a droplet generator configured to generate droplets from the liquid storage unit; anda droplet discharge line connected to the source gas injection line. 8. The gas concentration apparatus of claim 7, wherein the liquid comprises water. 9. The gas concentration apparatus of claim 7, wherein the droplet generator is a bubble-generating gas injection line or an ultrasonic vibrator immersed in the liquid. 10. A method of operating a gas concentration apparatus, comprising: providing a gas concentration module including a base component having a chamber accommodating carbon foam containing linear carbon materials, a gas inlet connected to the chamber, a gas outlet connected to the chamber, and a heat-transfer device disposed on at least one side of the base component;supplying a source gas including a target gas through the gas inlet to concentrate the target gas in the carbon foam;increasing a relative humidity in the chamber at least one of before and while supplying the source gas; andheating the chamber with the heat-transfer device, which causes discharging the target gas concentrated in the carbon foam. 11. The method of claim 10, wherein the linear carbon materials are irregularly arranged, and the carbon foam comprises a calcinated carbon layer coating outer surfaces of the linear carbon materials and a plurality ... | Kookmin University Industry Academy Cooperation Fo | A61B005/083 B01D053/02 A61B005/097 B01J020/20 B01J020/28 A61B005/08 B01D053/04
US09326725 | B2 | 2016-05-03 | Apparatus and method for human algometry | 2011-03-30 | An apparatus and method for performing human algometry are disclosed. They include a stimulator configured to apply electrical stimulation of variable intensity to an area of a patient's body, a monitoring device configured to measure a level of cortical activity in one or more regions of the patient's brain, and a microprocessor connected to the stimulator and the monitoring device that is configured to correlate the intensity of the electrical stimulation with the level of activity in the one or more regions of the patient's brain and to determine at least one of a measurement of pain intensity, a measurement of a sensory detection threshold (SDT), a measurement of a drug's analgesic impact, an indication of an onset of tolerance to a drug, an indication of an onset of analgesic-induced hyperalgesia, an indication of conditions of allodynia, a measurement of dose-response characteristics of pain management drugs, and a characterization of a pain condition. | 1. A system for objectively measuring pain comprising: a stimulator configured to apply electrical stimulation of variable intensity to an area of a patient's body;a monitoring device configured to measure a level of cortical activity at a plurality of different regions of the patient's brain including a primary somatosensory cortex and a dorsolateral prefrontal cortex; anda microprocessor connected to the stimulator and the monitoring device that is configured to correlate a first level of cortical activity at the primary somatosensory cortex of the plurality of different regions of the patient's brain with a second level of cortical activity at the dorsolateral prefrontal cortex of the plurality of different regions of the patient's brain to determine the relationship between an emotional component of pain and a nociceptive component of pain and remove the emotional component to derive an activity measurement value, andcorrelate the variable intensity of the electrical stimulation with the activity measurement value to determine, based on the correlation, an objective pain representation,wherein the system further includes:a pharmacologic-dispensing device that administers pain intervention drugs to the patient in accordance with the determined objective pain representation. 2. The system of claim 1, wherein the stimulator is configured to incrementally increase the intensity of the electrical stimulation until a sensory detection threshold (SDT) level of cortical activity is measured by the monitoring device. 3. The system of claim 1, wherein the stimulator is further configured to apply electrical stimulation to sensory nerve fibers by applying electrical stimulation at a first frequency in order to generate an action potential in Aβ fibers, a second frequency in order to generate an action potential in Aδ fibers, and a third frequency in order to generate an action potential in C fibers. 4. The system of claim 1, wherein the monitoring device utilizes at least one of near infrared spectroscopy (NIRS), electroencephalography (EEG), functional Magnetic Resonance imaging (fMRI), and near infrared imaging (NIRI) to measure the level of cortical activity in the plurality of different regions of the patient's brain. 5. The system of claim 4, wherein the monitoring device utilizes NIRS and a NIRS sensor with emitters and detectors to measure the level of cortical activity in plurality of different regions of the patient's brain. 6. The system of claim 4, wherein the monitoring device utilizes EEG with electrodes as sensors or EEG and NIRS with a combination of electrodes and at least one NIRS sensor as the sensors to measure the level of cortical activity in the plurality of different regions of the patient's brain. 7. The system of claim 1, wherein the microprocessor is further configured to accumulate a plurality of measurements of different levels of cortical activity in one or more regions of a plurality of first patients' brains in respo... | Children's National Medical Center | A61B005/04 A61B005/00 A61B005/0484 A61N001/36
US09326789 | B2 | 2016-05-03 | Flexible debulking catheters with imaging and methods of use and manufacture | 2014-06-13 | Catheters and methods for removing material from (or “debulking”) a body lumen while imaging the region comprised of the material are provided. The catheters can be used in body lumens, including but not limited to intravascular lumens such as coronary or peripheral arteries. The catheters include housings or other structures to mount or protect an imaging transducer. Generally, debulking catheters include a proximal portion, a distal portion having an opening, a cutting element which may be exposed through the opening to contact material in a body lumen and an imaging transducer with associated circuitry and display. The catheter debulks a body lumen when it is moved while the cutting element is in contact with the material in the lumen and the region comprised of the material is imaged before, after, or during catheter movement. | 1. A catheter for deployment in a vessel comprising: a control handle;an elongate proximal segment having proximal and distal ends, the proximal end being connected to the control handle;a distal tip assembly having proximal and distal ends;a tissue-removing element in the distal tip assembly;a housing having a proximal end portion comprising a first material, a distal end portion comprising a second material, and a body portion extending between the proximal and distal end portions, the body portion comprising a third material, the first and second materials being different from the third material, the proximal end portion being connected to the distal end of the elongate proximal segment, the distal end portion being connected to the proximal end of the distal tip assembly, the housing having a luminal wall with an inner surface defining a lumen, the luminal wall having at least one slot; anda transducer mounted at least partially within the lumen of the housing. 2. The catheter of claim 1 wherein an outer surface of the housing has at least one longitudinally extending rib on the outer surface of the luminal wall. 3. The catheter of claim 2 wherein the at least one longitudinally extending rib on the outer surface of the luminal wall is in the range of 5 to 8 ribs. 4. The catheter of claim 1 wherein the at least one slot in the luminal wall is two slots. 5. The catheter of claim 1 wherein the transducer is at a position spaced from the luminal wall to form a gap between the transducer and the luminal wall. 6. The catheter of claim 1 wherein the transducer is configured to have 360° image capability of the vessel through the luminal wall of the housing and through the at least one slot of the housing, a radial angle of the circumference of the housing imaged by the transducer through the at least one slot in the luminal wall of the housing is in the range of 60° to 180°. 7. The catheter of claim 6 wherein the radial angle is in the range of 120° to 150°. 8. The catheter of claim 7 wherein the radial angle is 130°. 9. The catheter of claim 4 wherein the transducer is configured to have 360° image capability of the vessel through the luminal wall of the housing and through the at least one slot of the housing, radial angle of the circumference of the housing imaged by the transducer through each slot in the luminal wall of the housing is in the range of 60° to 180°. 10. The catheter of claim 9 wherein the radial angle is in the range of 65° to 90°. 11. The catheter of claim 10 wherein the radial angle is 73°. 12. The catheter of claim 1 wherein the first and second materials comprise metal and the third material comprises polymer. 13. The catheter of claim 12 wherein the first material comprises stainless steel, the second material comprises tungsten and the third material comprises acrylonitrile butadiene styrene (ABS). 14. The catheter of claim 12 wherein the distal end of the elongate proximal segment connected to the proximal end... | Covidien Lp | A61B017/32 A61B017/3207 A61B005/00 A61B005/055 A61B008/08 A61B008/12 A61B008/00 A61B010/04 G01R033/28 A61B019/00 A61M025/00 A61M025/01 A61M025/06 G10K011/00 A61B017/00
US09332933 | B2 | 2016-05-10 | Method and apparatus for providing analyte sensor insertion | 2014-09-29 | Devices and methods for positioning a portion of a sensor at a first predetermined location, displacing the portion of the sensor from the first predetermined location to a second predetermined location, and detecting one or signals associated with an analyte level of a patient at the second predetermined location are disclosed. Also provided are systems and kits for use in analyte monitoring. | 1. A method, comprising: transcutaneously positioning an analyte sensor at a first predetermined location, the analyte sensor configured for fluid contact with interstitial fluid under a skin layer; andcoupling a data processing unit to a mounting unit such that the analyte sensor is displaced to a second predetermined location when the data processing unit is coupled to the mounting unit. 2. The method of claim 1, wherein transcutaneously positioning the sensor at the first predetermined location includes deploying the sensor at a first predetermined speed. 3. The method of claim 2, wherein displacing the sensor includes displacing the sensor to the second predetermined location at a second predetermined speed. 4. The method of claim 3, wherein the first predetermined speed is at a higher velocity than the second predetermined speed. 5. The method of claim 1, wherein displacing the sensor to the second predetermined location includes manually displacing the sensor. 6. The method of claim 1, further including retaining the sensor at the second predetermined location during a useful life of the sensor. 7. The method of claim 1, further including wirelessly transmitting signals generated by the analyte sensor to a remote location. 8. The method of claim 1, further including calibrating signals generated by the analyte sensor. 9. The method of claim 1, further including displaying data associated with a monitored analyte level based on signals generated by the analyte sensor. 10. The method of claim 1, further including processing signals generated by the analyte sensor with the data processing unit. 11. The method of claim 1, wherein the analyte sensor is transcutaneously positioned at the first predetermined location at a first angle relative to the skin layer, and further, wherein the analyte sensor is displaced to the second predetermined location at a second angle relative to the skin layer. 12. The method of claim 1, wherein transcutaneously positioning the analyte sensor at the first predetermined location includes coupling an introducer to the analyte sensor for transcutaneous positioning. 13. The method of claim 12, further including removing the introducer from the analyte sensor after transcutaneously positioning the analyte sensor at the first predetermined location. 14. An apparatus, comprising: a mounting unit for placement on a skin layer;an analyte sensor;an introducer coupled to the analyte sensor and configured to transcutaneously position the analyte sensor at a first predetermined location in fluid contact with interstitial fluid under the skin layer; anda data processing unit configured to couple to the mounting unit such that the analyte sensor is displaced to a second predetermined location when the data processing unit is coupled to the mounting unit. 15. The apparatus of claim 14, wherein the sensor is manually displaced to the second predetermined location. 16. The apparatus of claim 14, wherein the data proc... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/145 A61B005/1486
US09339233 | B2 | 2016-05-17 | Medical implantable lead with fixation detection | 2014-10-20 | A lead has an electrically controlled switch at a proximal end of a rotatable shaft. The electrically controlled switch has a first configuration to electrically connect a first electric conductor to a second electric conductor for electrically activating the pin and has a second configuration to electrically connect the first electric conductor to the rotatable shaft for electrically activating the helix. Proper fixation of the helix to an organ is determined by switching the electrically controlled switch to the first configuration to render the pin electrically active and to electrically deactivate the helix. Upon determination of proper fixation of the helix to the organ, the electrically controlled switch is switched to the second configuration to render the helix electrically active and to electrically deactivate the pin. | 1. A medical implantable lead adapted to be implanted into a human or animal body to monitor and/or control an organ inside the body, the lead comprising: an intermediate flexible lead part;a connector at a proximal portion of the intermediate flexible lead part;a tubular header at a distal portion of the intermediate flexible lead part;a rotatable shaft disposed within the tubular header, wherein the rotatable shaft is extendably and retractably arranged relative to the tubular header, and wherein the rotatable shaft has a passageway longitudinally extending from a proximal end to a distal end of the rotatable shaft;a first electric conductor electrically coupled to the connector;a helix at the distal end of the rotatable shaft, wherein the helix is electrically coupled to the rotatable shaft, and wherein the helix is configured for fixation to the organ upon rotation and extension of the rotatable shaft;a pin disposed partially within the central passageway at the distal end of the rotatable shaft, the pin being electrically insulated from the rotatable shaft;a second electric conductor disposed within the passageway of the rotatable shaft, the second electric conductor being electrically insulated from the rotatable shaft, and the second electric conductor coupled to the pin; anda mechanically controlled switch at the proximal end of the rotatable shaft, the mechanically controlled switch having a first configuration to electrically connect the first electric conductor to the second electric conductor for electrically activating the pin and having a second configuration to electrically connect the first electric conductor to the rotatable shaft for electrically activating the helix;wherein proper fixation of the helix to the organ is determined by switching the mechanically controlled switch with a stylet to the first configuration to render the pin electrically active and to electrically deactivate the helix; andwherein upon determination of proper fixation of the helix to the organ, the mechanically controlled switch is switched to the second configuration by removing the stylet to render the helix electrically active and to electrically deactivate the pin. 2. The medical implantable lead of claim 1, wherein a proximal end of the second electric conductor is coupled to the mechanically controlled switch and the distal end of the second electric conductor is coupled to the pin. 3. The medical implantable lead of claim 1, wherein the pin is coaxially arranged in relation to the helix and projects distally from a distal end of the rotatable shaft. 4. The medical implantable lead of claim 1, wherein the pin is separate from the helix. 5. The medical implantable lead of claim 1, wherein the pin is adapted to penetrate into tissue of the organ. 6. The medical implantable lead of claim 1, wherein a longitudinal length of the pin is less or equal to half a longitudinal length of the helix. 7. The medical implantable lead of claim 1, wherein th... | ST. JUDE MEDICAL AB | A61N001/00 A61B005/04 A61B005/00 A61N001/05
US09339193 | B2 | 2016-05-17 | Physiological adaptability system with multiple sensors | 2012-05-21 | According to an aspect of an embodiment, a method of measuring physiological adaptability of a subject is described. The method may include receiving, from a first sensor, a first data signal indicating a first biological function of the subject. The method may also include receiving, from a second sensor, a second data signal indicating a second biological function of the subject. The method may also include calculating an adaptability of the subject based on the first data signal and the second data signal. The method may also include reporting the calculated adaptability. | 1. A method of measuring physiological adaptability of a subject, the method comprising: receiving, from a first sensor, a first data signal indicating a first biological function of the subject;receiving, from a second sensor, a second data signal indicating a second biological function of the subject;calculating a stress adaptability of the subject based on the first data signal and the second data signal, comprising: determining, based on the first data signal from the first sensor and the second data signal from the second sensor, a single first average heart rate corresponding to time intervals when the subject is generally inhaling;determining, based on the first data signal from the first sensor and the second data signal from the second sensor, a single second average heart rate corresponding to time intervals when the subject is generally exhaling; andcalculating the stress adaptability based on a difference between the single first average heart rate and the single second average heart rate; andreporting the calculated stress adaptability,wherein the first biological function includes a respiratory function, the second biological function includes a cardiac function;wherein the method further comprises detecting a sequence of heart beats in the second data signal, each sequential pair of detected heart beats in the detected sequence of heart beats defining a start time and an end time of an associated time interval; andwherein the calculating the stress adaptability of the subject based on the first data signal and the second data signal further comprises determining time intervals of the sequential pairs of detected heart beats when the subject is generally inhaling and time intervals of the sequential pairs of detected heart beats when the subject is generally exhaling. 2. The method of claim 1, wherein: receiving the first data signal indicating the first biological function of the subject comprises receiving the first data signal indicating the first biological function of the subject that varies over time; andreceiving the second data signal indicating the second biological function of the subject comprises receiving the second data signal indicating the second biological function of the subject that varies over time in dependence on a variability of the first biological function. 3. The method of claim 1, wherein receiving the first and second data signals comprises receiving the first and second data signals while the subject is moving in an uncontrolled environment. 4. The method of claim 1, wherein determining time intervals of the sequential pairs of detected heart beats when the subject is generally inhaling and time intervals of the sequential pairs of detected heart beats when the subject is generally exhaling includes, for each time interval: determining a difference between a first instantaneous measurement of the respiratory function and a second instantaneous measurement of the respiratory function, the first instant... | FUJITSU LIMITED | A61B005/00 G06F019/00 A61B005/0205 A61B005/0456 A61B005/024 A61B005/08
US09345533 | B2 | 2016-05-24 | Catheter with bendable tip | 2012-05-14 | A medical probe includes a flexible insertion tube, having a distal end for insertion into a body cavity of a patient, and a distal tip, which is disposed at the distal end of the insertion tube and is configured to be brought into contact with tissue in the body cavity. A coupling member couples the distal tip to the distal end of the insertion tube and includes a tubular piece of an elastic material having a helical cut therethrough along a portion of a length of the piece. | 1. A method for performing a medical procedure, comprising: inserting into a body cavity of a patient a probe, which comprises a flexible insertion tube and a distal tip, which is disposed at a distal end of the insertion tube, a coupling member that couples the distal tip to the distal end of the insertion tube and comprises a tubular piece of an elastic material having a helical cut therethrough along a portion of a length of the piece, a position sensor within the distal tip, the position sensor being configured to generate a signal in response to a magnetic field, wherein the signal is indicative of a position of the distal tip, and a magnetic field generator within the distal end of the insertion tube for generating the magnetic field;generating the magnetic field;generating the signal in response to the magnetic field;analyzing the signal to determine if the distal tip is pressing against the tissue in the body cavity firmly to give the desired therapeutic or diagnostic result without causing undesired tissue damage; andif the distal tip is not pressing against the tissue in the body cavity firmly to give the desired therapeutic or diagnostic result without causing undesired tissue damage, repositioning the distal tip to bring the distal tip into contact with tissue in the body cavity. 2. The method according to claim 1, wherein inserting the probe comprises passing the probe through a blood vessel into a heart of the patient. 3. The method according to claim 1, further comprising ablating the tissue with which the distal tip is in contact. 4. The method according to claim 1 wherein the coupling member comprises a tubular piece of a superelastic alloy. | Biosense Webster, Inc. | A61B018/14 A61B005/05 A61B019/00 A61M025/00
US09326748 | B2 | 2016-05-03 | Clinical apparatuses | 2012-06-25 | A clinical apparatus includes a storage to store strain data, strain rate data, or acceleration data measured in a ROI (region of interest); and a data processor to calculate a stress, a stress tensor, a stress tensor component, an inertia, a mean normal stress, a pressure, a mechanical source, an elastic constant, a visco elastic constant, a viscosity, or a density of an arbitrary point within the ROI. The data processor calculates the stress, the stress tensor, the stress tensor component, the inertia, the mean normal stress, the pressure, or the mechanical source based on an equation representing a relation between (i) the stress, the stress tensor, the stress tensor component, the inertia, the mean normal stress, the pressure, or the mechanical source, (ii) the elastic constant, the visco elastic constant, the viscosity, or the density, and (iii) the measured strain data, the strain rate data, or the acceleration data. | 1. A clinical apparatus comprising: at least one sensor which is configured to sense one selected from a group consisting of an ultrasound signal, a magnetic nuclear resonance signal, an electromagnetic wave, and light;a storage which is configured to store at least one of strain tensor component data, strain rate tensor component data, and acceleration vector component data measured in a ROI (region of interest) set in a target by using said at least one sensor; anda data processor which is configured to calculate at least one of a stress and a pressure of at least one arbitrary point within the ROI by utilizing first reference data representing at least one reference value about at least one of the stress and the pressure and at least one reference position that has said at least one reference value, the pressure being a mean of normal stress tensor components, and each of the stress and the pressure being expressed using at least one product of at least one elastic constant and a volume strain, at least one product of at least one visco elastic constant and a volume strain rate, or at least one product of at least one viscosity and a volume strain rate;wherein said data processor is configured to linearly and directly calculate said at least one of the stress and the pressure while setting to unknown a Poisson's ratio or at least one of said at least one elastic constant, said at least one visco elastic constant, and said at least one viscosity, which expresses said at least one of the stress and the pressure, based on at least said first reference data and at least one equation representing a relation between (i) said at least one of the stress and the pressure, (ii) at least one of other stresses that is at least one stress tensor component expressed using at least one product of said at least one elastic constant or other elastic constants and the measured strain tensor component data, at least one product of said at least one visco elastic constant or other visco elastic constants and the measured strain rate tensor component data, or at least one product of said at least one viscosity or other viscosities and the measured strain rate tensor component data, (iii) when the arbitrary point within the ROI has a non-zero acceleration vector, at least one inertia vector component expressed using a product of a density and the measured acceleration vector component data, and (iv) when the arbitrary point within the ROI has a mechanical source, the mechanical source. 2. The clinical apparatus according to claim 1, wherein said data processor is configured to simultaneously calculate said at least one elastic constant or said other elastic constants, said at least one visco elastic constant or said other visco elastic constants, said at least one viscosity or said other viscosities, said density, said at least one stress tensor component, said at least one inertia vector component, or said mechanical source in addition to said at least one of the... |  | A61B008/08 G01S007/52 G01S015/89 A61B005/00 A61B005/01 G01S015/58
US09326704 | B2 | 2016-05-03 | Method and apparatus for determining sportsman jumps using fuzzy logic | 2012-09-12 | A method and apparatus for detecting a jump of a moving sportsman and separating the jump from other sportsman motion is described. Accelerometer data generated by a sportsman's motion is received in a computing device. The computing device is used to apply fuzzy logic membership functions to a plurality of parameters associated with the accelerometer data to detect a pattern associated with jumps. A subset of the data is identified as representing a jump based upon the detection of the pattern. The determination is used in the computing device to transform data, whereby jumps are separated from other sportsman motion. | 1. A method comprising: receiving, by a computing device, accelerometer data generated by a sportsman's motion from an accelerometer coupled to the sportsman and in communication with the computing device;applying, by the computing device, a plurality of fuzzy logic membership functions to a plurality of parameters associated with the accelerometer data to detect a pattern associated with a jump, wherein detection of the pattern associated with the jump includes: determining a fuzzy logic initial jump probability membership function;determining a fuzzy logic jump start probability membership function;determining a fuzzy logic landing shock probability membership function;determining a fuzzy logic false jump probability membership function; andcombining the jump start probability and landing shock probability membership functions together using fuzzy logic rules;identifying a subset of the accelerometer data representing the jump based upon the detected pattern; andtransforming the accelerometer data, by the computing device, by separating the identified subset of accelerometer data representing the jump from accelerometer data associated with other motion by the sportsman. 2. The method of claim 1, further comprising: receiving, by the computing device, global positioning sensor (GPS) measurements from a GPS device coupled to the sportsman and in communication with the computing device; andidentifying the subset of the accelerometer data representing the jump based at least in part on the received GPS measurements. 3. The method of claim 2, wherein the fuzzy logic membership functions are applied to the GPS measurements to detect the pattern associated with the jump. 4. The method of claim 1, wherein a time duration of the jump of the sportsman is further determined. 5. The method of claim 3, wherein acceptance and rejection are mapped as 2D regions on a membership function vs. a jumptime plane. 6. The method of claim 4, wherein a total norm of a 3x axis accelerometer is used. 7. The method of claim 4, wherein a projection of an accelerometer vector to the magnetic vector is used. 8. The method of claim 4, wherein accelerometer data is collected only in response to the computing device measuring a velocity of the sportsman above some predetermined threshold. 9. The method of claim 6, wherein the accelerometer data is corrected by an approximate rotational acceleration computed from gyro sensors coupled to the sportsman and in communication with the computing device. 10. The method of claim 1, wherein the accelerometer data is time tagged. 11. The method of claim 1, wherein the computing device comprises a smartphone with an accelerometer therein. 12. A non-transitory computer readable storage medium storing instructions that, when executed by a computing device, cause the computing device to: receive accelerometer data generated by a sportsman's motion from an accelerometer coupled to the sportsman and in communication with the computi... | AlpineReplay, Inc. | A61B005/11 A63B024/00 G06F019/00 A61B005/00 A43B003/00
US09333369 | B2 | 2016-05-10 | Systems and methods to reduce power deposition in tissue exposed to radio frequency electromagnetic fields | 2010-10-22 | An apparatus for exposing a region of interest of an object, animal or person to an alternating magnetic field has a source of radio-frequency electromagnetic radiation arranged to provide the alternating magnetic field in an exposure volume defined by the apparatus, and a shield arranged between the source of radio-frequency electromagnetic radiation and the exposure volume. The shield includes a material that has a sufficient thickness and arrangement to reduce power deposition to at least regions outside of the region of interest of the object, animal or person during exposure in the exposure volume. | 1. A method of exposing a region of interest of an object, animal or person to an alternating magnetic field, comprising: generating electromagnetic radiation having a predetermined frequency and amplitude;arranging said object, animal or person to be exposed to at least an alternating magnetic field component of said electromagnetic radiation;disposing the object, animal or person within an exposure volume defined by a rigid wall of a shield, such that the shield is disposed between the object, animal, or person and a source of the electromagnetic radiation, and such that the shield reduces power deposition to regions outside of said region of interest of said object, animal or person when being exposed to said alternating magnetic field; wherein said shield comprises a material having a sufficient thickness and arrangement to reduce power deposition to at least regions outside of said region of interest of said object, animal or person during exposure in said exposure volume, wherein said material of said shield comprises water, andwherein said source of radio-frequency electromagnetic radiation comprises an electrically conducting coil that winds around at least a portion of said exposure volume defined by said system to provide a solenoid having an inner diameter suitable to accommodate said object, animal or person, and wherein said water of said shield has a thickness of at least about 2% of the inner diameter of said coil and less than about 20% of said coil. 2. A method of exposing an object to an alternating magnetic field according to claim 1, wherein the shield comprises at least one of a liquid metal or metal particles suspended in the water. 3. A method of exposing an object to an alternating magnetic field according to claim 1, further comprising disposing magnetic particles in said object such that at least localized regions of said object are heated through interactions of said alternating magnetic field with said magnetic particles. 4. A method of exposing an object to an alternating magnetic field according to claim 1, wherein said generating electromagnetic radiation generates electromagnetic radiation with a frequency of at least 100 kHz. 5. A system for heating localized regions within a body, comprising: a source of radio-frequency electromagnetic radiation arranged to provide an alternating magnetic field in an exposure volume defined by said system;a shield having a rigid wall having an inner and an outer surface, wherein the inner surface of the rigid wall defines the exposure volume, and wherein the shield is arranged between said source of radio-frequency electromagnetic radiation and said exposure volume; anda plurality of magnetic particles that are suitable to be disposed in said body,wherein said shield comprises a material having a sufficient thickness and arrangement to attenuate electric field components of said electromagnetic radiation produced by said source by a predetermined amount prior to reaching said... | The Johns Hopkins University | A61N002/02 A61B005/055 A61B005/00 A61N001/40 G01R033/422
US09326914 | B2 | 2016-05-03 | Device for extracting, storing and/or processing blood or other substances of human or animal origin, and for applying blood compounds or other biological compounds | 2013-07-03 | Device (1a, 1b, 1c, 1d; 40) for extracting, storing and/or processing blood or other substances of human or animal origin, and for applying blood compounds or other biological compounds, which comprises a body (2; 41) inside which a piston (3; 42) is capable of moving longitudinally and is separable at least in part, where the body (2; 41) is provided with an internal space (9; 51) that may be connected to the outside by means of a conduit (6; 50) at a first end (4; 48) of the body (2; 41), where the body (2; 41) and the piston (3; 43) may be blocked longitudinally in at least one position to allow the creation of different degrees of vacuum. The device has multiple uses and is extremely versatile. | 1. A device for extracting, storing and/or processing blood or other substances of human or animal origin, and for applying blood compounds or other biological compounds, comprising: a hollow body;a piston arranged longitudinally movably inside said body;the body comprising a first end at which a conduit is provided connecting an internal space of the body located between the piston and the conduit with the outside of the body, and a second end comprising an opening through which the piston projects outward from the body, said second end being hermetically separated from the internal space by means of the piston, wherein:part of the piston is configured to be reversibly or irreversibly separated and removed from the rest of the device through said opening of said second end so that a remaining portion of the piston does not project out of the body, whereinthe piston comprises at least one blocking surface,the body comprises at least one blocking surface configured to come into contact with a blocking surface of the piston and thereby offer resistance to the piston advancing towards the first end while allowing the piston to freely move back towards the second end, wherethe device further comprises a perforable sealing member arranged at or proximate the conduit, providing an airtight closing of the conduit,wherein the piston is configured to move towards the second end of the body, such that a vacuum is formed within the internal space by the hermetical separation of the internal space and the second end by the piston, and by the airtight closing of the conduit by the perforable sealing member. 2. The device according to claim 1, wherein the sealing member is internal to the body. 3. The device, according to claim 1, wherein the sealing member is external and connectable to the first end. 4. The device according to claim 1, wherein the piston is capable of moving rotationally inside the body, and wherein at least one blocking surface of the body and at least one blocking surface of the piston are arranged so that blocking occurs only in certain rotational positions of the piston in relation to the body, the piston being free in relation to the body in the remaining rotational positions. 5. The device according to claim 4, wherein at least one blocking surface of the piston is comprised in a tab that projects radially from the piston and wherein at least one corresponding blocking surface of the body is comprised in a transversal seating surface, the body presenting internal walls, the distance of which to the longitudinal central axis decreases in the direction of rotation in which the piston must rotate to unblock from the body so that when the rotation of the piston is caused in said direction the tabs rotate in contact with the internal walls and are flexed until the piston is unblocked in relation to the body. 6. The device, according to claim 1, further comprising a handle on the second end of the body to enable the action of causing th... | BIOTECHNOLOGY INSTITUTE, I MAS D, S.L. | B01L099/00 A61J001/14 A61B005/15 A61B005/153 A61B005/154 A61M005/315
US09351644 | B2 | 2016-05-31 | Medical systems and methods | 2015-05-05 | This disclosure generally relates to medical systems and methods. In one aspect of the invention, a method includes determining a fluorescent light intensity at one or more points on each of multiple recorded images, and producing an image based on the determined fluorescent light intensity at the one or more points. | 1. A method, comprising: recording a first plurality of images of a plurality of points of at least one blood vessel, each of the first plurality of images comprising fluorescent light intensity information at each of the plurality of points of the at least one blood vessel;for each of the plurality of points of the at least one blood vessel, analyzing the first plurality of recorded images to determine a fluorescent light intensity at the point over a first period of time; andfor each of the plurality of points of the at least one blood vessel, displaying a first image that represents the fluorescent light intensity at the point over the first period of time. 2. The method of claim 1, wherein the at least one blood vessel comprises multiple blood vessels. 3. The method of claim 2, further comprising, after displaying the first image: recording a second plurality of images of the plurality of points of the multiple blood vessels, each of the second plurality of images comprising fluorescent light intensity information at each of the plurality of points of the multiple blood vessels;for each of the plurality of points of the multiple blood vessels, analyzing the second plurality of recorded images to determine a fluorescent light intensity at the point over a second period of time; andfor each of the plurality of points of the multiple blood vessels, displaying a second image that represents the fluorescent light intensity at the point over the second period of time. 4. The method of claim 3, further comprising simultaneously displaying the first and second images. 5. The method of claim 4, wherein the first and second images are displayed in a side-by-side fashion. 6. The method of claim 5, further comprising repairing an aneurysm at the multiple blood vessels. 7. The method of claim 6, further comprising: before recording the first plurality of images, injecting a fluorescent dye into the multiple blood vessels;after injecting the fluorescent dye into the multiple blood vessels, applying a clip to the aneurysm to clamp off the aneurysm; andafter applying the clip to the aneurysm, capturing the second set of images. 8. The method of claim 3, further comprising repairing an aneurysm at the multiple blood vessels. 9. The method of claim 8, wherein the first image is displayed before repairing the aneurysm at the multiple blood vessels, and the second image is displayed after repairing the aneurysm at the multiple blood vessels. 10. The method of claim 9, further comprising: before recording the first plurality of images, injecting a fluorescent dye into the multiple blood vessels;after injecting the fluorescent dye into the multiple blood vessels, applying a clip to the aneurysm to clamp off the aneurysm; andafter applying the clip to the aneurysm, capturing the second set of images. 11. The method of claim 8, further comprising: before recording the first plurality of images, injecting a fluorescent dye into the multiple blood vessels;af... | Carl Zeiss Meditec AG | G06K009/00 A61B005/00 A61B005/0275 A61B005/02
US10786157 | B2 | 2020-09-29 | Spectrophotometric sensor | 2015-11-20 | A near-infrared spectroscopy (NIRS) sensor assembly for measuring a characteristic of a biological tissue is provided. The NIRS sensor assembly includes a light source, at least one light detector, and a layer disposed within the sensor assembly. The light source is operable to emit light at one or more predetermined wavelengths. The at least one light detector has an active area for detecting light emitted by the light source and passed through the biological tissue. The light detector is operable to produce signals representative of the detected light. The layer disposed within the sensor assembly has at least one deflection element. | 1. A near-infrared spectroscopy (NIRS) sensor assembly for measuring a characteristic of a biological tissue, the NIRS sensor assembly comprising: a light source operable to emit light at one or more predetermined wavelengths;at least one light detector having an active area for detecting light emitted by the light source and passed through the biological tissue, and which detector is operable to produce signals representative of the detected light; anda pad having a top surface, a bottom surface, and at least one side surface extending between the top and bottom surfaces, at least one light source aperture extending from the top surface to the bottom surface, at least one light detector aperture extending from the top surface to the bottom surface, and at least one deflection element configured as a void disposed in the pad, the void having a length extending between a first lengthwise end and an opposite second lengthwise end, and the deflection element being open to at least one of the top or bottom surfaces, a first light blocking portion of the pad disposed between the first lengthwise end of the at least one deflection element and the at least one side surface, and a second light blocking portion of the pad disposed between the second lengthwise end of the at least one deflection element and the at least one side surface. 2. The sensor assembly of claim 1, wherein the at least one deflection element is independent of the at least one light source aperture and the at least one light detector aperture. 3. The sensor assembly of claim 2, wherein the at least one deflection element extends from the bottom surface to the top surface, thereby providing an open passage between the bottom surface and the top surface. 4. The sensor assembly of claim 3, wherein the at least one deflection element is configured as a slot. 5. A near-infrared spectroscopy (NIRS) sensor assembly for measuring a characteristic of a biological tissue, the NIRS sensor assembly comprising: a light source operable to emit light at one or more predetermined wavelengths;at least one light detector having an active area for detecting light emitted by the light source and passed through the biological tissue, and which detector is operable to produce signals representative of the detected light; anda pad having a top surface, a bottom surface, and at least one side surface extending between the top and bottom surfaces, at least one light source aperture extending from the top surface to the bottom surface, at least one light detector aperture extending from the top surface to the bottom surface, and at least one deflection element comprising a plurality of voids linearly aligned with one another, each of the voids extending into the pad from a first surface opening in the top surface, or extending into the pad from a second surface opening in the bottom surface, or extending between the first surface opening in the top surface and the second surface opening in the bottom surf... | Edwards Lifesciences Corporation | A61B005/00 A61B005/1455
US09339217 | B2 | 2016-05-17 | Analyte monitoring system and methods of use | 2012-11-21 | The present disclosure provides methods of processing data provided by a transcutaneous or subcutaneous analyte sensor utilizing different algorithms to strike a balance between signal responsiveness accompanied by signal noise and the introduction of error associated with that noise. The methods utilize the strengths of a lag correction algorithm and a smoothing algorithm to optimize the quality and value of the resulting data (glucose concentrations and the rates of change in glucose concentrations) to a continuous glucose monitoring system. Also provided are systems and kits. | 1. A method for monitoring an analyte using an analyte monitoring system, comprising: monitoring a data stream including a set of contiguous source data points related to analyte concentration;providing one or more sets of maximum lag corrected signals from the set of contiguous source data points, wherein each set of maximum lag corrected signals is generated utilizing correction terms based on time derivative estimates and historical source data points, and wherein parameters for the maximum lag correction minimize the correlation between an expected glucose error and time derivative estimates and minimize the correlation between the expected glucose error and a pre-determined array of historical source data points;providing a first one or more sets of maximum smoothed signals from the set of contiguous source data points, wherein each set of maximum smoothed signals is generated utilizing a first smoothing algorithm;providing a second one or more sets of maximum smoothed signals from the set of contiguous source data points, wherein each of the second set of maximum smoothed signals is generated utilizing a second smoothing algorithm;determining analyte concentration utilizing a weighted combination of the one or more sets of maximum lag corrected signals and the first and second one or more sets of maximum smoothed signals, wherein more weight is placed on the one or more sets of maximum lag corrected signals to determine the analyte concentration;determining a rate of change in the analyte concentration utilizing a weighted combination of the one or more sets of maximum lag corrected signals and the first and second one or more sets of maximum smoothed signals, where more weight is placed on the first and second one or more sets of maximum smoothed signals to determine the rate of change; andcontrolling administration of therapy based on the determined analyte concentration. 2. The method of claim 1, wherein the first smoothing algorithm and the second smoothing algorithm are different. 3. The method of claim 1, further comprising averaging the first one or more sets of maximum smoothed signals and the second one or more sets of maximum smoothed signals. | Abbott Diabetes Care Inc. | A61B005/00 A61B005/145 A61M005/172
US09348971 | B2 | 2016-05-24 | Two way short message service (SMS)-enabled blood glucose meter and related communications systems and methods | 2012-10-26 | A system and method for two way short message service (SMS)-enabled BGM and related communications is presented. The method includes receiving, by a processing device integrated with a blood glucose meter (BGM), a first short message service (SMS) message comprising a question, presenting the question at the BGM, receiving a response to the question, translating the response to the question into a second SMS message, translating blood glucose measurement information into a third SMS message, the blood glucose measurement information received from a test of a user by the BGM, and transmitting the second and third SMS messages to a receiver. | 1. A method comprising: receiving, by a processing device integrated with a blood glucose meter (BGM), a first short message service (SMS) message comprising a question;presenting the question at the BGM;receiving a response to the question;translating the response to the question into a second SMS message;translating blood glucose measurement information and control solution measurement information into a third SMS message, the blood glucose measurement information comprising a first measurement received from a test of a blood sample of a user by the BGM, and the control solution measurement information comprising a second measurement received from a test of a non-blood standard liquid medium solution by the BGM for calibration of blood glucose measurement accuracy of the BGM; andtransmitting the second and third SMS messages to a receiver. 2. The method of claim 1, wherein the third SMS message comprises information about a date of the test, a time of the test, and a measured blood glucose level measured by the test at the BGM. 3. The method of claim 1, wherein the question comprises at least one of a confirmation of personal information about the user, a request to change a frequency of the test of the user by the BGM, a request to perform an operation at the BGM, administration of the BGM, maintenance of the BGM, provisioning function of the BGM, or a request for the BGM to transmit the control solution measurement information. 4. The method of claim 1 further comprising: displaying the question on a visible screen integrated with the BGM. 5. The method of claim 1, wherein the third SMS message further comprises an error message. 6. The method of claim 1 further comprising: translating the first SMS message into a digital audio broadcasting format. 7. A method comprising: receiving, by a processing device, a question from a communication device of a medical provider, the question received via a secure connection between the processing device and the communication device of the medical provider;translating the question into a first short message service (SMS) message;transmitting the first SMS message to a receiving device integrated in a blood glucose meter (BGM);receiving, from a transmitting device integrated in the BGM, a second SMS message comprising a response to the question;receiving, from the transmitting device integrated in the BGM, a third SMS message comprising: blood glucose measurement information comprising a first measurement received from a test of a blood sample of a user by the BGM; andcontrol solution measurement information comprising a second measurement received from a test of a non-blood standard liquid medium solution by the BGM for calibration of blood glucose measurement accuracy of the BGM;translating the second and third SMS messages into a format that is compliant with the secure connection; andtransmitting, via the secure connection, the second and third SMS messages to the communication device of the medi... | ERP SYSTEMS CORP. | H04W004/00 G06F019/00 G08B001/08 A61B005/145 H04W004/14 H04L012/58 H04L029/08
US09342903 | B2 | 2016-05-17 | Method for generating image for PET attenuation correction from MR image and computer program | 2013-03-28 | When an image for PET attenuation correction is generated from an MR image, the MR image captured by MRI is segmented into regions according to pixel values. In a region in which a radiation attenuation coefficient is considered to be uniform, a radiation attenuation correction value is determined by referring to an existing radiation attenuation correction value table. In a region including multiple tissues having different radiation attenuation coefficients, a radiation attenuation correction value is determined by referring to a standard image. In such a manner, an image for PET attenuation correction in which tissues having similar pixel values in the MR image but different attenuation coefficients for radiation can be distinguished and that can accommodate individual differences and an affected area such as a space occupying lesion (for example, a cancer, abscess, or the like) and an organic defect is generated. | 1. A method for generating an image for PET attenuation correction from an MR image, comprising: segmenting an MR image captured by MRI into regions according to pixel values,determining a radiation attenuation correction value in a region in which a radiation attenuation coefficient is considered to be uniform by referring to an existing radiation attenuation correction value table, anddetermining a radiation attenuation correction value in a region including multiple tissues having different radiation attenuation coefficients by referring to a standard image. 2. The method for generating an image for PET attenuation correction from an MR image according to claim 1, wherein the standard image is deformed to an MR image of a subject when the standard image is referred to. 3. The method for generating an image for PET attenuation correction from an MR image according to claim 2, wherein the standard image is referred to by using an image of whole body tissue or an image of some tissues of the body tissue. 4. The method for generating an image for PET attenuation correction from an MR image according to claim 1, wherein the standard image is an image for PET attenuation correction, a CT image, or a UTE image. 5. The method for generating an image for PET attenuation correction from an MR image according to claim 1, wherein the MR image and the standard image are adjusted to each other in resolution. 6. The method for generating an image for PET attenuation correction from an MR image according to claim 1, wherein at least either one of the radiation attenuation correction value table and the standard image is modified according to personal information about the subject. 7. The method for generating an image for PET attenuation correction from an MR image according to claim 1, wherein subject's own radiation attenuation correction values or the image for PET attenuation correction, or the CT image or the UTE image is repeatedly used as at least either one of the radiation attenuation correction value table and the standard image. 8. A non-transitory computer readable medium storing a program for causing a computer to execute the method as set forth in claim 1. | NATIONAL INSTITUTE OF RADIOLOGICAL SERVICES SHIMADZU CORPORATION | G06K009/00 G06T011/00 A61B005/055 A61B005/00 G01T001/16 A61B006/03 G01R033/48 G06T007/00
US09345887 | B2 | 2016-05-24 | Local and non-local sensing for cardiac pacing | 2013-08-16 | Systems and methods for determining pacing timing intervals based on the temporal relationship between the timing of local and non-local cardiac signal features are described. A device includes a plurality of implantable electrodes electrically coupled to the heart and configured to sense local and non-local cardiac signals. Sense circuitry coupled to first and second electrode pairs senses a local cardiac signal via a first electrode pair and a non-local cardiac signal via a second electrode pair. Detection circuitry is used to detect a feature of the local signal associated with activation of a heart chamber and to detect a feature of the non-local signal associated with activation of the heart chamber. A control processor times delivery of one or more pacing pulses based on a temporal relationship between timing of the local signal feature and timing of the non-local signal feature. | 1. An implantable cardiac rhythm management device, comprising: a plurality of body-implantable electrodes configured to electrically couple to a heart, the plurality of implantable electrodes including at least a first pair of electrodes and a second pair of electrodes, at least one of the first pair of electrodes to be in direct contact with myocardial tissue for use to sense at least one local cardiac signal close to the contacted myocardial tissue such that a sensed local cardiac signal is representative of an activation signal close to the contacted myocardial tissue, the second pair of electrodes to be not in direct contact with myocardial tissue for use to sense non-local cardiac signals such a sensed non-local cardiac signal effectively provides a superposition of a number of activation signals occurring within the heart;sense circuitry coupled to the first and second electrode pairs and configured to sense the at least one local cardiac signal via the at least the first pair of electrodes and the non-local cardiac signal via the second electrode pair;detection circuitry configured to detect a delay interval between an activation of at least one heart chamber and a ventricular intracardiac electrogram peak and to detect a non-local QRS interval; anda control processor configured to:determine a temporal relationship between timing of the local signal feature and timing of the non-local signal feature, wherein the temporal relationship includes a ratio of a time length of the delay interval and a time length of the non-local QRS interval;determine an atrioventricular delay based on the temporal relationship; andtime delivery of one or more pacing pulses based on the temporal relationship and the atrioventricular delay. 2. The device of claim 1, comprising a pulse generator disposed in a body-implantable housing. 3. The device of claim 1, comprising the plurality of implantable electrodes configured to electrically couple to the heart being coupled to a housing and, at least the first pair of the plurality of electrodes configured to sense the at least one local cardiac signal predominantly indicating cardiac electrical activity within the at least one heart chamber and at least the second pair of the plurality of electrodes configured to sense the non-local cardiac signal comprising a superposition of cardiac electrical signals occurring within at least one ventricular heart chamber and associated with a cardiac contraction. 4. The device of claim 1, wherein the device is configured to operate on a beat-by-beat basis. 5. The device of claim 1, wherein the delay interval between an activation of at least one heart chamber and a ventricular intracardiac electrogram peak is a QLV interval. 6. An implantable cardiac rhythm management device, comprising: a plurality of body-implantable electrodes configured to electrically couple to a heart, the plurality of implantable electrodes including at least a first pair of electrodes and a second p... | Cardiac Pacemakers Inc. | A61N001/368 A61N001/365 A61B005/0452
US20160143583 | A1 | 2016-05-26 | Hydration Monitoring System | 2015-09-25 | A hydration monitoring system for the collection of data about fluid consumption and hydration levels of athletes during training or practice sessions. The system may also measure and analyze carbohydrate consumption. The system utilizes a hydration bottle containing a fluid and to measures the amount of fluid consumed in a given time interval and wirelessly transmit the measurements; and a scale configured to measure the weight of an athlete and wirelessly transmit the measurements. The system further utilizes a data communications hub configured to receive data comprising the measurements from the hydration bottle and scale and forward the data to a computer; and a computer configured to receive the data from the hub for analysis. The computer analyzes the data and calculates whether the athlete should consume more or less fluid and/or more or less carbohydrates. | 1. A hydration monitoring system comprising: a bottle, wherein the bottle is configured to measure an amount of fluid consumed by an athlete and wirelessly transmit data corresponding to the measurement;a fluid-loss device, wherein the fluid-loss device is configured to determine an amount of fluid lost by the athlete and wirelessly transmit data corresponding to the amount;a data communications hub, wherein the hub is configured to receive data from the bottle and the fluid-loss device;a computer, wherein the computer is configured to receive the data from the hub and determine whether the athlete should consume more fluid based on the data; anda display for displaying an output conveying the data and the determination. 2. The system of claim 1, wherein the fluid-loss device is a scale configured to measure the athlete's weight. 3. The system of claim 1, wherein the fluid-loss device comprises an input device and a processor, wherein the processor is configured to calculate the amount of fluid lost by the athlete based on a value inputted with the input device. 4. The system of claim 1, wherein the fluid-loss device comprises a sensor. 5. The system of claim 4, wherein the sensor is configured to be wearable by the athlete. 6. The system of claim 5, wherein the sensor includes a porous membrane configured to adhere to the athlete's skin, a microfluidics layer in contact with the porous membrane, a sensor module fluidly connected to the porous membrane, a memory, an antenna, and an outer textile layer. 7. The system of claim 6, wherein the microfluidics layer transports the fluid from the porous membrane to the sensor module, and the sensor determines a change of an ion concentration present in the fluid. 8. The system of claim 5, wherein the sensor includes a layer with a fluid-dependent characteristic. 9. The system of claim 8, wherein the characteristic is impedance or width. 10. The system of claim 4, wherein the sensor is carried by the athlete. 11. The system of claim 10, wherein the sensor comprises a temperature sensor, a moisture sensor, or a pressure sensor. 12. The system of claim 11, wherein the sensor comprises a pedometer. 13. The system of claim 1, wherein the bottle further comprises a display and at least part of the analysis from the computer is transmitted to the display on the bottle. 14. The system of claim 13, wherein the part of the analysis transmitted to the bottle provides an indication whether additional fluid should be consumed. 15. The system of claim 1, wherein the computer is portable. 16. The system of claim 1, wherein the bottle comprises a fluid measurement device to measure the amount of fluid consumed by the athlete, wherein the fluid measurement device is configured to measure the flow of the fluid. 17. The system of claim 16, wherein the fluid measurement device is an Archimedes screw turbine, a thermal mass flow meter, or two wires comprised of resistive material, wherein one wire is shielded... |  | A61B005/00 G01F001/68 G01F023/00 G01F001/10 A61B005/01 A61B005/11
US09345452 | B2 | 2016-05-24 | Ultrasound observation system and operation method of ultrasound observation system | 2015-01-07 | An ultrasound observation system includes: an ultrasound transducer that transmits ultrasound to a subject and receives ultrasound from the subject and generates an ultrasound signal; a displacement measurement circuit for elasticity image generation that measures a displacement of the subject based on the ultrasound signal; an elasticity modulus calculation circuit that calculates an elasticity modulus of the subject based on the measured displacement; a first automatic pressurization mechanism that generates a pressurization displacement in the subject; a signal waveform analysis circuit that analyzes a periodic motion based on a waveform of a signal that is based on a spontaneous displacement of the subject; and a pressurization mechanism control circuit that controls the first automatic pressurization mechanism so as to generate a pressurization displacement that reduces an influence of the spontaneous displacement of the subject at a time that an elasticity modulus is calculated based on the signal waveform analysis result. | 1. An ultrasound observation system, comprising: an ultrasound transducer that transmits ultrasound to a subject, receives the ultrasound that is reflected by the subject, and generates an ultrasound signal based on the ultrasound that is received;a displacement measurement circuit that measures a displacement of the subject based on the ultrasound signal, in order to generate an elasticity image of the subject;an elasticity modulus calculation circuit that calculates an elasticity modulus of the subject based on a measurement result of the displacement measurement circuit;a pressurization mechanism comprising a pressurization face that applies a pressing force to the subject to generate a pressurization displacement;a signal waveform analysis circuit that analyzes a periodic motion of the subject based on a waveform of a signal that is based on a spontaneous displacement of the subject; anda pressurization mechanism control circuit that, based on the periodic motion of the subject that is analyzed by the signal waveform analysis circuit, controls the pressurization mechanism so that the pressurization mechanism applies a pressing force to the subject and generates the pressurization displacement so as to reduce an influence of the spontaneous displacement of the subject at a time that the elasticity modulus is calculated by the elasticity modulus calculation circuit. 2. The ultrasound observation system according to claim 1, wherein: the signal waveform analysis circuit analyzes a waveform of a signal that is based on the spontaneous displacement of the subject, and acquires an analysis result including a local maximum point of the spontaneous displacement and a displacement cycle; andbased on the analysis result, the pressurization mechanism control circuit controls so as to cause the pressurization mechanism to perform pressurization at a timing that corresponds with the local maximum point of the spontaneous displacement in synchrony with the displacement cycle. 3. The ultrasound observation system according to claim 2, wherein the pressurization mechanism control circuit controls so as to cause the pressurization mechanism to perform pressurization at a same timing as the local maximum point of the spontaneous displacement. 4. The ultrasound observation system according to claim 3, further comprising: a displacement amount feedback circuit that calculates a displacement amount of the subject based on the ultrasound signal and outputs the displacement amount,wherein the pressurization mechanism control circuit further controls a pressurization displacement amount applied to the subject, based on the displacement amount that is obtained from the displacement amount feedback circuit. 5. The ultrasound observation system according to claim 4, wherein, based on the periodic motion of the subject that is analyzed by the signal waveform analysis circuit, the pressurization mechanism control circuit generates a first trigger signal for causing t... | OLYMPUS CORPORATION | A61B008/00 A61B008/08 A61B008/12 A61B005/0402 A61B008/02 A61B005/00
US09901265 | B2 | 2018-02-27 | Medical device | 2015-11-06 | A powered system includes a parent module including a first battery storing a charge and a child module including a second battery storing a charge and a controller. When the parent module and the child module are in electrical communication, the controller is configured to receive power from the parent module. When the charge in the parent module is depleted, the controller is configured to backfeed power from the child module to the parent module. | 1. A powered system, comprising: a parent module including a first battery storing a charge; anda child module including a second battery storing a charge and a controller,wherein when the parent module and the child module are in electrical communication, the controller is configured to receive power from the parent module;wherein, when the charge in the parent module is depleted, the controller is configured to backfeed power from the child module to the parent module;wherein the child module further includes one or more male connector pins configured to mate with a female connector, and wherein the controller is configured to de-energize the one or more male connector pins when the controller detects that the child module is not connected to the female connector;wherein the child module further includes a display, wherein the child module is configured to receive vital sign data, and wherein the child module is configured to display the vital sign data; andwherein the display includes an expected battery life display, the expected battery life display indicating the expected battery life including both the first battery and the second battery. 2. The powered system of claim 1, wherein the first battery comprises a nine-cell battery, and wherein the second battery comprises a two-cell battery. 3. The powered system of claim 1, wherein the powered system, when the child module is connected to a mains power source, is powered exclusively off the mains power source, and when the child module is unconnected to the mains power source, the powered system initially is powered exclusively off the first battery. 4. The powered system of claim 1, wherein the powered system, when the child module is connected to a mains power source, is powered exclusively off the mains power source, and when the child module is unconnected to the mains power source, the powered system initially is powered exclusively off the first battery. 5. A medical device, comprising: a parent module including a first battery;a child module including a second battery and a controller, the child module configured to receive vital signs data, wherein the parent module and the child module are in electrical communication, the controller is configured to receive power from the parent module; andwherein when a charge in the first battery is depleted, the controller is configured to backfeed power from the child module to the parent module; wherein the child module further includes one or more male connector pins configured to mate with a female connector, and wherein the controller is configured to de-energize the one or more male connector pins when the controller detects that the child module is not connected to the female connector;wherein the first battery comprises a nine-cell battery, and wherein the second battery comprises a two-cell battery;wherein the child module further includes a display, and wherein the child module is configured to display the vital sign data on the dis... | WELCH ALLYN, INC. | H02J001/00 A61B005/022 A61B005/0235 A61B005/00 G02F001/1333 H02J007/00 A61B005/0205 H05K001/02 A61B005/021 A61B050/13 H05K001/18
US09603774 | B2 | 2017-03-28 | Evidence-based acupuncture and moxibustion meridian-acupoint treating and detecting instrument supporting multiplexed output and method of use thereof | 2014-11-09 | A multipath-output-supporting evidence-based acupuncture and moxibustion acupoint therapeutic and detecting instrument and a using method thereof. The therapeutic and detecting instrument comprises a CPU core unit, a communication module, an audio module and a plurality of detecting, diagnosis and treatment circuits. The detecting, diagnosis and treatment circuit comprises an acupoint detecting unit and an acupoint acupuncture and moxibustion unit. The acupoint detecting unit comprises a resistance detecting circuit and an A/D digital converting circuit. The acupoint acupuncture and moxibustion unit comprises a D/A pulse current generating circuit and an amplifying circuit. The CPU core unit comprises an ARM processor, a memory, an auxiliary circuit and an evidence-based acupuncture and moxibustion therapeutic system, and the ARM processor is connected with the evidence-based acupuncture and moxibustion therapeutic system. | 1. An evidence-based acupunctural meridian-acupoint treating and detecting instrument supporting multiplexed output, comprising: a CPU core unit, a communication module, a human-machine interaction module, a debugging module, an audio module, and a plurality of detecting-diagnosing-treating circuits,wherein the CPU core unit is connected to each of the plurality of detecting-diagnosing-treating circuits,wherein each of the plurality of detecting-diagnosing-treating circuits comprises an acupoint detecting unit and an acupoint acupuncture unit; the acupoint detecting unit comprises a resistance detecting circuit and an A/D converting digital circuit; the acupoint acupuncture unit comprises a D/A pulse current generating circuit and an amplifying circuit; said CPU core unit comprises an ARM processor, a memory, an auxiliary circuit and an evidence-based acupunctural treatment system;wherein the resistance detecting circuit is connected to the ARM processor via the A/D converting digital circuit; the D/A pulse current generating circuit is connected to the amplifying circuit; the amplifying circuit is connected to the acupoint acupuncturing device; the ARM processor is connected to the auxiliary circuit, the audio module, the D/A pulse current generating circuit, the communication module, the memory, the evidence-based acupunctural treatment system, and the human-machine interaction module, andwherein the resistance detecting circuit measures an acupoint resistance, when the acupoint resistance is smaller than a predetermined resistance value, the D/A pluse current generating circuit generates pulse waves to the acupoint. 2. The evidence-based acupunctural meridian-acupoint treating and detecting instrument supporting multiplexed output according to claim 1, wherein said human-machine interaction module comprises buttons and a touch screen. 3. The evidence-based acupunctural meridian-acupoint treating and detecting instrument supporting multiplexed output according to claim 1, wherein said audio module comprises an audio conversion circuit and a loudspeaker, and the audio conversion circuit is connected to the ARM processor, and the loudspeaker. 4. The evidence-based acupunctural meridian-acupoint treating and detecting instrument supporting multiplexed output according to claim 1, wherein the pulse waves are continuous waves, dilatational waves, or intermittent waves. 5. The evidence-based acupunctural meridian-acupoint treating and detecting instrument supporting multiplexed output according to claim 1, wherein the communication module is coupled to the CPU core unit and comprises connections via WIFI, GPRS, and Internet. 6. The evidence-based acupunctural meridian-acupoint treating and detecting instrument supporting multiplexed output according to claim 1, wherein the D/A pluse current generating circuit automatically generates pulse waves to the acupoint when the acupoint resistance is smaller than the predetermined resistance value. | CHENGDU UNIVERSITY OF TRADITIONAL CHINESE MEDICINE CHENGDU SUNSHEEN TECHNOLOGY CO., LTD. | A61H039/00 A61H039/02 A61H039/06 A61B005/053 G06F019/00 A61B005/00
US20160135744 | A1 | 2016-05-19 | FABRIC-BASED PRESSURE SENSOR ARRAYS AND METHODS FOR DATA ANALYSIS | 2016-01-20 | A fabric-based pressure sensor array includes: (1) a first layer including M elongated conductive strips coated thereon; (2) a second layer including N elongated conductive strips coated thereon, the M elongated conductive strips extending crosswise relative to the N elongated conductive strips to define M×N intersections; and (3) a unitary textile sheet extending between the first layer and the second layer so as to overlap the M×N intersections, the textile sheet having a variable resistivity in response to applied pressure so as to define M×N pressure sensors at locations corresponding to the M×N intersections. | 1. A fabric-based pressure sensor array, comprising: a textile sheet having a first surface and a second surface, and including M elongated conductive strips coated on the first surface and N elongated conductive strips coated on the second surface,wherein the M elongated conductive strips extend crosswise relative to the N elongated conductive strips to define M×N intersections, and the textile sheet has a variable resistivity in response to applied pressure so as to define M×N pressure sensors at locations corresponding to the M×N intersections. 2. The fabric-based pressure sensor array of claim 1, wherein the textile sheet includes fibers coated with an electrically conductive material. 3. The fabric-based pressure sensor array of claim 2, wherein the fibers are coated with an electrically conductive polymer. 4. The fabric-based pressure sensor array of claim 1, wherein the sensor array is configured to detect pressure applied at multiple ones of the M×N pressure sensors by an object positioned on the textile sheet. 5. A fabric-based pressure sensor array, comprising: a textile sheet having a first surface and a second surface;a plurality of contact pads disposed adjacent to the first surface of the textile sheet; anda ground contact pad disposed adjacent to the second surface of the textile sheet and sized so as to at least partially overlap each of the plurality of contact pads to define a plurality of overlapping regions,wherein the textile sheet exhibits piezoresistivity so as to define a plurality of pressure sensors at locations corresponding to the plurality of overlapping regions. 6. The fabric-based pressure sensor array of claim 5, wherein the ground contact pad provides a common ground connection for the plurality of pressure sensors. 7. The fabric-based pressure sensor array of claim 5, wherein the plurality of contact pads are coated on the first surface of the textile sheet. 8. The fabric-based pressure sensor array of claim 5, further comprising a first layer disposed adjacent to the first surface of the textile sheet, and the plurality of contact pads are coated on the first layer. 9. The fabric-based pressure sensor array of claim 5, wherein the ground contact pad is coated on the second surface of the textile sheet. 10. The fabric-based pressure sensor array of claim 5, further comprising a second layer disposed adjacent to the second surface of the textile sheet, and the ground contact pad is coated on the second layer. 11. The fabric-based pressure sensor array of claim 5, wherein the ground contact pad is sized to provide flexibility to the fabric-based pressure sensor array. 12. A non-transitory computer-readable storage medium, comprising executable instructions to: receive a pressure distribution of an object as applied against an array of pressure sensors;select a pressure value from the pressure distribution as an outlier pressure value;adjust the outlier pressure value according to neighboring pressure val... |  | A61B005/00 G01L001/18
US20160128605 | A1 | 2016-05-12 | Apparatus for Calculating a Parameter of a Skin Surface Area | 2014-05-22 | Provided is a mobile autonomous apparatus (1) for measuring and analysing at least one parameter representing the fat of a skin surface area comprising at least one probe (10) comprising at least two pairs of electrodes intended to come into contact with the skin surface area defining a frrst pair of injecting electrodes (11, 12), and a second pair ofreceiving electrodes (13, 14) for measuring the response to the injection, an energy source (5) for producing, at the frrst pair of injecting electrodes (11, 12), a variable electrical signal with a current of intensity i included in the range [60 μA to 4000 μA] and a frequency included in the range [1 k Hz to 1000 kHz], a data processing and control unit (CPU) for determining the elements relative to the bio-impedance (Z) of said area and said at least one parameter representing the cellulite of the skin surface area. | 1. Autonomous mobile apparatus (1) for measuring and analyzing at least one parameter representing the fat of a skin surface area, comprising: At least one probe (10) comprising at least two electrode pairs intended to come into contact with the skin surface area defining a first pair of injecting electrodes (11, 12) and a second pair of receiving electrodes (13, 14) for measuring the response to the injection;A power source (5) for producing in the first pair of injecting electrodes (11, 12), a variable electrical signal with a variable current intensity (i) within the range of 60 μA to 4000 μA and with a frequency within the range of 1 kHz to 1000 kHz;A data control processing unit (CPU) for determining the information related to the bio-impedance (Z) of said area, with said at least one parameter representing the fat in the skin surface area. 2. Apparatus described in claim 1, in which said variable electrical signal with current intensity (i) is a sinusoidal alternating signal as a function of time. 3. Apparatus described in claim 1, in which said variable electrical signal with current intensity (i) is a rectangular signal, such as a square signal, as a function of time. 4. Apparatus described in claim 1, in which the parameter is the subcutaneous fat and the data control processing unit (CPU) makes it possible to determine the quantity of subcutaneous fat that is equal or proportional to: A*((Ret+Rinft)/2)−B In which Ret is the resistance extrapolated to a null frequency,Rinft is the resistance extrapolated to an infinite frequency,And A and B are constants. 5. Apparatus described in claim 1, in which the parameter is cellulite and the data control processing unit (CPU) makes it possible to determine the thickness of the cellulite being equal or proportional to: Ln (Rinft/C)/DIn which Rinft is the resistance extrapolated to an infinite frequency, and in which C and D are constants. 6. Apparatus described in claim 1, in which the data control processing unit (CPU) makes it possible to determine a characterization of the parameter representing cellulite and that comprises a natural integer on a scale between 0 and n, representing a stage of progression of the cellulite. 7. Apparatus described in claim 1, in which the data control processing unit (CPU) makes it possible to determine a characterization of the parameter representing cellulite, comprising a length of a hypodermal attachment for a unit length of the skin area, or the hypodermal attachment size for a unit surface area of the skin area. 8. Apparatus described in claim 1, in which the data control processing unit (CPU) makes it possible to determine a characterization of the parameter representing cellulite and being a surface roughness of the skin area. 9. Apparatus described in claim 1, in which the data control processing unit (CPU) makes it possible to determine a characterization of the parameter representing cellulite, which comprises a fat and water content percentage ... |  | A61B005/053 A61H015/00 A61H009/00 A61B005/00
US09351709 | B2 | 2016-05-31 | Image processing method and apparatus and program | 2014-08-19 | An image processing method for performing registration between a first medical image and a second medical image both being three-dimensional images is provided. The image processing method includes extracting a first vascular image in the first medical image and a second vascular image in the second medical image, detecting at least one vascular portion structure including a plurality of vascular portions close or connected to each other with respect to each of the first and second vascular images extracted, for each of a plurality of combinations of the at least one vascular portion structure in the first vascular image and the at least one vascular portion structure in the second vascular image, performing coordinate transformation on at least one of the first and second vascular images such that a deviation between the vascular portion structures becomes small, and calculating similarity in a prescribed area including the vascular portion structures between the first and second vascular images, and performing coordinate transformation for the registration using the coordinate transformation highest in the calculated similarity. | 1. An image processing method for performing registration between a first medical image and a second medical image both being three-dimensional images, the image processing method comprising: extracting a first vascular image in the first medical image and a second vascular image in the second medical image;detecting at least one vascular portion structure including a plurality of vascular portions close or connected to each other with respect to each of the first and second vascular images extracted;for each of a plurality of combinations of the at least one vascular portion structure in the first vascular image and the at least one vascular portion structure in the second vascular image, performing coordinate transformation on at least one of the first and second vascular images such that a deviation between the vascular portion structures becomes small, and calculating similarity in a prescribed area including the vascular portion structures between the first and second vascular images; andperforming coordinate transformation for the registration using the coordinate transformation highest in the calculated similarity. 2. A non-transitory computer-readable medium storing program code configured to cause a computer to perform: extracting a first vascular image in a first medical image and a second vascular image in a second medical image;detecting at least one vascular portion structure including a plurality of vascular portions close or connected to each other with respect to each of the first and second vascular images extracted by the extraction step;for each of a plurality of combinations of the at least one vascular portion structure in the first vascular image and the at least one vascular portion structure in the second vascular image, performing coordinate transformation on at least one of the first and second vascular images such that a deviation between the vascular portion structures becomes small, and calculating similarity in a prescribed area including the vascular portion structures between the first and second vascular images; andperforming coordinate transformation for registration between the first and second medical images using the coordinate transformation highest in the similarity. | General Electric Company | A61B008/00 A61B005/00 G06T007/00 A61B008/08 A61B005/02 A61B005/055 G01R033/56
US09964613 | B2 | 2018-05-08 | Shimming assistance unit, shimming assistance method, MRI apparatus and magnet apparatus | 2014-06-24 | The shimming work assistance unit performs singular value decomposition of a response matrix, which represents the relationship between an error magnetic field distribution and an adjusted magnetic moment placement distribution. From the multiple eigenmodes obtained, the eigenmodes are selected and added one by one in order from the eigenmode with the highest singular value, and the residual magnetic field error, which represents the fluctuation range of the difference between the magnetic field distribution, generated by the placement of the shimming magnetic moments corresponding to said eigenmode, and the error magnetic field distribution, is displayed on a display unit as a function graph of eigenmode order (line graph (1)). | 1. A shimming assistance unit comprising: a means for calculating an error magnetic field distribution of a difference between a measured magnetic field distribution or a magnetic field distribution calculated based on a temporary placement of a magnetomotive force, and a target magnetic field distribution;a means for calculating a singular value and a magnetic field intensity of each of eigenmodes by performing singular value decomposition of a response matrix that represents a relationship between the error magnetic field distribution and a placement distribution of a shimming magnetic moment that generates a magnetic field for cancelling the error magnetic field distribution;a means for calculating a residual magnetic field distribution of a difference between the magnetic field distribution generated by the magnetic field intensity of an eigenmode that is a target of calculation of the shimming magnetic moment and the error magnetic field distribution by adding a selected eigenmode to the target of calculation of the shimming magnetic moment while selecting an eigenmode one by one in order from the highest singular value from among the eigenmodes;a means for calculating the residual magnetic field error that represents a fluctuation range of the calculated residual magnetic field distribution as a function value of an order of the eigenmode that is a number put to an order of level of the singular value; anda means for displaying on a display device the calculated residual magnetic field error as a function graph of the order of the eigenmode. 2. The shimming assistance unit according to claim 1 further comprising: a means for calculating a sum total of the shimming magnetic moment placed corresponding to the eigenmode for the target of calculation of the shimming magnetic moment, as an index representing a physical amount of the shimming magnetic moment to be placed, and for displaying on a display device the calculated index as a function graph of the order of the eigenmode. 3. A shimming assistance method having a computer execute: calculating an error magnetic field distribution of a difference between a measured magnetic field distribution or a magnetic field distribution calculated based on a temporary placement of a magnetomotive force, and a target magnetic field distribution;calculating a singular value and a magnetic field intensity of each of eigenmodes by performing singular value decomposition of a response matrix that represents a relationship between the error magnetic field distribution and a placement distribution of a shimming magnetic moment that generates a magnetic field for cancelling the error magnetic field distribution;calculating a residual magnetic field distribution of a difference between the magnetic field distribution generated by the magnetic field intensity of an eigenmode that is a target of calculation of the shimming magnetic moment and the error magnetic field distribution by adding a selected eigenmode ... | Hitachi, Ltd. | G01V003/00 G01R033/3873 A61B005/055 G01R033/565 H01F007/20 G01R033/3875
US09717430 | B2 | 2017-08-01 | Real-time multi-functional ECG signal processing system, DSPE for the ECG signal processing system, and method thereof | 2014-06-12 | An electrocardiogram (ECG) signal processing system is provided. The ECG signal processing system comprises an analog-to-digital converter (ADC) configured to convert an input analog ECG signal into a digital ECG signal, and a digital signal processing engine (DSPE) coupled to the ADC to receive the digital ECG signal. The DSPE is configured to decompose and reconstruct the digital ECG signal. A dynamic system clock source is coupled to the ADC and the DSPE for dynamic signal sampling, the dynamic system clock source clocking the ADC and the DSPE at a first frequency f1 to detect one or more first parameters of the input analog ECG signal and at a second frequency f2 to detect one or more second parameters of the input analog ECG signal. | 1. An electrocardiogram (ECG) signal processing system, the ECG signal processing system comprising: an analog-to-digital converter (ADC) operable for sampling an input analog ECG signal at a first frequency f1 and a second frequency f2 to convert the input analog ECG signal into a digital ECG signal;a digital signal processing engine (DSPE) coupled to the ADC to receive the digital ECG signal, the DSPE being configured to sample the digital ECG signal at the first frequency f1 and the second frequency f2 to decompose and reconstruct the digital ECG signal; anda dynamic system clock source coupled to the ADC and the DSPE for dynamic adaptive signal sampling, the dynamic system clock source clocking the ADC and the DSPE at the first frequency f1 to detect one or more first parameters of the input analog ECG signal and at the second frequency f2 to detect one or more second parameters of the input analog ECG signal, wherein the DSPE comprises:a wavelet transformation (WT) unit, the WT unit comprising a plurality of scales, wherein the WT unit is adapted to decompose the digital ECG signal into a plurality of wavelets, each wavelet being output from one of the plurality of scales; anda plurality of signal processing blocks, each of the signal processing blocks coupled to one or more outputs of the plurality of scales and configured to receive and process the one or more wavelets from the respective outputs, wherein the signal processing blocks provide processing functions which differ from one another, wherein the plurality of signal processing blocks comprise a cardiac features extraction block, the cardiac features extraction block receiving a high pass filtering of a predetermined one of the plurality of scales, wherein the cardiac features extraction block is configured to process a QRS complex extraction when the dynamic system clock source is clocking the ADC and the DSPE at the first frequency f1, and to process P wave extraction and T wave extraction when the dynamic system clock source is clocking the ADC and the DSPE at the second frequency f2. 2. The ECG signal processing system of claim 1, wherein the ECG signal processing system is configured to detect a QRS complex when f1 is 500 Hz, and wherein the plurality of scales comprises six scales, the plurality of scales having 3 dB bandwidths of around 125 to 250 Hz; 62.5 to 125 Hz; 18 to 58.5 Hz; 8 to 27 Hz; 4 to 13.5 Hz; and 2 to 6.5 Hz for the first to the sixth scales respectively, when the dynamic system clock source is clocking at the first frequency f1. 3. The ECG signal processing system of claim 1, wherein the ECG signal processing system is configured to detect P and T waves when f2 is 250 Hz, and wherein the plurality of scales comprises six scales, the plurality of scales having 3 dB bandwidths of around 62.5 to 125 Hz; 18 to 58.5 Hz; 8 to 27 Hz, 4 to 13.5 Hz; 2 to 6.5 Hz; and 1 to 3.25 Hz for the first to the sixth scales respectively, when the dynamic system clock source is c... | Agency of Science, Technology and Research | H03M001/00 A61B005/04 A61B005/00 A61B005/0428 A61B005/0472 A61B005/0476 G01R019/25
US09955889 | B2 | 2018-05-01 | Registration maps using intra-cardiac signals | 2014-11-03 | Cardiac catheterization is facilitated by generating first and second electroanatomic maps of a heart of a subject and designating common spatial locations that correspond to first electrical events on the first electroanatomic map and second electrical events on the second electroanatomic map. The common spatial locations of the first electroanatomic map and the second electroanatomic map are aligned to establish an aligned map, and using the location data on the aligned map to guide a probe to a point of interest. | 1. A method, comprising of steps of: using a first system to generate a first electroanatomic map of a heart of a living subject;using a second system that is different to the first system to generate a second electroanatomic map of the heart;identifying areas with common electrical events in the first and second electroanatomic maps and finding in those areas at least three points in each of the first and second electroanatomic maps where the local activation times of points in the first and second electroanatomic maps do not differ by more than a first threshold value and where the respective distances between the points on each electroanatomic map do not exceed a second threshold value;using the at least three points to designate common spatial locations that correspond to first electrical events on the first electroanatomic map and second electrical events on the second electroanatomic map;aligning the common spatial locations of the first electroanatomic map and the second electroanatomic map to establish a set of aligned maps; andusing the set of aligned maps to guide a probe to a point of interest. 2. The method according to claim 1, further comprising the step of displaying the set of aligned maps. 3. The method according to claim 1, further comprising the step of: introducing a catheter into the heart to obtain electrical data for at least one of the first electroanatomic map and the second electroanatomic map. 4. The method according to claim 3, further comprising analyzing the electrical data to determine dominant frequencies at respective locations in the heart. 5. The method according to claim 3, further comprising analyzing the electrical data to determine phase information at respective locations in the heart. 6. The method according to claim 5, wherein analyzing the electrical data comprises converting or transforming a unit of measurement of the first electroanatomic map to be compliant with a unit of measurement of the second electroanatomic map. 7. The method according to claim 1, wherein at least one of the first electroanatomic map and the second electroanatomic map is obtained by body surface mapping. 8. The method according to claim 1, wherein the first electroanatomic map is obtained by body surface mapping and the second electroanatomic map is obtained using an intracardiac mapping catheter. 9. The method according to claim 1, wherein the first electroanatomic map and the second electroanatomic map are obtained using an intracardiac mapping catheter. 10. A data processing system comprising: a processor;a visual display screen; anda memory accessible to the processor storing programs and data objects therein, the programs including an electroanatomic map generator, an image registration program, an analysis program and a graphical user interface configured to present graphical information on the visual display screen, wherein execution of the programs cause the processor to perform the steps of:invoking the electr... | BIOSENSE WEBSTER (ISRAEL) LTD. | A61B005/044 A61B005/06 A61B005/04 A61B005/042 A61B005/00 A61B005/0452 G06F019/00 A61B005/046 A61B034/20
US20160139631 | A1 | 2016-05-19 | WRIST DEVICE EFFICIENCY | 2014-11-14 | A wrist device includes a curved body arranged to fit the wrist device around a user's wrist, an activity tracker circuitry configured to measure physical activity of the user of the wrist device, a processing circuitry configured to acquire physical activity-related measurement data from the activity tracker circuitry and to process the physical activity-related measurement data into physical activity-related metric characterizing the physical activity of the user, and a user interface element configured to display the physical activity-related metric. The user interface element includes a light emitting diode matrix layer including a plurality of light emitting diodes arranged in a matrix form, and a fluorescent layer configured to alter the wavelength of the light emitted by the light emitting diode matrix layer. | 1. A wrist device comprising: a curved body arranged to fit the wrist device around a user's wrist;an activity tracker circuitry configured to measure physical activity of the user of the wrist device;a processing circuitry configured to acquire physical activity-related measurement data from the activity tracker circuitry and to process the physical activity-related measurement data into physical activity-related metric characterizing the physical activity of the user; anda user interface element configured to display the physical activity-related metric, wherein the user interface element comprises:a light emitting diode matrix layer comprising a plurality of light emitting diodes arranged in a matrix form; anda fluorescent layer configured to alter the wavelength of the light emitted by the light emitting diode matrix layer. 2. The wrist device of claim 1, wherein the user interface element further comprises an optical conductor layer, and wherein the optical conductor layer is configured to enhance the contrast of the user interface element by blocking at least some of the scattered light from the light emitting diode matrix layer. 3. The wrist device of claim 2, wherein the optical conductor layer is arranged to be situated between the light emitting diode matrix layer and the fluorescent layer. 4. The wrist device of claim 2, wherein the optical conductor layer comprises a plurality of optical guides arranged in a matrix form, wherein the optical guides are made of a substantially transparent material to the wavelength of the plurality of light emitting diodes, and wherein the plurality of optical guides are arranged to be situated on top of the plurality of light emitting diodes so that the optical guides are aligned with the light emitting diodes. 5. The wrist device of claim 4, wherein the optical conductor layer further comprises a plurality of light guards in a matrix form, wherein the light guards are made of a substantially light blocking material, and wherein the plurality of light guards are arranged to be situated between the optical guides. 6. The wrist device of claim 5, wherein the light guards are further configured to reflect at least some of the light back to the optical guides. 7. The wrist device of claim 1, wherein the light emitting diode matrix layer comprises three layers: a first layer electrically coupled to an electrical terminal;a second layer comprising the plurality of light emitting diodes, wherein the plurality of light emitting diodes are electrically coupled with the first layer; anda third layer electrically coupled with plurality of light emitting diodes, wherein the second layer comprises electrical lead-troughs electrically coupling the first and the third layers. 8. The wrist device of claim 1, wherein the light emitting diode matrix layer is laminated inside a laminating structure. 9. The wrist device of claim 1, wherein the fluorescent layer is configured to increase the wavelength of the light... |  | G06F001/16 F21K099/00 G06F003/044 F21V008/00 A61B005/11 A61B005/00
US09913611 | B2 | 2018-03-13 | Garment integrated sensing system and method | 2015-11-10 | A system for monitoring biometric signals of a user comprising: a set of wireless sensor interfaces coupled to a garment, each of the wireless sensor interfaces comprising: 1) an electrode layer comprising a receiving region, 2) a positional identifier, associated with a position on the garment, and 3) a retention subsystem; a set of wireless sensor modules, each of the set of wireless sensor modules comprising: a contact region electrically coupleable to the receiving region of the electrode layer, a set of sensors configured to detect a set of biometric signal types, and a positional interrogator configured to identify the position associated with the corresponding wireless sensor interface; and a control module, communicatively coupled to the set of wireless sensor modules, wherein the control module queries a subset of the set of biometric signal types for transmission from each of the set of wireless sensor modules based on their positions. | 1. A biometric monitoring system, comprising: a set of wireless sensor interfaces coupled to a garment, each of the wireless sensor interfaces comprising: an electrode layer configured for electrical contact with a user during operation, the electrode layer comprising a receiving region,a positional identifier associated with a position on the garment, anda retention subsystem;a set of wireless sensor modules, each of the wireless sensor modules configured to be received by and removably coupled to the retention subsystem of a corresponding wireless sensor interface of the set of wireless sensor interfaces and comprising: a contact region electrically coupleable to the receiving region of the electrode layer of the corresponding wireless sensor interface,a set of sensors configured to detect a set of biometric signal types, anda positional interrogator configured to identify, upon interacting with the positional identifier of the corresponding wireless sensor interface, the position associated with the corresponding wireless sensor interface; anda control module communicatively coupled to the set of wireless sensor modules and configured to automatically select, for each wireless sensor module of the set of wireless sensor modules, a subset of the set of biometric signal types detected by an activated set of sensors corresponding to the wireless sensor module for transmission from the wireless sensor module, wherein the subset is selected based on a detected position of the wireless sensor module. 2. The system of claim 1, wherein the electrode layer further comprises a signal communication region, the signal communication region including a set of conductive leads, at least a portion of each of the conductive leads configured to electrically couple to the contact region of the corresponding wireless sensor module, and a set of biosensing contacts, coupled to the set of conductive leads of the signal communication region, the set of biosensing contacts composed of a conductive polymer, at least a portion of each of the biosensing contacts disposed adjacent to and in electrical contact with a skin region of the user. 3. The system of claim 1, wherein the retention subsystem comprises a fabric layer covering the electrode layer, wherein a portion of a periphery of the fabric layer is fixed to the garment to form a pocket comprising an opening. 4. The system of claim 1, wherein the positional identifier comprises a set of electronic circuit components configured to exhibit a known characteristic when an electrical current passes therethough, and wherein the positional interrogator is configured to induce the set of electrical components to exhibit the known characteristic. 5. The system of claim 1, wherein the positional identifier comprises a wireless identification integrated circuit configured to wirelessly provide an encoded signal. 6. The system of claim 1, wherein the positional identifier comprises an integrated circuit configured to pro... | MAD Apparel, Inc. | G08C019/22 A61B005/00
US09327076 | B2 | 2016-05-03 | Electronically and remotely controlled pill and system for delivering at least one medicament | 2013-08-26 | An electronically and remotely controlled pill (500) or medicament delivery system is provided. The pill (500) includes a housing (102); a medicament reservoir (104) for storing a medicament; an electronically controlled release valve or hatch (106) for dispensing one or more medicaments stored in the medicament reservoir (104) while traversing the gastrointestinal tract; control and timing circuitry (108) for opening and closing the valve (106); and a battery (109). The control and timing circuitry (108) opens an closes the valve (106) throughout a dispensing time period in accordance with a preset dispensing timing pattern which is programmed within the control and timing circuitry (108). RF communication circuitry receives control signals for remotely overriding the preset dispensing timing pattern, reprogramming the control and timing circuitry (108) or terminating the dispensing of the medicament within the body. The pill (500) includes an RFID tag (508) for tracking, identification, inventory and other purposes. | 1. A medicament delivery system for dispensing a medicament to a gastrointestinal tract of a body, said system comprising: a housing;an opening in the housing;a reservoir for storing said medicament within said housing, the reservoir being in fluid communication with the opening;a pressure mechanism disposed in the housing for applying pressure to the reservoir to direct the medicament toward the opening;control and timing circuitry programmed with a dispensing timing pattern and comprising a clock, wherein the timing circuitry is adapted to transmit a release signal according to the timing pattern and the clock independent of a sensed physiological parameter of the gastrointestinal tract and independent of a contemporaneous signal from outside the body; anda release controller for receiving the release signal and controlling a release of the medicament from the reservoir. 2. The system according to claim 1, further comprising at least one antenna and RF communication circuitry for receiving RF control signals to perform an action selected from the group consisting of reprogramming the control and timing circuitry with another dispensing timing pattern, terminating the dispensing of said medicament, dispensing another medicament, changing the dispensing timing pattern, and dispensing said medicament according to a person's vitals. 3. The system according to claim 2, further comprising a soluble membrane encapsulating said at least one antenna. 4. The system according to claim 1, further comprising a piezo-electric element capable of being vibrated at at least one predetermined frequency. 5. The system according to claim 4, further comprising a data structure for correlating said at least one predetermined frequency with an action, and wherein said action is selected from the group consisting of reprogramming the control and timing circuitry with another dispensing timing pattern, terminating the dispensing of said medicament, dispensing another medicament, changing the dispensing timing pattern, and dispensing said medicament according to a person's vitals. 6. The system according to claim 1, further comprising a transmitter. 7. The system according to claim 1, further comprising an RFID tag. 8. The system according to claim 1, wherein said control and timing circuitry comprises a start timer mechanism wherein said start timer mechanism transmits a signal to said timing circuitry for clocking a dispensing time period, and wherein said timing circuitry transmits a signal to said release controller for dispensing said medicament substantially according to said dispensing timing pattern during said dispensing time period. 9. The system according to claim 1, wherein said dispensing timing pattern correlates an approximate quantity of said medicament to be dispensed during each moment of time during a dispensing time period. 10. The system according to claim 1, further comprising another reservoir for storing another medicament. 11. The syst... | Medimetrics Personalized Drug Delivery | A61M005/00 A61M005/172 A61B005/00 A61M031/00 A61B005/07
US20160147941 | A1 | 2016-05-26 | MEDICAL SUPPORT SYSTEM | 2016-01-29 | In an initial emergency response step in which a patient is transported from a site of emergency to a hospital, when a paramedic starts treatment of the patient using a medical device such as a triage device or a vital sign measurement device, treatment start information is transmitted from the medical device to a medical support server of a medical support system. According to the treatment start information, the medical support server creates timeline information for managing medical care information of the patient along a time axis and starts management of the medical care information generated in the initial emergency response step. | 1. A medical support system comprising: a medical device used for treatment of patients; anda medical support server that is communicably connected to the medical device over a communication network, whereinthe medical device is used by a paramedic to treat the patient in an initial emergency response step in which the patient is transported from a site as a transport source to a medical facility as a transport destination, and the medical device includes a treatment start information transmission unit for transmitting treatment start information indicating that treatment for the patient starts to the medical support server upon the medical device being used, andthe medical support server includes a treatment start information reception unit for receiving the treatment start information, and a timeline creation unit for creating timeline information for managing medical care information of the patient along a time axis in the case where the treatment start information reception unit receives the treatment start information. 2. The medical support system according to claim 1, wherein a patient ID for identifying the patient is imparted to the treatment start information, andthe timeline creation unit creates one piece of timeline information for each patient ID. 3. The medical support system according to claim 2, wherein the patient ID is a temporary ID that is temporarily issued in the initial emergency response step. 4. The medical support system according to claim 2, wherein the medical device includes a patient ID reception unit for receiving the patient ID input manually or the patient ID transmitted from the medical support server, andthe patient ID received by the patient ID reception unit is imparted to the treatment start information. 5. The medical support system according to claim 2, wherein the treatment start information reception unit is adapted to receive the treatment start information from a plurality of the medical devices, andthe timeline creation unit creates the timeline information based on the initially received treatment start information among a plurality of pieces of the treatment start information including the same patient ID received by the treatment start information reception unit. 6. The medical support system according to claim 1, wherein the treatment start information reception unit receives first medical care information acquired from the patient by the medical device, as the treatment start information. 7. The medical support system according to claim 1, wherein the medical support server includes a medical care information registration unit for receiving the medical care information transmitted from the medical device and registering the received medical care information in the timeline information. 8. The medical support system according to claim 1, wherein at least one of a triage device for classifying the patient according to a degree of severity, a vital sign measurement device for measuring vital ... | FUJIFILM Corporation | G06F019/00 A61B005/145 A61G003/00 A61B008/00 A61B006/03 A61B006/00 A61B005/0205 A61B005/0402
US09875538 | B2 | 2018-01-23 | Reporting tool with integrated lesion stager | 2014-07-28 | A device and method integrates a lesion stager in a reporting tool. The method includes receiving a user input. The user input is image data including at least one variable value. The variable value is indicative of a variable parameter of the image data. The method includes associating each of the at least one variable value with a corresponding variable parameter to generate at least one variable-value pair. The method includes determining a computed output value as a function of select ones of the at least one variable-value pair. | 1. A method, comprising: receiving, by a processor, a user input, the user input being medical image data including at least one variable value, the variable value being a measurement of a feature of the medical image data indicative of a variable parameter of the medical image data, wherein the variable parameter is a required parameter for determining a computed output value;associating, by the processor, each of the at least one variable value with a corresponding variable parameter to generate at least one variable-value pair;determining, by the processor, whether at least one further variable-value pair is required to determine the computed output value;determining, by the processor, select ones of the at least one further required variable-value pair;requesting, by a user interface, for the select ones of the at least one further required variable-value pair;receiving, by the user interface, the select ones of the at least one further required variable-value pair; anddetermining, by the processor, the computed output value as a function of select ones of the at least one variable-value pair and select ones of the at least one further required variable-value pair. 2. The method of claim 1, further comprising: determining, by the processor, a context of the user input; andextracting, by the processor, select ones of the at least one variable value from the user input as a function of the context. 3. The method of claim 1, wherein the determination for at least one further variable-value pair is a function of at least one of predetermined settings and predetermined rules. 4. The method of claim 1, further comprising: integrating, by the processor, the computed output value in a report for the user input. 5. The method of claim 1, wherein the at least one variable value is included in the user input as at least one of an annotation, a free text, and a structured content. 6. The method of claim 1, further comprising: determining, by the processor, at least one further computed output value as a function of the computed output value. 7. The method of claim 1, wherein the medical image data is a magnetic resonance imaging (MRI), the at least one variable value is a width and length of a lesion, the variable parameter is dimensions of the lesion, and the computed output value is a Breast Imaging-Reporting and Data System (BI-RADS) score. 8. A reporting tool, comprising: a processor configured to receive a user input, the user input being medical image data including at least one variable value, the variable value being a measurement of a feature of the medical image data indicative of a variable parameter of the medical image data, wherein the variable parameter is a required parameter for determining a computed output value; anda memory arrangement configured to store the user input, the at least one variable value, and the variable parameter,wherein the processor is configured to associate each of the at least one variable value with a cor... | Koninklijke Philips N.V. | G06K009/00 G06T007/00 G06F019/00 A61B005/055 G06K009/46 G06K009/62 A61B005/05
US10046178 | B2 | 2018-08-14 | Automatic correlation modeling of an internal target | 2016-01-08 | A method and apparatus to automatically control the timing of an image acquisition by an imaging system in developing a correlation model of movement of a target within a patient. | 1. A method, comprising: automatically triggering, by a processing device, image acquisition of a plurality of pretreatment images of a target, wherein the plurality of pretreatment images are internal to a patient, wherein acquisition of each of the plurality of pretreatment images is triggered at a different time; andgenerating, by the processing device, a correlation model distinguishing between inspiration and expiration movements that maps the movements of an external marker to a target location of the target using the plurality of pretreatment images, wherein said automatically triggering comprises automatically triggering image acquisition at substantially evenly-distributed points of a periodic cycle, and wherein said automatically triggering each of the plurality of pretreatment images comprises:obtaining historical data of movement of the target in one or more previous periodic cycles;dividing the historical data into a plurality of regions based on magnitude of movement of the target;distinguishing between the inspiration and expiration movements of the historical data;establishing a plurality of model metric regions using the plurality of regions and the distinguished inspiration and expiration movement, wherein the plurality of model metric regions correspond to a plurality of phases of the periodic cycle;determining whether a current sample of the target movement is in a desired phase of the periodic cycle by categorizing the current sample into one of the plurality of model metric regions; andsending an imaging command to an imaging system to acquire the pretreatment image when the current sample is in the desired phase of the periodic cycle. 2. The method of claim 1, wherein said obtaining comprises acquiring a plurality of data points representative of a corresponding plurality of positions over time of an external marker associated with a patient, wherein the plurality of positions of the external marker define an external path of movement of the external marker, and the external path of movement defines the periodic cycle of the patient. 3. The method of claim 2, wherein said generating the correlation model comprises: identifying a path of movement of the target based on the plurality of data points and the plurality of pretreatment images at the substantially evenly-distributed points of the periodic cycle; anddeveloping the correlation model using the path of movement of the target. 4. The method of claim 2, further comprising updating the correlation model in response to an acquisition of a new data point. 5. The method of claim 1, further comprising: deriving a target position of the target based on the correlation model,sending a position signal associated with the target position to a beam generator controller; andcontrolling the beam generator controller to direct a beam at the target. 6. The method of claim 1, further comprising: periodically generating positional data about the target by automatically acquiring a... | ACCURAY INCORPORATED | G06K009/00 A61N005/10 A61B006/00 A61B005/11 A61B090/00 A61B034/10 A61B005/113 A61B005/00
US09339190 | B2 | 2016-05-17 | Charger with data transfer capabilities | 2006-02-15 | Gastric apparatus (18) is provided, including one or more sensors (70, 72, 74), adapted to generate respective sensor signals responsively to a gastrointestinal (GI) tract physiological parameter of a GI tract of a subject, and an implantable control unit (90), comprising a rechargeable battery and a first set of one or more transducers. The implantable control unit (90) is adapted to receive the sensor signals, using one or more of the transducers of the first set, and transmit data responsively to the sensor signals, using one or more of the transducers of the first set. An external control unit (200), including a power source and a second set of one or more transducers, is adapted to drive the power source to inductively transfer energy via one or more transducers of the second set, so as to recharge the battery, and receive the transmitted data, using one or more transducers of the second set. | 1. Apparatus, comprising: one or more sensors, adapted to generate respective sensor signals responsively to a physiological parameter of a subject;an implantable control unit, comprising a rechargeable battery and a first set of one or more transducers, the implantable control unit adapted to receive the sensor signals, and transmit data responsively thereto, using one or more of the transducers of the first set; andan external control unit, comprising a power source, a second set of one or more transducers, and a display screen, the external control unit adapted to: drive the power source to wirelessly transfer energy via one or more of the transducers of the second set,receive the transmitted data, using one or more of the transducers of the second set,be in connective interface with a remote service provider over a connection over which the external control unit is recharged,send to the remote service provider eating-related information including a number of meals consumed per time period by the subject,receive, from the remote service provider, subject-specific recommendations generated by the remote service provider based on an analysis of received eating-related information, anddisplay the subject-specific recommendations on the display screen,wherein the implantable control unit is adapted to receive the transmitted energy, using one or more of the transducers of the first set, and charge the battery using the energy. 2. The apparatus according to claim 1, wherein the physiological parameter includes a gastrointestinal (GI) tract physiological parameter of a GI tract of the subject, and wherein the one or more sensors are adapted to generate the respective sensor signals responsively to the GI tract physiological parameter. 3. The apparatus according to claim 2, wherein the GI tract physiological parameter includes a parameter indicative of electrical activity of the GI tract, and wherein the sensors are adapted to generate the signals responsively to the electrical activity parameter. 4. The apparatus according to claim 2, wherein the GI tract physiological parameter includes a parameter indicative of an impedance of the GI tract, and wherein the sensors are adapted to generate the signals responsively to the impedance parameter. 5. The apparatus according to claim 2, comprising an implantable stimulator coupled to the implantable control unit, wherein the implantable control unit is adapted to drive the stimulator to apply stimulation to the GI tract. 6. The apparatus according to claim 5, wherein the implantable control unit is adapted to configure the stimulation to treat diabetes of the subject. 7. The apparatus according to claim 5, wherein the implantable control unit is adapted to configure the stimulation to treat obesity of the subject. 8. The apparatus according to claim 5, wherein the external control unit is adapted to monitor, using the received data, information regarding the applied stimulation, the information sele... | METACURE LIMITED | A61B005/07 A61B005/00 A61B005/0488 A61F005/00 A61N001/36 G06F019/00 G06Q030/02 A61N001/362 A61N001/372
US09332948 | B2 | 2016-05-10 | System and method for patient positioning in cone-beam tomography | 2009-05-15 | A device for positioning a patient within an image volume of a cone-beam tomography system. The device includes a volume indicator adapted to indicate at least a front boundary of the image volume and having a horizontal indicator for horizontal alignment. The device also includes a head clamp adapted to position at least a portion of the head of the patient within the front boundary of the image volume indicated by the volume indicator. | 1. A device for positioning a patient, the device comprising: a volume indicator that includes: a boundary structure that delineates a portion of an outer boundary of an image volume, the image volume being a volume of space through which x-rays, emitted by an x-ray source, pass through at each of a plurality of scanning positions during a scanning operation, wherein a contour of the boundary structure matches a contour of the portion of the outer boundary of the image volume, anda structure for horizontal alignment that delineates another portion of the outer boundary of the image volume, wherein a contour of the structure for horizontal alignment matches a contour of the other portion of the outer boundary of the image volume; anda head clamp adapted to position at least a portion of a head of the patient within the outer boundary of the image volume,wherein the device is attached to a tomography system that includes the x-ray source. 2. The device as defined in claim 1, wherein the head clamp comprises: a plurality of ear tubes disposed at ends of the head clamp for restricting lateral movement of the head of the patient and for facilitating alignment of the head of the patient in a substantially horizontal plane relative to the structure for horizontal alignment, anda forehead alignment mechanism that includes a forehead pad for restricting forward movement of the head of the patient,wherein the plurality of ear tubes are positionable such that an axis through the plurality of ear tubes can be parallel to a plane defined by the structure for horizontal alignment. 3. The device as defined in claim 1, wherein the portions of the outer boundary of the image volume respectively delineated by the boundary structure and the structure for horizontal alignment correspond to a front boundary of the image volume. 4. The device as defined in claim 1, wherein the boundary structure is a parabolically shaped vertical structure, and wherein the outer boundary of the image volume defines an outer boundary for the patient to be positioned within in order for a patient's condyles to appear within a tomographic image. 5. The device as defined in claim 1, wherein a front part of the outer boundary is spherical. 6. The device as defined in claim 1, wherein a front part of the outer boundary is cylindrical. 7. The device as defined in claim 1, wherein the structure for horizontal alignment is a horizontal structure. 8. The device as defined in claim 1, wherein the head clamp is adjustable. 9. The device as defined in claim 1, further comprising: a measuring unit adapted to measure a position of the head clamp;a calculating unit adapted to calculate a position of the patient based on a result obtained by the measuring unit; anda directing unit adapted to direct re-positioning of the patient based on a result obtained by the calculating unit. 10. The device as defined in claim 1, further comprising: a measuring unit adapted to measure a position of the hea... | SIRONA DENTAL SYSTEMS, GmbH | A61B005/05 A61B006/14 A61B006/03 A61B006/04
US09354226 | B2 | 2016-05-31 | Transdermal systems, devices, and methods to optically analyze an analyte | 2010-12-16 | The invention provides transdermal optical analysis systems, test sensors, methods, and kits for determining the presence and/or concentration of at least one analyte in a fluid sample. The system includes a transdermal test sensor including an aqueous material including at least one analyte selective reagent and at least one optically active moiety. The optical system preferably uses fluorescent spectroscopy to correlate fluorescent emission or adsorption from a dye with the analyte concentration of the sample. An optical light source and/or detector may be housed with the aqueous material in a housing or external to the housing of the aqueous material. | 1. A non-invasive method for determining a concentration of at least one analyte in a fluid, the method comprising: providing a transdermal test sensor including an aqueous material and a reagent system within the aqueous material, wherein the aqueous material is held between a top adhesive and a bottom adhesive within a volume defined by a semipermeable membrane and an additional membrane, the bottom adhesive adapted to adhere the test sensor to tissue, the test sensor further including an optical screen in contact with the additional membrane and the top adhesive and adapted to allow light communication with the aqueous material;contacting porated tissue with the semipermeable membrane of the transdermal test sensor;allowing sufficient time for the fluid sample to traverse to an external surface of the porated tissue and enter the aqueous material external to the porated tissue, the reagent system including at least one analyte selective reagent, at least one optically active moiety, and a reference dye;applying at least one wavelength of light to the aqueous material;detecting at least one wavelength of light from the aqueous material; andcorrelating an alteration of the applied light to the detected light with the concentration of the at least one analyte in the fluid sample. 2. The method of claim 1, wherein the at least one optically active moiety includes a pH sensitive dye. 3. The method of claim 1, wherein the analyte selective reagent includes at least one receptor and one or more species that the analyte displaces from the at least one receptor. 4. The method of claim 1, wherein the analyte selective reagent includes at least one analyte binding moiety that undergoes a conformation change when contacted by the analyte. 5. The method of claim 1, wherein the analyte selective reagent includes at least one analyte binding moiety that undergoes a change in electron withdrawing ability when contacted by the at least one analyte. 6. The method of claim 1, wherein the optically active moiety includes a donor fluorescent dye and an acceptor fluorescent dye. 7. The method of claim 6, wherein the concentration of the at least one analyte is correlated with a decrease in acceptor dye fluorescence in the presence of the at least one analyte as compared to the acceptor dye fluorescence in the absence of the at least one analyte. 8. The method of claim 1, wherein the optically active moiety includes a donor fluorescent dye and a quencher. 9. The method of claim 1, wherein the at least one analyte is selected from the group consisting of glucose, cholesterol, glutamate, lactate, and combinations thereof. 10. The method of claim 1, wherein the at least one analyte is glucose. 11. The method of claim 1, wherein two or more analytes are detected. 12. The method of claim 1, wherein the fluid sample is interstitial fluid. 13. The method of claim 1, wherein the correlating of the alteration of the applied light to the detected light with the co... | ASCENSIA DIABETES CARE HOLDINGS AG | A61B005/00 G01N033/52 A61B005/145 A61B005/1455 G01N033/542 G01N033/66 G01N033/92
US09339238 | B2 | 2016-05-17 | Systems and methods for processing sensor data | 2012-05-16 | Systems and methods for processing sensor data are provided. In some embodiments, systems and methods are provided for calibration of a continuous analyte sensor. In some embodiments, systems and methods are provided for classification of a level of noise on a sensor signal. In some embodiments, systems and methods are provided for determining a rate of change for analyte concentration based on a continuous sensor signal. In some embodiments, systems and methods for alerting or alarming a patient based on prediction of glucose concentration are provided. | 1. A computer-implemented method for classifying a level of noise in a signal obtained from a continuous glucose sensor based on a signal strength, the method comprising executing on a computing system the steps of: receiving a signal from a continuous glucose sensor during a sensor session;determining a level of noise in the signal;classifying the determined level of noise in the signal into one of a plurality of different predetermined noise classifications during the sensor session based a plurality of noise ranges; andcontrolling an output of information during the sensor session responsive to the noise classification, wherein the information is selected from the group consisting of rate of change information, alarm information, and medicament delivery information,wherein the controlling comprises executing instructions to request user interaction, using a user interface of the computing system, with an insulin delivery system based on the noise classification. 2. The method of claim 1, further comprising adaptively determining the one or more noise thresholds for classification of the level of noise in the signal during the sensor session. 3. The method of claim 2, further comprising determining a signal strength of the signal from the continuous glucose sensor during the sensor session. 4. The method of claim 3, wherein the one or more noise thresholds are based on the signal strength of the signal. 5. The method of claim 1, wherein the output provides an indication of the noise classification to an external device based on the noise classification. 6. The method of claim 1, wherein the output provides an indication of the noise classification on a user interface based on the noise classification. 7. The method of claim 1, wherein the controlling further comprises selectively controlling a display of rate of change information on a user interface based on the noise classification. 8. The method of claim 1, wherein the controlling further comprises selectively triggering an alarm indicative of at least one of hypoglycemia, hyperglycemia, predicted hypoglycemia, or predicted hyperglycemia responsive based on the noise classification. 9. The method of claim 1, wherein the controlling further comprises selectively sending medicament therapy instructions based on the noise classification. 10. The system of claim 1, wherein controlling further comprises controlling medicament delivery based on the noise classification. 11. The method of claim 10, wherein controlling medicament delivery comprises controlling insulin delivery based on the noise classification. 12. The method of claim 1, wherein controlling further comprises instructing a closed loop delivery system to open a loop thereof. 13. The method of claim 1, further comprising determining the level of noise in the signal prior to the classifying. 14. A system for classifying a level of noise in a signal obtained from a continuous glucose sensor, the system comprising: a continuo... | DexCom, Inc. | A61B005/00 A61B005/145 A61B005/1486 A61B005/1495 A61M005/172 A61B005/01 G06F019/00
US09326687 | B2 | 2016-05-03 | Methods for fabrication of substrates for surface enhanced Raman spectroscopy | 2013-08-05 | Methods for fabricating metal nano-particle embedded enhancement substrates used for surface enhanced Raman spectroscopy (SERS) including ion implanting metal nano-particles into the substrate and etching the substrate to partially expose the metal nano-particles. The resulting material is useful as a SERS substrate for detection of molecules adsorbed on it by surface enhanced Raman spectroscopy. | 1. A method for fabricating substrates for surface enhanced Raman spectroscopy (SERS) comprising the steps of: implanting negatively or positively charged metal ions into a matrix material, andgrowing near-surface or subsurface metal nano-particles within the matrix material. 2. The method of claim 1, further comprising the step of: prior to the growing step, thermal annealing the implanted matrix. 3. The method of claim 1, further comprising the step of: subsequently to the growing step, selective surface etching the surface of the matrix to expose or partially expose the metal nano-particles. 4. The method of claim 3, wherein the etching is selected from the group consisting of chemical etching, electrochemical etching, ion etching, reactive ion etching, mechanical chemical polishing, and mixtures or combinations thereof. 5. The method of claim 1, wherein the implanting metal is selected from the group consisting of metals from the periodic table of elements including alkali metals, alkaline earth metals, transition metals, lanthanide metals, actinide metals or mixtures and combinations thereof. 6. The method of claim 1, wherein the implanting metal is selected from the group consisting of ruthenium, rhodium, palladium, silver, osmium, iridium, platinum, gold, and mixtures or combinations thereof. 7. The method of claim 1, wherein the implanting metal is selected from the group consisting of copper, ruthenium, rhodium, palladium, silver, osmium, iridium, platinum, gold, and mixtures or combinations thereof. 8. The method of claim 1, wherein the implanting metal is selected from the group consisting of silver, gold, copper, and mixtures or combinations thereof. 9. The method of claim 1, wherein the matrix material is selected from the group consisting of a semiconducting material, an insulating material, a conducting substrate, and mixtures or combinations thereof. 10. The method of claim 9, wherein the insulating material is a surface of a tip of an optical fiber. 11. The method of claim 1, wherein the implanting dose of the metal in the matrix material exceeds the solubility limit of the metal in the matrix material. 12. The method of claim 1, wherein the implanting step is carried out at room temperature or at an elevated temperature. 13. The method of claim 1, wherein the implanting beam current density is sufficient to trigger growth of metal nano-particles without post implantation high-temperature annealing. 14. The method of claim 1, wherein the nano particles ion implanted in the substrate surface includes particles having a largest diameter ranging between about 0.01 nm and about 200 nm; having shapes that are spherical or near spherical in shape, and having an inter-particle separation of less than 50 nm. 15. The method of claim 1, wherein the nano particles ion implanted in the substrate surface includes particles having a largest diameter ranging between about 1.0 nm and about 10 nm; having shapes that are spherical or... |  | A61B005/00 G01N021/65
US09332912 | B2 | 2016-05-10 | Electronic sphygmomanometer for measuring blood pressure based on arterial volume change | 2012-09-25 | Every pulse wave, a constant volume control unit performs servo control while updating a servo gain such that a difference between an arterial volume shown by an arterial volume signal that is detected by an arterial volume detection circuit and a control target value of the servo control is less than a control deviation at which a rate of change of the arterial volume relative to a change in cuff pressure is deemed to be constant. Because the control deviation is the difference between the arterial volume shown by the arterial volume signal and the control target value, a blood pressure decision unit decides, as a blood pressure, the cuff pressure sequentially detected in a period during which servo control is performed, by correcting the cuff pressure using the control deviation and the rate of change deemed to be constant. | 1. An electronic sphygmomanometer comprising: a cuff that is configured to be placed on a blood pressure measurement site;a pressure detection unit that detects a cuff pressure representing a pressure inside the cuff;a volume detection unit provided in the cuff that detects an arterial volume signal showing an arterial volume of the blood pressure measurement site in a process of changing the cuff pressure, wherein the arterial volume signal changes depending on a pulse wave associated with each heart beat;a cuff pressure adjustment unit that adjusts the cuff pressure by pressurization and depressurization; anda blood pressure measurement unit,wherein the blood pressure measurement unit comprises: a control target value detection unit that detects, as a control target value, a value of the arterial volume signal when an amplitude of the arterial volume signal detected by the volume detection unit indicates a maximum;a rate-of-change detection unit that sequentially detects a rate of change in the arterial volume in the process of changing the cuff pressure, based on the arterial volume signal detected by the volume detection unit;a control deviation detection unit that detects, as a control deviation, a difference between the value of the arterial volume signal and the control target value;a servo control unit that performs servo control on the cuff pressure adjustment unit using a servo gain, such that the value of the arterial volume signal matches the control target value; anda blood pressure decision unit that decides, as a blood pressure, the cuff pressure sequentially detected by the pressure detection unit in a period during which the servo control is performed, by correcting the cuff pressure using the control deviation and a rate of change that is deemed to be constant, andwherein the servo control unit updates the servo gain every pulse wave, such that the control deviation, which is the difference between the value of the arterial volume signal detected by the volume detection unit and the control target value, indicates a value that is less than a control deviation target value, which is the control deviation detected by the control deviation detection unit in a period during which the rate of change sequentially detected by the rate-of-change detection unit is deemed to be constant. 2. The electronic sphygmomanometer according to claim 1, wherein the blood pressure decision unit calculates a correction value, by dividing the control deviation by the rate of change that is deemed to be constant, and corrects the cuff pressure by adding the correction value to the cuff pressure. 3. The electronic sphygmomanometer according to claim 1, wherein the period during which the rate of change is deemed to be constant denotes a period during which a rate of change that is greater than or equal to a prescribed threshold is detected with respect to rates of change detected by the rate-of-change detection unit when the maximum value of the ampli... | OMRON HEALTHCARE Co., Ltd. | A61B005/02 A61B005/021 A61B005/024 A61B005/00
US09326715 | B1 | 2016-05-03 | OxyVu-1 hyperspectral tissue oxygenation (HTO) measurement system | 2012-12-26 | The invention is directed to a hyperspectral/multispectral system referred to as the OxyVu-1 system. The hyperspectral imaging technology performs spectral analysis at each point in a two-dimensional scanned a producing an image displaying information derived from the analysis. For the OxyVu-1 system, the spectral analytical methods determined in superficial tissues approximate values of oxygen saturation (HT-Sat), oxyhemoglobin levels (HT-oxy), and deoxyhemoglobin levels (HT-deoxy). The OxyVu-1 system displays the tissue oxygenation in a two-dimensional, color-coded image. The system contains a system console, a cart, system electronics, CPU, monitor, keyboard, pointing device and printer. The hyperspectral instrument head with support arm contains broadband illuminator, camera and spectral filter for collecting hyperspectral imaging cube. The single use OxyVu Check Pads and Targets are used to perform an instrument check prior to patient measurements. The OxyVu Target is placed within the intended field of view and is used as a fiduciary mark for image registration and for focusing. | 1. A hyperspectral/multispectral medical imaging system, comprising: an illumination system that is configured to illuminate a subject;an optical acquisition system that is configured to acquire reflected light from a subject;a central processing unit (CPU);memory; anda plurality of program modules stored in the memory and configured to be executed by the CPU, the plurality of program modules including: a calibration module, anda data acquisition module,wherein the calibration module includes instructions for: (A) calibrating illumination power and uniformity of the illumination system by a procedure comprising: (i) acquiring a hyperspectral/multispectral image of a calibration pad comprising a surface having a predetermined reflectance, the hyperspectral/multispectral image comprising a plurality of sub-images of the calibration pad illuminated by the illumination system, each respective sub-image in the plurality of sub-images acquired at a corresponding spectral band;(ii) subtracting a dark radiation image from the hyperspectral/multispectral image, the dark radiation image acquired previously when the optical acquisition system was blocked from light;(iii) determining, for each respective sub-image of the hyperspectral/multispectral image, the measured value of light reflectance across the calibration pad at the corresponding spectral band of the respective sub-image;(iv) determining a maximum deviation, for each respective sub-image of the hyperspectral/multispectral image, between the measured value of light reflectance across the calibration pad at the corresponding spectral band of the respective sub-image and an ideal value of light reflectance across the calibration pad at the corresponding spectral band of the respective sub-image;(v) comparing the maximum deviation for each respective sub-image of the hyperspectral/multispectral image to a first threshold value; and(vi) determining a success of the calibrating (A) wherein the calibrating (A) is deemed to have failed when the maximum deviation for a sub-image of the first hyperspectral/multispectral image of exceeds the first threshold value, andthe calibrating (A) is deemed to have passed when the maximum deviation for a sub-image of the hyperspectral/multispectral image does not exceed the first threshold value;(B) calibrating wavelength accuracy of the optical acquisition system; and(C) calibrating image focus of the optical acquisition system. 2. The hyperspectral/multispectral medical imaging system of claim 1, wherein the calibration pad comprises a colored region and wherein the instructions for calibrating wavelength accuracy of the optical acquisition system (B) include: subtracting a dark radiation image from the hyperspectral/multispectral image, the dark radiation image acquired previously when the optical acquisition system was blocked from light;determining, for each respective sub-image of the hyperspectral/multispectral image, the measured value of light reflectance ac... | Hypermed Imaging, Inc. | A61B005/1455 A61B005/1495
US09339225 | B2 | 2016-05-17 | Systems and methods for assessing sweat gland output | 2013-03-15 | Various method embodiments for performing a sweat test include triggering sweating in a test area of a test subject, drying the test area, and applying a test device on the test area. The test device is used to monitor sweat production of individual SGs within the test area to provide test results for each functional SG in the test area. The test results for each functional SG in the test area are recorded. | 1. A method for performing a sweat test; comprising: triggering sweating in a test area of a test subject;drying the test area, and applying a test device on the test area, wherein applying the test device includes applying a nano grid of nanowires on the test area;monitoring, using the test device, sweat production of individual sweat glands (SGs) within the test area to provide test results for each functional SG in the test area, including measuring a change in electrical characteristics of the nano grid of nanowires, wherein: the nano grid includes intersections of nanowires and includes air pockets configured to expose nanowires at each of the intersections, and the air pockets are configured to allow sweat water to enter and contact the nanowires; andmeasuring a change in electrical characteristics of the nano grid of nanowires includes for each of the intersections detecting whether an impedance between the nanowires has dropped significantly, indicating that sweat water has entered the pocket and is in contact with the nanowires at the intersection; andrecording the test results for each functional SG in the test area. 2. The method of claim 1, wherein triggering sweating in the test area includes introducing pilocarpine into dermis of the test subject using iontophoresis. 3. The method of claim 1, wherein: applying the test device further includes applying sensors on the test area with the nano grid of nanowires, wherein the sensors are configured to sense sweat content, the method further comprising evaluating sweat content using the sensors. 4. The method of claim 3, wherein applying sensors includes applying chemical field effect transistors (ChemFETs). 5. A system for performing a sweat test on a test subject, comprising: a test device configured to be applied to a test area on the test subject, wherein the test device is configured to detect sweat production of individual sweat glands (SGs) within the test area, wherein the test device includes a nano grid of nanoconductors and circuitry configured to test conductivity of the nano grid to detect sweat production of individual SGs in the test area, wherein: the nano grid of nanoconductors includes a first level of parallel nanoconductors and a second level of parallel nanoconductors, the first and second levels of nanoconductors crossing at a plurality of intersections, the first and second levels of nanoconductors generally separated by an insulator, the insulator including an air pocket at the intersections, each air pocket exposing one nanoconductor in the first level and one nanoconductor in the second level; andthe circuitry includes circuitry configured to test electrical impedance between two nanoconductors at each of the air pockets; anda processing system configured to analyze the detected sweat production, provide test results for each functional SG in the test area, and record the test results for each functional SG in the test area. 6. The system of claim 5, wherein ... | Daniel L. M Kennedy | A61B005/00
US09332915 | B2 | 2016-05-10 | System and method to identify sources associated with biological rhythm disorders | 2013-03-15 | In a system and method for identifying a driver of a source associated with a heart rhythm disorder, data are accessed from a plurality of sensors representing biological activity in the heart. A first region and a second region of the heart are identified as the source of the heart rhythm disorder. If the first region has repeating activation and controls the second region for at least a predetermined number of beats, the first region is identified as controlling the source of the heart rhythm disorder. | 1. A method of identifying a region associated with controlling a source of a heart rhythm disorder, the method comprising: accessing, by a computing device, data from a plurality of sensors representing biological activity in the heart;identifying, by the computing device, a first region and a second region of the heart comprising the source of the heart rhythm disorder, wherein the first region of the heart has repeating activation;determining, by the computing device, whether the first region controls the second region of the heart for at least a predetermined number of beats; andidentifying the first region as controlling the source when the first region controls the second region for at least the predetermined number of beats. 2. The method of claim 1, wherein a sequence of activation within the first region directly controls the second region. 3. The method of claim 1, wherein a sequence of activation within the first region directly controls the second region in a centrifugal pattern. 4. The method of claim 1, wherein a sequence of activation within the first region is rotational. 5. The method of claim 1, wherein a sequence of activation within the first region is rotational and includes outward spread to the second region. 6. The method of claim 1, wherein a sequence of activation is spatially located within the first region between the beats. 7. The method of claim 1, wherein a sequence of activation spatially moves within the first region between the beats. 8. The method of claim 1, wherein the first region has an area of less than 5 cm2. 9. The method of claim 1, further comprising determining stability of the repeating activation by computing indexes of centrifugal activation within the first region. 10. The method of claim 9, wherein the indexes of centrifugal activation is one or more of vectors of activation, linear velocity, correlation analysis or and spatial maps showing progressive spatial activation. 11. The method of claim 9, wherein the indexes of centrifugal activation are displayed on a map of sensor locations indicating higher and lower indexes of centrifugal activation. 12. The method of claim 1, further comprising determining persistence of the repeating activation by computing indexes of rotational activation within the first region. 13. The method of claim 12, wherein the indexes of rotational activation is one or more of progression over time in angles of rotational arc of activation sequence, sectors of rotation, traversal of a rotational perimeter, angular velocity and phase mapping. 14. The method of claim 13, wherein the indexes of rotational activation are displayed on a map of sensor locations indicating higher and lower indexes of rotational activation. 15. The method of claim 1, further comprising classifying the first region as a driver if the first region controls the second region for a plurality of the beats. 16. The method of claim 1, wherein the first region is a driver of the source. 1... | The Regents of the University of California Topera, Inc. The United States of America as Represented by the Department of Veterans Affairs, Office of the General Counsel | A61B005/024 A61B005/00 A61B005/042
US09327092 | B2 | 2016-05-03 | Ventilation mask with integrated piloted exhalation valve | 2013-03-04 | A mask for achieving positive pressure mechanical ventilation (inclusive of CPAP, ventilator support, critical care ventilation, emergency applications), and a method for a operating a ventilation system including such mask. The mask includes a piloted exhalation valve that is used to achieve the target pressures/flows to the patient. The pilot for the valve may be pneumatic and driven from the gas supply tubing from the ventilator. The pilot may also be a preset pressure derived in the mask, a separate pneumatic line from the ventilator, or an electro-mechanical control. The mask of the present invention may further include a heat and moisture exchanger (HME) which is integrated therein. | 1. A ventilation mask, comprising: a housing sized and configured to be positionable between a patient's nose and mouth in contact with the patient's nose, the housing defining an internal fluid chamber, at least one gas delivery lumen which fluidly communicates with the fluid chamber, and a valve pilot lumen; anda piloted exhalation valve attached to the housing and at least partially residing within the fluid chamber thereof, the exhalation valve comprising: a valve housing defining a valve chamber which fluidly communicates with the fluid chamber, the pilot lumen and ambient air; anda diaphragm disposed within the valve chamber, the diaphragm normally being biased to an open position wherein the fluid chamber fluidly communicates with ambient air via the valve chamber, and a closed position wherein fluid communication between the fluid chamber and ambient air is blocked thereby, the movement of the diaphragm from the open position to the closed position being facilitated by the selective pressurization of at least a portion of the valve chamber via the pilot lumen. 2. The ventilation mask of claim 1 wherein the housing further defines a pressure sensing lumen which fluidly communicates with the fluid chamber. 3. The ventilation mask of claim 1 wherein the housing is a resilient cushion including a spaced pair of hollow pillow portions which each fluidly communicate with the fluid chamber and are adapted to engage respective ones of the nostrils of the patient's nose. 4. The ventilation mask of claim 3 further comprising a heat and moisture exchanger disposed within the fluid chamber between the exhalation valve and the pillow portions of the cushion. 5. The ventilation mask of claim 3 wherein the valve housing comprises: a seat member; anda cap member attached to the seat member, the seat and cap members collectively defining the valve chamber;the diaphragm being captured between the seat and cap members in a manner wherein the diaphragm effectively segregates the valve chamber into a pilot section which fluidly communicates with the pilot lumen and an exhaust section which fluidly communicates with the fluid chamber in ambient air. 6. The ventilation mask of claim 5 wherein the diaphragm and the seat member define complimentary seating surfaces which are sized and configured relative to each other such that the movement of the diaphragm to the closed position facilitates the placement of the seating surfaces into sealed engagement with each other in a manner blocking fluid communication between the fluid chamber and ambient air. 7. The ventilation mask of claim 6 wherein the seat member includes a plurality of vents which are disposed therein and collectively define a fluid conduit between the fluid chamber and the exhaust section of the valve chamber. 8. The ventilation mask of claim 7 wherein the vents of the seat member reside within the fluid chamber of the cushion. 9. The ventilation mask of claim 8 further comprising a shield pl... | Breathe Technologies, Inc. | A61M011/00 A61M016/06 A61M016/20 A61M016/00 A61B005/087 A61B005/00 A61B005/097 A61B005/03 A61M016/08 A61M016/10
US09347894 | B2 | 2016-05-24 | Methods and systems for producing visible light and x-ray image data | 2011-08-31 | An imaging system may include a main enclosure having at least one access door that defines a substantially light-tight imaging compartment when the access door is in a closed position. An object platform provided within the main enclosure holds an object to be imaged. A camera system positioned on a first side of the object platform is operable to capture a visible light image of the object. A radiation detector positioned on the first side of the object platform is moveable from a first position to a second position across a field of view of the camera system. The radiation detector is operable to capture a radiographic image of the object by moving the radiation detector from the first position to the second position while detecting radiation from the object. | 1. An imaging system, comprising: a main enclosure having at least one access door, said main enclosure defining a substantially light-tight imaging compartment when the access door is in a closed position;an object platform provided within said main enclosure for holding an object to be imaged;a camera positioned on a first side of said object platform, said camera captures a visible light image of the object; anda radiation detector positioned on the first side of said object platform, said radiation detector being moveable across a field of view of said camera from a first position to a second position, said radiation detector captures a radiographic image of the object by moving said radiation detector from the first position to the second position while operating said detector to detect radiation from said object. 2. The imaging system of claim 1, wherein said radiation detector comprises an x-ray detector and wherein said imaging system further comprises an x-ray source positioned on a second side of said object platform, said x-ray source being operable to direct x-rays toward the second side of said object platform, said x-ray detector being operable to capture an x-ray image of the object by moving said x-ray detector from the first position to the second position while operating said x-ray source to direct x-rays toward the second side of said object platform, x-rays from said x-ray source passing through said object platform and the object positioned on said object platform before being incident on said x-ray detector. 3. The imaging system of claim 2, further comprising an image processing system operatively associated with said camera and said x-ray detector, said image processing system processing image data produced by said camera and said x-ray detector to produce a composite image, said composite image comprising portions of image data produced by said camera and image data produced by said x-ray detector. 4. The imaging system of claim 3, wherein the image data produced by said camera and said x-ray detector comprise electronic image data. 5. The imaging system of claim 1, further comprising a gantry assembly mounted to said imaging system and wherein said radiation detector is mounted to said gantry assembly so that said radiation detector is moveable along said gantry assembly in a longitudinal direction. 6. The imaging system of claim 5, wherein said gantry assembly further comprises: a first guide rail mounted to said imaging system so that said first guide rail extends along the longitudinal direction; anda second guide rail mounted to said imaging system, said second guide rail being in generally parallel, spaced-apart relation to said first guide rail. 7. The imaging system of claim 6, further comprising an actuator system operatively associated with said gantry assembly and said radiation detector, said actuator system operable to move said radiation detector along said gantry assembly from the first position to th... | Spectral Instruments Imaging, LLC | G01N023/04 A61B005/00 A61B006/00 G01N021/64 G01N021/76
US10182731 | B2 | 2019-01-22 | Medical electrode and limb clamp for an ECG device | 2014-05-30 | The present invention provides a medical electrode comprising a conductive metal base comprising a plate element and a boss formed on the plate element and a conductive support cylinder separate from the conductive metal base. The conductive support cylinder is rotatably mounted to the conductive metal base while remaining in electrical communication with said conductive metal base. The present invention also provides a limb clamp for an ECG device. According to the present invention, it is possible to prevent bending of the cable connecting with the medical electrode, thereby avoiding cable failure. | 1. A medical electrode comprising: a conductive metal base comprising a plate element and a boss formed on the plate element; anda conductive support cylinder separate from the conductive metal base;wherein the conductive support cylinder is mounted to the conductive metal base and is rotatable relative to the conductive metal base while remaining in electrical communication with the conductive metal base,wherein a recess is formed in one of the boss and a lower end of the conductive support cylinder; anda conductive bearing disposed within the recess,wherein the other one of the boss and the lower end of the conductive support cylinder is mounted into the conductive bearing and rotatable relative to the recess. 2. The medical electrode according to claim 1, further comprising an electrolyte gel applied to a lower surface of the plate element. 3. The medical electrode according to claim 1, wherein the conductive support cylinder comprises a through hole transverse to its longitudinal axis for receiving a cable. 4. The medical electrode according to claim 3, wherein the conductive support cylinder further comprises a threaded hole formed along its longitudinal axis and communicating with the through hole. 5. A limb clamp for an ECG device comprising: a first clamp portion;a second clamp portion;a spring element connecting the first clamp portion and the second clamp portion; anda medical electrode according to claim 1 mounted at the first clamp portion. 6. The limb clamp for an ECG device according to claim 5, wherein an outer thread is formed on the outer surface of the boss, the boss passes through a through hole formed in the first clamp portion and a nut is screwed onto the outer thread so that the medical electrode is mounted onto the first clamp portion. | KONINKLIJKE PHILIPS N.V. | A61B005/0408 A61B005/00 A61B005/0424 A61N001/04 H01R004/48 H01R004/56 A61B005/0402 A61B005/0428
US09591983 | B2 | 2017-03-14 | Wearable fetal monitoring system having textile electrodes | 2014-06-01 | A seamless, smart fetal monitoring garment and methods of using thereof. The system includes a knitted or interwoven garment having a multiplicity of conductive textile electrodes for sensing maternal and fetal electrical vital signals. The maternal and fetal electrical vital signals are selected from a group including maternal heart rate, fetal heart rate and electromyogram (EMG) activities including uterine activities. The method includes wearing the garment, acquiring electrical mixed common, maternal and fetal vital signals from surface region of a pregnant woman, using the plurality of textile electrodes, optimally weighted summing-up the acquired signals, analyzing the summed-up signals to thereby extract the maternal signal and the fetal signal, including determining their heart rates, and including detecting health hazards and in some embodiments, including detecting a uterine contraction sequence suggesting the need to be hospitalized for birth giving. | 1. A method for maternal and fetal monitoring comprising the steps of: a) wearing a knitted or interwoven smart maternal garment having a plurality of textile electrodes integrally knitted or interwoven therein, said textile electrodes being in communication flow with a processor;b) acquiring electrical mixed common, maternal and fetal electrical vital signals from a predetermined external surface region of a pregnant woman, using a plurality of textile electrodes integrally knitted or interwoven into a maternal garment;c) optimally weighted summing-up said acquired mixed maternal and fetal electrical vital signals to thereby form a summed-up maternal signal, having a substantially higher SNR than either of the acquired maternal and fetal electrical vital signals;d) analyzing the summed-up maternal signal to thereby extract the maternal signal from said summed-up maternal signal; ande) healing the acquired mixed maternal and fetal electrical vital signals, including deleting the extracted maternal signal from the respective acquired mixed maternal and fetal electrical vital signal, to thereby form a plurality of healed-maternal ECG signals;f) optimally weighted summing-up said healed-maternal ECG signals to thereby form a summed-up coherent fetal signal, having a substantially higher SNR than either of the healed-maternal ECG signals;g) analyzing said summed-up coherent fetal signal to thereby extract a fetal signal from said summed-up coherent fetal signal. 2. The method of claim 1 further comprising the step of: h) analyzing said summed-up coherent fetal signal thereby extract the EMG signal formed by electromyogram (EMG) activities including uterine activities, from said summed-up coherent fetal signal. 3. The method of claim 2, wherein said extracting and analyzing of said EMG signal from said summed-up electrical signal comprises the steps of: a) detecting the peaks in the maternal QRS complexes in said summed-up maternal signal;b) determining the boundary of each said detected maternal QRS complex;c) deleting each said detected maternal QRS complex from the respective acquired mixed maternal and fetal electrical vital signal and filling the gap formed therein, thereby forming a respective healed-maternal ECG signal;d) detecting the peaks of fetal QRS complexes in said summed-up coherent fetal signal;e) determining the boundary of each said detected fetal QRS complex;f) deleting each said detected fetal QRS complex from the summed-up coherent fetal signal and filling the gap formed therein, thereby forming a EMG summed-up signal; andg) analyzing the EMG summed-up signal. 4. The method of claim 3, wherein said analysis of said EMG summed-up signal includes performing a morphologically analysis to thereby determine if the uterine is contracting at a rate and amplitude that requires hospitalization for birth giving. 5. The method of claim 3, wherein said filling of said gap includes by using linear interpolation or spline. 6. The method of... | HEALTHWATCH LTD. | A61B005/0205 A61B005/08 A61B005/145 A61B005/0448 A41D001/20 A61B005/024 A61B005/0452 A61B005/0488 A61B005/01 A61B005/021 A61B005/0472 A61B005/11 A61B005/00
US10292592 | B2 | 2019-05-21 | Method and apparatus for optical recording of biological parameters in freely moving animals | 2015-11-12 | A system and method for monitoring biological parameters in freely moving animals that, in the illustrative embodiment, comprises a two-color optical measurement/recording system that is combined with a fluorescent protein reporter of cellular membrane potentials and a fluorescent protein that is insensitive to such cellular membrane potentials. The two wavelengths are used to un-mix physiological artifacts in recordings that occur in freely moving animals. | 1. A method comprising: delivering, to a target tissue, excitation light having a first wavelength and a second wavelength, wherein the target tissue contains a sensor and a reference;collecting optical domain responses from the sensor and the reference, the response from the sensor includes information pertaining to a biological parameter of interest as well as information pertaining to physiological artifacts and the response from the reference includes information pertaining to the physiological artifacts but not the biological parameter; andextracting the information pertaining to the biological parameter free of physiological artifacts. 2. The method of claim 1 wherein extracting a signal representative of a biological parameter further comprises unmixing the physiological artifacts from the response of the sensor using information contained in the response of the reference. 3. The method of claim 1 wherein delivering excitation light further comprises generating first light having the first wavelength from a first light source and generating second light having the second wavelength from a second light source. 4. The method of claim 3 and further comprising modulating the first light via a first modulation frequency. 5. The method of claim 4 and further comprising modulating the second light via the first modulation frequency and shifting phase of the second light relative to the first light. 6. The method of claim 3 and further comprising measuring the intensity of the first light and the second light. 7. The method of claim 1 wherein delivering excitation light to target tissue further comprises propagating the excitation light to a waveguide that is in optical communication with the target tissue. 8. The method of claim 7 wherein collecting optical domain responses from the sensor and the reference further comprises receiving the optical domain responses in the waveguide. 9. The method of claim 1 and further comprising separating the optical domain responses from the sensor and the reference. 10. The method of claim 1 and further comprising converting the optical domain responses from the sensor and the reference into electrical domain analogues thereof. 11. The method of claim 1 wherein extracting the information further comprises: generating an estimate of hemodynamic artifacts present in the response from the sensor; andgenerating a linear unmixing matrix for sorting the biological parameter of interest from the physiological artifacts using the estimate of hemodynamic artifacts. 12. The method of claim 11 wherein the linear unmixing matrix is generated using independent component analysis. 13. The method of claim 1 wherein the sensor is a first fluorescent protein and the reference is a second fluorescent protein. 14. The method of claim 1 wherein the biological parameter is membrane voltage of a neuron. 15. The method of claim 1 and further comprising delivering the sensor and the reference to the target tissue. 16. A... | The Board of Trustees of the Leland Stanford Junior University | A61B005/00
US11344217 | B2 | 2022-05-31 | NMR sensor and methods for rapid, non-invasive determination of hydration state or vascular volume of a subject | 2014-09-05 | The invention features methods for detecting the hydration state or vascular volume of a subject using a device capable of nuclear magnetic resonance (NMR) measurement. The methods involve exposing a portion of a tissue of the subject in vivo to a magnetic field and RF pulse from the device to excite hydrogen nuclei of water within the tissue portion, and measuring a relaxation parameter of the hydrogen nuclei in the tissue portion, the relaxation parameter being a quantitative measure of the hydration state or vascular volume of the subject as a whole. The invention also features devices and computer-readable storage media for performing the methods of the invention. | 1. A method of treating systemic dehydration or systemic overhydration in a subject, the method comprising (a) exposing a tissue portion comprising a lean tissue of the subject in vivo to a magnetic field and RF pulse from a nuclear magnetic resonance device to excite hydrogen nuclei of water within the tissue portion comprising an extracellular water compartment,(b) in the absence of a contrast agent, measuring a T2 relaxation time of the hydrogen nuclei in the tissue portion using the nuclear magnetic resonance device to produce T2 relaxation time data,(c) deconvoluting the T2 relaxation time data using multi-exponential regression analysis to produce a signal amplitude for the extracellular water compartment and determining the subject as (i) having systemic dehydration if the signal amplitude is decreased relative to a reference signal amplitude or (ii) having systemic overhydration if the signal amplitude is increased relative to the reference signal amplitude, wherein steps (a), (b), and (c) are non-invasive, and wherein the reference signal amplitude is produced by a reference measurement in the absence of a contrast agent comprising one or more prior measurements of the T2 relaxation time of hydrogen nuclei of water in the tissue portion of the subject in the state of systemic euhydration, and(d) treating systemic dehydration in the subject determined as having systemic dehydration or treating systemic overhydration in the subject determined as having systemic overhydration. 2. The method of claim 1, wherein the steps (a) and (b) are repeated two or more times. 3. The method of claim 1, wherein the subject has a disease or condition that increases the risk of having, or results from, dehydration or overhydration. 4. The method of claim 3, wherein the subject has a disease or condition selected from the group consisting of congestive heart failure (CHF), renal failure, liver cirrhosis, nephrotic syndrome, brain swelling, diabetes, staphylococcal infection, nephrolithiasis, diarrhea, colitis, preferably ulcerative colitis, pyelonephritis, cystic fibrosis, Huntington's disease, rotavirus infection, herpangina, salmonellosis, norovirus infection, pertussis, cryptosporidium infection, cholera, coma, and water intoxication; wherein the subject has, or is suspected of having, congestion or hemodilution: wherein the subject has at least one symptom of congestion selected from the group consisting of dyspnea, orthopnea, exercise intolerance, pathologic S3 or S4 heart sounds, tachycardia, tachypenia, jugular venous distention, peripheral edema, ascites, and increased filling pressures of the heart. 5. The method of claim 3, wherein the subject has congestive heart failure, liver failure, renal failure, Cushing's syndrome, or pulmonary congestion. 6. The method of claim 1, wherein the tissue portion is a peripheral body part. 7. The method of claim 1, wherein the method is performed on a tissue volume of at least 0.01 cm3. 8. The method of cl... | Massachusetts Institute of Technology | A61B005/055 G01R033/50 G01R033/383 A61B005/00 A61B005/02 G01R033/44 G01R033/46 A61B005/145
US09687145 | B2 | 2017-06-27 | Optical coherence elastography to assess biomechanics and detect progression of ocular and other tissues degenerative diseases | 2015-11-06 | An excitation force (internal or external) and phase-sensitive optical coherence elastography (OCE) system, used in conjunction with a data analyzing algorithm, is capable of measuring and quantifying biomechanical parameters of tissues in situ and in vivo. The method was approbated and demonstrated on an example of the system that combines a pulsed ultrasound system capable of producing an acoustic radiation force on the crystalline lens surface and a phase-sensitive optical coherence tomography (OCT) system for measuring the lens displacement caused by the acoustic radiation force. The method allows noninvasive and nondestructive quantification of tissue mechanical properties. The noninvasive measurement method also utilizes phase-stabilized swept source optical coherence elastography (PhS-SSOCE) to distinguish between tissue stiffness, such as that attributable to disease, and effects on measured stiffness that result from external factors, such as pressure applied to the tissue. Preferably, the method is used to detect tissue stiffness and to evaluate the presence of its stiffness even if it is affected by other factors such as intraocular pressure (IOP) in the case of cornea, sclera, or the lens. This noninvasive method can evaluate the biomechanical properties of the tissues in vivo for detecting the onset and progression of degenerative or other diseases (such as keratoconus). | 1. A method for measuring tissue stiffness and for differentiating tissue samples using optical coherence elastography, comprising: inducing elastic waves in the tissue samples;detecting properties of the waves using interferometry, low coherence interferometry, or optical coherence tomography at different measurement positions along the waves, wherein the detected properties include measured wave velocities and measured wave displacement amplitudes and/or using temporal analysis of the elastic wave displacement profile;determining elasticities of the tissue samples using the measured wave velocities dispersion, and attenuation;differentiating the tissue samples having different measured wave velocities;normalizing the measured wave displacement amplitudes for the tissue samples having similar measured wave velocities and needing further differentiation to produce normalized wave displacement data;using the normalized wave displacement data to identify tissue samples having faster wave attenuation and slower wave attenuation; andclassifying the tissue samples having faster wave attenuation as tissue sample having increased viscosity and reduced stiffness and the tissue samples having slower wave attenuation as tissue samples having reduced viscosity and increased stiffness. 2. The method of claim 1, wherein the tissue samples are ocular or any other soft or hard tissue samples. 3. The method of claim 1, wherein the step of inducing elastic waves is by directing a focused air-pulse on the tissue samples. 4. The method of claim 1, wherein the step of determining elasticities of the tissue samples is by calculating Young's modulus using the measured wave velocities. 5. The method of claim 1, wherein the step of normalizing the measured wave displacement amplitudes is by dividing the measured wave displacement amplitudes at the different measurement positions along the waves by the measured wave displacement amplitude at an excitation position. 6. The method of claim 1, wherein the step of using the normalized wave displacement data to identify tissue samples having faster wave attenuation and slower wave attenuation is by using a customized ratio having a formula of: r(ND1/ND2)=mean(ri)±std(ri),whereinri=ND1⁢⁢iND2⁢⁢iand ND1i and ND2i are normalized displacement data at an ith different measurement position for a first and a second tissue sample, wherein if r(ND1/ND2) is significantly greater than 1, the second tissue sample is identified as having faster wave attenuation and the first tissue sample is identified as having slower wave attenuation, andwherein if r is significantly less than 1, the first tissue sample is identified as having faster wave attenuation and the second tissue sample is identified as having slower wave attenuation. 7. The method of claim 6, wherein the step of using the normalized wave displacement data to identify tissue samples having faster wave attenuation and slower wave attenuation is repeated for different tissue... | UNIVERSITY OF HOUSTON SYSTEM | A61B003/10 A61B003/00 A61B003/16 A61B005/00
US09339204 | B2 | 2016-05-17 | Systems and methods for detecting ECG subwaveforms | 2015-03-19 | Systems and methods are provided to detect subwaveforms of an ECG waveform. Electrical impulses are detected between at least one pair of electrodes of two or more electrodes placed proximate to a beating heart and are converted to an ECG waveform for each heartbeat of the beating heart using the detector. One or more subwaveforms within P, Q, R, S, T, U, and J waveforms of the ECG waveform for each heartbeat or in an interval between the P, Q, R, S, T, U, and J waveforms that represent the depolarization or repolarization of an anatomically distinct portion of muscle tissue of the beating heart are detected using a signal processor. A processed ECG waveform that includes the one or more subwaveforms for each heartbeat is produced using the signal processor. The processed ECG waveform is received from the signal processor is displayed using a display device. | 1. An electrocardiography (ECG) system for detecting five ECG subwaveforms within the PR interval using electrodes configured to be placed on the skin of a patient, comprising: two or more electrodes configured to be placed on the skin of a patient, configured to partially surround a beating heart of the patient, and configured to receive electrical impulses from the beating heart;a detector that detects the electrical impulses from at least one pair of electrodes of the two or more electrodes and converts the electrical impulses to an ECG waveform for each heartbeat of the beating heart;a signal processor that receives the ECG waveform for each heartbeat from the detector, individually detects five ECG subwaveforms within the PR interval of the ECG waveform that represent the depolarization of the sinoatrial node (SAN), the atria (right atrium (RA) and left atrium (LA)), the atrioventricular node (AVN), the bundle of His (HIS), and the bundle branches (BB), respectively, of the beating heart using five predetermined frequency ranges experimentally found for the SAN, the RA and LA, the AVN, and the BB, combines the five ECG subwaveforms, and replaces the PR interval of the ECG waveform with the combined five ECG subwaveforms, producing a processed ECG waveform that includes the combined five individual ECG subwaveforms in place of the PR interval for each heartbeat; anda display device that displays the processed ECG waveform for each heartbeat. 2. The ECG system of claim 1, wherein the display device further displays the ECG waveform that includes the PR interval beneath the the processed ECG waveform that includes the combined five individual subwaveforms in place of the PR interval using the same time scale and displays lines from one or more subwaveforms of the five individual subwaveforms of the processed ECG waveform to the original PR interval of the ECG waveform to indicate the timing location of the one or more subwaveforms of the five individual subwaveforms in the PR interval of the ECG waveform. 3. The ECG system of claim 1, wherein the signal processor further calculates timing information about the five ECG subwaveforms, timing information about intervals between the five ECG subwaveforms, and timing information about the five ECG subwaveforms and their relation to P, Q, R, S, T, U, and J waveforms of the ECG waveform for each heartbeat, andthe display device further displays the timing information about the five ECG subwaveforms, the timing information about the intervals between the five ECG subwaveforms, and the timing information about the five ECG subwaveforms and their relation to the P, Q, R, S, T, U, and J waveforms of the ECG waveform for each heartbeat. 4. The ECG system of claim 1, further comprising a memory device that receives the ECG waveform and the processed ECG waveform from the signal processor. 5. The ECG system of claim 4, wherein the memory device further includes normal processed ECG waveform data for a no... | Guangren Chen | A61B005/0452 A61B005/04 A61B005/044
US09339216 | B2 | 2016-05-17 | Systems and methods for the screening and monitoring of inner ear function | 2013-03-15 | Methods and systems for monitoring or testing the hearing of a user using a portable unit include presenting a plurality of test frequencies for selection by a user and storing data about those frequencies, and the user's response to stimuli presented at those frequencies, and comparing the stored data to previously acquired hearing data. Using mathematical prediction algorithms, changes in baseline hearing values can be used to identify potential hearing issues, especially hearing issues associated with ototoxicity. | 1. A method for monitoring or testing the hearing of a user, the method comprising: presenting a plurality of test frequencies for selection by a user through a display or acoustically;receiving indications of respective selections of test frequencies by the user, including indications of a right ear or a left ear that is to be tested at the respective selected test frequencies;causing an audio test unit to output one or more test sounds at respective ones of the selected test frequencies to the indicated ear, the one or more test sounds being output at one or more stimulus levels;receiving an indication of whether the test sound was heard by the indicated ear at the selected test frequencies and, for instances where the test sound was heard, storing the stimulus level for which the test sound was heard;storing the highest set of test frequencies for which an indication was received that the test sound was heard by the indicated ear, the highest set of test frequencies for each ear being the sensitive range for ototoxicity for that respective ear;storing a baseline value for each frequency in the sensitive range for ototoxicity for each ear, the baseline value being based on the stimulus level at which the indication that the test sound was heard was received;monitoring the sensitive range for ototoxicity for each ear by performing one or more subsequent hearing tests and comparing the baseline values for the respective sensitive range for ototoxicity with stimulus levels obtained from additional hearing tests; andcalculating whether a change in the baseline values has occurred in the one or more subsequent hearing tests. 2. The method of claim 1, wherein the act of monitoring the sensitive range for ototoxicity comprises: causing an audio test unit to output additional test sounds at the test frequencies of the sensitive range of ototoxicity for each ear at one or more stimulus levels;receiving indications of whether the additional test sounds were heard by the indicated ear and, for instances where the additional test sounds were heard, storing the stimulus level for which the additional test sounds were heard;storing the highest set of test frequencies for which indications were received that the additional test sounds were heard by the indicated ear, the highest set of test frequencies for each ear being the current sensitive range for ototoxicity for that respective ear;storing a current baseline value for each frequency in the current sensitive range for ototoxicity for each ear, the current baseline value being based for the on the stimulus level at which the indications that the additional test sounds were heard was received;calculating whether a clinically relevant change in hearing has occurred; andindicating any changes between the baseline values and current baseline values for the sensitive range for ototoxicity and/or between the sensitive range for ototoxicity and the current sensitive range for ototoxicity. 3. The method of clai... | The United States of America As Represented By The Department Of Veterans Affairs | A61B005/00 A61B005/12
US09326703 | B2 | 2016-05-03 | Therapeutic manipulation device | 2011-03-10 | A manipulation device is programmable with different parameters associated with different manipulation exercises for different types of simulated patients. A measuring system determines the success or failure for any of the programmable manipulation exercises and tracks specific results that precisely track how close a student comes to specific target force values. This allows a student to more accurately track performance improvements and more effectively focus practice sessions on problem manipulation techniques. The measurement system not only measures the time and force values associated with the manipulation exercises but also measures angular displacement of the manipulation apparatus during the manipulation exercise. This allows the student to not only determine if a proper amount of force was applied during the manipulation exercise but also to determine if the force was maintained in the same direction and angle. | 1. An apparatus comprising: a housing including an elongate body and an elongate reciprocating body adapted to fit into the cavity of the elongate body;a biasing member adapted to fit within the axial cavity of the elongate body and further adapted to bias the elongate reciprocating body to a first position relative to the axis of the cavity;a sensing device disposed in the elongate body comprising an image sensor and a plurality of photo interrupters, the image sensor configured to capture a plurality of images of the biasing member in response to movement of the biasing member about the axis of the cavity and each of the plurality of photo interrupters configured to generate a corresponding interrupter signal in response to movement of the biasing member about the axis of the cavity; anda processor configured to quantify a therapeutic manipulation force applied during therapeutic manipulation based on the plurality of images and on the interrupter signals. 2. The apparatus according to claim 1 wherein the sensing device further comprises a gyroscope configured to generate an X-axis signal that identifies a rate of angular displacement along an X-axis of the housing. 3. The apparatus according to claim 2 wherein the sensing device is configured to generate a Z-axis signal that identifies a rate of angular displacement along a Z-axis of the housing. 4. The apparatus according to claim 3 wherein the processor determines an angular displacement rate of the apparatus during application of the therapeutic manipulation force according to both the X-axis signal and the Z-axis signal. 5. The apparatus according to claim 3 wherein the processor is configured to multiply the angular displacement rate by a predetermined time period to identify an incremental displacement angle and determine a maximum displacement of the housing during the therapeutic manipulation force by accumulating incremental displacement angles derived for previous time periods. 6. The apparatus according to claim 2 wherein the sensing device is configured to generate a Z-axis signal that identifies a rate of Y-axis angular displacement of the housing. 7. The apparatus according to claim 6 wherein the processor determines an amount of torque applied to the apparatus during application of the therapeutic manipulation force according to the Y-axis signal. | ORTHO-NEURO TECHNOLOGIES, INC. | A61B005/103 A61H001/00 G09B023/28 A61B005/00
US09339224 | B2 | 2016-05-17 | Event dosimeter devices and methods thereof | 2012-02-10 | A dosimetry device includes at least one sensor in a housing and a dosimetry processing device with a memory. The dosimetry processing device is coupled to the at least one sensor in the housing. The dosimetry processing device is configured to execute programmed instructions stored in the memory comprising: obtaining readings from the sensor; storing the readings; conducting an analysis of the stored readings to determine an injury risk assessment; and outputting at least one of the conducted analysis of the determined injury risk assessment or the stored readings. | 1. A dosimetry device for use in an ambient environment, the device comprising: a housing configured to be secured to at least one location on a person such that the person is wearing or carrying the housing in the ambient environment;at least one sensor at least partially disposed within the housing and accessible to the ambient environment, the at least one sensor being configured to derive a plurality of parametric signals from either a) acceleration readings generated from movement of the person, or b) pressure readings generated from perturbations propagating in the ambient environment, resulting from an identified event;at least one memory device disposed within the housing, the at least one memory device including a plurality of memory locations, the plurality of memory locations being configured to store personal medical history data of the person that is wearing or carrying the housing in the ambient environment, aggregate medical history data, and parametric data corresponding to the plurality of parametric signals;a control circuit disposed in the housing and coupled to the at least one sensor and the at least one memory device, the control circuit being configured to tailor and adjust an individualized injury threshold for the person based on the personal medical history data and the aggregate medical history data and store the tailored and adjusted individualized injury threshold prior to the identified event, and the control circuit being further configured to provide in real time an injury risk assessment corresponding to the identified event by determining whether one or more of the plurality of parametric signals exceeds the stored individualized injury threshold for the person; andat least one input/output (I/O) device coupled to the control circuit, the at least one I/O device being configured to provide the injury risk assessment in response to an assessment analysis request, whereby medical treatment of the person is guided according to the injury risk assessment that is provided by the I/O device, when the one or more of the plurality of parametric signals exceeds the stored individualized injury threshold for the person. 2. The device of claim 1, wherein the at least one sensor includes an inertial measurement sensor configured to measure an acceleration of the housing in real time, the plurality of parametric signals including at least one acceleration signal corresponding to the measured acceleration. 3. The device of claim 2, wherein the inertial measurement sensor includes a three-axis accelerometer, the at least one acceleration signal providing three-axis acceleration data in substantially real time. 4. The device of claim 2, wherein the inertial measurement sensor includes a gyroscope device, the at least one acceleration signal providing rotational acceleration data in substantially real time. 5. The device of claim 2, wherein the control circuit is configured to compare the measured acceleration with the indivi... | Rochester Institute of Technology Defense Advanced Research Projects Agency | A61B005/103 A61B005/16 A61B005/00 F41H001/02 F41H001/04 F42D005/00 A61B005/11
US09327190 | B2 | 2016-05-03 | Storage medium having stored thereon information processing program, and information procesing device | 2013-04-12 | A computer readable storage medium having stored thereon an information processing program executable by a computer of an information processing device which presents information corresponding to a biological signal acquired from a user, is provided. The information processing program causes the computer to function as biological signal acquiring means, motion/attitude information acquiring means, and presentation means. The biological signal acquiring means acquires the biological signal from the user. The motion/attitude information acquiring means acquires information about a motion or an attitude of the user from detecting means, in association with the biological signal acquired by the biological signal acquiring means. The presentation means performs predetermined presentation based on both the biological signal acquired by the biological signal acquiring means and the information acquired by the motion/attitude information acquiring means. | 1. A device, comprising: a touch input device;a vibration circuit;an accelerometer for generating accelerometer data indicative of aspects of motion of the device;wireless communication circuitry for receiving physiological data wirelessly transmitted from a physiological sensor; andmemory for storing the accelerometer data and the physiological data, whereinthe wireless communication circuitry wirelessly receives vibration circuit control signals for controlling the vibration circuit and wirelessly transmits the physiological data and the acceleration data to a computing device comprising a processing system for providing a presentation based on both the physiological data and the acceleration data. 2. The device according to claim 1, wherein the physiological data comprises heart rate data. 3. The device according to claim 1, wherein the accelerometer comprises a triaxial accelerometer. 4. A system comprising: a portable device comprising: a touch input device for receiving a touch input from a user bearing the portable device;an accelerometer for generating accelerometer data indicative of activity of the user;a gyroscope for generating gyroscope data indicative of activity of the user;a magnetic sensor for generating magnetic field data indicative of a detected magnetic field;a physiological sensor for generating physiological data associated with the user:memory for storing the accelerometer data, the gyroscope data and the physiological data;wireless communication circuitry for wirelessly communicating at least the accelerometer data, the gyroscope data and the physiological data; anda vibrator; anda computer device comprising: wireless communication circuitry for receiving the acceleration data, the gyroscope data and the physiological data from the portable device; anda processing system for presenting a presentation based on the received accelerometer data, the gyroscope data and the physiological data. 5. The system according to claim 4, wherein the wireless communication circuitry is configured for Bluetooth communication. 6. The system according to claim 4, wherein the physiological sensor generates physiological data associated with heart rate. 7. The system according to claim 4, wherein the portable device further comprises a linear arrangement of LEDs. 8. The system according to claim 4, wherein the portable device further comprises a light sensor. 9. The system according to claim 4, wherein the processing system of the computer device is configured to generate a vibrator control signal for controlling the vibrator of the portable device and the wireless communication circuitry of the computer device transmits the vibrator control signal to the portable device. | Nintendo, Co., Ltd. Tohoku University | A61B005/02 A63F013/20 A61B005/024 A61B005/00 A61M021/02 A63F013/40 A61B005/021 A61B005/026 A61B005/0402 A61B005/0488 A61B005/08 A61B005/1455 A61M021/00
US09341691 | B2 | 2016-05-17 | Short TE 3D radial sampling sequence for MRI | 2009-11-11 | A method for magnetic resonance imaging includes providing a radio frequency excitation pulse to a specimen. The pulse has a duration. The method includes, concurrent with providing the radio frequency excitation pulse, applying a first gradient having a first polarity. The method includes applying a readout gradient at a time after the duration. The readout gradient has inverse polarity relative to the first polarity. The method includes, concurrent with applying the readout gradient, acquiring magnetic resonance data from the specimen. The method includes generating an image based on the magnetic resonance data. | 1. A magnetic resonance imaging system comprising: a magnet operable to provide a magnetic field;a transmitter operable to transmit to a region within the magnetic field;a receiver operable to receive a magnetic resonance signal from the region; anda processor coupled to the transmitter and coupled to the receiver, the processor operable to cause the transmitter to provide, to the region, a radio frequency excitation pulse, a first gradient, and a readout gradient and operable to cause the receiver to generate magnetic resonance data based on the magnetic resonance signal, the processor configured to provide the radio frequency excitation pulse concurrent with the first gradient, the radio frequency excitation pulse having a duration and the first gradient having a first polarity, and after the duration, operable to provide the readout gradient, the readout gradient having an inverse polarity rely to the first polarity, and receive the magnetic resonance signal while providing the readout gradient, and the processor operable to generate an image based on the magnetic resonance data and wherein the image corresponds to a 3D radial image based on k-space data and wherein the first gradient corresponds to a spoke in the k-space data. 2. The magnetic resonance imaging system of claim 1, wherein the processor is operable to provide a time delay between the duration and the readout gradient. 3. The magnetic resonance imaging system of claim 1, wherein the radio frequency excitation pulse includes a frequency modulated pulse. 4. A method of operating a magnetic resonance scanner, the method comprising: providing a radio frequency excitation pulse to a specimen, the radio frequency excitation pulse having a duration;concurrent with providing the radio frequency excitation pulse, applying a first gradient having a first polarity, wherein applying the first gradient includes selecting a spoke in k-space data, the k-space data corresponding to a radial image;applying a readout gradient at a time after the duration, the readout gradient having inverse polarity relative to the first polarity;acquiring magnetic resonance data from the specimen while applying the readout gradient; andgenerating an image based on the magnetic resonance data. 5. The method of claim 4, wherein a magnitude of the readout gradient is substantially equal to a magnitude of the first gradient. 6. The method of claim 4, further including waiting a delay period after the duration and before applying the readout gradient. 7. The method of claim 4, wherein providing the radio frequency excitation pulse includes applying a frequency modulated pulse. 8. The method of claim 4, wherein providing the radio frequency excitation pulse includes applying at least one of a hyperbolic secant pulse or a sinc pulse. 9. The method of claim 4, wherein generating the image includes gridding. 10. The method of claim 4, wherein generating the image includes executing a Fourier transform. | Regents of the University of Minnesota | A61B005/05 G01R033/48 G01R033/46
US09331772 | B2 | 2016-05-03 | Quick re-connect diversity radio system for communicating patient data out of a shielded MR room | 2010-10-14 | An RF shielded room (14) which houses a magnetic resonance scanner (12) also houses a patient monitor (20) which includes a plurality of sensors (18) which collect physiological data about a patient undergoing a magnetic resonance examination. The monitor includes a transmitter (46) which transmits the physiological data on at least two frequency channels of a broadband frequency spectrum. A communication unit (26) located outside the RF shielded room includes first and second receivers (54, 58) which receive the physiological data from the monitor on the first and second frequency channels, respectively. A processor (62) combines physiological data received by the first and second receivers to. A display (32) displays the combined physiological data. A switch (56) switches antenna connections of the receivers (54, 58) responsive to detection of multipath fade. | 1. A system for communicating physiological data, the system comprising: a plurality of remote medical devices configured to sense physiological data about an associated patient disposed in a radio frequency (RF) shielded room RF shielded by an RF shield;a patient monitor configured to collect the physiological data from the remote medical devices and located within the RF shielded room, the patient monitor including: at least one transmitter which defines a wideband frequency spectrum divided by frequency into at least two frequency channels and which alternately transmits the physiological data alternately on the at least two frequency channels of a broadband frequency spectrum;a first communication unit located outside the RF shielded room, the first communication unit including: a first receiver configured to receive physiological data from the patient monitor on a first of the frequency channels;a second receiver configured to receive physiological data from the patient monitor on a second of the frequency channels;a first processor communicating with the first and second receivers to process physiological data received by the first and second receivers to generate a received physiological data stream;first and second antennas extending from outside the RF shielded room through the RF shield to be disposed at least partially inside the RF shielded room; anda switch configured to switch between a first configuration in which the first antenna feeds the first receiver and the second antenna feeds the second receiver and a second configuration in which the first antenna feeds the second receiver and the second antenna feeds the first receiver;wherein the first processor is further configured to operate the switch responsive to detection of a multipath destructive interference fade condition; anda display which displays the received physiological data stream. 2. The system according to claim 1, wherein the patient monitor further includes: an RF shielded room communication unit and controller configured to digitize the physiological data and assemble the digitized physiological data into data packets, each data packet including a current segment of physiological data and at least one previous segment of physiological data. 3. The system according to claim 2, wherein the first processor is further configured to: synchronize data packets received from the first and second receivers to generate the received physiological data stream with no redundant data. 4. The system according to claim 1, further including: a magnetic resonance (MR) scanner disposed in the RF shielded room and configured to image the patient associated with the remote medical devices. 5. The system according to claim 1, wherein the patient monitor includes: a display which displays the physiological data to an attendant in the RF shielded room. 6. A system for communicating physiological data, the system comprising: a patient monitor configured to be located within a radio ... | KONINKLIJKE PHILIPS N.V. | A61B005/00 H04B007/12 H04B007/06 H04B007/08 A61B005/055
US09345613 | B2 | 2016-05-24 | Regulatable intraoral mandibular advancement device, for preventing snoring and sleep apnoea | 2010-06-02 | The invention relates to a regulatable intraoral mandibular advancement device that can be used to prevent snoring and sleep apnoea. Said device comprises two splints connected by a central screw (1) inserted into a sheath (2) comprising an upper ring (3), said sheath being used as a guiding housing (4) for an upper bar (5) of an upper plate (6), and the lower part of the sheath comprising a longitudinal groove (8) enabling the course of an inner mobile threaded (10) ring (9) that is displaced by the screw (1), and on which a lower displacement ring (11) is arranged, open or closed, said ring being provided with a perpendicular hole (12) with elipsoidal conicity, in which the lower bar (13) fixed to the lower maxillary plate (14) is inserted. Said screw (1) is locked to the sheath (2) on its rear part. | 1. Intraoral regulatable mandibular advancement device, for preventing snoring and sleep apnoea, of the type used to join upper and lower maxillary plates adapted to conform the patient's maxillas and intended for prosthetic tasks in the field of odontology, whose purpose is to advance the mandible to avoid the obstruction of air flow through airways, the advancement device comprising: a screw housed in a polygonal cylindrical sheath that has on its front region an upper displacement ring which serves as a guiding house for an upper bar fixed to the upper maxillary plate, and wherein the sheath has a longitudinal groove for receiving an inner mobile ring having a thread connecting and being displaced by the screw, and on which a lower displacement ring is arranged said ring being provided with a perpendicular hole with ellipsoidal conicity that receives a lower bar that is fixed to the lower maxillary plate wherein said screw is blocked in the sheath in the rear part thereof and the upper and lower rings are prepared for giving mobility to the device on the upper and lower bars inserted in upper and lower plates, which have respectively opposing sliding planes. 2. Intraoral regulatable mandibular advancement device, for preventing snoring and sleep apnoea, according to claim 1, wherein a displacement mechanism is formed by the screw and the sheath, the screw and sheath each having small dimensions, approximately 15.5 mm in length and a frontal height of 6.5 mm, allowing by virtue of size their insertion in a central position between the maxillas. 3. Intraoral regulatable mandibular advancement device, for preventing snoring and sleep apnoea, according to claim 2, wherein the screw has a socket head suitable for a purpose-designed screwdriver. 4. Intraoral regulatable mandibular advancement device, for preventing snoring and sleep apnoea, according to claim 2, wherein the screw is retained in the polygonal cylindrical sheath directly or by a washer in a rear part thereof by radial riveting, laser welding or with a nut. 5. Intraoral regulatable mandibular advancement device, for preventing snoring and sleep apnoea, according to claim 1, wherein the upper and lower displacement rings have a specific design preventing blocking of the upper and lower bars during their course. 6. Intraoral regulatable mandibular advancement device, for preventing snoring and sleep apnoea, according to claim 5, wherein the upper and lower bars are preformed in their central part with a radial annulation of 14 to 20 degrees and show a longitudinal grooving coinciding with the dimensions of the upper ring open for dismounting of the plate. 7. Intraoral regulatable mandibular advancement device, for preventing snoring and sleep apnoea, according to claim 1, wherein the upper and lower maxillary plates have sliding platforms with opposing sliding planes. 8. Intraoral regulatable mandibular advancement device, for preventing snoring and sleep apnoea, according to clai... | LABORATORIO ORTOPLUS, S.L. | A61F005/56 A61F005/37 A61C005/14 A61F011/00 A61B019/00 A61C003/00 A61C009/00 A61C005/00 A61C003/06 A61B005/00 A61C007/36 A61C007/00 A61C007/08 A61F005/00 A61C007/06 A62B007/10 A62B018/02 A62B018/08
US09332913 | B2 | 2016-05-10 | System and method for discriminating hypervolemia, hypervolemia and euvolemia using an implantable medical device | 2011-12-21 | Techniques are provided for use by an implantable medical device or diagnostic sensor for detecting and discriminating euvolemia, hypervolemia and hypovolemia. In one example, the device detects a pressure signal within the patient representative of changes in cardiac pressure overall several cardiac cycles. The device generates separate time-domain and frequency-domain representations of the pressure signal and then discriminates among euvolemia, hypervolemia and hypovolemia within the patient based on an analysis of the time-domain and the frequency-domain representations of the signal. Depending upon the capabilities of the device, suitable warnings may be generated to alert the patient or caregiver. Diuretics or other medications can be titrated to address abnormal fluid conditions such as a fluid overload during hypervolemia. Techniques for detecting a pressure alternans pattern indicative of imminent decompensation are also described. | 1. A method for use with an implantable medical device for implant within a patient, the method comprising: applying a cardiac pressure detector to detect a pressure signal within the patient representative of changes in cardiac pressure during at least one cardiac cycle;applying a time-domain and frequency-domain processor to generate separate time-domain and frequency-domain representations of the pressure signal;applying a discrimination system to discriminate among euvolemia, hypervolemia and hypovolemia within the patient based on the time-domain and the frequency-domain representations of the pressure signal;wherein applying a discrimination system to discriminate euvolemia, hypervolemia and hypovolemia comprises: identifying an ejection period of the cardiac cycle within the time-domain representation of the pressure signal;assessing a shape of the time-domain representation of the pressure signal during the ejection period; anddiscriminating hypervolemia from hypovolemia and euvolemia based on the shape of the time-domain representation of the pressure signal during the ejection period; andapplying a controller unit to control at least one implantable medical device therapeutic function in response to applying a discrimination system to discriminate among euvolemia, hypervolemia and hypovolemia. 2. The method of claim 1 wherein applying a discrimination system to discriminate hypervolemia from hypovolemia and euvolemia comprises: generating an indication of hypervolemia if the time-domain representation of the pressure signal during the ejection period has a peaked shape; andgenerating an indication of lack of hypervolemia if the time-domain representation of the pressure signal during the ejection period has a more rounded shape. 3. The method of claim 1 wherein identifying the ejection period comprises: determining first order derivatives (dP/dt) and second order derivatives (d2P/dt2) of the time-domain representation of the pressure signal;identifying fiducial markers within the time-domain representation of the pressure signals based on the first and second order derivatives including: a maximum magnitude of dP/dt (dP/dt max); a time of dP/dt max; a minimum magnitude of dP/dt (dP/dt min); a time of dP/dt min; a peak magnitude of the pressure (P max); a time of P max; and a minimum magnitude of d2P/dt2; andidentifying the ejection period as the interval between the time of minimum d2P/dt2 after dP/dt max and the time of minimum d2P/dt2 before dP/dt min. 4. The method of claim 1 wherein assessing the shape of the time-domain representation of the pressure signal during the ejection period comprises generating a shape index representative of a degree of roundness to the time domain representation of the pressure signal during the ejection period. 5. The method of claim 4 wherein the shape index is generated based on one or more of a best-fit quadratic curve or a best-fit cosine curve. 6. A method for use with an implantable medical ... | PACESETTER, INC. | A61B005/02 A61B005/0215 A61N001/365 A61B005/00 A61N001/39
US09351655 | B2 | 2016-05-31 | Systems, devices, and methods for electrically coupling terminals to electrodes of electrical stimulation systems | 2009-08-21 | An implantable lead includes a first lead assembly with a distal tip and a medial end, a medial section with a first end and a second end, and a first intermediate assembly disposed between the first lead assembly and the first end of the medial section. The first lead assembly includes a plurality of external contacts and at least one conductive wire disposed in the first lead assembly. The at least one conductive wire extends from at least one external contact towards the medial end of the first lead assembly. The medial section includes a plurality of conductors extending from the first end to the second end. The first intermediate assembly includes a plurality of conductive elements. At least one of the conductive elements is configured and arranged to electrically couple the at least one conductive wire to at least one of the conductors. | 1. An electrical stimulating system comprising an implantable lead comprising: an elongated lead body having proximal end, a distal end, and a longitudinal length, the lead body comprising a first lead assembly disposed at one of the proximal end or the distal end of the lead body, the first lead assembly having a tip and an opposing medial end, the first lead assembly comprising a plurality of first external contacts, and at least one conductive wire disposed in the first lead assembly, the at least one conductive wire extending from at least one first external contact towards the medial end of the first lead assembly;a medial section with a first end and a second end, the medial section comprising a plurality of conductors extending from the first end to the second end; anda first intermediate assembly disposed between the first lead assembly and the first end of the medial section, the first intermediate assembly comprising a plurality of conductive elements, at least one of the conductive elements configured and arranged to electrically couple the at least one conductive wire of the first lead assembly to at least one of the conductors at the first end of the medial section;wherein the first lead assembly, the medial section, and the first intermediate assembly maintain an isodiametric profile along the longitudinal length of the lead body;wherein the lead body further comprises a second lead assembly electrically coupled to the second end of the medial section, the second lead assembly comprising a plurality of second external contacts, andat least one conductive wire disposed in the second lead assembly, the at least one conductive wire electrically coupling at least one second external contact disposed on the second lead assembly to at least one conductor at the second end of the medial section;wherein the first lead assembly is disposed at the distal end of the lead body and the second lead assembly is disposed at the proximal end of the lead body;wherein the plurality of first external contacts comprise a plurality of electrodes and the plurality of second external contacts comprise a plurality of terminals;a control module configured and arranged to electrically couple to the proximal end of the lead body, the control module comprising a housing, andan electronic subassembly disposed in the housing; anda connector for receiving the proximal end of the lead body, the connector having a proximal end, a distal end, and a longitudinal length, the connector comprising a connector housing defining a port at the distal end of the connector, the port configured and arranged for receiving, the proximal end of the lead body, anda plurality of connector contacts disposed in the connector housing, the connector contacts configured and arranged to couple to at least one of the plurality of terminals disposed on the proximal end of the lead body; anda lead extension coupling the lead body to the control module, the lead extension comprising an elonga... | Boston Scientific Neuromodulation Corporation | A61B005/042 A61N001/05 H01R024/58 H01R013/52 H01R107/00
US09839762 | B2 | 2017-12-12 | System and method for reducing stress levels using color, sound, and aroma therapy | 2015-11-20 | The present invention is directed toward a system and method utilizing synchronized colors, sounds, and aroma therapy. The visual, auditory, and aromatic stimulation is intended to alleviate stress and/or anxiety by relaxing and stimulating the brain, thereby correcting any neurophysical imbalances. Reduction in the levels of stress and/or anxiety is achieved through the use of the auditory, visual, and aromatic sequences which are individually specific and systematic. The system and method is preferably provided in multiple sessions with each session providing both a mechanism for providing treatment as well as assessment, which is used to form the basis for future treatments. | 1. A non-pharmacological method of reducing anxiety and stress levels within an individual comprising: assessing an individual's psychological status by subjecting said individual to a first treatment determining test, said first treatment determining test including administration of a questionnaire containing a plurality of pre-determined questions designed to assess said individual's emotional state;electronically analyzing data obtained from said first treatment determining test;determining a first treatment session to be administered to said individual, said first treatment session based on data obtained from said questionnaire;providing a first treatment session to said individual based on data obtained from said first treatment determining test, said first treatment session comprising:using an electronic display unit to expose said individual to a color treatment, said color treatment including exposure to specific colored light and one or more colored images based on a particular color treatment group; andusing a device for displaying sound to expose said individual to auditory notes, said auditory notes having preset frequencies for said particular color based treatment applied to said individual; andusing an aroma therapy dispensing device to expose said individual to one or more aromas, said one or more aromas chosen based on said color based treatment applied to said individual;said electronically analyzed data obtained from said first treatment determining test resulting in said treatment session generated having a specific sequence of said colored light, said one or more colored images, said auditory notes, or said one or more aromas, wherein said electronic display unit, said device for displaying sound, or said aroma therapy dispensing device are coordinated to display or dispense said colored light, colored images, auditory sounds or aromas based on said specific sequence determined. 2. The non-pharmacological method of reducing anxiety and stress levels within an individual according to claim 1 wherein said questionnaire includes a set of questions which, when answered in a predetermined manner, assign said individual to one of said color treatment groups. 3. The non-pharmacological method of reducing anxiety and stress levels within an individual according to claim 1 wherein said color treatment groups include a red color treatment group wherein exposure to colored light and one or more colored images is based on the color red, an orange color treatment group wherein exposure to colored light and one or more colored images is based on the color orange, a yellow color treatment group wherein exposure to colored light and one or more colored images is based on the color yellow, a green color treatment group wherein exposure to colored light and one or more colored images is based on the color green, a blue color treatment group wherein exposure to colored light and one or more colored images is based on the color blue, a purple ... | Harmonial Corporation | A61M021/02 A61B005/16 A61B005/053 A61B005/00 A61M021/00
US20160143532 | A1 | 2016-05-26 | WEARABLE DEVICE AND METHOD OF CONTROLLING THE SAME | 2015-11-20 | A wearable device and a control method thereof are provided. The wearable device includes a processor, a sensor module configured to obtain biometric information about a user wearing the wearable device, and a light emitting module configured to output visual information corresponding to the obtained biometric information, wherein the processor controls an output attribute of the visual information to be changed and output according to a change in the obtained biometric information. | 1. A wearable device comprising: a processor;a sensor module configured to obtain biometric information about a user wearing the wearable device; anda light emitting module configured to output visual information corresponding to the obtained biometric information,wherein the processor controls an output attribute of the visual information to be changed and output according to a change in the obtained biometric information. 2. The wearable device of claim 1, wherein the biometric information includes information on a pulse cycle of the user, and the visual information includes at least one color output by the light emitting module. 3. The wearable device of claim 2, wherein the output attribute includes an output period of a color output by the light emitting module or a kind of color output by the light emitting module. 4. The wearable device of claim 3, wherein when the pulse cycle is decreased, the processor controls the output period to be decreased and output according to the decreased pulse cycle. 5. The wearable device of claim 3, wherein when the pulse cycle is increased, the processor controls the output period to be increased and output according to the increased pulse cycle. 6. The wearable device of claim 3, wherein when the pulse cycle is increased or decreased, the processor controls a color output by the light emitting module to be output according to a predetermined output period. 7. A method of controlling a wearable device, the method comprising: obtaining biometric information about a user wearing the wearable device;outputting visual information corresponding to the obtained biometric information; andchanging an output attribute of the visual information according to a change of the obtained biometric information and outputting the visual information. 8. The method of claim 7, wherein the biometric information includes information on a pulse cycle of the user, and wherein the visual information includes at least one color output by the light emitting module. 9. The method of claim 8, wherein the output attribute includes an output period of a color output by the wearable device or a kind of color output by the wearable device. 10. The method of claim 9, further comprising: when the pulse cycle is decreased, decreasing the output period according to the decreased pulse cycle and outputting the visual information. 11. The method of claim 9, further comprising: when the pulse cycle is increased, increasing the output period according to the decreased pulse cycle and outputting the visual information. 12. The method of claim 9, further comprising: when the pulse cycle is increased or decreased, outputting a color output by the wearable device according to a predetermined output period. 13. A wearable device comprising: a light emitting module;a sensor module configured to detect a movement of a user wearing the wearing device;a communication module configured to receive information on a first distance set by the user; ... |  | A61B005/00
US20160125597 | A1 | 2016-05-05 | NOVEL MAGNETIC RESONANCE IMAGE TECHNIQUE FOR IMAGING AND EVALUATING COLLATERAL CIRCULATION | 2014-04-17 | The present invention relates to a novel magnetic resonance image technique for imaging collateral circulation in a human body. In the present invention, source data of a dynamic susceptibility contrast enhancement perfusion-weighted image (DSC-PWI) is subjected to an image-post-processing and an image obtained through the image-post-processing is then compared with an image obtained by digital subtraction angiography (DSA), thereby imaging the collateral circulation. The method according to the present invention can non-invasively, rapidly and easily visualize and evaluate the collateral circulation. Therefore, the method can provide information, which is important for understanding, treatment strategy determination and outcome prediction of a disease generated due to angiostenosis or occlusion. | 1. A method for imaging collateral circulation in the body, the method comprising: (a) image post-processing source data of dynamic susceptibility contrast-enhanced perfusion weighted imaging (DSC-PWI) as magnetic resonance imaging in the body; and(b) comparing images obtained through the image post-processing in step (a) with digital subtraction angiography (DSA) images to image collateral circulation. 2. The method of claim 1, wherein the image post-processing of the DSC-PWI source data in step (a) comprises step (a)-1 of subtracting a signal of an image obtained before the injection of a contrast agent from signals of images obtained after the injection of the contrast agent to obtain sequential images over time. 3. The method of claim 2, wherein the image post-processing of DSC-PWI source data in step (a) comprises step (a)-2 of merging adjacent consecutive image cuts in the sequential images obtained in step (a)-1. 4. The method of claim 3, wherein the merging of the adjacent consecutive image cuts is performed by merging two to ten image cuts. 5. The method of claim 1, wherein the comparing of the DSC-PWI images post-processed in step (b) and the DSA image is performed on images corresponding to the arterial, capillary, venous, or late venous phase. 6. The method of claim 5, wherein the DSC-PWI images in the arterial, capillary, venous, and late venous phases are composed of consecutive images, each of which is obtained by merging adjacent consecutive image cuts. 7. The method of claim 6, wherein the DSC-PWI images of the arterial phase are merged to construct a collateral circulation image in the early phase; the DSC-PWI images of the capillary phase are merged to construct a collateral circulation image in the mid phase; and the DSC-PWI images of the venous and late venous phases are merged to construct a collateral circulation image in the late phase. 8. The method of claim 1, wherein the collateral circulation is collateral circulation in the brain or heart region. 9. A method for creating a collateral circulation image, the method comprising: (a)′ obtaining R2* value at each voxel in dynamic susceptibility contrast-enhanced perfusion weighted imaging (DSC-PWI) source data by inserting the magnetic resonance signal intensity at the corresponding time, S(t); the baseline magnetic resonance signal intensity, So; and the echo time TE into the following equation: R2*(t)=−1/TE log [S(t)/So](b)′ summing all R2* values at every voxel on each cut, and plotting the change in the summated R2* value over time on a graph;(c)′ finding out the maximum summated R2* value from the graph, and setting an image cut corresponding to the maximum value as the midpoint of the mid-phase of a collateral circulation image; and(d)′ creating the collateral circulation image by merging adjacent images using the image cut, which has been set as the midpoint of the mid-phase of the collateral circulation image, as a reference. 10. A computer readable recordi... |  | G06T007/00 G06T005/50 G01R033/563 A61B005/026 G01R033/56
US20160120453 | A1 | 2016-05-05 | System for monitoring incontinence | 2014-11-04 | According to one aspect of the inventive concept there is provided a system for monitoring incontinence comprising: a urine sensitive circuit arranged to present a changed electrical characteristic when exposed to urine; a measurement circuit arranged to perform a measurement on a urine bladder of a wearer to determine at least one parameter which varies with a fill level of the urine bladder; and a processing circuit arranged to: determine whether the urine sensitive circuit has been exposed to urine, and in response to determining that the urine sensitive circuit has been exposed to urine, record data representing the at least one parameter determined by the measurement circuit. | 1. A system for monitoring incontinence comprising: a urine sensitive circuit arranged to present a changed electrical characteristic when exposed to urine;a measurement circuit arranged to perform a measurement on a urine bladder of a wearer to determine at least one parameter which varies with a fill level of the urine bladder; anda processing circuit arranged to: determine whether the urine sensitive circuit has been exposed to urine, andin response to determining that the urine sensitive circuit has been exposed to urine, record data representing the at least one parameter determined by the measurement circuit. 2. A system according to claim 1, wherein the processing circuit is further arranged to estimate a urine bladder fill level based on the at least one determined parameter. 3. A system according to claim 1, wherein the processing circuit is further arranged to determine a urine bladder fill threshold based on the at least one determined parameter represented by the recorded data. 4. A system according to claim 1, wherein the processing circuit is further arranged to determine a urine bladder fill threshold based on the at least one determined parameter represented by the recorded data and a previously determined urine bladder fill threshold. 5. A system according to claim 1, further comprising a sensor arranged to determine an orientation or a movement of the sensor, and the processing circuit being further arranged to, in response to determining that the urine sensitive circuit has been exposed to urine, record data representing an estimated movement and/or posture of the wearer based on an orientation and/or a movement determined by the sensor. 6. A system according to claim 5, wherein the processing circuit is further arranged to determine a urine bladder fill threshold based on the at least one determined parameter represented by the recorded data and the estimated movement and/or posture. 7. A system according to claim 5, wherein the measurement circuit is arranged to repeatedly determine at least one parameter which varies with a fill level of the urine bladder and the processing circuit is further arranged to compare the repeatedly determined parameters to one or more urine bladder fill thresholds, each threshold being associated with an orientation or a movement determined by the sensor. 8. A system according to claim 1, wherein the processing circuit is arranged to provide a signal which is indicative of a urine bladder fill level based on the at least one determined parameter. 9. A system according to claim 1, further comprising at least one or a combination of a microphone, a temperature sensor, a skin contact sensor, an accelerometer or an altimeter. 10. A system according to claim 1, wherein the measurement circuit and the processing circuit are galvanically connected. 11. A system according to claim 10, wherein the processing circuit and the urine sensitive circuit are galvanically separated. 12. A system accord... |  | A61B005/20 A61B008/00
US20160135891 | A1 | 2016-05-19 | SYSTEMS AND METHODS FOR IMAGING AND MANIPULATING TISSUE | 2015-09-11 | Exemplary embodiments of the present disclosure include systems and methods capable of imaging, manipulating, and analyzing tissue using light, including for example, coagulating and breaking the molecular bonds (e.g. cutting) tissue. | 1. A system comprising: a first light source configured to provide a signal for use in imaging tissue when the first light source is incident upon tissue;a second light source configured to coagulate tissue when the second light source is incident upon tissue; anda third light source configured to break molecular bonds of tissue when the third light source is incident upon tissue. 2. The system of claim 1 wherein the third light source is a diode laser seeded fiber amplified source. 3. The system of claim 2 wherein the diode seeded fiber amplified source is configured to emit energy in a range of wavelengths from 1800 nm to 2200 nm. 4. The system of claim 2 wherein the diode seeded fiber amplified source is configured to emit energy having a pulse profile, a pulse energy, and a pulse repetition rate. 5. The system of claim 4 wherein at least one of the pulse profile, pulse energy and pulse repetition rate can be controlled to adjust a tissue removal rate. 6. The system of claim 1 wherein the third light source is a tunable semiconductor laser seeded fiber amplified source. 7. The system of claim 6 wherein the tunable semiconductor laser seeded fiber amplified source is configured to emit energy in a range of wavelengths from 1800 nm to 2200 nm. 8. The system of claim 6 wherein the tunable semiconductor laser seeded fiber amplified source is configured to emit energy having a pulse profile, a pulse energy, and a pulse repetition rate. 9. The system of claim 8 wherein at least one of the pulse profile, pulse energy and pulse repetition rate can be controlled to adjust a tissue removal rate. 10. The system of claim 1 wherein the third light source is configured to break molecular bonds of tissue coagulated by the second light source when the third light source is incident upon tissue. 11. The system of claim 1 wherein the third light source is configured to alter the quaternary structure of proteins of tissue when the third light source is incident upon tissue. 12. The system of claim 1 wherein the first light source, the second light source and the third light source emit light through a single fiber at the same instance. 13. The system of claim 12 wherein the single fiber is a component of an endoscope or laproscope. 14. The system of claim 1 wherein the first light source, the second light source and the third light source emit light through a single fiber at different times. 15. The system of claim 14 wherein the single fiber is a component of an endoscope or laproscope. 16. The system of claim 1 wherein the signal obtained from the first light source is used to orient or position the second light source and the third light source. 17. The system of claim 16 wherein the signal obtained from the first light source is an input into a computer processor. 18. The system of claim 17 wherein the computer processor provides output data used to control an orientation or position of the second light source and the third light source. 19... |  | A61B018/22 A61B005/00
US20160135734 | A1 | 2016-05-19 | COMBINATION THERAPY FOR SLEEP DISORDERED BREATHING AND HEART FAILURE | 2015-11-18 | Methods and apparatus provide evaluation of a cardiac condition of a patient with sleep disordered breathing. One or more processors, such as processors associated with any of ICDs, CRTs and/or respiratory pressure therapy devices, may be configured to receive data representing one or more respiratory parameters for the patient. The processor(s) may receive cardiac-related patient data for the patient. The processor(s) may determine a presence of a smooth hemodynamic baseline from the received data representing one or more respiratory parameters for the patient. The processors may evaluate the cardiac condition of the patient from the received cardiac-related patient data based on the determined presence of the smooth hemodynamic baseline. | 1. A method for evaluating a cardiac condition of a patient with sleep disordered breathing comprising: receiving data representing one or more respiratory parameters for the patient;receiving cardiac-related patient data for the patient;determining in a processor a presence of a smooth hemodynamic baseline from the received data representing one or more respiratory parameters for the patient; andevaluating, with the processor, the cardiac condition of the patient from the received cardiac-related patient data based on the determined presence of the smooth hemodynamic baseline. 2. The method of claim 1 wherein the determining comprises detecting absence of sleep disordered breathing events for a predetermined period of time. 3. The method of claim 1 wherein the determining comprises: detecting sleep disordered breathing events; andcontrolling a respiratory therapy to alleviate the detected sleep disordered breathing events for a predetermined period of time. 4. The method of claim 1 further comprising collecting, with one or more sensors, data representing one or more hemodynamic parameters for the patient while the patient is within the smooth hemodynamic baseline. 5. The method of claim 4, wherein the evaluating is based on the data representing one or more hemodynamic parameters. 6. The method of claim 3 wherein the controlling comprises controlling a respiratory pressure therapy device to deliver respiratory pressure therapy for the patient. 7. The method of claim 6, wherein the respiratory pressure therapy comprises an adaptive servo-ventilation therapy. 8. The method of claim 2 wherein the sleep disordered breathing events comprise one or more of apnea, hypopnea, Cheyne-Stokes respiration, snoring and flow limitation. 9. The method of claim 3, wherein the controlling comprises controlling an implanted device to provide at least one of genioglossal nerve stimulation and phrenic nerve stimulation. 10. The method of claim 1, further comprising determining, based on the evaluating, efficacy of a cardiac treatment that has been provided to the patient. 11. The method of claim 10, wherein the cardiac treatment comprises one or more of administering a pharmaceutical cardiac drug and application of a cardiac resynchronization therapy. 12. The method of claim 10, further comprising titrating the cardiac treatment according to the cardiac condition of the patient. 13. The method of claim 12 wherein the titration of the cardiac treatment comprises determination of a diuretic dose. 14. The method of claim 12 wherein the titration of the cardiac treatment comprises an adjustment to a target ventilation of a respiratory pressure therapy. 15. The method of claim 14 wherein the respiratory pressure therapy is a pressure support ventilation therapy controlled in accordance with the target ventilation. 16. The method of claim 5 wherein the evaluating comprises outputting one or more hemodynamic parameters according to a time period of the smoo... |  | A61B005/00 A61M016/00 A61N001/36 A61B005/0205
US20160135712 | A1 | 2016-05-19 | SYSTEM AND METHOD FOR DETERMING THE POSITION OF A CATHETER | 2015-11-10 | A system for locating a catheter includes a catheter with a locatable tip and a hardware platform. The locatable tip can include a magnetic field generator. The hardware can include a plurality of sensor clusters and control circuitry. Each of the sensor clusters can measure a magnetic field generated by the magnetic field generator. The control circuitry can receive signals from the first, second, and third sensors clusters based on the measured magnetic fields and display a graphical representation for a position of the locatable tip. | 1. A catheter location device, comprising: one or more sensor clusters, each sensor cluster to measure a first set of voltage values when a locatable tip of a catheter is at a first position and a second set of voltage values when the locatable tip in at a second position; andcontrol circuitry to include a processor communicatively coupled to the one or more sensor clusters and a non-transitory computer-readable medium, the processor executing instructions stored on the non-transitory computer-readable medium to: calculate the second position based in part on the first position and in part on a change between the first set of voltage values and the second set of voltages, andstore the second position of the locatable tip to the non-transitory computer-readable medium. 2. The device of claim 1, wherein each sensor cluster includes at least two directional sensors. 3. The device of claim 2, the directional sensors each comprising a current block and first and second transducers. 4. The device of claim 3, the first transducer to measure a first voltage difference across a first dimension of the conductor block and the second transducer to measure a second voltage difference across a second dimension of the conductor block. 5. The device of claim 4, wherein the first and second dimensions of the conductor block are generally orthogonal with respect to each other. 6. The device of claim 4, the first set of voltage values to include the first and second voltage differences when the locatable tip is at the first position and the second set of voltage values to include the first and second voltage differences when the locatable tip is at the second position. 7. The device of claim 1, the processor executing instructions stored on the non-transitory computer-readable medium to communicate the second position relative to the first position via a user interface. 8. The device of claim 7, comprising a housing to enclose one or more of the plurality of sensor clusters, the processor, the non-transitory computer-readable medium, and the user interface. 9. The device of claim 5, wherein the plurality of sensor clusters are located above a chest cavity of a patient when the second position is calculated. 10. A catheter location system, comprising: a catheter with a locatable tip, the locatable tip including a magnetic field generator; anda hardware platform, the hardware platform to include: one or more sensor clusters, each of the sensor clusters to measure a magnetic field generated by the magnetic field generator, andcontrol circuitry to receive signals from the one or more sensor clusters based on the magnetic fields measured by the first, second, and third sensor clusters and display a graphical representation for a position of the locatable tip. 11. The catheter location system of claim 10, wherein the hardware platform is conductively coupled to the magnetic field generator, the magnetic field generator to include an electromagnet. 12. The cath... | AngioDynamics, Inc. | A61B005/06 A61B005/00
US09345434 | B2 | 2016-05-24 | Physiological signal measurement apparatus capable of automatically adjusting a measure position and method of utilizing the apparatus | 2014-06-23 | A physiological signal measurement apparatus is capable of automatically adjusting a measure position and suitable for installed on a support element to measure a physiological signal of a user. The physiological signal measurement apparatus includes a movable element, a physiological signal sensing element, a pressure sensing unit and a microcontroller unit. The movable element has a first pressure. The user exerts a second pressure on the physiological signal sensing element, and exerts a third pressure on the support element. The pressure sensing unit senses the first pressure, the second pressure and the third pressure to generate a first pressure signal, a second pressure signal and a third pressure signal. The microcontroller unit receives the physiological signals and the pressure signals, and controls the movable element by the pressure signals and the physiological signals, in order to increase the quality of signal measurement. | 1. A physiological signal measurement apparatus, capable of automatically adjusting a measure position and suitable for installed on a support element to measure a physiological signal of a user, the physiological signal measurement apparatus comprising: a movable element, having a first pressure inside thereof;a physiological signal sensing element, disposed on the surface of the movable element for sensing the physiological signal, the user exerting a second pressure on the physiological signal sensing element, and exerting a third pressure on the support element;a pressure sensing unit, electrically connected to the movable element and the physiological signal sensing element, for sensing the first pressure, the second pressure and the third pressure to generate a pressure signal; anda microcontroller unit, electrically connected to the physiological signal sensing element, the movable element and the pressure sensing unit, for receiving the pressure signal sent from the pressure sensing unit and the physiological signal sent from the physiological signal sensing element, wherein the microcontroller unit processes the pressure signal and the physiological signal with amplifying and filtering, and controls the movable element by means of the processed pressure signal and the physiological signal, in order to keep the physiological signal sensing element touching the user. 2. The physiological signal measurement apparatus of claim 1, wherein the physiological signal sensing element is a electrode for receiving the physiological signal of the user. 3. The physiological signal measurement apparatus of claim 1, wherein the movable element adjusts the first pressure by expansion and contraction. 4. The physiological signal measurement apparatus of claim 3, wherein the movable element is selected from one of an airbag and a spring. 5. The physiological signal measurement apparatus of claim 1, wherein the physiological signal is selected from one of Electrocardiography (ECG) signal, photoplethysmogram (PPG) signal, body temperature and blood flow. 6. The physiological signal measurement apparatus of claim 1, wherein the microcontroller unit has a signal processing circuit for amplifying and filtering the pressure signal sent from the pressure sensing unit and the physiological signal sent from the physiological signal sensing element. 7. The physiological signal measurement apparatus of claim 1, further comprising an user interface electrically connected to the microcontroller unit for displaying an image showing the processed pressure signal and the physiological signal. | National Taiwan University of Science and Technology | A61B005/02 A61B005/00 A61B005/0402 A61B005/026 A61B005/024
US10568585 | B2 | 2020-02-25 | Percutaneous coronary intervention (PCI) planning interface with pressure data and vessel data and associated devices, systems, and methods | 2015-11-12 | Devices, systems, and methods directed to evaluating a vessel of a patient are provided. The method includes outputting, to a display, a screen display including: a visual representation of a pressure ratio of pressure measurements obtained by first and second instruments positioned within a vessel while the second instrument is moved longitudinally through the vessel and the first instrument remains stationary within the vessel; and a visual representation of the vessel; receiving a user input to modify one of the visual representations of the pressure ratio and the vessel to simulate a therapeutic procedure; and updating the screen display, in response to the user input, including: modifying the selected one of the visual representation of the pressure ratio and the vessel based on the received user input; and correspondingly modifying the unselected one of the visual representation of the pressure ratio and the vessel. | 1. A system for evaluating a vessel of a patient, comprising: a first instrument sized and shaped for introduction into the vessel of the patient;a processing system communicatively coupled to the first instrument, a second instrument, and a display device, the processing system configured to: receive pressure measurements from the first instrument and the second instruments while the first instrument is moved longitudinally through the vessel and the second instrument remains stationary within the vessel;output, to the display device, a screen display including: a visual representation of a pressure ratio of pressure measurements received from the first and second instruments; anda visual representation of the vessel;receive a user input to modify one of the visual representations of the pressure ratio and the vessel to simulate a therapeutic procedure; andupdate the screen display, in response to the user input, by: modifying the selected one of the visual representation of the pressure ratio and the vessel based on the received user input; andcorrespondingly modifying the unselected one of the visual representation of the pressure ratio and the vessel. 2. The system of claim 1, wherein the visual representation of the vessel is an angiographic image. 3. The system of claim 2, wherein the processing system is configured to receive a user input by receiving a user input to insert a graphical representation of a stent in one of the visual representations of the pressure ratio and the vessel to simulate a percutaneous coronary intervention. 4. The system of claim 3, wherein the processing system is further configured to determine physiological parameters for a stent to be deployed in the vessel based on the characteristics of the graphical representation of the stent. 5. The system of claim 4, wherein: the physiological parameters include at least one of stent position, stent length, and stent diameter; andthe characteristics of the graphical representation of the stent include at least one of position, length, and diameter. 6. The system of claim 5, wherein the processing system is further configured to determine at least one of the stent position, stent length, stent diameter, the position of the graphical representation of the stent, the length of the graphical representation of the stent, and the diameter of the graphical representation of the stent based on at least one of the angiography data, the obtained pressure measurements, a pressure ratio calculated based on the obtained pressure measurements, and intravascular imaging data obtained within the vessel. 7. The system of claim 5, wherein the processing system is configured to: receive a user input further including receiving a user input to perform at least one of changing a length of the visual representation of the stent and changing a position of the visual representation of the stent in the visual representation of the pressure ratio;modify the selected one of the visual repres... | PHILIPS IMAGE GUIDED THERAPY CORPORATION | A61B005/00 G16H050/50 A61B005/0215 A61B006/00
US09333198 | B2 | 2016-05-10 | Treatment for cerebral palsy gait impairment | 2015-10-13 | A methylphenidate, particularly including dextro-threo-methylphenidate, is administered to a child to treat a gait impairment secondary to cerebral palsy, or a gait impairment resultant side effect from anti-cancer and anti-seizure drugs. The administered methylphenidate simultaneously treats speech and gait or limb impairments secondary to a disease or disorder. | 1. A pharmaceutical intervention for treating a gait impairment secondary to cerebral palsy in a child in need of such treatment, said pharmaceutical intervention comprises: a methylphenidate;whereby the gait impairment is diminished;and the gait impairment diminishment persists when the methylphenidate is no longer efficaciously present in the child. 2. The pharmaceutical intervention of claim 1, wherein the methylphenidate comprises dextro-threo-methylphenidate. 3. The pharmaceutical intervention of claim 1, wherein the methylphenidate comprises an analog of a methylphenidate. 4. The pharmaceutical intervention of claim 3, wherein the methylphenidate analog comprises a nonlinear lower alkyl phenidate. 5. The pharmaceutical intervention of claim 4, wherein the nonlinear lower alkyl phenidate comprises isopropylphenidate. 6. The pharmaceutical intervention of claim 1, wherein the GMFCS level number after administration of the pharmaceutical intervention is at least one level number less than before the intervention. 7. The pharmaceutical intervention of claim 1, wherein the methylphenidate consists essentially of dextro-threo-methylphenidate. 8. The pharmaceutical intervention of claim 6, wherein the methylphenidate consists essentially of dextro-threo-methylphenidate. 9. A pharmaceutical intervention for treating a gait impairment secondary to cerebral palsy in a child in need of such treatment, said pharmaceutical intervention comprises: a methylphenidate;wherein the gait impairment is diminished, and the gait impairment diminishment persists at a GMFCS level at least one level number less than prior to the intervention. 10. The pharmaceutical intervention of claim 9, wherein the methylphenidate comprises dextro-threo-methylphenidate. | GILROSE PHARMACEUTICALS, LLC | A61K031/435 A61K031/4458 A61B005/00 A61K031/137 A61K045/06 A61B005/055
US10485427 | B2 | 2019-11-26 | System and method for synthesis of impedance matching and signal converting material for all optical photo-acoustic detection | 2014-11-06 | A system and method is disclosed for non-contact detection and/or imaging and/or monitoring target subsurface tissue of at least one of human or animal anatomy. The system applies a first optical excitation signal to an outer tissue surface at a first location on the anatomy, which excites the target subsurface tissue to produce acoustic signals which are transmitted to an outer tissue surface of the anatomy. A gel-like, impedance matching and signal converting (IMASC) material layer is applied to the outer tissue surface at a second location on the anatomy. The IMASC material layer contains material elements which are able to influence characteristics of an optical signal impinging and reflected from the IMASC material, in accordance with acoustic signals that have been reflected from the target subsurface tissue, and which propagate into the IMASC material. An optical ultrasound detection system is used to process the reflected optical signals reflected from the IMASC material to provide information that may be used to provide an image of the target subsurface tissue. | 1. A system for non-contact detection and/or imaging and/or monitoring target subsurface tissue of a target, the system comprising: a housing;an optical excitation light generating system housed in the housing and forming a first light source for generating an optical excitation signal applied to an outer tissue surface at a first location on the target, which excites the target subsurface tissue to produce acoustic signals which are transmitted back to the outer tissue surface of the target;first and second gel-like, impedance matching and signal converting (IMASC) material portions which are adapted to be applied to the outer tissue surface at a pair of spaced apart second locations on the target, the IMASC material portions containing material elements embedded therein which are able to influence characteristics of an optical signal impinging and reflected from the IMASC material portions at the pair of spaced apart second locations on the target, in accordance with acoustic signals that have been reflected from the target subsurface tissue, and which propagate into the IMASC material portions at the pair of spaced apart second locations on the target;an optical heterodyne ultrasound detection system housed in the housing, and including a photodetector and a second light source, the second light source configured to emit at least one optical signal, which is split into a first emitted signal portion and a second emitted signal portion;the gel-like IMASC material portions operating to receive the first and second emitted signal portions at the second locations and to reflect back a pair of first reflected signal portions, where each one of the pair of first reflected signal portions is modified by the gel-like IMASC material in relation to the acoustic signals reaching the IMASC material at the second locations;the optical heterodyne ultrasound detection system further including being housed adjacent the optical excitation light generating system within the housing to enable the transmission of the optical excitation signal, the transmission of the first and second emitted signal portions, and the reception of the pair of first reflected signal portions, all from a single location represented by the housing;the photodetector being configured and positioned to receive both ones of the pair of first reflected signal portions for processing by the optical heterodyne detection system, the processing including analyzing resulting differences between the pair of first reflected signal portions and a second reflected reference signal portion, to generate electrical signals that relate to optical and physical properties of the subsurface target tissue; andwherein the IMASC material portions each have a viscosity high enough so that each one of the IMASC material portions at least one of: does not flow; orhas a minimal flow, outside the area where it is positioned; andwherein the IMASC material portions are each based on either a hydrogel or a lipogel t... | Lawrence Livermore National Security, LLC | A61B005/00 A61K049/22 A61K049/00
US10130248 | B2 | 2018-11-20 | Method for realizing ocular fundus photographs that are corrected for scattered light | 2014-05-15 | A method by which fundus photographs that are corrected for scattered light can be realised by using a digital fundus camera. The fundus is illuminated and a photograph of the fundus is realized. The photographed area is greater than the illuminated are. The scattered light intensity is determined and is used for correction. Measurement points for determining the scattered light intensity are defined in the non-illuminated area. Values of the scattered light intensity determined at the measurement points are averaged. Average-scattered light intensity thus determined is subtracted from the intensity values of the photograph to correct the photograph. | 1. A method for realizing ocular fundus images that are corrected for scattered light using a digital fundus camera, the method comprising: a) illuminating the ocular fundus by application of an illumination source located within the digital fundus camera and directed through optics of the digital fundus camera into an eve and onto the ocular fundus, andb) capturing a digital image of the ocular fundus through the optics of the digital fundus camera in which an imaged surface area of the ocular fundus is larger than an illuminated surface area of the ocular fundus,wherein in the digital image of the ocular fundus realized in b) a scattered light intensity is determined and is used for a correction, the method further comprisingc) defining, by application of the digital fundus camera, measurement points or a measurement area, the measurement points or the measurement area for determining the scattered light intensity being located in a non-illuminated area of the ocular fundus adjacent to or surrounding the illuminated area of the ocular fundus;d) in the case of measurement points in the non-illuminated area of the ocular fundus, eliminating outliers and after the outliers have been eliminated, averaging values of scattered light intensity determined at said measurement points to determine an averaged scattered light intensity by application of the digital fundus camera; ande) using the average scattered light intensity thus determined to correct said image of the ocular fundus realized in method step b) or to define the severity of a cataract by application of the digital fundus camera. 2. The method in accordance with claim 1, further comprising: a) illuminating the ocular fundus at an illumination angle β, and b) capturing a digital image of the ocular fundus at a field of view angle α, wherein β&lt;α. 3. The method in accordance with claim 1, further comprising, for determining the scattered light intensity in accordance with method step c), utilizing at least 5 of the measurement points. 4. The method in accordance with claim 3, further comprising, utilizing at least 10 of the measurement points. 5. The method in accordance with claim 3, further comprising utilizing at least 20 of the measurement points and distributing the measurement points as uniformly as possible in the non-illuminated area of the fundus. 6. The method in accordance with claim 1, wherein for defining the measurement points in accordance with method step c), it is necessary for both the non-illuminated and also the imaged area of the fundus either to be known or detected. 7. The method in accordance with claim 6, further comprising, during the determination of the non-illuminated area of the fundus, taking into account the ametropia of the eye to be examined. 8. The method in accordance with claim 6, further comprising detecting an edge of the non-illuminated area of the fundus using image processing methods. 9. The method in accordance with claim 8, wherein the im... | Carl Zeiss Meditec AG | A61B003/14 A61B003/00 A61B003/12 A61B005/00 A61B003/117
US09763594 | B2 | 2017-09-19 | Magnetic particle imaging devices and methods | 2015-12-16 | A magnetic particle imaging device is provided. The device includes a magnetic field source configured to produce a magnetic field having a non-saturating magnetic field region, an excitation signal source configured to produce an excitation signal in the non-saturating magnetic field region that produces a detectable signal from magnetic particles in the non-saturating magnetic field region, and a signal processor configured to convert a detected signal into an image of the magnetic particles. Aspects of the present disclosure also include methods of imaging magnetic particles in a sample, and methods of producing an image of magnetic particles in a subject. The subject devices and methods find use in a variety of applications, such as medical imaging applications. | 1. A magnetic particle imaging device, comprising: a magnetic field generating apparatus arranged proximate an imaging region of the magnetic particle imaging device, the magnetic field generating apparatus being configured to produce a magnetic field within the imaging region of the magnetic particle imaging device such that the magnetic field will have a field-free region (FFR) for an object under observation that contains a magnetic tracer the magnetic field generating apparatus being further configured to produce an excitation magnetic field that induces a signal from the magnetic tracer in the object under observation;a receiver arranged proximate the observation region, the receiver being configured to receive the signal from the magnetic tracer in the observation region;a radio-frequency (RF) shield configured to isolate the receiver and the imaging region therein; anda signal processor configured to be in communication with the receiver, the signal processor being configured to convert the signal into an image of the magnetic tracer. 2. The device of claim 1, wherein the magnetic field generating apparatus comprises at least two coaxially arranged magnetic field sources isolated outside of the RF shield and which create the FFR. 3. The device of claim 2, wherein the at least two magnetic field sources are permanent magnets. 4. The device of claim 2, wherein the magnetic field generating apparatus produces a magnetic field gradient ranging from 0.5 Tesla/meter to 30 Tesla/meter. 5. The device of claim 1, wherein the excitation magnetic field comprises a RF magnetic field. 6. The device of claim 5, wherein the excitation magnetic field further comprises an intermodulation excitation signal. 7. The device of claim 1, wherein the magnetic field generating apparatus further comprises a scanning magnetic field source arranged proximate the observation region, the scanning magnetic field source being configured to produce a scanning magnetic field that positions the FFR in the imaging region. 8. The device of claim 7, wherein the device is configured to produce a linearly varying signal with respect to a concentration of the magnetic tracer in the FFR. 9. The device of claim 1, wherein the receiver comprises a receiver coil with a Q factor of 100 or more. 10. The device of claim 1, wherein the receiver is configured to have a receive bandwidth ranging from 10 kHz to 1 MHz. 11. The device of claim 1, wherein the excitation magnetic field comprises an RF frequency source arranged within the RF shield and a low-frequency source arranged outside of the RF shield, wherein the low-frequency source provides a component of the excitation magnetic field that has a frequency lower than the RF frequency source. 12. The device of claim 1, wherein the RF shield is cooled. 13. The device of claim 1, wherein the RF shield is cooled using a circulating coolant. 14. The device of claim 13, wherein the circulating coolant is water. 15. The device of... | The Regents of the University of California | G01R033/02 A61B005/05 G01R033/10 G01R033/12 G01N027/72 A61B005/00
US09332924 | B2 | 2016-05-10 | Posture detection using thoracic impedance | 2011-12-13 | Physiological data, such as thoracic impedance data, can be obtained over a first time window to establish a baseline, or can be used to form one or more data clusters. Additional physiological data, such as thoracic impedance test data acquired over a later time window, can be obtained and compared to the baseline or data clusters to determine an indication of worsening heart failure. In an example, a quantitative attribute of one or more data clusters can be monitored and used to provide an indication of worsening heart failure. A posture discrimination metric can be obtained, such as using the physiological data obtained over the first time window. The additional physiological data, such as can be obtained over a second time window, can be compared to the posture discrimination metric to provide a patient posture status. | 1. A medical device, comprising: a processor, including: a first data input, configured to receive first thoracic impedance data measured using a first electrode configuration defining a first thoracic impedance vector, the first thoracic impedance data corresponding to multiple instances over a first time window that includes a patient posture change, and wherein the first thoracic impedance data corresponds to multiple patient postures;a second data input, configured to receive second thoracic impedance data measured using a different second electrode configuration defining a second thoracic impedance vector, the second thoracic impedance data corresponding to multiple instances over the same first time window, and wherein the second thoracic impedance data corresponds to the multiple patient postures; anda processor-readable medium, including instructions that, when performed by the processor, configure the medical device to: form a first function of the first thoracic impedance data against the second thoracic impedance data, the first function having multiple output values including information about the multiple patient postures;identify first and second data clusters using the multiple output values of the first function, wherein each of the first and second data clusters respectively corresponds to a different deemed posture state, and wherein at least one of the first and second data clusters spans an at least two-dimensional space;form a posture discrimination comparison metric using the first and second data clusters and without prior knowledge of the patient's posture status;obtain third and fourth thoracic impedance data using the first data input and the second data input, respectively; anddetermine a posture status using the third and fourth thoracic impedance data and the posture discrimination comparison metric. 2. The medical device of claim 1, wherein the processor-readable medium includes instructions that, when performed by the processor, configure the medical device to: determine a quantitative attribute of the first function; andprovide the posture status using the quantitative attribute. 3. The medical device of claim 2, wherein the processor-readable medium includes instructions that, when performed by the processor, configure the medical device to determine the quantitative attribute of the first function, the quantitative attribute including at least one of: a spread of a portion of the first function;a range of a portion of the first function;an area defined at least in part by a portion of the first function;a location or centroid of a portion of the first function;a distance between the locations or centroids of at least two different portions of the first function; ora different second function formed using at least two different portions of the first function. 4. The medical device of claim 2, wherein the first data input is configured to receive the third thoracic impedance data measured using the first electro... | Cardiac Pacemakers, Inc. | A61B005/053 A61B005/0205 A61B005/11 A61B005/00 A61B005/103 A61N001/365
US09346567 | B2 | 2016-05-24 | Disposable vial holder and method to prevent needle stick injuries | 2014-12-11 | A vial holder is provided to protect a user's hand from needle sticks. The holder includes a handle and a shield attached to a distal end of the handle. An opening is formed in the distal end of the handle to receive a vial. The vial is placed through the opening and into a passageway of the handle with the upper end of the vial exposed. A user grasps the handle which holds the vial in a stabilized manner. A needle may then safely approach the vial in which inadvertent slippage or movement of the needle results in contact with the needle against the shield and not contact with the user's hand. | 1. A vial holder comprising: a handle having an opening formed at distal end thereof, and a passageway communicating with said opening, said passageway extending a depth within a portion of the handle from said distal end toward a proximal end of said handle;a shield attached to a distal end of the handle;a flange connected to and protruding from said opening, and said flange extending radially inward into said passageway;said passageway is tapered commencing at a larger diameter at said opening to a smaller diameter towards said proximal end of said handle; andwherein said vial holder is made as a single member. 2. A vial holder, as claimed in claim 1, wherein: said shield is made from a needle impenetrable plastic having a first stiffness, and said handle is made from a plastic having a second lesser stiffness. 3. A vial holder, as claimed in claim 1, wherein: said shield is substantially circular shaped. 4. A vial holder, as claimed in claim 1, wherein: said shield is substantially planar. 5. A vial holder comprising: a handle having an opening formed at distal end thereof, and a passageway communicating with said opening, said passageway extending a depth within a portion of the handle from said distal end toward a proximal end of said handle;a shield attached to a distal end of the handle;a flange connected to said shield and extending from said opening away from said handle; andsaid passageway is tapered commencing at a larger diameter at said opening to a smaller diameter towards said proximal end of said handle; andwherein said vial holder is made as a single member. 6. A vial holder, as claimed in claim 5, wherein: said shield is made from a needle impenetrable plastic having a first stiffness, and said handle is made from a plastic having a second lesser stiffness. 7. A vial holder, as claimed in claim 5, wherein: said shield is substantially circular shaped. 8. A vial holder, as claimed in claim 5, wherein: said shield is substantially planar. 9. In combination, a vial holder and vial comprising: a handle having a passageway extending at least partially through said handle, said handle further having a proximal end and a distal end;said passageway having an interior tapered surface commencing with a larger diameter at an opening of said passageway at said distal end of said handle, to a smaller diameter towards said proximal end of said handle;a shield attached to said distal end of said handle;a vial secured in said passageway, said vial having a base and a sidewall extending from said base; andwherein a lower portion of said sidewall contacts said interior surface of said passageway, and an upper portion of said sidewall is spaced from and not in contact with said interior surface of said passageway. 10. A method of preventing needle stick injuries during use of a needle in withdrawing contents of a vial, the method comprising: providing a vial holder including a handle having a passageway therein, the handle having a proxim... | National Jewish Health | A61J001/14 A61J001/16 B65B003/00 A61M005/32 A61B019/02 A61B017/34 G21F005/018 A61B005/15 A61M005/178 A61B019/00 A61J001/20
US09344815 | B2 | 2016-05-17 | Method for augmenting hearing | 2014-02-11 | One variation of a method for augmenting sound includes: at a mobile computing device and a connected audio output device, outputting a tone in a hearing test; based on a response to the tone entered by a user, generating a hearing profile for the user and corresponding to the mobile computing device and the audio output device; receiving an audio signal; qualifying the audio signal as a particular audio type from a set of audio types; selecting, from a set of sound profiles, a particular sound profile corresponding to the particular audio type; at the mobile computing device, transforming the audio signal into a processed audio signal according to the particular sound profile and the hearing profile of the user; and outputting the processed audio signal at the connected audio output device. | 1. A method for augmenting sound, comprising: at a mobile computing device and a connected audio output device, outputting a tone in a hearing test, wherein outputting the tone includes outputting a first audible tone comprising a first frequency, recording a first volume adjustment for the first audible tone by the user, outputting a second audible tone comprising a second frequency, recording a second volume adjustment for the second audible tone by the user;based on a response to the tone entered by a user, generating a hearing profile for the user and corresponding to the mobile computing device and the connected audio output device, wherein generating the hearing profile comprises selecting a particular hearing model from a set of hearing models based on a difference between the first volume adjustment and the second volume adjustment, each hearing model in the set of hearing models comprising a hearing test result corresponding to a previous patient, generating the hearing profile for the user based on the particular hearing model result, and associating the hearing profile with an identifier of the mobile computing device, associating the hearing profile with an identifier of the connected audio output device, and storing the hearing profile in an account assigned to the user;receiving an audio signal at the mobile computing device;qualifying the audio signal as a particular audio type from a set of audio types;selecting, from a set of sound profiles, a particular sound profile corresponding to the particular audio type;at the mobile computing device, transforming the audio signal into a processed audio signal according to both the particular sound profile and the hearing profile of the user; andoutputting the processed audio signal at the connected audio output device. 2. The method of claim 1, wherein transforming the audio signal into the processed audio signal comprises retrieving the hearing profile from the account assigned to the user upon collection of the identifier of the mobile computing device and collection of the identifier of the connected audio output device. 3. The method of claim 1, wherein receiving the audio signal comprises receiving an analog signal from a microphone arranged in the connected audio device comprising a set of headphones, and wherein outputting the processed audio signal comprises outputting the processed audio signal through a speaker arranged in the set of headphones. 4. The method of claim 1, wherein qualifying the audio signal as the particular audio type comprises identifying the audio signal as an audio component of a phone call between the mobile computing device and a secondary device, wherein transforming the audio signal into the processed audio signal comprises attenuating a portion of the audio signal corresponding to a voice of the user and amplifying a portion of the audio signal received from the secondary device according to the particular sound profile. 5. The method of claim 1, fur... | Symphonic Audio Technologies Corp. | H04R025/00 A61B005/00 A61B005/12 H04M001/725 H04R029/00
US09326733 | B2 | 2016-05-03 | System and method for acquisition of biopotential signals with motion artifact reduction in real time operation | 2015-02-05 | A biopotential signal acquisition system comprising an analog readout unit configured to receive an analog biopotential signal and to extract an analog measured biopotential signal and an analog reference signal, and an ADC unit configured to provide a digital version of the analog measured biopotential signal and the analog reference signal. The system also includes a first digital filter unit comprising a cascaded integrator-comb filter configured to provide a first digital filtered version of the digital measured biopotential signal and the reference signal, and a second digital filter unit configured to calculate a digital motion artifact estimate based on the first digital filtered version signals. | 1. A biopotential signal acquisition system comprising: an analogue readout unit configured to receive an analogue biopotential signal and to extract an analogue measured biopotential signal and an analogue reference signal;a Sigma-Delta ADC unit configured to provide a digital measured biopotential signal from the analogue measured biopotential signal and a digital reference signal from the analogue reference signal;a first digital filter unit connected to an output of the Sigma-Delta ADC, and including a cascaded integrator-comb filter and a halfband decimation filter connected to the cascaded integrator-comb filter, wherein the first digital filter unit is configured to provide a first digital filtered version of the digital measured biopotential signal and the digital reference signal at an output of the cascaded integrator-comb filter, and a second digital filtered version of the digital measured biopotential signal and the digital reference signal at an output of the first digital filter unit; anda second digital filter unit connected to the output of the cascaded integrator-comb filter of the first digital filter unit, and configured to calculate a digital motion artifact estimate based on the first digital filtered version of the digital measured biopotential signal and the digital reference signal. 2. The biopotential signal acquisition system according to claim 1, wherein the second digital filter unit utilizes a digital adaptive filter. 3. The biopotential signal acquisition system according to claim 2, wherein the cascaded integrator-comb filter has a sampling frequency higher than a frequency band of the digital adaptive filter. 4. The biopotential signal acquisition system according to claim 2, wherein the digital adaptive filter is a Least Mean Square filter. 5. The biopotential signal acquisition system according to claim 1, wherein the analogue reference signal is an electrode-tissue impedance signal. 6. The biopotential signal acquisition system according to claim 1, wherein the analogue biopotential signal is an ECG signal. 7. The biopotential signal acquisition system according to claim 1, wherein the system is implemented on a single integrated circuit. 8. The biopotential signal acquisition system according to claim 1, wherein the calculated digital motion artifact estimate is converted to an analogue signal and fed to the analogue readout unit for motion artifact reduction. 9. The biopotential signal acquisition system according to claim 1, wherein the first digital filter unit further comprises a finite impulse response filter connected to the halfband decimation filter. 10. An electronic device comprising a biopotential signal acquisition system according to claim 1. 11. A method for acquisition of biopotential signals in a system according to claim 1, comprising: extracting, from the analogue biopotential signal, the analogue measured biopotential signal and the analogue reference signal;converting the analogue... | IMEC VZW | A61B005/04 A61B005/00 A61B005/0402 A61B005/0476 A61B005/0488
US09763619 | B2 | 2017-09-19 | Methods and devices for accurately classifying cardiac activity | 2016-01-15 | Methods, systems, and devices for signal analysis in an implanted cardiac monitoring and treatment device such as an implantable cardioverter defibrillator. In some examples, captured data including detected events is analyzed to identify likely overdetection of cardiac events. In some illustrative examples, when overdetection is identified, data may be modified to correct for overdetection, to reduce the impact of overdetection, or to ignore overdetected data. Several examples emphasize the use of morphology analysis using correlation to static templates and/or inter-event correlation analysis. | 1. An implantable cardiac stimulus (ICS) system comprising: a canister;a plurality of electrodes exposed external to the canister configured for implantation in a patient;operational circuitry disposed within the canister, operatively coupled to the plurality of electrodes, and configured to: capture cardiac signals occurring within a patient using one or more of the plurality of electrodes;detect electrical events in the captured cardiac signals;identify an event representation for one of the detected electrical events;compare the event representation to a template having a template fiducial point by identifying an event fiducial point of the event representation and aligning the template fiducial point and the event fiducial point to one another and calculating a first correlation between the event representation and the template when so aligned;again compare the event representation to the template, this time by selecting a misalignment of the template fiducial point and the event fiducial point to yield a second correlation;select a correlation result for the one of the detected electrical events from at least the first correlation or the second correlation whichever indicates greater correlation; anduse the correlation result to determine whether the patient has a cardiac condition which requires therapy; anda power source disposed within the canister and configured to power the operational circuitry;wherein the event fiducial point is a selected sample of the event representation, and the template fiducial point is a selected sample of the template. 2. The ICS of claim 1 wherein the operational circuitry is further configured to deliver therapy to the patient using the one or more of the plurality of electrode if, using the correlation result, the operational circuitry determines the patient has a cardiac condition that requires therapy. 3. The ICS of claim 1 wherein the operational circuitry is configured to use the correlation result to determine whether the patient has a cardiac condition which requires therapy by: calculating a plurality of calculation results using respective first and second correlation results for a plurality of detected electrical events;identifying likely inaccurate event detection;determining that the correlation result is lower than other, similarly calculated correlation results; anddetermining that the detected electrical event corresponding to the correlation result was an overdetection. 4. The ICS of claim 3 wherein the operational circuitry is configured to identify likely inaccurate event detection by observing whether the plurality of calculation results demonstrate a pattern indicating overdetection. 5. The ICS of claim 1 wherein the operational circuitry is configured to use the correlation result to determine whether the patient has a cardiac condition which requires therapy by comparing a series of correlation results to determine whether a monomorphic condition is occurring. 6. The ICS of claim 5... | Cameron Health, Inc. | A61B005/04 A61B005/00 A61B005/0468 A61B005/0452 A61N001/37 A61N001/365 A61B005/042 A61B005/046 A61B005/0464 A61N001/362 A61N001/39
US09339084 | B2 | 2016-05-17 | Time prediction system for the safe wearing of newly acquired footwear | 2012-01-25 | The device relates to a system for predicting the period for children to safely wear newly acquired footwear without the risk of damage to their growing feet, which is directly applicable in shoe shops (at their computer cash registers). The essence of the solution for a system for predicting the period for children to safely wear newly acquired footwear rests in the fact that it contains an input module measuring foot length and input of information about the age of the monitored child to which is linked a model for predicting the growth of the child's foot by applying the laws of growth, possibly also including genetic and local influences, connected to a comparative and inferential module establishing the predicted course of growth in foot length of the monitored child, to which is then linked an output module designating the nearest date for the necessary replacement of shoe size for the monitored child. | 1. A prediction system of the period for safely wearing newly acquired footwear for children without the risk of damaging their growing feet, the system comprising: an input module measuring sole length and input of information about the age of a monitored child and a date on which shoes were purchased for the monitored child, to which module is linked a prediction module of the growth of a child's feet implementing the laws of growth on the basis of the Carlberg growth model, including an empirical model originating from an analysis of child's foot length in connection with the age of children, performed on a group of at least 2000 children connected to a comparative and inferential module establishing a predicted course of sole length growth of the monitored child to which then is linked an output module designating a date when it will be necessary to stop using the purchased shoes. 2. The prediction system defined in claim 1, wherein the empirical model for growth of a child's foot length by the prediction module includes the foot length of the child's parents as another parameter. 3. The prediction system defined in claim 1, wherein the input module measuring foot size and the input of information about the age of the monitored child contains a submodule for data storage which is connected to the prediction module for growth of the child's feet. 4. The prediction system defined in claim 1, wherein the empirical model for the growth of children's feet implemented in the prediction module is continuously revised by the output submodule with genetic influences, the input of which include in particular the length parameters of the children's parents. 5. The prediction system defined in claim 1 wherein the growth model for children's feet implemented in the prediction module is continuously revised by inputs to the submodule of local influences connected to the submodule for data storage of the input module. 6. The prediction system defined in claim 1, wherein additional output of the comparative and inferential module establishing the predicted course of foot length growth of the monitored child are data about the lengths of the child's parent's feet. | UNIVERZITA TOMASE BATI VE ZLINE | A43D001/02 A61B005/107 A43B003/30 G06Q030/06 G06Q010/04 G06Q030/02
US09326716 | B2 | 2016-05-03 | Analyte monitoring device and methods of use | 2014-12-05 | In aspects of the present disclosure, a no coding blood glucose monitoring unit including a calibration unit is integrated with one or more components of an analyte monitoring system to provide compatibility with in vitro test strip that do not require a calibration code is provided. Also disclosed are methods, systems, devices and kits for providing the same. | 1. An apparatus, comprising: an analyte sensor interface configured to receive an in vitro analyte sensor, wherein the in vitro analyte sensor generates an analyte signal from a fluid sample; anda processor operatively coupled to a data receiver and the analyte sensor interface, the processor configured to recognize the in vitro analyte sensor without a user input when the in vitro analyte sensor is received in the analyte sensor interface and is a no-coding in vitro analyte sensor, and to retrieve a universal calibration parameter when the recognized in vitro analyte sensor is the no-coding in vitro analyte sensor. 2. The apparatus of claim 1, wherein the analyte sensor interface comprises a first contact structure and a second contact structure configured to contact the in vitro analyte sensor. 3. The apparatus of claim 1, wherein the in vitro analyte sensor is a calibration-adjusted in vitro analyte sensor. 4. The apparatus of claim 1, wherein the universal calibration parameter is stored in a memory operatively coupled to the processor, and includes a universal calibration code. 5. The apparatus of claim 4, wherein the universal calibration parameter stored in the memory is a standard calibration code, and the in vitro analyte sensor is manufactured to fit the standard calibration code. 6. The apparatus of claim 1, further including a data receiver operatively coupled to the processor and configured to receive data relating to analyte levels from an analyte sensor in fluid contact with interstitial fluid. 7. The apparatus of claim 6, wherein the analyte sensor comprises a plurality of electrodes including a working electrode, wherein the working electrode comprises an analyte-responsive enzyme and a mediator, and wherein at least one of the analyte-responsive enzyme and the mediator is chemically bonded to a polymer disposed on the working electrode. 8. The apparatus of claim 7, wherein the at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer. 9. The apparatus of claim 1, wherein the universal calibration parameter is a standard calibration code, and the in vitro analyte sensor is manufactured to fit the standard calibration code. 10. The apparatus of claim 6, wherein the analyte sensor is configured to monitor one or more of a glucose level, lactate level, or oxygen level. 11. An apparatus, comprising: an in vitro analyte sensor interface;one or more processors operatively coupled to the in vitro analyte sensor interface; anda memory operatively coupled to the one or more processors, the memory for storing instructions which, when executed by the one or more processors, causes the one or more processors to determine a type of in vitro analyte sensor received at the in vitro analyte sensor interface, to determine the type of in vitro analyte sensor without a user input when the in vitro analyte sensor is a no-coding in vitro analyte sensor, and to process one or more signals generated by the... | Abbott Diabetes Care Inc. | G01N033/00 A61B005/1495 A61B005/145 A61B005/1486 A61B005/00 A61B005/1477 A61B005/1473 G01N027/327
US09330455 | B2 | 2016-05-03 | Diagnostic support apparatus and diagnostic support method | 2013-10-30 | The diagnostic support apparatus includes a feature quantity calculation unit vectorizing a test image into shift-invariant feature quantities, a base representation unit transforming the shift-invariant feature quantities of the test image into a linear combination of base vectors with first coefficients, the base vectors being calculated from a plurality of vectors representing shift-invariant feature quantities of a plurality of normal structural images including no lesion site, a lesion determination unit determining that the test image includes a lesion site when a difference between the first coefficients and second coefficients is greater than a determination threshold value, the second coefficients being used to transform shift-invariant feature quantities calculated from one of the normal structural images into a linear combination of the base vectors, and an output unit outputting a result of the determination by the lesion determination unit. | 1. A diagnostic support apparatus comprising: a non-transitory memory that stores a program; anda hardware processor that executes the program and causes the diagnostic support apparatus to operate as:a feature quantity calculation unit configured to vectorize a test image into shift-invariant feature quantities, the test image being an image in which presence of a lesion site is unknown;a base representation unit configured to transform the shift-invariant feature quantities of the test image into a linear combination of base vectors with first coefficients, the base vectors being represented by eigen vectors calculated through principal component analysis of a plurality of vectors representing shift-invariant feature quantities of a plurality of normal structural images that include no lesion site;a storage unit in which second coefficients are stored, the second coefficients being used to transform shift-invariant feature quantities calculated from one of the plurality of normal structural images into a linear combination of the base vectors;a lesion determination unit configured to determine that the test image includes a lesion site when a difference between the first coefficients obtained from a result of the transforming by the base representation unit and the second coefficients stored in the storage unit is greater than a determination threshold value; andan output unit configured to output a result of the determination by the lesion determination unit. 2. The diagnostic support apparatus according to claim 1, wherein the feature quantity calculation unit is configured to calculate a plurality of shift-invariant feature quantities from a pixel of interest in the test image and calculate a test image feature quantity vector having the calculated shift-invariant feature quantities as vector elements,the base representation unit is configured to calculate the first coefficients and calculate a test coefficient vector having the calculated first coefficients as vector elements, the first coefficients being used to represent the test image feature quantity vector as the linear combination of the base vectors, the base vectors being calculated from a plurality of image feature quantity vectors each having, as vector elements, shift-invariant feature quantities calculated from pixels in each of the normal structural images, the normal structural images each being a medical image of a normal structure,the diagnostic support apparatus further comprisinga nearest neighbor vector detection unit configured to detect a normal coefficient vector that is most similar to the test coefficient vector as a nearest neighbor vector from among a plurality of normal coefficient vectors each having second coefficients as vector elements, the second coefficients being used to represent each of the image feature quantity vectors as a linear combination of the base vectors, andthe lesion determination unit is configured to compare a distance between the test coef... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | G06T007/00 A61B006/00 A61B005/055 A61B006/03 A61B008/08
US09352165 | B2 | 2016-05-31 | Method and apparatus for verifying discriminating of tachycardia events in a medical device having dual sensing vectors | 2014-04-17 | A method and medical device for detecting a cardiac event that includes sensing cardiac signals from a plurality of electrodes, the plurality of electrodes forming a first sensing vector sensing a first interval and a second sensing vector simultaneously sensing a second, determining, for each of the first interval and the second interval, whether each beat of the plurality of beats is one of a match beat and a non-match beat, determining whether each beat of the plurality of beats is one of a high confidence beat and a low confidence beat, determining, for each of the first interval and the second interval, the number of beats determined to be both a non-match beat and a high confidence beat is greater than a non-match threshold, and determining whether to deliver therapy for the cardiac event in response to identifying of each of the first interval and the second interval as being one of shockable and not shockable. | 1. A method of detecting a cardiac event in a medical device, comprising: sensing cardiac signals from a plurality of electrodes, the plurality of electrodes forming a first sensing vector sensing a first interval of the cardiac signal comprising a plurality of beats during a predetermined time period and a second sensing vector simultaneously sensing a second interval of the cardiac signal comprising a plurality of beats during the predetermined time period;determining, for each of the first interval and the second interval, whether each beat of the plurality of beats is one of a match beat and a non-match beat;determining whether each beat of the plurality of beats is one of a high confidence beat and a low confidence beat;determining, for each of the first interval and the second interval, the number of beats determined to be both a non-match beat and a high confidence beat is greater than a non-match threshold;identifying each of the first interval and the second interval as being one of shockable and not shockable in response to the number of beats identified as being both a non-match beat and a high confidence beat; anddetermining whether to deliver therapy for the cardiac event in response to identifying of each of the first interval and the second interval as being one of shockable and not shockable. 2. The method of claim 1, wherein determining whether each beat of the plurality of beats is one of a match beat and a non-match beat comprises: determining a fourth order difference signal corresponding to the beat;determining a template alignment point and a beat alignment point in response to the fourth order difference signal;aligning the template and the beat in response to the template alignment point and the beat alignment point; anddetermining a morphology match metric between the aligned template and the beat. 3. The method of claim 2, wherein determining whether each beat of the plurality of beats is one of a high confidence beat and a low confidence beat comprises: determining a number of pulses associated with the beat being less than a narrow pulse threshold; andcomparing the number of pulses associated with the beat being less than a narrow pulse threshold to a narrow pulse count threshold. 4. The method of claim 3, wherein determining a number of pulses associated with the beat being less than a narrow pulse threshold comprises determining, for each pulse associated with a beat, whether both a pulse width is less than a pulse width threshold and a pulse amplitude is greater than a pulse amplitude threshold. 5. The method of claim 2, wherein determining whether each beat of the plurality of beats is one of a match beat and a non-match beat comprises: comparing a slope of the fourth order difference signal to a maximum slope threshold;determining pulses associated with the beat in response to the comparing;determining a pulse amplitude for each pulse; anddetermining a maximum pulse amplitude in response to the determined ampl... | Medtronic, Inc. | A61B005/04 A61N001/38 A61N001/39
US10258748 | B2 | 2019-04-16 | Vein scanner with user interface for controlling imaging parameters | 2015-12-22 | A portable handheld vein-image-enhancing device broadly includes: a first laser emitting a beam of light at a first wavelength; a second laser emitting a beam of light at a second wavelength; a scanner to scan the beams of light onto the target surface; a photo detector to receive the reflected vein image and convert the image into a signal, for use by the second laser and scanner to project the image onto the target surface; and a user interface and electronic circuitry to permit adjustments to one or more of the following imaging parameters: a vein size parameter to set a vein size to be imaged by the device; a field of view size; a size for a field of high resolution for the projected image; a brightness of the projected image; and a laser output intensity parameter setting a depth of penetration by the first wavelength for the imaging. | 1. A portable handheld vein-image-enhancing device, for use in selective imaging of subcutaneous veins of a target surface, said device comprising: a first laser configured to emit a beam of light at a first wavelength;a second laser configured to emit a beam of light at a second wavelength being different than said first wavelength;a scanner configured to scan said beams of said first and second wavelengths, of light onto the target surface;a photo detector configured to receive said first wavelength of light reflected from the target surface as a contrasted vein image formed by differential absorption and reflection by subcutaneous veins and, surrounding tissue therein; said photo detector configured to convert said contrasted image into a signal; said second laser configured to receive said signal, and said second laser and said scanner thereby configured to project said contrasted vein image onto the target surface using said, second wavelength of light, tor the veins of said contrasted vein image to be projected onto the target surface directly over the imaged subcutaneous veins;a user interface and electronic circuitry configured to permit adjustment of one or more imaging parameters for said selective imaging by said device;wherein said one or more imaging parameters comprises a vein size parameter to set a vein size segment to be imaged by said device. 2. The portable handheld vein-image-enhancing device according to claim 1, further comprising a memory, and wherein said user interface and electronic circuitry is further configured to permit said adjusted one or more imaging parameters to be stored in said memory. 3. The portable handheld vein-image-enhancing device according to claim 1, wherein one or more imaging parameters further comprises a field of view parameter to set a size of the field of view to be imaged by said device. 4. The portable handheld vein-image-enhancing device according to claim 1, wherein one or more imaging parameters further comprises a resolution parameter to set a size of a field of high, resolution for said projected image. 5. The portable handheld vein-image-enhancing device according to claim 1, wherein one or more imaging parameters further comprises a brightness projection parameter to set a brightness of said projected image by said second laser. 6. The portable handheld vein-image-enhancing device according to claim 1, wherein one or more imaging parameters further comprises a laser output intensity parameter to set a depth to be penetrated by said first laser for Said imaging by said device. 7. The portable handheld vein-image-enhancing device according to claim 1, wherein said scanner comprises one or more mirrors configured to oscillate in two orthogonal directions. 8. The portable handheld vein-image-enhancing device according to claim 1, wherein said first wavelength of light comprises an infrared wavelength of light, and said second wavelength of light comprises a visible red wavelength of ... | AccuVein, Inc. | A61M005/42 A61B005/00 A61B005/153 A61B005/1172 A61B005/1171 A61B005/15 A61B090/98 A61B005/024
US09591991 | B2 | 2017-03-14 | Position detection system | 2016-01-28 | A position detection system for detecting a position of a capsule medical device is provided. The capsule medical device has a coil constituting a part of a resonance circuit to generate an AC magnetic field. The position detection system includes: sensing coils that detect the AC magnetic field and output detection signals; and a calculation unit that calculates the position of the capsule medical device based on the detection signals. The calculation unit: determines a sensing coil having a maximum amplitude of the AC magnetic field as a reference coil among the sensing coils based on amplitude values of the AC magnetic field detected by the sensing coils, the amplitude values being calculated based on the detection signals; and determines a polarity of each sensing coil based on a first detection signal from the reference coil and a second detection signal from each sensing coil other than the reference coil. | 1. A position detection system for detecting a position of a capsule medical device, the capsule medical device having: a coil constituting a part of a resonance circuit to generate an AC magnetic field by flowing of current; a power source for supplying power to the coil; and a capsule casing, the coil and the power source being provided inside the casing, the position detection system comprising: a plurality of sensing coils configured to detect the AC magnetic field generated by the coil and to output detection signals; anda processor comprising hardware, the processor being configured to: calculate the position of the capsule medical device based on the plurality of detection signals output from the plurality of sensing coils,determine a first sensing coil having a maximum amplitude value of the AC magnetic field as a reference coil among the plurality of sensing coils based on amplitude values of the AC magnetic field detected by the plurality of sensing coils, the amplitude values being calculated based on the plurality of detection signals; anddetermine a polarity of each of the sensing coils based on a first detection signal output from the reference coil and a second detection signal output from a second sensing coil other than the reference coil among the plurality of sensing coils. 2. The position detection system according to claim 1, wherein the processor is further configured to determine the polarity of each of the sensing coils based on a comparison result between a phase of the first detection signal and a phase of the second detection signal. 3. The position detection system according to claim 2, wherein if a phase difference between the first detection signal and the second detection signal is equal to or smaller than ¦π/2¦, the processor is further configured to determine that the polarity of the second sensing coil is equal to the polarity of the reference coil, andif a phase difference between the first detection signal and the second detection signal is larger than¦π/2¦, the processor is further configured to determine that the polarity of the second sensing coil is opposite to the polarity of the reference coil. 4. The position detection system according to claim 3, wherein the processor is further configured to: calculate the position of the capsule medical device at predetermined time intervals, andset the polarity of the first sensing coil determined as the reference coil to be the same polarity as a determination result for the first sensing coil obtained by previous determination of the polarity. 5. The position detection system according to claim 4, wherein if the determination result for the first sensing coil is not obtained by the previous determination of the polarity, the processor is further configured to set the polarity of the first sensing coil determined as the reference coil to be positive. 6. The position detection system according to claim 1, wherein the processor is further configured to correct a ... | OLYMPUS CORPORATION | A61B005/00 A61B005/06 A61B001/04
US10368936 | B2 | 2019-08-06 | Systems and methods for selecting, activating, or selecting and activating transducers | 2015-11-16 | A graphical path, which may illustrate a proposed transducer-activation path, may be defined in accordance with first user input, which places a first user input element in an activated state, motion-based user input, and second user input, which places the first user input element in a deactivated state. The graphical path may be displayed as including transducer graphical elements, between-transducer graphical elements, or both, and may be displayed among at least one graphical representation of intra-cardiac information. Activation of transducers corresponding to transducer graphical elements included in the graphical path may be initiated during or after completion of the definition of the graphical path. Such activation may include the transmission of energy sufficient for tissue ablation along an ablation path corresponding to the graphical path. | 1. A transducer activation system comprising: a data processing device system;an input-output device system communicatively connected to the data processing device system; anda memory device system communicatively connected to the data processing device system and storing a program executable by the data processing device system, the program comprising:display instructions configured to cause the input-output device system to display a graphical representation of at least intra-cardiac information;input-processing instructions configured to:cause reception of first user input via the input-output device system and, in response to receiving the first user input, place a first user input element in an activated state;cause reception of second user input via the input-output device system and, in response to receiving the second user input, place the first user input element in a deactivated state; andcause reception of motion-based user input via the input-output device system;path definition instructions configured to cause definition of a graphical path including a first location on the graphical path defined according to a first parameter set associated with the first user input, a second location on the graphical path defined according to a second parameter set associated with the second user input, and an elongate path portion of the graphical path defined according to a path traced by the motion-based user input; andactivation instructions configured to cause activation of a transducer-based device system, initiated during or after completion of the definition of the graphical path, to transmit energy sufficient for tissue ablation along an ablation path corresponding to the graphical path,wherein the display instructions are configured to cause the input-output device system to display the graphical path including each of the first location, the second location, and the elongate path portion among the graphical representation of the intra-cardiac information. 2. The transducer activation system of claim 1 wherein the program further comprises: sampling instructions configured to cause sampling of data by each of one or more transducers of the transducer-based device system, a portion of the transducer-based device system including the one or more transducers positionable in a cardiac chamber during the sampling; andgeneration instructions configured to cause generation of the intra-cardiac information based at least in part on the sampled data, the sampled data sampled from each of a plurality of locations in the cardiac chamber, and whereinthe generation instructions are configured to cause mapping of each of a plurality of parts of the intra-cardiac information to a respective one of the plurality of locations in the cardiac chamber, and the display instructions are configured to cause the input-output device system to display the plurality of parts of the intra-cardiac information with a first spatial relationship that is consistent with... | KARDIUM INC. | A61B018/12 A61B019/00 A61B005/00 A61B005/01 A61B005/042 A61B018/00 A61B018/14 A61B034/10
US20160143589 | A1 | 2016-05-26 | BIOLOGICAL MEASUREMENT DEVICE AND BIOLOGICAL MEASUREMENT METHOD | 2015-10-26 | A biological measurement device includes a phase modulation type interferometer, a first stage, a light transmitting plate, and a second stage. The first stage is able to retain a target which is a biological body. The light transmitting plate is disposed between the first stage and the interferometer in a direction of an optical axis of measurement light of the interferometer. The second stage supports the light transmitting plate. The first stage is configured to expand by a pressure of a gaseous body in the direction of the optical axis of the measurement light of the interferometer and to press the target against the light transmitting plate. | 1. A biological measurement device, comprising: a phase modulation type interferometer;a first stage capable of retaining a target which is a biological body;a light transmitting plate disposed between the first stage and the interferometer in a direction of an optical axis of measurement light of the interferometer; anda second stage supporting the light transmitting plate,wherein the first stage is configured to expand by a pressure of a gaseous body in the direction of the optical axis of the measurement light of the interferometer, and to press the target against the light transmitting plate. 2. The biological measurement device according to claim 1, wherein rigidity of the first stage is lower than rigidity of the second stage. 3. The biological measurement device according to claim 1, wherein the first stage is a balloon which expands by the gaseous body. 4. The biological measurement device according to claim 1, wherein the first stage retains the target in a state that the target is opened in a direction orthogonal to the optical axis of the measurement light of the interferometer. 5. The biological measurement device according to claim 1, wherein the light transmitting plate is an optical window capable of transmitting the measurement light from the interferometer, andthe first stage is capable of sandwiching the target between the first stage and the optical window. 6. The biological measurement device according to claim 5, wherein the optical window protrudes from the second stage towards the first stage. 7. The biological measurement device according to claim 1, wherein the first stage is capable of deforming in response to a volume fluctuation caused by biological activity of the target. 8. The biological measurement device according to claim 1, wherein the target is a human body, andthe pressure of the gaseous body is 30 mmHg or greater and 80 mmHg or less. 9. The biological measurement device according to claim 1, wherein the gaseous body is air. 10. The biological measurement device according to claim 1, wherein the second stage is a jig including an upper plate, a lower plate and a plurality of support portions connecting the upper plate to the lower plate,the second stage is disposed on a mechanical stage configured to move in three perpendicular directions in accordance with control signals of a controller,the first stage is disposed on an upper surface of the lower plate of the second stage, andthe light transmitting plate protrudes from a lower surface of the upper plate of the second stage. 11. A biological measurement method for measuring a target by using a biological measurement device including a phase modulation type interferometer, a first stage, a light transmitting plate disposed between the first stage and the interferometer in a direction of an optical axis of measurement light of the interferometer, and a second stage supporting the light transmitting plate, the method comprising: retaining the target wi... |  | A61B005/00 G01B011/14 G01B009/02
US20160128562 | A1 | 2016-05-12 | Apparatus and Method of Determining an Eye Prescription | 2014-07-02 | Eye prescriptions may be determined by providing a simple, easy to use, portable device with a specially configured targeting light source that aligns the eye, mitigates accommodation, and provides accurate results. Unlike stationary, closed view autorefractors, this device typically is portable, self-usable, relatively inexpensive, enabling more widespread use across the world. | 1-49. (canceled) 50. An optical method comprising: providing an optical apparatus having a proximal port and a distal port that together form a visual channel, the distal port at least in part defining an optical axis;aligning a living being's eye with the proximal port, the eye viewing through the distal port to target indicia exterior of the apparatus;determining orientation of the eye relative to the optical axis; andgenerating a cue as a function of the orientation of the eye relative to the optical axis, the cue being perceptible to a human being. 51. The optical method as defined by claim 50, wherein the cue includes at least one of a visual cue, an acoustic cue and a mechanical cue. 52. The optical method as defined by claim 50, wherein generating the cue comprises producing a sound that changes in amplitude, as a function of the orientation of the eye relative to the optical axis. 53. The optical method as defined by claim 50, wherein generating the cue comprises producing a sound that changes in frequency, as a function of the orientation of the eye relative to the optical axis. 54. The optical method as defined by claim 50, wherein generating the cue comprises producing a sound that changes in duty cycle, as a function of the orientation of the eye relative to the optical axis. 55. The optical method as defined by claim 50, wherein determining the orientation of the eye comprises using wavefront aberrometry techniques. 56. The optical method as defined by claim 50, wherein: the optical apparatus comprises a plurality of target light sources, each target light source being configured to selectively generate the target indicia; the method further comprising:alternatively illuminating one or more of the target light sources to direct the eye toward alignment with the optical axis. 57. The optical method as defined by claim 50, wherein: the optical apparatus comprises a target light source that is configured to produce a light beam for forming the target indicia; the method further comprising:moving the target light source to direct the eye toward alignment with the optical axis. 58. The optical method as defined by claim 50, wherein the eye moves toward alignment with the optical axis in response to generation of a cue. 59. An optical apparatus for measuring ocular aberrations of an eye of a living being, the apparatus comprising: a proximal port and a distal port that together form a visual channel from the proximal port through the distal port, the distal port at least in part defining an optical axis;a cue generator configured to generate a cue perceptible to a human being;orientation logic operatively coupled to the cue generator and the visual channel and configured to determine orientation of the living being's eye relative to the optical axis and to cause operation of the cue generator as a function of the orientation. 60. The optical apparatus as defined by claim 59, wherein the cue generator is configured to produce at ... |  | A61B003/103 A61B005/00 A61B003/18 A61B003/00 A61B003/10
US20160120451 | A1 | 2016-05-05 | Cartridge | 2016-01-13 | A cartridge for insertion into a meter comprises: a plurality of cartridge brackets disposed on an inner wall of the cartridge; a spindle mounted so as to be rotatable within the cartridge and movable longitudinally in first and second directions within the cartridge; and a plurality of testing members each arranged to be supported temporarily by at least one of the plurality of cartridge brackets, each of the testing members including a hole through which the spindle is located, the spindle having a plurality of spindle brackets disposed on an outer surface thereof, each spindle bracket being movable between first and second positions wherein in the first position the spindle bracket's distal end is at a first distance radially from the longitudinal axis of the spindle and wherein in the second position the spindle bracket's distal end is at a greater distance radially from the longitudinal axis of the spindle, and wherein when a given one of the spindle brackets is in the second position its distal end is located against a surface of one of the testing members so as to prevent movement of the testing member in the second longitudinal direction relative to the spindle. When the spindle advances in the second longitudinal direction relative to the cartridge, the spindle brackets are configured to move from the second to the first position so as to allow the spindle brackets to move past the testing members in the second longitudinal direction relative to the testing members, the spindle brackets being configured to return to the second position after moving past the testing members such that each spindle bracket is located against a surface of a different one of the testing members to the testing member against which it was originally located and such that the spindle brackets communicate force from the spindle to the testing members in the first longitudinal direction to move the testing members in the first longitudinal direction as the spindle is moved in the first longitudinal direction. | 1.-15. (canceled) 16. An apparatus, comprising: a cartridge having a cavity;a plurality of cartridge brackets coupled to an inner wall of the cartridge, wherein the plurality of cartridge brackets have the form of protrusions that extend into the cavity of the cartridge;a spindle mounted within the cartridge and movable longitudinally in a first direction and a second direction within the cartridge;a plurality of spindle brackets coupled to the spindle, wherein the plurality of spindle brackets have the form of protrusions that extend into the cavity of the cartridge; andwherein the plurality of cartridge brackets are configured to move in the first direction in response to the spindle moving in the first direction and configured to resist movement in the second direction in response to the spindle moving in the second direction, and wherein the plurality of spindle brackets are configured to move in a first direction in response to the spindle moving in the second direction and configured to resist movement in a second direction in response to the spindle moving in the first direction. 17. The apparatus of claim 16, wherein the spindle is rotatable. 18. The apparatus of claim 16, further comprising: a plurality of testing members each arranged to be supported temporarily by at least one of the plurality of cartridge brackets. 19. The apparatus of claim 18, wherein each of the plurality of testing members includes a hole through which the spindle is located. 20. The apparatus of claim 18, wherein at least a distal end of each cartridge bracket may be configured to deform towards the inner wall of the cartridge so as to allow movement of the plurality of testing members in the first direction with respect to the cartridge. 21. The apparatus of claim 18, wherein, when the spindle moves in the second direction, a distal end of each of the plurality of cartridge brackets is located against a lower surface of one of the plurality of testing members so as to prevent movement of the testing member in the second direction relative to the spindle. 22. The apparatus of claim 18, wherein when the spindle moves in the second direction, the plurality of spindle brackets are configured to move past the plurality of testing members in the second direction relative to the plurality of testing members. 23. The apparatus of claim 7, wherein after moving past the plurality of testing members, each of the plurality of spindle brackets is located against a lower surface of a different one of the plurality of testing members and such that the plurality of spindle brackets communicate force from the spindle to the testing members in the first direction to move the testing members in the first direction as the spindle is moved in the first direction. 24. The apparatus of claim 16, wherein the cartridge comprises a biasing arrangement configured to bias the spindle in the first direction. 25. The apparatus of claim 16, wherein the cartridge comprises a spindle ... |  | A61B005/151 G01N033/66 A61B005/157 A61B005/15 A61B005/145
US20160135752 | A1 | 2016-05-19 | SMALL DATA AGGREGATOR FOR PERSONAL HEALTH MANAGEMENT | 2014-11-19 | One embodiment provides a method, including: identifying, using a processor of an electronic device, at least one user associated with a user of the electronic device; receiving, using a network component of the electronic device, health-related information associated with the at least one user; generating, using a processor of the electronic device, a notification based on the health-related information; and providing, using an output device of the electronic device, the notification to the user of the electronic device. Other aspects are described and claimed. | 1. A method, comprising: identifying, using a processor of an electronic device, at least one user associated with a user of the electronic device;receiving, using a network component of the electronic device, health-related information associated with the at least one user;generating, using a processor of the electronic device, a notification based on the health-related information; andproviding, using an output device of the electronic device, the notification to the user of the electronic device. 2. The method of claim 1, wherein the at least one user is identified using data selected from the group consisting of a contact list of the electronic device, a calendar appointment of the electronic device, and geographic location data. 3. The method of claim 1, wherein the notification comprises an indication that the at least one user is ill. 4. The method of claim 1, further comprising: sensing, using a sensor, health-related data associated with the user of the electronic device; andcommunicating, using a network component of the electronic device, the health-related data associated with the user of the electronic device to a remote device. 5. The method of claim 4, wherein the health-related data associated with the user is selected from the group consisting of activity level data, meeting attendance data, sleep cycle data, data derived from a wearable device, and dietary data. 6. The method of claim 4, further comprising storing, using a memory of the electronic device, the health-related data associated with the user of the electronic device. 7. The method of claim 4, further comprising: receiving, using an input device of the electronic device, reporting configuration data; andconfiguring, using a processor, the communicating based on the reporting configuration data. 8. The method of claim 1, further comprising: receiving, from another electronic device, health-related data associated with the user of the electronic device. 9. The method of claim 8, wherein the receiving health-related data associated with the user of the electronic device comprises receiving data derived from an environmental sensor. 10. The method of claim 1, wherein the notification is proactively communicated to the user of the electronic device. 11. An electronic device, comprising: an output device;a network component;a processor; anda memory storing instructions executable by the processor to:identify at least one user associated with a user of the electronic device;receive, using the network component of the electronic device, health-related information associated with the at least one user;generate a notification based on the health-related information; andprovide, using the output device, the notification to the user of the electronic device. 12. The electronic device of claim 11, wherein the at least one user is identified using data selected from the group consisting of a contact list of the electronic device, a calendar appointment of the electronic ... |  | A61B005/00 A61B005/117 A61B005/11
US20160135706 | A1 | 2016-05-19 | Medical Premonitory Event Estimation | 2015-11-13 | A system and method for medical premonitory event estimation includes one or more processors to perform operations comprising: acquiring a first set of physiological information of a subject, and a second set of physiological information of the subject received during a second period of time; calculating first and second risk scores associated with estimating a risk of a potential cardiac arrhythmia event for the subject based on applying the first and second sets of physiological information to one or more machine learning classifier models, providing at least the first and second risk scores associated with the potential cardiac arrhythmia event as a time changing series of risk scores, and classifying the first and second risk scores associated with estimating the risk of the potential cardiac arrhythmia event for the subject based on the one or more thresholds. | 1. A medical premonitory event estimation system, comprising: a non-transitory computer-readable storage medium in communication with one or more processors and having instructions stored thereon which, when executed by the one or more processors, cause the one or more processors to perform operations comprising: for a plurality of time periods, calculating an event estimation of risk score associated with a potential medical event for a subject occurring within the associated time period based at least partly on physiological parameter data of the subject. 2. The system of claim 1, wherein the physiological parameter data comprises ECG data. 3. The system of claim 1, wherein the physiological parameter data comprises at least one of blood pressure data, heart rate data, thoracic impedance data, pulse oxygen level data, respiration rate data, heart sound data, lung sound data, and activity level data. 4. The system of claim 1, wherein the potential medical event comprises a cardiac event. 5. The system of claim 4, wherein the cardiac event comprises at least one of an ectopic beat, a run of ectopic beats, a ventricular tachycardia, a bradycardia, asystole, and a T-wave abnormality. 6. The system of claim 1, wherein the potential medical event comprises at least one of a plurality of medical events, an increase in a rate of medical events, and/or an increase in an intensity of medical events. 7. The system of claim 1, wherein the potential medical event is defined in a multidimensional parameter space comprising the physiological parameter data and at least one other type of physiological parameter data and/or demographic data of the subject. 8. The system of claim 1, wherein the one or more processors perform operations comprising: calculating a plurality of different event estimation of risk scores associated with the potential medical event for the subject within the associated time period based at least partly on the physiological parameter data. 9. The system of claim 1, wherein the one or more processors perform operations comprising: calculating a plurality of different event estimation of risk scores associated with a plurality of different potential medical events for the subject within the associated time period based at least partly on the physiological parameter data. 10. The system of claim 1, wherein the calculating the event estimation of risk score comprises calculating a criticality score indicating a significance of the potential medical event with respect to at least one other potential medical event. 11. The system of claim 1, wherein the calculating the event estimation of risk score comprises calculating a confidence score including a probability that the potential medical event occurs within the associated time period. 12. The system of claim 1, wherein the one or more processors perform operations comprising: determining that the event estimation of risk score satisfies one or more event estimation of risk thresh... |  | A61B005/04 A61B005/0205 A61B005/0464 A61B005/0468 A61B005/046 A61B005/00 A61N001/39 A61B008/08 A61B005/021 A61B005/085 A61B007/04 A61B007/00 A61B005/11 A61B005/053 A61B005/0472 A61B005/0456 A61B005/1455
US09763629 | B2 | 2017-09-19 | Medical device with context-specific interfaces | 2014-11-07 | A medical device includes: a processor; a display; and memory encoding instructions that, when executed by the processor, cause the processor to create an interface on the display, the interface including: a login area including one or more prompts for information identifying a user of the medical device; a profile selection area including two or more profiles offered by the device, wherein each of the profiles is a series of tasks to be performed by the medical device, and a description of each profile is provided in the profile selection area; a vital signs area programmed to display a plurality of vital signs associated with a patient; and an additional parameters area programmed to capture one or more additional parameters associated with the vital signs displayed in the vital signs area. | 1. A medical device comprising: a processor;a display; andmemory encoding instructions that, when executed by the processor, cause the processor to create multiple, sequential interfaces on the display for sequential capture of information, the sequential interfaces including: a first interface, including: a login area including one or more prompts for information identifying a user of the medical device; anda profile selection area including two or more profiles offered by the medical device, wherein each of the two or more profiles is a series of tasks to be performed by the medical device, and wherein the profile selection area comprises: a separate area within the profile selection area for selection of a respective one of the two or more profiles;a name of the respective one of the two or more profiles; anda description of what is accomplished using the respective one of the two or more profiles;a second interface programmed to display data associated with multiple captured vital signs of a patient and including a first next button; anda third interface including multiple additional parameter areas and a second next button, wherein one of the multiple additional parameters areas records an early warning score for the patient, and wherein at least one of the multiple additional parameter areas includes a skip button. 2. The medical device of claim 1, wherein the login area includes prompts for a clinician identifier and a password. 3. The medical device of claim 1, wherein the user must provide the information identifying the user before selecting one of the two or more profiles in the profile selection area. 4. The medical device of claim 1, wherein the second interface comprises a vital signs interface that is loaded once the user has provided the information identifying the user and selected one of the two or more profiles. 5. The medical device of claim 4, wherein the first interface further comprises a patient identifier area that prompts for identification of a patient before the vital signs interface is provided. 6. A medical device comprising: a processor;a display; andmemory encoding instructions that, when executed by the processor, cause the processor to create a first interface and a second interface on the display,wherein the first interface comprises: a login area including one or more prompts for information identifying a user of the medical device; anda profile selection area including two or more profiles offered by the medical device, wherein each of the two or more profiles is a series of tasks to be performed by the medical device, and wherein each of the two or more profiles comprises a profile name and a description of profile functionality, andwherein the second interface comprises: a vital signs area programmed to display data associated with a plurality of vital signs for a patient and including a first next button; andan additional parameters area including multiple additional parameters, a second next button an... | WELCH ALLYN, INC. | A61B005/00 G06F019/00
US09339237 | B2 | 2016-05-17 | Continuous transdermal monitoring system and method | 2015-07-09 | Various embodiments of methods and systems for continuous transdermal monitoring (“CTM”) are disclosed. One exemplary method for CTM begins by monitoring an output signal from an accelerometer. The accelerometer output signal may indicate acceleration and deceleration of a body part of a user, such as the user's wrist. Based on the accelerometer output signal, it may be determined that the body part of the user has decelerated to a minimum, e.g., substantially zero. With a determination that the body part has decelerated to the minimum, e.g., substantially zero, or has not accelerated beyond the minimum, e.g., substantially zero, the method may determine a reading from a pulse oximeter associated with the accelerometer. Advantageously, the pulse oximetry reading, or a reading from other sensors associated with the accelerometer, may be optimally accurate as motion artifact may be minimized. The pulse oximetry reading may be recorded for later query and/or rendered for the benefit of the user. | 1. A method for minimizing the effects of motion artifact on sensor readings taken during continuous transdermal monitoring of a body part in motion, the method comprising: monitoring, with a processor, an output signal from an accelerometer, wherein the output signal indicates acceleration and deceleration of the body part of a user as the body part is in motion;determining, with the processor, from the output signal that the body part of the user has decelerated to a predetermined minimum before becoming momentarily stationary to begin acceleration;based on the determination that the body part of the user has decelerated to the predetermined minimum, determining, with the processor, a plurality of readings from a sensor, wherein determining the plurality of readings from the sensor occurs over a time window that includes the body part of the user becoming momentarily stationary; andrecording in a memory device the determined plurality of readings from the sensor. 2. The method of claim 1, further comprising rendering an average of the determined plurality of readings from the sensor to the user via a display. 3. The method of claim 1, wherein the sensor is a pulse oximeter configured to: pulse a red light source for approximately 50 microseconds; anddelay approximately 450 microseconds before pulsing an infrared light source approximately 50 microseconds. 4. The method of claim 3, wherein the plurality of readings from the pulse oximeter indicates a blood oxygen saturation level of the user. 5. The method of claim 3, wherein the pulse oximeter is a reflective pulse oximeter. 6. The method of claim 1, wherein the accelerometer is a 3-axis accelerometer. 7. The method of claim 1, wherein the body part of the user is a wrist. 8. A system for minimizing the effects of motion artifact on sensor readings taken during continuous transdermal monitoring of a body part in motion, the system comprising: a sensor package comprising a processor, an accelerometer, a sensor and a memory device, the sensor package configured to: monitor an output signal from the accelerometer, wherein the output signal indicates acceleration and deceleration of the body part of a user as the body part is in motion;determine from the output signal that the body part of the user has decelerated to a predetermined minimum before becoming momentarily stationary to begin acceleration;based on the determination that the body part of the user has decelerated to the predetermined minimum, determine a plurality of readings from the sensor, wherein the determined plurality of readings from the sensor occurs over a time window that includes the body part of the user becoming momentarily stationary; andrecord the determined plurality of readings from the pulse oximeter in the memory device. 9. The system of claim 8, further comprising a hub device configured to render an average of the determined plurality of readings from the sensor to the user via a display. 10. The system of c... |  | A61B005/02 A61B005/00 A61B005/024 A61B005/1455 A61B005/11 A61B005/01 A61B005/0205 A61B005/021 A61B005/053
US10194825 | B2 | 2019-02-05 | Methods and apparatuses for the localization and treatment of disease such as cancer | 2015-12-17 | Embodiments of the present invention provide methods of detecting disease, methods of treating disease using targeted hyperthermia, methods of treating disease using targeted chemical agents, methods of treating disease comprising accurate measurements of the efficacy of treatments. The effect of nanoparticles on magnetic fields can be used to determine the location of a disease, and a measure of the number of cells characteristic of the disease. This location and measure can be used to guide therapy, and provide information regarding the most effective therapy to be applied. The same nanoparticles can be used to facilitate hyperthermia treatments, and to allow targeted application of chemical therapeutic agents. | 1. A method for determining the location and number of cells characteristic of a disease comprising: (1) introducing to the cells in vivo a plurality of superparamagnetic magnetic nanoparticles conjugated with biocompatible features that cause the nanoparticles to interact with the cells such that particles that interact with the cells are inhibited from rotating as rapidly as nanoparticles that do not interact with the cells, (2) after step (1), subjecting the region of the cells of interest to a magnetic field imparting energy to the nanoparticles to magnetize the nanoparticles, (3) measuring a residual magnetic field in the region of the cells during a time after the net magnetic moments of nanoparticles that do not interact with the cells have decayed and before the net magnetic moments of nanoparticles that do interact with the cells cells have decayed, and performing at least one of (4) determining the number of cells characteristic of the disease from the residual magnetic field, and (5) determining the location of the cells characteristic of the disease by the location of the residual magnetic field. 2. A method as in claim 1, wherein step (3) measuring a residual magnetic field comprises attenuating or removing the applied magnetic field and measuring the decay of a residual magnetic field after the applied magnetic field has been attenuated or removed. 3. A method for treating disease, comprising (a) determining the location and number of cells characteristic of the disease according to claim 1, and (b) applying a therapy to the patient. 4. A method as in claim 3, wherein step (b) comprises applying hyperthermia to the region of the cells. 5. A method as in claim 4, wherein applying hyperthermia comprises applying an oscillating magnetic field to the region of the cells, wherein the oscillating magnetic field generates heat by its effect on the nanoparticle/feature conjugates. 6. A method according to claim 3, further comprising (c) determining a measure of the number of cells characteristic of the disease after treatment, then repeating steps (b) and (c) until a desired number of cells characteristic of the disease have been eliminated from the tumor. 7. A method as in claim 3, wherein step (b) comprises (b1) providing delivery packages, wherein a delivery package comprises a magnetic nanoparticle, a drug that harms cells characteristic of the disease, and an inhibitor that inhibits action of the drug; (b2) introducing a plurality of delivery packages into a patient; (b3) applying a magnetic field to encourage the delivery packages to the location of a tumor; (b4) treating the patient such that the inhibitor is overcome and drug acts on the cells characteristic of the disease. 8. A method as in claim 7, wherein the inhibitor comprises a coating that encapsulates the drug, and wherein treating the patient so that the inhibitor is overcome comprises applying a magnetic field such that the delivery packages generate heat, motion, o... | Imagion Biosystems Inc. | A61B005/05 C07K016/40 C07K016/28 C07K016/30 C07K016/18 A61K041/00 A61B005/04 A61B005/06 A61B005/00 A61K049/18 A61B090/00
US09344546 | B2 | 2016-05-17 | Fitness activity related messaging | 2014-07-28 | In one embodiment, a method for generating a message to a friend of a user is provided, comprising: processing activity data of a first user measured by an activity monitoring device to update a value of an activity metric for the first user; identifying a change in an inequality relationship between the value of the activity metric for the first user and a value of the activity metric for a second user; in response to identifying the change in the inequality relationship, prompting the first user to generate a message to the second user. | 1. A method comprising: receiving, over a network by at least one server computer, activity data of a first user measured by an activity monitoring device;processing, by synchronization logic of the at least one server computer, the activity data of the first user to update a value of an activity metric for the first user;in response to updating the value of the activity metric for the first user, identifying, by the at least one server computer, a change in an inequality relationship between the value of the activity metric for the first user and a value of the activity metric for a second user, the change in the inequality relationship being defined by the value of the activity metric for the first user changing, from being less than the value of the activity metric for the second user, to being greater than the value of the activity metric for the second user;in response to identifying the change in the inequality relationship, triggering a messaging logic of the at least one server computer to prompt the first user via a first user device to generate a message to the second user for display on a second user device. 2. The method of claim 1, wherein the inequality relationship between the value of the activity metric for the first user and the value of the activity metric for the second user defines a relative ranking of the first user and the second user on a competitive leaderboard for the activity metric. 3. The method of claim 2, wherein the change in the inequality relationship defines a change in the relative ranking of the first user and the second user on the competitive leaderboard for the activity metric. 4. The method of claim 3, wherein the triggering of the message logic of the at least one server computer to prompt the first user to generate a message includes identifying the change in the relative ranking of the first user and the second user. 5. The method of claim 1, wherein the triggering of the message logic of the at least one server computer to prompt the first user to generate a message to the second user includes triggering a notification to the first user, the notification providing access to an interface configured to receive input from the first user to define content of the message to the second user. 6. The method of claim 5, wherein the notification is a push notification to a mobile device. 7. The method of claim 5, wherein the interface provides at least one option to send a predefined message to the second user. 8. The method of claim 1, wherein the activity metric is defined by one or more of the following: stepcount, elevation climbed, distance traveled, active amount of time. 9. A method comprising: receiving, over a network by at least one server computer, activity data of a first user measured by an activity monitoring device;processing, by synchronization logic of the at least one server computer, the activity data of the first user to update a value of an activity metric for the first user;in resp... | Fitbit, Inc. | H04L012/58 H04M001/725 G08B003/10 G08B005/36 G08B006/00 A61B005/00 G01C022/00 G06F003/041 G06F015/00 G06F019/00 H04L029/06 H04L029/08 G08B021/18 A61B005/0205 A61B005/021 A61B005/024 A61B005/11 A61B005/1455 A61B005/22
US09339348 | B2 | 2016-05-17 | Devices, systems, and methods for assessing a vessel | 2012-04-30 | Embodiments of the present disclosure are configured to assess the severity of a blockage in a vessel and, in particular, a stenosis in a blood vessel. In some particular embodiments, the devices, systems, and methods of the present disclosure are configured to assess the severity of a stenosis in the coronary arteries without the administration of a hyperemic agent. | 1. A method of evaluating a vessel of a patient, comprising: introducing a first intravascular pressure-sensing instrument into the vessel of the patient proximal of a stenosis of the vessel;introducing a second intravascular pressure-sensing instrument into the vessel of the patient distal of the stenosis of the vessel;receiving, at a computing device of an intravascular processing system in communication with the first and second intravascular pressure-sensing instruments, proximal pressure measurements obtained by the first intravascular pressure-sensing instrument without application of a hyperemic agent during a cardiac cycle of the patient with the first intravascular pressure-sensing instrument positioned proximal of the stenosis;receiving, at the computing device of the intravascular processing system in communication with the first and second intravascular pressure-sensing instruments, distal pressure measurements obtained by the second intravascular pressure-sensing instrument without application of a hyperemic agent during the cardiac cycle of the patient with the second intravascular pressure-sensing instrument positioned distal of the stenosis;receiving, at the computing device of the intravascular processing system, an ECG signal for the cardiac cycle of the patient;selecting, by the computing device of the intravascular processing system, a diagnostic window within the cardiac cycle of the patient, wherein the diagnostic window encompasses only a portion of the cardiac cycle of the patient and wherein a starting point of the diagnostic window is determined based on a peak pressure measurement of at least one of the received proximal pressure measurements and the received distal pressure measurements and an ending point of the diagnostic window is determined based on the ECG signal;identifying, by the computing device of the intravascular processing system a plurality of the proximal pressure measurements obtained during the diagnostic window from the received proximal pressure measurements based on the selected diagnostic window, wherein the plurality of proximal pressure measurements obtained during the diagnostic window are a subset of the received proximal pressure measurements;identifying, by the computing device of the intravascular processing system, a plurality of the distal pressure measurements obtained during the diagnostic window from the received distal pressure measurements based on the selected diagnostic window, wherein the plurality of distal pressure measurements obtained during the diagnostic window are a subset of the received distal pressure measurements;calculating, by the computing device of the intravascular processing system, a pressure ratio between an average of the plurality of distal pressure measurements obtained during the diagnostic window and an average of the plurality of proximal pressure measurements obtained during the diagnostic window;identifying a treatment option based on the calculated press... | Imperial Colege of Science, Technology and Medicine Volcano Corporation Medsolve Limited | A61B005/02 A61B019/00 A61B005/0215
US09332893 | B2 | 2016-05-10 | Delivery of biological compounds to ischemic and/or infarcted tissue | 2014-08-05 | The delivery of biological compounds to ischemic and/or infarcted tissue are described herein where such a system may include a deployment catheter and an attached imaging hood deployable into an expanded configuration. In use, the imaging hood is placed against or adjacent to a region of tissue to be imaged in a body lumen that is normally filled with an opaque bodily fluid such as blood. A translucent or transparent fluid, such as saline, can be pumped into the imaging hood until the fluid displaces any blood, thereby leaving a clear region of tissue to be imaged via an imaging element in the deployment catheter. Additionally, any number of therapeutic tools can also be passed through the deployment catheter and into the imaging hood for performing any number of procedures on the tissue for identifying, locating, and/or accessing ischemic and/or infarcted tissue. | 1. A method of identifying a tissue region of interest for treatment, comprising: expanding a non-inflatable fluid barrier, defining an open area which is in fluid communication with an environment external to the fluid barrier and projecting distally from a deployment catheter, against the tissue region of interest;urging a transparent fluid into the open area such that an opaque fluid is displaced therefrom;visualizing the tissue region of interest through the transparent fluid; andassessing at least one parameter of the tissue region. 2. The method of claim 1 further comprising intravascularly advancing the deployment catheter prior to expanding. 3. The method of claim 1 wherein expanding comprises self-expanding the fluid barrier from a low-profile delivery configuration into an expanded deployed configuration. 4. The method of claim 1 wherein urging a transparent fluid comprises pumping the transparent fluid into the open area through a fluid delivery lumen defined through the deployment catheter. 5. The method of claim 4 wherein pumping the transparent fluid comprises urging saline, plasma, water, or perfluorinated liquid into the open area such that blood is displaced. 6. The method of claim 1 wherein visualizing comprises imaging the tissue region via an imaging element positioned within or adjacent to the expanded fluid barrier. 7. The method of claim 1 wherein assessing comprises visually inspecting the tissue region for indications of ischemic or infarcted regions of tissue. 8. The method of claim 1 wherein assessing comprises measuring a partial pressure of oxygen or carbon dioxide present in the tissue region. 9. The method of claim 1 wherein assessing comprises measuring a temperature of the tissue region. 10. The method of claim 1 wherein assessing comprises measuring for a presence of a biomarker indicative of tissue injury. 11. The method of claim 10 wherein the biomarker is selected from the group consisting of troponin T, troponin I, troponin C; creatine phosphokinase, MB fraction (CKMB), myoglobin, lactate dehydrogenase (LDH), and combinations thereof. 12. The method of claim 1 wherein assessing comprises detecting differences in electrical current or electrical potential in the tissue region. 13. The method of claim 1 wherein assessing comprises measuring relative tissue hardness or deflection over the tissue region. 14. The method of claim 1 further comprising identifying injured tissue based upon an assessment of the at least one parameter. 15. The method of claim 14 further comprising treating the injured tissue with a bioactive agent. 16. The method of claim 15 further comprising injecting or infusing the bioactive agent upon or into the injured tissue. 17. The method of claim 15 wherein the bioactive agent is selected from the group consisting of stem cells, vascular endothelial growth factor, fibroblast growth factor, pro-angiogenic cytokines, tumor necrosis factor alpha, interleukin 8, SIKVAV, and neuro... | Intuitive Surgical Operations, Inc. | A61M031/00 A61B001/00 A61B001/12 A61B005/00 A61B001/015 A61M025/00
US09332972 | B2 | 2016-05-10 | Aspiration biopsy apparatus and method | 2012-03-29 | An aspiration apparatus includes a barrel having a needle receiving end and a piston receiving end. The barrel has an inner diameter and an outer diameter. The piston receiving end of the barrel has substantially the same inner diameter as the inner diameter of the barrel. The barrel has an opening therein through a sidewall of the barrel. The opening is between the needle receiving end and the piston receiving end. The opening is distant from the needle receiving end. The aspiration apparatus also includes a needle sealingly attached to the needle receiving end of the barrel, and a piston having an end which substantially seals to the inner diameter of the barrel. The piston is capable of a number of positions including a first position near the needle receiving end, a third position past the opening in the barrel. | 1. An aspiration apparatus comprising: a barrel having an inner diameter and an outer diameter, the barrel further comprising: a needle receiving end; anda piston receiving end having substantially the same inner diameter as the inner diameter, the barrel having an opening therein through a sidewall of the barrel at a position between the needle receiving end and the piston receiving end;a needle sealingly attached to the needle receiving end of the barrel;a piston having an end which substantially seals to the inner diameter of the barrel, the piston having a first position near the needle receiving end wherein moving the piston away from the first position results in a vacuum being formed within the barrel, and wherein moving the piston to a third position past the opening in the barrel releases the vacuum; anda mechanical apparatus to controllably move the piston between a first position and a second position, wherein the mechanical apparatus further comprises: an inner sleeve removably attached to the piston;an outer sleeve capturing the barrel, the outer sleeve including a lever to enable moving of the inner sleeve and attached piston with respect to the outer sleeve. 2. The aspiration apparatus of claim 1 further comprising a spring positioned between the inner sleeve and the outer sleeve. 3. The aspiration apparatus of claim 2 wherein the inner sleeve has an inner shoulder for receiving one end of the spring and the outer sleeve includes an outer shoulder for receiving the other end of the spring. 4. The aspiration apparatus of claim 1 wherein the inner sleeve is attached to the piston and moves with the piston, the inner sleeve including a primary latching surface, and the outer sleeve attached to the barrel, the outer sleeve including a lever rotatably attached to the outer sleeve and having a primary latch hook for holding the inner sleeve in a position where the piston is in a first position with respect to the barrel. 5. The aspiration apparatus of claim 1 wherein the inner sleeve is attached to the piston and moves with the piston, the inner sleeve including a primary latching surface and a secondary latching surface, and the outer sleeve attached to the barrel, the outer sleeve including a first lever rotatably attached to the outer sleeve and having a primary latch hook for holding the inner sleeve in a position where the piston is in a first position with respect to the barrel, and including a second lever for latching the secondary latching surface to hold the piston at a position with respect to the barrel intermediate the first and third position. 6. The aspiration apparatus of claim 5, wherein the first lever and the second lever are rotatably attached to the same axis. 7. The aspiration apparatus of claim 5, wherein the first lever and the second lever are rotatably attached to the outer sleeve, and wherein the first lever and the second lever rotate about the same axis. 8. The aspiration apparatus of claim 1 wherein t... | GREER MEDICAL, INC. | A61B005/00 B65D081/00 A61B010/02
US09345397 | B2 | 2016-05-24 | Optical sensing system for cochlear implant surgery | 2011-09-21 | A sensing system for implant surgery includes an insertion device for moving an implant into a narrow cavity in a patient's body. A sensor measures the distance from an end of the insertion device to anatomic surfaces at a distance from the end of the insertion device. An optical coherence tomography (OCT) system integrates the sensor and produces OCT images, which can be quantified to distance measurements. The system is particularly useful for cochlear implant surgery. | 1. A sensing system for implant surgery, comprising: an insertion device or stylet configured for moving a cochlear implant into a narrow cavity in a patient's body;an optical sensor for measuring distance from an end of the insertion device to anatomic surfaces at a distance from the end of the insertion device; andan optical coherence tomography (OCT) system connected to the optical sensor,wherein the optical sensor is incorporated within and at a distal end of the insertion device, andwherein a distal end of the optical sensor provides a reference signal for the OCT system. 2. The sensing system of claim 1, wherein the optical sensor comprises an optical waveguide to transmit and collect a laser beam of light. 3. The sensing system of claim 2, wherein the optical waveguide comprises an optical fiber. 4. The sensing system of claim 1, wherein the OCT system is a Fourier domain common path OCT (FD-CPOCT) system, and wherein the OCT system includes a computer processor for processing OCT images from the optical sensor to anatomic surfaces and converting the OCT images to distances. 5. The sensing system of claim 1, further comprising a display device for displaying OCT images or distance information to a surgeon. 6. The sensing system of claim 1, further comprising an auditory device for providing auditory information representing distance information. 7. The sensing system of claim 1, wherein said optical sensor includes facets for creating forward and side viewing probe beams so as to measure multiple distances. 8. The sensing system of claim 7, wherein the multiple distances are measured simultaneously. 9. The sensing system of claim 7, wherein the multiple distances are measured sequentially. 10. The sensing system of claim 1, further comprising multiple optical sensors to measure multiple distances between multiple anatomical locations. 11. The sensing system of claim 1, wherein the distance measurements are directed to a control system of a robotic device for cochlear implant surgery. 12. The sensing system of claim 1, wherein the optical sensor is an ultrasound sensor. 13. The sensing system of claim 1, wherein the insertion device is a stylet. 14. The sensing system of claim 1, wherein the insertion device is a sheath. 15. The sensing system of claim 14, wherein the optical sensor is an ultrasound sensor. 16. The sensing system of claim 1, wherein the OCT system produces A-mode OCT images. 17. The sensing system of claim 1, wherein the OCT system produces B-mode OCT images. 18. The sensing system of claim 1, wherein the OCT system produces C-mode OCT images. 19. A sensing system for implant surgery, comprising: a cochlear implant adapted to be positioned into a narrow cavity during implant surgery;an optical sensor for measuring distance from an end of the optical sensor to anatomical surfaces at a distance from the end of the optical sensor; andan optical coherence tomography (OCT) system connected to the optical sensor... | The Johns Hopkins University | A61B006/00 A61B001/227 A61B001/00 A61B005/00 A61F011/00 A61N001/36
US09326711 | B2 | 2016-05-03 | Optical perfusion sensor detector | 2008-06-30 | A reflectance optical perfusion sensor may include at least one light source and a plurality of detector elements arranged in a planar or nonplanar configuration, such as a three-dimensional array. The detector elements may sense light emitted by the at least one light source and reflected by a blood mass of a patient, such as blood within a blood vessel. In some examples, the detector elements may be arranged such that photodetection surfaces of at least two of the detector elements are nonparallel. In addition to or instead of the nonplanar arrangement of detector elements, an optical perfusion sensor may include a detector array including a plurality of detector elements at least partially surrounding a light source. Varying the location and orientations of detector elements may help increase a quantity of light emitted by the at least one light source and reflected toward the optical perfusion sensor by blood. | 1. An implantable medical device comprising: an outer housing surface including a first portion configured having a planar profile and a second portion of the surface configured having a curvilinear profile;a light source disposed on the first portion having an emitter surface oriented in a first plane; anda plurality of detector elements disposed adjacent to the light source with at least a first of the detector elements being arranged on the first portion and at least a second of the detector elements being arranged on the second portion, each of the detector elements having a photodetection surface and the photodetection surface of the first detector element being oriented in the first plane and the photodetection surface of the second detector element being oriented in a second plane that is different from the first plane, wherein each detector element of the plurality of detector elements is configured to sense a reflection of light that is emitted by the emitter surface and reflected onto the respective detector element. 2. The implantable medical device of claim 1, wherein the plurality of detector elements are arranged in a three-dimensional array of detector elements. 3. The implantable medical device of claim 1, wherein at least a first detector element is arranged in a first spatial position and a second detector element is arranged in a second spatial position that is different from the first spatial position in at least two dimensions. 4. The implantable medical device of claim 1, further comprising a housing that defines a surface comprising a parabolic cross-section, wherein at least some of the plurality of detector elements are coupled to the surface. 5. The implantable medical device of claim 1, wherein at least one of the detector elements comprises a photodiode. 6. The implantable medical device of claim 1, wherein the plurality of detector elements are symmetrically arranged around the light source. 7. The implantable medical device of claim 1, further comprising a housing, wherein the light source and the detector elements are arranged on a common side of the housing. 8. The implantable medical device of claim 7, wherein the housing defines a recess and the light source and the detector elements are disposed in the recess. 9. The implantable medical device of claim 7, wherein the housing defines the surface having the planar profile and the curvilinear profile. 10. The implantable medical device of claim 7, wherein the housing defines a first surface and a second surface, and the light source is coupled to the first surface and at least one of the detector elements of the detector array is coupled to the second surface. 11. The implantable medical device of claim 1, wherein the plurality of detector elements at least partially surround the light source. 12. The implantable medical device of claim 11, wherein the plurality of detector elements surround the light source. 13. The implantable medical device of claim ... | Medtronic, Inc. | A61B005/145 A61B005/1459 A61B005/00 A61B005/1455 A61N001/08 A61N001/37
US09326718 | B2 | 2016-05-03 | Blood glucose measurement for small blood volume | 2010-07-28 | A system is proposed for detection of at least one analyte in a body fluid, in particular for detection of blood glucose. The system is designed to generate a sample of the body fluid and to transfer at least some of the sample to at least one test element, in particular a test panel. The system is designed such that a time period between the generation of the sample and the application to the test element is less than 1 second, preferably less than 500 ms. | 1. A system for detecting at least one analyte in a body fluid, the system comprising: at least one test element comprising a test panel configured to detect the at least one analyte;at least one lancet configured to generate a body fluid sample of less than 1 μl;an actuator configured to incite the at least one lancet;at least one transfer device configured to transfer a whole or partial volume of the body fluid sample generated by the at least one lancet to the at least one test element in a transfer time of less than 1 second; anda measurement device configured to detect the at least one analyte, configured to detect an actual sample volume of the body fluid sample generated by the at least one lancet including a volume of the body fluid sample transferred to the at least one test element, configured to correct a measured value of the at least one analyte taking into account the actual sample volume, and configured to at least partially compensate changes in concentration to the measured value of the at least one analyte taking into account evaporation effects dependent on sample volume, wherein the measurement device comprises at least one sensor for detecting the actual sample volume. 2. The system of claim 1, wherein the transfer device comprises at least one capillary, the at least one lancet comprising at least partially a capillary, the at least one lancet comprising a partially opened capillary, or the at least one lancet comprising a wholly or partially integrated capillary and a wholly or partially integrated at least one test element. 3. The system of claim 2, wherein the transfer device is the at least one lancet comprising the wholly or partially integrated capillary and the wholly or partially integrated at least one test element. 4. The system of claim 2, wherein the transfer device comprises at least one hydrophilic coating. 5. A system for detecting at least one analyte in a body fluid, the system comprising: at least one test element comprising a test panel configured to detect the at least one analyte;at least one lancet configured to generate a body fluid sample of less than 1 μl;an actuator configured to incite the at least one lancet;at least one capillary configured to transfer a whole or partial volume of the body fluid sample generated by the at least one lancet to the at least one test element in a transfer time of less than 1 second, wherein the at least one capillary has a ratio of fillable capillary length to capillary diameter of less than 100; anda measurement device configured to detect the at least one analyte, configured to detect an actual sample volume of the body fluid sample generated by the at least one lancet including a volume of the body fluid sample transferred to the at least one test element, configured to correct a measured value of the at least one analyte taking into account the actual sample volume, and configured to at least partially compensate changes in concentration to the measured value... | Roche Diabetes Care, Inc. | A61B005/157 B65D081/00 A61B005/15 A61B005/145
US09326753 | B2 | 2016-05-03 | Ultrasonograph | 2013-07-11 | An ultrasonograph that can automatically optimize the sweep rate of Doppler images and M-mode images according to the heart rate of a test subject. The ultrasonograph includes means for sending an ultrasonic beam into body tissue, a means for receiving an ultrasonic signal reflected off the body tissue and a blood flow, a means for constructing a cross-sectional image of the body tissue from the received signal, a means for performing phase detection on an ultrasonic Doppler blood flow signal, a means for calculating a frequency component of the Doppler blood flow signal, and a means for performing a sweeping display of the calculated frequency component as Doppler images in a time series. The ultrasonograph has a heart rate measuring means, an optimum sweep rate calculating means and an optimum sweep rate setting means. | 1. An ultrasonograph to send an ultrasonic beam into body tissue and acquiring an ultrasonic Doppler signal based on an ultrasonic signal reflected off the body tissue and a blood flow, comprising: an ultrasonic blood flow signal filtering unit that eliminates a low frequency component that is a part of the Doppler signal and that is generated by movement of the body tissue;a frequency analysis unit that calculates a frequency component of the ultrasonic Doppler signal from which the low frequency component has been eliminated by the ultrasonic blood flow signal filtering unit;a display control unit that performs in a time series a sweeping display of the frequency component as Doppler images calculated by the frequency analysis unit;a Doppler trace calculation unit that generates a Doppler trace waveform based on the frequency component calculated by the frequency analysis unit;a heart rate measurement unit that measures a heart rate of a test subject based on the Doppler trace waveform at a predetermined timing;a sweep rate calculation unit that calculates a sweep rate based on the heart rate obtained by the heart rate measurement unit; anda sweep rate control unit that sets the sweep rate of the Doppler images based on the sweep rate calculated by the sweep rate calculation unit. 2. The ultrasonograph according to claim 1, further comprising: an operation unit that can be operated by an operator,wherein the sweep rate control unit sets the sweep rate of the Doppler images based on an operation of the operator. 3. The ultrasonograph according to claim 2, wherein the operation unit is capable of setting how many heart beats are represented by the Doppler images in one screen. 4. The ultrasonograph according to claim 1, wherein the sweep rate control unit automatically executes a sweep rate setting in a case in which a mode including a display of the Doppler images has been transitioned to. 5. The ultrasonograph according to claim 1, wherein the heart rate measurement unit calculates a time required for each heart beat from the Doppler trace waveform obtained by the Doppler trace calculation unit, and calculates the heart rate based on a time required for each heart beat. 6. A processor of an ultrasonograph that performs a sweeping display of Doppler images in a time series based on a frequency component of an ultrasonic Doppler signal from which a low frequency component generated by movement of tissue has been eliminated, comprising: a Doppler trace calculation unit that generates a Doppler trace waveform based on the frequency component;a heart rate measurement unit that measures a heart rate of a test subject at a predetermined timing based on the Doppler trace waveform;a sweep rate calculation unit that calculates a sweep rate based on the heart rate obtained by the heart rate measurement unit; anda sweep rate control unit that sets a sweep rate of Doppler images for which the ultrasonograph performs a sweeping display based on a sweep ... | KONICA MINOLTA, INC. | A61B008/00 A61B008/08 A61B008/06 A61B005/024 A61B005/0456 A61B005/00
US09339594 | B2 | 2016-05-17 | Bladder evacuation system | 2013-06-11 | The evacuation system employs a piston to deliver fluid for irrigating and aspirating a body cavity, such as a bladder. A pivotal screen is disposed across the flow path of the fluid to move out of the fluid flow path during irrigation of the bladder and to move back into the flow path during aspiration of the bladder. The screen screens out particulate matter from the return flow for depositing into the container of the system. | 1. A bladder evacuation system comprising a housing having a cavity, an inlet communicating with said cavity for passage of fluid therebetween and an exit port opposite said inlet and communicating with said cavity for passage of fluid therebetween;means for selectively expelling a fluid through said inlet into said cavity and drawing fluid from said cavity through said inlet;a container removably secured to said housing in communication with said cavity; anda screen mounted in said cavity of said housing, said screen being movable between a first position disposed over said container and out of a flow path for the fluid between said inlet and said exit port and a second position disposed across said flow path whereby in response to passage of fluid in said flow path from said inlet to said exit port, said screen is moved into said first position to allow fluid to pass through said flow path and in response to drawing of fluid from said cavity through said inlet, said screen is moved into said second position to allow fluid from said exit port to be drawn into said cavity and said inlet while screening particulate matter therefrom for depositing into said container. 2. A bladder evacuation system as set forth in claim 1 further comprising a tube having a proximal end in communication with said exit port and a distal end for communication with a bladder of a patient. 3. A bladder evacuation system as set forth in claim 1 wherein said housing includes a cylinder extending longitudinally from said inlet and said means includes a piston reciprocally mounted in said cylinder. 4. A bladder evacuation system as set forth in claim 1 wherein said means includes a syringe having a cylinder in communication with said inlet of said housing and a piston reciprocally mounted in said cylinder for selectively expelling a fluid therefrom and drawing in fluid thereto. 5. A bladder evacuation system as set forth in claim 1 wherein said screen has a plurality of slots therein. 6. A bladder evacuation system as set forth in claim 1 wherein said screen has a shape complementary to said cavity and includes a distal flap portion and a plurality of parallel walls extending proximally from said flap portion and defining slots therebetween for passage of fluid therethrough. 7. A bladder evacuation system as set forth in claim 6 wherein at least some of said walls have a distal face extending at an obtuse angle from said flap portion. 8. A bladder evacuation system as set forth in claim 1 further comprising a spring mounted in said housing for biasing said screen from said first position into said second position. 9. A bladder evacuation system comprising a container;a tube having a proximal end in communication with said container and a distal end for communication with a bladder of a patient;a syringe having a cylinder in communication with said container and said proximal end of said tube and a piston reciprocally mounted in said cylinder for selectively expellin... |  | A61M001/00 A61B005/15 A61M003/02
US09345574 | B2 | 2016-05-24 | Force-based heart valve sizer | 2012-12-06 | A valve sizer for determining an appropriate replacement valve size when performing a heart valve replacement procedure is provided. In one version the valve sizer has a hollow shaft with proximal and distal ends. A movable sizing element couples to the distal end of the shaft and is radially expandable between first, contracted and second, expanded positions. An actuator assembly on a handle includes an actuator coupled to a clutch member via a ball-spring-detent clutch. A rod extends through the shaft and maintains a fixed distance between the handle and a distal hub in the sizing element. Movement of the actuator causes axial movement of the shaft, thereby causing radial expansion of sizing petals relative to the hub. The clutch slips when a predetermined reaction force from the surrounding valve annulus is met by the petals. | 1. A method of sizing a patient's cardiac valve annulus, comprising: providing a valve sizer having a shaft with a proximal end and a distal end and an expandable sizing element coupled to the distal end of the shaft, the sizing element being radially expandable between a first retracted position and a second expanded position, the valve sizer further having an actuator assembly comprising an actuator that moves relative to a handle, a clutch ring mounted at the proximal end of the shaft and coupled for rotation to the actuator via a clutch mechanism, and a stationary rod extending through at least a portion of the shaft, wherein the shaft is connected to the clutch ring so that movement of the actuator transmits through the clutch mechanism to the clutch ring and shaft and causes axial movement of the shaft, wherein the axial movement of the shaft causes radial expansion of the sizing element, wherein the actuator comprises an actuator ring and the sizing element includes a central hub, where the stationary rod is fixed with respect to both the handle and the central hub, and wherein the clutch ring is connected via a screw thread to the handle so that rotation of the clutch ring causes axial movement thereof and of the shaft;inserting the valve sizer in the first retracted position into the patient so that the sizing element is positioned within the valve annulus; androtating the handle until the clutch mechanism of the actuator assembly begins to slip indicating that the sizer has applied a desired outward force to the annulus. 2. The method of claim 1, wherein the clutch mechanism comprises a plurality of bearings biased by springs into detents. 3. The method of claim 2, wherein the plurality of bearings and the springs are held within the clutch ring and the detents are formed on an inner surface of the actuator ring. 4. The method of claim 1, wherein the sizing element includes a plurality of sizing petals radially expandable between the first and second positions. 5. The method of claim 4, wherein the plurality of sizing petals move in a plane substantially perpendicular to a longitudinal axis defined by the shaft. 6. The method of claim 4, wherein the shaft moves axially and contacts and pivots a lever for each of the petals, wherein pivoting of the levers causes radial expansion of the petals. 7. The method of claim 4, wherein the shaft moves axially and connects to a camming member that directly contacts and causes radial expansion of the petals. 8. The method of claim 4, wherein the plurality of petals defines a cylindrical annulus portion and an outwardly-extending flange on a proximal end of the cylindrical annulus portion. 9. The method of claim 8, wherein the outwardly-extending flange has an axially undulating peripheral shape. 10. A method of sizing a patient's anatomical orifice, comprising: providing a valve sizer having: a proximal handle having an actuator;a shaft extending distally from the handle;a surgical sizing ... | Edwards Lifesciences Corporation | A61F002/24 A61B005/107 A61B019/00
US10338166 | B2 | 2019-07-02 | Magnetic resonace imaging apparatus and control method therefor | 2015-01-29 | In order to provide a magnetic resonance imaging apparatus capable of calculating the amount of absorption of electromagnetic waves into the object with high accuracy, the magnetic resonance imaging apparatus includes: a calculation means for calculating the amount of absorption of electromagnetic waves into the object upon the emission of RF pulses in a part of the object or a bed position where imaging is scheduled; a means for setting imaging conditions, in which the calculated amount of absorption satisfies conditions of the specified value of the amount of absorption of electromagnetic waves, from the relationship between the calculated amount of absorption and the specified value of the amount of absorption of electromagnetic waves; and a bed control device that controls a top plate according to the set imaging conditions. | 1. A magnetic resonance imaging apparatus, comprising: a bed including a top plate for moving an object placed thereon;a magnetic field generation device that generates a magnetic field in an imaging space where imaging of the object is performed;an irradiation coil that generates RF pulses to be emitted to the object;an image generation device that detects an NMR signal generated by the object and generates an MRI image based on the detected NMR signal;an input and output device configured to input an imaging position or imaging conditions of the object or to display the imaging position or the imaging conditions;a storage device that stores data regarding an absorption rate of electromagnetic waves; anda CPU that performs operations that includes: calculates an amount of absorption of electromagnetic waves of the object upon emission of the RF pulses at the input imaging position based on the data regarding the absorption rate of electromagnetic waves read from the storage device,determines whether the calculated value satisfies conditions of a specified value regarding absorption of electromagnetic waves,controls an imaging operation at the imaging position in accordance with data indicating the imaging conditions or the imaging position determined that the amount of absorption of electromagnetic waves satisfies the conditions of the specified value,measures the amount of absorption of electromagnetic waves of the object as a measured amount based on a RF pulse emitted in imaging,calculates data regarding an absorption rate of electromagnetic waves of the object as specific absorption rate of an individual object based on the measured amount, andrecalculates the amount of absorption of electromagnetic waves of the object based on the specific absorption rate of the individual object. 2. The magnetic resonance imaging apparatus according to claim 1, wherein: the input and output device includes a display,when a first imaging position and a second imaging position are input as the imaging position, the first and second imaging positions are displayed on the display of the input and output device,the CPU performs further operations that includes: reads data regarding the absorption rate of electromagnetic waves at the first and second imaging positions from the storage device, and calculates the amount of absorption of electromagnetic waves at the first and second imaging positions based on the data regarding the absorption rate of electromagnetic waves at the first and second imaging positions,determines whether the calculated amount of absorption satisfies the conditions of the specified value for each of the first and second imaging positions, and displays an imaging position where the conditions of the specified value are not satisfied on the display when the calculated amount of absorption does not satisfy the conditions of the specified value, andcontrols the imaging operation based on imaging conditions in a state in which the calculated ... | HITACHI, LTD. | G01R033/00 G01R033/28 A61B005/055 A61B005/00 G01R033/54 G01R033/58
US10299730 | B2 | 2019-05-28 | Host of physiological detection system and physiological detection device with high accuracy | 2015-06-02 | There is provided a physiological detection system including a physiological detection device and a host. The physiological detection device is configured to transmit a physiological data series to the host according to a local oscillation frequency. The host is configured to calculate a physiological value according to the physiological data series and determine a correction parameter according to a receiving data parameter and a reference data parameter, wherein the correction parameter is configured to correct the physiological value, process the physiological data series or adjust the local oscillation frequency of the physiological detection device. | 1. A host of a physiological detection system, the host comprising a clock generator and configured to receive a physiological data series sent from a physiological detection device, wherein the physiological data series is sent according to a local oscillation frequency of the physiological detection device,calculate a physiological value according to the physiological data series,calculate a ratio between a receiving data parameter associated with the received physiological data series and a reference data parameter associated with a host oscillation frequency provided by the clock generator of the host to determine a correction parameter, andcorrect the calculated physiological value according to the correction parameter when the local oscillation frequency of the physiological detection device is different from the host oscillation frequency. 2. The host as claimed in claim 1, wherein the physiological value is a heart rate. 3. The host as claimed in claim 1, wherein the reference data parameter is a reference oscillating number of the host oscillation frequency of the host within a predetermined time interval, the receiving data parameter is a receiving data number of the host receiving the physiological data series sent from the physiological detection device within the predetermined time interval, andthe correction parameter is associated with a ratio of the reference oscillating number and the receiving data number. 4. The host as claimed in claim 1, wherein the host is configured to transmit the correction parameter every predetermined correction interval. 5. The host as claimed in claim 4, wherein the transmitted correction parameter is configured to adjust the local oscillation frequency of an adjustable oscillator of the physiological detection device. 6. The host as claimed in claim 1, wherein each physiological data of the physiological data series is a detected value of a single pixel or a sum of detected values of a plurality of pixels of a CMOS sensor. 7. A host of a physiological detection system, the host comprising a clock generator and configured to receive a physiological data series sent from a physiological detection device, wherein the physiological data series is sent according to a local oscillation frequency of the physiological detection device,calculate a ratio between a receiving data parameter associated with the received physiological data series and a reference data parameter associated with a host oscillation frequency provided by the clock generator of the host so as to determine a correction parameter,process the received physiological data series according to the correction parameter, andcalculate a physiological value according to the processed physiological data series to improve accuracy of calculating the physiological value when the local oscillation frequency of the physiological detection device is different from the host oscillation frequency. 8. The host as claimed in claim 7, wherein the phys... | PIXART IMAGING INC. | A61B005/00 A61B005/024 A61B005/1455
US10507321 | B2 | 2019-12-17 | Multilayer structure and method of manufacturing a multilayer structure | 2015-11-25 | In some examples, a medical device system a thin film including at least one electrically conductive track extending between at least one electrode and at least one electrical contact, a first and second polymer layer; wherein, at a portion of the thin film between the at least one electrode and the at least one electrical contact, the first polymer layer and second polymer layer surround the at least one electrically conductive track; and at least one discrete ceramic member located between the first and second polymer layers at a portion of the thin film between the at least one electrode and the at least one electrical contact, wherein the at least one discrete ceramic member does not surround the at least one conductive track, and wherein the at least one discrete ceramic member is configured to increase adhesion between the first polymer layer and second polymer layer. | 1. A medical device system configured to at least one of deliver electrical stimulation to a patient or sense electrical activity of the patient, the system comprising: a thin film including: at least one electrode,at least one electrical contact,at least one electrically conductive track extending between the at least one electrode and the at least one electrical contact,a first polymer layer;a second polymer layer, wherein, at a portion of the thin film between the at least one electrode and the at least one electrical contact, the first polymer layer and second polymer layer surround the at least one electrically conductive track; andat least one discrete ceramic member located between the first and second polymer layers at the portion of the thin film between the at least one electrode and the at least one electrical contact, wherein the at least one discrete ceramic member does not surround the at least one conductive track at the portion of the thin film between the at least one electrode and the at least one electrical contact, wherein the first polymer layer and second polymer layer surround the at least one discrete ceramic member at the portion of the thin film between the at least one electrode and the at least one electrical contact, wherein the at least one discrete ceramic member is in direct contact with the first polymer layer and the second polymer layer, and wherein the at least one discrete ceramic member is configured to increase adhesion between the first polymer layer and second polymer layer. 2. The system of claim 1, wherein the first polymer layer and the second polymer layer each comprise a medical grade polymer. 3. The system of claim 2, wherein the medical grade polymer comprises medical grade poly(p-xylylene) polymer. 4. The system of claim 1, wherein the at least one discrete ceramic member comprises at least one of SiN, SiOx, Si-Carbide, alumina, or oxynitride. 5. The system of claim 1, wherein each discrete ceramic member of the at least one discrete ceramic member is not in contact with the at least one electrode and the at least one electrical contact. 6. The system of claim 1, further comprising a ceramic layer between the at least one electrically conductive track and the first polymer layer, and between the electrically conductive track and the second polymer layer, wherein the ceramic layer surrounds the at least one electrically conductive track at the portion of the thin film between the at least one electrode and the at least one electrical contact. 7. The system of claim 6, wherein the ceramic layer comprises at least one of SiN, SiOx, Si-Carbide, alumina, or oxynitride. 8. The system of claim 6, wherein the ceramic layer and the at least one discrete ceramic member are formed of substantially a same ceramic composition. 9. The system of claim 1, wherein the at least one discrete ceramic member and the at least one electrically conductive track do not directly contact each other. 10. The system o... | Medtronic Bakken Research Center B.V. | A61N001/05 A61L031/02 A61L031/04 A61B005/04
US09854984 | B2 | 2018-01-02 | Optical detection unit and biological information detection device | 2016-01-08 | An optical detection unit includes a light emitting portion 150 that emits light toward a target object, a light receiving portion 140 that receives light from the target object, and a light blocking member 70 that blocks light from being incident to at least the light receiving portion 140. The light blocking member 70 includes a light blocking wall 100 that is formed by performing sheet metal processing on a metal, and is provided between the light emitting portion 150 and the light receiving portion 140 so as to block light from the light emitting portion 150 from being incident to the light receiving portion 140. The light blocking wall 100 is formed of a first metal surface 71 of the light blocking member 70. | 1. An optical detection unit comprising: a light emitting portion that emits light toward a target object;a light receiving portion that receives light from the target object; anda light blocking member between the light emitting portion and the light receiving portion,wherein the light blocking member is made of a metal,wherein the light blocking member is formed of a first metal wall, the first metal wall defining a first thickness, and wherein the first metal wall is between the light emitting portion and the light receiving portion,wherein the light blocking member is provided with a second metal wall and a third metal wall that are provided in a direction perpendicular to the first metal wall, the second metal wall defining a second thickness and the third metal wall defining a third thickness, and wherein the second metal wall and the third metal wall are spaced from each other a first distance,wherein the first metal wall defines a first length configured to cover ends of the second and third metal walls, and wherein the first length is greater than the second thickness, the third thickness, and the first distance combined; andwherein the first metal wall and the second metal wall are provided so as to be adjacent to each other via a first gap region, and the first metal wall and the third metal wall are provided so as to be adjacent to each other via a second gap region. 2. The optical detection unit according to claim 1, wherein the light blocking member further includesa fourth metal wall that is provided in a direction intersecting the first metal wall and blocks light from being incident to the light receiving portion, andwherein the fourth metal wall is provided with a diaphragm which restricts light from the target object on an optical path between the target object and the light receiving portion. 3. The optical detection unit according to claim 2, wherein the fourth metal wall couples to the first metal wall and defines an aperture. 4. The optical detection unit according to claim 3, wherein the aperture in the fourth metal wall is semi-circular. 5. The optical detection unit according to claim 3, comprising a fifth metal wall coupled to the fourth metal wall, wherein the first metal wall and the fifth metal wall couple to opposite sides of the fourth metal wall. 6. The optical detection unit according to claim 1, wherein a distance LD between the light emitting portion and the light receiving portion satisfies LD&lt;3 mm. 7. The optical detection unit according to claim 6, wherein the distance LD satisfies 0.3 mm&lt;LD&lt;2.5 mm. 8. The optical detection unit according to claim 1, further comprising: a board on which the light emitting portion, the light receiving portion, and the light blocking member are mounted,wherein the light blocking member further includesfirst and second projections that are engaged with holes of the board in order to fix the light blocking member to the board, andwherein the first and second proj... | Seiko Epson Corporation | A61B005/1455 A61B005/024 A61B005/0205 A61B005/00
US09333348 | B2 | 2016-05-10 | Apparatus for determining optimum treatment parameters | 2015-06-22 | A diagnostic device in an electrical interferential treatment regime applies different pulse frequencies in one selection of electrode placement and determines the change of blood flow in response to the change in pulse frequency in an attempt to determine an optimum or workable set of parameters. The device includes electronic storage to record a series of tests and their results and a communication link. The communication link enables a professional caregiver to determine whether the parameter determinations, made at a location remote from the professional's office, such as at the residence of the patient, were conducted and were conducted correctly. The communication link also enables the professional to determine whether the treatments were conducted and conducted correctly. | 1. A system for determining and applying treatment parameters for a trial electrical interferential treatment regimen conducted on a patient situated at a first location, comprising: pairs of electrodes adapted for external temporary attachment, respectively, to opposite limbs of a patient situated at said first location and at least one electronic device at said first location for delivering a series of trial electrical interferential treatments to the patient through the electrodes, and, for a pair of electrode placements, each electrical interferential treatment comprises electrical interferential treatment parameters of at least one beat frequency;a sensor configured to measure at least one biological function responsive to autonomic nervous system activity of the patient caused by each electrical interferential treatment parameter delivered;said at least one electronic device responsive to said sensor and adapted to convert the at least one measured biological function to a numerical value;said at least one electronic device further adapted to compare said numerical value to a predetermined numerical value indicative of a desired biological function;said at least one electronic device further adapted to store data of said numerical values, said treatment parameters of said series of trial electrical interferential treatments and said treatment parameters of said at least one electrical interferential treatment that corresponds to a numerical value that is in excess of said predetermined numerical value;said at least one electronic device further adapted to electronically transmit said stored data to at least one other electronic device situated at a second location away from the immediate vicinity of the first location;said at least one electronic device further adapted to allow selection of said treatment parameters of the at least one electrical interferential treatment that produced said numerical value in excess of the predetermined value and apply the selected treatment parameters through the electrodes to the patient in a subsequent selected treatment regimen;said at least one electronic device being adapted for viewing said numerical values; andsaid at least one other electronic device being adapted for viewing said numerical values, said treatment parameters of said series of trial electrical interferential treatments and said treatment parameters of said identified electrical interferential treatment that produced said numerical value in excess of the predetermined value and determining from said stored data whether said treatment parameters of said identified electrical interferential treatment produced a numerical value in excess of the predetermined value. 2. The system of claim 1 wherein the sensor comprises a digit temperature sensor. 3. The system of claim 2 wherein said digit temperature sensor is adapted to measure a patient's digit temperature prior to delivering the series of trial electrical interferential treatments an... |  | A61N001/36 A61B005/00 A61N001/08 A61N001/32 G06F019/00
US09351683 | B2 | 2016-05-31 | Wrinkle detection method, wrinkle detection device and recording medium storing wrinkle detection program, as well as wrinkle evaluation method, wrinkle evaluation device and recording medium storing wrinkle evaluation program | 2014-03-21 | A wrinkle detection method comprises the steps of obtaining face data of a subject; setting one or more analytical regions on the obtained face data; producing plural pieces of wrinkle component extracted data by extracting wrinkle components extending in one direction or in plural directions at different angles from each other in an angle range in which wrinkles tend to extend and which is set in advance for the set one or more analytical regions; producing plural pieces of wrinkle component emphasized data corresponding to the pieces of wrinkle component extracted data, respectively, by emphasizing the wrinkle components; producing composite data by combining the pieces of wrinkle component emphasized data; and detecting a wrinkle component having an intensity equal to or greater than a predetermined threshold value from the composite data as a wrinkle of the subject. | 1. A wrinkle detection method, comprising the steps of: storing in a memory in advance angle ranges in which wrinkles tend to extend and which are set for respective portions of a face;obtaining face data of a subject;setting one or more analytical regions on the obtained face data;reading out from the memory an angle range corresponding to each of the portions at which the one or more analytical regions have been set;processing each of the one or more analytical regions with a plurality of wrinkle direction edge filters of plural directions at different angles in the read out angle range to produce plural pieces of wrinkle component extracted data by extracting wrinkle components extending in one direction or in plural directions at different angles from each other in the read out angle ranges;producing plural pieces of wrinkle component emphasized data corresponding to the pieces of wrinkle component extracted data, respectively, by emphasizing the wrinkle components;producing composite data by combining the pieces of wrinkle component emphasized data; anddetecting a wrinkle component having an intensity equal to or greater than a predetermined threshold value from the composite data as a wrinkle of the subject. 2. The wrinkle detection method according to claim 1, wherein the face data is obtained by converting image data acquired by imaging a face of the subject into a luminance image and detecting the face from the luminance image, and the wrinkle detection method comprises the steps of: setting the one or more analytical regions on the detected face;producing a plurality of wrinkle component extracted images as the plural pieces of wrinkle component extracted data by, from the set one or more analytical regions, extracting wrinkle components extending in one direction or in plural directions at different angles in an angle range in which wrinkles tend to extend and which is set in advance for the one or more analytical regions;producing a plurality of wrinkle component emphasized images as the plural pieces of wrinkle component emphasized data and corresponding to the wrinkle component extracted images, respectively, by emphasizing the wrinkle components;producing a composite image as the composite data by combining the wrinkle component emphasized images; anddetecting a wrinkle component having an intensity equal to or greater than a predetermined threshold value from the composite image as the wrinkle of the subject. 3. The wrinkle detection method according to claim 2, wherein the wrinkle component emphasized images are obtained by reducing wrinkle orthogonal components orthogonal to the wrinkle components in the wrinkle component extracted images using a plurality of wrinkle orthogonal component extracted images obtained by processing the one or more analytical regions with a plurality of wrinkle orthogonal direction edge filters respectively orthogonal to the wrinkle direction edge filters. 4. The wrinkle detection method according ... | FUJIFILM Corporation | G06T007/00 A61B005/00 A45D044/00
US10123748 | B2 | 2018-11-13 | Active patient risk prediction | 2014-10-30 | Electronic health records of a plurality of patients are received. A risk prediction model for a disease based on the electronic health records of the plurality of patients is created. An electronic health record of an original patient is received. A neighboring group of patients of the plurality of patients is identified, wherein the neighboring group of patients is two or more patients similar to the original patient. An ordering of the two or more patients of the neighboring group of patients is received, wherein the ordering of the two or more patients of the neighboring group of patients is based upon how similar each patient of the two or more patients is to the original patient. The risk prediction model is updated based on the ordering of the two or more patients of the neighboring group of patients. | 1. A method for updating a patient risk prediction model, the method comprising: receiving electronic health records of a plurality of patients;creating, by one or more computer processors, a risk prediction model for a disease based on the electronic health records of the plurality of patients;receiving, by one or more computer processors, an electronic health record of an original patient;identifying, by one or more computer processors, a neighboring group of patients of the plurality of patients, wherein the neighboring group of patients is two or more patients similar to the original patient;generating, by one or more computer processors, a question as to a relative ordering of the two or more patients of the neighboring group of patients with respect to the original patient;receiving an answer to the question as to the relative ordering of the two or more patients of the neighboring group of patients, wherein the ordering of the two or more patients of the neighboring group of patients is based upon how similar each patient of the two or more patients is to the original patient; andupdating, by one or more computer processors, the risk prediction model based on the answer to the question as to the relative ordering of the two or more patients of the neighboring group of patients. 2. The method of claim 1, further comprising: estimating, by one or more computer processors, the risk that the original patient will suffer from the disease based upon the updated risk prediction model. 3. The method of claim 1, wherein the ordering of the patients of the neighboring group of patients is a complete ordering including all patients of the neighboring group of patients or a partial ordering including at least one of the patients of the neighboring group of patients. 4. The method of claim 1, wherein the risk prediction model for a disease based on the plurality of patients is selected from a group consisting of: constrained least square problem, a linear system of equations, or a quadratic program. 5. The method of claim 1, wherein the risk prediction model is updated until a patient similarity converges in a constrained similarity process. 6. The method of claim 1, wherein the risk prediction model is updated more than one time. 7. The method of claim 1, wherein the electronic health record includes one or more of the following: demographics; medical history; medication and allergies; immunization status; laboratory test results; radiology images; vital signs; and personal statistics. 8. A computer program product for updating a patient risk prediction model, the computer program product comprising: one or more computer readable storage media and program instructions stored on the one or more computer readable storage media, the program instructions comprising: program instructions to receive electronic health records of a plurality of patients;program instructions to create a risk prediction model for a disease based on the electronic health ... | International Business Machines Corporation | G16H050/00 A61B005/00 G16H050/30 G16H050/70 G06F019/00 A61B005/055
US09339210 | B2 | 2016-05-17 | Device for obtaining a vital sign of a subject | 2014-04-22 | A device for obtaining a vital sign of a subject despite motion of the subject, in particular for discriminating a vital sign such as a respiratory information signal from noise in a projection based vital signs registration, comprises an interface that receives a set of image frames of a subject. An analysis unit determines the amount of direction changes and/or the time distances between direction changes within a region of interest in a subset of image frames comprising a number of image frames of said set. An evaluation unit determines if said region of interest within said subset of image frames comprises a vital sign information and/or noise by use of the determined amount of direction changes and/or the time distances for said subset of image frames. A processor determines the desired vital sign of the subject from said region of interest within said subset of image frames. | 1. A device for obtaining a vital sign of a subject, comprising: an interface configured to receive a set of image frames of a subject,one or more processors configured to: analyze a subset of image frames of a region of interest in a subject comprising a plurality of contiguous image frames to determine direction changes indicating changes of a direction of motion appearing within said region of interest and at least one of a number of directions per unit time and times between contiguous direction changes,determine two projection profiles in two predetermined orthogonal directions from said contiguous subset of image frames and determine the number of direction changes per unit time from said projection profiles,decide whether said region of interest of said contiguous subset of image frames comprises vital sign information in response to physical motion direction changes appearing at substantially the same moment in time in both projection profiles,extract the vital sign of the subject from said region of interest within said contiguous subset of image frames. 2. A device for obtaining a vital sign of a subject comprising: an interface that receives a subset of contiguous image frames of a subject,an ROI selection unit that selects a region of interest within said contiguous subset of image frames by: determining two projection profiles in two predetermined orthogonal directions from said contiguous subset of image frames,dividing said projection profiles into projection areas,determining a frequency of movement direction changes in the contiguous subset and/or times between contiguous direction changes within said projection areas,selecting the two or more adjacent projection areas having a smallest frequency of direction changes and/or largest times between contiguous direction changes, said selected projection areas representing the region of interest;an evaluation unit that determines if said region of interest within said contiguous subset of image frames comprises vital sign information and/or noise by use of the determined frequency of direction changes and/or the times between contiguous direction changes for said contiguous subset of image frames, anda processor that determines the vital sign of the subject from said region of interest within said contiguous subset of image frames if it is determined that said region of interest within said contiguous subset of image frames comprises vital sign information. 3. The device as claimed in claim 2, wherein said evaluation unit is configured to determine if the determined frequency of direction changes is lower than a predetermined threshold indicating that said region of interest within said contiguous subset of image frames comprises a vital sign information. 4. The device as claimed in claim 2, further comprising a control unit that controls a number of image frames of said subset of image frames used for determining the frequency of direction changes. 5. The device as claimed in cla... | KONINKLIJKE PHILIPS N.V. | G06K009/00 A61B005/08 A61B005/00 A61B005/024 A61B005/11 G06T007/20 A61B005/113
US10687729 | B2 | 2020-06-23 | Apparatus and method for estimating a value of a physiological characteristic | 2015-11-09 | There is provided a measurement apparatus for estimating a value of a physiological characteristic of a subject. The apparatus comprises a radio-frequency (RF) power generating apparatus, a detector, and a controller in communication with the RF power generating apparatus. The RF power generating apparatus comprises a transmit antenna for emitting RF radiation; an RF signal generator; and a transmission line connecting the RF signal generator to the transmit antenna. The detector is arranged to measure the variation of at least one parameter correlated with attenuation of the emitted RF radiation, during the measurement period. The controller is arranged to: cause the RF power generating apparatus to emit RF radiation during a measurement period; receive measurements of the variation of the at least one parameter from the detector; and calculate a value of a physiological characteristic of a subject based on the received measurements. | 1. A measurement apparatus, comprising: a transmit antenna for emitting radio frequency (RF) radiation, wherein the transmit antenna is placed in proximity to a body part of a subject such that a variation in an amount of fluid present in the body part of the subject causes a variation in an impedance of the transmit antenna;an RF signal generator;a transmission line connecting the RF signal generator to the transmit antenna;a detector configured to measure a variation of a parameter correlated with the variation in the impedance of the transmit antenna during a measurement period; anda controller, in communication with the RF signal generator and the detector, and configured to: cause the RF signal generator to provide an RF signal to the transmit antenna such that the transmit antenna emits the RF radiation during the measurement period;receive measurements of the variation of the parameter during the measurement period from the detector; andcalculate a value of a physiological characteristic of the subject based on the measurements,wherein the parameter correlated with the variation in the impedance of the transmit antenna is a level of RF power reflected by an interface between the transmit antenna and the transmission line. 2. The measurement apparatus of claim 1, wherein the detector comprises at least one directional coupler connected to the transmission line. 3. The measurement apparatus of claim 1, wherein the measurement apparatus comprises a smartphone. 4. The measurement apparatus of claim 1, further comprising: a sensor part configured to be attached to the subject; anda separate reader part configured to wirelessly communicate with the sensor part using the RF radiation. 5. The measurement apparatus of claim 1, wherein the RF signal generator is configured to generate the RF radiation having a frequency of one or more of: 2.4 GHz, 2.5 GHz, 5 GHz, and 13.5 MHz. 6. The measurement apparatus of claim 4, wherein the sensor part comprises: the RF power generating apparatus, the detector, and a first wireless communication apparatus. 7. The measurement apparatus of claim 6, wherein the first wireless communication apparatus comprises a near-field communications (NFC) apparatus or a radio-frequency identification (RFID) apparatus. 8. The measurement apparatus of claim 4, wherein the separate reader part comprises the controller and a second wireless communication apparatus. 9. The measurement apparatus of claim 8, wherein the second wireless communication apparatus comprises a near-field communications (NFC) apparatus or a radio-frequency identification (RFID) apparatus. 10. A method, comprising: placing a transmit antenna in proximity to a body part of a subject such that a variation in an amount of fluid present in the body part causes a variation in an impedance of the transmit antenna, wherein the variation in the impedance of the transmit antenna is used to measure a physiological characteristic of the subject;causing a radio... | KONINKLIJKE PHILIPS N.V. | A61B005/05 A61B005/00 A61B090/98 A61B005/0265 A61B005/053 A61B005/024
US09333373 | B2 | 2016-05-10 | Methods for delivery of optical signals to tissue for the treatment of a disease or condition | 2015-11-18 | Embodiments described herein provide methods for treating various conditions and diseases using an optical signal. In one or more embodiments an apparatus is providing having an optical window, which is used to deliver an optical signal to provide stimulation to one or more tissue sites in the body such as the brain, optic nerve, eye, ganglia, spine, or other like site. The optical signals can be used to treat a variety of neurological diseases and conditions including epilepsy, migraine headaches and chronic pain. In particular applications the optical signals can be used to treat, inhibit or prevent epileptic or other neurological seizures by providing an optical input to a foci or surrounding tissue in the brain causing the seizure. The optical signal may also be combined with an electrical signal to produce an aggregate effect in tissue for treating the disease or condition such as a neurological disease or condition. | 1. A method for stimulating neural tissue to treat a neurological condition, the method comprising: providing a stimulator apparatus configured to operate in or on the neural tissue, the stimulator apparatus including a substrate comprising a transparent dielectric material and including an optical window coated on an exterior portion by a transparent polymeric layer comprising a compound configured to minimize attachment of cells to the optical window so as to maintain an optical transparency of the optical window for an extended period of time, the stimulator apparatus configured to generate one or more optical stimulation signals; anddelivering an optical stimulation signal generated by the stimulator apparatus to the neural tissue through the optical window to treat the neurological condition. 2. The method of claim 1, wherein the compound comprises an eluting compound configured to minimize attachment of cells to the optical window, the method further comprising: eluting the compound from the coating for an extended period of time so as to maintain optical transparency of the optical window. 3. The method of claim 2, wherein the compound is eluted from the coating for a period up to ten years so to maintain the transparency of the optical window. 4. The method of claim 1, wherein the optical stimulation signal is delivered to brain tissue to treat the neurological condition. 5. The method of claim 1, wherein at least a portion of the stimulator apparatus is implanted in the skull. 6. The method of claim 1, wherein at least a portion of the stimulator apparatus is implanted in the brain. 7. The method of claim 6, wherein the optical stimulation signal is further delivered to allow for measurement of changes in optical properties of the brain tissue. 8. The method of claim 1, wherein the optical stimulation signal is configured to treat an epileptic seizure. 9. The method of claim 8, wherein the optical stimulation signal is configured to inhibit or prevent the epileptic seizure. 10. The method of claim 8, wherein the optical stimulation signal is delivered to a foci in the brain causing the epileptic seizure. 11. The method of claim 10, wherein the optical stimulation signal is delivered to tissue surrounding a foci in the brain causing the epileptic seizure. 12. The method of claim 1, further comprising: delivering an electrical stimulation signal to brain tissue to treat the neurological condition. 13. The method of claim 12, wherein the electrical stimulation signal is generated by the stimulator. 14. The method of claim 12, wherein the electrical stimulation signal is synchronized with the optical stimulation signal. 15. The method of claim 12, wherein the electrical stimulation signal and the optical stimulation signal produce an aggregate effect in the neural tissue to treat the neurological condition. 16. The method of claim 1, wherein the optical stimulation signal is configured to treat a migraine headache. 17. The me... | InCube Labs, LLC | A61N001/00 A61B005/02 A61N005/06 A61N001/05 A61N001/36 A61N001/375
US11045108 | B2 | 2021-06-29 | Magnetic resonance imaging receive coil assembly | 2014-11-26 | A radiofrequency receive coil assembly can include a first conductive loop and a second conductive loop electrically connected at a node. The first and second conductive loops can extend into a treatment beam region of the radio frequency receive coil assembly through which one or more beams of ionizing radiation pass. The first conductive loop and the second conductive loop can overlap each other to provide electromagnetic isolation and/or can use a common conductor combined with a shared capacitor to provide electromagnetic isolation, with the shared capacitor or other electrical components, as well as any conductive loop overlaps, being positioned outside of the treatment beam region. These features can, among other possible advantages, minimize and homogenize attenuation of the beams of ionizing radiation by the radiofrequency receive coil assembly. | 1. A system comprising: an instrument configured to emit a treatment beam; and a radiofrequency receive coil assembly for use in a magnetic resonance imaging system, the radiofrequency receive coil assembly comprising: a first conductive loop that extends into a treatment beam region of the radiofrequency receive coil assembly through which the treatment beam passes; and a second conductive loop that extends into the treatment beam region, the second conductive loop being electrically connected to the first conductive loop at a node, wherein the node is in the treatment beam region, and the second conductive loop overlapping the first conductive loop outside of the treatment beam region. 2. The radiofrequency receive coil assembly of claim 1, wherein the first conductive loop and the second conductive loop do not overlap each other in the treatment beam region such that the treatment beam passes through only a layer of conductive material in traversing the treatment beam region. 3. The radiofrequency receive coil assembly of claim 1, further comprising a capacitor shared by the first conducting loop and the second conducting loop. 4. The radiofrequency receive coil assembly of claim 3, wherein the first conducting loop and the second conducting loop share a common conductor and wherein the capacitor provides electrical isolation between the first conducting loop and the second conducting loop. 5. The radiofrequency receive coil assembly of claim 1, wherein the treatment beam region is free of any capacitors. 6. The radiofrequency receive coil assembly of claim 1, further comprising a substrate that supports the first conductive loop, the second conductive loop, and an insulating region. 7. The radiofrequency receive coil assembly of claim 6, wherein at least one of the insulating layer and the substrate comprises polyimide. 8. The radiofrequency receive coil assembly of claim 6, further comprising a third conductive loop, wherein the first and second conductive loops are disposed on a first side of the substrate, and the third conductive loop is disposed on a second side of the substrate. 9. The radiofrequency receive coil assembly of claim 8, wherein the third conductive loop extends into the treatment beam region. 10. The radiofrequency receive coil assembly of claim 1, wherein the first conductive loop and/or second conductive loop comprises at least one of copper, silver, and aluminum. 11. The radiofrequency receive coil assembly of claim 1, wherein a linear distance separates centers of the first conducting loop and the second conducting loop, the linear distance being sufficient to cause a mutual inductance between the first conducting loop and the second conducting loop to be below 5% of the inductance of the first conducting loop and the second conducting loop. 12. The radiofrequency receive coil assembly of claim 11, wherein the linear distance is sufficient to cause the mutual inductance between the first conducting loop and ... | VIEWRAY TECHNOLOGIES, INC. | A61B005/055 A61N005/10 G01R033/34 G01R033/48 G01R033/3415
US11376151 | B2 | 2022-07-05 | Gravity independent medical drainage systems | 2015-11-13 | Provided herein are medical drainage systems comprising structures to prevent blockages within drainage tubing, especially medical drainage tubing, for example caused by debris or airlocks. The novel systems of the invention employ, in some embodiments, elements of the Archimedian screw. The Archimedian screw elements comprise a rotating conveyer present in the tubing which promotes continuous flow of material through the tubing and prevents the formation of dependent loops and other blockages. The systems of the invention may be employed in diverse contexts, including urinary catheter drainage systems and thoracostamy systems. The systems of the invention, being gravity independent, may also be employed for the drainage or a movement of materials in space and other low gravity environments. | 1. A medical drainage device, comprising a section of flexible tubing, having a proximal end configured to be disposed towards a patient and a distal end disposed towards a collection vessel, wherein the tubing is configured to allow air and fluid to flow through the tubing and wherein the tubing is subject to formation of air-locks from bends or loops in the tubing;a source of material, which said material enters the proximal end of the tubing;a flexible screw conveyer disposed within the lumen of the tubing and configured topromote unobstructed flow of material through the tubing for drainage of the material from the proximal end of the tubing to the distal end of the tubing by mechanically preventing obstructions and airlocks in the tubing; anda motor which rotates the screw conveyer;whereinthe flexible screw conveyer comprises a planar screw surface and wherein the planar screw surface comprises slits or holes in the planar screw surface, wherein such slits or holes are configured to allow air ingress in order to maintain the tubing at ambient air pressure. 2. The device of claim 1, wherein the source of material is a urinary catheter and the material is urine. 3. The device of claim 1, wherein the source of material is a thoracostamy tube and the material is discharge from the thoracic cavity. 4. The device of claim 1, wherein the flexible screw conveyer comprises a continuous plane screw. 5. The device of claim 1, wherein the flexible screw conveyer comprises a solid screw. 6. The device of claim 1, wherein the flexible screw conveyer comprises a polymeric material. 7. The device of claim 6, wherein the polymeric material is selected from a group consisting of polypropylene, polyethylene, silicone rubber, polyurethane, epoxy resin, polymethylmethacrylate, polyamides, polyimides, elastomers, fluorocarbons, nylons, polyvinyl, poly olefin and closed cell resin foam. 8. The device of claim 1, wherein the flexible screw conveyer comprises a central shaft of one material and a planar screw portion of another material. 9. The device of claim 2, wherein the tubing extends substantially from a drainage outflow port of the urinary catheter to the collection vessel. 10. The medical drainage device of claim 1, wherein the flexible screw conveyer spans substantially the entire length of the section of flexible tubing. 11. The medical drainage device of claim 1, wherein the flexible screw conveyers comprises a central shaft; andwherein the motor is present at the distal end of the tubing and is connected to the central shaft of the flexible screw conveyer. 12. A method of moving material through a medical drainage systemwherein the medical drainage system comprises flexible tubing;wherein the tubing is configured to allow air and fluid to flow through the tubing and wherein the tubing is subject to formation of air-locks from bends or loops in the tubing;the method comprising the actions ofconnecting the proximal end of the tubing to a source ... | - | A61M027/00 A61F005/44 A61M001/00
US10448840 | B2 | 2019-10-22 | Apparatus for generating data output containing physiological and motion-related information | 2016-01-15 | A method of generating a data string containing physiological and motion-related information includes sensing physical activity of a subject via at least one motion sensor attached to the subject, sensing physiological information from the subject via at least one photoplethysmography (PPG) sensor attached to the subject, and processing signals from the at least one motion sensor and signals from the at least one PPG sensor into a serial data string of physiological information and motion-related information. A plurality of subject physiological parameters can be extracted from the physiological information, and a plurality of subject physical activity parameters can be extracted from the motion-related information. The serial data string is parsed out such that an application-specific interface can utilize the physiological information and motion-related information for an application that generates statistical relationships between subject physiological parameters and subject physical activity parameters in the physiological information and motion-related information. | 1. A photoplethysmography (PPG) sensor module, comprising: at least one optical emitter configured to direct optical energy to a region of a body of a subject wearing the sensor module;at least one optical detector configured to sense optical energy absorbed, scattered and/or reflected by the body of the subject and generate a PPG signal;at least one motion sensor configured to sense footstep motion of the subject and to generate a motion signal; andat least one processor in electronic communication with the at least one optical detector and the at least one motion sensor, wherein the at least one processor is configured to: process the PPG signal from the at least one optical detector and the motion signal from the at least one motion sensor to remove footstep motion artifacts from the PPG signal;generate a blood flow pulse signal from the processed PPG signal, wherein the blood flow pulse signal has reduced motion artifacts from footstep noise;process the blood flow pulse signal to extract a plurality of subject physiological parameters, wherein the plurality of subject physiological parameters comprise at least subject heart rate and subject heart rate variability information, and wherein at least one subject physiological parameter is extracted using a peak finding algorithm;process the motion signal to extract a plurality of subject physical activity assessments; andgenerate a multiplexed serial data output of the extracted subject physiological parameters and the extracted subject physical activity assessments, wherein the multiplexed serial data output is parsed out such that an application-specific interface can utilize the physiological parameters and the physical activity assessments for an application. 2. The PPG sensor module of claim 1, wherein the PPG sensor module comprises light transmissive material in optical communication with the at least one optical emitter, wherein the light transmissive material comprises at least one optical lens in optical communication with the at least one optical emitter, and wherein the PPG sensor module is configured to be integrated into a wearable device worn at a limb or digit of the subject. 3. The PPG sensor module of claim 2, wherein the PPG sensor module is integrated within a single circuit board of the wearable device. 4. The PPG sensor module of claim 1, wherein the at least one optical emitter comprises a plurality of optical emitters wherein at least one of the plurality of optical emitters is configured to direct optical energy in a direction that is substantially perpendicular to a direction of optical energy emitted by another one of the plurality of optical emitters. 5. The PPG sensor module of claim 4, wherein at least one of the plurality of optical emitters emits optical energy at a different optical wavelength than another one of the plurality of optical emitters. 6. The PPG sensor module of claim 4, further comprising light transmissive material in optical communication with... | Valencell, Inc. | A61B005/00 A61B005/0205 G06F019/00 A61B005/024 H04R001/10 A61B005/11 A61B005/1455 A61B005/0476 A61B005/16 F21V008/00 A61B005/021 A61B005/08 A61B005/01 A61B005/026 A61B005/145 H04R001/02 G16H040/63 G16H040/67 G16H050/20 G16H030/40 A61B005/0295 A61B005/091
US09549681 | B2 | 2017-01-24 | Matrix-based patient signal analysis | 2014-11-18 | Disclosed herein is a framework for facilitating patient signal analysis. In accordance with one aspect, the framework generates a signature cycle matrix corresponding to a single cycle of the patient signal data by segmenting the patient signal data into multiple portions. The signature cycle matrix may further be segmented into sub-matrices corresponding to different levels. Properties of the sub-matrices may then be determined to generate diagnostic results. | 1. A system for patient signal analysis, comprising: a patient monitor configured to acquire a patient signal data from a patient; anda computer system communicatively coupled to the patient monitor, wherein the computer system includesa non-transitory memory device,a processor device in communication with the memory device, anda signal analysis unit stored in the memory device and operative with the processor device to perform steps including generating a signature cycle matrix corresponding to a single cycle of the patient signal data by segmenting the patient signal data into multiple portions,segmenting the signature cycle matrix into sub-matrices corresponding to different levels,determining properties of the sub-matrices,determining a sequential energy index based on the properties of the sub-matrices, wherein the sequential energy index comprises a sum of squared magnitudes corresponding to the sub-matrices, andcharacterizing an abnormality based on the properties and the sequential energy index. 2. The system of claim 1 wherein the signal analysis unit is further operative with the processor device to generate the signature cycle matrix by segmenting each of the multiple portions into sub-portions,generating wave vectors based on values extracted from the sub-portions, andgenerating the signature cycle matrix based on the wave vectors. 3. The system of claim 2 wherein the values comprise voltage amplitude values. 4. The system of claim 3 wherein the voltage amplitude values comprise a maximum voltage amplitude value of a corresponding portion. 5. The system of claim 2 wherein the values comprise time durations. 6. The system of claim 1 wherein the sub-matrices comprise: a first level sub-matrix that represents existence of a P wave in the cycle;a second level sub-matrix that represents both rising edges of the P wave and a Q wave in the cycle;a third level sub-matrix that represents rising edges of the P wave, the Q wave and an R wave in the cycle;a fourth level sub-matrix that represents information from the P wave, the Q wave, the R wave and an S wave in the cycle; anda fifth level sub-matrix that represents information from the P wave, the Q wave, the R wave, the S wave and a T wave in the cycle. 7. The system of claim 1 wherein the properties comprise trace or determinant values. 8. The system of claim 7 wherein the squared magnitudes comprise magnitudes of trace values corresponding to the sub-matrices. 9. The system of claim 7 wherein the squared magnitudes comprise magnitudes of determinant values corresponding to the sub-matrices. 10. The system of claim 1 wherein the abnormality comprises cardiac arrhythmia or myocardial ischemia. | SIEMENS MEDICAL SOLUTIONS USA, INC. | A61B005/04 A61B005/0432 A61B005/0456 A61B005/0472 A61B005/044 A61B005/00
US20160120414 | A1 | 2016-05-05 | Monitoring Activity of a User in Locomotion on Foot | 2016-01-12 | A device supported by a user while the user is in locomotion on foot during an outing to automatically measure amounts of time taken by the user to complete respective distance intervals. During the outing, data representing the automatically measured amounts of time may be automatically stored in memory of the device. During the outing, the device may display at least one of an average pace of the user since a most recently completed distance interval, an average speed of the user since the most recently completed distance interval, and a projected time to complete a current distance interval based on monitored performance since the most recently completed distance interval. The device may determine a remaining distance to be traveled to complete a specified goal distance for the outing and a calculated fraction of the specified goal distance for the outing. | 1. A unitary apparatus configured to be worn by a user, comprising: a unitary body, further comprising: a processor;a user interface;a sensor in operative communication with the processor, and configured to capture motion data indicative of an elevation of the user; anda non-transitory computer-readable medium comprising computer-executable instructions that when executed by the processor perform at least: receiving the motion data from the sensor;identifying, from the motion data, a change in elevation of the user over a period of time;estimating a heart rate of the user by correlating to the identified change in elevation of the user over the period of time; andgenerating a graphical representation of the estimated heart rate and outputting the graphical representation to the user interface. 2. The unitary apparatus of claim 1, wherein rates of changes in elevation correlated to heart rates are stored in the non-transitory computer-readable medium. 3. The unitary apparatus of claim 2, wherein the unitary apparatus is calibrated by: measuring heart rates of the user at different elevations. 4. The unitary apparatus of claim 3, wherein the measured heart rates are received from a heart rate monitor during a calibration process. 5. The unitary apparatus of claim 1, wherein the user interface is configured to display real-time activity information associated with motion of the user. 6. The unitary apparatus of claim 5, wherein the real-time activity information displayed by the user interface includes the estimated heart rate of the user. 7. The unitary apparatus of claim 5, wherein the real-time activity information displayed by the user interface includes an average speed of the user. 8. The unitary apparatus of claim 1, wherein the unitary apparatus is configured to be worn on a wrist of the user. 9. The unitary apparatus of claim 8, wherein the unitary apparatus is further configured to calculate an estimated heart rate of a second user using received second motion data for the second user. 10. An apparatus, comprising: a processor;an interface;a sensor configured to capture motion data indicative of an elevation of a user; andmemory storing computer-readable instructions that, when executed by the processor, cause the apparatus to: receive the motion data from the sensor;identify, from the motion data, a change in elevation of the user over a period of time;estimate a heart rate of the user by correlating to the identified change in elevation of the user over the period of time; andoutput the estimated real-time heart rate to the interface. 11. The apparatus of claim 10, wherein the computer-readable instructions, when executed by the processor, are configured to calibrate the apparatus by: measuring, using a hear rate monitor, heart rates of the user at a plurality of different elevations. 12. The apparatus of claim 10, wherein the computer-readable instructions, when executed by the processor, are further configured to: calculate ... |  | A61B005/0205 A61B005/00
US20160120434 | A1 | 2016-05-05 | WIRELESS PHYSIOLOGICAL MONITORING DEVICE AND SYSTEMS | 2015-10-30 | The present disclosure relates to a cardiac monitoring system and methods for using such a system. Preferred embodiments detect and record cardiac information via a wearable device, then extract data features from the recorded cardiac information. The extracted data features may then be analyzed and used in clinical diagnosis. | 1-4. (canceled) 5. A wearable physiological monitoring device for monitoring physiological signals in a mammal, the device comprising: a housing forming a watertight enclosure, the housing integral with two flexible wings extending from the housing, the flexible wings configured to adhere to the skin of the mammal;wherein the flexible wings comprise electrodes configured to measure a physiological signal;a processor contained within the housing, the processor configured to continuously record and extract a plurality of features from the physiological signal; anda transmitter contained within the housing, the transmitter configured to transmit the features and a plurality of references to the locations of the features to a computing device, the computing device configured to identify a plurality of positions of identified signals from the extracted features and references. 6. The device of claim 5, wherein the plurality of features are extracted during an extraction window. 7. The device of claim 5, wherein the features comprise R peak information. 8. The device of claim 7, wherein the R peak information comprises R peak locations. 9. The device of claim 5, wherein the identified signals comprise heart rhythms requiring a clinical intervention. 10. The device of claim 9, wherein clinical intervention comprises one of a change in dosage regimen, a surgical procedure, and a surgical implantation. 11. The device of claim 9, wherein the clinical intervention comprises continuing a therapy. 12. The device of claim 5, wherein the computing device is a server. 13. The device of claim 5, wherein the positions of the identified signals are re-transmitted from the computing device to the processor. 14. The device of claim 13, wherein the processor is configured to read the positions of the identified signals and access an area of memory corresponding to the positions of the identified signals to infer a rhythm segment. 15. The device of claim 14, wherein the transmitter is further configured to transmit the rhythm segment. 16. The device of claim 15, wherein the transmitter is configured to transmit the rhythm segment only if the rhythm segment meets a confidence determination, the confidence determination comprising exceeding a defined threshold. 17. The device of claim 5, wherein the computing device is configured to be in communication with a smartphone. 18. The device of claim 5, wherein the computing device is a smartphone. 19. The device of claim 5, wherein the processor is configured to collect a secondary physiological signal and the transmitter is configured to transmit the secondary physiological signal to the computing device. 20. The device of claim 19, wherein the secondary physiological signal comprises a motion signal. 21. The device of claim 19, wherein the secondary physiological signal comprises acoustic noise 22. The device of claim 19, wherein the processor is further configured to collect a tertiary physiological signa... |  | A61B005/0408 A61B005/04 A61B005/00 A61B007/04 A61B005/0456 A61B005/0205
US20160128616 | A1 | 2016-05-12 | RETRACTABLE FLUID COLLECTION DEVICE | 2016-01-15 | Retractable medical device including a body having an open back end and a front end. A movable member is arranged within the body. The movable member has a back end and a front end. A needle holding member is movable to a retracted position within the movable member when the movable member experiences at least one of: the front end of the movable member is caused to expand generally radially; the back end of the movable member is caused to contract generally radially; the front end of the movable member is caused to expand generally radially when the back end of the movable member is caused to contract generally radially; and axial movement caused by a cap closing off the back end of the body which engages a tapered surface of the back end of the movable member. | 1. A retractable medical device comprising: a body having an open back end and a front end;a movable member arranged within the body, the movable member having a back end and a front end;a needle holding member movable to a retracted position within the movable member when the movable member experiences at least one of: the front end of the movable member is caused to expand generally radially;the back end of the movable member is caused to contract generally radially;the front end of the movable member is caused to expand generally radially when the back end of the movable member is caused to contract generally radially;axial movement caused by a cap closing off the back end of the body which engages a tapered surface of the back end of the movable member. 2. The device of claim 1, wherein the body and the movable members are at least one of generally cylindrical and generally tubular. 3. The device of claim 1, wherein the movable member is sized and configured to at least partially receive therein one of a fluid collection tube and an evacuated blood collection receptacle. 4. The device of claim 1, wherein an outer peripheral surface of the needle holding member frictionally engages with an inner surface of the front end of the movable member. 5. The device of claim 1, wherein an outer peripheral surface of the needle holding member comprises at least one projection which engages with at least one recess arranged on an inner surface of the front end of the movable member. 6. The device of claim 1, wherein an outer peripheral surface of the needle holding member comprises at least one recess which engages with at least one projection arranged on an inner surface of the front end of the movable member. 7. The device of claim 1, wherein the body comprises a device arranged within the body for limiting axial movement of the movable member within the body. 8. The device of claim 1, further comprising a cap structure and arranged to close off a back end of the body. 9. The device of claim 1, further comprising a cap having a portion which contacts the back end of the movable member. 10. The device of claim 1, further comprising a cap having tapered surfaces which contact tapered surfaces of the back end of the movable member. 11. The device of claim 1, further comprising a cap structured and arranged to lock to the back end of the body. 12. The device of claim 1, further comprising a cap structured and arranged to cause the needle holding member to move to the retracted position when the cap is locked to the body. 13. The device of claim 1, further comprising a cap structured and arranged to cause the needle holding member to move to the retracted position when the cap closes off the back end of the body. 14. The device of claim 1, further comprising a cap structured and arranged to cause the needle holding member to move to the retracted position when the cap engages with a back end of the movable member. 15. The device of claim 1, whe... | STAT MEDICAL DEVICES, INC. | A61B005/15
US20160120459 | A1 | 2016-05-05 | Wound Acknowledgement Interface | 2015-10-29 | Systems, tools, and methods for providing wound data and/or information to a remote location for clinical assessment and acknowledgement are disclosed. | 1. A system for capturing an image of a wound area of a subject, comprising: a. an instrument adapted to capture image data of a wound area of a subject;b. a wand attached to the instrument; andc. a wand tip adapted to be placed against at least one of a surface of the subject's skin adjacent the wound area and a surface of the wound area. 2. The system of claim 1, wherein the instrument comprises a camera. 3. The system of claim 2, wherein the wand is removably secured to the camera. 4. The system of claim 1, wherein the wand tip comprises an analgesic compound or other medicament. 5. The system of claim 4, wherein the wand and wand tip are integral. 6. The system of claim 4, wherein the wand tip is removable from the wand. 7. The system of claim 1, wherein the wand comprises a single point contact device adapted to enable a user to manipulate the relative position of the instrument with respect to the wound area in a repeatable manner without moving the wand tip. 8. A kit, comprising the system of claim 1. 9. The kit of claim 8 further comprising a resealable package comprising a plurality of removable, disposable sterile wand tips. 10. The kit of claim 8 further comprising instructions for capturing image data. 11. A system for providing wound information to a remote location, comprising: a. an instrument adapted to capture image data of a wound area of a subject;b. a processing unit associated with the instrument adapted to store, process, and/or transfer the image data to a remote location; andc. an interface adapted to receive the image data and display an image comprising the image data,wherein the image provides information on the interface to enable a user to clinically assess the wound area of the subject based on wound area appearance, wound area size, wound area condition, wound area topography, and wound type. 12. The system of claim 11, wherein the image comprises image data comprising two or more of wound area appearance, wound area size, wound area temperature, and wound area topography. 13. The system of claim 11, wherein the image provided on the interface comprises a scale along X and Y axes of the image. 14. The system of claim 11 further comprising an image standard for distance calibration of the instrument. 15. The system of claim 14, wherein the processing unit is adapted to calculate wound area size of the image based on the distance calibration of the instrument. 16. A method of providing wound information to a remote location, comprising: a. capturing an image of a wound area of a subject using an instrument, wherein the image provides information on a remote interface to enable a remote user to clinically assess the wound area of the subject based on at least wound area appearance, wound area size, wound area condition, wound area topography, and wound type;b. storing the image in a processing unit;c. transferring the image to a remote location;d. displaying the image on the remote interface; ande. recor... |  | A61B005/00 A61B005/107 A61B005/01
US09345408 | B2 | 2016-05-24 | Oral screening toothbrush | 2014-03-15 | An oral screening toothbrush is provided for screening for oral diseases such as oral cancer and dental caries. The toothbrush emits blue light into the oral cavity and captures an image of the reflection of blue light. Image processing techniques are used to analyze the images to screen for oral diseases. Further, the toothbrush is comprised in a system including a data transfer medium (i.e. “smartphone”) and the Cloud, which allows for data transfer between multiple platforms from the toothbrush. The oral screening toothbrush and extended system provide for in-home oral screening and interaction with health professionals. | 1. An oral screening toothbrush, comprising: a handle;a brush head having a plurality of bristles;a neck that is arranged to be distal to the handle and proximal to the brush head;a data processing unit that is configured to store and process data;a transceiver that is configured to transmit data;at least one light-emitting diode that is configured to emit light into the oral cavity; andat least one photo detector that is located on an external surface of the toothbrush and configured to receive the emitted light that is reflected from the surface of the oral cavity and transmit photographic data. 2. The oral screening toothbrush of claim 1, wherein the transceiver is further configured to receive data. 3. The oral screening toothbrush of claim 1, wherein at least one light-emitting diode is configured to emit blue light at a wavelength of 450-495 nanometers. 4. The oral screening toothbrush of claim 1, wherein the data processing unit is configured to control the intensity and activation of at least one light-emitting diode. 5. The oral screening toothbrush of claim 1, wherein multiple light-emitting diodes are configured in a cluster. 6. The oral screening toothbrush of claim 1, wherein the brush head comprises at least one light-emitting diode. 7. The oral screening toothbrush of claim 1, wherein the handle comprises at least one light-emitting diode. 8. The oral screening toothbrush of claim 1, wherein the neck comprises at least one light-emitting diode. 9. The oral screening toothbrush of claim 8, further comprising a waveguide that is configured to deliver light from at least one light-emitting diode to the exterior of the oral screening toothbrush. 10. The oral screening toothbrush of claim 1, wherein the photo detector is selected from the group consisting of an active pixel sensor, a charge-coupled device, a reverse-biased LED, a photoresistor, a photovoltaic cell, a phototransistor, and a quantum dot photo detector, or any combination thereof. 11. The oral screening toothbrush of claim 1, wherein the photo detector is an active pixel sensor comprising a complementary metal-oxide-semiconductor (CMOS) sensor. 12. The oral screening toothbrush of claim 1, wherein the photo detector is configured to couple to at least one optical fiber. 13. The oral screening toothbrush of claim 1, wherein the photo detector transmits photographic data to the data processing unit that is configured to process the photographic data for oral conditions. 14. The oral screening toothbrush of claim 13, wherein the oral condition is selected from the group consisting of oral cancer and dental caries, or any combination thereof. 15. An oral screening system, comprising: an oral screening toothbrush having a handle, a brush head with a plurality of bristles, a data processing unit that is configured to store and process data, a transceiver that is configured to transmit data, at least one light-emitting diode that is configured to emit light into the... | BEAM TECHNOLOGIES, LLC | A61B005/00 A46B015/00
US09345416 | B2 | 2016-05-24 | Means for securing a catheter into a vessel | 2014-02-27 | An electrophysiology catheter, e.g., a coronary sinus catheter, for insertion into a cardiac vessel, such as the coronary sinus, includes a handle and a catheter shaft coupled at one end to the handle. The catheter shaft has a distal end and an anchor is associated with the catheter shaft and is movable between a deployed position and a collapsed position. In the deployed position, the anchor extends radially outward from an outer surface of the catheter shaft for contacting a wall and temporarily anchoring the catheter shaft within the coronary sinus. The catheter also includes an actuator for causing deployment and collapsing of the anchor upon manipulation of the actuator. | 1. An electrophysiology catheter for insertion into a cardiac vessel comprising: a handle;a catheter shaft coupled at one end to the handle;an anchor comprising: a plurality of wires arrayed around the catheter shaft; anda collar disposed around the catheter shaft and slideable along the catheter shaft, the collar attached to a first end of each of the plurality of wires,wherein the plurality of wires are movable between a deployed position and a collapsed position based on movement of the collar relative to the catheter shaft,wherein in the deployed position, the plurality of wires extend radially outward from an outer surface of the catheter shaft for contacting a wall of the cardiac vessel and temporarily anchoring the electrophysiology catheter shaft within the cardiac vessel while permitting blood flow between the plurality of wires and through the cardiac vessel, andwherein the anchor serves as a recording electrode;a plurality of electrodes that are disposed along the catheter shaft, wherein the anchor is disposed between a first set of electrodes and a second set of electrodes;a flexible elongate mandrel that extends within the catheter shaft and is coupled at one end to an actuator and at an opposite end to the collar, wherein the mandrel is configured to carry electrical signals from the anchor; andthe actuator for causing deployment and collapsing of the plurality of wires within the cardiac vessel. 2. The catheter of claim 1, wherein the electrophysiology catheter comprises a coronary sinus catheter and the cardiac vessel comprises a coronary sinus. 3. The catheter of claim 1, wherein a second end of each of the plurality of wires is attached to the catheter shaft. 4. The catheter of claim 1, wherein the plurality of wires are electrically isolated from the plurality of electrodes that are disposed along the catheter shaft. 5. The catheter of claim 4, further comprising a plurality of separate electrical wires routed to the plurality of wires of the anchor. 6. The catheter of claim 5, wherein the plurality of wires are insulated from one another by an insulation coating. 7. The catheter of claim 1, wherein the actuator is a slide actuator disposed within the handle, wherein a sliding action of the actuator is translated through the mandrel to the collar. 8. The catheter of claim 7, further including an actuator lock mechanism for locking the actuator in first and second positions that correspond respectively to the anchor being in a fully deployed position and a fully collapsed as well as intermediate positions therebetween. 9. The catheter of claim 1, wherein the plurality of electrodes disposed along the catheter shaft are recording electrodes. 10. The catheter of claim 9, wherein the recording electrodes are configured to be operably connected to a recording device. 11. The catheter of claim 9, wherein the recording electrodes are configured to be operably connected to a 3D mapping system. 12. The catheter of claim 1, wh... | Boston Scientific Scimed Inc. | A61B005/042 A61B005/00 A61M025/04
US09348322 | B2 | 2016-05-24 | Smart device including biometric sensor | 2014-08-19 | An apparatus can include a wrist worn device configured to be worn on a wrist of a user. The apparatus can include a controller. The apparatus can include a power supply. The apparatus can include a light emitter that can emit light from a user side of the wrist worn device to a wrist of the user. The apparatus can include a light detector that can detect light reflected from the wrist of the user from the first light emitter and can send a detector signal to the controller. The detector signal can be based on the detected light. The apparatus can include a lens coupled to a user side of the wrist worn device external to the light emitter and light detector. The lens can include an opaque section. The lens can also include light transmissive section that transmits light from the light emitter to the user. | 1. A user portable apparatus comprising: a wrist worn device configured to be worn on a wrist of a user;a controller coupled to the wrist worn device;a power supply coupled to the controller;a light emitter coupled to the controller, the light emitter configured to emit light from a user side of the wrist worn device to a wrist of the user;a light detector coupled to the controller, the light detector configured to detect light reflected from the wrist of the user from the first light emitter, the light detector configured to send a detector signal to the controller, where the detector signal is based on the detected light; anda lens coupled to a user side of the wrist worn device external to the light emitter and the light detector, the lens including an opaque section, and the lens including a light transmissive section that transmits the light emitted from the light emitter to a wrist of the user. 2. The apparatus according to claim 1, wherein the wrist worn device comprises one of a watch and a smart watch. 3. The apparatus according to claim 1, wherein the light transmissive section comprises a first light transmissive section that transmits light from the light emitter to the wrist of the user, andwherein the lens includes a second light transmissive section separate from the first light transmissive section, where the second light transmissive section transmits light reflected from the wrist of the user to the light detector. 4. The apparatus according to claim 3, wherein the lens includes a third light transmissive section separate from the first light transmissive section and the second light transmissive section;wherein the light emitter comprises: a first light emitter configured to emit light through the first light transmissive section;a second light emitter configured to emit light through the third light transmissive section; andwherein the light detector is configured to detect light emitted from the first light emitter and the second light emitter. 5. The apparatus according to claim 1, further comprising a display coupled to the wrist worn device, the display configured to display at least clock information,wherein the controller sends a user biometric signal to the display based on the detector signal, andwherein the display displays biometric information based on the biometric signal. 6. The apparatus according to claim 1, comprising a wireless transmitter configured to transmit a wireless communication signal based on the detector signal generated by the light detector. 7. The apparatus according to claim 3, wherein the wireless communication signal comprises at least one of a wireless near field communication signal, a wireless local area network communication signal, and a cellular communication signal. 8. The apparatus according to claim 1, further comprising a charging coil surrounding the light emitter and the light detector, the charging coil configured to charge the power supply. 9. The apparatus according to c... | Google Technology Holdings LLC | G01N033/48 G04G021/02 H04B005/00 H02J007/02 G01J001/04 A61B005/1455 A61B005/024
US09351758 | B2 | 2016-05-31 | Methods and apparatuses for fluoro-less or near fluoro-less percutaneous surgery access | 2015-04-03 | A needle access assembly and method for obtaining percutaneous needle access with little or no fluoroscopy. The method can include selecting a target for percutaneous access, directing a laser guide at a desired needle-insertion angle and in line with the selected target, aligning the needle access assembly with the laser, and inserting the needle into the target. | 1. A needle access device configured for insertion into a patient with reduced fluoroscopy, the device comprising: a needle connected to a hub portion, the hub portion comprising: an opaque cap portion positioned at a proximal end of the hub such that light from a proximal light source is first incident on the opaque cap portion;a non-opaque body portion positioned between the opaque cap portion and the needle, the non-opaque body portion being at least partially cylindrical and having an interior space forming an inner illumination chamber within the non-opaque body portion, the inner illumination chamber of the non-opaque body portion comprising a translucent, lateral wall portion through which light from the proximal light source can be visualized;a proximal-facing, reflective surface positioned within the inner illumination chamber and at a distal portion of the inner illumination chamber, the proximal-facing, reflective surface comprising a curvature that is convex in a proximal direction; anda channel having an axis and extending through the opaque cap portion and comprising an inlet and an outlet, the inlet positioned at a proximal surface of the hub portion and the outlet positioned distally of the inlet, the channel having a length to width ratio such that the inner illumination chamber of the non-opaque body portion only illuminates when the proximal light source is aligned with the axis of the channel,wherein light from the proximal light source reflects off the curvature of the proximal-facing, reflective surface and toward the translucent, lateral wall portion to illuminate the inner illumination chamber only when the proximal light source is aligned with the channel, such that the proximal light is visible to a user through the translucent, lateral wall portion of the inner illumination chamber. 2. The needle access device of claim 1, wherein the channel has a uniform diameter from the inlet to the outlet. 3. The needle access device of claim 1, wherein the channel has a diameter that is less than about 20% of an outer diameter of the opaque cap portion. 4. The needle access device of claim 1, wherein an outer diameter of the non-opaque body portion is less than an outer diameter of the opaque cap portion. 5. The needle access device of claim 1, wherein the channel has a diameter that is less than or equal to an outer diameter of the needle. 6. The needle access device of claim 1, wherein the reflective surface comprises a reflective coating. 7. The needle access device of claim 1, wherein the reflective surface comprises a dome reflector. 8. The needle access device of claim 1, further comprising at least two concentric circles disposed on the proximal surface of the hub portion. 9. The needle access device of claim 1, further comprising a crosshatch disposed on the proximal surface of the hub portion. 10. The needle access device of claim 1, wherein the hub portion further comprises a luer connector configured to connect ... | FACULTY PHYSICIANS AND SURGEONS OF LOMA LINDA UNIVERSITY SCHOOL OF MEDICINE: LOMA LINDA UNIVERSITY | A61B017/34 A61B005/06 A61B001/00 A61B001/307 A61B019/00 A61B001/018 A61B017/00 A61M025/00 A61B001/06 A61B008/08 A61B006/00 A61M025/10 A61B006/12
US09333085 | B2 | 2016-05-10 | Patient selectable knee arthroplasty devices | 2013-09-30 | Disclosed herein are methods and devices for repairing articular surfaces in a knee joint. The articular surface repairs are customizable or highly selectable for each patient and geared toward providing optimal fit and function. Kits are also provided to enable customized repairs to be performed. | 1. A method of generating an articular repair system including an articular implant having an inner surface for attachment to a bone of a patient's joint at an implantation site and an outer surface for providing an articular surface, the method comprising: deriving a shape of cartilage and/or subchondral bone surface from electronic image data of the patient's joint;determining a desired shape for at least a portion of the outer surface of the articular implant at least in part from the derived shape of the cartilage;selecting the articular implant from a library of implants such that at least a portion of the outer surface of the articular implant approximates the desired shape, wherein at least a portion of the inner surface of the articular implant is configured to engage with a chamfer cut on the bone of the patient's joint; andproviding a surgical tool in the articular repair system, wherein the surgical tool is created based, at least in part, on the derived shape of the cartilage and/or subchondral bone surface, and configured to facilitate the implantation of the articular implant into the patient's joint. 2. The method of claim 1, wherein the patient's joint is a knee joint, wherein the articular implant is a tibial implant. 3. The method of claim 1, wherein the patient's joint is a knee joint, wherein the articular implant is a femoral implant that has only one condylar portion. 4. The method of claim 3, wherein the condylar portion is a first condylar portion and the articular implant further comprises a second condylar portion, wherein the first and second condylar portions are disposed on opposite sides of the trochlear portion. 5. The method of claim 4, wherein a bone-facing implant surface of the trochlear portion bone is configured to abut at least a portion of uncut trochlear bone upon implantation of the implant. 6. The method of claim 4, wherein the bone-facing implant surface of the trochlear portion bone is configured to abut at least a portion of cut trochlear bone upon implantation of the implant. 7. The method of claim 1, wherein the library includes implants of various sizes, shapes and/or thicknesses. 8. The method of claim 1, wherein the selected articular implant is further manipulated to match the patient's joint. 9. The method of claim 8, wherein the selected articular implant is further manipulated by mechanical abrasion. 10. The method of claim 8, wherein the selected articular implant is further manipulated by, at least in part, an automated process. 11. The method of claim 8, wherein the selected articular implant is further manipulated by, at least in part, a manual process. 12. A method of generating an articular repair system including an articular implant having an inner surface for attachment to a bone of a patient's joint at an implantation site and an outer surface for providing an articular surface, the method comprising: deriving an articular shape from electronic image data of the patient's ... | ConforMIS, Inc. | A61F002/28 A61F002/02 A61F002/38 A61F002/30 G06T007/00 A61B005/00 A61F002/46
US10441182 | B2 | 2019-10-15 | Magnetic resonance imaging apparatus and magnetic resonance imaging method | 2014-07-08 | Provided is a technique in MRI to efficiently suppress downstream blood in a specific region in a blood vessel having a slow flow velocity, such as the portal vein. For this purpose, a plurality of Beam Sat pulses are applied so as to equally suppress signals of blood flowing into a desired imaging region from a desired blood vessel during the period from applying an IR pulse to starting main imaging. Downstream blood in a specific region in a blood vessel having a slow flow velocity, such as the portal vein, can be suppressed efficiently by determining application conditions of the plurality of Beam Sat pulses that achieve the above based on a flow velocity of blood in a desired blood vessel and T1 of the said blood. | 1. A magnetic resonance imaging apparatus including: an imaging sequence generation section configured to apply imaging conditions to a predetermined pulse sequence to generate an imaging sequence; andan imaging section configured to collect echo signals from an imaging region according to the imaging sequence to reconstruct an image from the echo signals,wherein the predetermined pulse sequence includes an IR (Inversion Recovery) pulse and a plurality of two-dimensional selective high-frequency pulses,wherein the imaging conditions include application conditions of the plurality of two-dimensional selective high-frequency pulses,wherein the imaging sequence generation section includes an application condition determination part that determines whether or not set application conditions suppress blood signals that are the echo signals from blood in a targeted blood vessel,wherein in response to the echo signals are suppressed, the set application conditions are determined as applied application conditions that are the application conditions to be applied to the imaging sequence, andwherein in response to the echo signals are not suppressed, the set application conditions changed to suppress the blood signals are determined as the applied application conditions. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the application conditions include a total application time and an application interval of the plurality of two-dimensional selective high-frequency pulses as well as an amplitude and a pulse width of each of the plurality of two-dimensional selective high-frequency pulses. 3. The magnetic resonance imaging apparatus according to claim 2, wherein the application condition determination part determines: the total application time so that the total application time does not exceed a predetermined period determined by a T1 recovery time of blood while suppressing the echo signals from the blood in the targeted blood vessel;the application interval so that the blood flowing into the imaging region is continuously suppressed;the amplitude so as not to exceed a limitation by SAR (Specific Absorption Rate); andthe pulse width according to the determined amplitude while maintaining an application area of a two-dimensional selective high-frequency pulse. 4. The magnetic resonance imaging apparatus according to claim 2, further including: a body motion detection section that detects a displacement amount by periodic body motion of an object after the peak; andan adoption/rejection determination section that determines the adoption or rejection of the echo signals collected by the imaging section according to the relationship between the displace amount and an application start timing of a two-dimensional selective high-frequency pulse specified by the applied application conditions,wherein the imaging sequence is a body motion synchronization sequence, andwherein the imaging section reconstructs the image only from the echo ... | Hitachi, Ltd. | A61B005/026 A61B005/055 A61B005/00 G01R033/483 G01R033/56 G01R033/565 G01R033/567 A61B005/113
US10022091 | B2 | 2018-07-17 | User interface for artifact removal in an EEG | 2016-01-09 | A method and system for a user interface for artifact removal in an EEG is disclosed herein. The invention allows an operator to select a plurality of artifacts to be automatically removed from an EEG recording using a user interface. The operator pushes a button on the user interface to apply a plurality of filters to remove the plurality of artifacts from the EEG and generate a clean EEG for viewing. | 1. A method for displaying electroencephalogram (EEG) data, the method comprising: with a processor performing artifact reduction on an original EEG signal comprising a first plurality of channels to generate a processed continuous EEG report, the processed continuous EEG report comprising a second plurality of channels, wherein performing artifact reduction comprises partitioning the original EEG signal into a plurality of epochs, performing artifact reduction on the plurality of epochs to generate a plurality of artifact reduced epochs, and combining the plurality of artifact reduced epochs to generate a processed EEG recording, wherein each of the plurality of artifact reduced epochs overlaps an adjacent artifact reduced epoch to produce a continuous processed EEG recording without discontinuities in the processed EEG recording, wherein the plurality of artifact reduced epochs are combined using a weighted average wherein each of the plurality of artifact reduced epochs and the adjacent artifact reduced epoch are combined using a weighted average wherein the weight is proportional to a ratio of distances to epoch centers wherein a value from each of the plurality of artifact reduced epochs is weighted higher for an overlapping portion nearer a center of each of the plurality of artifact reduced epochs and a value from the adjacent artifact reduced epoch is weighted higher for an overlapping portion nearer a center of the artifact reduced epoch;overlaying the processed continuous EEG report on the original EEG report to generate a combined EEG report, wherein an x-axis of the processed continuous EEG report is aligned with an x-axis of the original EEG report, wherein a y-axis of the processed continuous EEG report is aligned with a y-axis of the original EEG report; anddisplaying the combined EEG report on a display screen of a monitor, the monitor in communication with the processor, wherein the processed EEG report is visually distinctive from the original EEG report, wherein an activity at a specific time on one channel of the first plurality of channels of the original EEG report is identifiable on a corresponding channel of the second plurality of channels of the processed EEG report at the specific time, wherein the amplitudes of the signals for the original EEG report are contained within each of the plurality of first channels in order to prevent overlapping of the signals and the amplitudes of the signals for the processed EEG report are contained within each of the plurality of second channels in order to prevent overlapping of the signals. 2. The method according to claim 1 wherein the artifact reduction is for at least one of muscle artifact, eye movement artifact, electrical artifact, heartbeat artifact, tongue movement artifact, and chewing artifact. 3. The method according to claim 1 wherein the combined EEG report comprises the processed EEG report having a first color and the original EEG report having a second color differe... | Persyst Development Corporation | A61B005/04 A61B005/00 A61B005/0476 A61B005/0478 G16H040/63
US09345426 | B2 | 2016-05-24 | Purification of glucose concentration signal in an implantable fluorescence based glucose sensor | 2013-03-29 | Methods, sensors, and systems for determining a concentration of glucose in a medium of a living animal are disclosed. Determining the glucose concentration may involve emitting excitation light from a light source to indicator molecules, generating a raw signal indicative of the amount of light received by a photodetector, purifying and normalizing the raw signal, and converting the normalized signal to a glucose concentration. The purification may involve removing noise (e.g., offset and/or distortion) from the raw signal. The purification and normalization may involve tracking the cumulative emission time that the light source has emitted the excitation light and tracking the implant time that has elapsed since the optical sensor was implanted. The purification and normalization may involve measuring the temperature of the sensor. The purification, normalization, and conversion may involve using parameters determined during manufacturing, in vitro testing, and/or in vivo testing. | 1. A method of determining a concentration of glucose in a medium of a living animal using an optical sensor implanted in the living animal, the method comprising: emitting, using a light source of the optical sensor, excitation light to indicator molecules of the optical sensor, the indicator molecules having an optical characteristic responsive to the concentration of glucose;generating, using a photodetector of the optical sensor, a raw signal indicative of the amount of light received by the photodetector, wherein the light received by the photodetector includes glucose-modulated light emitted by the indicator molecules and at least one of excitation light emitted by the light source and non-glucose modulated light emitted by the indicator molecules;tracking, using circuitry of the optical sensor, the cumulative emission time that the light source has emitted the excitation light;tracking, using circuitry of the optical sensor, the implant time that has elapsed since the optical sensor was implanted in the living animal;adjusting, using circuitry of the optical sensor, the raw signal to compensate for offset and distortion based on the tracked cumulative emission time and the tracked implant time;converting, using circuitry of the optical sensor, the adjusted signal into a measurement of glucose concentration in the medium of the living animal; andconveying, using an inductive element of the optical sensor, the measurement of glucose concentration. 2. The method of claim 1, further comprising: measuring, using a temperature sensor of the optical sensor, a temperature of the optical sensor;correcting, using circuitry of the optical sensor, the raw signal to compensate for temperature sensitivity of the light source based on the measured temperature. 3. The method of claim 1, wherein the non-glucose modulated light emitted by the indicator molecules comprises light emitted by distortion producing indicator molecule subspecies. 4. The method of claim 3, wherein the distortion producing indicator molecule subspecies include oxidated species, and adjusting the raw signal comprises: calculating the light emitted by the oxidated species based on the tracked cumulative emission time and the tracked implant time; andsubtracting the calculated light emitted by the oxidated species from the raw signal. 5. The method of claim 4, wherein the distortion producing indicator molecule subspecies include photo-activated oxidated species, and adjusting the raw signal comprises: calculating the light emitted by the photo-activated oxidated species based on the tracked cumulative emission time and the tracked implant time; andsubtracting the calculated light emitted by the photo-activated oxidated species from the raw signal. 6. The method of claim 3, wherein the distortion producing indicator molecule subspecies include thermal degradation product species, and adjusting the raw signal comprises: calculating the light emitted by the thermal degradation produ... | Senseonics, Incorporated | A61B005/1455 A61B005/145 A61B005/00 A61B005/1459
US09332357 | B2 | 2016-05-03 | Uncomfortable sound pressure estimation system, uncomfortable sound pressure estimation apparatus, uncomfortable sound pressure estimation method, and computer program thereof | 2013-11-15 | An exemplary uncomfortable sound pressure estimation system includes: a sound stimulation scheduling section configured to determine points in time of outputting a group of sound stimulations so that times of measurement for event-related potentials in response to the second sound and third sound arrive after a predetermined noise-lingering time has elapsed since a next point of occurrence of periodic noise; an event-related potential characteristic amount extraction section configured to receive the points in time of outputting the sound stimulation group, and extract a characteristic amount of an event-related potential of the electroencephalogram signal based on each point of outputting the group of sound stimulations as a starting point; and an uncomfortable sound pressure determination section configured to determine an uncomfortable sound pressure for the frequency of the sound stimulation group. | 1. An uncomfortable sound pressure estimation system comprising: one or more memories; andcircuitry which in operation is configured to:measure an electroencephalogram signal of a user;present a sound stimulation group including a first sound, a second sound, and a third sound to the user, the first sound, the second sound, and the third sound having a same frequency and having sound pressures in a predetermined range;extract, from the electroencephalogram signal in a predetermined analysis zone defined based on a point of presenting at least one of the second sound and the third sound to the user as a starting point, a characteristic amount of event-related potential associated with at least one of the second sound and the third sound;determine an uncomfortable sound pressure for the frequency of the sound stimulation group, based on the characteristic amount extracted;extract, from the electroencephalogram signal of the user, a point of occurrence of at least one of a plurality of noises constituting periodic noises;predict a next point of noise occurrence from the point of occurrence of the at least one noise extracted; anddetermine an output timing for the sound stimulation group so that an analysis zone for at least one of the second sound and third sound begins after a predetermined time has elapsed since the predicted next point of noise occurrence. 2. The uncomfortable sound pressure estimation system of claim 1, wherein the circuitry is configured to output the first sound, the second sound, and the third sound so as to consecutively decrease in sound pressure. 3. The uncomfortable sound pressure estimation system of claim 1, wherein the circuitry is configured to extract an N1-P2 amplitude or a wavelet-coefficient related characteristic amount. 4. The uncomfortable sound pressure estimation system of claim 1, wherein the circuitry is configured to subject the electroencephalogram signal of the user to differentiation, and if a result of the differentiation is equal to or greater than a predetermined threshold value, determines that noise has occurred. 5. The uncomfortable sound pressure estimation system of claim 1, wherein the circuitry is configured to calculate a period of periodic noise occurrences, and based on an average value or weighted average value of occurrence intervals between a past plurality of points of noise occurrence, predicts a next point of noise occurrence. 6. The uncomfortable sound pressure estimation system of claim 1, wherein the circuitry is configured to set an output timing for the sound stimulation group so that the analysis zone does not fall into a time slot which begins from 100 ms before the predicted next point of noise occurrence and spans 250 ms, the analysis zone being a zone from 0 ms to 300 ms after the third sound is output. 7. The uncomfortable sound pressure estimation system of claim 1, wherein the circuitry further is: configured to determine the frequency of the sound stimulation group... | Panasonic Intellectual Property Management Co., Ltd. | H04R025/00 A61B005/0484 A61B005/12
US09351640 | B2 | 2016-05-31 | Personal emergency response (PER) system | 2013-11-04 | A method of method of automatically requesting assistance for a patient by verbally communicating with a user through a base unit; wearing a mobile device with a plurality of different sensors on the patient to detect patient motion, wherein the mobile device transmits data to the base unit; and determining if a fall has occurred and requesting help through the base unit without a user pressing a button to call for help. | 1. A method of automatically requesting assistance for a patient, comprising: verbally communicating with a user through a base unit;wearing a mobile device with a plurality of different sensors on the patient to detect patient motion, wherein the mobile device transmits data to the base unit, wherein the patient motions form a sequence of elemental motions;identifying each elemental motion of a sequence of elemental motions of the accelerometer; andidentifying the activity of the patient, comprising matching the sequence of identified elemental motions with a library of stored sequences of elemental motions, wherein each stored sequence of elemental motions corresponds with an activity; anddetermining if a fall has occurred and requesting help through the base unit without a user pressing a button to call for help. 2. The method of claim 1, comprising determining a location of the mobile device. 3. The method of claim 1, comprising providing an indoor location determination system based on WiFi or Bluetooth to supplement a global positioning system (GPS). 4. The method of claim 1, comprising estimating an indoor location of the mobile device based on a radio frequency transmission. 5. The method of claim 1, comprising distinguishing between falls and other types of normal body movement. 6. The method of claim 1, comprising notifying an emergency service for assistance automatically without a user pressing a button to call for help. 7. The method of claim 1, comprising detecting a fall, waiting for a predetermined period before sending an alert to a call center which contacts the patient and sending local emergency services if needed. 8. The method of claim 1, comprising capturing movement activities and comparing the movement activities to fall patterns and determining if a fall has occurred. 9. The method of claim 1, comprising sending fall alert data and location to a call center. 10. The method of claim 9, wherein the call center calls a user on the mobile device and sends help as needed to a current location for the mobile device. 11. The method of claim 10, wherein the mobile device sends accelerometer pattern matching workload to a remote server to conserve battery power. 12. A method of automatically requesting assistance for a patient, comprising: verbally communicating with a user through a base unit;wearing a mobile device with a plurality of different sensors on the patient to detect patient motion, wherein the mobile device transmits data to the base unit;identifying each elemental motion by:generating an acceleration signature based on sensed acceleration of the device;matching the acceleration signature with at least one of a plurality of stored acceleration signatures, wherein each stored acceleration signatures corresponds with a type of motion;identifying the type of motion of the device based on the matching of the acceleration signature with the stored acceleration signature; anddetermining if a fall has occurred an... | KONINKLIJKE PHILIPS N.V. | G08B001/08 A61B005/00 G06F019/00 G08B021/04 A61B005/01 A61B005/0205 A61B005/0402 A61B005/0476 A61B005/0488 A61B005/11 A61B007/04 A61B008/00 G08B021/02 H04M003/51 A61B008/08
US09351648 | B2 | 2016-05-31 | Implantable medical device electrode assembly | 2012-08-24 | An implantable medical device (IMD) includes an electrode that forms a first snap-fit attachment area and an insulator that forms a through-hole, a second snap-fit attachment area and a third snap-fit attachment area. The second snap-fit attachment area mates with the first snap-fit attachment area of the electrode. The IMD further includes a body including an elongated conductive housing and a feedthrough wire extending therefrom. The body forms a fourth snap-fit attachment area on one end that mates with the third snap-fit attachment area of the insulator such that the feedthrough wire extends through the through-hole of the insulator. The housing encloses at least one of a battery, a sensor, and an electronic circuit. The insulator functions to electrically isolate the electrode from the housing of the body. | 1. An implantable medical device comprising: an electrode that forms a first snap-fit attachment area;an insulator that forms a through-hole, a second snap-fit attachment area and a third snap-fit attachment area, wherein the second snap-fit attachment area mates with the first snap-fit attachment area of the electrode; anda body including an elongated conductive housing and a feedthrough wire extending therefrom, wherein the body forms a fourth snap-fit attachment area on one end that mates with the third snap-fit attachment area of the insulator such that the feedthrough wire extends through the through-hole of the insulator,wherein the housing encloses at least one of a battery, a sensor, or an electronic circuit, andwherein the insulator functions to electrically isolate the electrode from the housing of the body. 2. The implantable medical device of claim 1, wherein the electrode forms a first hole that receives the feedthrough wire of the body, andwherein the electrode forms a second hole that receives an injection of an adhesive, the implantable medical device further comprising the adhesive. 3. The implantable medical device of claim 2, wherein the adhesive fills cavities between the electrode, the insulator and the body to secure the electrode to the body. 4. The implantable medical device of claim 2, wherein the insulator forms cavities to receive the injection of the adhesive via the second hole of the electrode. 5. The implantable medical device of claim 1, further comprising a weld joint between the feedthrough wire and the electrode. 6. The implantable medical device of claim 1, wherein the housing of the body has a substantially cylindrical shape. 7. The implantable medical device of claim 1, wherein the sensor is a pressure sensor, wherein the housing encloses the pressure sensor and the electronic circuit. 8. The implantable medical device of claim 1, wherein the housing encloses at least one of the sensor and the electronic circuit, andwherein the implantable medical device further includes the battery connected to the body on an opposite end of the housing relative to the electrode. 9. The implantable medical device of claim 8, wherein the battery includes a battery feedthrough serving as a battery terminal for the battery, wherein the battery feedthrough is electrically connected to the components enclosed within the housing. 10. The implantable medical device of claim 9, wherein the battery terminal is a first battery terminal, wherein the battery includes a conductive battery case serving as a second battery terminal of the battery, wherein battery case is electrically and mechanically coupled to the housing of the body. 11. The implantable medical device of claim 1, wherein the implantable medical device includes the sensor, wherein the sensor selected from a group consisting of: a pressure sensor;a cardiac electrogram sensor;an ECG sensora fluid flow sensor;perfusion sensoran oxygen sensor;a carbon dioxide sensor... | Medtronic, Inc. | A61B005/00 A61B005/0215 A61N001/372 A61B005/0205 A61B005/0452
US09335269 | B2 | 2016-05-10 | Optical control device, control device, optical scope, and scanning optical device | 2012-06-05 | An optical control device is provided in a scanning optical device that applies light emitted from a light source to an observation target as spot light that is applied in a spot-like shape, and detects reflected light from the observation target while scanning with the spot light, the optical control device including an intensity enhancement section that enhances intensity of light within a specific wavelength band included in a wavelength band of white light emitted from a white light source, an irradiation section that applies specific wavelength band-enhanced white light to the observation target, the specific wavelength band-enhanced white light being white light for which intensity of light within the specific wavelength band is enhanced, and a light detection section that detects reflected light from the observation target when the irradiation section applies the specific wavelength band-enhanced white light to the observation target. | 1. A scanning endoscope comprising: an optical control section that applies light emitted from a light source to an observation target as spot light that is applied in a spot-like shape, and detects reflected light from the observation target while scanning with the spot light; andan image processing section that generates a first image that corresponds to white light and a second image that corresponds light within a specific wavelength band based on the reflected light detected by the optical control section;the optical control section comprising: an intensity enhancement section that enhances intensity of the light within the specific wavelength band that corresponds to the second image and is included in a wavelength band of white light emitted from a white light source;an irradiation section that applies specific wavelength band-enhanced white light to the observation target, the specific wavelength band-enhanced white light being white light for which intensity of light within the specific wavelength band is enhanced by the intensity enhancement section; anda light detection section that detects the reflected light from the observation target when the irradiation section applies the specific wavelength band-enhanced white light to the observation target, and generates spectral data, andthe image processing section generates the second image based on the spectral data that corresponds to the specific wavelength band for which the intensity of light is enhanced, and generates the first image based on weighted spectral data calculated from the spectral data that corresponds to the specific wavelength band for which the intensity of light is enhanced and the spectral data that corresponds to a wavelength band other than the specific wavelength band while setting a weight applied to the spectral data that corresponds to the specific wavelength band to be smaller than a weight applied to the spectral data that corresponds to a wavelength band other than the specific wavelength band. 2. The scanning endoscope as defined in claim 1, the intensity enhancement section acquiring the specific wavelength band-enhanced white light by increasing intensity of light within the specific wavelength band included in the wavelength band of the white light emitted from the white light source. 3. The scanning endoscope as defined in claim 2, the intensity enhancement section acquiring first white light and second white light from the white light source, acquiring light within the specific wavelength band from the first white light, and synthesizing the light within the specific wavelength band and the second white light to acquire the specific wavelength band-enhanced white light. 4. The scanning endoscope as defined in claim 3, a first white light source and a second white light source being provided as the white light source, andthe intensity enhancement section acquiring the first white light from the first white light source, and acquiring the second white... | OLYMPUS CORPORATION | G01N021/64 A61B001/06 A61B001/07 A61B005/00 A61B001/04 A61B001/045 A61B001/00
US09585616 | B2 | 2017-03-07 | Determining treatment compliance using speech patterns passively captured from a patient environment | 2014-11-17 | Methods and systems are described for monitoring patient speech to determine compliance of the patient with a prescribed regimen for treating for a brain-related disorder. Patient speech is detected with an audio sensor at the patient location, and speech data is transmitted to a monitoring location. The audio sensor and other components at the patient location may be incorporated into, or associated with, a cell phone, computing system, or stand-alone microprocessor-based device, for example. Patient speech is processed at the patient location and/or monitoring location to identify speech parameters and/or patterns that indicate whether the patient has complied with the prescribed treatment regimen. Patient identity may be determined through biometric identification or other authentication techniques. The system may provide a report to an interested party, for example a medical care provider, based on whether (and/or the extent to which) the patient has complied with the prescribed treatment regimen. | 1. A system comprising: at least one audio sensor for sensing at least one audio signal including spontaneous speech from a patient at a patient location, the patient having a brain-related disorder and a prescribed treatment regimen for treating at least one aspect of the brain-related disorder;signal processing circuitry for detecting the spontaneous speech in the at least one audio signal and generating speech data including data indicative of whether the patient has complied with the prescribed treatment regimen based upon the detected spontaneous speech; andat least one transmitting device for transmitting a speech data signal containing the speech data including data indicative of whether the patient has complied with the prescribed treatment regimen from the patient location to a receiving device at a monitoring location. 2. The system of claim 1, wherein the signal processing circuitry includes: patient identification circuitry configured to determine a presence of the patient from at least one identity signal sensed at the patient location;wherein the signal processing circuitry is configured to detect the spontaneous speech from the patient based at least in part on the determination of the presence of the patient by the patient identification circuitry. 3. The system of claim 2, wherein the at least one identity signal includes at least a portion of the at least one audio signal, wherein the patient identification circuitry is configured to analyze the at least one audio signal to determine the presence of the patient by identifying at least a portion of the at least one audio signal that resembles known speech of the patient, and wherein the signal processing circuitry is configured to detect the spontaneous speech from the patient by identifying speech data corresponding to presence of the patient detected from the at least one audio signal. 4. The system of claim 2, wherein the at least one identity signal includes an image signal received from an imaging device at the patient location, wherein the patient identification circuitry is configured to analyze the image signal to determine the presence of the patient, and wherein the signal processing circuitry is configured to detect the spontaneous speech from the patient by identifying speech data corresponding to presence of the patient detected from the image signal. 5. The system of claim 4, wherein the patient identification circuitry is configured to analyze the image signal to determine the presence of the patient through at least one of facial recognition and gait analysis. 6. The system of claim 2, wherein the at least one identity signal includes a biometric signal from at least one biometric sensor at the patient location, wherein the patient identification circuitry is configured to analyze the biometric signal to determine the presence of the patient, and wherein the signal processing circuitry is configured to detect the spontaneous speech from the patient by identif... | Elwha LLC | G10L015/00 A61B005/00 G06F019/00 G06Q050/24 G10L025/66 A61B005/16
US09326705 | B2 | 2016-05-03 | Method and system for monitoring physiological and athletic performance characteristics of a subject | 2010-08-26 | The present invention is directed to systems and methods for monitoring characteristics of a subject. A system according to an exemplary embodiment of the invention includes a sensor subsystem including at least one respiratory sensor disposed proximate to the subject and configured to detect a respiratory characteristic of the subject, wherein the sensor subsystem is configured to generate and transmit at least one respiratory signal representing the respiratory characteristic, and at least one physiological sensor disposed proximate to the subject and configured to detect a physiological characteristic of the subject, wherein the sensor subsystem is configured to generate and transmit at least one physiological signal representing the physiological characteristic, and a processor subsystem in communication with the sensor subsystem, the processor subsystem being configured to receive at least one of the at least one respiratory signal and the at least one physiological signal. | 1. A fitness monitoring system for monitoring a subject engaged in a physical activity, the system comprising: a sensor subsystem including a first sensor and a second sensor each positioned at the subject's torso, wherein the first and second sensors are responsive to changes in distance therebetween, wherein the sensor subsystem is configured to generate and transmit a distance signal representative of the distance between the first and second sensors;a spatial sensor configured to generate and transmit a spatial signal, wherein the spatial sensor is configured to detect movement of the subject; anda processor subsystem in communication with the sensor subsystem and the spatial sensor and including a benchmark spatial signal and a benchmark distance signal, the processor subsystem being configured to receive the distance signal and the spatial signal, wherein the processor subsystem is configured to compare the spatial signal and the distance signal to the respective benchmark spatial signal and benchmark distance signal to assess respiratory and spatial characteristics of the subject and to classify a type of fitness activity of the subject; anda physiological sensor, wherein the physiological sensor is configured to monitor at least one of electrical activity of the brain, electrical activity of the heart, pulse rate, blood oxygen saturation level, skin temperature, EMG, ECG, EEG, and core temperature. 2. A fitness monitoring system for monitoring a subject engaged in a physical activity, the system comprising: a sensor subsystem including a first sensor and a second sensor each positioned at the subject's torso, wherein the first and second sensors are responsive to changes in distance therebetween, wherein the sensor subsystem is configured to generate and transmit a distance signal representative of the distance between the first and second sensors;a spatial sensor configured to generate and transmit a spatial signal, wherein the spatial sensor is configured to detect movement of the subject;a monitoring subsystem configured to receive the distance signal, wherein the processor subsystem is configured to process the distance signal to obtain a signal that is representative of a respiratory parameter, wherein the monitoring subsystem is configured to store the received signal, and wherein the monitoring subsystem is configured to display a representation of the respiratory parameter; anda processor subsystem in communication with the sensor subsystem and the spatial sensor and including a benchmark spatial signal and a benchmark distance signal, the processor subsystem being configured to receive the distance signal and the spatial signal, wherein the processor subsystem is configured to compare the spatial signal and the distance signal to the respective benchmark spatial signal, benchmark distance signal, or a stored received signal to assess respiratory and spatial characteristics of the subject and to classify a type of fitness activit... | ADIDAS AG | G06F019/00 A61B005/00 A61B005/113 A63B024/00 A61B005/0205 A61B005/024 A61B005/0402 A61B005/0488 A61B005/08 A61B005/11 A61B005/1455
US09339214 | B2 | 2016-05-17 | Body movement detection device | 2012-06-19 | A body movement detection device is provided with a main body, a display unit, and a control unit, the control unit including an ascent/descent method detection unit that detects whether a user who is wearing or carrying the main body is ascending/descending using a lift device, a counting unit that counts the number of times that the user ascends/descends using the lift device and the number of times that the user ascends/descends without using the lift device, based on the detection result of the ascent/descent method detection unit, and a display control unit that causes the display unit to display each of the values counted by the counting unit. The method by which the user ascends/descends can thereby be clearly displayed. | 1. A body movement detection device comprising a main body, a display unit and a control unit, the control unit including: a movement detector that detects that a user who is wearing or carrying the main body is ascending or descending floors using a lift device, or using stairs;an electronic counter for counting: separately for different number of floors changed in one movement: (1) a number of times the user used the lift device to ascend or descend floors, and (2) a number of times the user used the stairs to ascend or descend floors, based on a detection result of the movement detector; anda display module, operable on a processor, that instructs the display unit to display, separately for different number of floors: (1) a number of times the user used the lift device to ascend or descend floors, and (2) a number of times the user used the stairs to ascend or descend floors. 2. The body movement detection device according to claim 1, wherein the counter counts the number of times that the user ascends or descends using the lift device and the number of times that the user ascends or descends without using the lift device, per ascent and descent by (1) setting an ascending flag when the user is ascending and counting the number of ascending steps only while the ascending flag is set and (2) setting a descending flag when the user is descending and counting the number of descending steps only while the descending flag is set. 3. The body movement detection device according to claim 1, wherein the counter counts the number of times that the user ascends or descends using the lift device and the number of times that the user ascends or descends without using the lift device, per number of floors that the user has ascended or descended. 4. The body movement detection device according to claim 1 further comprising: an acceleration sensor for detecting acceleration of the main body; andan atmospheric pressure sensor for detecting an ambient atmospheric pressure of the main body,wherein the movement detector detects whether the user is ascending or descending using the lift device, based on a detection result from the acceleration sensor and a detection result from the atmospheric pressure sensor. | OMRON HEALTHCARE CO., LTD. | A61B005/11 A61B005/00
US09339191 | B2 | 2016-05-17 | Optical measurement apparatus | 2009-06-15 | Provided is a near-infrared spectroscopy apparatus using a phosphor. [Object]: Information in a specimen is observed by using light without contacting an optical fiber and an electronic circuit with the specimen. [Means for Solution]: A small and lightweight phosphor is contacted with the specimen to measure the fluorescence intensity at a separate position. | 1. An optical measurement apparatus comprising: one or more light irradiators, each of which irradiates light with a wavelength of 200 nm to 1500 nm onto a living object at an incident light position on a surface of the living object, respectively;one or more luminescent cells, each of which is filled with a luminescent material and has a first surface which is disposed directly on the surface of the living object at a light detection position spaced apart from the incident light position on the surface of the living object, respectively;an optical detector which detects light emitted from the luminescent material; anda shield member which covers at least another surface of each of the one or more luminescent cells other than the first surface that is in contact with the surface of the living object, respectively, and which shields the luminescent material from stray light having the wavelength of the light from the one or more light irradiators, respectively, and a part of the shield member of each of the one or more luminescent cells transmits light emitted from the luminescent material,wherein the optical detector is disposed on the same side of the living object as the one or more light irradiators with respect to the surface of the living object on which each of the one or more luminescent cells is disposed, andwherein the optical detector measures a luminescence intensity of the light emitted from the luminescent material of each of the one or more luminescent cells, respectively, while the light is irradiated by the one or more light irradiators, respectively, and absorption information of the incident light in the living object is measured from the luminescence intensity measured by the optical detector. 2. The optical measurement apparatus according to claim 1, wherein the one or more luminescent cells have a triangular shape and the optical detector is an image-pickup device. 3. The optical measurement apparatus according to claim 1, wherein a marker is disposed at the incident light position, and the marker is irradiated with light by the one or more light irradiators from a position apart from the living object, and a mechanism of tracking the marker even when the living object moves is provided. 4. The optical measurement apparatus according to claim 1, further comprising: a mechanism of correcting the detected luminescence intensity based on a positional relation between the optical detector or the light irradiators and the living object. 5. The optical measurement apparatus according to claim 1, wherein light emitted from the luminescent material has a wavelength different than the wavelength of light from the one or more light irradiators. 6. The optical measurement apparatus according to claim 1, wherein the shield member covers plural surfaces of each of the one or more luminescent cells other than the first surface that is in contact with the surface of the living object. 7. The optical measurement apparatus according to ... | Hitachi, Ltd. | F21V009/16 A61B005/00 A61B005/02 A61B005/026 G01N021/47 G01N021/64 G01N021/359 G01N021/31
US09345890 | B2 | 2016-05-24 | Systems and methods for managing noise in implantable medical devices | 2010-06-15 | Embodiments of the invention are related to managing noise in sensed signals in implantable medical devices, amongst other things. In an embodiment the invention includes a method for processing electrical signals obtained from a patient including gathering a first set of electrical signals using an implantable medical device, filtering to provide a second set of electrical signals, the second set including frequencies above a threshold frequency, and estimating the amount of noise present in the first set of electrical signals based on the magnitude of the second set. In an embodiment, the invention includes a medical device configured to gather a first set of electrical signals, filter the first set to provide a second set of electrical signals including frequencies above a threshold frequency, and estimate the amount of noise present in the first set based on the magnitude of the second set. Other embodiments are also included herein. | 1. A method for processing atrial electrical signals obtained from a patient comprising: gathering a first set of atrial electrical signals within the patient using an implantable medical device during an atrial pacing threshold test, the first set of electrical signals comprising a set of samples of electrical activity from a discrete sample time window spanning a pacing event, wherein the sample time window starts in response to an electrical stimulation p acing pulse;filtering the first set of atrial signals to provide a second set of electrical signals, the second set of electrical signals comprising frequencies above a cutoff frequency;estimating the amount of noise present in the first set of atrial electrical signals based on the magnitude of the second set of electrical signals;calculating a noise threshold value according to a capture detection threshold;discarding the first set of atrial electrical signals if the estimated amount of noise exceeds the noise threshold value;calculating an updated noise adjusted pace artifact threshold based on the estimated amount of noise if the noise does not exceed the noise threshold value;comparing the updated noise adjusted pace artifact threshold to the noise threshold value; andwhen the updated noise adjusted pace artifact threshold is less than the noise threshold value, classifying the electrical signal as loss of capture or fusion. 2. The method of claim 1, wherein one or more of the initial filtered results making up the second set of electrical signals are not considered in estimating the amount of noise. 3. The method of claim 1, further comprising identifying the first set of electrical signals as indicative of noise if the estimated amount of noise exceeds a threshold value. 4. The method of claim 3, further comprising incrementing a noise counter if the first set of electrical signals is identified to be indicative of noise. 5. The method of claim 4, further comprising terminating an automatic threshold testing procedure if the noise counter exceeds a predetermined value. 6. The method of claim 1, wherein filtering the first set of signals is performed with a Chebyshev high pass filter at 68 Hz. 7. The method of claim 1, the first set of electrical signals comprising a set of samples of atrial myocardial electrical activity taken over a 100 ms time window. 8. The method of claim 1, wherein estimating the amount of noise present in the first set of electrical signals based on the magnitude of the second set of electrical signals comprises calculating a rectified average value for the magnitude of the second set of electrical signals. 9. A medical device comprising: a processor circuit; and a memory circuit operatively connected to the processor circuit; the medical device configured togather a first set of electrical signals from a patient, the first set of electrical signals comprising a set of samples of electrical activity from a discrete sample time window spanning a pacing eve... | Cardiac Pacemakers, Inc. | A61N001/37 A61B005/0428 A61B005/0432 A61B005/00 G06K009/00
US09326683 | B2 | 2016-05-03 | Systems and methods for detection of wound fluid blood and application of phototherapy in conjunction with reduced pressure wound treatment system | 2011-04-18 | Wound fluid blood detection systems and methods are described that are operable in conjunction with reduced pressure wound treatment (RPWT) systems, as well as ancillary therapy and monitoring systems applied concurrently with RPWT systems. The blood detection monitor operates by optically characterizing the content of wound fluids to the extent of identifying percentage blood content. This identification relies upon the transmission of select wavelengths of light across a volume of wound fluid to a photo detector (connected to signal processing instrumentation) capable of quantifying the absorption characteristics of the fluid. The detection components may be implemented in conjunction with either a fluid flow conduit (i.e. the reduced pressure tubing directing fluid away from the wound dressing) or more directly in association with the materials that comprise the wound dressing positioned within the wound bed itself. In addition, the present invention is configured to operate in conjunction with blood gas monitoring, systems operating with the RPWT. Finally, the components of the systems that serve to illuminate the wound bed, lend themselves to use in conjunction with illumination in discreet electromagnetic wave wavelengths suitable for the application of phototherapy. | 1. A blood detection system operable with a reduced pressure wound treatment (RPWT) system including a conduit for conducting reduced pressure to a wound dressing and withdrawing wound fluid from a wound, comprising: an electromagnetic (EM) wave source oriented to at least partially illuminate an interior portion of the conduit, wherein wound fluid is present within the conduit during operation of the RPWT system;wherein said EM wave source comprises at least one light wave guide integrated into the conduit walls, the light wave guide having a first end oriented to at least partially illuminate the interior portion of the conduit and a second end for receiving EM waves originating adjacent at least one LED remote from the conduit; andan electromagnetic (EM) wave detector positioned to be receptive of attenuated EM waves originating from said source and attenuated after passing through the interior portion of the conduit, said attenuation resulting from a partial and wavelength-specific absorption measurable and indicative of a level of blood within said wound fluid;wherein said EM wave detector comprises a light wave guide integrated into the conduit walls, the light wave guide having a first end positioned generally opposite of said EM wave source for receiving attenuated EM waves, and a second end for transmitting the attenuated EM waves and terminating adjacent a solid state photo detector remote from the conduit. 2. The blood detection system of claim 1 further comprising a microprocessor electrically coupled to the EM wave source and the EM wave detector and operable to sequentially pulse the EM wave source to effect absorption by the wound fluid and attenuation measurements by the EM wave detector configured to provide attenuation measurements to the microprocessor, wherein the microprocessor provides an output signal indicating the level of blood within the wound fluid flowing through the conduit. 3. The blood detection system of claim 2, further comprising an alert indicator device, said microprocessor further in signal communication with said alert indicator device and programmed to trigger said alert indicator device upon said signal data indicating a new level of EM wave absorption outside a predetermined range. 4. The blood detection system of claim 2, further comprising an RPWT control device, said microprocessor further in signal communication with said RPWT control device and programmed to operate said RPWT control device upon said signal data indicating a level of EM wave absorption outside a predetermined range, said RPWT control device connected to said RPWT system and serving to modify the operation of said RPWT system. 5. The blood detection system of claim 1, wherein said EM wave source comprises a plurality of light sources, each light source configured to illuminate the wound fluid in a distinct range of spectral wavelengths. 6. The blood detection system of claim 5, wherein the distinct range of spectral wavelengths i... | KCI Licensing, Inc. | A61M001/00 A61B005/00 A61B005/145 A61B005/1468 A61M031/00 A61M005/20 A61M035/00 A61M027/00 A61F002/00 B65D081/00 A61N005/06
US09339332 | B2 | 2016-05-17 | Neuromodulation catheters with nerve monitoring features for transmitting digital neural signals and associated systems and methods | 2013-08-30 | Neuromodulation catheters with nerve monitoring features for transmitting digital neural signals and associated systems and methods are disclosed herein. A neuromodulation catheter configured in accordance with some embodiments of the present technology can include, for example, a handle and an elongated shaft attached to the handle. The shaft can have a proximal portion and a distal portion configured to be moved within a lumen of a blood vessel of a human patient. The neuromodulation catheter can further include an array of contacts at the distal portion of the shaft and a digitizer at the handle or the shaft. The contacts can be configured to detect analog neural signals from within the blood vessel. The digitizer can be configured to receive the analog neural signals from the contacts, digitize the analog neural signals into digital neural signals, and transmit the digital neural signals to a read/write module external to the patient. | 1. A neuromodulation catheter, comprising: a handle; an elongated shaft attached to the handle, the shaft having a distal portion and a proximal portion, wherein at least the distal portion is configured to be moved within a lumen of a blood vessel of a human patient; an array of contacts at the distal portion of the shaft, wherein the contacts are configured to detect analog neural signals from nerves along the blood vessel while the contacts are within the blood vessel; and a digitizer at the handle of the shaft, wherein the digitizer is configured to receive the analog neural signals from the contacts, digitize the analog neural signals into digital neural signals comprising electroneurogram (ENG) signals, and transmit the digital neural signals to a read/write module external to the patient, wherein the digitizer is configured to filter signals received from the contacts to differentiate the ENG signals from electromyogram (EMG) signals before transmitting the digital neural signals to the read/write module. 2. The neuromodulation catheter of claim 1 wherein the digitizer is configured to be detachable from the handle for reuse with other neuromodulation catheters. 3. The neuromodulation catheter of claim 1 wherein the digitizer is configured to filter signals received from the contacts using an amplifier assembly having a quasi-tripole (QT), true-tripole (TT), or adaptive-tripole (AT) arrangement to detect electroneurogram (ENG) signals. 4. The neuromodulation catheter of claim 1 wherein the digitizer comprises an analog to digital circuit. 5. The neuromodulation catheter of claim 1 wherein the digitizer has a cross-sectional dimension of about 9-25 mm2. 6. The neuromodulation catheter of claim 1 wherein the digitizer further comprises a telemetry module configured to wirelessly transmit the digitized signals from within the patient to the read/write module. 7. The neuromodulation catheter of claim 1 wherein the digitizer is inductively coupled to the read/write module. 8. The neuromodulation catheter of claim 1, further comprising at least one energy delivery element at the distal portion of the elongated shaft, wherein the energy delivery element is configured to deliver neuromodulation energy to a target site within the blood vessel of the patient. 9. The neuromodulation catheter of claim 8 wherein the energy delivery element comprises an RF electrode. 10. The neuromodulation catheter of claim 1 wherein the contacts are further configured to deliver neuromodulation energy to the nerves along the blood vessel of the patient and detect the analog neural signals before or after the delivery of the neuromodulation energy. 11. The neuromodulation catheter of claim 1, further comprising a support member at the distal portion of the shaft, wherein the support member has a spiral shape and is configured to contact an interior wall of the blood vessel when the support member is in a deployed state, wherein the contacts are spaced apart f... | Medtronic Ardian Luxembourg S.a.r.l. | A61B018/14 A61B017/32 A61B018/02 A61B005/04 A61B005/00 A61B017/00 A61B018/00 A61N007/02
US10512429 | B2 | 2019-12-24 | Discrimination of cheyne-stokes breathing patterns by use of oximetry signals | 2015-12-23 | Methods and apparatus provide Cheyne-Stokes respiration (“CSR”) detection based on a blood gas measurements such as oximetry. In some embodiments, a duration, such as a mean duration of contiguous periods of changing saturation or re-saturation occurring in an epoch taken from a processed oximetry signal, is determined. An occurrence of CSR may be detected from a comparison of the duration and a threshold derived to differentiate saturation changes due to CSR respiration and saturation changes due to obstructive sleep apnea. The threshold may be a discriminant function derived as a classifier by an automated training method. The discriminant function may be further implemented to characterize the epoch for CSR based on a frequency analysis of the oximetry data. Distance from the discriminant function may be utilized to generate probability values for the CSR detection. | 1. A computer implemented method of detecting an occurrence of Cheyne-Stokes respiration in one or more programmed processors comprising: accessing blood gas data representing a measured blood gas signal from an oximeter;determining a duration of one or more contiguous periods of changing saturation of a blood gas from the blood gas data;determining a presence of a peak in a predetermined frequency range for de-saturation and re-saturation cycles of the blood gas data;detecting the occurrence of Cheyne-Stokes respiration from (a) a comparison of the determined duration and a threshold derived to differentiate saturation changes due to Cheyne-Stokes respiration and saturation changes due to obstructive sleep apnea, and (b) a comparison of the determined peak to a further threshold;generating an output indicating the detected occurrence of Cheyne-Stokes respiration; anddetermining a quality indicator for determining a confidence level associated with detecting the occurrence of Cheyne-Stokes respiration,wherein the processor determines the quality indicator by determining a number of samples where the measured blood gas signal drops below a predetermined percentage threshold. 2. The method of claim 1 wherein the one or more contiguous periods of changing saturation comprises re-saturation periods and the measured blood gas signal comprises an oximetry signal. 3. The method of claim 2 wherein the determined duration comprises a mean period length and wherein the detecting indicates an occurrence when the mean period length exceeds the threshold. 4. The method of claim 3 wherein the threshold comprises a discriminant function. 5. The method of claim 4 wherein the detecting the occurrence further comprises determining a distance from the threshold and comparing the distance to a further threshold. 6. The method of claim 1 further comprising processing the blood gas data to remove artifact data. 7. The method of claim 1 further comprising measuring the blood gas with the oximeter. 8. The method of claim 1, wherein the predetermined percentage threshold is 10%. 9. The method of claim 1, wherein the quality indicator is a function of a ratio of the number of samples where the measured blood gas signal drops below the predetermined percentage threshold and a total number of samples considered. 10. A computer implemented method of detecting an occurrence of Cheyne-Stokes respiration in one or more programmed processors comprising: accessing blood gas data representing a measured blood gas signal from an oximeter;determining a duration of one or more contiguous periods of changing saturation of a blood gas from the blood gas data;determining a presence of a peak in a predetermined frequency range for de-saturation and re-saturation cycles of the blood gas data;detecting the occurrence of Cheyne-Stokes respiration from (a) a comparison of the determined duration and a threshold derived to differentiate saturation changes due to Cheyne-Stokes respira... | ResMed Pty Ltd | A61B005/1455 A61B005/08 A61B005/00 A61B005/087 G06K009/00 G06K009/62 A61M016/00
US09339218 | B2 | 2016-05-17 | Concentration measuring device | 2014-07-11 | A concentration measuring device having a spot-application region, and a configuration capable of measuring the concentration of an object component in the sample liquid dropped on the spot-application region, the concentration measuring device including: a light irradiation device, which is provided on a side below the spot-application region, for emitting an irradiation light having a wavelength absorbable by a coloring matter in the dropped sample liquid. The irradiation light passes through the spot-application region, thereby the sample liquid dropped on the spot-application region is irradiated with the irradiation light from the side below the spot-application region. As a result, a non-covered region in the spot-application region can be easily found by visual observation. The spot-application region is composed of, for example, a through hole formed on the supporting film and a separation membrane that blocks the through hole. | 1. A concentration measuring device having a spot-application region for dropping a sample liquid into, and a configuration capable of measuring the concentration of an object component in the sample liquid dropped on the spot-application region, the concentration measuring device is configured: to move the spot-application region with the dropped sample liquid to a position different from the position at which the sample liquid was dropped, andto measure the concentration of an object component in the sample liquid dropped on the spot-application region based on an enzyme electrode method, by an electrode type sensor provided on a side below the moved spot-application region,the concentration measuring device comprising:a light irradiation device, which is provided on a side below the spot-application region at the position at which the sample liquid is dropped, for emitting visible light having a wavelength absorbable by a coloring matter in the dropped sample liquid,wherein the light irradiation device is placed such that the visible light passes through the spot-application region, thereby the spot-application region and the sample liquid dropped on the spot-application region are irradiated with the visible light from the side below the spot-application region. 2. The concentration measuring device according to claim 1, wherein the spot-application region comprises a through hole provided in a supporting film and a separation membrane for blocking the through hole, the through hole being a hole determining the peripheral shape of the spot-application region, and the separation membrane being a porous thin film allowing penetration of only a liquid component in the sample liquid. 3. The concentration measuring device according to claim 1, wherein a plurality of sample liquids having different colors are to be dropped, and the light irradiation device is configured to emit plural kinds of lights corresponding to the respective colors and having wavelengths absorbable by respective coloring matters in the sample liquids. 4. The concentration measuring device according to claim 1, wherein the sample liquid is blood or a red liquid of processed blood, and the visible light is a green light. 5. The concentration measuring device according to claim 1, wherein the sample liquid is a standard concentration test solution for calibration of the concentration measuring device for the concentration measurement, the standard concentration test solution has a blue color, and the visible light is a red light. | HORIBA, Ltd. | G01N033/48 A61B005/145 A61B005/1486 B01L003/00 C12Q001/00 G01N033/66 G01N027/327 G01N021/03
US09351643 | B2 | 2016-05-31 | Systems and methods for optical measurement for in-situ surgical applications | 2014-01-27 | A system and method is presented for performing optical measurements, including a light source configured to emit a light beam, a first pattern generator defining a first longitudinal axis and configured to project a first generated pattern, and a second pattern generator defining a second longitudinal axis and configured to project a second generated pattern. The first and second generated patterns have different angular divergency. The first pattern generator is a diffractive circle pattern generator, whereas the second pattern generator is a diffractive cross pattern generator. Adjustment of the first and second generated patterns with respect to each other cause the system to serve as an optical ruler for performing the optical measurements when the first and second generate patterns overlap or coincide with each other at certain points. | 1. A system for performing optical measurements, the system comprising: a light source configured to emit a light beam;a first pattern generator defining a longitudinal axis and configured to project a first generated pattern, the first pattern generator configured to receive a portion of the light beam; anda second pattern generator placed along the longitudinal axis defined by the first pattern generator and configured to project a second generated pattern, the second pattern generator longitudinally spaced from the first pattern generator to define a gap therebetween, and the second pattern generator configured to receive a remaining portion of the light beam;wherein the first and second generated patterns have different angular divergency. 2. The system according to claim 1, wherein the light source is a laser. 3. The system according to claim 1, wherein the first pattern generator is a diffractive circle pattern generator, whereas the second pattern generator is a diffractive cross pattern generator. 4. The system according to claim 1, wherein the first pattern generator is located between the light source and the second pattern generator. 5. The system according to claim 1, wherein the first pattern generator is positioned at a first distance with respect to the light source and the second pattern generator is positioned at a second distance with respect to the light source, the second distance being greater than the first distance. 6. The system according to claim 1, wherein the first and second generated patterns are projected onto a target site in an overlapping manner. 7. The system according to claim 6, wherein the first generated pattern is a circle and the second generated pattern is a cross-hair that is adjusted with respect to the circle based on the distance of the system relative to the target site. 8. The system according to claim 1, wherein adjustment of the first and second generated patterns with respect to each other cause the system to serve as an optical ruler for performing the optical measurements when the first and second generated patterns overlap. 9. The system according to claim 1, wherein the light source is coaxial with one of the first or second pattern generators. 10. The system according to claim 1, wherein the system is mounted on a surgical device. 11. The system according to claim 1, wherein the second generated pattern is a diffractive cross pattern having a plurality of tick marks thereon configured to provide a positional relationship with the first generated pattern. 12. The system according to claim 1, wherein at least one of the first and second pattern generators is a moving diffractive element. 13. The system according to claim 12, wherein the moving diffractive element is a lead screw translational motion mechanism. 14. A surgical instrument comprising: a handle portion;a body portion extending distally from the handle portion and defining a longitudinal axis;an end effector assembly dis... | Covidien LP | A61B005/00 A61B018/22 A61B019/00 A61B017/00 A61B001/00 A61B005/107 G01B011/02 G01B011/25
US09326706 | B2 | 2016-05-03 | Hearing profile test system and method | 2013-08-27 | Examples of systems and methods for profiling the hearing ability of a consumer are disclosed. One example includes a personal computer and a handheld device configured to produce calibrated acoustic output at suprathreshold levels above 20 db HL, and at step levels of 10-20 decibels, and presented test frequency bands across an audiometric frequency range from 400 to 8000Hz. The consumer's minimal audibility levels are registered, and a hearing profile score is presented to indicate hearing ability and hearing aid candidacy. In some embodiments, band-limited natural sounds are presented. Systems and methods disclosed herein, with considerations for noise present in the consumer's environment, allow for rapid calibrated hearing profiling, using a standard personal computer and minimal hardware, thus particularly suited for self-testing outside clinical environments such as at home or the office. | 1. A method of administering a hearing evaluation for a consumer, the method comprising: generating test audio signals at predetermined levels by an audio generator;delivering a sequence of calibrated acoustic test stimuli, representing the test audio signals, in each of three or more test frequency bands within an audiometric frequency range, to the ear of the consumer at suprathreshold sound levels of at least 20 dB HL, and wherein a step level for consecutive calibrated test stimuli at each test frequency band is at least 10 dB;registering, using a computerized hearing test system, minimum audibility levels of the consumer within the suprathreshold sound levels delivered at each test frequency band; andpresenting a computed hearing profile score to the consumer by the computerized hearing test system, wherein the computed hearing profile score is based on the consumer's minimum audibility level only within the suprathreshold sound levels delivered at or above 20 dB HL in step sizes of at least 10 dB at each test frequency band. 2. The method of claim 1, wherein the computerized hearing test system comprises a personal computer. 3. The method of claim 2, wherein the personal computer is selected from the group consisting of a smart phone and a tablet. 4. The method of claim 2; further comprising providing online customer support to the consumer by a remote customer support computer system with connection to the personal computer via the Internet. 5. The method of claim 4, wherein the online customer support comprises streaming customer support audio to the personal computer from the customer support computer system. 6. The method of claim 2, wherein the hearing profile score is presented through a display of the personal computer. 7. The method of claim 1, wherein the hearing test system comprises an earphone configured to deliver the calibrated acoustic test stimuli to an ear of the consumer. 8. The method of claim 7, wherein the earphone is configured to attenuate ambient background noise by at least 5 dB across the audiometric frequency range of 500 to 6,000 Hz. 9. The method of claim 7, wherein the earphone comprises an ear canal insert configured to occlude the ear canal and attenuate ambient background noise. 10. The method of claim 1, wherein a maximum of the computed heating profile score is between four and six. 11. The method of claim 1, wherein the computerized hearing test system is connected online to any of a remote server and a remote database. 12. The method of claim 1, wherein the remote server hosts a hearing profile software application configured to administer the hearing evaluation for the consumer. 13. The method of claim 12, further comprising executing the hearing profile software application from a personal computer. 14. The method of claim 1, further comprising saving the hearing profile score at a remote database. 15. The method of claim 1, wherein the calibrated acoustic test stimuli include a chirp sou... | iHear Medical, Inc. | A61B005/12 A61B005/00 H04R025/00
US09326734 | B1 | 2016-05-03 | Method of compressing electrocardiography signal and electrocardiography signal sensing apparatus thereof | 2015-02-04 | A method of compressing an electrocardiography (ECG) signal and an ECG signal sensing apparatus thereof are provided. The method of compressing the ECG signal includes the following steps. A current first characteristic wave is identified and captured from the ECG signal. A set of difference is obtained by subtracting a previous first characteristic wave from the current first characteristic wave. A first encoding process is executed on the set of difference, so as to obtain a first bit stream. A compressed ECG data is generated, wherein the compressed ECG data at least comprises the first bit stream. | 1. A method for compressing an electrocardiography signal, adapted to an electrocardiography signal sensing apparatus, and the method for compressing the electrocardiography signal comprising: identifying and capturing a current first characteristic wave from the electrocardiography signal;obtaining a set of difference by subtracting a previous first characteristic wave from the current first characteristic wave;executing a first encoding process on the set of difference to obtain a first bit stream;generating a compressed electrocardiography data comprising the first bit stream in response to the first bit stream;obtaining at least one residual wave by subtracting the current first characteristic wave from the electrocardiography signal;executing a second encoding process on the at least one residual wave to obtain a second bit stream; andgenerating the compressed electrocardiography data comprising the first bit stream and the second bit stream in response to the first bit stream and the second bit stream. 2. The method for compressing the electrocardiography signal as claimed in claim 1, wherein the first encoding process comprises a Huffman coding process. 3. The method for compressing the electrocardiography signal as claimed in claim 1, wherein the second encoding process comprises a differential coding process and a Huffman coding process. 4. The method for compressing the electrocardiography signal as claimed in claim 1, further comprising: obtaining a start information and a peak information related to the current first characteristic wave and an offset information related to a current second characteristic wave following the current first characteristic wave from the electrocardiography signal; andgenerating the compressed electrocardiography data comprising the first bit stream, the second bit stream, the start information, the peak information and the offset information in response to the first bit stream, the second bit stream, the start information, the peak information and the offset information. 5. An electrocardiography signal sensing apparatus, comprising: a sensor, sensing an electrocardiography signal;a capturing unit, coupled to the sensor, and identifying and capturing a current first characteristic wave from the electrocardiography signal;a first calculation unit, coupled to the capturing unit and the sensor, and obtaining a set of difference by subtracting a previous first characteristic wave from the current first characteristic wave;a first encoding unit, coupled to the first calculation unit, and executing a first encoding process on the set of difference to obtain a first bit stream;a package unit, coupled to the capturing unit and the first encoding unit, and generating a compressed electrocardiography data comprising the first bit stream in response to the first bit stream;a second calculation unit, coupled to the capturing unit, and obtaining at least one residual wave by subtracting the current first characteri... | Sunplus Technology Co., Ltd. | A61B005/04 A61B005/00 A61B005/0402
US10531820 | B2 | 2020-01-14 | Device, system and method for determining the concentration of a substance in the blood of a subject | 2015-11-16 | In a device, system and method for determining the concentration of a substance, such as the oxygen saturation in the blood of a subject, the influence of specular reflection is reduced or removed. An input unit is configured to receive detection signals reflected back from a skin area of the subject in response to irradiation of the skin area by a radiation signal. A signal extraction unit is configured to extract at least two photo-plethysmography, PPG, signals at two different wavelengths from the detection signals. A processing unit is configured to compute the concentration of a selected substance in the blood of the subject based on the PPG signals taking into account a skin tone of the subject. | 1. A device for determining a concentration of a substance in blood of a subject, comprising: an input unit configured to receive detection signals indicative of radiation reflected back from a skin area of the subject in response to irradiation of the skin area,a signal extraction unit configured to extract at least two photo-plethysmography, PPG, signals at two different wavelengths from the detection signals, anda processing unit configured to obtain skin tone information indicative of a skin tone of the subject and compute the concentration of a desired substance in the blood of the subject based on the PPG signals, wherein the processing unit is further configured to compensate the computation for the skin tone of the subject based on the skin tone information, wherein the processing unit is further configured to at least one of: use as a correction factor for correcting the denominator of the first normalized pulsatility at a wavelength in the red spectrum a fraction of the DC level of the PPG signal at a wavelength in the infrared spectrum,use as a correction factor a fraction in the range of 5% to 15% of the DC level of the PPG signal at a wavelength in the infrared spectrum; andcompute arterial blood oxygen concentration from the PPG signals byC1-C2·ACR/DCRACI⁢⁢R/DCI⁢⁢R·(1-S)1-S·DCI⁢⁢R/DCRwhereACR/DCR is the normalized pulsatility at a wavelength in the red spectrum,ACIR/DCIR is the normalized pulsatility at a wavelength in the infrared spectrum,DCR is the DC level of the PPG signals at a wavelength in the red spectrum,DCIR is the DC level of the PPG signals at a wavelength in the infrared spectrum,C1 and C2 are predetermined calibration constants, andS is an estimate of the relative specular reflection contained in the DC level of the PPG signals. 2. The device as claimed in claim 1, wherein the processing unit is further configured to form a ratio of a first normalized pulsatility at a first wavelength and a second normalized pulsatility at a second wavelength for the computation of the concentration, wherein the denominator of the first and/or the second normalized pulsatility is corrected by a correction factor. 3. The device as claimed in claim 1, wherein the processing unit is configured to use a value in the range of 5% to 15% for S. 4. The device as claimed in claim 1, further including: a non-contact remote photoplethysmography (rPPG) camera configured to generate images of the skin area of the subject. 5. The device as claimed in claim 4, wherein the processing unit is further configured to adapt the computation to an average intensity of the radiation reflected back from the skin area at the two different wavelengths. 6. A device for determining a concentration of a substance in blood of a subject, comprising: an input unit configured to receive detection signals reflected back from a skin area of the subject in response to irradiation of the skin area by a radiation signal,a signal extraction unit configured to extract a... | KONINKLIJKE PHILIPS N.V. | A61B005/1455 A61B005/1495 A61B005/00
US11457865 | B2 | 2022-10-04 | Wearable sports monitoring equipment for measuring heart rate or muscular activity and relating method | 2015-11-17 | The invention relates to wearable electronic equipment and method for measuring heart rate or muscular activity of a person. The equipment comprises one or more heartbeat or muscular activity sensors for providing a heartbeat or muscular activity signal, respectively, and a motion sensor for providing a motion signal corresponding to movement of the person. In addition, there are provided processing means for detecting from the heartbeat or muscular activity signal first events corresponding to individual heartbeats or muscle activations, respectively, and from the motion signal second events corresponding to movement changes stronger than a predefined threshold. The processing means are further programmed to form a first time series of said first events, and to determine the heartbeat or muscular activity parameter using the first time series. According to the invention, the processing means are programmed to omit from said first time series those first events that coincide with and/or take place at the same frequency as the second events. The invention helps to remove motion-induced artifacts from ECG and EMG signals for more accurate sports performance monitoring and characterization. | 1. Wearable electronic equipment for determining a heartbeat or muscular activity parameter of a person during a physical performance, comprising: a heartbeat or muscular activity sensor configured to provide a heartbeat or muscular activity signal, respectively;a motion sensor configured to provide a motion signal corresponding to movement of the person; anda processing unit configured to:detect during said physical performance from the heartbeat or muscular activity signal first events corresponding to characteristic features of individual heartbeats or muscle activations, respectively, and to record a first series of said first events during a time span in said physical performance,detect during said physical performance from the motion signal acceleration events which are stronger above a predefined threshold as second events, and to record a second series of said second events during said time span in said physical performance,generate an incomplete series of heartbeats by selectively removing, during said physical performance, such first events from said first series of events in said time span which coincides with a second event in said second series of events, anddetermine a heart rate of the person based on the incomplete series of heartbeats. 2. The wearable electronic equipment according to claim 1, wherein the processing unit is further configured to disregard periods of the heartbeat or muscular activity signal that coincide with second events when recording said first series of events. 3. The wearable electronic equipment according to claim 1, wherein the processing unit is further configured to: record said first series of said first events detected as a binary series, where one binary value represents detected events and the opposite binary value represents a non-existence of such events,record said second series of said second events detected as a binary series, where one binary value represents detected events and the other binary value represents a non-existence of such events, andfilter frequency components of said first series of events that are in common with at least one frequency component of the second series of events, in order to form a filtered first series of events. 4. The wearable electronic equipment according to claim 3, wherein the processing unit is further configured to carry out filtering of frequency components by: performing a Discrete Fourier Transform (DFT) of said first binary series,performing a Discrete Fourier Transform (DFT) of said second binary series, andfiltering common frequency components of the first and second binary series from said first series of events. 5. The wearable electronic equipment according to claim 1, wherein the motion sensor is an acceleration sensor. 6. The wearable electronic equipment according to claim 1, wherein the wearable electronic equipment is configured to determine a muscular activity parameter and the muscular activity parameter includes cadence and/or one or ... | Suunto Oy | A61B005/00 A61B005/0205 A61B005/11 A61B005/0245 A61B005/25 A61B005/296 A61B005/389
US20160128597 | A1 | 2016-05-12 | THIN PLANAR BIOLOGICAL SENSOR | 2015-09-04 | A thin planar biological sensor is provided, comprising at least one flexible electrode used for contacting a subject's skin and measuring biological signal. The flexible electrode is made of silicon-silver-based material, in which a ratio of silicon to silver-compound is about 7:3. A conductive plate is coupled to the flexible electrode. A buffer layer covers the conductive plate and the flexible electrode. A printed circuit board is used to receive the measured biological signal. By employing the novel proposed biological sensor, it is advantageous of measuring biological signals not only EEG signals but also ECG and EMG signals. Besides being thin, planar, flexible and non-allergy attributing a low cost to the proper fabrications makes the proposed sensor characterized with a potential for becoming an important tool of medical measurement. | 1. A thin planar biological sensor, which applies to measuring biological signals of a subject, comprising at least one flexible electrode used for contacting said subject's skin and measuring biological signals of said subject, wherein said flexible electrode is made of silicon-silver-based material, in which a ratio of silicon to silver-compound is 7:3;at least one conductive plate, electrically coupled to said flexible electrode and disposed below said flexible electrode;a buffer layer, disposed below said conductive plate such that said conductive plate is disposed between said buffer layer and said flexible electrode; anda printed circuit board, disposed below said buffer layer and electrically coupled to said conductive plate and said flexible electrode for receiving biological signals of said subject. 2. The thin planar biological sensor according to claim 1, wherein said silver-compound further comprises silver and silicon-dioxide (SiO2), in which a ratio of silver to silicon-dioxide is 1:9. 3. The thin planar biological sensor according to claim 1, wherein said conductive plate is made of metal. 4. The thin planar biological sensor according to claim 1, wherein said buffer layer is made of soft silicon so as to cover said flexible electrode and said conductive plate. 5. The thin planar biological sensor according to claim 1, further comprising a signal transmitting and receiving element, which is connected to said printed circuit board through at least one wire so as to receive biological signals of said subject and to transmit said biological signals to a remote end for further analyses. 6. The thin planar biological sensor according to claim 5, wherein said signal transmitting and receiving element is connected to said remote end through a wireless connection. 7. The thin planar biological sensor according to claim 5, further comprising a housing disposed outside said signal transmitting and receiving element, and said housing is made of anti-electrostatic and anti-electromagnetic material. 8. The thin planar biological sensor according to claim 7, wherein said housing is made by an injection molding process. 9. The thin planar biological sensor according to claim 1, wherein said printed circuit board is selected to be a flexible printed circuit (FPC). 10. The thin planar biological sensor according to claim 1, wherein said biological signals of said subject can be selected from a group consisting of EEG, ECG, EMG, and EOG signals. 11. The thin planar biological sensor according to claim 1, further disposed on an elastic belt, in which said elastic belt surrounds said subject's head such that said flexible electrode contacts said subject's head much more adhesively. 12. The thin planar biological sensor according to claim 1, wherein a number of said flexible electrode can be one or more than one, and a number of said conductive plate connected to said flexible electrode can be one or more than one. |  | A61B005/0478 A61B005/0496 A61B005/0488 A61B005/00 A61B005/0402
US20160143590 | A1 | 2016-05-26 | ELECTROCARDIOGRAPHY SYSTEM | 2015-11-19 | Embodiments include an electrocardiography system that includes or is connected to a sensor arrangement, wherein the sensor arrangement includes surface electrodes that record electric signals. The electrocardiography system includes a signal processing unit, with filters and amplifiers, which is electrically connected to the sensor arrangement. The electrocardiography system includes a signal evaluation unit connected via a first signal input to the signal processing unit. The signal evaluation unit processes one or more prepared ECG signals provided by the signal processing unit in order to recover an electrocardiogram signal and to display the electrocardiogram signal. The signal evaluation unit includes a second signal input that receives secondary signals obtained by body sensors or monitors, wherein the signal evaluation unit processes the prepared ECG signals depending on the secondary signals. | 1. An electrocardiography system comprising: a sensor arrangement, wherein the sensor arrangement comprises surface electrodes that record electric signals;a signal processing unit electrically connected to the sensor arrangement; and,a signal evaluation unit connected via a first signal input to the signal processing unit;wherein the signal evaluation unit is configured to process one or more prepared electrocardiograph (ECG) signals provided by the signal processing unit in order to obtain an electrocardiogram signal, anddisplay the electrocardiogram signal or output the electrocardiogram signal to be displayed,wherein the signal evaluation unit comprises a second signal input that receives one or more secondary signals obtained via at least one secondary sensor, wherein the at least one secondary sensor comprise one or more body sensors or monitors and,wherein the signal evaluation unit is further configured to process the prepared ECG signals in accordance with the secondary signals. 2. The electrocardiography system according to claim 1, wherein the signal evaluation unit is further configured to form or to adapt a compensation function in accordance with the secondary signals and to apply the compensation function to the prepared ECG signal. 3. The electrocardiography system according to claim 1, wherein the sensor arrangement further comprises the at least one secondary sensor, wherein the at least one secondary sensor differs from the surface electrodes, wherein the at least one secondary sensor is configured to detect a physical variable and delivery the one or more secondary signals, andthe signal evaluation unit further comprises an adaptive filter with a transfer function, wherein the transfer function is one or more of influenced by and controlled by the one or more secondary signals. 4. The electrocardiography system according to claim 3, wherein the transfer function is configured to compensate for a transfer function between ECG signals recorded outside an MRI scanner and ECG signals recorded within magnetic fields of the MRI scanner. 5. The electrocardiography system according to claim 3, wherein the signal evaluation unit is further configured to form an initial transfer function for the adaptive filter, wherein a change of an ECG measurement of a patient introduced in the MRI scanner is detected to form the initial transfer function. 6. The electrocardiography system according to claim 3, wherein the at least one secondary sensor is further configured to record an intracardial ECG, and wherein the one or more secondary signals comprise an intracardial ECG. 7. The electrocardiography system according to claim 3, wherein the at least one secondary sensor is an acoustic sensor configured to record heart sounds. 8. The electrocardiography system according to claim 3, wherein the at least one secondary sensor is a pulse oximeter configured to measure a pulse. 9. The electrocardiography system according to claim 1, wherein th... | BIOTRONIK SE & CO. KG | A61B005/00 A61B005/044 A61B005/055 A61B008/06 A61B005/042 A61B005/024 A61B005/026 A61B005/04 A61B007/02
US09329161 | B2 | 2016-05-03 | Monitoring of the functionality of a converter of a breath analysis apparatus | 2012-06-27 | A device for the gas analysis of a gas mixture includes a converter configured to convert a first gas component into a target gas component, a sensor system configured to detect the target gas component or another component of the gas mixture after the conversion by the converter, and an evaluating apparatus. The converter is further configured to change the concentration of a second gas component or to cause a conversion to a second target gas component. The sensor system is configured to determine the concentration of the second gas component or of the second target gas component. The evaluating apparatus is configured to determine a value for the aging of the converter on the basis of the concentration of the second gas component or of the second target gas component. | 1. An apparatus for the gas analysis of a gas mixture, comprising: a converter configured (i) to convert a first gas component into a first target gas component, and (ii) to modify a concentration of a second gas component or to cause conversion in a second target gas component;a sensor system configured (i) to detect the first target gas component or of another component of the gas mixture after the conversion by the converter, and (ii) to determine the concentration of the second gas component or of the second target gas component; andan evaluation device configured to ascertain a value of the ageing of the converter with the aid of the concentration of the second gas component or of the second target gas component. 2. The apparatus as claimed in claim 1, wherein the converter carries out conversion of the second gas component into a second target gas component. 3. The apparatus as claimed in claim 1, wherein the converter is further configured to cause storage of the second gas component. 4. The apparatus as claimed in claim 1, wherein the first gas component is nitrogen monoxide and the first target gas component is nitrogen dioxide. 5. The apparatus as claimed in claim 1, wherein the evaluation device is further configured to determine the value of the ageing of the converter from the fact that, when after the gas mixture enters, the concentration of the second gas component or of the second target gas component is greater or less than a threshold value. 6. The apparatus as claimed in claim 1, wherein the evaluation device is further configured to determine the value of the ageing of the converter with the aid of a gradient of the second gas component or of the second target gas component. 7. An asthma breath analysis apparatus comprising: an apparatus for the gas analysis of a gas mixture, including a converter configured (i) to convert a first gas component into a first target gas component, and (ii) to modify a concentration of a second gas component or to cause conversion in a second target gas component,a sensor system configured (i) to detect the first target gas component or of another component of the gas mixture after the conversion by the converter, and (ii) to determine the concentration of the second gas component or of the second target gas component, andan evaluation device configured to ascertain a value of the ageing of the converter with the aid of the concentration of the second gas component or of the second target gas component. | Robert Bosch GmbH | A61B005/08 G01N033/497 G01N033/00
US09352058 | B2 | 2016-05-31 | Method of preparing iron oxide nanoparticles coated with hydrophilic material, and magnetic resonance imaging contrast agent using the same | 2012-02-06 | The present invention relates to a method of preparing biocompatible iron oxide nanoparticles by coating iron oxide nanoparticles having improved magnetism via annealing treatment using salt particles with a hydrophilic material and to a magnetic resonance imaging (MRI) contrast agent including the iron oxide nanoparticles prepared thereby. Among hydrophilic materials, carboxymethyl dextran (CM-dextran) is the most efficient in terms of stabilizing the annealed iron oxide nanoparticles and exhibiting contrast effects. | 1. A method of preparing iron oxide nanoparticles coated with a hydrophilic material, comprising: a) annealing a power mixture comprising salt particles and magnetic iron oxide nanoparticles coated with a hydrophobic organic material, thus obtaining iron oxide nanoparticles having no organic material; andb) coating a surface of the iron oxide nanoparticles having no organic material with a hydrophilic material,wherein the annealing in a) is performed under air conditions. 2. The method of claim 1, wherein the hydrophobic organic material coated to the iron oxide nanoparticles in a) is at least one selected from the group consisting of a C4˜C25 fatty acid, a C4˜C25 aliphatic alcohol, and a C4˜C25 aliphatic amine. 3. The method of claim 2, wherein the fatty acid is oleic acid. 4. The method of claim 1, wherein the salt particles in a) is at least one selected from the group consisting of sodium sulfate (Na2SO4), sodium bicarbonate (NaHCO3), potassium bicarbonate (KHCO3), calcium bicarbonate (Ca(HCO3)2), sodium chloride (NaCl), potassium chloride (KCl) and calcium chloride (CaCl2). 5. The method of claim 1, wherein the salt particles in a) have a diameter of 1˜500 μm. 6. The method of claim 1, wherein the magnetic iron oxide nanoparticles in a) are obtained by reacting an iron complex comprising iron as a center atom and a C4˜C25 organic acid group (carboxylate) attached thereto as a ligand with a C4˜C25 fatty acid. 7. The method of claim 6, wherein the iron complex is iron oleate. 8. The method of claim 6, wherein the fatty acid is oleic acid. 9. The method of claim 1, wherein the magnetic iron oxide nanoparticles in a) have a diameter of 1˜100 nm. 10. The method of claim 1, wherein the magnetic iron oxide nanoparticles and the salt particles in a) have a weight ratio ranging from 1:10 to 1:1000. 11. The method of claim 1, wherein the annealing in a) is performed by carrying out heating at 400˜600° C. for 1˜10 hr in an air atmosphere. 12. The method of claim 1, wherein, after annealing in a), removing the salt particles is additionally performed. 13. The method of claim 12, wherein the removing the salt is performed using centrifugation. 14. The method of claim 12, wherein the temperature is decreased to room temperature before the removing the salt after annealing in a). 15. The method of claim 1, wherein the hydrophilic material in b) is a dextran derivative, polyacrylic acid, starch, silica, polyethyleneglycol (PEG) or PEG-phosphate (PO-PEGs). 16. The method of claim 15, wherein the dextran derivative is diethylaminoethyl-dextran (DEAE-dextran), dextran sulfate (DS) or carboxymethyl dextran (CM-dextran). 17. The method of claim 1, wherein removing the salt are simultaneously performed with the coating in b). | HANWHA CHEMICAL CORPORATION | A61B005/055 B05D003/00 A61K049/18 C01G049/02 C09C001/24 B82Y030/00 C09C003/08
US09326731 | B2 | 2016-05-03 | Method of measuring an artefact removed photoplethysmographic (PPG) signal and a measurement system | 2012-03-26 | A method of measuring an artifact removed photoplethysmographic (PPG) signal and a measurement system for measuring an artifact removed photoplethysmographic (PPG) signal are provided. The method comprises obtaining a first set of PPG signals from a plurality of detectors at respective measurement sites using a first illumination; obtaining a second set of PPG signals from the plurality of detectors using a second illumination; obtaining at least two pairs of PPG signals, each pair comprising one PPG signal from the first set and one PPG signal from the second set, and for each pair, computing an artifact reference signal to obtain a candidate PPG signal; and choosing one of the candidate PPG signals as the artifact removed PPG signal. | 1. A method of producing an artefact removed photoplethysmographic (PPG) signal, the method comprising: obtaining a first set of PPG signals from a plurality of detectors at respective measurement sites using a first illumination;obtaining a second set of PPG signals from the plurality of detectors using a second illumination;obtaining at least two pairs of PPG signals, each pair comprising one PPG signal from the first set and one PPG signal from the second set;computing an artefact reference signal based on the each pair of PPG signals using a vector subspace approach by using, for the pairs of PPG signals, artefact reference signals at an angle of between0∼π2to a PPG signal in the respective pair of PPG signals in the vector subspace and separating the used artefact reference signals into a plurality of subspaces, further wherein from each of said plurality of subspaces, an artefact reference signal selected at the centre of the respective subspace is applied to a filter with one PPG signal from the first set and one PPG signal from the second set to reduce a vector space of artefact reference signals until an optimum artefact reference signal is identified and applied to the filter with one PPG signal from the first set and one PPG signal from the second set as a result of which a candidate PPG signal is obtained from which the artefact has been removed by the filter; choosing one of the candidate PPG signals as the artefact removed PPG signal, wherein the step of choosing one of the candidate PPG signals comprises using one or more parameters to compare the candidate PPG signals; andestimating vital signs using the chosen candidate PPG signal. 2. The method as claimed in claim 1, wherein the parameters comprise an averaged standard deviation of maximum peak points and minimum peak points of each candidate PPG signal. 3. The method as claimed in claim 1, wherein the parameters comprise a cross-correlation of two separated segments of each candidate PPG signal. 4. The method as claimed in claim 1, further comprising applying a force at the plurality of detectors such that signal amplitudes of the first set of PPG signals and/or the second set of PPG signals are maximum. 5. The method as claimed in claim 4, comprising computing an area under curve measurement using waveforms of the PPG signals to determine whether the signal amplitudes are maximum. 6. The method as claimed in claim 5, wherein the computing an area under curve measurement comprises sampling each waveform and filtering each sample, further wherein the area under curve measurement is a summation of absolute values of the filtered samples. 7. The method as claimed in claim 1, further comprising providing the first and second illumination using a light emitting diode. 8. The method as claimed in claim 7, wherein the light emitting diode is positioned such that motion artefacts are minimized in the first set of PPG signals and/or the second set of PPG signals. 9. The method a... | NITTO DENKO CORPORATION | A61B005/1455 A61B005/00 A61B005/024 A61B005/0295 A61B005/026
US20160143530 | A1 | 2016-05-26 | Smart Patient Monitoring | 2015-11-24 | A patient monitoring method wherein data is received from a transducer by an electronic device, the electronic device receives input from the patient regarding timing of various events, the data is processed to incorporate the event timing data, and the data is then transmitted to a terminal. The terminal is accessible by, e.g., a healthcare provider. The transducer is used to perform medical diagnostic tests, such as fetal non-stress tests, electrocardiography, Holter monitoring, and electroencephalography. In the event that an emergency is identified, the healthcare provider can take immediate action. The utilization of an electronic device, e.g., a smartphone, by patients not only ensures ease of use and compliance with testing, but also conveniently allows patients to take an active role in their healthcare. | 1. A patient monitoring method, comprising the steps of: recording signals collected from a transducer;receiving an event input from a user;processing recorded data with the event input;graphing the recorded data as a function of time to generate a data graph;graphing the event input as a function of time to generate an event graph;overlaying the data graph with the event graph; andtransmitting the data graph to a terminal. 2. The method of claim 1, further comprising the step of tagging the event input with a date and a time. 3. The method of claim 1, wherein the transducer is configured to receive electrical impulses from a brain. 4. The method of claim 1, wherein the transducer is configured to receive electrical impulses from a heart. 5. The method of claim 1, wherein the transducer is configured to generate sound waves. 6. The method of claim 1, wherein the event input includes a list of user-selectable events. 7. A non-transitory computer-readable medium storing a logic that, when executed by a processor of a computer, causes the computer to perform a method comprising the steps of: recording data collected from the transducer;receiving an event input from a user;processing recorded data with the event input;graphing the recorded data as a function of time to generate a data graph;graphing the event input as a function of time to generate an event graph;overlaying the data graph with the event graph; andtransmitting processed data to a terminal. 8. The method of claim 7, further comprising the step of tagging the event input with a date and a time. 9. The method of claim 7, wherein the transducer is configured to receive electrical impulses from a brain. 10. The method of claim 7, wherein the transducer is configured to receive electrical impulses from a heart. 11. The method of claim 7, wherein the transducer is configured to generate sound waves 12. The method of claim 7, wherein the event input includes a list of user-selectable events. 13. A computer device with memory stored instructions which, when executed by a processor, causes the processor to perform the method comprising the steps of: recording data collected from the transducer;receiving an event input from a user;processing recorded data with the event input;graphing the recorded data as a function of time to generate a data graph;graphing the event input as a function of time to generate an event graph;overlaying the data graph with the event graph; andtransmitting processed data to a terminal. 14. The method of claim 13, further comprising the step of tagging the event input with a date and a time. 15. The method of claim 13, wherein the transducer is configured to receive electrical impulses from a brain. 16. The method of claim 13, wherein the transducer is configured to receive electrical impulses from a heart. 17. The method of claim 13, wherein the transducer is configured to generate sound waves. 18. The method of claim 13, wherein the event input inclu... |  | A61B005/00 A61B005/0402 A61B005/0476
US20160143544 | A1 | 2016-05-26 | System for Calculating Biological Information Under Exercise Load, Biological Information Calculation Method, and Portable Information Terminal | 2013-06-27 | A system for calculating biological information under exercise load, including a sensor device for measuring a motion of a user, a portable heart rate sensor for measuring a heart rate of a user, and a server including a biological information calculation function, wherein the system measures a motion of a user under exercise load on the user, measures the heart rate of the user after the stop of the exercise load, measures a temporal difference between the stop of the exercise load and measurement of the heart rate, estimates a drop in the heart rate based on the temporal difference and finds an estimated heart rate immediately before the stop of exercise of the user, and calculates biological information of the user under the exercise load based on the motion of the user under the exercise load and the estimated heart rate. | 1. A system for calculating biological information under exercise load, comprising: a sensor device for measuring a motion of a user;a portable heart rate sensor for measuring a heart rate of a user; anda server comprising a biological information calculation function,wherein the system:measures a motion of a user under exercise load on the user;measures the heart rate of the user after the stop of the exercise load;measures a temporal difference between the stop of the exercise load and measurement of the heart rate;estimates a drop in the heart rate based on the temporal difference and finds an estimated heart rate immediately before the stop of exercise of the user; andcalculates biological information of the user under the exercise load based on the motion of the user under the exercise load and the estimated heart rate. 2. The system for calculating biological information under exercise load according to claim 1, wherein the server comprises heart rate drop data corresponding to a temporal difference between the stop of the exercise load and measurement of the heart rate in order to estimate a drop in the heart rate based on the temporal difference. 3. The system for calculating biological information under exercise load according to claim 2, wherein a time zone in which the heart rate is measurable is within one minute after the stop of the exercise load. 4. The system for calculating biological information under exercise load according to claim 2, wherein the heart rate drop data is created based on post-exercise heart rate change data previously collected from a plurality of subjects. 5. The system for calculating biological information under exercise load according to claim 2, wherein the heart rate drop data is created per user by learning data previously measured per user. 6. The system for calculating biological information under exercise load according to claim 2, wherein the portable heart rate sensor is a portable terminal comprising a built-in camera, andcontinuously shoots the skin of a finger or the face of the user by the camera, and detects a periodical change in color of the skin due to blood flow conversion thereby to measure the heart rate. 7. The system for calculating biological information under exercise load according to claim 2, wherein the sensor device comprises an acceleration sensor capable of measuring a motion or an orientation of a person, and an RTC for holding current time information or calendar information. 8. The system for calculating biological information under exercise load according to claim 7, wherein the server comprises:a heart rate correction means for correcting a measurement result of the heart rate after the stop of the exercise load into a heart rate under the exercise load based on the heart rate drop data and the temporal difference;an exercise characteristic amount calculation means for calculating the exercise characteristic amount of the user under the exercise load based on a motio... |  | A61B005/0205 A61B005/00
US20160129217 | A1 | 2016-05-12 | HYPNOTIC SYSTEM AND METHOD FOR THE SAME | 2015-04-02 | A hypnotic system includes a physiological parameter detection device, an audio information comparison device, a communication device and an audio device. The physiological parameter detection device detects the physiological parameter of a user for generating a physiological parameter signal. The audio information comparison device receives the physiological parameter signal, locates a respective audio information from an internal audio-comparing table, and then generates a corresponding audio signal for transmitting to the communication device. The audio device receives the audio signal forwarded by the communication device and then plays the corresponding audio information matching user's physiological parameter so as to serve a hypnotic purpose to the user. | 1. A hypnotic system, comprising: a physiological parameter detection device for detecting at least one physiological parameter of a user so as thereby to generate and then forward a physiological parameter signal;an audio information comparison device, further comprising: a first transmission module, communicatively coupled with the physiological parameter detection device for receiving and then forwarding the physiological parameter signal;a comparison and storage module, storing an audio-comparing table of physiological parameters, further including a plurality of audio information; anda first processing module, electrically coupled with the first transmission module and the comparison and storage module, being to receive the physiological parameter signal so as thereby to locate at least one of the plurality of the audio information in the audio-comparing table of physiological parameters stored in the comparison and storage module that matches the at least one physiological parameter of the user and thus further to generate and then forward an audio signal;a communication device, further comprising a second transmission module communicatively coupled with the first transmission module of the audio information comparison device for receiving and then forwarding the audio signal; andan audio device for receiving the audio signal, further comprising a third transmission module and an audio module, the third transmission module being communicatively coupled with the second transmission module of the communication device to receive and then forward the audio signal, the audio module being electrically coupled with the third transmission module to receive the audio signal and then play the corresponding audio information. 2. The hypnotic system of claim 1, wherein the first processing module of the audio information comparison device judges the physiological parameter to determine a sleep cycle information corresponding to the user so as thereby to locate the audio information that matches the physiological parameter from the audio-comparing table of physiological parameters. 3. The hypnotic system of claim 1, wherein the physiological parameter is one of a brain wave, a heartbeat, a breath and a myoelectricity. 4. The hypnotic system of claim 1, wherein the audio information comparison device is a network server. 5. The hypnotic system of claim 1, wherein the communication device is one of a mobile phone and a tablet computer. 6. The hypnotic system of claim 1, wherein the audio device is one of a earphone and a speaker. 7. A hypnotic method, comprising the steps of: (a) detecting at least one first physiological parameter of a user so as thereby to generate and then transmit a corresponding first physiological parameter signal;(b) receiving the first physiological parameter signal so as thereby to locate at least one first audio information matching the at least one first physiological parameter from an audio-comparing table of physiologic... |  | A61M021/02 A61B005/00 A61B005/024 A61B005/0476 A61B005/0488
US20160139051 | A1 | 2016-05-19 | APPARATUS AND METHOD FOR DISTINGUISHING BETWEEN DIFFERENT TISSUE TYPES USING SPECIFIC RAMAN SPECTRAL REGIONS | 2014-07-11 | A portable apparatus and method for distinguishing between different tissue types, such as normal tissue, necrotic tissue, and tumor tissue are provided, where the apparatus includes a housing and a plurality of Raman spectrometers disposed within the housing, each spectrometer having a different spectral region. A processor is provided in communication with the plurality of spectrometers, the processor analyzing output from the plurality of spectrometers to identify the tissue type of the tissue sample. A method of selecting the spectral regions which provide a desired combined classification accuracy for determining the tissue type is also provided. | 1. A portable apparatus for distinguishing between different tissue types in a tissue sample, the apparatus comprising: a housing;a plurality of Raman spectrometers disposed within the housing, each spectrometer having a different spectral region; anda processor in communication with the plurality of spectrometers, the processor analyzing output from the plurality of spectrometers to identify the tissue type of the tissue sample. 2. The apparatus of claim 1, wherein the different tissue types include normal tissue, necrotic tissue, and tumor tissue. 3. The apparatus of claim 1, wherein the plurality of spectrometers includes five spectrometers. 4. The apparatus of claim 1, wherein each spectral region does not exceed 20 wavenumbers. 5. The apparatus of claim 1, wherein the spectrometers have a first spectral region including wavenumbers 1657-1660, a second spectral region including wavenumbers 1153-1172, a third spectral region including wavenumbers 1002-1004, a fourth spectral region including wavenumbers 1106-1123, and a fifth spectral region including wavenumbers 1254-1268. 6. The apparatus of claim 1, wherein the different tissue types include normal grey matter brain tissue, necrotic brain tissue, and glioblastoma tumor tissue. 7. The apparatus of claim 6, wherein the processor identifies a Raman spectrum for necrotic brain tissue as being characterized by an increased protein content and a decreased lipid content compared to normal brain tissue. 8. The apparatus of claim 6, wherein the processor identifies a Raman spectrum for glioblastoma as being characterized by a decreased lipid content, a decreased cholesterol content, and an increased nucleic acid content compared to normal grey matter brain tissue. 9. The apparatus of claim 6, where the processor identifies a Raman spectrum for glioblastoma as being characterized by an increased lipid content, and increased nucleic acid content, and a decreased protein content compared to necrotic brain tissue. 10. The apparatus of claim 1, further comprising a probe connected to the housing, and a light source disposed within the housing for illuminating the tissue sample via the probe. 11. The apparatus of claim 10, wherein the apparatus further includes a tracking system to detect a position of the probe in real time relative to an anatomical landmark associated with the tissue sample. 12. A method for distinguishing between different tissue types in a tissue sample, the method comprising: providing a portable apparatus having a housing, a light source disposed within the housing, and a plurality of Raman spectrometers disposed within the housing, each spectrometer having a different spectral region;illuminating the tissue sample using the light source;receiving light from the tissue sample with the plurality of spectrometers; andanalyzing output from the plurality of spectrometers to identify the tissue type of the tissue sample. 13. The method of claim 12, wherein the different tissu... |  | G01N021/65 A61B005/1455 A61B005/145 A61B005/00
US20160143547 | A1 | 2016-05-26 | RAPID RATE-ESTIMATION FOR CELL PHONES, SMART WATCHES, OCCUPANCY, AND WEARABLES | 2015-09-24 | Techniques for respiratory and metabolic monitoring in mobile devices, wearable computing, security, illumination, photography, and other applications may use a phosphor-coated broadband white LED to produce broadband light, which may be transmitted along with ambient light to a target (e.g., ear, face, wrist, or the like). Some scattered light returning from a target may be passed through a spectral filter to produce multiple detector regions sensitive to a different waveband and/or wavelength range, and the detected light may be analyzed to determine a measure of a respiratory rate or effort. In the absence of LED light, ambient light may be sufficient illumination for analysis. The disclosed techniques may provide identifying features of type or status of a tissue target (e.g., respiratory rate, heart rate, heart rate variability, heart function, lung function, fat content, hydration status, confirmation of living tissue, and the like). | 1. A method, comprising: detecting a repetitive event using light;detecting another repetitive event using the light; andgenerating an output measure that is a function of an estimated rate of the repetitive event and the another repetitive event, the output measure being based, at least in part, on a time interval elapsed between the repetitive event and the another repetitive event. 2. The method of claim 1, wherein the light is backscattered from a subject. 3. The method of claim 1, wherein the light is transmitted through a subject. 4. A method, comprising: noninvasively detecting broadband light returning after interaction with a subject, the broadband light being further separated into one or more wavebands;determining a timing of a repetitive event using the broadband light detected;determining another timing of another repetitive event using the broadband light detected; andgenerating an output measure, the output measure being a function of an estimated rate of the repetitive event, the output measure based on the timing and the another timing. 5. A method, comprising: noninvasively detecting broadband light, the broadband light returning after interacting with a subject, the broadband light being further separated into one or more wavebands;using a waveform analysis to detect a timing of a repetitive event using the broadband light detected; andgenerating an output measure that is a function of an estimated rate of the repetitive event and another repetitive event, the output measure being based on the waveform analysis. 6. The method of claim 5, further comprising generating illumination using a solid-state, broadband light source. 7. The method of claim 5 wherein the waveform analysis comprises time-frequency transformation. 8. The method of claim 5, wherein the waveform analysis comprises a modal frequency analysis. 9. The method of claim 5, wherein the output measure is a respiratory rate. 10. The method of claim 5, wherein the output measure is a heart rate. 11. The method of claim 5, wherein the output measure is a calorie use rate. | AliphCom | A61B005/024 A61B005/00 A61B005/08
US09743868 | B2 | 2017-08-29 | Circuitry to allow low current operation of a device capable of determining a blood property | 2014-11-20 | Systems, methods, and devices of the various embodiments provide a device capable of determining a blood property based at least in part on the measurement of the amount of light received by a receiver circuit with a limited current capacity by charging a capacitor with a low voltage power supply and intermittently discharging the capacitor through a light emitting diode. In some embodiments the device may be a pulse oximeter capable of taking blood oxygen readings. In some embodiments the device may be a heart rate monitor to determine a heart rate based on an amount of light passed through tissue. The various embodiments may enable pulse oximeters and/or heart rate monitors to be incorporated into small unobtrusive body patches, while enabling the pulse oximeters to operate from a power output equivalent to that of a small coin cell battery or printed cell battery. | 1. A device, comprising: a capacitor configured to be charged by a low voltage power source;at least one light-emitting diode (LED) connected to the capacitor by a switch;a receiver circuit configured to measure light emitted by the at least one LED after the light passes through tissue of a patient, wherein the receiver circuit comprises a first photodetector; anda processor connected to the switch and the receiver circuit, wherein the processor is configured with processor executable instructions to perform operations comprising: determining a time of a pulse maxima and a time of a pulse minima based on measurements of light received by the first photodetector of the receiver circuit;determining a derivative estimate of a sampled pulse curve of a group of pulse measurements taken at the determined time of the pulse maxima, wherein the group of pulse measurements are based on the measurements of light received by the first photodetector of the receiver circuit;adjusting a time delay between the determined time of the pulse maxima and the determined time of the pulse minima based on the determined derivative estimate of the sampled pulse curve of the group of pulse measurements taken at the determined time of the pulse maxima;controlling the switch to provide charge from the capacitor to the at least one LED to cause the at least one LED to emit light at the expiration of the adjusted time delay;controlling the switch to isolate the capacitor from the at least one LED to cause the at least one LED to stop emitting light after a period of time;measuring an amount of light received by the receiver circuit after the period of time; anddetermining a blood property based at least in part on the measurement of the amount of light received by the receiver circuit after the period of time. 2. The device of claim 1, wherein the processor is configured with processor-executable instructions to perform operations such that determining the blood property based at least in part on the measurement of the amount of light received by the receiver circuit after the period of time comprises determining a heart rate based at least in part on the measurement of the amount of light received by the receiver circuit after the period of time. 3. The device of claim 1, wherein the processor is configured with processor-executable instructions to perform operations such that determining the blood property based at least in part on the measurement of the amount of light received by the receiver circuit after the period of time comprises determining a blood oxygen level based at least in part on the measurement of the amount of light received by the receiver circuit after the period of time. 4. The device of claim 3, wherein the processor is configured with processor-executable instructions to perform operations further comprising: determining when the capacitor is charged by the low voltage power source to a predetermined voltage; andcontrolling the switch to provide ch... | QUALCOMM Incorporated | A61B005/1455 A61B005/00 A61B005/024 A61B005/026
US09339661 | B2 | 2016-05-17 | Medical device adjusting operation when used with non-authenticated patient parameter collecting accessory | 2015-04-03 | Embodiments are directed to a medical device, such as a defibrillator, for use with an accessory capable of collecting a parameter of a patient. The medical device is capable of at least performing a basic functionality, an advanced functionality, and of defibrillating the patient. The medical device includes an energy storage module within a housing for storing an electrical charge that is to be delivered to the patient for the defibrillating. The medical device includes a processor structured to determine whether a data set received from the accessory confirms or not a preset authentication criterion about the accessory. Although when the accessory is coupled to the housing the medical device is capable of the defibrillating and the basic functionality, the medical device is capable of the advanced functionality only when the accessory is coupled to the housing and it is determined that the preset authentication criterion is confirmed. Embodiments also include methods of operation and a programmed solution. | 1. A system for a medical device for use with an accessory capable of collecting a parameter of a patient, the medical device capable of at least one of a basic functionality, an advanced functionality, and of defibrillating the patient, the system comprising: a data storage device in the accessory configured to store a data set;a reader structured to read a data set from the data storage device;a processor coupled to the reader and structured to determine whether the data set from the data storage device confirms a preset authentication criterion about the accessory,in which the defibrillating and the basic functionality of the medical device are enabled and the advanced functionality of the medical device is disabled when the accessory is coupled to the medical device and it is determined that the preset authentication criterion is not confirmed, andin which the medical device is capable of the defibrillation, the basic functionality and the advanced functionality when the accessory is coupled to the medical device and it is determined that the preset authentication criterion is confirmed. 2. The system of claim 1, in which the data storage device comprises a programmable read-only memory, an erasable programmable read-only memory, a flash memory, a radio frequency identification tag, or a bar code. 3. The system of claim 1, in which the accessory comprises an electrode. 4. The system of claim 3, in which the electrode comprises an ECG electrode. 5. The system of claim 1, in which the accessory is selected from a group consisting of a non-invasive blood pressure hose, a pressure cuff, one or more leads, a temperature measuring device, a capnography device, a respiratory device, a communication device, and a physiological parameter sensor. 6. The system of claim 1, in which the reader comprises an radio frequency identification reader, a bar code reader, or an infrared reader. 7. The system of claim 1, in which the parameter includes one of capnography, pulse oximetry, non-invasive blood pressure, ECG with three or more leads, invasive blood pressure, temperature, heart rate, respiration rate and CPR performance monitoring. 8. The system of claim 1, in which the advanced functionality of the medical device is enabled for a limited time after it is determined that the preset authentication criterion is not confirmed. 9. The system of claim 1, in which a query is caused to be transmitted by the processor to the accessory, andthe data set is received responsive to the query. 10. The system of claim 1, in which the data set is received responsive to attaching the accessory to the housing. 11. The system of claim 1, in which the data set is received responsive to power-on of the device. 12. The system of claim 1, in which the received data set is encrypted,a decryption algorithm is applied by the processor to determine whether the received data set confirms the authentication or does not confirm the authentication. 13. The system of claim... | PHYSIO-CONTROL, INC. | A61N001/39 A61B005/01 A61B005/021 A61B005/0215 A61B005/0408 A61B005/08 A61B005/1455 G06F017/30 G06F021/60 H01M002/34 H01M006/50 H01M010/42 H01M016/00 H02J007/00 H02J007/34
US09341464 | B2 | 2016-05-17 | Method and apparatus for sizing and fitting an individual for apparel, accessories, or prosthetics | 2012-10-02 | A system for sizing and fitting an individual for apparel, accessories, or prosthetics includes at least one energy emitter configured to emit energy onto a field-of-view that contains an individual, and at least one energy sensor configured to capture reflected energy from within the field-of-view. A spatial measurement module calculates spatial measurements of a surface portion of the body of the individual when the individual is either stationary or moving about in real-time, based on data from the energy sensor. | 1. A method of determining a spatial distance of curvature along an unclosed curve that tracks a body surface of an individual for sizing and fitting the individual for apparel, accessories, or prosthetics, the method comprising: capturing data corresponding to the individual positioned within a field-of view of at least one energy emitter and at least one energy sensor, the at least one energy sensor configured to capture energy reflected from the individual and other objects within the field-of-view, wherein the data corresponds to an incomplete representation of a surface of the body of the individual;generating depth data for the field-of-view based on the captured data from the at least one energy sensor; andcalculating one or more spatial measurements of a portion of a body surface of the individual based on the generated depth data for the field-of-view for sizing and fitting the individual for apparel, accessories, or prosthetics, comprising: determining an estimated position of each of one or more body landmarks based at least in part on the generated depth data;determining an outline of the single body portion that is bounded by the estimated position(s), wherein said determining provides an incomplete representation of the surface of the portion of the body surface in question; selecting two or more terminuses of the outline of the single body portion;determining a spatial distance of curvature connecting at least two of the two or more terminuses using at least one of heuristic technique or approximation; andstoring the spatial distance of curvature as a measurement of the single body portion. 2. The method according to claim 1, wherein the data captured corresponding to the individual within the field-of-view is captured in real-time while the individual moves about in real-time. 3. The method of claim 1, wherein the data captured corresponding to the individual within the field-of-view is captured while the individual remains substantially motionless for a predetermined amount of time. 4. The method according to claim 3, wherein the individual is positioned so that a face of the individual is presented to the at least one energy sensor, with arms of the individual extended outwardly and away from opposite corresponding sides of the individual. 5. The method according to claim 1, wherein the at least one energy emitter emits at least one of a pattern of infrared light, a pattern of laser light, a pattern of ultraviolet light, a pattern of visible light, a pattern of X-rays, a pattern of microwaves, a pattern of radio waves, a pattern of sound waves, a pattern of ultrasound energy, or a pattern of thermal energy. 6. The method according to claim 1, wherein at least one additional energy sensor is positioned so as to view substantially the same field-of-view as the at least one energy sensor, and wherein the at least one energy sensor and the at least one additional energy sensor acquire a substantially non-overlapping spectra of ... | ATLAS5D, INC. | G01B011/24 G01B011/00 A61F002/50 A61B005/00 G06T007/60 A41H001/02 A61B005/107
US20160143805 | A1 | 2016-05-26 | CPR ASSISTANCE SYSTEM AND CPR MONITORING METHOD | 2015-11-20 | A CPR assistance system comprises a sensor for making a physiological measurement to generate a sensor signal which conveys information concerning the point in time at which venous return is completed. Time points are then determined at which CPR compression cycles should be administered, and output information is provided relating to the determined time points. The output information may be visual or audible advice to a person giving CPR or it may be control information for controlling an automated CPR system. | 1. A CPR assistance system, comprising: a sensor (70) for making a physiological measurement to generate a sensor signal which conveys information concerning the point in time at which venous return is completed; anda processor (74) for processing the sensor signal to determine time points at which CPR compression cycles should be administered based on the point in time at which venous return is completed; andan output device (76,78; 74) providing output information relating to the determined time points. 2. A system as claimed in claim 1, wherein the sensor signal conveys information which indicates when blood flow returns to zero after a CPR compression cycle. 3. A system as claimed in claim 1 or 2, wherein the sensor (70) comprises a blood flow sensor such as an external ultrasound flow sensor. 4. A system as claimed in claim 3, wherein the processor (74) is adapted to determine the time at the flow first becomes positive after a negative peak of blood flow. 5. A system as claimed in any preceding claim, wherein the sensor (70) comprises a pressure sensor for measuring the right atrial pressure or a plethysmographic sensor. 6. A system as claimed in any preceding claim, wherein the processor (74) is adapted to determine: the time at which the physiological signal remains between an upper and lower limit for a certain amount of time or the time at which the derivative of the physiological signal remains between an upper and lower limit for a certain amount of time; orthe time at which the physiological signal has its maximum value after the completion of a chest compression. 7. A system as claimed in any preceding claim, wherein the processor (74) is adapted to apply a delay after the sensed completion of the venous return to compensate for a delay between the physical location of the sensor and the actual completion of venous return. 8. A system as claimed in claim 7, wherein the processor (74) is adapted to determine the delay by measuring the delay between the start of a CPR compression cycle and the start of the response of the physiological measurement sensor signal to that CPR compression cycle. 9. An emergency care monitor comprising a CPR assistance system as claimed in any preceding claim, wherein the processor (74) is adapted to provide: an output warning signal when a CPR compression is started more than a predetermined time before and/or after the determined time points at which CPR compression cycles should be administered; oran indication of a suitable compression rate corresponding to the determined time points or previously determined time points. 10. An automated CPR device comprising: a chest compression system (80); anda CPR assistance system as claimed in any one of claims 1 to 8 for controlling the timing of operation of the chest compression system. 11. A CPR monitoring method, comprising: (60) making a physiological measurement to generate a sensor signal which conveys information concerning the point in time at... |  | A61H031/00 A61B005/021 A61B008/06
US20160142478 | A1 | 2016-05-19 | COMMUNICATION MANAGEMENT SYSTEM | 2014-06-23 | [PROBLEM] To provide a communication management system dynamically forming a communication network between users based on user's health condition and living behaviors to support more active communication between the users relating to the health or life.;[SOLVING MEANS] A computer system according to an embodiment is a communication management system configured to provide communication network information for users, and includes a storage apparatus configured to store for each of the users physiological information and/or living behavior information of the user, acquired dynamically in daily life, and a control apparatus configured to produce the communication network information being associated with other users to a first user based on the physiological information and/or living behavior information of the first user. The control apparatus is configured to produce the communication network information by using each of the dynamically acquired physiological information and/or living behavior information of the first user as a network formation element. | 1. A communication management system connected to a plurality of information terminal apparatuses and configured to provide communication network information for communication of users between the plurality of information terminal apparatuses, comprising: a storage apparatus configured to store for each of the users physiological information and/or living behavior information of the user, acquired dynamically in daily life; anda control apparatus configured to produce the communication network information being associated with other users to a first user based on the physiological information and/or living behavior information of the first user,wherein the control apparatus is configured to produce the communication network information by using each of the dynamically acquired physiological information and/or living behavior information of the first user as a network formation element. 2. The communication management system according to claim 1, wherein the control apparatus is configured to extract, from the physiological information acquired dynamically in each of users corresponding to network constituent members of the communication network information, first physiological information acquired in preceding measurement and second physiological information acquired in present measurement for each of the users, and to produce the communication network information including a status change representing a change from the first physiological information to the second physiological information. 3. The communication management system according to claim 1 or 2, wherein the control apparatus is configured to produce, based on definition information for grouping the physiological information into a plurality of types, the communication network information including an display area of each of the types, each user being mapped to an associated one of the display areas according to the type. 4. The communication management system according to any one of claims 1 to 3, wherein the control apparatus is configured to use the living behavior information acquired dynamically in each of users corresponding to network constituent members of the communication network information to produce daily life rhythm information including the living behavior information arranged in temporal sequence for each of the users, and the control apparatus is configured to produce, by using the produced life rhythm information as a network formation element, the communication network information in which the life rhythm information of each of the plurality of users is displayed. 5. The communication management system according to claim 4, wherein the control apparatus is configured to determine a life rhythm pattern of the life rhythm information of the first user based on definition information for grouping the life rhythm information into predetermined patterns, and the control apparatus is configured to produce life rhythm change information in which the life rhythm informati... |  | H04L029/08 G06F019/00 A61B005/00
US09326689 | B2 | 2016-05-03 | Thermally tagged motion tracking for medical treatment | 2012-05-08 | Motion tracking is performed with a thermal pattern within a patient. A pattern of different temperature is created in tissue, such as warming up tissue in a checkerboard pattern. The temperature pattern is used over time to track motion of the tissue. The tracked motion may be used to treat the tissue throughout at least part of a periodic cycle. | 1. A method of motion tracking in medical treatment, the method comprising: transmitting ultrasound to a plurality of locations in a pattern in tissue within a patient;heating the tissue in the pattern in response to the transmitting;scanning, with an imaging system, the tissue of the patient after the heating;detecting first spatial distribution of temperature in the tissue at a first time, the first spatial distribution including the pattern;detecting second spatial distribution of the temperature in the tissue at a second time, the second spatial distribution including the pattern; anddetermining a change in position between the first time and the second time from the first spatial distribution and the second spatial distribution of the temperature caused by the pattern. 2. The method of claim 1 further comprising: applying treatment to the tissue as a function of the change in position. 3. The method of claim 1 wherein transmitting comprises generating a plurality of beams to form the pattern as intersecting lines below a surface of a skin of the patient. 4. The method of claim 1 wherein heating comprises increasing the temperature at locations on the pattern by at least 2 degrees. 5. The method of claim 1 wherein scanning comprises scanning with an ultrasound or magnetic resonance system as the imaging system, the scanning comprising thermometry imaging. 6. The method of claim 1 wherein detecting the first spatial distribution and the second spatial distribution comprises calculating thermal strain. 7. The method of claim 1 wherein detecting the first spatial distribution and the second spatial distribution comprises identifying locations with elevated temperature. 8. The method of claim 1 wherein determining the change in the position comprises determining translation and rotation. 9. The method of claim 1 wherein determining the change comprises correlating the first spatial distribution with the second spatial distribution with different relative offsets and selecting the offset with a greatest correlation. 10. The method of claim 1 further comprising: displaying an image of the tissue with the imaging system. 11. The method of claim 1 further comprising: scanning, with the imaging system, the tissue of the patient prior to the heating; anddetermining a difference between the scanning prior to the heating and the scanning after the heating;wherein determining the change in position comprises determining the change as a function of the difference. 12. The method of claim 11 wherein the scanning prior to the heating comprises scanning during a first breathing cycle, and wherein the scanning after the heating comprises scanning during a second breathing cycle, the scanning prior and after each comprising acquiring a sequence of frames of data through the first breathing cycle and the second breathing cycle, respectively. 13. In a non-transitory computer readable storage medium having stored therein data representing instructions ... | Siemens Medical Solutions USA, Inc. | A61B005/05 A61B005/01 A61N007/02 A61B005/055 A61B005/11 A61B008/08 A61B017/00 A61B018/00
US09352130 | B2 | 2016-05-31 | Torque device for a guidewire | 2014-03-14 | A torque device for attaching to and selectively gripping or securing and releasing a guidewire to permit rotational and longitudinal manipulation of the guidewire to steer and control the guidewire. In one embodiment the torque device comprises a nose portion, a neck portion, and a base portion with varying outer diameters for improved rotational control. | 1. A medical torque device comprising: a nose portion having a first outer circumference;a neck portion having a second outer circumference, the second outer circumference being smaller than the first outer circumference; anda base portion having a third outer circumference which is greater than the second outer circumference;wherein a distal end of the neck portion is connected to a proximal end of the nose portion and the base portion with an inner lumen which extends through the nose portion, neck portion, and base portion and through which a guidewire passes,wherein the second outer circumference has a continuous non-spherical shape such that the guidewire is completely enclosed and no portion of the guidewire is exposed within the neck portion. 2. The medical torque device of claim 1, wherein the nose portion includes a beveled distal end opposite to an end which is connected to the neck portion. 3. The medical torque device of claim 1, wherein the inner lumen at a beveled distal end of the nose portion includes a conical front end recess which has a larger diameter than the lumen at a region where the nose portion is connected to the neck portion. 4. The medical torque device of claim 1, wherein the second outer circumference has a polygonal shape having at least three sides and at least three vertices where the sides meet. 5. The medical torque device of claim 1, wherein the base portion comprises a base top portion and a base bottom portion, wherein an urging means is configured to urge the base top portion toward the base bottom portion and with the guidewire positioned therebetween, so as to hold the guidewire in place in a closed position. 6. The medical torque device of claim 5, wherein the base top portion and the base bottom portion are configured so that an operator is able to apply a counter force which causes the base top portion to be moved away from the base bottom portion so that the urging force applied to the guidewire is removed and to cause the base top portion and the base bottom portion to move from the closed position to the open position. 7. The medical torque device of claim 1, wherein the base portion has a conical outer circumference wherein a distal end of the base portion at a region where the base portion is connected to the neck portion has a smaller circumference than a proximal end of the base portion. 8. A medical torque device comprising: a first component comprising: a nose portion having a first outer circumference;a neck portion having a second outer circumference, the second outer circumference being smaller than the first outer circumference;a base bottom portion having a guidewire-locking groove where a guidewire positioned within the base bottom portion is exposed; andan inner lumen which extends through the nose portion, neck portion and base bottom portion which houses the guidewire;a base top portion including a guidewire-locking ridge which is capable of applying a pinching force on the gui... | Adrian Moger | A61B005/00 A61M025/00 A61M025/09
US09339239 | B2 | 2016-05-17 | Methods and devices for optimization of magnetic resonance imaging protocols | 2013-09-10 | MRI techniques are widely and successfully applied in medicine and biophysics because MRI provides good contrast between different soft tissues without ionizing radiation. MRI protocols are optimized in four aspects: imaging parameters, k-space strategies, RF system calibration and contrast inhomogeneity correction. The signal intensities of normal and disease tissues are simulated, for example, using Bloch equations for an imaging sequence with tissue MR parameters. The relationships between imaging parameters and tissue contrasts are calculated using the numerically simulated signal intensities. The optimal imaging parameters and/or k-space strategies are determined to maximize image contrast and minimize artifacts with acceptable spatial-temporal resolution based on characterization of the imaging hardware. The RF system is optionally calibrated to improve the accuracy of the imaging parameters and reduce inter-scanner variability. Additionally, contrast-to-noise inhomogeneity caused by transmit field and receive sensitivity is optionally corrected by optimal flip angle and measured receiver sensitivity. | 1. A method for optimizing k-space strategies for MR acquisition, comprising: setting independent image parameters, wherein the independent image parameters are independent of a k-space strategy in an imaging sequence;simulating, using a computing device, at least one respective image quality metric for each of a plurality of k-space acquisitions, wherein the image quality metric comprises a contrast metric;optimizing the k-space strategy, wherein optimizing the k-space strategy comprises using the simulated image quality metrics for each of the k-space acquisitions to select a sampling order in k-space center that maximizes the contrast metric and minimizes an image artifact caused by the k-space strategy; andacquiring at least one MR image using the optimized k-space strategy, wherein selecting a sampling order in k-space center further comprises:filling a k-space zero line using a k-space acquisition having a maximum contrast metric, wherein the contrast metric comprises at least one of a contrast, contrast efficiency, CNR or CNR efficiency, anddesigning the sampling order to minimize the image artifact caused by the k-space strategy, wherein the image artifact caused by the k-space strategy comprises at least one of signal inhomogeneity, SNR inhomogeneity, contrast inhomogeneity, CNR inhomogeneity, signal loss, geometry distortion or image ghost. 2. The method of claim 1, wherein optimizing the k-space strategy further comprises selecting a sampling order for a predetermined k-space trajectory. 3. The method of claim 2, wherein the k-space trajectory comprises at least one of a rectilinear, radial, echo planar imaging, spiral, projection reconstruction, random k-space, under-sampled k-space, or partial k-space sampling trajectory, and the sampling order comprise at least one of a sequential, centric, interleave, reverse, or random sampling order. 4. The method of claim 1, wherein simulating, using a computing device, at least one respective image quality metric for each of a plurality of k-space acquisitions further comprises using Bloch equations for the imaging sequence and MR parameters of a tissue. 5. The method of claim 1, wherein the imaging sequence comprises at least one of a gradient echo sequence, echo planar sequence, spin echo sequence, or combinations thereof with or without magnetization preparation. 6. The method of claim 1, wherein the image quality metric further comprises at least one of signal intensity, SNR or SNR efficiency, and the contrast metric comprises at least one of a contrast, contrast efficiency, CNR or CNR efficiency. 7. The method of claim 1, wherein optimizing the k-space strategy further comprises optimizing the k-space strategy under limited conditions, wherein the limited conditions comprise at least one of a predetermined k-space trajectory, predetermined sampling order or predetermined imaging parameters. 8. The method of claim 2, wherein optimizing the k-space strategy further comprises selecting... | Ohio State Innovation Foundation | G06K009/00 A61B005/00 G01R033/54 A61B005/055 G01R033/48 G01R033/58
US09943228 | B2 | 2018-04-17 | Functional skin patch | 2015-11-11 | A functional skin patch having a first surface and a second surface opposite the first surface is provided. The functional skin patch includes a functional unit having a thermo harvester and an antenna unit. The thermo harvester has a first terminal thermally connected to the first surface and a second terminal. The antenna unit has a first terminal thermally connected to the second terminal of the thermo harvester and a second terminal thermally connected to the second surface. The antenna unit has a stacked layer structure including, in this sequence, a metal layer thermally connected to the second terminal of the thermo harvester, a ferrite layer thermally connected to the metal layer, and an antenna layer thermally connected to the ferrite layer. | 1. A functional skin patch, comprising: a first surface of the functional skin patch;a second surface of the functional skin patch and opposite the first surface;a functional unit disposed in the functional skin patch and comprising a thermo harvester, the thermo harvester having a first terminal thermally connected to the first surface and a second terminal; andan antenna unit disposed in the functional skin patch and having a first terminal thermally connected to the second terminal of the thermo harvester and a second terminal thermally connected to the second surface,wherein the antenna unit has a stacked layer structure comprising, in this sequence, a metal layer thermally connected to the second terminal of the thermo harvester, a ferrite layer thermally connected to the metal layer, and an antenna layer thermally connected to the ferrite layer. 2. The functional skin patch of claim 1, wherein the metal layer is a flexible metal foil having a thickness in a range between 5 μm and 1000 μm. 3. The functional skin patch of claim 1, wherein the ferrite layer is a flexible ferrite foil having a thickness in a range between 5 μm and 1000 μm. 4. The functional skin patch of claim 1, wherein the antenna layer includes an antenna pattern printed on the ferrite layer. 5. The functional skin patch of claim 1, wherein the metal layer, the ferrite layer, and the antenna layer are in direct contact with each other. 6. The functional skin patch of claim 1, wherein the antenna layer further comprises a metal dot layer printed on the ferrite layer for enhancing the surface area of the antenna layer. 7. The functional skin patch of claim 1, wherein the area of the first and the second surfaces of the functional skin patch are in a range of 20 cm2 to 500 cm2. 8. The functional skin patch of claim 1, further comprising: a heat collector unit having a first terminal thermally connected to the first surface and a second terminal thermally connected to the first terminal of the thermo harvester; anda thermally isolating layer,wherein the thermo harvester is embedded in the thermally isolating layer. 9. The functional skin patch of claim 1, wherein the functional unit comprises an RF-circuit connected to the antenna layer via a connection plug extended through the metal layer and the ferrite layer. 10. The functional skin patch of claim 1, wherein the functional unit comprises a sensor unit configured to sense a body health parameter including at least one of a body temperature, a body pulse frequency, an electrocardiogram recording, an electro encephalogram recording, a body function, a blood sugar value, a blood pressure, and a blood heparin value. 11. The functional skin patch of claim 1, wherein the thermo harvester comprises an energy storage unit configured to store harvested body energy. 12. The functional skin patch of claim 1, wherein the functional unit comprises a processing unit configured to process measurement data of the sensor unit and t... | Infineon Technologies AG | H01Q001/27 A61B005/00 H01L035/32 A61B005/0205 H01Q001/22 H01Q007/00 A61B005/01 A61B005/024 A61B005/021 A61B005/0476 A61B005/145
US09326818 | B2 | 2016-05-03 | Double loop lasso with single puller wire for bi-directional actuation | 2015-06-08 | A catheter has a distal assembly with at least one loop, if not two, with ring electrodes. The distal assembly has an elongated support member covered by an extruded form having two lumens, one sized for nonslip tight fit with the support member. A single continuous puller wire for bidirectional deflection is pre-bent into two long portions and a U-shape bend therebetween. The U-shape bend is anchored at a distal end of a deflectable section which is reinforced by at least one washer having at least two holes, each hole axially aligned with a respective lumen in the deflectable section. Each hole is centered with a lumen so that each puller wire portion therethrough is straight and subjected to tensile force only. A proximal end of the support member is flattened and serrated to provide a better bonding to the distal end of the deflectable section. | 1. A catheter comprising: an elongated tubular catheter body with at least one lumen extending there through;a deflectable intermediate section having proximal and distal ends and at least two off-axis lumens extending there through, the proximal end of the deflectable intermediate section being attached to a distal end of the catheter body; anda distal assembly distal of the deflectable intermediate section, the distal assembly comprising: a tubular structure having a generally straight proximal region attached to the intermediate section, a generally circular main region generally transverse and distal to the proximal region;an elongated support member extending through the generally circular main region of the distal assembly, the elongated support member having shape memory;a non-conductive cover over at least the main region of the distal assembly; andat least one electrode carried on the generally circular main region of the distal assembly;at least first and second washers at the distal end of the deflectable intermediate section, each of the first and second washers having at least two off-axis lumens, each axially aligned with a respective off-axis lumen of the deflectable intermediate section;a puller wire having two proximal portions and a mid-portion transverse to the two proximal portions, each proximal portion extending through a respective off-axis lumen of the deflectable intermediate section and a respective off-axis lumen of each of the first and second washers, the mid-portion being distal of the first and second washers and extending between the two off-axis lumens of the first and second washers; anda control handle comprising a deflection mechanism adapted to act on at least one proximal portion of the puller wire anchored in the control handle to deflect the deflectable intermediate section. 2. The catheter of claim 1, further comprising a connector section between the distal assembly and the deflectable intermediate section. 3. The catheter of claim 2, wherein a proximal end of the support member is received in the connector section. 4. The catheter of claim 2, wherein an interior of the connector section is filled with a bonding material. 5. The catheter of claim 4, wherein a proximal end of the support member is potted in the bonding material. 6. The catheter of claim 4, further comprising a location sensor potted in the bonding material. 7. The catheter of claim 6, wherein the location sensor is surrounded in heat shrink tubing. 8. The catheter of claim 1, wherein a proximal end of the support member is flattened. 9. The catheter of claim 1, wherein a proximal end of the support member has an uneven surface. 10. The catheter of claim 1, further comprising a first compression coil surrounding one of the two proximal puller wire portions and a second compression coil surrounding the other of the two proximal puller wire portions. 11. The catheter of claim 10, wherein the first compression coil has a first distal... | BIOSENSE WEBSTER (ISRAEL) LTD. | A61B005/042 A61B018/14 A61B005/00 A61M025/01 A61B018/00 A61B019/00 A61M025/06
US09326662 | B2 | 2016-05-03 | Probe and endoscope | 2013-11-04 | A probe is a probe insertable through a channel provided in an endoscope insertion portion. The probe is provided with a plurality of accommodation portions for accommodating a plurality of coil units having a predetermined length and a predetermined height, and a plurality of axial portions. Each accommodation portion includes a base portion having a length substantially equal to the predetermined length in a longitudinal direction of the probe, and two wall portions provided to stand upright at both ends of the base portion in the longitudinal direction at a height substantially equal to the predetermined height from the base portion, and the base portion and the wall portions are integrally formed. The plurality of axial portions extend from the wall portions in the longitudinal direction and connect between the plurality of accommodation portions. | 1. A probe insertable through a channel provided in an endoscope insertion portion or incorporated in the endoscope insertion portion, the probe comprising: a plurality of elongated accommodation members, each comprising: one or more accommodation portions, each of the one or more accommodation portions accommodating an electronic part, where each of the electronic parts having a predetermined length and a predetermined height, each of the one or more accommodation portions having a base portion having a length in a longitudinal direction substantially equal to the predetermined length of the electronic part and a wall portion provided at each of both ends of the base portion in the longitudinal direction having a height from the base portion substantially equal to the predetermined height of the electronic part, the base portion and the two wall portions being integrally formed so as to define a concave portion for engagingly accommodating the electronic part from a direction orthogonal to an axis of the accommodation member; andan axial portion that extends in the longitudinal direction from at least one of the wall portions to connect to at least one other of the plurality of accommodation members,wherein, in a state where the plurality of the accommodation members are connected together, a normal direction with respect to a bottom surface of each base portion that contacts each electronic part is the same, andwherein in a state where the electronic part is accommodated in the accommodation portion, a width of the base portion in a direction orthogonal to the axis of the accommodation member as the each electronic part is viewed from the normal direction is equal to or less than a width of the each electronic part in the direction orthogonal to the axis of the accommodation member. 2. The probe according to claim 1, wherein the wall portions have a tapered portion in the longitudinal direction of the probe. 3. The probe according to claim 1, wherein the probe is configured by connecting the plurality of accommodation members formed by being united with one or two or more axial portions. 4. The probe according to claim 3, wherein a connection in a connection portion is performed by causing a protruding portion provided at one end of one accommodation member to engage with a hole portion or a concave portion provided at one end of the other accommodation member connected to the one accommodation member. 5. The probe according to claim 4, wherein a normal direction with respect to a plane on which the protruding portion in the each accommodation member is formed or a plane on which the hole portion or the concave portion is formed is the same for the plurality of accommodation members. 6. The probe according to claim 1, wherein the plurality of accommodation members and the plurality of axial portions are each separate members. 7. The probe according to claim 1, wherein a contact surface of the base portion that comes into contact with the ... | OLYMPUS CORPORATION | A61B001/00 A61B008/12 A61B005/06 A61B008/00 A61B001/018 A61B001/04
US09326692 | B2 | 2016-05-03 | Blood pressure measurement device and blood pressure measurement method | 2010-11-16 | A blood pressure measurement device is characterized in that the said device is designated to arithmetically apply to the computed blood pressure value a correction value that is in accordance with the difference in atmospheric pressure between the position of the blood pressure transmitting part and the position of the heart of the subject wearing the blood pressure transmitting part as obtained from the sensor signal, on the basis of the internal pressure variation of the blood pressure transmitting part. | 1. A blood pressure measurement device to measure blood pressure of a subject, the blood pressure measurement device comprising: a blood pressure transmitting part;a pressure transmission channel configured to be selectively switched between an air pressure state in which atmospheric pressure is introduced to the pressure transmission channel and an internal pressure state in which an internal pressure of the blood pressure transmitting part is introduced to the pressure transmission channel;a sensor coupled to the pressure transmission channel to measure air pressure in the pressure transmission channel in the air pressure state at first and second positions and to measure an internal pressure variation in the pressure transmission channel in the internal pressure state at the second position;a counting unit configured to count a sensor signal of the sensor with a first frequency in the air pressure state, and to count the sensor signal of the sensor with a second frequency in the internal pressure state, the second frequency being higher than the first frequency;an instruction component configured to instruct the subject to move the blood pressure transmitting part to the first position adjacent to a heart of the subject, in order to measure the air pressure by the sensor at the first position;a measurement switching unit configured to switch a pressure state for measurement between the internal pressure state and the air pressure state;a blood pressure measuring part configured to compute a blood pressure value on the basis of the internal pressure variation obtained from the sensor signal counted in the counting unit; anda blood pressure value correction unit configured to arithmetically apply to the computed blood pressure value a correction value that is in accordance with the difference in the air pressure between the second position and the first position as obtained from the sensor signal of the sensor counted in the counting unit. 2. The blood pressure measurement device according to claim 1, wherein a precision of measuring the air pressure at the second position and the first position through the use of the sensor in the air pressure state is more precise than a precision of measuring the internal pressure variation through the use of the sensor in the internal pressure state. 3. A blood pressure measurement device to measure blood pressure of a subject, the blood pressure measurement device comprising: a sensor configured to measure air pressure at first and second positions while a state for measurement is in an air pressure state in which atmospheric pressure is measured by the sensor, and to measure an internal pressure variation at the second position while a state for measurement is in an internal pressure state in which an internal pressure of a blood pressure transmitting part is measured by the sensor;a counting unit configured to count a sensor signal of the sensor with a first frequency in the air pressure state, and to c... | Seiko Epson Corporation | A61B005/0215 A61B005/022
US09326719 | B2 | 2016-05-03 | Measurement of neural functionality using phase sensitive optical coherence reflectometry | 2012-12-19 | Optical methods, devices, and systems for noninvasively detecting transient surface displacements in a neuron are disclosed. Methods, devices, and systems provided may employ a phase-sensitive optical low coherence reflectometer. In addition, surface displacements due to action potential propagation in neural tissues may be detected in some embodiments using back-reflected light. According to some embodiments, exogenous chemicals or reflect ion coatings are not required. Transient neural surface displacement of less then 1 nm in amplitude and 1 ms in duration may be detected and may be generally coincident with action potential arrival to the optical measurement site. The systems and methods may be used for noninvasive detection of various neuropathies such as retinal neuropathies. They may also be useful in detecting the effects of various pharmacological agents. | 1. A device for detecting neural activity beam comprising: a. a beam combiner configured to combine light;b. a broadband source optically linked to the beam combiner;c. an aiming beam source optically linked to the beam combiner;d. a separator linked to the beam combiner and configured to separate light into a reference beam and a sample beam; reference optics linked separator, wherein the reference optics includes a phase modulator coupled to a rapid scanning delay line; sample optics operably linked to the separator;e. a recombiner optically linked to the reference optics and sample optics configured to combine light;f. an optical spectrometer optically connected to the recombines and configured to detect a plurality of spectral interference densities from the recombiner; andg. a processor processes detected spectral interference densities to produce transient surface displacement information of a neuron, wherein the processor transforms the detection of the transient surface displacement information into an indicator of neural activity in the neuron. 2. The device of claim 1, further comprising a neural stimulator. 3. The device of claim 2, wherein the neural stimulator is selected from the group consisting of an optical stimulator, an electrical stimulator, a chemical stimulator, and combinations thereof. 4. The device of claim 1, wherein the system has a surface displacement sensitivity of 0.1 nanometers. 5. The device of claim 1, wherein the system has a surface displacement temporal resolution of 0.1 microseconds. 6. The device of claim 1, wherein the transient surface displacement information includes a Δp that is calculated from a differential phase (Δφ) sensitivity of the detected spectral interference densities and a center wavelength of the broadband source (λo) using the equation: Δp=(λo/4π)Δφ. 7. The device of claim 1, wherein the transient surface displacement information includes a displacement of one or more neural cell surfaces. 8. The device of claim 1, wherein the processor transforms the detection of inappropriate neural activity into an indicator of a neuropathy in the neuron. 9. The device of claim 1, wherein the processor measures a transient retardation change (ΔR) by a measured intensity changes (ΔI) during an action potential propagation due to the transient retardation change using the equation: ΔR=½R (ΔI/I), where R is a retardation due to a resting birefringence and I is a resting intensity of light. 10. The device of claim 1, wherein the processor measures changes in a form-birefringence by the transient surface displacement information in response to a rapid repetitive stimuli. 11. The device of claim 1, wherein the rapid scanning delay line comprises a collimator coupled to a diffraction grating, and a lens coupled to the diffraction grating and a mirror on a galvanometer; wherein a separation between the diffraction grating and the lens is adjusted by the rapid scanning delay line to reduce a width of a... | Board of Regents, The University of Texas System | A61B006/00 A61B005/00 A61N001/00 A61N005/06 A61N001/36
US09351676 | B2 | 2016-05-31 | Electrochemical sensor, lancet, and bodily fluid measuring apparatus | 2011-10-28 | An electrochemical sensor includes a base plate provided with a concave part formed on one of surfaces thereof, a fluid channel formed so that a bottom part of the concave part and the other one of the surfaces of the base plate are communicated with each other, a plurality of electrodes formed on the concave part; a reagent fixed on the electrodes, a cover which covers the concave part, and an air channel which causes the inside and outside of the concave part to be communicated with each other. | 1. An electrochemical sensor, comprising: a base plate having a top surface and a bottom surface opposite of the top surface, the base plate includes a top concave part formed on the top surface;a fluid channel formed in the base plate, the fluid channel being a through-hole formed at a center of the top concave part in planar view, and formed in a direction that is orthogonal to the base plate to communicate a bottom part of the top concave part with the bottom surface opposite of the bottom part of the top concave part;a plurality of electrodes formed on the top concave part;a reagent fixed on the electrodes, having a ring-shape and disposed such that an inner edge of the ring-shape surrounds an opening of the through-hole;a cover which is directly provided on the base plate on which the plurality of electrodes are formed and is adapted to cover the top concave part; andan air channel disposed on the cover adapted to provide communication between an inside space existing between the top surface of the top concave part and the cover and an outside space above the cover. 2. The electrochemical sensor according to claim 1, wherein an outer edge shape in planar view of the top concave part is a triangle, a trapezoid, or a circle. 3. The electrochemical sensor according to claim 1, wherein the air channel includes at least one air hole formed in the cover. 4. The electrochemical sensor according to claim 3, wherein a planar view shape of the top concave part is formed in a triangle, and the air channel includes three air holes formed respectively at positions corresponding to apex portions of the triangle of the top concave part. 5. The electrochemical sensor according to claim 3, wherein a planar view shape of the top concave part is formed in a circle, and the air channel includes said at least one air hole formed on the cover and disposed so as to overlap with the fluid channel in a planar view state of the base plate. 6. The electrochemical sensor according to claim 1, further comprising: a pair of second concave parts formed around the top concave part,wherein the plurality of electrodes include:a first electrode pattern in which an electrode extending from the top concave part to one of the pair of second concave parts and an electrode lead part are formed integrally;a second electrode pattern which is insulated from the first electrode pattern, and in which an electrode extending from the top concave part to the other one of the pair of second concave parts and an electrode lead part are formed integrally;the cover has through-holes corresponding to positions of the second concave parts; andterminals for charging voltage to electrodes are inserted into the second concave parts through the through-holes of the cover. 7. The electrochemical sensor according to claim 1, wherein the bottom surface of the base plate is recessed inward. 8. The electrochemical sensor according to claim 1, wherein one of the plurality of electrodes is formed t... | ARKRAY, Inc. | A61B005/00 A61B005/1468 A61B005/15 A61B005/145 A61B005/1486 G01N027/327
US09352207 | B2 | 2016-05-31 | Action detection and activity classification | 2012-02-21 | Activities, actions and events during user performance of physical activity may be detected using various algorithms and templates. Templates may include an arrangement of one or more states that may identify particular event types and timing between events. Templates may be specific to a particular type of activity (e.g., types of sports, drills, events, etc.), user, terrain, time of day and the like. | 1. A system, comprising: an activity processing system having at least a first processor and a first memory, the first memory storing computer-executable instructions that, when executed by the at least a first processor, cause the activity processing system to: receive identification of an athletic activity to be performed;determine, based on the received identification of the athletic activity to be performed, a sensor subscription;transmit, to a sensor system, the determined sensor subscription;a sensor system including a plurality of sensors, at least a second processor and at least a second memory storing computer-executable instructions that, when executed, cause the sensor system to: receive, from the plurality of sensors, raw athletic activity data associated with a potential corresponding event during performance of the athletic activity;process the raw athletic activity data to smooth and filter the received athletic activity data;transmit, from the sensor system to the activity processing system, the processed athletic activity data, transmitting the processed athletic activity data to the activity processing system including controlling, by the sensor system, at least one of a type of data and an amount of data transmitted based on the determined sensor subscription;the activity processing system further including instructions that, when executed, cause the activity processing system to: compare the received processed athletic activity data to a plurality of physical action templates to determine whether a physical activity of at least one physical action template of the plurality of physical action templates has occurred, the comparing including:determining a plurality of required states associated with the physical activity of the at least one physical action template and forming a portion of the at least one physical action template;for each of the plurality of required states, determine a state-specific set of one or more constraints defining match characteristics of a corresponding required event,wherein a state-specific set of one or more constraints for a first state of the plurality of required states includes a relative constraint defining a required relationship between the first state and a second state of the plurality of required states,wherein the state-specific set of one or more constraints further includes a scoring distribution for determining a match score for the one or more constraints;determine, based on the comparison, the match score for the one or more constraints;determine, based on the match score, whether the physical activity has occurred; andtransmit a result of the determining whether the physical activity has occurred to one or more devices. 2. The system of claim 1, wherein the state-specific set of one or more constraints further includes a first parameter indicating whether a first characteristic of the corresponding event is to be scored. 3. The system of claim 2, wherein the state-specific set of... | NIKE, Inc. | A63F013/00 A63B071/06 A63B024/00 A61B005/00 A61B005/11 G06F019/00 G09B019/00
US09351660 | B2 | 2016-05-31 | Biomedical electrode configuration for suppressing movement artifact | 2012-04-02 | A biomedical sensor for detecting EMG signals may include a non-conducting fixed framework supporting a symmetrical arrangement of four electrode surfaces, configured as two signal detection contacts, each with a respective associated signal reference contact. The mechanical and electrical configuration act together to suppress movement artifact. | 1. A biomedical sensor for sensing phenomena of a muscle, comprising: a substrate; a plurality of conductive areas arranged to be spaced along a length of the substrate and configured as signal detection contacts each with a proximate, distinct respective associated signal reference contact, forming a plurality of signal detection/reference contact pairs spaced along the length, wherein each of the conductive areas is formed in the shape of at least one of: a disc attached to and coplanar with the substrate and a dome whose cross-sectional profile exhibits a concave curved shape attached to and projecting from the substrate and wherein each of the conductive areas is formed from a conductive plastic and wherein each one in the plurality of signal detection/reference contact pairs is aligned in a direction transversely to the length and in parallel with respect to each other so that each signal detection/reference contact pair produces movement induced artifacts in unison in response to relative movement between the plurality of signal detection/reference contact pairs and the muscle; and a differential sensing circuit that receives input from each of the signal detection/reference contact pairs and provides an output that is an arithmetic difference between the signals detected at detection contacts measured with respect to the reference contacts, the reference contacts being directly electrically coupled together and to a reference input of the differential sensing circuit, and the detection contacts being provided at separate inputs to the differential sensing circuit, wherein each signal detection/reference contact pair is configured to react in response to the relative movement, wherein at least one common electrical component caused by the relative movement is cancelled out by the differential sensing circuit, and wherein the at least one common electrical component caused by the relative movement includes the movement induced artifacts that are suppressed by the differential sensing circuit. 2. The biomedical sensor as in claim 1, wherein the substrate is rigid. 3. The biomedical sensor as in claim 1, further comprising: a framework coupled to the substrate, wherein the substrate is a body-directed electrically insulating bottom surface of the framework. 4. The biomedical sensor as in claim 1, wherein the plurality of signal detection/reference contact pairs are spaced apart so that the signal detection contact is aligned to be on the same side of each pair. 5. The biomedical sensor as in claim 1, wherein each of the conductive contact areas are formed of substantially identical material composition, size and shape. 6. The biomedical sensor as in claim 1, wherein the substrate is contoured in a concave shape to surround each of the conductive areas. 7. The biomedical sensor as in claim 1, further comprising: a framework coupled to the substrate, wherein the substrate is a body-directed electrically insulating bottom surface of the fra... | Altec, Inc. | A61B005/04 A61B005/0488 A61B005/0492 A61B005/00
US09339209 | B2 | 2016-05-17 | Body-worn monitor for measuring respiratory rate | 2010-04-19 | The invention provides a system for measuring respiratory rate (RR) from a patient. The system includes an impedance pneumography (IP) sensor, connected to at least two electrodes, and a processing system that receives and processes signals from the electrodes to measure an IP signal. A motion sensor (e.g. an accelerometer) measures at least one motion signal (e.g. an ACC waveform) describing movement of a portion of the patient's body to which it is attached. The processing system receives the IP and motion signals, and processes them to determine, respectfully, frequency-domain IP and motion spectra. Both spectra are then collectively processed to remove motion components from the IP spectrum and determine RR. For example, during the processing, an algorithm determines motion frequency components from the frequency-domain motion spectrum, and then using a digital filter removes these, or parameters calculated therefrom, from the IP spectrum. | 1. A method for measuring respiratory rate from a patient, comprising the following steps: (a) measuring a first impedance pneumography signal from the patient with an impedance pneumography sensor connected to at least two electrodes, the first impedance pneumography signal representing a time-dependent capacitance change in the patient's torso;(b) measuring a first motion signal with a motion sensor connected to the patient's torso, the first motion signal corresponding to breathing-induced motion, posture of the patient, and degree of motion of the patient; and(c) transmitting the first impedance pneumography signal and the first motion signal to a processing system comprising a microprocessor and collectively processing the first motion signal and the first impedance pneumography signal using the processing system to determine a set of coefficients associated with a digital adaptive filter, wherein the set of coefficients are determined as a function of both the first impedance pneumography signal and the first motion signal;(d) storing the set of coefficients in a computer memory;(e) measuring a second impedance pneumography signal from the patient with the impedance pneumography sensor, the second motion signal corresponding to breathing-induced motion, posture of the patient, and degree of motion of the patient;(f) measuring a second motion signal with the motion sensor, the second impedance pneumography signal representing a time-dependent capacitance change in the patient's torso;(g) transmitting the second impedance pneumography signal and the second motion signal to the processing system and analyzing the second motion signal using the processing system to determine a motion parameter;(h) when the motion parameter exceeds a pre-determined threshold, using the processing system to process the set of coefficients and the second impedance pneumography signal to determine a first processed impedance pneumography signal and analyze the first processed impedance pneumography signal to determine the respiratory rate;(i) when the motion parameter does not exceed the pre-determined threshold, using the processing system to digitally filter the second impedance pneumography signal with a bandpass filter, derivatize the digitally filtered IP waveform to provide a second processed impedance pneumography signal, and analyze the processed impedance pneumography signal using the processing system to determine the respiratory rate; and(j) displaying the respiration rate on a display screen operably connected to the processing system. 2. The method of claim 1, wherein the digital adaptive filter is calculated from an impulse response function. 3. The method of claim 2, wherein the impulse response function is calculated from one of a finite impulse response function and an autoregressive moving average model. 4. The method of claim 3, wherein an order of the digital adaptive filter calculated from the impulse response function is between 20 and 30. ... | SOTERA WIRELESS, INC. | A61B005/08 A61B005/05 A61B005/113 A61B005/0402 A61B005/1455 A61B005/00
US10085715 | B2 | 2018-10-02 | Ultrasonic scanning system | 2015-05-21 | An ultrasonic scanning system includes a host device, an ultrasonic probe, and at least two electrodes. The ultrasonic probe is electrically connected to the host device for performing an ultrasonic scanning operation toward a detecting region to generate a scanning signal and to transmit the scanning signal to the host device. The at least two electrodes are disposed on the ultrasonic probe and are connected to the host device for detecting a physiological signal of the detecting region and transmitting the physiological signal to the host device. | 1. An ultrasonic scanning system comprising: a host device;an ultrasonic probe electrically connected to the host device for performing an ultrasonic scanning operation toward a detecting region to generate a scanning signal and transmitting the scanning signal to the host device;a first sleeving member having a frame portion and an arm portion, the frame portion having an opening to expose an ultrasonic signal detecting surface of the ultrasonic probe, the arm portion extending from the frame portion toward the ultrasonic probe and being detachably connected to the ultrasonic probe, the frame portion further having a first frame structure and a second frame structure, the first frame structure being telescopically connected to the second frame structure to form the opening cooperatively with the second frame structure, and the first frame structure being slidable relative to the second frame structure for adjusting an area of the opening; andat least two electrodes attached to the frame portion and connected to the first frame structure and the second frame structure respectively, the at least two electrodes being connected to the host device for detecting a physiological signal of the detecting region and transmitting the physiological signal to the host device. 2. The ultrasonic scanning system of claim 1, wherein the arm portion has at least two elastic arms, and the at least two elastic arms are located at two sides of the ultrasonic probe to clamp the ultrasonic probe cooperatively. 3. The ultrasonic scanning system of claim 2, wherein each elastic arm has an attachment end detachably attached to the ultrasonic probe. 4. The ultrasonic scanning system of claim 1 further comprising: a signal transmitting device disposed on the arm portion and electrically connected to the at least two electrodes for connecting to the host device. 5. The ultrasonic scanning system of claim 4, wherein the signal transmitting device comprises: a connection port disposed on the arm portion and electrically connected to the at least two electrodes; anda signal transmitting cable electrically connected to the host device, the signal transmitting cable having a plug, the plug being inserted into the connection port to establish signal transmission between the host device and the at least two electrodes. 6. The ultrasonic scanning system of claim 1, wherein the first frame structure has at least one first arm portion and a first ear portion, the second frame structure has at least one second arm portion and a second ear portion, the first ear portion extends from the at least one first arm portion toward the at least one second arm portion, the second ear portion extends from the at least one second arm portion toward the at least one first arm portion, and the at least one first arm portion and the at least one second arm portion are movably disposed through the second ear portion and the first ear portion respectively so as to make the first frame structure s... | Qisda Corporation | A61B008/00 A61B005/053 A61B005/0488 A61B008/14 A61B005/0402 A61B008/08
US09974427 | B2 | 2018-05-22 | Handle remote control for use with bronchoscopy navigation system | 2014-11-14 | A surgical system includes an extended working channel, a magnetic field generator, a sensor, a workstation, a computer readable recording medium, and a remote control handle. The extended working channel is sized for insertion into the working channel of a bronchoscope and to receive one or more instruments therethrough. The sensor is associated with a distal end of the extended working channel for detecting a magnetic field generated by the magnetic field generator. The workstation includes a display and is operably connected to the sensor. The computer readable recording medium is associated with the workstation for storing a software program that enables association of the sensed magnetic field with a location of the sensor and that generates one or more images including one or more fields for presentation on the display. The remote control handle enables manipulation of the one or more fields generated by the software. | 1. A surgical system comprising: an extended working channel sized for insertion into a working channel of a bronchoscope and to receive one or more instruments therethrough, the extended working channel having a proximal handle;a magnetic field generator;a locatable guide extendable through the extended working channel having a sensor associated with a distal end of the extended working channel for sensing a magnetic field generated by the magnetic field generator;a workstation including a display and operably connected to the sensor;a non-transitory computer readable recording medium associated with the workstation, the computer readable recording medium storing a software program that enables association of the sensed magnetic field with a location of the sensor and that generates one or more images including one or more fields for presentation on the display; anda remote control handle having a distal portion directly releasably coupled to the proximal handle of the extended working channel, the remote control handle in communication with the workstation enabling manipulation of the one or more fields generated by the software and presented on the display. 2. The surgical system of claim 1, wherein the remote control handle includes a plurality of directional buttons to navigate between fields presented on the display by the software. 3. The surgical system of claim 1, wherein the remote control handle includes a selection control button enabling selection of one or more options presented in the one or more fields on the display by the software. 4. The surgical system of claim 1, wherein the remote control handle includes a joystick to navigate between the one or more fields presented on the display by the software. 5. The surgical system of claim 4, wherein the joystick is vertically depressible to enable selection of one or more options presented in the one or more fields on the display by the software. 6. The surgical system of claim 4, further comprising a hand selection switch enabling left or right handed operation orienting movements of the software with selection. 7. The surgical system of claim 1, wherein the remote control handle includes a trackball to navigate between the one or more fields presented on the display by the software. 8. The surgical system of claim 7, wherein the trackball is vertically depressible to enable selection of one or more options presented in the one or more fields on the display by the software. 9. The surgical system of claim 8, wherein the remote control handle includes an indicia of location on the trackball. 10. The surgical system of claim 1, wherein the remote control handle is in wireless communication with the workstation. 11. The surgical system of claim 1, wherein the remote control handle and the handle of the extended working channel can be secured together fixing the extended working channel relative to the sensor. 12. The surgical system of claim 11, wherein a portion of the hand... | COVIDIEN LP | A61B001/00 A61B001/018 A61B001/267 A61B005/06 A61B010/04 A61B034/20 A61B017/00 A61B034/00
US10912480 | B2 | 2021-02-09 | Sensor system and process for measuring electric activity of the brain, including electric field encephalography | 2014-06-20 | A sensor system and process for measuring electromagnetic activity of a brain are provided. The system and process employ a sensor assembly having a plurality of electrodes arranged in a closely spaced arrangement and a processor to determine a weighted average of the signals indicative of an electric field generated by electromagnetic activity of the brain. The system provides a medical body area network of a subject including one or more of the sensor assemblies and one or more additional sensors, which may be within a smartphone or other wearable device. | 1. A sensor assembly for measuring electromagnetic activity of a brain of a subject, comprising: a support plate having first and second sides;a plurality of electrodes mounted on and protruding from the first side of the support plate and supported for orientation toward and electrical contact with the subject's scalp, the plurality of electrodes arranged in a spaced array on the support plate, one of the plurality of electrodes comprising a reference electrode, each remaining electrode of the plurality of electrodes comprising an active electrode, the support plate configured to be attached to the subject with the plurality of electrodes oriented for electrical contact with the subject's scalp;a plurality of amplifiers mounted on the second side of the support plate, each amplifier of the plurality of amplifiers in electrical communication with the reference electrode and in electrical communication with an associated one of the active electrodes to receive signals from the reference electrode and the associated ones of the active electrodes indicative of electric field activity of the brain and to amplify potential differences between each associated one of the active electrodes and the reference electrode; anda microcontroller in communication with the plurality of amplifiers to receive signals from the plurality of amplifiers representative of the amplified potential differences between the reference electrode and the active electrodes, the microcontroller including a processor operative to determine, using the amplified potential difference signals, a tangential electric field component indicative of an electric field generated by electromagnetic activity of the brain;wherein the microcontroller is further operative to determine a weighted average from a sum of potential signals from each active electrode relative to the reference electrode multiplied by a component of a distance to the reference electrode, the sum divided by a further sum of the square of the distance components. 2. The sensor assembly of claim 1, wherein the microcontroller is further operative to determine a weighted average from a sum of potential signals from each electrode divided by a number of the plurality of electrodes. 3. The sensor assembly of claim 1, wherein the microcontroller is further operative to determine a weighted average from a sum of potential differences from each electrode relative to the reference electrode multiplied by a component of a distance to a center of the arrangement of the plurality of electrodes, the sum divided by a sum of the distance to the center of the arrangement to the fourth power. 4. The sensor assembly of claim 1, wherein the microcontroller is further operative to determine a higher order derivative of an electric potential. 5. The sensor assembly of claim 1, further comprising a transceiver for sending and receiving signals between the microcontroller and an external device. 6. The sensor assembly of claim 1, further c... | Northeastern University | A61B005/0478 A61B005/00 A61B005/04 A61B005/11
US20160120436 | A1 | 2016-05-05 | Psychological Evaluation and Methods of Use | 2016-01-12 | The psychological impact of entertainment material, visual objects, brands and advertising, commercial communication, the response to an individual presenting a message, or to an individual seeking public office can be assessed using methods employing measurements of SSVEP or SSVER phase increase and gaze tracking of subjects in a variety of ways. Various psychological states may be analyzed in order to accurately predict success and to enable early modification in development stages of products or communication. | 1-181. (canceled) 182. A method of quantitatively assessing at least one psychological response of at least one subject to at least one stimulus including the steps of: (a) presenting the at least one stimulus to the at least one subject, the at least one stimulus having a sequence of audio, visual and/or audiovisual features that occur as a function of time;(b) obtaining, during the presentation of the at least one stimulus, electroencephalogram (EEG) signals and/or magnetoencephalogram (MEG) signals from at least one pre-determined scalp site on the at least one subject;(c) calculating steady state visually evoked potential (SSVEP) phase increase and/or steady state visually evoked response (SSVER) phase increase from said EEG and/or MEG signals to obtain output signals representing pre-determined psychological states of the at least one subject occurring as a function of time;(d) optionally combining the output signals from a plurality of subjects to obtain pooled output signals; and(e) analyzing the output signals and/or pooled output signals to quantitatively assess the at least one subject's at least one psychological response to the at least one stimuli. 183. The method of claim 182 wherein the at least one stimulus is selected from the group consisting of entertainment material, commercial communication, at least one character and combinations thereof. 184. The method of claim 182, further including the step of presenting a semantic probe simultaneously with step 182 (a). 185. The method of claim 182, further including the following steps: presenting at least one stimulus different from the at least one stimulus in (a) to the at least one subject; and waiting for a pre-determined period of time. 186. The method of claim 184 wherein reference stimulus is presented to the at least one subject. 187. The method of claim 186 further including the steps of claim 182 (b)-(e), except that the EEG and/or MEG signals obtained, and the quantitative assessment analyzed is for the reference stimulus. 188. The method of claim 182 wherein the at least one stimulus has a set of pre-determined parameters. 189. The method of claim 187, further including the steps of tracking the gaze position of at least one eye of the at least one subject during the claim 182 (a) step; and determining the difference in SSVEP and/or SSVER phase from step 182 (c) and the SSVEP and/or SSVER phase calculated for the reference stimulus according to claim 187. 190. The method of claim 182 wherein said at least one pre-determined scalp site is selected to quantitatively assess at least one psychological state selected from the group consisting of visual attention to detail, visual attention to global features, multi-modal attention to detail, desirability, multi-modal attention to global features, emotional intensity, long-term memory encoding, attraction-repulsion, engagement, likeability, behavioral intent, and combinations thereof. 191. The method of claim 182 inclu... | Neuro-Insight Pty. Ltd. | A61B005/0484 A61B005/00 A61B005/04
US20160135709 | A1 | 2016-05-19 | METHOD AND DEVICE FOR A BIOELECTRIC IMPEDANCE ANALYSIS (BIA) OF THE BODY OF A PERSON | 2014-06-24 | To perform a bioelectrical impedance measurement (BIA) of the body of a person several current pulse giving input electrodes and several voltage drop output electrodes need to be connected with the surface of the body. The voltage drop between the output electrodes of individual body sections is detected by the measurement. The measured values obtained are converted in combination with for example statistical data to a body structure and visceral risk analysis. In order for a BIA belly, BIA belly hand and BIA-hand-hand, when four hand electrodes and four abdominal electrodes are combined, a person can do a measurement in a simple manner, even without the often considerable effort of wiring the body, according to the invention the measuring template in the form of a bicycle handlebar (2). On its underside in the two middle template sectors (6, 7) it contains four abdomen electrodes and on its upper side (4) in the handles (8, 9) each contains two hand electrodes (12, 12′). The person holds it in a given arm position in a physiological simple manner with reproducible accuracy in a defined position and accordingly pressed against the abdomen. | 1. Method for the bioelectrical impedance measurement (BIA) of the body of a person, wherein a plurality of a current pulse giving input electrodes and a plurality of the voltage drop measured sensitive output electrodes with the surface of the body brought into contact and between these electrodes the bioelectrical impedance of the body metrological recorded and the measured values obtained with stored, for example, statistical data converted to a body structure—and visceral risk analysis, characterized in that via a constructed like the form of bicycle handlebars, with a measurement module (1) connected measuring template (2), which in the middle stencil portions (6, 7) respectively to the left and right have an input electrode (13) and an output electrode (14) and at the handles (8, 9) of the outer stencil portions (15, 16) respectively left and right, an input electrode (12) and an output electrode (12′), is performed by the person self-sufficient in a simple manner with reproducible accuracy simultaneously a BIA-belly, a BIA abdominal hand and a BIA-hand-hand, wherefore the person holds the measurement template (2) at the handles (8, 9)—with thumb and index finger contact to the local electrodes (12, 12′) and their middle part of the template (6, 7) with the electrodes (13, 14) in a defined position held against the stomach. 2. The method of claim 1, characterized in a way to secure the defined position during the measurement the measuring template (2) having a predetermined arm position and is aligned with its center at the navel (21). 3. The method of claim 1 or 2, characterized in a way that the measurement is carried out with a frequency of 50 kHz AC pulses and without additional mounting or wiring on the body and the measurements obtained with measuring module (1) with an integrated computer or calculating unit are used to generate a body structure and visceral risk analysis. 4. The method according to claim 3, characterized in a way that the measurement be carried out with a further and/or several measuring frequencies between 1-1000 kHz AC pulses. 5. A device for performing the method according to one or more of claims 1 to 4, characterized by: one with a measurement module (1) connected, measuring template in the shape of a bike handlebar (2), which's band-shaped base body (3) is ergonomically shaped in a way that the measuring template (2) with it's at the outer template sections (15, 16) arranged handgrip (8, 9) can be held in only a predetermined arm position suitable to the stomach and pressed against it;in the electrodes (12, 12′) arranged in the handgrip (8, 9), while holding the measuring template (2′) each thumb inevitably gets in contact to the output electrode (12′) and each index finger to the input electrode (12);on the underside (5) of the measuring template (2) in the middle part of the template (6, 7) left and right from the middle part of the template arranged input and output electrodes (13, 14). 6. The method o... |  | A61B005/053 A61B005/00
US09351685 | B2 | 2016-05-31 | Stacked adhesive optical sensor | 2014-01-21 | An optical sensor, having a cover layer, an emitter disposed on a first side of the cover, a detector disposed on the first side of said cover, and a plurality of stacked independent adhesive layers disposed on the same first side of the cover, wherein the top most exposed adhesive layer is attached to a patient's skin. Thus, when the sensor is removed to perform a site check of the tissue location, one of the adhesive layers may also be removed and discarded, exposing a fresh adhesive surface below for re-attachment to a patient's skin. The independent pieces of the adhesive layers can he serially used to extend the useful life of the product. | 1. A sensor comprising: a base layer configured to attach to a tissue location on a patient;an emitter coupled to the base layer and positioned in a first aperture formed in the base layer;a detector coupled to the base layer and positioned in a second aperture formed in the base layer; anda plurality of adhesive layers coupled to the base layer, wherein the adhesive layers comprise respective cutouts, and wherein the respective cutouts are configured to be aligned with one another when the plurality of adhesive layers are coupled to the base layer, and wherein the cutout of each adhesive layer of the plurality of stacked adhesive layers comprises a ring shape that surrounds both the emitter and the detector such that there is no adhesive present in the direct optical path between the emitter and the detector. 2. The sensor of claim 1, wherein the base layer comprises a major axis and a minor axis, and wherein the emitter and detector are aligned along the major axis. 3. The sensor of claim 2, comprising a cable extending from the base layer and coupled to the emitter and the detector, wherein the cable extends from the major axis. 4. The sensor of claim 2, wherein the plurality of adhesive layers comprise respective release tabs, and wherein the release tabs are positioned at a location along the major axis. 5. The sensor of claim 2, wherein the respective cutouts comprise corresponding major axes aligned with the major axis of the base layer. 6. The sensor of claim 1, wherein the sensor comprises a pulse oximetry sensor. 7. The sensor of claim 1, wherein the sensor comprises a forehead sensor. 8. The sensor of claim 1, wherein the sensor comprises a plurality of emitters. 9. The sensor of claim 1, wherein the sensor comprises a plurality of detectors. 10. A sensor comprising: a conformable layer comprising a plurality of apertures;a plurality of optical components associated with the sensor and positioned in the respective apertures; anda plurality of stacked independent nonmetallic adhesive layers disposed on the conformable layer, wherein the plurality of adhesive layers comprise a nonadhesive surface and an adhesive surface, and wherein the adhesive surface of at least one adhesive layer is in direct contact with the nonadhesive surface of another adhesive layer and wherein the plurality of adhesive layers comprise respective cutouts sized and shaped to expose a ring shape region on the conformable layer larger than an area corresponding to the optical components and the plurality of apertures such that there is no adhesive present in the direct optical path between the optical components. 11. The sensor of claim 10, wherein the optical components are covered by a compliant material. 12. The sensor of claim 11, wherein the compliant material is an optically transparent material. 13. The sensor of claim 10, wherein the conformable layer comprises a foam. 14. The sensor of claim 10, wherein a perimeter of the cutouts comprises an ... | Covidien LP | A61B005/1455 A61B005/00 A61B005/0205 A61B005/024 G01N021/49 G01N021/31
US09349303 | B2 | 2016-05-24 | Automated caregiving device with prompting based on caregiver progress | 2014-05-12 | A device for assisting a caregiver in delivering therapy to a patient, the device comprising a user interface configured to deliver prompts to a caregiver to assist the caregiver in delivering therapy to a patient; at least one sensor configured to detect the caregiver's progress in delivering the therapy, wherein the sensor is other than an electrode in an electrical contact with the body; a memory in which a plurality of different prompts are stored; a processor configured to determine which of the different prompts should be selected for delivery based on the progress detected by the sensor. | 1. A cardiac resuscitation system for assisting a caregiver in delivering electrical resuscitation therapy to a patient using disposable therapy electrodes that are stored in packaging and removed and applied to the skin of the patient's thorax, the system comprising a user interface configured to deliver prompts to a caregiver to assist the caregiver in delivering electrical resuscitation therapy to a patient;at least one photoelectric sensor configured to detect the caregiver's progress in removing a therapy electrode from its packaging and applying it to the patient;a memory in which a plurality of different prompts are stored;a processor configured to determine which of the different prompts should be selected for delivery based on the progress detected by the sensor. 2. The system of claim 1 wherein there are a plurality of photosensors configured to detect the caregiver's progress in delivering the therapy, including at least one photosensor on each electrode. 3. The system of claim 1 wherein the processor is configured to vary the time at which prompts are delivered based on the progress detected by the sensor. 4. The system of claim 1 further comprising one or more additional sensors in electrical contact with the body and configured to detect the caregiver's progress in delivering the therapy. 5. The system of claim 2 wherein the photosensor on each electrode comprises a photoemitter/photosensor. 6. The system of claim 5 in which a reflective surface is provided on the interior of the electrode packaging in the vicinity of each of the photosensors when each electrode is in the position in which it is stored in the packaging, wherein the reflective surface reflects light emitted by the photoemitter and the reflected light is sensed by the photosensor. 7. The system of claim 6 in which the processor determines that the packaging has been opened based on an increase in sensed light. 8. The system of claim 6 in which the processor determines that an electrode has been peeled from a liner to which it was attached in the packaging by recognizing a reduction in reflected photoemitter light sensed at the photosensor. 9. The system of claim 5 in which the processor determines that the packaging has been opened based on an increase in sensed light. 10. The system of claim 5 in which the processor determines that an electrode is still in an unopened package by determining that the light level sensed by the photosensor on the electrode is below a threshold. 11. The system of claim 1 in which the photosensor and processor are configured to detect whether the electrode has been applied to clothing rather than to skin. 12. The system of claim 11 in which detection of an intermediate level of reflected photoemitter light is used by the processor to determine that the electrode has been applied to a surface. 13. The system of claim 12 in which detection of whether the surface to which the electrode is applied is skin or clothing is accomplish... | ZOLL Medical Corporation | A61N001/38 G09B023/28 A61B005/00 A61N001/39 G09B005/06
US09327119 | B2 | 2016-05-03 | Electrostimulation system | 2012-07-27 | An electrostimulation system of the invention includes a signal generator, a docking assembly, and at least one transmission assembly. The signal generator is configured to produce at least one signal. The docking assembly includes a main housing; at least one mechanical retaining element operatively connected to the main housing that mechanically secures the signal generator to the docking assembly; and, at least one electrical docking element mechanically connected to the main housing that electrically connects the docking assembly to the signal generator. The at least one transmission assembly is adapted to fit on a portion of a user's anatomy, including at least one transmission assembly connection element (TACE) adapted to transmit a signal from the docking assembly to the user. The signal generator is electrically and mechanically detachably connected to the docking assembly during use. | 1. An electrostimulation system comprising: a) a signal generator configured to produce at least one signal;b) a docking assembly comprising: i. a main housing;ii. at least one mechanical retaining element operatively connected to said main housing to mechanically secure said signal generator to said docking assembly; and,iii. at least one electrical docking element mechanically connected to said main housing for electrical connection of said docking assembly to said signal generator; andc) at least one transmission assembly adapted to fit on a portion of a user's anatomy, comprising at least one transmission assembly connection element (TACE) adapted to transmit a signal from said docking assembly to the user,wherein said signal generator is electrically and mechanically detachably connected to said docking assembly during use, and, wherein i. said at least one transmission assembly comprises a plurality of signal transmission elements, said plurality of signal transmission elements including a set of signal transmission elements being at selected anatomic locations at an inner surface of said transmission assembly and at least one other signal transmission element spaced from said anatomic locations;ii. said TACE comprises conductive fabric or other conductive material positioned on said transmission assembly and electrically connected to said signal transmission elements; and,iii. said docking assembly further comprises at least one docking assembly connection element (DACE) electrically connectable to said TACE so as to provide an electrical connection of said at least one docking assembly to said TACE. 2. The electrostimulation system of claim 1 wherein, said at least one transmission assembly comprises a conductive glove including said plurality of signal transmission elements, said plurality of signal transmission elements including said set of signal transmission elements being at said selected anatomic locations at an inner surface of said conductive glove, said anatomic locations being fingertip locations, said plurality of signal transmission elements including and at least one other signal transmission element spaced from said fingertip locations. 3. The electrostimulation system of claim 2 wherein said set of signal transmission elements comprises at least one active signal transmission element and at least one other signal transmission element comprises a dispersive signal transmission element. 4. The electrostimulation system of claim 3 wherein said at least one active signal transmission element is positioned on the fingertips and said at least one other dispersive transmission element is positioned on a forearm portion of the conductive glove. 5. The electrostimulation system of claim 1 wherein, said at least one transmission assembly comprises a knee orthotic device including said plurality of signal transmission elements, said plurality of signal transmission elements including at least one signal transmission element being... | VISION QUEST INDUSTRIES INCORPORATED | A61B005/04 A61N001/00 A61N001/36 A61N001/04 A61F005/01 A61N001/372 A61N001/20 A61N001/22
US20160143542 | A1 | 2016-05-26 | Minimally Invasive Optical Photoacoustic Endoscopy with a Single Waveguide for Light and Sound | 2015-11-23 | An endoscopic device for photoacoustic imaging, including a multimode optical fiber having a distal end and a proximal end, a light source to provide a light beam to the proximal end of the multimode optical fiber, a transducer to capture acoustic waves that are emitted from the proximal end of the multimode optical fiber, and a processing device to generate a photoacoustic image based on data from the captured acoustic waves captured by the transducer, wherein the distal end of the multimode optical fiber is configured to be inserted into a sample, the sample generating the acoustic waves by a photoacoustic effect. | 1. An endoscopic device for photoacoustic imaging, comprising: a multimode waveguide having a distal end and a proximal end;a light source to provide a light beam to the proximal end of the multimode waveguide;a transducer to capture acoustic radiation that is emitted from the proximal end of the multimode waveguide; anda processing device to generate a photoacoustic image based on data from the captured acoustic radiation captured by the transducer,wherein the distal end of the multimode waveguide is configured to be inserted into a sample, the sample generating the acoustic radiation by a photoacoustic effect. 2. The endoscopic device according to claim 1, further comprising: an acousto-optic coupler to direct the light beam from the light source to the proximal end of the multimode waveguide and also configured to direct the acoustic radiation received from the proximal end of the multimode waveguide to the transducer. 3. The endoscopic device according to claim 1, wherein the light source includes: a laser light to generate the light beam, anda laser optics to focus the light beam generated by the laser light onto the proximal end of the multimode waveguide. 4. The endoscopic device according to claim 1, wherein the multimode waveguide includes: a fluid core, anda cladding forming a layer around the fluid core,wherein the fluid core is configured to guide the acoustic radiation from the distal end to the proximal end, and the cladding is configured to guide light of the light beam from the proximal end to the distal end of the multimode waveguide. 5. The endoscopic device according to claim 1, wherein the multimode waveguide includes: a fluid core, anda cladding forming a layer around the fluid core,wherein the fluid core is configured to guide the acoustic radiation from the distal end to the proximal end, and the fluid core is also configured to guide light of the light beam from the proximal end to the distal end of the multimode waveguide. 6. The endoscopic device according to claim 1, wherein the multimode waveguide includes: a fluid core arranged at the distal end of the multimode waveguide,a fiber-optic hydrophone arranged at the proximal end of the multimode waveguide and extending throughout the multimode waveguide but for the distal end, anda cladding forming a layer around the fluid core and the fiber-optic hydrophone. 7. The endoscopic device according to claim 1, wherein the transducer is also configured to capture fluorescent radiation that is emitted from the proximal end of the multimode waveguide, in addition to the acoustic radiation, the fluorescent radiation generated by the sample. 8. A method to generate a photoacoustic image from a sample with a multimode waveguide, the multimode waveguide penetrating into the sample such that a distal end of the multimode waveguide faces an area of the sample under test inside the sample, the method comprising the steps of: radiating a proximal end of the multimode waveguide wit... |  | A61B005/00 A61B001/06 A61B001/07 A61B001/00
US20160129343 | A1 | 2016-05-12 | REHABILITATIVE POSTURE AND GESTURE RECOGNITION | 2014-06-12 | A kinetic rehabilitation system comprising: a kinetic sensor comprising a motion-sensing camera; and a computing device comprising: (a) a non-transient memory comprising a stored set of values of rehabilitative gestures each defined by a time series of spatial relations between a plurality of theoretical body joints, and wherein each time series comprises: initial spatial relations, mid-gesture spatial relations and final spatial relations, and (b) a hardware processor configured to continuously receive a recorded time series of frames from said motion-sensing camera, wherein each frame comprises a three-dimensional position of each of a plurality of body joints of a patient, wherein said hardware processor is further configured to compare, in real time, at least a portion of the recorded time series of frames with the time series of spatial relations, to detect a rehabilitative gesture performed by the patient. | 1. A kinetic rehabilitation system comprising: a kinetic sensor comprising a motion-sensing camera to detect a patient's gestures; anda computing device comprising: (a) a non-transient memory comprising a stored set of values of rehabilitative gestures each defined by a time series of spatial relations between a plurality of theoretical body joints, wherein each rehabilitative gesture comprises gesture phases including at least an initial gesture phase, a mid-gesture phase and a final gesture phase, and wherein each time series of spatial relations for a rehabilitative gesture comprises: initial spatial relations, mid-gesture spatial relations and final spatial relations, and(b) a hardware processor configured to: (i) automatically translate a therapy plan provided for the patient to a video game level, (ii) continuously receive a recorded time series of frames from said motion-sensing camera during a game play, wherein each frame comprises a three-dimensional position of each of a plurality of body joints of a patient, (iii) compare, in real time, at least a portion of the recorded time series of frames with the time series of spatial relations, to detect an initial gesture phase, a mid-gesture phase and a final gesture phase of a rehabilitative gesture performed by the patient, and (iv) provide real time feedback to the patient during the game play regarding the performed rehabilitative gesture. 2. The system according to claim 1, wherein each of said time series of spatial relations further comprises one or more range values for each of at least one of said spatial relations. 3. The system according to claim 1, wherein said time series of spatial relations further comprises one or more range values for the transition time between each of at least one of said spatial relations. 4. The system according to claim 1, wherein said spatial relations each comprise angles between vectors formed in a three-dimensional space by said theoretical body joints. 5. The system according to claim 1, wherein said spatial relations comprise distances in a three-dimensional space between said theoretical body joints. 6. The system according to claim 1, wherein said motion-sensing camera is configured to yield said recorded time series of frames at a frame rate of 20 frames per second or more. 7. The system according to claim 1, wherein said motion-sensing camera is configured to yield said recorded time series of frames at a frame rate of 30 frames per second or more. 8. The system according to claim 1, wherein said motion-sensing camera is configured to yield said recorded time series of frames at a frame rate of 40 frames per second or more. 9. The system according to claim 1, wherein said hardware processor is further configured to convert said three-dimensional position in said recorded time series of frames to angles between vectors formed in a three-dimensional space by said theoretical body joints. 10. The system according to claim 1, wherein said ... |  | A63F013/21 A61B005/00 A61B005/11 G06K009/00 H04N005/77
US20160128599 | A1 | 2016-05-12 | HANDHELD EMG STIMULATOR DEVICE WITH ADJUSTABLE SHAFT LENGTH | 2016-01-14 | A nerve stimulating device with an adjustable length is shown and described. The device includes a handle and an elongated shaft with a stimulating electrode. The stimulating electrode can be selectively positioned to a plurality of distances relative to the handle. The device allows nerves in nerve regions at multiple positions relative to the surface of the patient's body to be stimulated with a single device. | 1. A handheld nerve stimulating device, comprising: a handle having a proximal end and a distal end spaced apart along a handle length axis, the handle further having a hollow interior;an elongated shaft having a proximal end and a distal end spaced apart along an elongated shaft length axis parallel to the handle axis, the elongated shaft further having an externally insulated portion and an exposed electrode at the distal end of the elongated shaft, wherein the handheld nerve stimulating device has an adjustable length defined between the proximal end of the handle and the distal end of the elongated shaft, the exposed electrode has a fixed length along the elongated shaft length axis, and the proximal end of the elongated shaft is movable within the hollow interior of the handle to adjust the length of the handheld nerve stimulating device such that the exposed electrode is selectively and securely repositionable at a plurality of positions along the handle axis with respect to the distal end of the handle, and the proximal end of the elongated shaft is not extendable from the proximal end of the handle. 2. The device of claim 1, further comprising a lock that engages a portion of the handle, wherein the lock has a tightened configuration and a loosened configuration, such that when the lock is in the tightened configuration, the elongated shaft is immovable with respect to the handle, and when the lock is in the loosened configuration, the shaft is movable with respect to the handle along the handle axis. 3. The device of claim 1, wherein the electrode is a ball-tip electrode. 4. The device of claim 1, wherein the elongated shaft is selectively movable to a plurality of indexed positions, and each indexed position corresponds to a fixed distance between the electrode and the handle. 5. The device of claim 4, wherein when the elongated shaft moves from one indexed position to another indexed position, an indexed movement sound is emitted. 6. The device of claim 1, further comprising a lead wire, wherein the lead wire extends from the proximal end of the handle, the distal end of the handle is located between the electrode and the proximal end of the handle, and the electrode is in electrical communication with the lead wire. 7. The device of claim 1, wherein the elongated shaft has a proximal end and a distal end and includes a shaft stop surface between the proximal end of the shaft and the distal end of the shaft, the handle has a proximal end and a distal end and includes a handle stop surface between the proximal end of the handle and the distal end of the handle, such that the engagement of the shaft stop surface and the handle stop surface retains the shaft to the handle. 8. The device of claim 1, wherein the shaft is selectively detachable from the handle. 9. The device of claim 1, wherein the handle is angled. 10. A handheld nerve stimulating kit, comprising the hand held nerve stimulating device of claim 1 and at least one re... |  | A61B005/0492 A61B005/00 A61B005/04
US09986932 | B2 | 2018-06-05 | QT interval determination methods and related devices | 2014-11-17 | Described herein is a system and method of automatically monitoring QT intervals in a patient based on one or more EKG signals received from attached monitoring devices. Each EKG signal is analyzed to detect attributes of the first and second EKG signals, including QRS onset information, QRS peak information, and T-wave offset information. A QT interval is calculated based on QRS onset information derived from the first EKG signal and T-wave offset information derived from the second EKG signal. The calculated QT interval is compared to thresholds to detect elongation of the QT interval and an alert is generated in response to a detected elongated QT interval. | 1. A monitoring system for measuring QT intervals in patients, the monitoring system comprising: a first monitoring device adhered to a patient at a first location that monitors a first electrocardiogram (EKG) signal associated with the patient;a second monitoring device adhered to a patient at a second location different than the first location that monitors a second EKG signal associated with the patient, and;a QT interval measurement module configured to receive the first EKG signal from the first monitoring device and the second EKG signal from the second monitoring device, wherein the QT interval measurement module: identifies first QRS complex attributes based only on the first EKG signal, wherein identified first QRS complex attributes include onset of the QRS complex (iRon_a) and QRS complex peak (IRpk_a);identifies second QRS complex attributes based only on the second EKG signal, wherein identified second QRS complex attributes include QRS complex peak (IRpk_b) and T-wave offset (iToff_b);selects qualified beat segments comprised of a plurality of beats suitable for QT interval measurement within the received first and second EKG signals;measures QT intervals within the qualified beat segments based on the identified first and second QRS complex attributes, based on the equation iQT=(iRpk_a−iRon_a)+(iToff_b-Rpk_b), wherein iQT is the QT interval; anddetecting a long QT interval based on a comparison of the measured QT interval to a threshold value. 2. The monitoring system of claim 1, wherein locating the T-wave offset for each beat includes utilizing and identified heart rate of the patient and a look-up table to estimate a T-wave offset window in which to search for the T-wave offset. 3. The monitoring system of claim 2, wherein minimum and maximum values within the defined T-wave offset window are identified and utilized to detect the T-wave offset with respect to each beat. 4. The monitoring system of claim 1, wherein the QT interval measurement module generates an alert in response to detected long QT intervals. 5. The monitoring system of claim 4, wherein the QT interval measurement module provides generated alerts and the received first EKG signal to a remote monitoring center. 6. The monitoring system of claim 1, wherein both the first monitoring device and the second monitoring device receiving timing information from an external source, wherein first and second EKG signals include timestamps based on the received timing information. 7. The monitoring system of claim 6, wherein the first EKG signal and the second EKG signal are aligned based on timestamps associated with both EKG signals. 8. The monitoring system of claim 1, wherein qualified beats within the first EKG signal are selected and a search is conducted within the second EKG signal to locate a qualified segment of beats that can be aligned with the qualified beats within the first EKG signal. 9. The monitoring system of claim of 8, wherein alignment between q... | Medtronic Monitoring, Inc. | A61B005/04 A61B005/0472 A61B005/00 A61B005/0245 A61B005/0452
US10327978 | B2 | 2019-06-25 | Therapeutic undergarments for the treatment of functional gastrointestinal disorders including irritable bowel syndrome | 2015-09-25 | An apparatus, system, and method for treating symptoms of IBS. The apparatus includes a base material comprising a material that fits to the body of a user, an elastic portion capable of applying compression across the abdomen of a user, and a tension adjusting mechanism for adjusting the compression applied to the abdomen in order to treat the symptoms of IBS. A system for treating IBS may include a processor configured to receive a first set of data from biosensors configured to measure contractions in the bowel, receive a second set of data from a galvanic skin response sensor configured to measure electrical conductance of the skin; and, analyze patterns between the first and second sets of data. | 1. Apparatus for treating symptoms of Irritable Bowel Syndrome (IBS), comprising: a base material comprising a material that fits to the body of a user;an elastic portion capable of applying compression across the abdomen of a user; anda tension adjusting mechanism for adjusting the compression applied to the abdomen in order to treat the symptoms of IBS, wherein the tension adjusting mechanism comprises at least one strap having a first end extending from a first position on the base material and a second end that releasably fastens at a second position on the base material such that the strap extends around at least a portion of the base material to adjust the compression to apply a targeted pressure to a sigmoid and descending colon of the user. 2. The apparatus of claim 1, wherein the apparatus comprises an undergarment. 3. The apparatus of claim 1, wherein the tension adjusting mechanism comprises multiple connector strips attached to the base material and corresponding receiving strips that enable the connector strips to be fastened in an adjustable manner to adjust the compression applied to the abdomen. 4. The apparatus of claim 1, wherein the tension adjusting mechanism comprises an elastic strap attached to the base material at an angle such that the elastic strap wraps from a first position around the back of the user to a second position located higher than the first position. 5. The apparatus of claim 1, further comprising: a pouch configured to be positioned over the abdomen of the user, wherein the tension adjustment mechanism adjusts the compression of the pouch against the user. 6. The apparatus of claim 5, wherein the pouch is configured to be positioned over a sigmoid colon of the user. 7. The apparatus of claim 5, further comprising: multiple straps extending from the pouch, wherein the tension adjusting mechanism connects to the multiple straps to adjust the compression of the pouch to the abdomen of the user. 8. The apparatus of claim 7, wherein the tension adjusting mechanism comprises a crank attached to the base material at a side opposite the pouch. 9. The apparatus of claim 5, wherein the pouch is configured to receive at least one selected from a group consisting of a shaped removable insert, an inflatable bladder, a heat pack, a cold pack, an electro-stimulator, a biosensor that measures physiological activity, and a wireless transmitter. 10. The apparatus of claim 1, wherein the tension adjusting mechanism comprises multiple plastic rings or metal clasps that enable a hook to fasten at multiple locations in order to adjust the compression. 11. The apparatus of claim 1, wherein the apparatus comprises three horizontal sections, an upper and lower section comprising the base material, and a middle section comprising a material that is less elastic than the base material. 12. The apparatus of claim 11, wherein the tension adjusting mechanism comprises a cinch. 13. The apparatus of claim 1, further comprising... | COLOWRAP, LLC. | A61B005/00 A61B005/11 A61F005/02 A61F005/03 A61F005/34 A61F007/00 A61F007/02 A61H007/00 A61N001/36 A61B005/053
US09333087 | B2 | 2016-05-10 | Herniated disc repair | 2014-12-22 | Systems for minimally invasive disc augmentation include an anulus augmentation component and a nucleus augmentation component. Both are suited for minimally invasive deployment. The nucleus augmentation component restores disc height and/or replaces missing nucleus pulposus. The anulus augmentation component shields weakened regions of the anulus fibrosis and/or resists escape of natural nucleus pulposus and/or the augmentation component. Methods and deployment devices are also disclosed. Method of supporting and augmenting a nucleus pulposus by inserting a flexible biocompatible material into the disc space using an anchoring means are also provided. | 1. A method of facilitating treatment of an intervertebral disc in a spine comprising: introducing a distal end of an implant delivery device beyond at least a first outer layer of an anulus of an intervertebral disc, the distal end configured to deliver an implant within the intervertebral disc;delivering an implant in the intervertebral disc using the implant delivery device,wherein the implant is configured to be held along an innermost layer of the anulus in a nuclear space of the intervertebral disc;inserting at least one fixation device into a vertebral body surrounding the intervertebral disc;connecting the at least one fixation device to the implant; andremoving the implant delivery device. 2. The method of claim 1, wherein the at least one fixation device comprises a suture. 3. The method of claim 2, wherein the suture is resorbable. 4. The method of claim 1, wherein the at least one fixation device comprises a woven or web-like material. 5. The method of claim 1, wherein the at least one fixation device comprises a connection member that is knotted. 6. The method of claim 1, wherein the at least one fixation device comprises a resorbable material. 7. The method of claim 1, wherein the at least one fixation device comprises at least one connection member and at least one anchor portion. 8. The method of claim 7, wherein the at least one connection member comprises a suture or wire. 9. The method of claim 7, wherein the at least one anchor portion comprises at least one barb. 10. The method of claim 1, wherein the implant is cap-shaped, bar-shaped, plate-shaped, or bead-shaped. 11. The method of claim 1, wherein the at least one fixation device is affixed to said vertebral body with a cement or surgical adhesive. 12. The method of claim 1, wherein the at least one fixation device comprises a screw anchor, and wherein said anchor is screwed into the vertebral body. 13. The method of claim 1, wherein the implant delivery device carries said implant. 14. The method of claim 1, wherein the implant delivery device comprises a handheld grasping tool. 15. The method of claim 1, wherein the implant delivery device comprises a delivery tube. 16. A method of facilitating treatment of an intervertebral disc in a spine comprising: providing an implant delivery system comprising an implant delivery device, at least one fixation device, and an implant, the implant delivery system configured for: (i) introducing a distal end of the implant delivery device beyond at least a first outer layer of an anulus of an intervertebral disc, the distal end configured to deliver the implant within the intervertebral disc;(ii) delivering the implant in the intervertebral disc using the implant delivery device,wherein the implant is configured to be held along an innermost layer of the anulus in the space occupied by the nucleus; and(iii) inserting the at least one fixation device into a vertebral body surrounding the intervertebral disc. 17. The metho... | Intrinsic Therapeutics, Inc. | A61F002/44 A61B005/107 A61B017/32 A61B017/3207 A61B017/70 A61F002/46 A61B005/00 A61B017/00 A61B017/22 A61B017/34 A61B019/00 A61F002/28 A61F002/30
US10188352 | B2 | 2019-01-29 | Acne-affected skin determination method and acne-affected skin determination device | 2015-12-21 | There is provided a acne-affected skin determination method including: setting an upper region and a lower region so as to vertically divide the face of a subject into two regions; acquiring a measurement value of the amount of moisture in the lower region; acquiring a measurement value of the amount of moisture in the upper region; calculating a target moisture value indicating the level of the amounts of moisture distributed in one of the upper region or the lower region and a reference moisture value indicating the level of the amounts of moisture distributed in a reference region including the other of the upper region and the lower region, based on the measurement values of the amounts of moisture in the upper region and the lower region; obtaining the degree of change in moisture of the target moisture value with respect to the reference moisture value; and determining the ease of generation of acne in the face of the subject based on the obtained degree of change in moisture. | 1. An acne-affected skin determination method comprising: setting an upper region and a lower region so as to vertically divide the face of a subject into two regions,acquiring amount of moisture of a plurality of measurement sites in the lower region;acquiring amount of moisture of a plurality of measurement sites in the upper region;calculating an average value of the amounts of moisture of the plurality of measurement sites in one of the upper region or the lower region as a target moisture value, setting a region including the other of the upper region and the lower region as a reference region and calculating an average value of the amounts of moisture of the plurality of measurement sites;obtaining a difference in moisture of the target moisture value with respect to the reference moisture value; anddetermining the ease of generation of acne in the face of the subject based on the difference in moisture obtained. 2. The acne-affected skin determination method according to claim 1, wherein the target moisture value indicates the average value of the amounts of moisture of the plurality of measurement sites in the lower region and the reference moisture value indicates the average value of the amounts of moisture of the plurality of measurement sites in the upper region. 3. The acne-affected skin determination method according to claim 1, wherein the target moisture value indicates the average value of the amounts of moisture of the plurality of measurement sites in the lower region and the reference moisture value indicates the average value of the amounts of moisture of the plurality of measurement sites in a region of the entirety of the face. 4. The acne-affected skin determination method according to claim 1, wherein, with respect to a half section of the face of the subject in which the face is divided into a left region and a right region by a central vertical line, the half section of the face is divided horizontally into three regions using a first vertical line which passes through the corner of the mouth from an inner corner of an eye and a second vertical line which passes through an outer corner of the eye, and the half section of the face is divided vertically into six regions using a first horizontal line which vertically divides the forehead into two regions, a second horizontal line which passes through an upper end portion of the eye, a third horizontal line which passes through a lower end portion of the eye, a fourth horizontal line which passes through an upper end portion of the mouth, and a fifth horizontal line which passes through a lower end portion of the mouth, so as to divide the half section of the face into sixteen regions except for the eye and the mouth,wherein the regions positioned on an upper side of the third horizontal line are set to regions A1 to A8, the regions positioned between the third horizontal line and the fourth horizontal line are set to regions A9 to A11 from the central portion to the sid... | FUJIFILM Corporation | A61B005/053 A61B005/00 G06F019/00
US20160120454 | A1 | 2016-05-05 | System for monitoring incontinence | 2014-11-04 | According to one aspect of the inventive concept there is provided a system for monitoring incontinence comprising: a urine sensitive circuit arranged to present a changed electrical characteristic when exposed to urine; a measurement circuit arranged to perform a measurement on a urine bladder of a wearer to determine at least one parameter which varies with a fill level of the urine bladder; a sensor arranged to determine an orientation and/or a movement of the sensor; and a processing circuit arranged to: determine whether the urine sensitive circuit has been exposed to urine; estimate an amount of urine released on to the urine sensitive circuit; andin response to determining that the urine sensitive circuit has been exposed to urine, record data representing said at least one parameter determined by the measurement circuit, an estimated movement and/or posture of the wearer based on an orientation and/or a movement determined by the sensor, and an estimate of the amount of urine released on to the urine sensitive circuit. | 1. A system for monitoring incontinence comprising: a urine sensitive circuit arranged to present a changed electrical characteristic when exposed to urine;a measurement circuit arranged to perform a measurement on a urine bladder of a wearer to determine at least one parameter which varies with a fill level of the urine bladder;a sensor arranged to determine an orientation and/or a movement of the sensor; anda processing circuit arranged to: determine whether the urine sensitive circuit has been exposed to urine,estimate an amount of urine released on to the urine sensitive circuit, andin response to determining that the urine sensitive circuit has been exposed to urine, record data representing said at least one parameter determined by the measurement circuit, an estimated movement and/or posture of the wearer based on an orientation and/or a movement determined by the sensor, and an estimate of the amount of urine released on to the urine sensitive circuit. 2. A system according to claim 1, wherein the changed electrical characteristic of the urine sensitive circuit is indicative of the amount of urine released on to the urine sensitive circuit. 3. A system according to claim 1, wherein the urine sensitive circuit is arranged to present an electrical response when exposed to urine which is indicative of an amount of urine released on to the urine sensitive circuit, and wherein the processing circuit is arranged to determine that the urine sensitive circuit has been exposed to urine by detecting the electrical response of the urine sensitive circuit and to estimate the amount of urine released on to the urine sensitive circuit based on the electrical response of the urine sensitive circuit. 4. A system according to claim 3, wherein the changed electrical characteristic or the electrical response of the urine sensitive circuit includes one or more selected from the group comprising: a changed resistance, a changed capacitance, a changed inductance, a changed impedance, a changed resonance frequency, a changed voltage generated by the urine sensitive circuit, a changed current generated by the urine sensitive circuit, a changed resonance frequency of the urine sensitive circuit. 5. A system according to claim 4, wherein the processing circuit is arranged to estimate the amount of urine released on to the urine sensitive circuit by determining a size of the change. 6. A system according to claim 1, wherein the urine sensitive circuit includes a portion arranged to generate a current when exposed to urine to power a transmission of a signal from the urine sensitive circuit, and wherein the processing circuit is arranged to record said data in response to receiving the signal from the urine sensitive circuit. 7. A system according to claim 6, wherein the signal transmitted from the urine sensitive circuit is indicative of the amount of urine released on to the urine sensitive circuit, and wherein the processing circuit is arranged to estimate ... |  | A61B005/20 A61B005/145 A61B005/11 A61B005/01 A61B005/07 A61B007/02 A61B007/04 A61B005/1473 A61B005/00
US20160135741 | A1 | 2016-05-19 | APPARATUSES FOR HOME USE IN DETERMINING TISSUE WETNESS | 2014-07-01 | Compact and lightweight, non-invasive apparatuses to determine tissue wetness/hydration based on the frequency responses of regions of the tissue below a sensor of the apparatus. Described herein are compact and lightweight apparatuses having a sensor with an array of electrodes that is directly connected or connectable to control circuitry to attach to the back of the sensor, which can be worn by a patient. The control circuitry may include a multiplexer (MUX) coordinating the reciprocal selection of drive and sensing electrodes, and a one or more constant current sources. Methods of using these devices to detect tissue wetness are also described. | 1. A compact and lightweight device for detecting tissue wetness, the device comprising: a flexible sensor having a front and a back;an array of electrodes arranged on the front of the sensor;control circuitry configured to apply a constant current at a plurality of frequencies to the array of electrodes, the control circuitry comprising: a constant current source,a multiplexer adapted to select electrodes from the array of electrodes to act as a current source electrode, a current sink electrode, and pairs of voltage sensing electrodes,a controller connected to the multiplexer and the constant current source and adapted to drive current between different combinations of current source and current sink electrodes and to record voltages from the sensing electrodes; anda connector on the sensor adapted to connect the control circuitry to the array of electrodes. 2. A compact and lightweight device for detecting tissue wetness, the device comprising: a sensor having a front and a back;an array of electrodes arranged on the front of the sensor;control circuitry configured to apply a constant current at a plurality of frequencies, the control circuitry comprising: a constant current source comprising a wideband digital to analog converter,a multiplexer adapted to select electrodes from the array of electrodes to act as a current source electrode, a current sink electrode, and pairs of voltage sensing electrodes, anda controller connected to the multiplexer and the constant current source and adapted to control the multiplexer to sequentially drive current between different combinations of current source and current sink electrodes and to record sensed voltages; anda connector on the sensor adapted to connect the control circuitry to the array of electrodes. 3. A compact and lightweight device for detecting tissue wetness, the device comprising: a sensor having a front and a back;an array of electrodes arranged on the front of the sensor;control circuitry configured to apply a constant current at a plurality of frequencies, the control circuitry comprising: a first constant current source;a second constant current source that is 180 degrees out of phase with the first constant current source;a multiplexer adapted to select electrodes from the array of electrodes to act as a current source electrode, a current sink electrode, and pairs of voltage sensing electrodes, anda controller connected to the multiplexer and the first and second constant current sources, and adapted to control the multiplexer to sequentially drive current from the first constant current source on the current source electrode and current from the second constant current source on the current sink electrode and to record voltages sensed on the voltage sensing electrode pairs;a connector on the sensor adapted to connect the control circuitry to the array of electrodes. 4. The device of claim 1, wherein the constant current source comprises a digital to analog converter configured ... |  | A61B005/00 A61B005/053
US20160146832 | A1 | 2016-05-26 | METHODS AND COMPOSITIONS FOR DIAGNOSIS AND PROGNOSIS OF RENAL INJURY AND RENAL FAILURE | 2014-06-05 | It is an object of the present invention to provide a combination of a functional assessment of renal function together with biomarker results in order to improve assessment of patient at risk of or having, an acute kidney injury. A loop diuretic such as furosemide inhibits luminal active chloride transport throughout the thick ascending limb of Henle, thereby preventing sodium reabsorption and resulting in natriuresis and increased urine flow. Loop diuretic-induced increases in urine output might be a method to assess the integrity of the renal tubular function in the setting of early AKI, and so a kidney's response, or lack thereof, to a diuretic challenge as a clinical assessment of tubular function can identify patients with severe tubular injury before it is clinically apparent (e.g. a rise in creatinine). | 1. A method of evaluating a patient for acute kidney injury progression, comprising: performing a biomarker assay on a body fluid sample obtained from the patient to provide an assay result, wherein the biomarker(s) in the biomarker assay are correlated to a risk of acute kidney injury;correlating the assay result to the patient's risk of acute kidney injury relative to a risk in a predetermined population of individuals;classifying the patient's suitability for receiving a diuresis stress evaluation using the patient's risk of acute kidney injury;if the patient is classified as suitable, performing the diuresis stress evaluation on the patient to provide a diuresis-induced urine output value; andcorrelating the diuresis-induced urine output value to a risk of acute kidney injury progression. 2. The method of claim 1, wherein the biomarker assay comprises measuring a concentration of one or more biomarkers selected from the group consisting of Metalloproteinase inhibitor 2, Thrombospondin-1, Antileukoproteinase, Insulin-like growth factor-binding protein 7, Metalloproteinase inhibitor 4, Metalloproteinase inhibitor 1, Hyaluronic acid, Transmembrane glycoprotein NMB, Follistatin, Hepatocyte growth factor, Tumor necrosis factor receptor superfamily member 6, Growth-regulated alpha protein, C—C motif chemokine 24, Metalloproteinase inhibitor 3, C—X—C motif chemokine 6, Tumor necrosis factor receptor superfamily member 11B, Cystatin-C, Beta-2-microglobulin, Serum albumin, Clusterin, Interleukin-8, Neutrophil gelatinase-associated lipocalin, Interleukin-2 receptor alpha chain, Hepatitis A virus cellular receptor 1, Chitinase-3-like protein 1, Serum creatinine, and Hyaluronic acid. 3. The method of claim 2, wherein the biomarker assay comprises measuring a concentration of Metalloproteinase inhibitor 2. 4. The method of claim 3, wherein the biomarker assay comprises measuring a concentration of Insulin-like growth factor-binding protein 7. 5. The method of claim 4, wherein classifying the patient's suitability for receiving a diuresis stress evaluation further comprises using a baseline urine output value for the patient. 6. The method claim 4, wherein the classifying the patient's suitability for receiving a diuresis stress evaluation further comprises using a baseline change in serum creatinine for the patient. 7. The method of claim 4, wherein the classifying the patient's suitability for receiving a diuresis stress evaluation further comprises determining if the patient has experienced 6 hours of oliguria defined as a urine output of &lt;0.5 ml/kg/hour, or a 0.3 mg/dL rise in serum creatinine, or an increase of 150-200% above a baseline serum creatinine concentration. 8. The method of claim 4, wherein the classifying step comprises identifying the patient as suitable to receive a diuresis stress evaluation if the patient's relative risk is greater than 1.0. 9. The method of claim 4, wherein the classifying step comprises identifying the pat... |  | G01N033/68 A61B005/145 A61K031/341 G01N033/70 G01N033/493
US10006882 | B2 | 2018-06-26 | Biosensing system and methods using electron-ionic mechanisms at fluid-sensor interfaces | 2015-11-20 | An example biosensor is provided and includes a semiconductor sensing element, a first electrode and a second electrode located on a first plane of the sensing element with a first electric field being applied thereacross, a third electrode located on a second plane of the sensing element parallel to and removed from the first plane with a second electric field being applied across the first electrode and the third electrode perpendicular to the first electric field, and a dielectric substrate having a first portion that constrains a fluid including an analyte on a surface of the sensing element, and a second portion that facilitates dielectric separation of the fluid from the electrodes. The mutually perpendicular electric fields facilitate adjusting a height of a fluid-sensor interface comprising an electrical double layer in the fluid enabling detection and characterization of the analyte. | 1. A biosensor, comprising: a semiconductor sensing element;a first conducting electrode and a second conducting electrode located on a first plane of the sensing element, wherein a first electric field is applied across the first electrode and the second electrode;a third conducting electrode located on a second plane of the sensing element parallel to and removed from the first plane, wherein a second electric field is applied across the first electrode and the third electrode perpendicular to the first electric field; anda dielectric substrate having a first portion adjacent a second portion, the first portion adjacent the sensing element and comprising pores, the first portion facilitating reception and containment of a fluid comprising an analyte on a surface of the sensing element distal from the third electrode, the surface defining an interface between the first portion and the sensing element, and the second portion facilitating dielectric separation of the fluid from the first electrode, the second electrode and the third electrode, wherein the mutually perpendicular first electric field and the second electric field facilitate adjusting a height of a fluid-sensor interface comprising an electrical double layer (EDL) in the fluid enabling detection and characterization of the analyte;wherein the first portion of the substrate facilitates reception and containment of the fluid in the pores. 2. The biosensor of claim 1, wherein the first electric field is generated from direct current (DC) voltage applied across the first electrode and the second electrode, wherein the second electric field is generated from alternating current (AC) voltage applied across the first electrode and the third electrode. 3. The biosensor of claim 1, wherein the fluid-sensor interface further comprises a surface of the sensing element. 4. The biosensor of claim 3, wherein the fluid-sensor interface further comprises a layer of linker molecules bound to the surface of the sensing element. 5. The biosensor of claim 4, wherein the fluid-sensor interface further comprises a layer of capture probes bound to the layer of linker molecules. 6. The biosensor of claim 5, wherein the fluid-sensor interface further comprises blocking molecules bound to some of the linker molecules. 7. The biosensor of claim 1, wherein a first electrical response is measured across the first electrode and the third electrode, wherein a second electrical response is measured across the first electrode and the third electrode, wherein a comparison between the first electrical response and the second electrical response provides a measure of a signal-to-noise ratio (SNR) of the biosensor. 8. The biosensor of claim 7, wherein the first electric field is adjusted to change a Debye length of the EDL generated in the fluid, wherein the Debye length affects the SNR of the biosensor. 9. The biosensor of claim 7, wherein a concentration of the analyte is estimated from the second electrical r... | EnLiSense, LLC | G01N027/447 G01N033/543 A61B005/1477 A61B005/1468
US09332927 | B2 | 2016-05-10 | Fluid heat transfer system for animal diagnostic bed and methods of using the same | 2014-03-05 | A diagnostic system is disclosed, and comprises a diagnostic device and a fluid heat transfer system. The diagnostic device has an associated bed for supporting a portion of a subject during a diagnostic procedure. The fluid heat transfer system comprises a fluid circulation component, a heater component, and a fluid line. The heater component includes a heater element and a mixing valve, and the fluid circulation component is fluidly coupled to the heater component so that a fluid can be pumped from the fluid circulation component to the heater component. The fluid line extends from the heater to the fluid circulation component, and the mixing valve is positioned near the bed so that heat can be rapidly transferred through the fluid line to alter a body temperature of the subject. | 1. A diagnostic system, comprising: an MRI device and an associated bed for supporting a portion of a subject during a diagnostic procedure;a fluid heat transfer system, comprising: a fluid circulation component;a heater component including a heater element and a mixing valve, the fluid circulation component fluidly coupled to the heater component so that a fluid can be pumped from the fluid circulation component to the heater component; anda fluid line extending from the heater to the fluid circulation component, the mixing valve of the heater positioned near the bed so that heat can be rapidly transferred through the fluid line to alter a body temperature of the subject. 2. The diagnostic system of claim 1, wherein the mixing valve comprises an inlet for the fluid at a first temperature and an inlet for the fluid at a second, higher temperature. 3. The diagnostic system of claim 2, wherein the mixing valve comprises an outlet for the fluid at a temperature intermediate the first temperature and second temperature. 4. The diagnostic system of claim 1, wherein the bed comprises a frame and a plurality of protrusions extending therefrom. 5. The diagnostic system of claim 1, wherein the fluid line is configured to interengage one or more protrusion of the plurality of protrusions to maintain the fluid line in position near the bed. 6. The diagnostic system of claim 1, wherein the fluid circulation component comprises a pump. 7. The diagnostic system of claim 1, further comprising a control interface electrically coupled with one or more of fluid circulation component and heater component. 8. The diagnostic system of claim 7, wherein the control interface is electrically coupled with the mixing valve. 9. The diagnostic system of claim 8, wherein the control interface can cause a flow rate through one of a first inlet and a second inlet of the mixing valve to change. 10. A diagnostic system, comprising: a bed configured for use with an MRI device and to support a portion of a subject during a diagnostic procedure; anda fluid heat transfer system, comprising: a first fluid line fluidly coupled with a heating element;a second fluid line;a mixing valve disposed downstream of the heating element and having a first fluid inlet corresponding to the first fluid line and a second fluid inlet corresponding to the second fluid line, the mixing valve having an outlet for a mixed fluid having a temperature higher than a temperature of fluid in the second fluid line; anda third fluid line extending away from the mixing valve, the mixing valve positioned near the bed such that the mixed fluid rapidly transfers heat from the third fluid line to a subject on the bed. 11. The diagnostic system of claim 10, wherein a first fluid pump is fluidly coupled with the first fluid line. 12. The diagnostic system of claim 11, wherein a second fluid pump is fluidly coupled with the second fluid line. 13. The diagnostic system of claim 11, further comprising a contro... | SA Instruments, Inc. | A61B005/05 A61B005/055
US20160147794 | A1 | 2016-05-26 | ULTRASONIC DIAGNOSTIC APPARATUS, MEDICAL IMAGE PROCESSING APPARATUS, AND MEDICAL IMAGE PROCESSING METHOD | 2015-11-19 | According to one embodiment, an ultrasonic diagnostic apparatus includes a plurality of types of databases for detecting a contour position of a predetermined region of an object, input interface circuitry, and detection circuitry. The input interface circuitry designates a database, of the plurality of databases, which is desired by a user. The detection circuitry detects a contour position of the predetermined region on an ultrasonic image as an input image by using the designated database. | 1. An ultrasonic diagnostic apparatus comprising: a plurality of types of databases for detecting a contour position of a predetermined region of an object;input interface circuitry configured to designate a database, of the plurality of databases, which is desired by a user; anddetection circuitry configured to detect a contour position of the predetermined region on an ultrasonic image as an input image by using the designated database. 2. The apparatus of claim 1, wherein the plurality of types of databases include a definition of a contour position of the predetermined region determined by a standard ultrasonic diagnostic professional and a definition of a contour position of the predetermined region determined based on an image different from an ultrasonic image. 3. The apparatus of claim 1, wherein the plurality of types of databases include a contour position of the predetermined region recognized by a specific ultrasonic diagnostic expert and an average position of a contour position of the predetermined region recognized by a plurality of standard ultrasonic diagnostic professionals. 4. The apparatus of claim 1, wherein the plurality of types of databases include a contour position of the predetermined region obtained from a reference image having a resolution different from that of the input image. 5. The apparatus of claim 4, wherein the reference image comprises a magnetic resonance imaging image, an X-ray computed tomographic image, or an ultrasonic image having a higher spatial frequency than the input image. 6. The apparatus of claim 1, further comprising learning circuitry configured to update at least one of the plurality of types of databases by learning an operation by a teacher, wherein the plurality of types of databases are constructed for each hospital facility which has performed the learning or for the each teacher. 7. The apparatus of claim 1, wherein the input interface circuitry designates at least two of the plurality of types of databases, and the detection circuitry detects a contour position of the predetermined region by using the designated at least two databases. 8. The apparatus of claim 1, wherein the input image comprises a slice image, M-mode image, or Doppler waveform image of the predetermined region. 9. A medical image processing apparatus comprising: a storage unit configured to store an input image including a predetermined region of an object and a plurality of types of databases for detecting a contour position of the predetermined region of the object;input interface circuitry configured to designate a knowledge-based dictionary including a database, of the plurality of types of databases, which is desired by a user; anddetection circuitry configured to detect a contour position of the predetermined region on a predetermined input image by using the designated knowledge-based dictionary. 10. The apparatus of claim 9, wherein the input image comprises an ultrasonic image, a magnetic resonance... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | G06F017/30 A61B008/08 G06T007/00 A61B005/055 A61B006/03 A61B008/00 A61B008/14
US09329694 | B2 | 2016-05-03 | Preemptive machine learning-based gesture recognition | 2014-08-07 | A system and method for detecting a viewing gesture with respect to a wrist-worn device employ a logistic-regression model to pre-learn gesture metrics for on events. An output model is produced for deployment on a consumer device, allowing real-time gesture detection with high accuracy and low latency. | 1. A method of recognizing a viewing gesture made by a user of a user-wearable device having a three-dimensional (“3D”) accelerometer and a viewable display, the method comprising: providing one or more tester-wearable devices to be worn by testers, each tester-wearable device having at least a 3D accelerometer and a selection element;collecting accelerometer data and selection-element data as each of the one or more tester-wearable devices is worn, each tester selecting the selection element to signify on events when the display is desired to turn on;calculating at least one metric based on accelerometer data surrounding each on event and employing the at least one metric as a feature for logistic regression;generating a potential logistic-regression model describing a relationship between tester motion and tester viewing of the viewable display; andtraining, validating, and testing the potential logistic-regression model to generate an output model that includes a vector of metric means, a vector of metric standard deviations, and a vector of metric weights. 2. The method of claim 1 wherein the accelerometer data surrounding each on event are biased towards samples prior to the event. 3. The method of claim 1 further comprising loading the output model onto the user device. 4. The method of claim 3 further comprising collecting accelerometer data at the user device and checking for on events based on the output model. 5. The method of claim 4 further comprising powering the viewable display on when an on event is detected. 6. The method of claim 1 wherein calculating at least one metric based on accelerometer data surrounding each on event further comprises identifying as negative events any gestures with respect to which the testers did not select the selection element to signify an on event. 7. The method of claim 1 wherein training, validating, and testing the logistic-regression model to generate an output model comprises employing separate portions of the collected accelerometer data and selection-element data for each of training, validating, and testing. 8. The method of claim 7 wherein the respective portions of the collected accelerometer data and selection-element data for each of training, validating, and testing are about 60%, 20%, and 20%. 9. The method of claim 1 wherein the training, validating, and testing of the logistic-regression model to generate an output model are executed in MATLAB™. 10. A wrist-worn user device comprising: a viewable display having a powered on mode and a powered off mode;a three-dimensional accelerometer; anda controller, the controller being configured to maintain the viewable display in the powered off mode and to sample data from the accelerometer and to detect, based on the sampled accelerometer and a pre-loaded pre-trained logistic-regression model, whether the user intends to view the viewable display, and to change the viewable display to the powered on mode if it is determined that the ... | Google Technology Holdings LLC | G06F003/01 A61B005/00 G06N007/02
US09341694 | B2 | 2016-05-17 | Method and magnetic resonance system for distortion correction in magnetic resonance imaging | 2012-06-08 | In a method for distortion correction in spiral magnetic resonance imaging, a first MR data set is acquired by scanning raw data space along a spiral trajectory beginning at a first point. A first complex MR image is determined from the first MR data set, which includes first phase information for image points of the first MR image. A second MR data set is acquired by scanning raw data space along the spiral trajectory beginning at a second point that differs from the first point. A second complex MR image is determined from the second MR data set, which includes second phase information for image points of the second MR image. A geometric distortion for image points of the first or second MR image is determined from the first and second phase information, for example with a PLACE method. | 1. A method for correcting distortion in spiral magnetic resonance imaging wherein raw magnetic resonance data are acquired with modulation of at least two gradients respectively in two spatial directions comprising: in a computer, automatically entering a first magnetic resonance raw data set, acquired after a single radio-frequency excitation, into k-space along a spiral trajectory in k-space, beginning at a first point of k-space;in said computer, automatically determining a first complex MR image from said first MR raw data set, said first complex MR image comprising first phase information for respective image points of said first MR image;in said computer, automatically entering a second MR raw data set, also acquired after a single radio-frequency excitation, into k-space along said spiral trajectory, beginning at a second point of k-space along said spiral trajectory that differs from said first point;in said computer, automatically determining a second complex MR image from said second MR raw data set, said second complex MR image comprising second phase information for respective image points of said second MR image;in said computer, automatically determining a geometric distortion for respective image points of said first MR image or said second MR image from said first and second phase information; andin said computer, automatically correcting said geometric distortion in said first MR image or in said second MR image, to produce at least one corrected MR image, and making said at least one corrected MR image available in electronic form at an output of said computer, as a data file. 2. The method as claimed in claim 1 comprising entering said first and second MR raw data sets into k-space along said spiral trajectory with a predetermined radius that decreases along said spiral trajectory. 3. The method as claimed in claim 1 comprising entering said first and second MR raw data sets into k-space along said spiral trajectory with a predetermined radius that increases along said spiral trajectory. 4. The method as claimed in claim 1 comprising entering said first and second MR raw data sets into k-space with a variable speed. 5. The method as claimed in claim 1 comprising entering said first and second MR raw data sets into k-space with a variable density. 6. The method as claimed in claim 1 comprising determining said geometric distortion by using a PLACE method to generate at least one of a field map and a displacement map. 7. A method for correcting distortion in radial magnetic resonance imaging wherein raw magnetic resonance data are acquired with modulation of at least two readout gradients respectively in two spatial directions comprising: in a computer, automatically entering a first magnetic resonance raw data set, acquired after a single radio-frequency excitation, into k-space along multiple radial trajectories in k-space respectively proceeding at different angles through a center of k-space, beginning at a first point... | Siemens Aktiengesellschaft | G06K009/40 G01R033/34 G01R033/565 A61B005/055 G01R033/48 G01R033/561
US09339208 | B2 | 2016-05-17 | Tracheal tube with pressure monitoring lumen and method for using the same | 2010-01-18 | According to various embodiments, methods and systems for determining pressure in the lungs may employ tracheal pressure measurements. The tracheal pressure measurements may be obtained through a pressure monitoring lumen associated with a tracheal tube. Such systems may include a purging or flushing mechanism to keep the pressure monitoring lumen free of any obstructions. The flushing mechanism may utilize fluids delivered through the pressure monitoring lumen at a predetermined point in the breathing cycle and/or at a predetermined pressure relative to the respiratory gases. The resulting pressure measurements may be used to determine a more accurate estimate of lung pressure, which in turn may be used to control a ventilator and provide breathing assistance to a patient. | 1. A method comprising: measuring a tracheal pressure of a patient's trachea with a pressure transducer in fluid communication with a pressure monitoring lumen, wherein the pressure monitoring lumen is associated with a tracheal tube;receiving information related to a respiratory gas pressure over a breathing cycle of a patient;determining a baseline pressure associated with a minimum pressure point over the course of the breathing cycle of the patient;delivering a purging fluid as a bolus of purging fluid timed to coincide with the minimum pressure point over the course of the breathing cycle of the patient through the pressure monitoring lumen at a purging pressure greater than the respiratory gas pressure; anddetermining a corrected tracheal pressure of the patient's trachea that accounts for the bolus of purging fluid based at least in part on the baseline pressure, wherein determining the corrected tracheal pressure comprises substituting the measured pressure during the bolus delivery with the baseline pressure. 2. The method of claim 1, further comprising determining a second corrected tracheal pressure at a different point over the course of the breathing cycle based on a difference between a purging pressure and the tracheal pressure of the patient's trachea. 3. The method of claim 1, comprising determining the baseline pressure in the pressure monitoring lumen at exhalation and using the baseline pressure to determine the corrected tracheal pressure. 4. The method of claim 3, wherein the baseline pressure is a pressure left in the lungs after an end of exhalation. 5. The method of claim 1, wherein delivering the purging fluid is controlled by instructions encoded within a processor. 6. The method of claim 1, further comprising determining a second corrected tracheal pressure at a different point over the course of the breathing cycle by subtracting a purging pressure from the pressure in the pressure monitoring lumen. | COVIDIEN LP | A61B005/08 A61M016/04 A61M016/00
US09332969 | B2 | 2016-05-10 | Fluid flow control apparatus and patient fluid sampling method | 2009-01-06 | A fluid flow control apparatus used in medical fluid sampling. The apparatus includes a valve assembly, first and second ports, and a cap. The valve assembly comprises a body rotatably assembled to a housing and defining a passageway and at least two recesses defined by an interior surface of the housing and the body which is formed separate from the passageway. The ports extend from the housing and are configured for connection to medical devices. The valve assembly is operable to interchangeably align the passageway and the recesses with the ports. Finally, the cap is connected to the housing, and forms an opening for coupling to a specimen container. | 1. A fluid flow control apparatus used in medical fluid sampling, comprising: a valve assembly comprising a housing and a body, the housing having a cavity between a top surface and a first opening of the housing, the body having a top portion and a bottom portion, the body rotatably retained within the cavity such that the bottom portion terminates above the housing first opening and the top portion extends above the housing top surface, the body comprising a passageway through a horizontal axis of the body and defining at least two diametrically opposed circumferential recesses along respective vertical axes of the body parallel to the axis of rotation, the at least two diametrically opposed circumferential recesses defined by an interior circumferential surface of the housing and respective recessed portions of an exterior surface of the body, that are fluidly isolated from the passageway;the housing having (i) a first port and a second port, each port extending from the housing and configured for connection to medical devices, and (ii) a ridge on the top surface of the housing comprising first and second protrusions, the first protrusion being engaged by a stem portion in a first flow position, and the body being rotatable such that the stem portion engages a second protrusion in a second flow position; anda cap integrally connected to the first opening of the housing, the cap forming a chamber for coupling to a specimen container, the chamber having cross-sectional width that is greater than the cross-sectional width of the first opening,wherein the chamber of the cap is configured to be fluidly connected to the first and second ports through the at least two diametrically opposed circumferential recesses when two of the at least two diametrically opposed circumferential recesses are aligned with the first and second ports,wherein the valve assembly is operable to interchangeably align the passageway and two of the at least two diametrically opposed circumferential recesses with the first and second ports. 2. The fluid flow control apparatus of claim 1, wherein the valve assembly comprises a stopcock valve. 3. The fluid flow control apparatus of claim 1, wherein the first port is a female fitting and the second port is a male fitting. 4. The fluid flow control apparatus of claim 1, wherein the first port and second port are luer fittings. 5. The fluid flow control apparatus of claim 1, wherein the valve assembly is configured to provide at least two flow positions. 6. The fluid flow control apparatus of claim 5, wherein the protrusion corresponds to each of the at least two flow positions of the valve assembly. 7. The fluid flow control apparatus of claim 5, wherein a first flow position of the valve assembly includes the passageway fluidly connecting the first and second ports. 8. The fluid flow control apparatus of claim 5, wherein a second flow position of the valve assembly defines a direction of fluid flow from the first port, th... | CareFusion 207, Inc. | A61C017/06 A61C017/14 A61M001/00 F16K011/085 E03B007/07 F16K037/00 F17D003/00 A61B005/00 B65D081/00 A61B010/00 A61M039/22 F16K011/083
US09333298 | B2 | 2016-05-10 | Apparatus and method for automatically supplying insulin based on amount of exercise | 2013-03-14 | An apparatus and method for automatically supplying insulin based on exercise amount are provided. The apparatus includes a blood sugar measurer for measuring a blood sugar level of a patient; a controller for, comparing the measured blood sugar level with a predetermined blood sugar level, if the measured blood sugar level is greater than or equal to the predetermined blood sugar level, determining whether a current time is within a blood sugar change time zone, if the measured blood sugar level is equal to or higher than the predetermined blood sugar level, and if the current time does is not within the blood sugar change time zone, acquiring exercise amount information about the patient and determining a dose of insulin to be injected inject into the patient based on the acquired exercise amount information, if the current time does not fall within the blood sugar change time zone; and an insulin injector for injecting the determined dose of insulin into the patient. | 1. An apparatus for automatically supplying insulin based on exercise amount, the apparatus comprising: a blood sugar measurer for measuring a blood sugar level of a patient;a controller for: comparing the measured blood sugar level with a predetermined blood sugar level,if the measured blood sugar level is greater than or equal to the predetermined blood sugar level, determining whether a current time is within a blood sugar change time zone,if the current time is within the blood sugar change time zone, determining a dose of insulin to inject into the patient based on the measured blood sugar level, andif the current time is not within the blood sugar change time zone, acquiring exercise amount information about the patient and determining the dose of insulin to inject into the patient based on the acquired exercise amount information; andan insulin injector for injecting the determined dose of insulin into the patient. 2. The apparatus of claim 1, further comprising an exercise amount measurer for acquiring the exercise amount information about the patient. 3. The apparatus of claim 1, wherein the controller further receives the exercise amount information from a mobile terminal through short-range wireless communication. 4. The apparatus of claim 1, further comprising a memory for storing a blood sugar change record of the patient and living pattern information about the patient. 5. The apparatus of claim 4, wherein, if the measured blood sugar level is greater than or equal to the predetermined blood sugar level, the controller further determines whether the current time falls within the blood sugar change time zone by reading the blood sugar change record. 6. The apparatus of claim 1, wherein the controller further determines whether the patient is exercising based on the acquired exercise amount information and, if the patient is exercising, determines the dose of insulin based on the acquired exercise amount information. 7. The apparatus of claim 4, wherein the controller further determines whether the exercise was planned based on the living pattern information about the patient, determines a blood sugar change based on the exercise amount information if the exercise was not planned, and if the blood sugar change is above a threshold range, warns the patient to stop exercising. 8. The apparatus of claim 7, wherein the controller further measures another blood sugar level of the patient through the blood sugar measurer and determines a dose of insulin based on the measured blood sugar level, after warning the patient to stop the exercise. 9. A method for automatically supplying insulin based on exercise amount in an automatic insulin supply apparatus, the method comprising: receiving a blood sugar level of a patient measured by a blood sugar sensor;comparing the measured blood sugar level with a predetermined blood sugar level;if the measured blood sugar level is greater than or equal to the predetermined blood sugar level, determ... | Samsung Electronics Co., Ltd | A61M005/172 A61K038/28 A61B005/00 G06F019/00 A61B005/145 A61B005/11 A61M005/142
US09326823 | B2 | 2016-05-03 | Real-time tracking and navigation system and method for minimally invasive surgical procedures | 2013-05-02 | Telemetrical control of a robotic interventional device for minimally invasive surgical procedure is based on an operative interaction between a tracking sub-system, MRI sub-system, navigation sub-system and the robotic interventional device. The tracking sensor sub-system is integrated with the interventional device to produce tracking information corresponding to the robotic interventional device location in the operative site. The navigation sub-system integrates the tracking information with the real-time images of the operative site produced by the MRI sub-system, and displays the integrated information to a user, to enable the telemetrical control of the interventional device for performing an intended procedure (biopsy, tissue resection, etc.). The navigation sub-system, based on the integrated real-time tracking information and real-time images, calculates and dynamically updates coordinates of subsequent imaging slices. | 1. A system for telemetrical control of an interventional device for minimally invasive surgical procedure, comprising: a Magnetic Resonance Imaging (MRI) scanner sub-system operated to produce substantially in real-time MR images of an operative site comprising a target of interest;an interventional device disposed at said operative site within said MRI scanner sub-system;a tracking sub-system operatively coupled to said MRI scanner sub-system and operating concurrently therewith, wherein said tracking sub-system is configured to automatically and continually produce tracking information corresponding to a location of said interventional device within said operative site, said tracking information being produced based on gradient fields of said MRI scanner sub-system;a processor controlled navigation sub-system operatively coupled to said tracking sub-system, said MRI scanner sub-system, and said interventional device, wherein said navigation sub-system is configured to:(a) integrate said tracking information with said real-time MR images,(b) display said substantially real-time MR images of said operative site generated by said imaging MRI scanner sub-system, and(c) fluidly interface with a user to present, in a continuous fashion, said real-time MRI images and said tracking information, and to receive the user's commands to telemetrically control said interventional device to reach said target of interest in an optimal manner for an intended interaction therewith based on at least said substantially in real-time MR images and said tracking information. 2. The system of claim 1, further comprising a processor controlled control sub-system operatively coupled to said navigation sub-system and said interventional device, wherein said control sub-system is configured to generate control signals responsive to the user's commands input into said navigation sub-system and to apply said control signals to said interventional device to control the motion thereof in said operative site relative to said target of interest. 3. The system of claim 2, wherein said tracking sub-system further includes a sensor sub-system integrated with said interventional device, wherein said sensor sub-system includes at least one passive magnet field sensor (PMFS) positioned in proximity to a tip of said interventional device. 4. The system of claim 3, wherein said sensor sub-system further includes at least one second PMFS positioned at a predetermined location along the length of said interventional device. 5. The system of claim 3, wherein said imaging MRI scanner sub-system further generates images based on an imaging pulse sequences, wherein said tracking sub-system uses predetermined tracking gradient pulses embedded in said imaging pulse sequences to induce currents in said sensor sub-assembly, and wherein said induced currents are processed to determine the position and orientation of said at least one PMFS. 6. The system of claim 5, further comprising pre-ca... | University of Maryland, College Park University of Maryland, Baltimore | A61B005/05 A61B019/00 A61B010/02 A61B017/00 A61B018/20 A61B006/03 A61B008/08 A61N007/02 A61B005/055 A61B018/04 A61B018/14 A61B018/00
US10292646 | B2 | 2019-05-21 | Mobile health care device and operating method thereof | 2015-04-29 | Provided are a health care device and a method of operating the health care device. The method includes: detecting, by the health care device, biosignal information of a user through skin of the user; obtaining health status information of the user by using the detected biosignal information; providing a task of the user for improving a health status, according to the obtained health status information; and monitoring whether the user achieves the provided task. | 1. A method of operating a health care device, the method comprising: detecting, by a sensor of the health care device, biosignal information of a user through skin of the user;obtaining, by the health care device, health status information of the user based on the detected biosignal information;obtaining, by the health care device, in response to the health status information being within a reference range, a task of the user to cause the health status information to be outside of the reference range;providing, by the health care device, the task; andmonitoring, by the health care device, whether the user achieves the provided task,wherein the monitoring comprises: receiving, from the user via an input to a user interface of the health care device, a user response regarding the provided task,calculating, by the health care device, an achievement rate of the provided task based on the received user response, andcalculating, by the health care device, a correlation value that indicates a degree of an effect of the task on a change in the health status by comparing previous health status information and current health status information only when the achievement rate is equal to or above a predetermined value. 2. The method of claim 1, wherein the reference range includes an abnormal range corresponding to abnormal health status information. 3. The method of claim 1, wherein the providing the task comprises: displaying, by the health care device, one or more details of the provided task; andmodifying at least one of the displayed one or more details according to an input of the user to the user interface of the health care device. 4. The method of claim 3, wherein the modifying comprises, in response to a modification according to the input of the user being outside a predetermined modifiable range, providing a notification indicating that the modification is outside the predetermined modifiable range. 5. The method of claim 3, wherein the modifying comprises, in response to a modification according to the input of the user being outside a predetermined modifiable range, determining a value of the one or more details to be a boundary value of the predetermined modifiable range. 6. The method of claim 3, wherein the one or more details comprise at least one of a type, a cycle, an execution amount, a start time, and an execution duration of the provided task. 7. The method of claim 1, wherein the monitoring further comprises providing an execution start notification regarding the provided task. 8. The method of claim 1, further comprising, in response to the calculated correlation value being below a reference value, stopping the monitoring of whether the user achieves the provided task. 9. The method of claim 1, wherein the biosignal information is detected in a non-invasive manner. 10. The method of claim 1, wherein: the biosignal information comprises at least one of blood glucose, cholesterol, bio-impedance, electrocardiogram (ECG), ball... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/00 A61B005/0205 A61B005/145 A61B005/021 A61B005/0404 A61B005/11 A61B005/0488 A61B005/053 A61B005/024
US09339219 | B2 | 2016-05-17 | Devices, systems, and methods related to analyte monitoring and management | 2015-11-30 | Analyte monitoring devices, systems, and methods are provided that relate to: enabling different application features on a data processing device for analyte monitoring devices with different analyte monitoring features; programming analyte monitoring devices in advance; personalizing an analyte monitoring device; graphically representing a remaining insulin level in a user body; and graphically representing analyte measurement related data for on-demand readings; protecting access to feature of an analyte monitoring device. | 1. A method for selectively displaying features of a first analyte monitoring device and a second analyte monitoring device, comprising: coupling a first analyte monitoring device or a second analyte monitoring device to a computer via a communication interface, wherein the first analyte monitoring device comprises a first plurality of analyte monitoring features selected from performing blood glucose measurements, performing blood ketone measurements, performing carbohydrate measurements, performing bolus insulin calculations, performing basal insulin calculations, and performing device personalization; and the second analyte monitoring device comprises a second plurality of analyte monitoring features selected from performing blood glucose measurements, performing blood ketone measurements, performing carbohydrate measurements, performing bolus insulin calculations, performing basal insulin calculations, and performing device personalization;receiving, with a processor of the computer, an indication that the first analyte monitoring device is communicatively coupled to the computer via the communication interface;determining, with the processor, a first plurality of application features corresponding with the first plurality of analyte monitoring features for the first analyte monitoring device;initiating, with the processor, a user interface application and display the first plurality of application features on a display of the computer;receiving, with the processor, an indication that the second analyte monitoring device is communicatively coupled to the computer via the communication interface;determining, with the processor, a second plurality of application features corresponding with the second plurality of analyte monitoring features for the second analyte monitoring device, and the second plurality of application features: (a) includes an application feature that is absent from the first plurality of application features, or(b) does not include an application feature present in the first plurality of application features; andmodifying, with the processor, the user interface application to display on the display the application feature that is absent from the first plurality of application features, if a determination is made that the second plurality of application features include an application feature that is absent from the first plurality of application features, ormodifying, with the processor, the user interface application to no longer display on the display the application feature that is present in the first plurality of application features but not in the second plurality of application features, if a determination is made that the second plurality of application features does not include the application feature present in the first plurality of application features. 2. The method of claim 1, wherein the second plurality of application features does not include an application feature present in the first plurality of applica... | Abbott Diabetes Care Inc. | C12M001/34 A61B005/145 A61B005/00 A61B005/15
US09326730 | B2 | 2016-05-03 | Breast implant with covering and analyte sensors responsive to external power source | 2014-06-10 | Breast implants including sensor modules and related methods are described herein. Breast implants include those with: a shell configured to be substantially filled with a viscous material; at least one fluid-permeable cover, the cover completely enveloping the shell; and a plurality of sensor modules attached to the shell and positioned at a distance from each other, each of the sensor modules oriented to detect one or more analytes in a fluid between the shell and the cover, wherein each of the plurality of sensor modules includes a unique identifier and is configured to utilize energy transmitted from an external source. | 1. A method of monitoring information from a breast implant, comprising: receiving first information from a first sensor module attached to a shell of a breast implant within an individual, wherein the first information includes a first sensor module identifier and sensor data from both a first sensor unit configured to detect a first analyte and a second sensor unit configured to detect a second analyte within the first sensor module;receiving second information from a second sensor module attached to the shell of the breast implant within the individual, wherein the second information includes a second sensor module identifier and sensor data from both a first sensor unit configured to detect a first analyte and a second sensor unit configured to detect a second analyte within the second sensor module;forming an initial record from the first information and the second information;calculating deviation limits regarding both the first analyte and the second analyte from the initial record;setting deviation parameters based on the deviation limits and a predetermined set of standards;saving the initial record and the deviation parameters in memory in a computing device;receiving third information from the first sensor module attached to the shell of the breast implant within the individual, wherein the third information includes the first sensor module identifier and the sensor data from both the first sensor unit and the second sensor unit within the first sensor module;receiving fourth information from the second sensor module attached to the shell of the breast implant within the individual, wherein the fourth information includes the second sensor module identifier and the sensor data from both the first sensor unit and the second sensor unit within the second sensor module;updating the initial record with the third information and the fourth information;saving an updated record in the memory in the computing device;comparing the updated record to the initial record and to the deviation parameters; andindicating if the updated record is within the deviation parameters of the initial record for both the first analyte and the second analyte. 2. The method of claim 1, comprising: receiving fifth information from the first sensor module attached to the shell of the breast implant within the individual, wherein the fifth information includes the first sensor module identifier and the sensor data from both the first sensor unit and the second sensor unit within the first sensor module;receiving sixth information from the second sensor module attached to the shell of the breast implant within the individual, wherein the sixth information includes the second sensor module identifier and the sensor data from both the first sensor unit and the second sensor unit within the second sensor module;updating the initial record with the fifth information and the sixth information;saving the updated record in memory in the computing device;comparing the update... | Elwha LLC | A61F002/12 A61B005/00 A61B005/07
US09339223 | B2 | 2016-05-17 | Device and method for determining analyte levels | 2013-12-30 | Devices and methods for determining analyte levels are described. The devices and methods allow for the implantation of analyte-monitoring devices, such as glucose monitoring devices, that result in the delivery of a dependable flow of blood to deliver sample to the implanted device. The devices comprise a unique microarchitectural arrangement in the sensor region that allows accurate data to be obtained over long periods of time. | 1. A device for measuring a glucose concentration in a host, the device comprising: an implantable sensor operably connected to an electronic circuit and configured to continuously measure a signal indicative of a glucose concentration in a host for a period of time; anda membrane located over at least a portion of the sensor, wherein the membrane is configured to control a flux of oxygen and glucose, wherein the membrane comprises a product of a reaction of a compound comprising an isocyanate group with a compound comprising a hydroxyl group and/or a compound comprising an amine group;wherein the device during in vitro testing is capable of exhibiting, at a glucose concentration of 400 mg/dL, no more than a 10% drop in sensor output in response to a change of dissolved oxygen concentration from a pO2 of 150 mm Hg to 30 mm Hg. 2. The system of claim 1, wherein the product is polyurethane. 3. The system of claim 1, wherein the product is polyurea. 4. The system of claim 1, wherein the product is a polyurethaneurea copolymer. 5. The system of claim 1, wherein the membrane comprises a layer that is permeable to hydrogen peroxide, wherein the layer restricts passage of ascorbic acid therethrough. | DexCom, Inc. | A61B005/00 A61B005/1486 A61B005/145 C12Q001/00 A61B005/1495
US10295975 | B2 | 2019-05-21 | Processing board, medical diagnostic apparatus, and method of controlling the medical diagnostic apparatus | 2015-09-11 | A processing board mounted on a medical diagnostic apparatus, the medical diagnostic apparatus, and a method of controlling the medical diagnostic apparatus are provided. The processing board includes a sensor configured to sense at least one selected from the group consisting of an installation position, an installation direction, and a power consumption of the processing board. The processing board further includes a controller configured to set an identifier (ID) of the processing board based on an output signal of the sensor. | 1. A processing board mounted on a medical diagnostic apparatus, the processing board comprising: a sensor configured to sense, as an output value, at least one selected from the group consisting of an installation position, an installation direction, and a power consumption of the processing board;a storage configured to store one or more identifier (ID) values for identifying the processing board in the medical diagnosis apparatus, along with one or more output values of the sensor, the one or more output values respectively corresponding to the one or more ID values;a communication interface; anda controller configured to: find, among the one or more ID values that are stored, an ID value corresponding to the output value that is sensed;set the ID value that is found, to identify the processing board in the medical diagnosis apparatus; andcontrol the communication interface to transmit, to another processing board, the ID value that is set. 2. The processing board according to claim 1, wherein the communication interface comprises a controller area network (CAN) communication interface configured to transmit the ID value that is set, to the other processing board via a CAN communication network. 3. The processing board according to claim 1, wherein the sensor comprises at least one selected from the group consisting of an accelerometer, a magnetic sensor, a height sensor, and a current sensor. 4. The processing board according to claim 3, wherein the accelerometer is configured to sense a direction of gravity of the processing board with a strength greater than a threshold. 5. The processing board according to claim 1, further comprising a switch configured to output a switch value, based on a turn-on or turn-off input. 6. The processing board according to claim 1, wherein the medical diagnostic apparatus is one among an ultrasonic diagnostic apparatus, a magnetic resonance imaging apparatus, and a computed tomography apparatus. 7. The processing board according to claim 1, wherein the controller is further configured to control the communication interface to transmit, to the other processing hoard, a firmware update request of the processing board, along with the ID value that is set. 8. The processing board according to claim 7, wherein the communication interface is further configured to receive, from the other processing board, a firmware update file of the processing board, based on the firmware update request that is transmitted. 9. A medical diagnostic apparatus comprising: a processing board mounted on the medical diagnostic apparatus, the processing board being configured to: sense, as an output value, at least one selected from the group consisting of an installation position, an installation direction, and a power consumption of the processing board;store one or more identifier (ID) values for identifying the processing board in the medical diagnosis apparatus, along with one or more output values of a sensor, the one or mor... | SAMSUNG ELECTRONICS CO., LTD. | A61B006/04 G05B019/02 G05B019/042 A61B006/00 A61B005/055 A61B008/00
US09983053 | B2 | 2018-05-29 | Optical sensor device | 2016-01-21 | An optical sensor device includes a light emitter for emitting, to a living body, lights having two wavelengths and blinking at a predetermined frequency, and a light receiver for receiving the lights from the living body. The light receiver outputs first and second detection signals corresponding to the respective wavelengths. A filter circuit extracts, from the first and second detection signals, modulation signals that are obtained with amplitude modulation of signals of the predetermined frequency. The modulation signals are amplified by a post-amplifier and are taken into an arithmetic processing unit after being converted to digital signals by an AD converter. The arithmetic processing unit calculates DC components and AC components of the first and second detection signals by employing the modulation signals converted the digital signals. | 1. An optical sensor device, comprising: a light emitter configured to periodically emit light at a predetermined frequency toward a measurement target;a light receiver configured to detect at least one of reflected light and diffused light from the measurement target, and to output a modulated electrical signal based on the detected light;a filter circuit configured to filter the modulated electrical signal at the predetermined frequency and to output a modulation signal;an amplifier configured to amplify the modulation signal;an AD converter configured to convert the modulation signal to a digital signal;an arithmetic processing unit coupled to the AD converter and configured to calculate a DC component and an AC component of the electrical signal based on the digital signal; anda separation circuit including the filter circuit and configured to separate the modulation signal and the DC component of the electrical signal,wherein the AD converter is further configured to convert the DC component of the electrical signal to a DC component digital signal, andwherein the arithmetic processing unit is further configured to convert the modulation signal and the DC component digital signal based on an amplification factor of the amplifier, such that an amplitude ratio between the modulation signal and the DC component digital signal reaches a state before the amplification by the amplifier, and to calculate the AC component of the electrical signal based on the modulation signal. 2. The optical sensor device according to claim 1, wherein the arithmetic processing unit is further configured to calculate the DC component attributable to extraneous light noise based on a signal contained in the DC component of the electrical signal and obtained during an off-period of the light emitter. 3. An optical sensor device, comprising: a light emitter configured to emit light toward a measurement target;a light receiver configured to detect at least one of reflected light and diffused light from the measurement target, and output an electrical signal based on the detected light;a separation circuit configured to separate the electrical signal into a DC component and an AC component;an amplification circuit configured to amplify the DC component and the AC component, at first and second amplification factors, respectively;an AD converter configured to convert the DC component and the AC component to first and second digital signals, respectively; andan arithmetic processing unit configured to restore the DC component and the AC component of the electrical signal by converting the DC component and the AC component output as the first and second digital signals such that an amplitude ratio between the DC component and the AC component output as the digital signals obtains a state before the amplification by the amplification circuit. 4. An optical sensing method, comprising: emitting light, by a light emitter, at a predetermined frequency towards a measurement ta... | MURATA MANUFACTURING CO., LTD. | H03F003/08 G01J001/44 G01J003/427 G01N021/25 G01N021/31 A61B005/1455 A61B005/024 A61B005/00
US10403401 | B2 | 2019-09-03 | Medical apparatus with selectively enabled features | 2015-11-18 | A medical system includes a medical apparatus, a computer, a user input device, and at least one feature in communication with and controlled by the computer. The computer is in communication with the user input device, which is configured and arranged to allow a user to purchase the use of the feature. The computer is configured to enable the use of the feature after the user purchases use of the feature. | 1. A medical system comprising: a medical patient support apparatus, the medical patient support apparatus forming a patient support with a patient support surface for supporting a patient thereon and including a control system configured to control said patient support, said control system including a computer, and the patient support having a plurality of inactive hardware features, the inactive hardware features not having resident software;a user input device in communication with the computer;andthe plurality of inactive hardware features in communication with the computer, the user input device configured and arranged to allow a user to select one of the plurality of inactive hardware features as a user selected feature and seek authorization for the use of the user selected feature of the patient support, and after receiving authorization the computer configured to download the software appropriate for enabling the user selected feature and enable the use of the user selected feature after the software is downloaded to change the user selected feature to an active feature, and the computer disabling the active feature when the user selected feature is no longer authorized. 2. The patient support apparatus according to claim 1, wherein the user input device is configured and arranged to allow a user to make payment, for example, through a credit card, or to initiate billing to the user or to a third party. 3. The patient support apparatus according to claim 1, wherein the inactive features comprise at least one feature selected from the group consisting of a vital signs sensor, a load cell for monitoring patient activity, a bed exit load cell, a pressure sensor for monitoring interface pressure at a support surface of the patient support, a moisture sensor, a temperature sensor, an optical sensor, a sensor for predictive analytics, and a depth sensor for measuring envelopment into the support surface of the patient support. 4. The patient support apparatus according to claim 1, wherein the inactive features include an entertainment device. 5. The patient support apparatus according to claim 4, wherein the entertainment device comprises a device selected from the group consisting of a TV, a tablet, a phone, an electronic game, a noise cancellation device, speakers, and a held-held music player. 6. The patient support apparatus according to claim 1, wherein the inactive features include a treatment device or a diagnostic device. 7. The patient support apparatus according to claim 6, wherein the plurality of inactive features includes a treatment device, the treatment device comprising a device selected from the group consisting of a turning device, a deep vein thrombosis device (DVT device), a percussion and/or vibration device, a temperature management device, a vacuum assisted device, and a hyperbaric device for delivering oxygen. 8. The medical system according to claim 1: wherein the user input device is configured and arranged to ... | Stryker Corporation | G06F019/00 A61B005/0402 A61H007/00 G16H040/63 A61G007/018 A61H023/02 A61H023/00 A61G007/05 G07F017/00 A61B005/00 A61G007/012 A61G007/005 A61B005/0205 A61H009/00
US09330240 | B2 | 2016-05-03 | Method and system of identifying bodily imbalances | 2013-09-09 | In one embodiment, a method of identifying bodily imbalances comprises presenting a mind map to a subject; asking the subject a question based on the mind map; performing muscle response testing; and, identifying the imbalance. The mind map may be presented in printed form, using an electronic means, or presented verbally. The mind map further comprises six (6) categories: pathogenic imbalances, structural imbalances, nutritional imbalances, circuitry imbalances, toxic imbalances, and energetic imbalances; and a series of subcategories within each category. | 1. A method of identifying bodily imbalances using a mind map, the method comprising: presenting a mind map to a subject, wherein the mind map comprises one or more categories of health and wellness and one or more bodily imbalances;asking the subject one or more questions based on the mind map;performing muscle response testing on the subject after each question has been asked; andbased upon the subject's muscle response, identifying the corresponding bodily imbalance. 2. The method of claim 1, wherein the mind map is presented in printed form. 3. The method of claim 1, wherein the mind map is presented using an electronic means. 4. The method of claim 1, wherein the mind map is presented verbally. 5. The mind map of claim 1 comprising at least six categories for identification. 6. The categories of claim 5 comprising pathogenic imbalances, structural imbalances, nutritional imbalances, circuitry imbalances, toxic imbalances, and energetic imbalances. 7. The printed form of claim 2, comprising tabbed, or otherwise noted, pages in one or more books and/or binders. 8. The electronic means of claim 3, comprising a computing device, personal computers, tablets, smartphones, or PDAs. 9. A method of identifying bodily imbalances using a mind map, the method comprising: presenting an interactive mind map to a subject via a computing device, wherein the mind map comprises one or more categories of health and wellness and one or more bodily imbalances;asking the subject one or more questions based on the mind map;performing muscle response testing on the subject after each question has been asked;interacting with the mind map on the computing device based upon the subject's muscle response; andidentifying the corresponding bodily imbalance. 10. The interactive mind map of claim 9 comprising at least one link from a general category to a series of subcategories. 11. The at least one link of claim 10 comprising a shortcut to a local file/folder or a hyperlink to an online webpage, graphic, or chart. 12. The method of claim 9, wherein the computing device comprises a personal computer, tablet, smartphone, or PDA. 13. The method of claim 9, wherein the one or more questions elicit a “yes” or “no” response from the subject. 14. The method of claim 9, wherein interacting with the mind map comprises selecting, highlighting, clicking, or otherwise activating the underlying link to the subcategory. 15. The at least one category of claim 10 comprising pathogenic imbalances, structural imbalances, nutritional imbalances, circuitry imbalances, toxic imbalances, and/or energetic imbalances. 16. A method of identifying bodily imbalances using a mind map, the method comprising: presenting an interactive mind map to a subject via a computing device, wherein the mind map comprises one or more illustrated categories of health and wellness and one or more bodily imbalances;asking the subject a series of questions based on the one or more categories in the mi... |  | A61B005/00 G06F019/00 A61B005/11
US20160135768 | A1 | 2016-05-19 | SYSTEMS AND METHODS FOR DISPLAYING A PHYSIOLOGIC WAVEFORM | 2014-11-17 | A method and system for displaying a physiologic waveform. The method and system acquire positron emission tomography (PET) coincidence event data of an object of interest. The method and system further select a subset of the PET coincidence event data corresponding to a time window and apply a multivariate data analysis technique to the subset of the PET coincidence event data. The method and system also generate a physiologic waveform based on the multivariate data analysis, and display the physiologic waveform on a display. | 1. A method for displaying a physiologic waveform, said method comprising: acquiring positron emission tomography (PET) coincidence event data of an object of interest;selecting a subset of the PET coincidence event data corresponding to a time window;applying a multivariate data analysis technique to the subset of the PET coincidence event data;generating a physiologic waveform based on the multivariate data analysis; anddisplaying the physiologic waveform on a display. 2. The method of claim 1, wherein the multivariate data analysis technique further comprises applying a principal component analysis (PCA) to the subset of the PET coincidence event data. 3. The method of claim 1, further comprising selecting a new subset of the PET coincidence event data corresponding to an adjusted time window; repeat the applying operation based on the new subset of the PET coincidence event data;generating an updated physiologic waveform based on the multivariate data analysis from the new subset and includes the physiologic waveform;displaying the updated physiologic waveform. 4. The method of claim 1, wherein the physiologic waveform is based on a principal component derived from the multivariate data analysis technique. 5. The method of claim 1, further comprising calculating a metric for each PC, wherein the multivariate data analysis technique generates a plurality of PC for the subset of the PET coincidence event data. 6. The method of claim 5, further comprising selecting one of the PC based on the metric value relative to the other PC metric values, wherein the physiologic waveform is based on the selected PC. 7. The method of claim 1, wherein the physiologic waveform corresponds to at least one of respiratory movement or cardiovascular movement of the object of interest. 8. The method of claim 1, further comprising continually adjusting the time window to include a larger subset of the PET coincidence data relative to the subset; and repeating the applying, generating and displaying operations based on each adjusted time window. 9. The method of claim 1, wherein the time window is based on an anatomical range of the patient corresponding to the physiologic movement. 10. The method of claim 1, further comprising adjusting an acquisition time based on the physiologic waveform. 11. The method of claim 1, wherein the acquiring operation is based on a continuous table motion (CTM) acquisition during a PET scan. 12. The method of claim 11, wherein a velocity of a motorized table during the CTM is adjusted based on the physiologic waveform. 13. A positron emission tomography (PET) imaging system comprising: a data acquisition controller configured to acquire PET coincidence event data from a detector ring assembly;a multivariate data analysis module (MDAM) communicatively coupled to the data acquisition controller, the MDAM configured to select a subset of the PET coincidence event data corresponding to a time window;apply a multivariate data an... |  | A61B006/00 A61B005/00 A61B005/113 A61B006/04 A61B006/03 A61B005/0205
US10582878 | B2 | 2020-03-10 | Systems and methods for providing visual feedback of touch panel input during magnetic resonance imaging | 2014-05-09 | Systems and methods of are disclosed for providing visual feedback to a subject during magnetic resonance imaging, where the visual feedback is associated with input provided by the subject to a magnetic resonance compatible touch panel. A video camera is employed to record video images of the interaction between the subject and the touch panel, and the video images are processed to generate a real-time video signal including a rendering of the input provided to the touch panel and the interaction between the subject's hands and the touch panel. The real-time video signal is provided to the subject as visual feedback, and is displayed within a time duration that is sufficiently fast to avoid the detection of the visual feedback as an error signal with the subject's brain in relation to the sense of proprioception. A measurement of the force applied to the touch panel by the subject may be recorded and employed when rendering the real-time video. The systems and methods may be employed for a wide range of diagnostic and therapeutic procedures involving magnetic resonance imaging. | 1. A method of providing visual feedback to a subject during magnetic resonance imaging, wherein the visual feedback is associated with input provided to a touch panel, the method comprising: recording input provided by the subject to a magnetic resonance imaging compatible touch panel while the subject is positioned within a magnetic resonance imaging scanner;recording, with a video camera, video images of the interaction between the subject and the touch panel while the subject resides within the magnetic resonance imaging scanner, such that the video images include one or more hands of the subject;processing the video images and the input provided to the touch panel to generate a real-time video signal comprising: a rendering of the input provided to the touch panel by the subject; andan image of the one or more hands, showing a real-time position of the one or more hands relative to the input provided to the touch panel; anddisplaying the real-time video signal on a display device viewable by the subject in real time while the subject resides within the magnetic resonance imaging scanner, thereby providing the subject with visual feedback of the real-time position of the one or more hands relative to the input provided to the touch panel. 2. The method according to claim 1 further comprising recording one or more magnetic resonance images of the subject while recording the input from the subject. 3. The method according to claim 1 further comprising: receiving, in real-time, in addition to the video images, a measurement of a force applied by the subject to the touch panel; andrendering the input in the real-time video signal according to the measured force. 4. The method according to claim 3 wherein the thickness of a line drawn by the subject is rendered in real-time according to the measured force. 5. The method according to claim 3 wherein the duration during which a given pixel is displayed is associated with the force applied to the touch panel at a location of the given pixel. 6. The method according to claim 3 further comprising: recording the force as a function of time, thereby obtaining a time-dependent measured force; andprocessing the time-dependent measured force to compute one or more measures. 7. The method according to claim 6 wherein the one or more measures include one or more of: a total non-contact time, a total contact time, a percentage of contact vs. non-contact time, and an average applied force when entering input. 8. The method according to claim 6 wherein the time-dependent measured force is employed as a covariate of interest for modelling brain activity based on functional magnetic resonance images. 9. The method according to claim 1 wherein the input is provided by a stylus held by the subject, and wherein the video signal further comprises an image of the stylus showing a real-time position of the stylus relative to the touch panel. 10. The method according to claim 9 further comprising: processing the... | SUNNYBROOK RESEARCH INSTITUTE BAYCREST CENTRE FOR GERIATRIC CARE | A61B005/00 A61B005/055 G01R033/48 G01R033/28 G16H040/63
US09339663 | B2 | 2016-05-17 | Wearable cardiac defibrillator system emitting CPR prompts | 2014-10-30 | A wearable cardiac defibrillator (“WCD”) system may include a support structure that a patient can wear, an energy storage module that can store an electrical charge, and a discharge circuit that can discharge the electrical charge through the patient so as to shock him or her, while the patient is wearing the support structure. Embodiments may actively take into account bystanders, both to protect them from an inadvertent shock, and also to enlist their help. In some embodiments the WCD system includes a speaker system that transmits a sound designed to assist a bystander to perform CPR. Optionally CPR chest compressions received by the patient can be further detected, and feedback can be given. In embodiments, a WCD system may include a user interface that can be controlled to output CPR prompts tailored to a skill level of the bystander. | 1. A wearable cardiac defibrillator (WCD) system, comprising: a support structure configured to be worn by a patient;an energy storage module configured to be coupled to the support structure and to store an electrical charge;a discharge circuit configured to be coupled to the energy storage module and configured to discharge the electrical charge through the patient;one or more motion detectors coupled to the support structure, at least one of the motion detectors being coupled to the support structure at such a point as to be near a back of the patient while the support structure is worn by the patient; anda speaker system coupled to the support structure and configured to output a transmitted sound having substantially periodic contents designed to assist a bystander to perform CPR chest compressions on the patient while the support structure is worn by the patient, andin which the performed CPR chest compressions are detected by the one or more motion detectors andthe transmitted sound further communicates feedback regarding the detected CPR chest compressions. 2. The WCD system of claim 1, in which the transmitted sound further communicates a request to not remove the support structure from the patient. 3. The WCD system of claim 1, in which the feedback refers to a depth or to a rate of the detected CPR chest compressions. 4. The WCD system of claim 1, further comprising: a screen, andin which the feedback is given also by visual representations on the screen. 5. The WCD system of claim 1, in which at least one of the motion detectors is coupled to the support structure at such a point as to be near a sternum of the patient, while the support structure is worn by the patient. 6. The WCD system of claim 1, further comprising: an impedance sensing module, andin which the performed CPR chest compressions are detected by the impedance sensing module. 7. The WCD system of claim 1, in which the transmitted sound further communicates a request for the bystander to no longer touch the patient. 8. A wearable cardiac defibrillator (WCD) system, comprising: a support structure configured to be worn by a patient;an energy storage module configured to be coupled to the support structure and to store an electrical charge;a discharge circuit configured to be coupled to the energy storage module and configured to discharge the electrical charge through the patient while the support structure is worn by the patient; anda user interface configured to be coupled to the support structure and to receive a usage input from a bystander, the user interface including a speaker system configured to output a transmitted sound having substantially periodic contents designed to assist a bystander to perform CPR chest compressions on the patient, an aspect of the contents being different responsive to a received usage input. 9. The WCD system of claim 8, in which the usage input relates to a skill of the bystander. 10. The WCD system of claim 8, in which the tr... | West Affum Holdings Corp. | A61N001/00 A61N001/39 A61B005/00 A61B005/11 A61B005/053
US10168430 | B2 | 2019-01-01 | Wireless devices and systems for tracking patients and methods for using the like | 2015-11-13 | Disclosed are apparatuses, systems, and methods for tracking patients that suffer from dementia. The disclosed apparatus is a wearable device capable of micro-tracking through Bluetooth Low Energy technology and capable of macro-tracking through GPS technology. The device may additionally include sensors to monitor other information such as the health of the patient or the patient's surrounding environment. The disclosed systems utilize the disclosed device in an overall system for tracking patients. These systems teach how the device interacts with the other components of the system (e.g., signal beacons, wireless transmitters, a central processing unit, mobile computing devices) to provide an integrated system to tracking the location and monitoring the well being of the patient. Finally, methods for tracking patients that use the disclosed devices and systems are disclosed. | 1. A wearable device for tracking a patient comprising: a processor;a power source electronically connected to the processor;at least one antenna communicatively connected to the processor;a Bluetooth Low Energy (BLE) chip communicatively connected to the processor; anda Global Positioning System (GPS) chip communicatively connected to the processor; the GPS chip being configured to receive GPS location information;the at least one antenna being configured to transmit data to at least one signal beacon when the wearable device is inside of a predetermined distance from the at least one beacon, andthe at least one antenna being configured to transmit information based on the GPS location information when the wearable device is outside of a predetermined distance from the at least one beacon. 2. The wearable device of claim 1, further including at least one sensor to detect health-related information of the patient. 3. The wearable device of claim 2, wherein the at least one sensor includes at least one of a heart rate monitor, a thermometer, an accelerometer, and a gyroscope. 4. The wearable device of claim 1, further including at least one sensor to detect environmental conditions. 5. The wearable device of claim 4, wherein the at least one sensor includes at least one of a thermometer, a light sensor, and a microphone. 6. The wearable device of claim 1, further comprising a lock configured to block removal of the wearable device by the patient. 7. The wearable device of claim 6, wherein the lock is disengaged by at least one of at least one magnet, RFID technology, and a wireless signal. 8. The wearable device of claim 1, further comprising a notifier device configured to notify the patient of any alerts sent wirelessly to the tracking device. 9. The wearable device of claim 8, wherein the notifier is at least one of a light, a speaker, and a vibration motor. 10. A system for tracking patients comprising: a wearable tracking device configured to be worn by a patient wirelessly connected to a central processing unit, the device including: a processor;a power source electronically connected to the processor;at least one antenna communicatively connected to the processor;a BLE chip communicatively connected to the processor; anda GPS receiver communicatively connected to the processor;at least one beacon configured to transmit BLE signals to and receive BLE signals from the wearable tracking device within a predetermined range of the at least one beacon and configured to wirelessly transmit data received from the wearable tracking device to the central processing unit; andat least one computing device wirelessly configured to be connected to the central processing unit capable of displaying data received by the central processing unit the at least one computing device having a user interface. 11. The system for tracking patients of claim 10, wherein the wearable tracking device is wirelessly connected to the central processing unit. 12. ... | - | G01S019/14 G01S019/31 A61B005/0205 A61B005/00 H04W004/80 A61B005/11 G01S019/16
US09339232 | B2 | 2016-05-17 | Medical implantable lead with fixation detection | 2014-10-20 | A lead has a helix at a distal end of a rotatable shaft. Prior to implantation of the lead, the helix has a first configuration comprising an electrically active surface coated with an electrically insulative layer of a biologically dissolvable material to render the helix electrically inactive. Upon fixation of the helix into an organ, the helix has the first configuration and a pin is in electrical contact with the organ to detect proper fixation of the helix into the organ. Subsequent to fixation of the helix into the organ, the helix has a second configuration comprising an electrically active surface exposed upon dissolving of the electrically insulative layer of the biologically dissolvable material to render the helix electrically active. | 1. A medical implantable lead adapted to be implanted into a human or animal body to monitor and/or control an organ inside the body, the lead comprising: an intermediate flexible lead part; a connector at a proximal portion of the intermediate flexible lead part; a tubular header at a distal portion of the intermediate flexible lead part; a rotatable shaft disposed within the tubular header, wherein the rotatable shaft is extendably and retractably arranged relative to the tubular head; an electric conductor electrically coupling the connector to the rotatable shaft; a helix at a distal end of the rotatable shaft, wherein the helix is electrically coupled to the rotatable shaft, and wherein the helix is configured for fixation to the organ upon rotation and extension of the rotatable shaft; and a pin at a distal end of the rotatable shaft, wherein the pin is electrically coupled to the rotatable shaft and is electrically coupled to the helix, and wherein the pin is configured for electrical contact to the organ upon rotation and extension of the rotatable shaft; wherein the pin, the rotatable shaft and the helix are arranged in a unipolar configuration; wherein prior to implantation of the lead, the helix having a first configuration comprising an electrically active surface coated with an electrically insulative layer of a biologically dissolvable material to render the helix electrically inactive; wherein upon fixation of the helix into the organ, the helix having the first configuration and the pin being in electrical contact with the organ to detect proper fixation of the helix into the organ; and wherein subsequent to fixation of the helix into the organ, the helix having a second configuration comprising the electrically active surface exposed upon dissolving of electrically insulative layer of the biologically dissolvable material to render the helix electrically active. 2. The medical implantable lead of claim 1, wherein the pin is coaxially arranged in relation to the helix and projects distally from the distal end of the rotatable shaft. 3. The medical implantable lead of claim 1, wherein the pin is separate from the helix. 4. The medical implantable lead of claim 1, wherein the pin is adapted to penetrate into tissue of the organ. 5. The medical implantable lead of claim 1, wherein a longitudinal length of the pin is less or equal to half a longitudinal length of the helix. 6. The medical implantable lead of claim 1, wherein the pin is adapted to penetrate into the tissue. 7. The medical implantable lead of claim 1, wherein the pin has a same electric potential as the helix. 8. The medical implantable lead of claim 1, wherein the pin is disposed centrally on the distal end of the rotatable shaft. | ST. JUDE MEDICAL AB | A61N001/00 A61B005/04 A61B005/00 A61N001/05
US10062958 | B2 | 2018-08-28 | Wearable electronic patch with antenna enhancement | 2014-11-21 | A wearable electronic patch with an enhanced radio antenna includes an antenna, radio circuitry, a base portion, a distal portion, and intermediate portion, and a spacer, configured to raise the antenna away from the base portion, and thus away from a wearer to improve radiation properties of the antenna. The spacer may be sized and shaped to expand from a compressed state, as it may be when the patch is packaged, to an expanded state that raises the antenna when it is applied to a wearer. Other aspects, embodiments, and features are also claimed and disclosed. | 1. A wearable electronic patch for biological monitoring of a wearer, comprising: an antenna; a base portion having a first surface configured to face the wearer and a second surface is opposite the first surface of the base portion; a distal portion spaced away from and opposed to the base portion and having a first surface configured to face the second surface of the base portion, wherein the antenna is secured to the distal portion, wherein the base portion is configured to be closer to the wearer than the distal portion when the wearable electronic patch is worn by the wearer; an intermediate portion joining the base portion and the distal portion; and a spacer configured to be disposed between the second surface of the base portion and the first surface of the distal portion when the wearable electronic patch is worn by a wearer, wherein the spacer is sized and shaped to expand from a compressed state for moving the antenna away from the base portion. 2. The wearable electronic patch of claim 1, wherein when compressed the spacer maintains a space between the base portion and the antenna. 3. The wearable electronic patch of claim 1, wherein the spacer is formed of a foam material configured to expand after the base portion is secured to the wearer. 4. The wearable electronic patch of claim 1, wherein the spacer has a dielectric constant of approximately 1.0. 5. The wearable electronic patch of claim 1, wherein the base portion, the distal portion and the intermediate portion form a continuous printed circuit board. 6. The wearable electronic patch of claim 1, wherein a first end of the intermediate portion attached to the distal portion is folded over a second end of the intermediate portion attached to the base portion. 7. The wearable electronic patch of claim 1, further comprising an adhesive layer securing at least one of the base portion to the spacer and the spacer to the distal portion. 8. The wearable electronic patch of claim 1, wherein the antenna is a circuit element printed onto the distal portion. 9. The wearable electronic patch of claim 1, further comprising radio circuitry coupled to the antenna, wherein the radio circuitry is disposed between the base portion and the distal portion. 10. The wearable electronic patch of claim 1, further comprising a sensor attached to at least one of the base portion and the distal portion, wherein the sensor is configured to measure one or more physical characteristics. 11. The wearable electronic patch of claim 1, further comprising a body detection circuit configured to detect when the wearable electronic patch is worn by the wearer. 12. The wearable electronic patch of claim 1, further comprising a battery at least partially disposed between the base portion and the distal portion. 13. The wearable electronic patch of claim 1, further comprising a removable packaging covering at least one of the base portion and the distal portion, wherein removing of the packaging releases th... | QUALCOMM Incorporated | H05K005/00 G11B033/02 H01Q001/27 A61B005/01 A61B005/024 A61B005/0402 A61B005/145 A61B005/00 H01Q009/04 H05K003/30 G06K019/07 G06K019/077 A61B005/021 A61B005/0478 A61B005/053 A61B005/1455 A61B005/1468 H01Q009/26
US10413271 | B2 | 2019-09-17 | Method, system and device for auscultation | 2015-11-17 | The present disclosure provides a medical/healthcare system which comprises an electronic stethoscope and a medical host under a medical/healthcare environment. The electronic stethoscope comprises an auscultation module for recording a sound from internal body of a patient to form an auscultation data, a reader for obtaining a patient information from the patient, a processor module to receive and associate the patient information to the auscultation data to form an auscultation file, a storage for storing the auscultation file, a display module to display the auscultation file for selection, a control input interface for selecting the auscultation file for playback, an output module for playing the auscultation file by outputting the sound corresponding to the auscultation file when selected, and a power module to provide electricity to all the others. | 1. An electronic stethoscope, comprising: a reader module for obtaining a patient information from a patient;as chestpiece for obtaining sounds from the patient;a control input interface for receiving a user identification (ID) and a user password;a processor module electrically connected to the reader module, the chestpiece, and the control input interface for receiving and storing the patient information, the user ID and the user password, forming an auscultation data according to the sounds, associating the patient information to an auscultation data to form an auscultation file and performing user verification according to the user ID and the user password; andan output module electrically connected to the processor module for outputting the sounds according to the auscultation file,wherein the control input interface is further configured for sending a playback signal to the processor module, the processor module is further configured for sending an outputting signal to the output module based on the playback signal, the output module is further configured for playing the auscultation file when receiving the outputting signal,the control input interface is further configured for receiving an input corresponding to a body area image of the patient and a plurality of positions within the body area image, and sending a selection signal to the processor module based on the input the processor module is further configured for generating an auscultation position data based on the selection signal, the auscultation file comprises the auscultation position data,the processor module is further configured for storing a plurality of playback modes, the output module is further configured for outputting the sounds under one of the playback modes, and the playback modes have different sound frequency ranges, andthe playback signal is associated with the user ID and the user password, the body area image of the patient, the plurality of positions within the body area image, the selection signal, the auscultation position data and the playback modes. 2. The electronic stethoscope according to claim 1, further comprising a display module configured for displaying a playlist containing the auscultation file. 3. The electronic stethoscope according to claim 1, wherein the chestpiece comprises a microphone and a diaphragm. 4. The electronic stethoscope according to claim 1, herein the patient information comprises a patient ID, a medical history, a date of birth, a height, a weight, a blood type, a gender, an age, an emergency contact, or any combination thereof. 5. The electronic stethoscope according to claim 1, wherein the reader module is a barcode reader configured for reading at least one of an one-dimensional (1D) barcode and a two-dimensional (2D) barcode. 6. The electronic stethoscope according to claim 1, wherein the output module is a phone jack. 7. The electronic stethoscope according to claim 1, wherein the output module is further configure... | iMEDI PLUS Inc. | A61B007/04
US09345487 | B2 | 2016-05-24 | Precision bone drill and method of use | 2014-02-04 | Disclosed is a drilling device and a method for the drilling of human or animal bone tissue. The method comprises the following steps; (a) providing a drill having an electrically conductive bone cutter and an electrically conducting nosepiece assembly spaced apart from the electrically conductive bone cutter and acting as a counter electrode; (b) placing the electrically conductive bone cutter and nosepiece assembly into contact with the bone tissue to be drilled; (c) measuring the electrical impedance between the bone cutter tip and the nosepiece or the current through the cutter motor; (d) commencing drilling and moving the bone cutter toward the bone surface; (e) starting a step counter when the impedance reduces or the cutter motor current increases; (f) incrementing the step counter with every step of the vertical drive motor until the step count reaches the prespecified value thereby controlling the depth of the hole drilled. | 1. A method of drilling through bone comprising the steps of: (i) providing a drilling device having an electrically conductive bone cutter and an electrically conducting nosepiece assembly spaced from the conductive bone cutter, and acting as a counter electrode;(ii) placing the conductive bone cutter and nosepiece assembly into contact with the bone tissue to be drilled; and(iii) commencing the drilling of the bone tissue; said drilling device further comprising:(a) a housing that includes a drilling assembly and an electrically conductive nosepiece assembly to be placed in contact with bone tissue to be drilled;(b) a bone cutter rotated by an electric motor which is mobile in translation in the housing and wherein the bone cutter is electrically conductive and is adapted to be moved into and out of the bone tissue being drilled;(c) a printed circuit board control module electrically connected to the drilling assembly for controlling the translation of the bone cutter in the housing and into and out of the bone tissue being drilled;(d) a sensor, electrically connected to the drilling assembly and to the control module, for detecting a change in an electrical impedance between the bone cutter and the surface of the bone tissue being drilled; wherein upon detection, by the sensor, of a change in the electrical impedance corresponding to a transition between air (no load; high electrical impedance) and bone tissue surface, sends a signal to the control module begins incrementing a step counter; and(e) wherein the direction of the vertical drive motor is automatically reversed when the step counter reaches a prespecified count, thereby withdrawing the bone cutter from the bone tissue being drilled, and thereby controlling the depth of the drilled hole. 2. The method of claim 1, wherein the bone cutter is moved toward and away from the bone by a vertical drive motor that comprises a stepper motor. 3. The method of claim 2, wherein the stepper motor further includes a stepping power supply and settable counter control. 4. The method of claim 3, wherein the counter control starts counting the steps once a pre-set trigger impedance is sensed between the bone cutter and the electrically conductive impedance sensing electrode or the nosepiece of the housing. 5. The method of claim 4, wherein the drilling device further includes a vertical drive motor step counter pre-set to stop the forward motion of the bone cutter after reaching the desired count, and thus the desired bone cutter depth. 6. The method of claim 1, wherein the drilling device further comprises a replaceable bone cutter and nosepiece assembly that are attached to the drill device to provide a disposable drilling and sensing drill-stabilizing unit. 7. The method of claim 1, wherein the drilling device further comprises a drill electrical current load monitor for depth control of the bone cutter. 8. The method of claim 1, wherein the drilling device further comprises a forward motion ... | Path Scientific, LLC | A61B017/00 A61B017/16 A61B005/053 A61B017/20 A61B018/00
US09336594 | B2 | 2016-05-10 | Cardiac pulse rate estimation from source video data | 2014-03-07 | What is disclosed is a system and method for estimating cardiac pulse rate from a video of a subject being monitored for cardiac function. In one embodiment, batches of overlapping image frames are continuously received and processed by isolating regions of exposed skin. Pixels of the isolated regions are processed to obtain a time-series signal per region and a physiological signal is extracted from each region's time-series signals. The physiological signal is processed to obtain a cardiac pulse rate for each region. The cardiac pulse rate for each region is compared to a last good cardiac pulse rate from a previous batch to obtain a difference. If the difference exceeds a threshold, the cardiac pulse rate is discarded. Otherwise, it is retained. Once all the regions have been processed, the retained cardiac pulse rate with a minimum difference becomes the good cardiac pulse rate for comparison on a next iteration. | 1. A method for continuous estimation of cardiac pulse rate from video of a subject being monitored for cardiac function in a non-contact, remote sensing environment, the method comprising: repeatedly receiving batches of time-sequential image frames captured by at least one video imaging device of a subject being monitored for cardiac function; andprocessing successive batches of image frames, comprising:A) isolating, in said image frames, at least two regions of exposed skin of said subject where a signal corresponding to said subject's cardiac function was registered by at least one imaging channel of said video imaging device;B) for each of said isolated regions, performing: (i) processing pixels of said isolated region to obtain a time-series signal for this region;(ii) extracting, from said time-series signals, a physiological signal corresponding to said subject's cardiac function;(iii) processing said physiological signal to obtain a cardiac pulse rate for this region;(iv) comparing said cardiac pulse rate for this region to a last good cardiac pulse rate for said subject determined by having processed a previous batch to obtain a difference therebetween, on a first iteration said comparison being performed using an initial cardiac pulse rate determined for said subject; and(v) in response to said difference exceeding a threshold level, discarding this region's cardiac pulse rate, otherwise retaining this region's cardiac pulse rate; andC) in response to having performed steps A and B, determining that a retained cardiac pulse rate with a minimum difference between a last good cardiac pulse rate for a previous batch to be a next good cardiac pulse rate for a current batch, said next good cardiac pulse rate being used for comparison purposes on a next iteration. 2. The method of claim 1, wherein said video imaging device is any of: a color video camera, a monochrome video camera, an infrared video camera, a multispectral video imaging device, a hyperspectral video camera, and a hybrid camera comprising any combination hereof. 3. The method of claim 1, wherein pixels associated with said regions of exposed skin are isolated in said image frames using any of: pixel classification, object identification, thoracic region recognition, color, texture, spatial features, spectral information, pattern recognition, face detection, facial recognition, and a user input. 4. The method of claim 1, wherein said regions of exposed skin may or may not partially overlap. 5. The method of claim 1, wherein said batch of image frames partially temporally overlaps with the image frames of a previous batch. 6. The method of claim 1, wherein said threshold used for determining whether a difference between this region's cardiac pulse rate and the last good cardiac pulse rate for said subject is a function of any of: an age of said subject, a health of said subject, and a state of the subject's cardiac function. 7. The method of claim 1, further comprising fi... | Xerox Corporation | G06K009/00 G06T007/00 A61B005/00 A61B005/024
US09333071 | B2 | 2016-05-10 | Breast implant with regionalized analyte sensors and internal power source | 2014-06-11 | Breast implants including sensor modules and related methods are described herein. Breast implants include those with: a shell configured to be substantially filled with a viscous material; a plurality of projections extending from an external surface of the shell, the projections forming a plurality of compartments adjacent to the external surface of the shell; at least one fluid-permeable cover attached to the projections, the cover completely enveloping the shell and the plurality of projections; a plurality of sensor modules attached to the shell, each of the sensor modules oriented to detect one or more analytes in a fluid within one of the plurality of compartments, wherein each of the plurality of sensor modules includes a unique identifier; and at least one power source operably attached to the plurality of sensor modules. | 1. A method of monitoring information from a breast implant, comprising: sending, from a remote device, a query signal to at least one transmission unit attached to a breast implant in vivo, the at least one transmission unit attached to at least two sensor modules attached to the breast implant with a distance between the at least two sensor modules;receiving, from a remote device, a response signal from the at least one transmission unit attached to the breast implant, the response signal including positional information from each of the at least two sensor modules;processing, in a computing device, the response signal to identify information from the at least two sensor modules; andidentifying, for each of the at least two sensor modules, a detection result and a position on the breast implant. 2. The method of claim 1, wherein the detection result for each of the one or more sensor modules includes a positive result, a negative result, or a null result. 3. The method of claim 1, comprising: initiating a display of the detection result and the position on the breast implant for at least one of the sensor modules. 4. The method of claim 1, comprising: initiating a graphic display of the breast implant, the graphic display including a position for each of the sensor modules and the detection result for at least one of the sensor modules. 5. A method of monitoring information from a breast implant, comprising: sending a signal from at least one transmission unit attached to a breast implant in vivo, the signal including information regarding at least two sensor modules attached to the breast implant with a distance between the at least two sensor modules, the at least one transmission unit attached to the at least two sensor modules;receiving, at a remote device, the signal from the at least one transmission unit attached to the breast implant;processing, in a computing device, the signal to identify information from the at least two sensor modules; andidentifying, for each of the at least two sensor modules, a detection result and a position on the breast implant. 6. The method of claim 5, wherein the detection result for each of the one or more sensor modules includes a positive result, a negative result, or a null result. 7. The method of claim 5, comprising: initiating a display of the detection result and the position on the breast implant for at least one of the sensor modules. 8. The method of claim 5, comprising: initiating a graphic display of the breast implant, the graphic display including a position for each of the sensor modules and the detection result for at least one of the sensor modules. | Elwha LLC | A61F002/12 G01N033/574 G01N033/52 A61B005/00 A61B005/145 A61B005/07 A61B005/1473
US09554752 | B2 | 2017-01-31 | Skeletal measuring means | 2016-01-27 | An integrated imaging system is invented for creating an optimal imaging session by importing information in real time from several sources and using that information to automatically and continuously adjust the parameters of the imaging session so as to create the optimal session for the prescribed testing session | 1. A system measuring skeletal joint motion in a subject comprising: a) a motion device means adapted and configured to continuously move a joint of the subject, the motion device means comprising: a platform base means, anda motion platform means further comprising a static platform means connected to an upper surface of the platform base means, a movable platform means connected to at least one of the static platform means or an upper surface of the platform base means, wherein the static platform means is adjacent the movable platform means wherein movement of the movable platform means is achieved in operation by a motor in communication with the moveable platform means;b) an imaging device in communication with the motion device means adapted and configured to obtain imaging data; andc) a computing system adapted and configured to analyze the obtained imaging data to generate an instruction and then communicate the instruction to at least one of the motion device means and the imaging device wherein the computing system is configurable to receive information from at least one of the motion device means and the imaging device during an imaging session;analyze the received information during the imaging session;determine whether a target anatomy is within a field of view during the imaging session;evaluate a position of the imaging device, a position of the motion device means, a radiographic imaging technique, and a geometric configuration of a collimator;determine whether a change to one or more of a position of the imaging device, a position of the motion device means, a radiographic imaging technique, and a geometric configuration of a collimator will bring the target anatomy within the field of view; andinstruct at least one of the motion device means and imaging device to change an imaging environment automatically and continuously during the imaging session to maintain the target anatomy within a field of view by one or more of changing the position of the imaging device, changing the position of the motion device means, changing the radiographic imaging technique, and changing the geometric configuration of the collimator. 2. The system of claim 1, wherein the imaging device is an X-ray tube and image intensifier with dosage control. 3. The system of claim 1, wherein the imaging device is a magnetic resonance scanner. 4. The system of claim 1, wherein the platform is a laterally moveable platform. 5. The system of claim 1, wherein the movable platform means is situated on a support means which lies on the upper surface of the platform base means. 6. The system of claim 5, wherein the imaging device is an X-ray tube and image intensifier with dosage control. 7. The system of claim 6, wherein the imaging device is a magnetic resonance scanner. 8. The system of claim 1, further comprising a control arm means for driving movement of the moveable platform means. 9. The system of claim 1 further comprising a processing system. 10. T... | ORTHO KINEMATICS, INC. | G06T007/00 A61B005/00 A61B005/11 G06T007/20 A61B005/055 A61B006/04 A61B006/00
US09355214 | B2 | 2016-05-31 | False clinical alarm rate reduction | 2014-06-13 | A clinical facility communication network connects a notification system to a plurality of patient monitors and to a plurality of clinician communication devices. Each of the patient monitors operate to detect physiological characteristics associated with a patient and to send event messages to the notification system. The notification system operates to receive the event messages and to determine which of the event messages include information of sufficient clinical significance to warrant generating and sending an alarm to a clinician. The event messages that are determine not to include clinically significant information are ignored. | 1. A method for identifying a clinically significant event, comprising: setting a value in a virtual call point for each of a clinically significant event duration, a clinical event count period and a clinical event count threshold;receiving at the virtual call point a plurality of clinical events associated with a patient the duration of at least some of which are detected to be less than the value set for the clinically significant event duration, and maintaining a current count of the clinical events having a duration that is less than the value set for the clinically significant event duration; anddetermining that the current count of clinical events having a duration that is less than the value set for the clinically significant event duration is equal to the count threshold value prior to the event count period expiring and determining that the patient is experiencing a clinically significant event. 2. The method of claim 1, further comprising the virtual call point deactivating an alarm suspend function that when active operates to prevent the generation of an alarm when the virtual call point detects a clinical event having a duration that is less than the value set for the clinically significant event duration. 3. The method of claim 2, further comprising the virtual call point generating and sending an alarm to a clinician that includes an indication that the patient is experiencing a clinically significant event. 4. The method of claim 1, wherein the clinically significant event duration value defines a minimum period of time that a clinical event is active to be considered clinically significant. 5. The method of claim 1, wherein the event duration value and the clinical event count period value is a predetermined period of time. 6. The method of claim 1, wherein the clinical event count threshold value is an integer value. 7. The method of claim 1, wherein the virtual call point is comprised of one or more virtual call point instances. 8. The method of claim 7, wherein each of the one or more virtual call point instances is comprised of a clinical event detection function, an alarm suspend function, a clinical event counting function and an alarm generation function. 9. The method of claim 1, wherein each one of the plurality of clinical events is received from a patient monitoring device and is comprised of clinical event information relating to the physiological state of a patient. 10. The method of claim 9, wherein the clinical event information is any one or more of a heart rate, blood pressure, respiratory rate, blood oxygen, and body temperature. 11. An event notification system, comprising: a virtual call point having a clinical event detector, a clinical event counter, and alarm generation logic;wherein the virtual call point receives a plurality of clinical events associated with a patient, the clinical event detector determining that a least some of the clinical events having a duration that is less than a pre-set... | Globe Star, Inc. | G08C019/00 G08B021/00 G08C019/16 A61B005/04 A61H001/00 A61B017/00 G06F019/00 A61B005/00
US20160148531 | A1 | 2016-05-26 | SYSTEMS AND METHODS FOR COORDINATING MUSCULOSKELETAL AND CARDIOVASCULAR OR CEREBROVASCULAR HEMODYNAMICS | 2014-11-25 | Described herein are systems and methods for favorably coordinating a timing relationship between a musculoskeletal activity cycle and a cardiac cycle of a user. A method may include repetitively detecting a signal that correlates to a blood volume in the user; determining an actual value of the signal that varies with the timing relationship; computing a trend of the actual value of the signal; and adjusting the movement guidance based on the trend of the actual value. A system may include a prompt device configured to provide recurrently a movement guidance to the user for guiding performance of the rhythmic musculoskeletal activity; a sensor configured to provide a signal that correlates to a blood volume in the user; and a processor configured to determine an actual value of the signal that varies with the timing relationship and to adjust the movement guidance based on the trend of the actual value. | 1. A method for favorably coordinating a timing relationship between an MSKC of a rhythmic musculoskeletal activity of a user and a CC of the user, the method comprising: recurrently providing a movement guidance from a prompt device to the user for guiding performance of the rhythmic musculoskeletal activity; andrepetitively: detecting a signal, using a sensor, that correlates to a cyclically-varying arterial blood volume in a tissue of the user;determining an actual value of a measured characteristic of the signal that varies with the timing relationship between the MSKC and the CC of the user;computing a trend of the actual value of the measured characteristic using a processor; andadjusting the movement guidance based on the trend of the actual value so as to cause the actual value of the measured characteristic to approach a relative preferred value of the measured characteristic. 2. The method of claim 1, wherein the movement guidance includes at least one of a recurrent audible, visual, or tactile prompt. 3. The method of claim 1, wherein the detecting step comprises using as the sensor technology one of photoplethysmography, impedance plethysmography, laser-Doppler blood flow, acoustic sensing, and arterial tonometry. 4. The method of claim 1, wherein the measured characteristic of the signal that varies with the timing relationship between the MSKC and the CC of the user includes at least one of a pulse amplitude, a measure of relative peak to valley signal waveform curvature, a measure of signal waveform peak curvature, a measure of signal waveform valley curvature, a measure of signal waveform complexity, and a measure of an asymmetry of the signal waveform. 5. The method of claim 1, wherein the relative preferred value of the measured characteristic is one of a threshold crossing of an increasing trend, a threshold crossing of a decreasing trend, a local maximum, and a local minimum of the trend of the actual value of the measured characteristic. 6. The method of claim 2, wherein a heart rate of the user is substantially an integer multiple of the prompt rate. 7. The method of claim 1, further comprising: detecting, using one or more sensors, signals that correlate to a heart rate of the user and a musculoskeletal activity cycle rate (MSKR) of the user;determining, using the processor, the heart rate of the user and an MSKR of the user by processing the one or more signals;specifying a maximum allowable absolute difference between the heart rate and the MSKR; andexecuting, using the processor, steps of claim 3 only when an absolute value of a difference between the MSKR and heart rate is less than, or less than or equal to, a specified allowable difference. 8. The method of claim 7, further comprising: specifying a target MSKR; andrecurrently providing a prompt from the prompt device when the absolute value of the difference between the MSKR and the target MSKR is greater, or greater than or equal to, the specified allowable di... |  | G09B019/00 A61B005/0476 A61B005/00 A61B005/0205
US20160133160 | A1 | 2016-05-12 | Rescue Performance Metric | 2015-12-28 | An external defibrillator system, the system comprising: one or more sensors configured to detect one or more parameters associated with performance of CPR; one or more wearable computing devices configured to provide feedback to a user about CPR performed by the user based on the one or more parameters associated with the performed CPR; and an external defibrillator comprising: a memory configured to store instructions; and a processor to execute the instructions to perform an operation comprising: analyzing the one or more parameters to determine a CPR performance metric indicative of an overall performance of the user. | 1. An external defibrillator system, the system comprising: one or more sensors configured to detect one or more parameters associated with performance of CPR;one or more wearable computing devices configured to provide feedback to a user about CPR performed by the user based on the one or more parameters associated with the performed CPR; andan external defibrillator comprising: a memory configured to store instructions; anda processor to execute the instructions to perform an operation comprising: analyzing the one or more parameters to determine a CPR performance metric indicative of an overall performance of the user. 2. The system of claim 1, wherein the one or more wearable computing devices include at least one of the following: a tablet computing device, a headset, wearable glasses, and a smartphone. 3. The system of claim 1, wherein at least one of the one or more wearable computing devices directs feedback to one predetermined user. 4. The system of claim 1, wherein the one or more wearable computing devices includes one or more of a watch or wearable glasses. 5. The system of claim 4, wherein the wearable computing device includes wearable glasses having one or more lenses. 6. The system of claim 5, wherein the visual summary is displayed on at least one lens of the wearable glasses. 7. The system of claim 1, wherein generating the CPR performance metric includes generating data reflecting measurements of two or more actions performed on the victim. 8. The system of claim 1, wherein the CPR performance metric is based, at least in part, on weighted values applied to the one or more parameters. 9. The system of claim 8, wherein the weighted values are applied to at least one of the following: rate of chest compressions, depth of chest compressions, and fraction based on a guideline. 10. The system of claim 1, wherein the wearable computing device is configured to provide a visual summary including the CPR performance metric within five minutes of cessation of the CPR by the rescuer. 11. The system of claim 1, wherein the CPR performance metric is based on a function that weights each of the parameters according to weights determined using a neural network. 12. The system of claim 1, wherein the CPR performance metric is based on a function that weights each of the parameters according to weights determined using a linear regression technique. 13. The system of claim 1, wherein the wearable computing device is configured to receive input data from the user. 14. The system of claim 1, comprising a heads-up device to provide feedback to a user about CPR performed by the user based on the one or more parameters associated with the performed CPR. 15. An external defibrillator system, the system comprising: one or more sensors configured to detect one or more parameters associated with performance of CPR;a heads-up device configured to provide feedback to a user about CPR performed by the user based on the one or more parameters... |  | G09B023/28 A61B005/0402 A61N001/39
US20160143571 | A1 | 2016-05-26 | METHOD AND SYSTEM FOR DETERMINING PSYCHOLOGICAL DISORDER CONDITION IN A PERSON AND PROVIDING ASSISTANCE THEREFOR | 2015-02-10 | This technology relates to a method and system for determining psychological disorder condition in a person and providing assistance to the person with the psychological disorder. One or more sensors are placed on the person to monitor the activities of the person. The one or more sensors provide sensor data associated with the monitored activities of the person to a computing unit. The computing unit detects the activities and identifies a recurring activity. For the recurring activity, the computing unit determines probability parameters and based on the probability parameters the probability of the psychological condition is determined. For providing assistance to the person with psychological disorder condition, the computing compares the activity with the disorder activities stored in the computing unit. If a match is found, then one or more indications are provided to the person to indicate that the activity has already been performed. | 1. A method for determining psychological disorder condition in a person, the method comprising: receiving, by a computing unit, sensor data at predefined time intervals from one or more sensors;identifying, by the computing unit, one or more activities performed by the person at each predefined time interval based on the sensor data;identifying, by the computing unit, at least one recurring activity in the one or more activities;determining, by the computing unit, a probability value and a weighed ratio for each of one or more probability parameters associated with the at least one recurring activity;generating, by the computing unit, one or more weighted probability parameters based on the probability value and the weighed ratio; anddetermining, by the computing unit, probability of the psychological disorder condition in the person based on each of the one or more weighted probability parameters. 2. The method as claimed in claim 1, wherein the one or more sensors are placed on the person. 3. The method as claimed in claim 1, wherein the sensor data is at least one of location of the person, one or more actions performed by the person and one or more objects in the location. 4. The method as claimed in claim 1, wherein the at least one recurring activity is identified by comparing the one or more activities with each other. 5. The method as claimed in claim 1, wherein identifying the one or more activities comprises: correlating, by the computing unit, each of the sensor data;comparing, by the computing unit, the correlated data with a list of one or more predefined activity data; andidentifying, by the computing unit, the one or more activities corresponding to the matched one or more predefined activity data. 6. The method as claimed in claim 1, wherein the one or more probability parameters are similarity between the one or more activities, frequency of the recurring activities, one or more non-similar activities performed between the recurring activities and similarity in location of the person while performing the recurring activities. 7. The method as claimed in claim 1 further comprising generating, by the computing unit, a report for providing information associated with the one or more activities and the probability of the psychological disorder condition for each of the one or more activities. 8. A method of providing real-time assistance to a person with psychological disorder condition, the method comprising: receiving, by a computing unit, sensor data from one or more sensors;identifying, by the computing unit, one or more activities performed by the person based on the sensor data;comparing, by the computing unit, the one or more activities with a list of one or more predefined disorder activities;providing, by the computing unit, one or more indications in real-time for assisting the person with psychological disorder based on the comparison. 9. The method as claimed in claim 8, wherein the one or more sensors are placed... |  | A61B005/16 A61B005/0488 A61B005/11
US09351645 | B2 | 2016-05-31 | Photoacoustic measurement apparatus and probe unit for use therewith | 2013-10-25 | A photoacoustic measurement apparatus which includes a diffuser plate for diffusing laser light inputted from the upstream side of an optical system, a lens system for focusing the laser light diffused by the diffuser plate, and an optical fiber cable which includes an optical fiber having a core and cladding structure and a light energy resistant structure at an end portion on the side of the lens system, and is disposed such that the laser light focused by the lens is inputted from one end, in which the light projection unit projects the laser light guided by the optical fiber cable as the measuring light. In a photoacoustic measurement performed by guiding the measuring light using an optical fiber, high energy measuring light can be guided while inhibiting end face damage of the optical fiber. | 1. A photoacoustic measurement apparatus, comprising: a diffuser plate which diffuses measuring light inputted from the upstream side of an optical system;a lens system which focuses the measuring light diffused by the diffuser plate;an optical fiber cable which includes an optical fiber having a core and cladding structure and a light energy resistant structure at an end portion on the side of the lens system, and is disposed such that the measuring light focused by the lens system is inputted from one end;a light projection unit which projects the measuring light guided by the optical fiber cable into a subject;an electroacoustic transducer unit which detects a photoacoustic wave generated in the subject by the projection of the measuring light and converts the photoacoustic wave into an electrical signal; anda signal measurement unit which performs a signal measurement based on the electrical signal,wherein the signal measurement unit includes an acoustic image generation unit which generates a photoacoustic image with respect to the electrical signal of the photoacoustic wave. 2. The photoacoustic measurement apparatus of claim 1, wherein: the lens system focuses the measuring light such that a minimum beam diameter of the measuring light is not less than ½ of a core diameter of the optical fiber; andthe optical fiber cable is disposed such that the measuring light is inputted with a beam diameter of the measuring light being not less than ½ of the core diameter of the optical fiber. 3. The photoacoustic measurement apparatus of claim 2, wherein a divergence angle φ of the measuring light when incident on the diffuser plate, a diffusion angle θ of the diffuser plate, a focal length f of the lens system, and a core diameter d of the optical fiber satisfy a formula (1) given below:d2≤2.4·f·tan((ϕ2)2+(θ2)2).(1) 4. The photoacoustic measurement apparatus of claim 1, wherein the lens system has a focal length of 4 to 40 mm. 5. The photoacoustic measurement apparatus of claim 1, wherein: the diffuser plate is a holographic diffuser; andthe holographic diffuser has a diffusion angle of 0.2 to 2.0°. 6. The photoacoustic measurement apparatus of claim 1, wherein: the diffuser plate is a light shaping diffuser in which tiny concave lenses are arranged randomly on one surface of a substrate; andthe light shaping diffuser has a diffusion angle of 0.2 to 2.0°. 7. The photoacoustic measurement apparatus of claim 1, wherein the optical fiber has a core diameter of 100 to 600 μm. 8. The photoacoustic measurement apparatus of claim 1, wherein the optical fiber cable further comprises a covering member which covers the optical fiber such that a side surface portion of the optical fiber near the end face is exposed. 9. The photoacoustic measurement apparatus of claim 8, wherein: the optical fiber has the core and cladding structure in which the cladding is comparable in damage threshold energy density to the core; andthe covering member covers the optic... | FUJIFILM Corporation | A61B008/08 A61B005/00 A61B001/07
US09345394 | B2 | 2016-05-24 | Medical apparatus | 2014-04-04 | A medical apparatus includes: a storage section that stores three-dimensional image information in a subject, the three-dimensional image information being acquired in advance; an image pickup unit that acquires an optical image in the subject; a luminal organ extracting section that extracts image information of a three-dimensional shape of a specific luminal organ in the subject; a position correspondence control section that sets correspondence between two coordinate systems; a feature information acquisition section that acquires feature information in the image information of the extracted three-dimensional shape of the specific luminal organ; an associated image generating section that generates an image in which the feature information is associated with the image information of the three-dimensional shape of the specific luminal organ; a control section that performs control to display a piece of the feature information in the vicinity of the position information of the image pickup unit, etc. | 1. A medical apparatus comprising: a storage device configured to store three-dimensional image information of a subject, the three-dimensional image information being acquired in advance;an image pickup device configured to acquire an optical image inside the subject; anda processor comprising hardware, wherein the processor is configured to implement: a luminal organ extracting section configured to extract image information of a three-dimensional shape of a specific luminal organ in the subject from the three-dimensional image information;a position correspondence control section configured to generate position correspondence information in an initialization state in which correspondence is set between position information of a predetermined position in the specific luminal organ in a first coordinate system and position information of a distal end portion that is set to the predetermined position in a second coordinate system, thereby setting correspondence between position information in the image information of the three-dimensional shape of the specific luminal organ in the first coordinate system and position information of the image pickup device in the second coordinate system;a feature information acquisition section configured to acquire, as feature information, position information related to a branch region where a lumen branches in the image information of the three-dimensional shape of the specific luminal organ extracted by the luminal organ extracting section in a region where the image pickup device is inserted based on the position information of the image pickup device inserted into the specific luminal organ in the first coordinate system, the position information of the image pickup device being caused to have correspondence by the position correspondence control section based on the position information of the image pickup device in the second coordinate system;an associated image generating section configured to generate an image in which the position information related to the branch region where the lumen branches is associated with the image information of the three-dimensional shape of the specific luminal organ by superposedly displaying the position information related to the branch region where the lumen branches on the image information of the three-dimensional shape of the specific luminal organ;a determination section configured to determine whether the position information of the image pickup device in the first coordinate system, which is caused to have correspondence by the position correspondence control section based on the position information of the image pickup device in the second coordinate system, indicates the image pickup device in the first coordinate system is positioned inside or outside the specific luminal organ in the image information of the three-dimensional shape in the first coordinate system; anda control section configured to perform control to display, on a display, a piece of the posit... | OLYMPUS CORPORATION | A61B001/05 A61B001/00 A61B001/07 A61B005/00 A61B001/267
US09326713 | B2 | 2016-05-03 | System and method for detecting and measuring ethyl alcohol in the blood of a motorized vehicle driver transdermally and non-invasively in the presence of interferents | 2012-08-03 | This invention relates to a system and method that can quickly and accurately detect and measure ethyl alcohol in the blood of a motorized vehicle driver transdermally and non-invasively within minutes of ethanol consumption and in the presence of interferents. The system includes an array of sensors embedded into the steering mechanism of a motorized vehicle, a data base of patterns produced through empirical testing of various analytes in various combinations and concentrations, neural net based pattern recognition algorithms to ascertain the driver's transdermal alcohol concentration and a database derived from human testing, correlating the driver's transdermal alcohol concentration with the driver's blood alcohol concentration. The detection system is integrated with a motor vehicle decision module which can prevent the operation of a motorized vehicle by a driver whose BAC exceeds a preset limit. | 1. A method for detecting and measuring bloodstream ethanol comprising the steps of: providing an array of sensors adapted to process the pre-perspiratory vapor emanating from a person, the pre-perspiratory vapor containing one or more analytes;each individual sensor in the array of sensors producing a sensor output signal in response to the one or more analytes present; at least one chemicapacitor sensor producing an identifiable sensor output signal to an analyte of ethanol; andat least two chemicapacitor sensors, of the same technology as said at least one chemicapacitor sensor, in the sensor array producing different sensor output signals in response to an identical analyte;storing sensor output signal array patterns corresponding to one or more output signals of the array of sensors as stored output signal array patterns in a pattern recognition library;receiving individual sensor output signals from the array of sensors, at a computing device, as an array of received individual sensor output signals exhibiting a received individual sensor output signal array pattern;producing a computer comparison between the received individual sensor output signal array pattern to the stored output signal array patterns stored in the pattern recognition library; anddetermining a transdermal alcohol content measurement as a function of the comparison. 2. The method of claim 1 further comprising the steps of: determining the transdermal alcohol content measurement through neural network analysis of the comparison. 3. The method of claim 1, wherein the output of the sensor is a function of change of capacitance in the presence of a target analyte. 4. The method of claim 1, wherein at least one sensor in the array of sensors is a sensor for ambient humidity. 5. The method of claim 4, wherein at least one sensor in the array is a sensor for ambient temperature. 6. The method of claim 1, wherein the transdermal alcohol content is determined by determining that a portion of the output of the array of sensors corresponds to interferents and subtracting the portion of the signal attributable to interferents from the output of the array of sensors. 7. The method of claim 1, wherein the method is performed within a vehicle, the vehicle comprising an operational control, at least the sensors being disposed in the operational control, the operational control being required to come in contact with a user of the vehicle for operation of the vehicle, and further comprising a vehicle decision module, the method further comprising the step of receiving a signal at the vehicle decision module, and the vehicle decision module preventing movement of the vehicle as a function of the transdermal alcohol content measurement. 8. The method of claim 7 wherein said operational control is a steering wheel. 9. The method of claim 8 wherein said pattern recognition library and computing device are disposed on said operational control. 10. The method of claim 8 further compris... | Sober Steering Sensors Canada Inc. | A61B005/1477 A61B005/145 A61B005/00
US09341601 | B2 | 2016-05-17 | Elastography device and method | 2009-12-02 | An elastography device for quantitatively and/or qualitatively measuring the viscoelastic properties of any medium, includes a deformable tube for positioning the ultrasonic transducer located at the end of the deformable tube and for holding it so as to ensure, during at least one measurement, perpendicularity and contact between the ultrasonic transducer and the medium. | 1. An elastography device for quantitatively and/or qualitatively measuring the viscoelastic properties of any medium, said device comprising: a vibration generator configured to generate at least one low-frequency elastic wave;an ultrasonic transducer configured to observe a propagation of said at least one low-frequency elastic wave;a trigger configured to trigger at least one measurement,said at least one measurement being performed by generating the at least one low-frequency elastic wave in a medium, andemitting and acquiring simultaneously to the generation of the at least one low-frequency elastic wave high-frequency ultrasounds using said ultrasonic transducer to observe the propagation of said at least one low-frequency elastic wave;a deformable tube positioned between said vibration generator and said ultrasonic transducer, the deformable tube being freely movable to spatially move without angular limitation the ultrasonic transducer and having a ductility that allows the deformable tube to hold the ultrasonic transducer in position after movement of the deformable tube, wherein the deformable tube is configured to hold said ultrasonic transducer located at the end of said deformable tube so as to ensure, during at least one measurement, perpendicularity and contact between said ultrasonic transducer and the medium, anda support attached to the vibration generator and configured to spatially hold the ultrasonic transducer. 2. The device according to claim 1, comprising a spring system located between said ultrasonic transducer and said deformable tube. 3. The device according to claim 1, comprising at least one rod capable of transmitting the low-frequency elastic waves from said vibration generator to said ultrasonic transducer. 4. The device according to claim 1, comprising a gripping tube. 5. The device according to claim 1, wherein said vibration generator is formed by a vibrating table. 6. The device according to claim 1, wherein said vibration generator is formed by a loudspeaker-type secondary device. 7. The devices according to claim 1, wherein said device includes a fine-tuning system configured to rotationally and/or translationally adjust the position of said ultrasonic transducer. 8. The device according to claim 1, wherein said vibration generator is located at the end of said deformable tube where the ultrasonic transducer is located. 9. The device according to claim 1, wherein said vibration generator is located at the opposite end of the deformable tube where the ultrasonic transducer is located so that said vibration generator is directly connected to said opposite end of the deformable tube. 10. The device according to claim 1, wherein said device includes at least one high-frequency ultrasonic imaging array. 11. The device according to claim 1, wherein said device is coupled to an organic synchronization device. 12. The device according to claim 1, wherein said ultrasonic transducer has an active diameter o... | ECHOSENS | A61B008/14 G01N029/22 A61B005/00 A61B008/00 A61B008/08 G01N029/04
US09326682 | B2 | 2016-05-03 | Systems, processes and software arrangements for evaluating information associated with an anatomical structure by an optical coherence ranging technique | 2013-01-07 | Software systems, arrangements and processes for evaluating an image associated with at least one portion of an anatomical structure are provided. For example, first information associated with the at least one portion of the anatomical structure second information associated with the at least one portion of the anatomical structure can be received. Third information can be generated by determining a relationship between the first information and the second information. Further, the image can be evaluated using a predetermined pathological scoring criteria and the third information. | 1. A process for evaluating at least one image associated with at least one portion of an anatomical structure, comprising: receiving first information associated with the at least one portion of the anatomical structure;receiving second information associated with the at least one portion of the anatomical structure;with a computer processing device, generating third information by determining a relationship between the first information and the second information; andwith the computer processing device, evaluating the at least one image using a predetermined pathological scoring criteria and the third information, wherein the predetermined pathological scoring criteria includes at least one of (i) a surface maturation score measured by patterns of reflectance of nuclear density or (ii) a glandular irregularity score. 2. The process according to claim 1, wherein at least one of the first information or the second information is associated with a light remitted from the at least one portion. 3. The process according to claim 2, wherein the light is reflected from the at least one portion. 4. The process according to claim 2, wherein the light is fluorescence. 5. The process according to claim 1, wherein the at least one portion is provided in a living subject. 6. The process according to claim 1, wherein the at least one portion of the anatomical structure is situated on a microscope slide. 7. The process according to claim 6, wherein the slide is stained with at least one of Hematoxylin and Eosin, Masson's Trichrome, Papanicolaou's stain, Diff-Qwik, or Periodic Acid Shiff. 8. The process according to claim 1, wherein the first information and the second information are provided for approximately the same location of the at least one portion of the anatomical structure. 9. The process according to claim 1, wherein the third information is obtained based on physical and chemical structures associated with the first information and the second information. 10. The process according to claim 1, wherein the image is associated with a light remitted from the at least one portion. 11. The process according to claim 10, wherein the light is reflected from the at least one portion. 12. The process according to claim 10, wherein the light is fluorescence. 13. The process according to claim 10, wherein the at least one portion is provided in a living subject. 14. The process according to claim 1, wherein at least one of the first information or the second information is obtained by an optical coherence tomography system. 15. The process according to claim 1, wherein at least one of the first information or the second information is obtained by a spectrally encoded confocal microscopy system. 16. The process according to claim 1, wherein the anatomical structure resides below the skin. 17. The process according to claim 16, wherein at least one of the first information or the second information is obtained by a confocal microscopy system. 18... | The General Hospital Corporation | G06K009/00 A61B005/00 A61B010/02 G06T007/00 G06T011/00 A61B010/00 A61B017/00 G06F019/00
US09351141 | B2 | 2016-05-24 | Headset computer with handsfree emergency response | 2013-07-18 | Operating conditions for a headset computer are used to detect a situation where an emergency response mode should be activated. In emergency response mode, the headset computer may then broadcast location, live (real time) audio and video/data streams, and record other information concerning a possible theft and/or other situation of interest to law enforcement and/or security services. | 1. A method of controlling operation of a headset computer comprising: receiving a hands-free gesture input from an original user, exclusively indicative of an impending emergency situation;in response to the received certain input from the original user, activating an emergency response mode whereby standard headset computer functions remain active for use by either the original user or an alternate user, but the following actions are also performed:capturing at least one of an audio or video stream, the capturing occurring as a result of entering the emergency response mode, without requiring a user input specifically associated with capturing one or more of the audio or video stream;determining location information;identifying one or more wireless devices within a proximity of the headset computer, automatically pairing the headset computer with at least one of the identified one or more wireless devices, and copying personal information from paired ones of the identified one or more wireless devices; andforwarding at least one of the captured audio or video stream, the copied personal information, and the location information to at least one of a service provider, law enforcement agency, or security organization. 2. The method of claim 1, wherein the one or more functions further include comparing a captured voice audio data against one or more template voices to determine voice biometric identification data of the one or more persons in the vicinity. 3. The method of claim 1, further including comparing a captured image against one or more template images to determine image biometric identification data of one or more persons in the vicinity, and forwarding the determined biometric identification data. 4. The method of claim 1, wherein the one or more functions further include overriding any password changes made by the alternate user. 5. The method of claim 1, wherein the received certain input from the original user includes a spoken phrase, head movement, or hand gesture. 6. The method of claim 1, wherein the received certain input from the original user includes an image, sound, geographical position, general position, orientation, atmospheric condition, pressure, health condition, environment, energy, acceleration, altitude, motion, velocity, speed, or light including visible light, infrared light and ultraviolet light. 7. The method of claim 1, wherein the one or more functions are performed in a clandestine manner, such that the alternate user is not notified that the one or more functions are performed. 8. A device that controls operation of a headset computer comprising: a receiving unit that receives a hands-free gesture input from an original user, exclusively indicative of an impending emergency situation;a response activation unit configured to, in response to the received certain input from the original user, activate an emergency response mode whereby standard headset computer functions remain active, for use by either ... | Kopin Corporation | A61B005/00 G06F003/01 G06F003/042 H04M001/05 H04M001/60 H04M001/725 H04R005/033 H04W004/22 G06F003/03 G02B027/01
US09347932 | B2 | 2016-05-24 | Device and method for breath analysis using acoustic resonance flow rate | 2011-10-03 | The present invention describes a device and method for measuring flow rate from air flow using acoustic resonance. It is shown that the device can be used to measure breath flow rate and biomarker concentration within human breath. | 1. A device for measuring gas flow for chemical analysis, comprising: at least one acoustic generator;an acoustic chamber connected to said acoustic generator, wherein said acoustic chamber comprises at least one gas flow inlet and at least one gas flow outlet configured to produce a back-pressure dependent on gas flow;a chemical sensor coupled to the acoustic chamber, the chemical sensor configured to measure an analyte in the gas flow in response to the gas flow having an acoustic frequency in a specific acoustic frequency range; andan acoustic sensor coupled to the acoustic chamber, the acoustic sensor configured to measure the acoustic frequency produced by the gas flow and to provide an indication of the gas flow being in the specific acoustic frequency range. 2. The device of claim 1, wherein said acoustic generator comprises an acoustic flute, a single reed instrument, or a whistle. 3. The device of claim 1, wherein the chemical sensor comprises a nitric oxide (NO) sensor. 4. The device of claim 1, wherein the acoustic sensor is configured to measure the acoustic frequency of gas having a gas flow rate in the range of 2000 milliliter per minute to 3000 milliliter per minute. 5. The device of claim 1, wherein the acoustic sensor comprises at least one of a microphone and an A/D converter. 6. The device of claim 5, wherein the acoustic sensor is operably linked to a processor for measurement of air flow rate using acoustic resonance. 7. The device of claim 6, wherein said processor comprises at least one of a computer and a cell phone. 8. The device of claim 1, arranged to measure exhaled-breath from a patient. 9. A method for measuring exhaled breath air flow comprising: receiving exhaled breath flow into a flow sensor housing forming an acoustic chamber, the flow sensor housing including an acoustic generator connected to an inlet, the flow sensor housing, housing a chemical sensor;generating a back-pressure in the acoustic chamber;producing an acoustic signal in the acoustic chamber, wherein said acoustic signal is modulated by the back-pressure;detecting said acoustic signal with an acoustic sensor;analyzing said acoustic signal to determine if the exhaled breath flow is in a specific acoustic frequency range; andmeasuring an analyte in the exhaled breath flow via a chemical sensor, in response to a determination that the exhaled breath flow is in the specific acoustic frequency range. 10. The method of claim 9, where the chemical sensor comprises an NO sensor. 11. The method of claim 9, wherein analyzing the acoustic signal to determine if the exhaled breath flow is in a specific acoustic frequency range includes analyzing the acoustic signal via a cell phone. 12. The method of claim 9, wherein analyzing the acoustic signal to determine if the exhaled breath flow is in a specific acoustic frequency range includes analyzing the acoustic signal via a computer. 13. The method of claim 9, further comprising producing the acousti... | ARIZONA BOARD OF REGENTS, A BODY CORPORATE OF THE STATE OF ARIZONA ACTING FOR AND ON BEHALF OF ARIZONA STATE UNIVERSITY | G01N033/497 A61B005/087 A61B005/083
US09332928 | B2 | 2016-05-10 | Method and apparatus to synchronize a location determination in a structure with a characteristic of the structure | 2009-04-14 | An volume of a patient can be mapped with a system operable to identify a plurality of locations and save a plurality of locations of a mapping instrument. The mapping instrument can include one or more electrodes that can sense a voltage that can be correlated to a three dimensional location of the electrode at the time of the sensing or measurement. Therefore, a map of a volume can be determined based upon the sensing of the plurality of points without the use of other imaging devices. An implantable medical device can then be navigated relative to the mapping data. | 1. A method of synchronizing a position determination within a volume of a patient and a physiological cycle of the patient, comprising: determining at least a first portion and a second portion of a physiological cycle of the patient that relate to a surface position of a structure of the patient;collecting a plurality of map data points based on tracking a mapping instrument at a plurality of three-dimensional locations in the patient at least by detecting an injected current during the determined physiological cycle;classifying the collected plurality of map data points as in the first portion of the physiological cycle or the second portion of the physiological cycle, wherein a first sub-plurality of the collected plurality of map data points is classified as in the first portion of the physiological cycle and a second sub-plurality of the collected plurality of map data points is classified as in the second portion of the physiological cycle; andafter classifying the collected plurality of map data, only displaying a representation of a shape of the surface with one of: the first sub-plurality of collected and classified map data points that relates to a first shape of the surface of the structure during the first portion of the physiological cycle, orthe second sub-plurality of collected and classified map data points that relates to a second shape of the surface of the structure during the second portion of the physiological cycle. 2. The method of claim 1, further comprising: saving the classified collected plurality of map data; andsuccessively displaying with a display device for viewing by a user sequentially only the representation of the first sub-plurality of collected and classified map data points and then only the representation of the second sub-plurality of collected and classified map data points;wherein in the physiological cycle the first portion is immediately prior to the second portion. 3. The method of claim 1, wherein determining the physiological cycle includes determining the first portion of the physiological cycle and the second portion of the physiological cycle of a cardiac cycle and a respiratory cycle; wherein the cardiac cycle is determined with an electrocardiogram and the respiratory cycle is determined by measuring an impedance between a first electrode and a second electrode connected to the patient. 4. The method of claim 3, further comprising: identifying a plurality of portions of both of the first and the second portions of the physiological cycle. 5. The method of claim 1, wherein determining the physiological cycle includes connecting an electrocardiogram with the patient to measure the systole and diastole of the patient. 6. The method of claim 5, wherein displaying includes: displaying the map data classified as in systole and displaying the map data classified as diastole to illustrate motion of the heart, wherein the heart is the structure. 7. The method of claim 6, further comprising: deter... | Medtronic, Inc. | A61B005/05 A61B005/0402 A61B005/06 A61B005/0432 A61B005/044 A61B005/042 A61B005/053 A61B019/00 A61B005/0428 A61B008/06 A61B008/00 A61B017/00 A61B017/22
US09339557 | B2 | 2016-05-17 | Photosensitizer-metal nanoparticle complex and composition containing the complex for photodynamic therapy or diagnosis | 2009-10-08 | Provided are a photosensitizer-metal nanoparticle complex and a composition for photodynamic therapy or diagnosis having the same. The complex includes a photosensitizer, a metal nanoparticle, and a backbone linking the photosensitizer with the metal nanoparticle. The backbone has a polypeptide substrate capable of being specifically degraded by a protease. When the complex is administered to a patient, fluorescence and production of reactive oxygen species from the conjugated photosensitizers are inhibited in normal tissues due to the resonance energy transfer between the photosensitizer and metal nanoparticles, but in tumor tissues, fluorescence and production of reactive oxygen species from the released photosensitizers are activated, thereby effectively destroying the tumor tissues. In addition, the selective fluorescence in the tumor tissues can further improve accuracy of tumor diagnosis using the protease-activatable photosensitizer-metal nanoparticle complex. | 1. A photosensitizer-metal nanoparticle complex, which is represented by the following Formula (“Pro-Leu-Gly-Val-Arg-Gly-Lys” disclosed as SEQ ID NO: 17):wherein M is a metal nanoparticle, PS is a photosensitizer having a carboxyl group, andthe lysine (Lys) is linked with the photosensitizer by an amide bond. 2. The complex according to claim 1, wherein the photosensitizer is a porphyrin-based compound or a non-porphyrin compound in the form of a free base or metal complex. 3. The complex according to claim 2, wherein the photosensitizer is pyropheophorbide a. 4. The complex according to claim 1, wherein the metal nanoparticle has gold, silver, copper, platinum, palladium, nickel, iron, or a combination of at least two thereof. 5. The complex according to claim 1, wherein the metal nanoparticle is a gold nanoparticle. 6. The complex according to claim 1, wherein the metal nanoparticle has the form of a sphere, rod, pyramid, star, or core-shell. 7. A composition for photodynamic therapy or diagnosis, comprising a photosensitizer-metal nanoparticle complex according to claim 1 and a pharmaceutically available carrier. 8. A photosensitizer-metal nanoparticle complex, which is represented by Formula 1 (“Pro-Leu-Gly-Val-Arg-Gly” disclosed as SEQ ID NO: 15): | NATIONAL CANCER CENTER | A61K051/00 A61K049/00 A61B005/00 A61K039/00 C07D345/00 C07B047/00 C07D225/00 C07D295/00 A61K047/48 A61K041/00 C12N009/00
US09337831 | B2 | 2016-05-10 | Capacitive wire sensing for furniture | 2012-01-09 | A system and method for incorporating presence-sensing technology into furniture is provided. More particularly, the invention relates to incorporating a conductive medium into the periphery of a furniture item, such as an adjustable bed, for detecting the presence of a user or other body. The bed includes a mattress with a covering material and a tape edge surrounding the perimeter of the mattress. A capacitive wire is coupled the tape edge, and is adapted to have a voltage based on proximity of an object to the capacitive wire. The bed also includes a processor coupled to the capacitive wire. The processor receives information provided by the capacitive wire, determines when a change in voltage satisfies a threshold, and indicates whether presence is detected. Based on a determination of presence, or lack of presence, a variety of corresponding features of the adjustable bed may be activated. | 1. An adjustable bed, comprising: a mattress support that comprises a plurality of support panels, at least one of said support panels movable relative to the other ones of said support panels to thereby adjust the bed, wherein the adjustable bed comprises a head end, a foot end opposite and parallel to the head end, a first side perpendicular to the head end and the foot end, and a second side perpendicular to the head end and the foot end;a mattress resting on top of the mattress support, said mattress separate from the mattress support that the mattress is resting on top of, said mattress having a covering material disposed over at least a top surface of the mattress;at least one capacitive component coupled to the bed, said coupling to the bed comprising the at least one capacitive component on a perimeter of at least one of the plurality of support panels to provide a defined sensing area based on a position of the at least one capacitive component around the perimeter, said perimeter of at least one of the plurality of support panels comprising at least a portion of an outer edge of the first side, at least a portion of an outer edge of the second side, and at least a portion of an outer edge of one or more of the head end and the foot end, wherein the at least one capacitive component is adapted to have a voltage based on proximity of an object to the at least one capacitive component, and further wherein the at least one capacitive component is configured to detect user occupancy with respect to the mattress; anda processor coupled to the at least one capacitive component, the processor adapted to receive information provided by the at least one capacitive component and to determine that a change in voltage satisfies a threshold. 2. The bed of claim 1, wherein the at least one capacitive component is directly coupled to at least a portion of a perimeter of a top of at least one of the plurality of support panels. 3. The bed of claim 1, wherein the at least one capacitive component is directly coupled to at least a portion of a perimeter of a bottom of at least one of the plurality of support panels. 4. The bed of claim 1, wherein the at least one capacitive component is coupled to at least a portion of a tape edge surrounding a perimeter of the top surface of the mattress, the tape edge coupled to the covering material. 5. The bed of claim 1, wherein the at least one capacitive component is interwoven with the covering material. 6. The bed of claim 1, wherein the at least one capacitive component comprises a wire. 7. The bed of claim 1, wherein the at least one capacitive component is coupled to at least one spring inside the mattress. 8. The bed of claim 1, wherein the processor comprises a control enclosure coupled to a bottom of at least one of the plurality of support panels. 9. The bed of claim 1, wherein the processor receives information directly from the at least one capacitive component. 10. The bed of claim 1, wherein d... | L&P Property Management Company | G01R027/26 H03K017/955 G01N027/22 A61B005/11 A61B005/00
US09326684 | B2 | 2016-05-03 | Magnetic enhancement in determination of physiological blood parameters | 2011-11-08 | Systems and methods for measuring a physiological parameter of blood in a patient are provided herein. An example method includes establishing with a magnetic array, an first magnetic field along tissue of the patient inserted into the magnetic array and a second magnetic field perpendicular to the tissue, emitting optical signals into further tissue of the patient during at least a first alignment of the magnetic array around the tissue, detecting characteristics of the optical signals, and identifying a value of a physiological parameter based on at least the characteristics of the optical signals. | 1. A system for measuring a physiological parameter of blood in a patient, the system comprising: a magnetic array having an opening therein configured to have tissue of the patient inserted therein and simultaneously establish in the tissue an axial magnetic field and a radial magnetic field perpendicular to the axial magnetic field;an emitter configured to emit a first optical signal into the tissue during a first angular configuration between the emitter and the magnetic array, and emit a second optical signal into the tissue during a second angular configuration between the emitter and the magnetic array the second angular configuration being different from the first angular configuration;a detector configured to detect characteristics of the first optical signal and the second optical signal;a user interface; anda processing module configured to: calculate a differential of the characteristics of the first and the second optical signals;identify a value of a physiological parameter from the differential; anddisplay the physiological parameter on the user interface. 2. The system of claim 1, wherein the magnetic array comprises a central hole configured to have the tissue of the patent inserted therein, and wherein the axial magnetic field and the radial magnetic field are established within the central hole. 3. The system of claim 1, wherein the physiological parameter comprises a total hemoglobin parameter of the blood of the patient. 4. The system of claim 1, wherein the characteristics of the first optical signal and the second optical signal comprise an amplitude of the first optical signal and the second optical signal after scattering through the tissue. 5. The system of claim 1, wherein the magnetic array comprises a composite array of permanent magnets forming a central hole, wherein the permanent magnets are organized so both the axial magnetic field and the radial magnetic field are observed in the central hole. 6. The system of claim 1, wherein the tissue comprises a digit of the patent, and wherein the digit is inserted to a first distance into a central hole of the magnetic array for application of the axial magnetic field and the radial magnetic field at the first distance along the digit, and wherein the first optical signal and the second optical signal are each applied at a second distance along the digit. 7. The system of claim 1, wherein the first angular configuration between the emitter and the magnetic array is correlated to a first intensity of the first optical signal and a second intensity of the second optical signal, wherein the second angular configuration between the emitter and the magnetic array is correlated to a second intensity of the first optical signal and a first intensity of the second optical signal, and wherein the first intensity is greater than the second intensity. 8. The system of claim 1, wherein the emitter is configured to emit the first optical signal and the second optical signal over ... | Covidien LP | A61B005/1455 A61B005/00 A61B005/055
US09351666 | B2 | 2016-05-31 | Anthropometric measuring device | 2014-01-21 | A measuring device with a measuring tape and a one-piece handle. The handle includes a main body with a fixed reference, an anchor portion with a movable reference and a resilient element. The measuring tape is secured to the anchor portion. The resilient element allows the anchor portion to move with respect to the main body when the measuring tape is under the tension. The resilient element is configured so that the movable reference comes into alignment with the fixed reference when the measuring tape is under proper tension. The fixed reference may also provide a reference from which to take measurements. The resilient element may include a pair of nonlinear segments that straighten under tension. The main body may include a pair of fixed references disposed on opposite sides of the measuring tape, and the anchor portion may include a pair of corresponding movable references. | 1. An anthropometric measuring device for taking measurements under a defined force comprising: a one-piece handle including a fixed portion, a movable portion and a resilient portion joining said fixed portion to said movable portion, said fixed portion, said movable portion and said resilient portion being integrally formed, said handle including first and second fixed references associated with said fixed portion and a movable reference associated with said movable portion, said first fixed reference and said second fixed reference being disposed on opposite sides of said movable portion, said resilient portion including first and second resilient elements disposed toward opposite ends of said movable portion, said first and second resilient elements being configured to undergo elastic deformation that allows said movable reference to be moved into alignment with said first fixed reference and said second fixed reference under the defined force; anda measuring tape having at least one of graduations and markings, said measuring tape secured to said movable portion, wherein force applied to said measuring tape is transferred to said first and second resilient elements to cause elastic deformation of said resilient portion and movement of said movable portion with respect to said fixed portion, whereby said movable reference is movable into alignment with said first fixed reference and said second fixed reference by application of the defined force to said measuring tape in an appropriate direction. 2. The measuring device of claim 1 wherein said first and second fixed references provide references for determining proper application of force and a reference for taking a measurement based on said at least one of graduations and markings. 3. The measuring device of claim 1 wherein said first and second resilient elements having a non-linear shape in a direction associated with the defined force. 4. The measuring device of claim 3 wherein said movable portion includes an anchor bar with opposite ends, said resilient elements being joined to said anchor bar toward opposite ends, said resilient elements be spaced apart to accommodate said measuring tape therebetween. 5. The measuring device of claim 4 wherein said movable reference is a visual indication associated with said anchor bar. 6. The measuring device of claim 4 wherein said movable reference includes a first movable reference at one end of said anchor bar and a second movable reference at an opposite end of said anchor bar. 7. The measuring device of claim 6 wherein said fixed portion includes spaced apart first and second arms, said fixed portion including said first fixed reference on said first arm and said second fixed reference on said second arm. 8. The measuring device of claim 7 wherein said first fixed reference and said second fixed reference are disposed on opposite sides of said measuring tape. 9. The measuring device of claim 3 wherein said fixed portion, said movable p... | Wolverine World Wide, Inc. | G01B003/10 A61B005/107 A43D001/02 A61B005/00 G01B003/00
US09706951 | B2 | 2017-07-18 | Methods for assessing vaginal atrophy | 2016-01-25 | An array of methods for assessing vaginal atrophy are disclosed. The methods may be used alone or in combination with a treatment or as part of a kit. | 1. A method for identifying urogenital changes associated with vaginal atrophy, the method comprising: a. measuring pH at the labia majora, labia minora, introitus, or combinations thereof;b. determining brush sensitivity at the introitus;c. assessing glycogen amount at the introitus;d. analyzing epithelial cells sampled at the introitus;e. assessing protein amount at the introitus or the labia majora or the labia minora;f. assessing metabolites amounts at the introitus or the labia majora or the labia minora;g. assessing histamine amount at the introitus or the labia majora or the labia minora;h. transcriptomic heat mapping at the labia majora or the labia minora or the introitus; andi. evaluating gene expression at the labia majora using a gene probe set. 2. The method of claim 1, wherein the method of measuring pH further comprises using a pH probe. 3. The method of claim 1, wherein the changes in gene expression determined by the gene probe set are limited to gene expressions associated with collagen expression. 4. The method of claim 1, wherein the changes in gene expression determined by the gene probe set are limited to cell cycle progression. 5. The method of claim 1, wherein the method is used to assess the extent of vaginal atrophy prior and post a treatment regime. 6. The method of claim 5, wherein the treatment shows an effect of reducing the pH relative to atrophy by 0.1 to 2.0 pH units. 7. The method of claim 5, wherein the treatment shows an effect of reducing pH relative to atrophy by 0.3 units to 1.5 units. 8. The method of claim 5, wherein the treatment results in the introitus having a pH of 3.5 to 5.8 units. 9. The method of claim 5, wherein the treatment is a urogenitally or vaginally applied emollient comprised of estrogen, isoflavones, niacinamide, hyaluronin, panthenol, fatty oils, buffered acids, and mixtures thereof. 10. The method of claim 5, wherein the treatment is an oral or dermal delivery system comprised of estrogen. 11. The method of claim 1, wherein analyzing epithelial cells at the introitus further comprises measuring the abundance of rete pegs in the hymenal ring, wherein the method further comprises a treatment, and wherein the treatment shows an effect of increasing the rete abundance by 10 to 100%. 12. The method of claim 1, wherein analyzing epithelial cells at the introitus further comprises measuring the thickness of the hymenal epithelia, wherein the method further comprises a treatment, and wherein the treatment shows an effect of increasing the thickness relative to atrophy by 10 to 300%. | The Procter & Gamble Company | A61B005/1477 A61B005/145 G01N033/68 A61B005/00 G01N033/50
US10517275 | B2 | 2019-12-31 | Method, a device and a system for determining a state of an animal | 2014-06-16 | A device (5) attached to the neck (6) of an animal (2) comprises an accelerometer (17) which produces first and second signals indicative of the raised and lowered state of the head (7) of the animal (2) and movement of the animal (2). A microprocessor (20) in the device (5) processes the first and second signals to determine if the animal is ruminating, resting, feeding or in a highly active state during respective second predefined time periods of approximately 15 minutes duration. Data indicative of the ruminating, resting, feeding and the highly active state of the animal is stored by the microprocessor (20) in the device (5) and is periodically wirelessly communicated to a base station computer which further processes the data to determine various health states of the animal. | 1. A method for determining a state of an animal, the method comprising: providing a sensing means for sensing two conditions of the animal, a first one of the sensed conditions comprising an orientation of a head of the animal, and a second one of the sensed conditions comprising a level of activity of the animal, the sensing means comprising an accelerometer configured to produce signals indicative of acceleration to which the accelerometer is subjected along two axes perpendicular to each other,attaching the accelerometer to the animal to produce signals indicative of the first and second sensed conditions,providing a signal processor configured to read the signals from the accelerometer indicative of the first and second sensed conditions,operating the signal processorto read the signals from the accelerometer indicative of the first and second sensed conditions of the animal,to determine at least one state of the animal from a plurality of states as being ruminating in response tothe read signals indicative of the first one of the sensed conditions being indicative of the head of the animal being in a raised state,two or more frequency components of the read signals indicative of the second one of the sensed conditions lying within a predefined frequency range indicative of ruminating, andan average value of magnitude values of the frequency components of the read signals of the second one of the sensed conditions lying within the predefined frequency range indicative of ruminating being within a predefined range of magnitude values indicative of ruminating. 2. A method as claimed in claim 1 in which one of the determined states of the animal is feeding, and the state of the animal is determined as feeding in response to the read signal indicative of the first one of the sensed conditions being indicative of the head of the animal being in a lowered state, and the second one of the sensed conditions indicative of the level of activity of the animal being indicative of a low level of activity indicative of feeding. 3. A device for determining the state of an animal, the device comprising a sensing means for monitoring two conditions of the animal and for producing signals indicative of the sensed conditions, the sensing means comprising an accelerometer configured to produce signals indicative of acceleration to which the accelerometer is subjected along two axes perpendicular to each other, the accelerometer being configured for attaching to the animal to produce the signals indicative of the sensed conditions, and a signal processor configured to read the signals from the accelerometer, and to carry out the method as claimed in claim 1 for determining the state of the animal. 4. A system for determining a state of an animal, the system comprising a device as claimed in claim 3 for determining the state of an animal, and a remote computing means for receiving data communicated wirelessly by the device. 5. A method as claimed in claim 2 i... | DAIRYMASTER | A01K029/00 A01K011/00 A61B005/00 A61B005/11
US20160128647 | A1 | 2016-05-12 | Medical Device With Enhanced Viewing Mode | 2014-11-07 | A medical device includes: a processor; a display; and memory encoding instructions that, when executed by the processor, cause the processor to create an interface on the display, the interface including: a plurality of vital signs data positioned on the interface, with each vital sign being displayed in an enlarged font and being listed in order of importance; and a control for accessing other functionality of the medical device. | 1. A medical device comprising: a processor;a display; andmemory encoding instructions that, when executed by the processor, cause the processor to create an interface on the display, the interface including: a plurality of vital signs data positioned on the interface, with each vital sign being displayed in an enlarged font and being listed in order of importance; anda control for accessing other functionality of the medical device. 2. The medical device of claim 1, wherein the vital signs data is positioned vertically on the interface. 3. The medical device of claim 2, wherein the vital signs data is listed from most important to least important. 4. The medical device of claim 1, wherein a font size for one or more of the vital signs data is varied. 5. The medical device of claim 1, wherein a font color for one or more of the vital signs data is varied. 6. The medical device of claim 1, wherein only numerical data is displayed on the medical device. 7. The medical device of claim 1, wherein the interface is displayed as a lock screen. 8. The medical device of claim 7, wherein the lock screen is displayed automatically when the medical device is inactive for a period of time. 9. A medical device comprising: a processor;a display; andmemory encoding instructions that, when executed by the processor, cause the processor to create a lock screen interface on the display, the lock screen interface including: a plurality of vital signs data positioned on the lock screen interface, with each vital sign being displayed in an enlarged font and being listed in a vertical orientation in order of importance; anda control for unlocking the medical device. 10. The medical device of claim 9, wherein a font size for one or more of the vital signs data is varied. 11. The medical device of claim 10, wherein a font color for one or more of the vital signs data is varied. 12. The medical device of claim 11, wherein only numerical data is displayed on the medical device. 13. The medical device of claim 9, wherein a font color for one or more of the vital signs data is varied. 14. The medical device of claim 9, wherein only numerical data is displayed on the medical device. 15. The medical device of claim 9, wherein the lock screen interface is displayed automatically when the medical device is inactive for a period of time. 16. The medical device of claim 9, wherein the vital signs data is listed in two columns. 17. A medical device comprising: a processor;a display; andmemory encoding instructions that, when executed by the processor, cause the processor to create a lock screen interface on the display, the lock screen interface including: a plurality of vital signs data positioned on the lock screen interface, with each vital sign being displayed in an enlarged font and being listed in a vertical orientation in order of importance, wherein a font size for one or more of the vital signs data is varied, and wherein a font color for one or more of the... |  | A61B005/00
US20160140864 | A1 | 2016-05-19 | SYSTEM AND METHOD FOR MANAGING BINGE EATING DISORDERS | 2016-01-21 | Systems, methods, devices, and applications for managing a binge eating disorder (BED). A smartphone and server application may be provided for monitoring and treating BED. Ecological Momentary Assessment (EMA) may be used for automated and/or manual acquisition of data related to triggers associated with BED. Risk factors may be captured in real-time to provide data to both patients and clinicians. For manual data acquisition, the user may be prompted to input emotional states, urges to binge, eating behavior, and binge episodes. For automated acquisition, time, place (e.g. via GPS), medicine adherence, and physical activity may be recorded. The user may be provided with alerts and/or interventions when it is determined that they are at risk of binging. | 1. A system for managing a binge eating disorder, the system comprising: a first mobile device, comprising a processor, a memory, and a display, configured to receive input from a user regarding eating habits and psychological information;at least one peripheral device, operatively coupled to the first mobile device, configured to record data and transmit information to the first mobile device;the first mobile device further configured to transmit user information and the recorded data from the peripheral device, to a server; andthe first mobile device configured to generate and display, to the user, notifications regarding the users eating habits. 2. The system of claim 1, wherein the notifications are generated based on the input or the recorded data. 3. The system of claim 1, further comprising a clinical portal device configured to access the information stored in the server. 4. The system of claim 1, further comprising a second mobile device, configured to communicate with the first mobile device via a social networking application. 5. The system of claim 1, wherein the at least one peripheral device comprises an activity band, a web-connected scale, a smart pill bottle, a smart pill bottle cap, or a geolocation device. 6. The system of claim 1, wherein the notifications are generated based on an analysis of the input or the recorded data and other input or recorded data received from other users. 7. A mobile device for managing a binge eating disorder, comprising: a processor,a memory,a display, andan input interface;the input interface configured to receive user behavioral data and user context data;the user behavioral data including information regarding a mood, a physical state, an urge to binge, a binging episode, a meal, a social interaction, an interpersonal conflict, or a desire for guidance;the user context data including information about a user location, activity, body weight, body composition, or medication compliance; andwherein the device is configured to display an alert relating to the binge eating disorder based on the user behavioral data, the user context data, or both the user behavioral data and the user context data. 8. The mobile device of claim 7, wherein the alert comprises an intervention related to the behavioral data or the context data. 9. The mobile device of claim 7, wherein the input interface is configured to receive manual input of the behavioral data or the context data. 10. The mobile device of claim 7, wherein the input interface is configured to receive the context data from a sensor. 11. The mobile device of claim 10, wherein the sensor comprises an activity monitor, a scale, a body composition monitor, a medication monitoring device, a smart pill bottle, a smart pill bottle cap, or a geolocation device. 12. The mobile device of claim 10, wherein the input interface is configured to receive the context data automatically. 13. The mobile device of claim 7, further configured to display the al... | Drexel University | G09B019/00 A61B005/00 A61B005/11
US20160128583 | A1 | 2016-05-12 | SYSTEM AND APPARATUS COMPRISING A MULTISENSOR GUIDEWIRE FOR USE IN INTERVENTIONAL CARDIOLOGY | 2016-01-20 | A system and apparatus comprising a multisensor guidewire for use in interventional cardiology, e.g. for Transcatheter Valve Therapies (TVT), comprises a plurality of optical sensors for direct measurement of cardiovascular parameters, e.g. transvalvular blood pressure gradients and flow. A conventional outer coil contains a shaped core wire having a cross-section defining helical grooves extending along its length, which accommodate optical fibers and optical sensors within a diameter Dcore of the core wire. Advantageously, the diameter and material of the core wire provides the guidewire with sufficient stiffness for use as a support guidewire for valve replacement, e.g. Transcatheter Aortic Valve Implantation (TAVI), while accommodating multiple sensors and optical fibers within a guidewire of outside diameter ≦0.89 mm. An optical connector couples the guidewire to a control system. Optionally, the guidewire includes a contact force sensor; a pre-formed tip; and/or a separable micro-connector for proximal mounting of over-the-guidewire components. | 1. A multisensor guidewire for measuring blood pressure concurrently at multiple locations during a minimally invasive intravascular or cardiac intervention, comprising: a tubular covering layer comprising a flexible coil (coil), the coil having a length extending between a proximal end and a distal end, an outside diameter of ≦1 mm, a core wire extending within the coil from the proximal end to the distal end, and the distal end comprising a flexible distal tip;a plurality of optical sensors and a plurality of optical fibers; a sensor end of each optical fiber being attached and optically coupled to an individual one of the plurality of optical sensors;the core wire having an external surface with a cross-sectional profile defining a plurality of grooves extending along a length of the core wire, each groove accommodating an individual optical fiber and a respective optical sensor within a diameter Dcore of the core wire and providing a sensor arrangement with said plurality of optical sensors positioned at respective sensor locations within a distal end portion of the guidewire;a proximal end of each of the plurality of optical fibers being coupled to an optical input/output connector at the proximal end of the guidewire for connection to an optical control system; andthe plurality of optical sensors of the sensor arrangement including at least two optical pressure sensors at respective sensor locations spaced apart lengthwise along a length of said distal end portion. 2. The multisensor guidewire of claim 1, wherein the plurality of grooves are symmetrically spaced around the core wire. 3. The multisensor guidewire of claim 2, wherein the grooves are helical grooves. 4. The multisensor guidewire of claim 3, wherein the helical grooves have a pitch of at least 25 mm (1 inch). 5. The multisensor guidewire claim 1, wherein the grooves have a depth that accommodates each optical fiber and optical sensor within a respective groove of the core wire without protruding beyond the diameter Dcore of the core wire. 6. The multisensor guidewire of claim 5, wherein, at sensor positions in the distal end portion, the grooves are enlarged to accommodate optical sensors having an external diameter greater than that of the optical fibers. 7. The multisensor guidewire of claim 1, wherein surfaces of the core wire defining the grooves are radiused to no less than a minimum radius Rmin for formation of the grooves by a wire-drawing process. 8. The multisensor guidewire of claim 1, wherein each optical fiber is adhesively bonded to the core wire within its respective groove at a point adjacent the sensor location. 9. The multisensor guidewire of claim 1, wherein the coil has an outside diameter of ≦0.89 mm (≦0.035 inch). 10. The multisensor guidewire of claim 1, for use as a support guidewire for transcatheter valve replacement, wherein the core wire comprises a medical grade stainless steel alloy and the diameter Dcore of the core wire provides the guide... |  | A61B005/0215
US20160128591 | A1 | 2016-05-12 | MEASUREMENT SYSTEM OF RESPIRATION-RELATED SIGNALS AND METHOD OF OPERATING THE SAME | 2015-01-14 | A measurement system of respiration-related signals includes a plurality of electrodes, an ECG processing device, and an analysis device. The ECG processing device acquires a first analog differential signal and a second analog differential signal through the electrodes and generates a Lead-I ECG signal and a Lead-II ECG signal which is orthogonal to the Lead-I ECG signal after processing the two analog differential signals. The analysis device executes a trigonometric function to calculate angle change values of a cardiac axis based on the acquired Lead-I ECG signal and the Lead-II ECG signal, and further generate a cardiac axis angle change waveform according to the continuous angle change values. Accordingly, the cardiac axis angle change waveform is suitable for the prediction of respiratory conditions of the participant because of the high positive correlation between the frequency of the cardiac axis angle change waveform and the human respiratory frequency. | 1. A measurement system of respiration-related signals, comprising: a plurality of electrodes configured to measure a first analog differential signal and a second analog differential signal of a participant;an ECG processing device connected to the electrodes and configured to generate a Lead-I digital signal and a Lead-II digital signal according to the first analog differential signal and the second analog differential signal; an included angle being between the Lead-I digital signal and the Lead-II digital signal, and the included angle being greater than or equal to 75° and less than or equal to 105°;an analysis device connected to the ECG processing device and configured to execute a function conversion according to the Lead-I digital signal and the Lead-II digital signal, and calculate a plurality of angle change values of a cardiac axis; anda display device connected to the analysis device to generate an angle change waveform of a respiration-related signal according to the angle change values. 2. The measurement system of respiration-related signals in claim 1, wherein the included angle is 90°. 3. The measurement system of respiration-related signals in claim 1, wherein the analysis device is configured to execute the function conversion according to a ratio between the Lead-I digital signal and the Lead-II digital signal; wherein the function conversion is executed by a tangent function conversion of the trigonometric function conversion. 4. The measurement system of respiration-related signals in claim 3, wherein the Lead-I digital signal is generated by projecting an electrocardiac signal on a first vector, and the analysis device is configured to calculate the Lead-I digital signal to acquire a first projection; the Lead-II digital signal is generated by projecting the electrocardiac signal on a second vector, and the analysis device is configured to calculate the Lead-II digital signal to acquire a second projection; wherein the angle change value is equal to:Θ=tan-1firstprojection71secondprojection72when the first vector is orthogonal to the second vector. 5. The measurement system of respiration-related signals in claim 4, wherein the analysis device is configured to calculate the first projection and the second projection according to an amplitude area or a peak amplitude value of the Lead-I digital signal and the Lead-II digital signal. 6. The measurement system of respiration-related signals in claim 4, wherein the electrodes comprises a first electrode, a second electrode, and a third electrode, and the electrodes are configured on an inside surface of a physical measurement clothing. 7. The measurement system of respiration-related signals in claim 6, wherein the first electrode is configured at a location which is above a cardiac horizontal line and the right of a cardiac vertical line, wherein the cardiac horizontal line is horizontal through the participant's heart and the cardiac vertical line is verticall... |  | A61B005/04 A61B005/00 A61B005/0408 A61B005/0205 A61B005/044
US20160135688 | A1 | 2016-05-19 | OBJECT INFORMATION ACQUIRING APPARATUS | 2015-11-12 | Provided is an object information acquiring apparatus including: a holding member that has a curvature and holds an object; a conversion element that receives an acoustic wave propagating from the object via the holding member and outputs an electrical signal; a driving unit that moves at least one of the conversion element and the holding member so that an acoustic wave reception surface of the conversion element is not parallel to a tangential plane of the holding member at a position at which a normal line of the acoustic wave reception surface drawn from the conversion element crosses the holding member; and an information processor that acquires specific information inside the object using the electrical signal. | 1. An object information acquiring apparatus comprising: a holding member configured to have a curvature and hold an object;a conversion element configured to receive an acoustic wave propagating from the object via the holding member and output an electrical signal;a driving unit configured to move at least one of the conversion element and the holding member so that an acoustic wave reception surface of the conversion element is not parallel to a tangential plane of the holding member at a position at which a normal line of the acoustic wave reception surface drawn from the conversion element crosses the holding member; andan information processor configured to acquire specific information inside the object using the electrical signal. 2. The object information acquiring apparatus according to claim 1, wherein the driving unit moves at least one of the conversion element and the holding member so that the acoustic wave reception surface is inclined toward an outer side of the holding member. 3. The object information acquiring apparatus according to claim 1, further comprising: a probe in which a plurality of the conversion elements are arranged in directions of long and short sides thereof so that the acoustic wave reception surfaces are flush with each other. 4. The object information acquiring apparatus according to claim 3, wherein the driving unit moves the probe so that the acoustic wave reception surfaces are inclined about a rotation axis parallel to the long side. 5. The object information acquiring apparatus according to claim 1, wherein the information processor performs a conversion process on the specific information obtained by reconstructing the electrical signal according to an angle between the acoustic wave reception surface of the conversion element and a scanning plane when at least one of the conversion element and the holding member is driven by the driving unit. 6. The object information acquiring apparatus according to claim 1, wherein the driving unit moves the conversion element in relation to the holding member so that the normal line of the acoustic wave reception surface changes sequentially. |  | A61B005/00
US20160143514 | A1 | 2016-05-26 | SPHERICAL MECHANISM FOR MAGNETIC MANIPULATION | 2014-06-12 | A magnetic manipulation device (16) can include a housing (20), a spherical magnetic body (22) contained within the housing (20), and a plurality of sensors to detect the direction of the magnetic dipole of the spherical magnetic body (22). The spherical magnetic body (22) can be rotatable about a sphere axis of rotation which is (24a, 24b, 24c) can be in contact with the spherical magnetic body (22) to rotate the spherical magnetic body (22) about the sphere axis of rotation. | 1. A magnetic manipulation device, comprising: a housing;a spherical magnetic body contained within the housing, the spherical magnetic body being rotatable about a sphere axis of rotation which is omnidirectionally variable; anda plurality of rotators in contact with the spherical magnetic body and adapted to rotate the spherical magnetic body about the sphere axis of rotation. 2. The device of claim 1, wherein the plurality of rotators includes three rotators. 3. The device of claim 1, wherein the plurality of rotators are omniwheels. 4. The device of claim 1, wherein the plurality of rotators includes three omniwheels, wherein an axis of rotation of each of the three omniwheels is oriented approximately orthogonal to one another 5. The device of claim 1, further comprising a magnetic field sensor proximal to the spherical magnetic body to measure a dipole orientation of the spherical magnetic body. 6. The device of claim 5, wherein the magnetic field sensor is coupled to the housing. 7. The device of claim 1, wherein the spherical magnetic body comprises a spherical permanent magnet. 8. The device of claim 1, wherein the spherical magnetic body comprises a non-spherical permanent magnet encapsulated in a spherical structure. 9. The device of claim 1, wherein each of the plurality of rotators include a plurality of compressible support arms. 10. The device of claim 9, wherein each of the compressible support arms include a plurality of opposing cutouts to form a serpentine leveraged support arm that is compressible along a support axis extending a length of the support arm. 11. The device of claim 10, further comprising a compliant material substantially filling the cutouts. 12. The device of claim 1, wherein the device is part of a system for manipulation of a magnetic capsule endoscope, the system further comprising a robotic manipulator supporting the housing and a manipulation processor for causing movements of the robotic manipulator and the plurality of rotators. 13. A system for manipulation of a magnetic capsule endoscope, comprising: a magnetic manipulation device comprising a spherical magnetic body and a plurality of rotators in contact with the spherical magnetic body to rotate the spherical magnetic body in a desired direction;a robotic arm supporting the magnetic manipulation device, said robotic arm being movable along at least two axes;a magnetic field sensor proximal to the spherical magnetic body to measure a magnetic dipole direction of the spherical magnetic body; anda processor for causing movements of the robotic arm and the plurality of rotators based on a location of the magnetic capsule endoscope and the magnetic dipole direction of the spherical magnetic body. 14. The system of claim 13, wherein the magnetic field sensor comprises a plurality of magnetic field sensors. 15. The system of claim 14, wherein the plurality of magnetic field sensors comprises three sensors in close proximity to each other orien... |  | A61B001/00 A61B034/30 A61B005/06 A61B001/04
US09354122 | B2 | 2016-05-31 | Zero-heat-flux, deep tissue temperature measurement system | 2011-11-17 | A zero-heat-flux, deep tissue temperature measurement system measures internal body temperature by way of a probe having a heater and thermal sensors arranged in a zero-heat-flux construction. The measurement system includes control mechanization that determines heater and skin temperatures based upon data obtained from the probe and uses those temperatures to calculate a deep tissue temperature. The measurement system includes a signal interface cable having a connector where a probe can be releasably connected to the system. The cable and attached connector are a removable and replaceable part of the system, separate from the probe. The measurement system provides an output signal imitating a standard input signal configuration used by other equipment. | 1. A zero-heat-flux temperature measurement system for measuring deep tissue temperature using a probe with first and second flexible substrate layers sandwiching a layer of thermally insulating material and a tab with electrical connection pads, in which a heater and a first thermal sensor are disposed on the first substrate layer, a second thermal sensor and a programmable memory device are disposed on the second substrate layer, the system including: a controller with a signal connector jack and an emulator;a probe signal interface cable with first and second ends;a first connector attached to the first end of the probe signal interface cable for detachably connecting to the tab;a second connector attached to the second end of the probe signal interface cable for being inserted into and removed from the signal connector jack in the controller;in which the probe signal interface cable and the first and second connectors are a single integrated element separate from the probe,in which the controller includes probe control logic, and the zero-heat-flux temperature system further includes an information switch having a first state in which the information switch is operative to connect thermistor signals from the probe signal interface cable to the probe control logic and a second state in which the information switch is operative to connect programmable memory device information from the probe signal interface cable to the control logic. 2. The zero-heat-flux temperature measurement system of claim 1, including a heater switch operative to switch a pulse-width-modulated drive signal through the probe signal interface cable to the heater. 3. The zero-heat-flux temperature system of claim 1, including a heater switch operative to switch a pulse-width-modulated drive signal through the probe signal interface cable to the heater. 4. The zero-heat-flux temperature system of claim 3, in which the first state of the information switch blocks the transfer of programmable memory device signals from being transferred through the probe signal interface cable and the second state of the information switch enables the transfer of programmable memory device signals through the probe signal interface cable. 5. The zero-heat-flux temperature system of claim 4, further including an emulation unit with an emulation output jack, and an emulation output cable connected to the emulation output jack. 6. The zero-heat-flux temperature system of claim 5, in which the emulation unit emulates a YSI-400 thermistor. 7. A deep tissue temperature measurement system, including: a zero-heat-flux measurement probe with a heater, a first thermal sensor operative to sense a heater temperature, a second thermal sensor operative to sense a skin temperature, and a programmable memory device, and a connector interface;a processing unit with a probe signal connector and an emulator jack;a probe signal interface cable with first and second ends;a first connector attached to the fi... | 3M INNOVATIVE PROPERTIES COMPANY | A61B005/01 G01K001/16
US09330457 | B2 | 2016-05-03 | Device, method, and program for analyzing a magnetic resonance image using phase difference distribution | 2012-09-26 | An image analysis device for analyzing a magnetic resonance image obtained from a living body includes a phase difference distribution creating unit configured to create a phase difference distribution of a magnetic resonance image obtained from a predetermined area of the living body, a fitting unit configured to fit the phase difference distribution created by the phase difference distribution creating unit with a plurality of function groups, and a verifying unit configured to verify normality of the living body included in the predetermined area, based on the magnetic susceptibility of the tissue included in the predetermined area determined on the basis of the parameters of the plurality of function groups fit to the phase difference distribution by the fitting unit. | 1. An image analysis device for analyzing a magnetic resonance image obtained from a living body, comprising: a central processing unit (CPU) that: creates a phase difference distribution of a magnetic resonance image obtained from a predetermined area of the living body;fits the phase difference distribution with a plurality of function groups; andverifies normality of the living body included in the predetermined area, based on the magnetic susceptibility of the tissue included in the predetermined area determined on the basis of the parameters of the plurality of function groups fit to the phase difference distribution,wherein a horizontal axis is set as a phase value and a vertical axis is set as a data number in the phase difference distribution. 2. The image analysis device of claim 1, wherein the plurality of function groups includes at least one Gauss distribution. 3. The image analysis device of claim 1, wherein the plurality of function groups includes a first and second Gauss distributions,the first Gauss distribution has a larger standard deviation than the second Gauss distribution, andthe CPU further creates a magnetic resonance image enhancing a pixel having a phase in which the first Gauss distribution is larger than the second Gauss distribution. 4. An image analysis method for analyzing a magnetic resonance image obtained from a living body, comprising: creating a phase difference distribution of a magnetic resonance image obtained from a predetermined area of the living body;fitting the phase difference distribution with a plurality of function groups; andverifying normality of the living body included in the predetermined area, based on the magnetic susceptibility of the tissue included in the predetermined area determined on the basis of the parameters of the plurality of function groups fit to the phase difference distribution,wherein a horizontal axis is set as a phase value and a vertical axis is set as a data number in the phase difference distribution. 5. A non-transitory computer readable medium storing an image analysis program for implementing a function of analyzing a magnetic resonance image obtained from a living body in a computer, wherein the program comprises functionality to: create a phase difference distribution of a magnetic resonance image obtained from a predetermined area of the living body;fit the phase difference distribution with a plurality of function groups; andverify normality of the living body included in the predetermined area, based on the magnetic susceptibility of the tissue included in the predetermined area determined on the basis of the parameters of the plurality of function groups fit to the phase difference distribution,wherein a horizontal axis is set as a phase value and a vertical axis is set as a data number in the phase difference distribution. 6. The image analysis device of claim 2, wherein the plurality of function groups includes a first and second Gauss distributions,the ... | NATIONAL UNIVERSITY CORPORATION KUMAMOTO UNIVERSITY | G06K009/00 G06T007/00 A61B005/055 A61B005/00 G01R033/56 G01R033/24
US09339234 | B2 | 2016-05-17 | System and method for measuring physiological parameters | 2012-05-15 | A measuring method for measuring a physiological parameter via a measuring system comprising a test strip, an auxiliary measuring device and an electronic device having an application program and an analog-to digital converting unit is provided. The measuring method includes steps of coupling the auxiliary measuring device between the test strip and the electronic device to form a loop; the electronic device executing the application program to provide a first analog signal, and transferring the first analog signal to the auxiliary measuring device via the loop; the auxiliary measuring device applying a voltage to the test strip according to the first analog signal, and causing the test strip to generate a second analog signal; the analog-to-digital converting unit converting the second analog signal into a digital signal via the application program; and the application program calculating the digital signal to estimate the physiological parameter. | 1. A measuring method for measuring a physiological parameter via a measuring system comprising a test strip, an auxiliary measuring device and an electronic device having an application program, a correction datum, digital-to-analog converting (DAC) unit and an analog-to digital converting (ADC) unit, comprising steps of: coupling the auxiliary measuring device to the electronic device through an audio port, and the auxiliary measuring device forms a loop between the test strip and the electronic device;the electronic device executing the application program to initialize a measuring procedure by generating a first digital signal, wherein said first digital signal is an audio digital signal with a specific frequency;the DAC unit converting the first digital signal to a first analog signal, and transferring the first analog signal to the auxiliary measuring device via the loop;the auxiliary measuring device applying a voltage to the test strip according to the first analog signal, and causing the test strip to generate a second analog signal;the ADC unit converting the second analog signal into a second digital signal via the application program; andthe correction datum and the application program are used to correct and calculate the second digital signal to estimate the physiological parameter. 2. A method as claimed in claim 1, wherein when the auxiliary measuring device receives a power from the electronic device, the auxiliary measuring device converts the first analog signal into the voltage. 3. A method as claimed in claim 1, wherein when the auxiliary measuring device includes a battery, the auxiliary measuring device switches a first voltage mode to a second voltage mode in response to a third analog signal to cause the battery to provide the voltage. 4. A method as claimed in claim 1, wherein the physiological parameter comprises at least one of physiological parameters associated with a blood sugar concentration and a cholesterol concentration. 5. A method as claimed in claim 1, wherein the first analog signal is an audio signal. 6. A method for measuring a physiological parameter, comprising steps of: providing an electronic device having an analog-to-digital converting (ADC) unit, a digital-to-analog converting (DAC) unit, and an application program and a correction datum;generating a first digital signal having a specific frequency by the application program to initialize a measuring procedure, wherein the first digital signal is an audio digital signal;converting the first digital signal to a first analog signal by the DAC unit, wherein the first analog signal is an audio analog signal;transferring the first analog signal to an auxiliary measuring device through an audio port;converting the first analog signal into a voltage by a power source control unit;generating a second analog signal associated with the physiological parameter based on the voltage;transferring the second analog signal to the electronic device through the ... | BIONIME CORPORATION | A61B005/00 A61B005/145 A61B005/15
US09332909 | B2 | 2016-05-10 | Fluoroscopy apparatus | 2013-07-16 | Provided is a fluoroscopy apparatus including a fluorescence-image generating portion that generates a fluorescence image; an identifying portion that identifies a position of a high-luminance region in the fluorescence image; a storage portion that stores the position of the high-luminance region; a detecting portion that detects an amount of change in a physical quantity, which can possibly act as a cause of changes in a property of the high-luminance region, starting from a time at which the position of the high-luminance region is identified by the identifying portion; a confidence-level calculating portion that calculates a confidence level of the property of the high-luminance region based on the detected amount of change; and a display-image generating portion that generates a display image in which the display mode at the position of the high-luminance region is set in accordance with the confidence level. | 1. A fluoroscopy apparatus comprising: a light source that radiates excitation light onto an observation subject;a fluorescence-image generating portion that generates a fluorescence image by capturing fluorescence generated at the observation subject due to irradiation with the excitation light from the light source;an identifying portion that identifies a position of a high-luminance region that has a luminance value equal to or greater than a predetermined threshold in the fluorescence image generated by the fluorescence-image generating portion;a storage portion that stores the position of the high-luminance region identified by the identifying portion;a detecting portion that detects an amount of change in a physical quantity, which can possibly act as a cause of changes in a property of the high-luminance region, starting from a time at which the position of the high-luminance region is identified by the identifying portion;a confidence-level calculating portion that calculates a confidence level of the property of the high-luminance region identified by the identifying portion based on the amount of change detected by the detecting portion; anda display-image generating portion that generates a display image in which the display mode at the position of the high-luminance region stored in the storage portion is set in accordance with the confidence level calculated by the confidence-level calculating portion. 2. The fluoroscopy apparatus according to claim 1, wherein the detecting portion detects, as the amount of change in the physical quantity, an elapsed time since the time at which the position of the high-luminance region is identified by the identifying portion. 3. The fluoroscopy apparatus according to claim 1, wherein the light source radiates reference light onto the observation subject together with the excitation light;a reference-image generating portion that generates a reference image by capturing return light returning from the observation subject due to the irradiation with the reference light from the light source is provided; andthe detecting portion detects, as the amount of change in the physical quantity, an amount of movement of the observation subject in the reference image generated by the reference-image generating portion since the time at which the position of the high-luminance region is identified by the identifying portion. 4. The fluoroscopy apparatus according to claim 2, wherein the confidence-level calculating portion calculates the confidence level in accordance with a function that decreases with an increase in the amount of change in the physical quantity. 5. The fluoroscopy apparatus according to claim 3, wherein the confidence-level calculating portion calculates the confidence level in accordance with a function that decreases with an increase in the amount of change in the physical quantity. 6. The fluoroscopy apparatus according to claim 2, wherein the confidence-level calculating portion calcu... | OLYMPUS CORPORATION | H04N007/18 A61B005/00 A61B001/04 A61B001/00 A61B001/06
US20160128595 | A1 | 2016-05-12 | MEANS AND METHOD FOR THE DETECTION OF CARDIAC EVENTS | 2016-01-19 | A system for detecting cardiac events and alerting the patient of the cardiac event issued by an implanted medical device. A processor detects a cardiac event and issues an activation signal to a patient alerting mechanism which emits an alarm during a first time interval followed by an alarm-off period of time. Reminder alarms are issued periodically following the alarm-off period of time. Reminder alarms are issued during reminder on-periods of time which are shorter than the first time interval. The alarm-off period and the time between periodic reminder alarms are greater than the alarm first time interval. A disabling device permits disabling of reminder alarms. | 1. A system for detecting cardiac events in a human patient, the system including: at least two electrodes for obtaining an electrical signal from the patient's heart; a processor configured to detect at an emergency cardiac event, by processing the electrical signal from the patient's heart; a patient alerting mechanism connected to the processor for receiving an activation signal from the processor dependent on said processor detection of said cardiac event, the patient alerting mechanism being activated by said activation signal, the patient alerting mechanism configured for: responsive to said activation signal, emitting an alarm signal having a first number of pulses within a first period of time defining an initial alarm on-period; the initial alarm-on period being followed by a respective alarm off-period of time devoid of pulses, the patient alerting mechanism being further configured to emit a reminder alarm signal periodically during a predetermined reminder time period following the initial alarm off-time period, wherein each of the reminder alarms occurs during a reminder alarm on-period that is shorter than said first period of time, and wherein the alarm-off period and the time between periodic reminder alarms are both greater than said first period of time, said patient alerting mechanism including a disabling mechanism for disabling said reminder alarm. 2. The system of claim 1 where the periodic reminder time period is greater than 15 minutes. 3. The system of claim 1 where the periodic reminder time period is less than 5 hours. 4. The system of claim 1 where the alarm off-period of time is at least twice as long as the alarm on-period. 5. The system of claim 1 where the alarm on-period is less than 10 minutes. 6. The system of claim 1 where at least a portion of the system is implanted in the human patient. 7. The system of claim 1 where the cardiac events detected by the system include acute myocardial infarction. 8. The system of claim 1 where the cardiac events detected by the system include ischemia at an elevated heart rate. 9. The system of claim 1 where the cardiac events detected by the system include low heart rate. 10. The system of claim 1 where the cardiac events detected by the system include excessively high heart rate. 11. The system of claim 1 where the cardiac events detected by the system include arrhythmias. |  | A61B005/0464 A61B005/00 A61B005/0245 A61B005/042
US10973430 | B2 | 2021-04-13 | Fully reconfigurable modular body-worn sensors | 2015-11-12 | A system is disclosed including: one or more electroencephalography (EEG) sensor nodes in a fully reconfigurable sensor network configured to detect electrical signals indicating activity of a brain of a subject and to transmit signal information relating to the detected electrical signals, each EEG sensor node embodying an analog front end (AFE) circuit, and a centralized command control node (CCN) in the sensor network in communication with the one or more EEG sensor nodes and configured to receive the transmitted signal information from the one or more EEG sensor nodes and to perform a processing technique on the received information resulting in processed signal data, wherein the one or more EEG sensor nodes are configured to detect the electrical signals without a driven right leg (DRL) circuit in the sensor network. | 1. A system comprising: one or more electroencephalography (EEG) sensor nodes in a reconfigurable sensor network configured to detect electrical signals indicating activity of a brain of a subject and to transmit signal information relating to the detected electrical signals, each EEG sensor node embodying an analog front end (AFE) circuit;a centralized command control node (CCN) in the sensor network in communication with the one or more EEG sensor nodes and configured to receive the transmitted signal information from the one or more EEG sensor nodes and to perform a processing technique on the received information resulting in processed signal data; andone or more other sensor nodes, other than EEG sensor nodes, connected in the sensor network, in communication with the CCN, and configured to detect other signals and to transmit information relating to the detected other signals to the CCN;wherein the one or more EEG sensor nodes are configured to detect the electrical signals without a driven right leg (DRL) circuit in the reconfigurable sensor network;further wherein the one or more EEG sensor nodes and the one or more other sensor nodes in the sensor network are configured to transmit data among themselves, and to perform data processing on data received from another sensor node; andfurther wherein the reconfigurable sensor network is adapted to allow plug-and-play reconfiguration of the one or more EEG sensor nodes during deployment, wherein said sensor nodes comprise plug-and-play modules. 2. The system of claim 1, wherein the one or more EEG sensor nodes are in communication with the one or more other sensor nodes. 3. The system of claim 1, wherein the one or more EEG sensor nodes and the one or more other sensor nodes are each configured to be worn by the subject. 4. The system of claim 1, wherein the other signals are associated with physiological data of the subject. 5. The system of claim 1, wherein the other signals are associated with nonphysiological data of the subject. 6. The system of claim 1, wherein the CCN is wirelessly in communication with the one or more EEG sensor nodes, the one or more other sensor nodes, or both. 7. The system of claim 1, wherein the one or more EEG sensor nodes are further configured to locally perform a processing technique on the detected electrical signals resulting in processed signal data and to transmit the processed signal data to the CCN. 8. The system of claim 1, wherein the AFE circuit includes, in order, a first amplifier, a plurality of filters, and a second amplifier. 9. The system of claim 1, wherein the sensor network is configured according to an inter-integrated circuit (I2C) topology. 10. The system of claim 1, wherein the CCN is further configured to control an aspect of the sensor network and to control an aspect of the one or more EEG sensor nodes. 11. The system of claim 1, wherein the CCN is further configured to dynamically control the one or more EEG sensor nodes so ... | - | A61B005/04 A61B005/0476 A61B005/00
US09352163 | B2 | 2016-05-31 | Electrocardiogram identification | 2015-08-20 | A defibrillating system includes a processor coupled to a memory. The processor and the memory are configured to identify a treatment event associated with treatment of a victim with the defibrillating system, and transmit a representation of a portion of an ECG signal associated with the identified treatment event. In some cases, the processor and the memory are configured to identify the portion of the ECG signal associated with the identified treatment event. In some cases, the portion of the ECG signal is of a predetermined length of time having a start time and an end time based on a time associated with the identified treatment event. | 1. A defibrillating system comprising: a processor coupled to a memory, the processor and the memory configured to: determine that treatment of a patient using the defibrillating system is completed; andupon determining that treatment of the patient using the defibrillating system is complete, provide, for a user, a representation of a portion of a physiological signal obtained by the defibrillating system during treatment of the patient. 2. The defibrillating system of claim 1, wherein determining that treatment of the patient is completed comprises detecting that the defibrillating system has been disconnected from the patient. 3. The defibrillating system of claim 2, wherein detecting that the defibrillating system has been disconnected from the patient comprises detecting a change in impedance in a circuit that includes the defibrillating system and the patient. 4. The defibrillating system of claim 1, wherein the physiological signal comprises an electrocardiogram (ECG) signal. 5. The defibrillating system of claim 4, wherein the processor and memory are further configured to identify a portion of the ECG signal analyzed by the defibrillating system in determining whether to deliver treatment to the patient; and wherein providing a representation of a portion of a physiological signal comprises providing the identified portion of the ECG signal. 6. The defibrillating system of claim 4, wherein the processor and memory are further configured to provide an identifier associated with the portion of the ECG signal. 7. The defibrillating system of claim 6, wherein the identifier associated with the portion of the ECG signal comprises a time associated with the portion of the ECG signal. 8. The defibrillating system of claim 6, wherein the identifier associated with the portion of the ECG signal comprises an indication of whether the portion of the ECG signal represents a shockable cardiac rhythm or a non-shockable cardiac rhythm. 9. The defibrillating system of claim 6, wherein the identifier associated with the portion of the ECG signal comprises an identifier of a type of cardiac rhythm represented by the identified portion of the ECG signal. 10. The defibrillating system of claim 1, wherein the physiological signal comprises one or more of an SpO2 signal, a CO2 signal, a ventilation signal, and a chest compression signal. 11. The defibrillating system of claim 1, wherein the processor and memory are further configured to provide audio information associated with the portion of the physiological signal. 12. The defibrillating system of claim 1, wherein providing a representation of the portion of the physiological signal comprises displaying the representation of the portion of the physiological signal on a display interface of the defibrillating system. 13. The defibrillating system of claim 1, wherein providing a representation of the portion of the physiological signal comprises transmitting the representation of the portion of t... | ZOLL Medical Corporation | A61B005/04 A61N001/39 A61B005/0432 A61B005/00 A61B005/046 A61B005/0464 A61B005/0404 A61B005/0408
US09454709 | B2 | 2016-09-27 | Method for rapid-multicontrast brain imaging | 2014-11-12 | A method for providing an image of a subject is provided. The plurality of Cartesian points is divided into a first group and a second group, wherein all of the Cartesian points in the first group have a kr that is less than or equal to kr of all of the Cartesian points in the second group. The second group is divided into N subgroups. An inversion recovery radio frequency is applied to the subject. K-space data for the first group of Cartesian points for M different time periods is acquired, wherein k-space data is acquired for only one sub-group sequentially and cyclically for the M different time periods. For each time period i of the M different time periods, acquired k-space data for the first group of Cartesian points at time period i and k-space data from N consecutive subgroups of Cartesian points. | 1. A method for providing an image of a subject in a magnetic resonance imaging MRI system, comprising: defining a plurality of Cartesian points on a ky-kz plane with a center, wherein each of the plurality of Cartesian points has a kr, wherein kr is a distance from a Cartesian point of the plurality of Cartesian points to the center;dividing the plurality of Cartesian points into a first group (Group A) and a second group (Group B), wherein all of the Cartesian points in the first group have a kr that is less than or equal to kr of all of the Cartesian points in the second group;dividing the second group into N subgroups, wherein N is a whole number greater than or equal to two, wherein the subgroups have an incoherent point spread function;applying an inversion recovery (IR) radio frequency to the subject such that a net longitudinal magnetization in the subject is substantially inverted and begins to relax back to equilibrium;acquiring k-space data for the first group of Cartesian points for M different time periods, wherein k-space data is acquired for only one sub-group for each time period of the M different time periods, and wherein k-space data is acquired for each subgroup sequentially and cyclically for the M different time periods;for each time period i of the M different time periods, using acquired k-space data for the first group of Cartesian points at time period i and k-space data from N consecutive subgroups of Cartesian points where one of the N consecutive subgroups is the subgroup that is acquired during time period i to create image data of the subject for first time period i. 2. The method, as recited in claim 1, further comprising: for each time period i for the M different time periods, using the image data for time period i to create an image for time period i; anddisplaying the image for time period. 3. The method, as recited in claim 1, wherein the image data from a plurality of time periods are combined to create a relaxation map. 4. The method, as recited in claim 1, wherein the first plurality of Cartesian points are within a radial sector of the circle, and wherein the Cartesian points within the first group are all Cartesian points in the radial sector at or within a radius of R from the center of the circle and wherein the Cartesian points within the second group are all Cartesian points in the radial sector at or beyond the radius of R from the center of the circle. 5. The method, as recited in claim 1, wherein the Cartesian points within the first group are Cartesian points in the circle at or within a radius of r from the center of the circle and wherein the Cartesian points within the second group are Cartesian points in the circle at or beyond the radius of r from the center of the circle. 6. The method, as recited in claim 1, wherein Cartesian points in each subgroup are dispersed as much as possible. 7. The method, as recited in claim 1, wherein the Cartesian points in each subgroup are chosen by at l... | The Board of Trustees of the Leland Stanford Junior University | G06K009/00 G06K009/46 G06T007/00 A61B005/00 A61B005/055 G01R033/48 G01R033/56
US10219735 | B2 | 2019-03-05 | Intraoral moisture measuring device | 2014-02-19 | An intraoral moisture measuring device includes: a swing member that swings with respect to a main body about a predetermined swing center; a moisture amount detection unit provided at a tip of the swing member, for detecting a moisture amount by being directly or indirectly abutted against a measurement site in a mouth; and a biasing member for biasing the swing member in one of swing directions. Consequently, the intraoral moisture measuring device is capable of measuring intraoral moisture in a simple and highly-accurate manner. | 1. An intraoral moisture measuring device comprising: a swing member that swings with respect to a main body about a predetermined swing center;a capacitance sensor provided at a tip of the swing member, the capacitance sensor detecting a moisture amount by being directly or indirectly abutted against a measurement site in a mouth;a biasing member for biasing the swing member in one of a plurality of swing directions;the swing member comprising: a fixed part positioned close to the main body that swings about the predetermined swing center; anda detachable part detachably attached to the fixed part and provided with the capacitance sensor; anda support member supporting a rear side of the measurement site and extending from the main body, the support member being arranged to face the swing member and being disposed at positions to interpose the measurement site in cooperation with the swing member;wherein the swing member further includes an operation unit extended on a side opposite to the capacitance sensor with respect to the swing center, and the swing member can be swung by operating the operation unit with a hold holding the main body. 2. The intraoral moisture measuring device according to claim 1, wherein the swing center is disposed in the main body. 3. The intraoral moisture measuring device according to claim 1, comprising a strain gauge, the strain gauge detecting a pressing force to the measurement site by the capacitance sensor. 4. The intraoral moisture measuring device according to claim 3, wherein the strain gauge is disposed inside the main body. 5. The intraoral moisture measuring device according to claim 3, comprising a controller, the controller controlling the capacitance sensor and the strain gauge, and wherein the controller determines if the pressing force detected by the strain gauge falls within an appropriate range, and wherein the controller starts moisture amount measurement when the pressing force detected by the strain gauge falls within the appropriate range. 6. The intraoral moisture measuring device according to claim 5, wherein the controller stops the moisture amount measurement when the pressing force detected by the strain gauge falls outside the appropriate range. 7. The intraoral moisture measuring device according to claim 1, wherein the detachable part is provided with a first electrode and a second electrode included in the capacitance sensor and wirings connected to the first electrode and the second electrode, and the first electrode, the second electrode, and the wirings are formed by electrically-conductive thin films disposed in the detachable part. 8. The intraoral moisture measuring device according to claim 7, wherein the wirings are covered with a shield constituted by the electrically-conductive thin films. 9. The intraoral moisture measuring device according to claim 1, wherein the support member includes a protrusion protruded toward the rear side of the measurement site, and the pr... | KABUSHIKIKAISHA RAIFU | A61B005/00 A61B005/053
US09345410 | B2 | 2016-05-24 | Diagnostic and optimization using exercise recovery data | 2014-03-03 | An apparatus, such as a cardiac function management system can detect heart sounds following a sensed transition in physical activity level, such as from an elevated physical activity level to rest. A technique can include systems, methods, machine-readable media, or other techniques that can include identifying a physical activity level transition, receiving a heart sound signal, determining characteristics of the heart sound and subject physiologic activity to provide an indication, such as a heart failure status indication. | 1. A system comprising: a sensor interface circuit configured to obtain information indicative of both (1) a physical activity of a subject and (2) a heart sound of the subject from a mechanical sensor, and to provide a resulting sensor signal;a signal processor circuit, coupled to the sensor interface circuit to receive the sensor signal, the signal processing circuit configured to, using information about the sensor signal: detect a first physical activity level transition from a first elevated activity level to a less elevated activity level;detect a second physical activity level transition from a different second elevated activity level to the same or different less elevated activity level;identify heart sounds corresponding to multiple cardiac cycles following the detected first and second physical activity level transitions;determine respective heart sound characteristics for the identified heart sounds; anddetermine a physiologic indicator using a trend of the determined heart sound characteristics. 2. The system of claim 1, wherein the respective heart sound characteristics include indications of an energy of at least one of an S1, S2, S3, or S4 heart sound. 3. The system of claim 1, wherein the signal processing circuit is configured to determine the respective heart sound characteristics by determining heart timing intervals that are based on the identified heart sounds relative to their respective physical activity level transitions. 4. The system of claim 1, wherein the signal processing circuit is configured to: determine a first heart sound characteristic in response to the first physical activity level transition and classify the first heart sound characteristic as being associated with the first physical activity level transition;determine a second heart sound characteristic in response to the second physical activity level transition and classify the second heart sound characteristic as being associated with the second physical activity level transition; anddetermine the physiologic indicator using information about a trend of the first and second heart sound characteristics. 5. The system of claim 4, wherein the signal processing circuit is configured to determine the physiologic indicator using one of a slope or an intercept of the trend of the first and second heart sound characteristics. 6. The system of claim 1, wherein the respective heart sound characteristics for the identified heart sounds includes respective heart sound timing intervals for the identified heart sounds. 7. The system of claim 6, wherein the heart sound timing intervals include pre-ejection period (PEP) intervals corresponding to the first and second physical activity level transitions, and wherein the signal processing circuit is configured to determine the physiologic indicator using the PEP intervals. 8. The system of claim 1, further comprising the mechanical sensor. 9. The system of claim 8, wherein the mechanical sensor, the sensor interfac... | Cardiac Pacemakers, Inc. | A61B005/0205 A61B005/024 A61B005/11 A61B005/00
US09345899 | B2 | 2016-05-24 | Vector switching in an implantable cardiac stimulus system | 2014-08-06 | The implantable cardiac treatment system of the present invention is capable of choosing the most appropriate electrode vector to sense within a particular patient. In certain embodiments, the implantable cardiac treatment system determines the most appropriate electrode vector for continuous sensing based on which electrode vector results in the greatest signal amplitude, or some other useful metric such as signal-to-noise ratio (SNR). The electrode vector possessing the highest quality as measured using the metric is then set as the default electrode vector for sensing. Additionally, in certain embodiments of the present invention, a next alternative electrode vector is selected based on being generally orthogonal to the default electrode vector. In yet other embodiments of the present invention, the next alternative electrode vector is selected based on possessing the next highest quality metric after the default electrode vector. In some embodiments, if analysis of the default vector is ambiguous, the next alternative electrode vector is analyzed to reduce ambiguity. | 1. A method of cardiac signal analysis in an implantable cardiac stimulus device comprising: implanting, into the patient, cardiac stimulus system comprising a canister and a lead assembly, the lead assembly including a plurality of electrodes thereon and the canister including at least one electrode thereon for use in cardiac activity sensing;identifying a first electrode pair defining a first sensing vector by observing a cardiac signal metric for a plurality of available sensing vectors and using the first electrode pair as a default electrode pair when observing cardiac activity of the patient;determining whether cardiac signals captured using the default electrode pair enable formation of a template representative of normal sinus rhythm for the patient; and, if not,changing the default electrode pair to a different pair of electrodes defining a sensing vector other than the first sensing vector. 2. A method of cardiac signal analysis in an implantable cardiac treatment device, the implantable cardiac treatment device including at least first, second and third electrodes coupled to operational circuitry for the device, the method comprising: configuring the device by sensing signals using at least first and second sense vectors defined by first and second pairs of electrodes chosen from the first, second and third electrodes;calculating signal metrics for the first and second vectors;selecting a default sense vector by analysis of the signal metric;using the default sense vector for sensing of cardiac signals;analyzing a first signal captured from the default sense vector to determine whether: a normal sinus rhythm has been sensed;an arrhythmia has been sensed; orambiguity is present; anddetermining that ambiguity is present and selecting a different sense vector from the default sense vector and analyzing a second signal captured from the different sense vector in order to resolve the ambiguity. 3. The method of claim 2 further comprising finding that the ambiguity cannot be resolved using the different sense vector and reinstating the default sense vector. 4. The method of claim 2 further comprising finding that the ambiguity can be resolved using the different sense vector and reconfiguring the device to use the different sense vector as the default sense vector thereafter. 5. The method of claim 2 further comprising: after the ambiguity is identified, analyzing the default sense vector to determine whether the default sense vector provides a suitable signal for cardiac signal analysis; anddetermining the default sense vector does not provide a suitable signal and reconfiguring the device to use another vector as the default sense vector. 6. The method of claim 2, wherein the device comprises a housing and at least one of the electrodes is located on the housing. 7. The method of claim 6, further comprising generating a cardiac stimulus using an electrode pair including the electrode on the housing. 8. A method of operating an impla... | CAMERON HEALTH, INC. | A61N001/00 A61N001/39 A61B005/0424 A61N001/362 A61B005/00 A61N001/37
US10376157 | B2 | 2019-08-13 | Systems and methods for determining respiration information using phase locked loop | 2016-01-21 | A patient monitoring system may receive a physiological signal such as a photoplethysmograph (PPG) signal that exhibits frequency and amplitude modulation based on respiration. A phase locked loop may generate a frequency demodulated signal and an amplitude demodulated signal from the PPG signal. An autocorrelation sequence may be generated for each of the frequency demodulated signal and the amplitude demodulated signal. The autocorrelation sequences may be combined and respiration information may be determined based on the combined autocorrelation sequence. | 1. A method for determining respiration rate using a patient monitor, the method comprising: receiving a photoplethysmograph (PPG) signal;generating, using the patient monitor, an oscillator signal;generating, using the patient monitor, a phase-shifted oscillator signal based on a phase difference between the PPG signal and the oscillator signal;generating, using the patient monitor, an amplitude demodulated signal by multiplying the phase-shifted oscillator signal and the PPG signal;determining, using the patient monitor, the respiration rate based on the amplitude demodulated signal; anddisplaying, using a display, the respiration rate. 2. The method of claim 1, wherein the patient monitor is configured to receive the PPG signal from a pulse oximetry sensor, wherein the pulse oximetry sensor detects light attenuated by the subject. 3. The method of claim 1, wherein a product signal generated by multiplying the phase-shifted oscillator signal and the PPG signal is low pass filtered before generating the amplitude demodulated signal. 4. The method of claim 1, wherein determining the respiration rate comprises: generating an autocorrelation sequence based on the amplitude demodulated signal; anddetermining the respiration rate based on the autocorrelation sequence. 5. The method of claim 1, wherein the PPG signal is band pass filtered based on a heart rate associated with the PPG signal. 6. The method of claim 1, wherein the oscillator signal is generated using voltage controlled oscillator or a modulo-N counter. 7. The method of claim 1, wherein the modulo-N counter has a length controlled based on the phase difference between the PPG signal and the oscillator signal. 8. The method of claim 1, wherein the PPG signal comprises a fixed length analysis window. 9. The method of claim 1, wherein the oscillator signal is an approximation of the PPG signal. 10. The method of claim 1, wherein the phase-shifted oscillator signal matches phase and frequency of the PPG signal. 11. A patient monitor for determining respiration rate, the patient monitor comprising: an interface configured to receive a photoplethysmograph (PPG) signal;an oscillator configured to generate an oscillator signal;a processor configured to:generate a phase-shifted oscillator signal based on a phase difference between the PPG signal and the oscillator signal;generate an amplitude demodulated signal by multiplying the phase-shifted oscillator signal and the PPG signal; anddetermine respiration rate based on the amplitude demodulated signal; anda display configured to display the respiration information. 12. The patient monitor of claim 11, wherein the interface is configured to receive the PPG signal from a pulse oximetry sensor, wherein the pulse oximetry sensor detects light attenuated by the subject. 13. The patient monitor of claim 11, wherein a product signal generated by multiplying the phase-shifted oscillator signal and the PPG signal is low pass filtered before ge... | NELLCOR PURITAN BENNETT IRELAND | A61B005/0205 A61B005/08 A61B005/0295 A61B005/00 A61B005/1455
US09345409 | B2 | 2016-05-24 | Blood pressure measurement apparatus and control method for blood pressure measurement apparatus | 2015-07-14 | A blood pressure measurement apparatus includes a sensor, an input section, and a circuit. The sensor is configured to measure blood flow speed in a target blood vessel. The input section is configured to input blood pressure from a measurement apparatus continuously measuring the blood pressure. The circuit is configured to estimate the blood pressure on the basis of the blood flow speed by referencing a correlation formula that indicates correlation characteristics of the blood pressure from the measurement apparatus and the blood flow speed from the sensor. | 1. A blood pressure measurement apparatus comprising: a sensor configured to measure blood flow speed in a target blood vessel;an input section configured to input blood pressure from a measurement apparatus continuously measuring the blood pressure; anda circuit configured to estimate the blood pressure based on the blood flow speed by referencing a correlation formula that indicates correlation characteristics of the blood pressure from the measurement apparatus and the blood flow speed from the sensor. 2. The blood pressure measurement apparatus according to claim 1, wherein the circuit is further configured to correct the correlation formula by utilizing the blood pressure from the measurement apparatus. 3. The blood pressure measurement apparatus according to claim 2, wherein the correlation formula has a plurality of parameters, andthe circuit rederives the correlation characteristics by recalculating all of the values of the plurality of parameters. 4. The blood pressure measurement apparatus according to claim 2, wherein the circuit is further configured to determine a blood flow state of the target blood vessel using results of measuring of the blood flow speed, and to determine a necessity of correction of the correlation characteristics by comparing a reference blood flow state of the target blood vessel which is set in advance and the blood flow state which is determined. 5. The blood pressure measurement apparatus according to claim 4, wherein the circuit is configured to measure the blood flow speed in a plurality of positions which are different positions in a diameter direction in the target blood vessel, and to determine distribution or change trend of the blood flow speed in a transverse direction of the target blood vessel, which indicates blood flow speed distribution, as the blood flow state using the results of the measuring of the blood flow speed. 6. The blood pressure measurement apparatus according to claim 5, wherein the circuit is further configured to determine the necessity of correction of the correlation characteristics by comparing undulations of the blood flow speed distribution of the reference blood flow state as reference undulations and undulations of the blood flow speed distribution as gauging undulations, which is determined. 7. The blood pressure measurement apparatus according to claim 4, wherein the circuit is further configured to determine whether the correction is necessary based on the results of the comparison. 8. The blood pressure measurement apparatus according to claim 1, wherein the correlation formula is a linear function. 9. The blood pressure measurement apparatus according to claim 1, wherein the correlation formula is a non-linear function. 10. The blood pressure measurement apparatus according to claim 4, wherein the circuit is further configured to determine the necessity of correction of the correlation formula when a difference between the reference blood flow state and the b... | Seiko Epson Corporation | A61B005/05 A61B005/021 A61B008/04 A61B005/0285 A61B008/06 A61B008/00
US09339612 | B2 | 2016-05-17 | Tissue penetration device | 2008-12-16 | A tissue penetration device and method of using same. The tissue penetration device may optionally include sampling and analyzing functions, which may be integrated. An embodiment provides control of a lancet used for sampling blood. Electric field coils or solenoids may drive the lancet using electromagnetic force. Advancement and retraction of a lancet may be controlled by a feedback loop monitoring the position and velocity of the lancet embodiments of the lancet driver can be configured to follow a predetermined tissue lancing profile. Embodiments of the invention include a lancet and method for using a lancet to maintain the patency of the wound tract once the lancet has cut into the skin. | 1. A method for creating a puncture wound in the epidermis of a patient's body part to obtain a body fluid sample, comprising: providing a fluid sampling device with a housing and a cartridge positionable in the housing that includes a plurality of penetrating members;creating a skin opening in the epidermis with one of said plurality of penetrating members at a puncture site to a depth below the puncture site during a skin-opening step;using a processor to allow said one of said plurality of the penetrating members to settle at the depth of the puncture site for a sufficient amount of time at a body fluid sampling site below the epidermis;using the processor to customize a penetrating member driver velocity profile by scaling or modifying the penetrating member driver velocity profile based on user input information and in response to said customizing, the processor modulating power from a power supply to the penetrating member driver through an amplifier;withdrawing said one of said plurality of the penetrating members from the puncture site;maintaining the patient's skin in position at the body fluid sampling site until an adequate volume of blood has been collected; andproviding a signal to the patient from the fluid sampling device that the patient's skin may be lifted from the sampling site after said one of said plurality of the penetrating members has been withdrawn, the signal being either a visual or audio signal. 2. The method of claim 1, wherein said one of said plurality of the penetrating members is decelerated at an end of a skin-opening step and then accelerated again in the sample collection step. 3. The method of claim 2, wherein said one of said plurality of the penetrating members is stopped at the end of the skin-opening step. 4. The method of claim 3, wherein said one of said plurality of the penetrating members is retracted at the end of the skin-opening step. 5. The method of claim 4, wherein the penetrating member is completely withdrawn from the epidermis at the end of the skin-opening step. 6. The method of claim 1, further comprising: after inserting said one of said plurality of the penetrating members into the skin opening to a second depth,retracting said one of said plurality of the penetrating members into a collection position where a tip of said one of said plurality of the penetrating members remains within the skin. 7. The method of claim 6, further comprising: decelerating the penetrating member as said penetrating member approaches the collection position. 8. A method of using a penetrating member apparatus to create a puncture wound in an epidermis of a patient's skin at a sampling site to obtain a body fluid sample, the penetrating member apparatus including a cartridge with a plurality of penetrating members, a drive for advancing and retracting said plurality of penetrating members, and a housing having an opening from which said plurality of penetrating members are accessible for contacting a bod... | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/00 A61M005/32 A61B005/15 A61B005/155 A61B005/151
US09338583 | B2 | 2016-05-10 | Method for energy efficient body sensor network discovery | 2011-07-14 | When monitoring a patient in a healthcare environment, a location of a battery-powered mobile aggregator (MA) sensor (22), which is mounted on a patient with one or more other patient-mounted sensors (12, S1, S2, S3), is monitored. The MA (22) maintains a personal area network (PAN) (24) for the sensors (12, S1, S2, S3) and wirelessly transmits signals from patient mounted sensors (12, S1, S2, S3) to a wired network (18). When the MA (22) comes within range of a mains-powered fixed aggregator (FA) (16), the MA (22) is informed of the availability of the FA (16), transfers patient data communication duties to the FA (16), and shuts down its PAN to conserve battery power. If the patient moves out of range of the FA (16), then the MA (22) regenerates the PAN (24) and resumes wireless patient data transmission. | 1. A method of transferring a patient monitoring sensor group between fixed and mobile modes of communication with a wired healthcare network, comprising: detecting the sensor group, which comprises one or more patient-mounted sensors and a mobile aggregator sensor, within communication range of a fixed aggregator module;informing the mobile aggregator sensor that it and the one or more patient-mounted sensors are within range of the fixed aggregator module;when the fixed aggregator module is within range, communicating status information from the sensor group to a wired network via the fixed aggregator module while bypassing the mobile aggregator sensor; andterminating a mobile aggregator personal area network (MA-PAN), which is used for communication between the sensor group and the mobile aggregator sensor, when the sensor group is within range of the fixed aggregator module in order to conserve battery power at the mobile aggregator sensor. 2. The method according to claim 1, further including, when the mobile aggregator sensor is out of range of the fixed aggregator module: establishing the mobile aggregator personal area network (MA-PAN) at the mobile aggregator sensor when the mobile aggregator sensor is out of range of the fixed aggregator;receiving the status information from the one or more patient sensors at the mobile aggregator sensor via the MA-PAN; andtransmitting the status information from the mobile aggregator sensor to the wired network and bypassing the fixed aggregator module. 3. The method according to claim 2, wherein the mobile aggregator sensor transmits the status information to the wired network over a wireless local area network (WLAN) communication link. 4. The method according to claim 1, wherein detecting the sensor group includes: monitoring, at the fixed aggregator module, a plurality of wireless local access network (WLAN) radio channels in order to detect WLAN transmissions from the mobile aggregator sensor. 5. The method according to claim 1, wherein detecting the sensor group includes: establishing a communication link between the wired network and each of the mobile aggregator sensor and the fixed aggregator module; andmonitoring, at the wired network, a location of the mobile aggregator sensor relative to the fixed aggregator module using at least one of:a global positioning system (GPS) and a map of the of the healthcare facility; anda map of the healthcare facility and an asset tracking system. 6. The method according to claim 1, wherein detecting the sensor group includes: transmitting a beacon signal from the fixed aggregator module;receiving, via the wired network, location information from the mobile aggregator sensor indicating that the mobile aggregator sensor is within communication range of the fixed aggregator module. 7. The method according to claim 1, wherein detecting the sensor group includes: monitoring, at the fixed aggregator module, a plurality of short range radio channels in order ... | KONINKLIJKE PHILIPS N.V. | H04W016/14 H04W004/00 A61B005/00 G06F019/00 H04W048/16 A61B005/0205 A61B005/11 H04W084/00 H04W084/22
US09351647 | B2 | 2016-05-31 | Indication-based worsening HF alert | 2013-04-08 | This document discusses, among other things, receiving a user selection of a heart failure symptom, receiving a user selection of an abnormal psychological condition, receiving information about a physiological patient status parameter, and determining a heart failure status using the received information. | 1. A system, comprising: a user input configured to: receive a user selection of a heart failure symptom to indicate whether a heart failure symptom is or has been present, wherein the heart failure symptom includes at least one of dyspnea, edema, or weight change; andreceive a user selection of a heart failure (HF) comorbidity to indicate whether an HF comorbidity is or has been present; anda heart failure (HF) status processor, coupled to the user input, the HF status processor configured to determine a heart failure status using information about the user selection of the heart failure symptom and information received from a physiological patient status parameter detector, the information received from the physiological patient status parameter detector including information about a first physiological patient status parameter and a different second physiological patient status parameter;wherein the HF status processor configured to determine the heart failure status is further configured to controllably adjust the heart failure status determination, wherein controllably adjusting the heart failure status determination comprises: adjusting a first weight for the information about the first physiological patient status parameter relative to a second weight for the information about the different second physiological patient status parameter; andusing the information obtained from the first and second physiological patient status parameters in conjunction with the first and second weights. 2. The system of claim 1, comprising a heart failure (HF) status indicator configured to provide a heart failure status indication using information from the HF status processor. 3. The system of claim 2, wherein the HF status processor is configured to controllably adjust at least one of a sensitivity, a specificity, a positive predictive value, or a negative predictive value of the heart failure condition determination using information about the user selection of the heart failure symptom. 4. The system of claim 1, wherein the first physiological patient status parameter includes at least one of a heart rate, a respiration rate, a respiration timing, a blood pressure, a lung tidal volume, a physical activity level, a weight, an intrathoracic impedance, a heart sound timing, or a heart sound magnitude. 5. The system of claim 4, wherein the user input configured to receive the heart failure symptom indicator is further configured to receive an indication that edema is or has been present; and wherein the physiological patient status parameter includes the intrathoracic impedance. 6. The system of claim 4, wherein the user input configured to receive the heart failure symptom indicator is further configured to receive an indication that dyspnea is or has been present; and wherein the first physiological patient status parameter includes at least one of the respiration rate, the respiration timing, the lung tidal volume, or the physical activity level. 7.... | Cardiac Pacemakers, Inc. | A61B005/0205 A61B005/00 A61B005/021 A61B005/053 A61B005/085 A61B005/16
US09355216 | B2 | 2016-05-31 | Using IPTV as health monitor | 2013-10-16 | An IPTV can receive a user name and present a screen enabling a user by means of a remote control to select health information categories, inputting numeric parameters into the selected categories which can be uploaded to an Internet server for analysis and viewed by the user for subsequent monitoring/charting. | 1. Method, comprising: receiving from a user interface, input indicating an identity of a person;receiving from direct user input and/or from a biological sensor with the person, input representing a physical parameter of the person;storing a time associated with the numeric input along with the input representing the physical parameter;outputting for display a visualization the person's history of the measurable physical parameter; andoutputting for display at least one selectable parameter selectable to input numbers representative thereof, the at least one selectable parameter including “blood pressure” and/or “blood glucose”, wherein the visualization includes upper and/or lower limit lines representing upper and/or lower “safe” limits of the parameter being displayed. 2. The method of claim 1, comprising presenting a message automatically on the display after a predetermined period of time during which the visualization appears on the display advising that the numeric input is not within guidelines. 3. The method of claim 2, comprising automatically contacting a medical caregiver responsive to the numeric input not being within guidelines. 4. The method of claim 3, comprising enabling a user to disable and enable automatically contacting a medical caregiver. 5. Apparatus comprising: processor configured for controlling a display and for communicating with the Internet through a network interface;the processor configured for accessing instructions on a computer readable storage medium for configuring the processor for executing logic including:receiving first input indicating an identity of a user;receiving second input representing a physical parameter of the user;storing a time associated with the second input;presenting on the display a visualization of the user's history of the measurable physical parameter; andoutputting for presentation on the display at least one selectable parameter selectable to input numbers representative thereof, the at least one selectable parameter including “blood pressure” and/or “blood glucose”, wherein the visualization includes a graph of time versus parameter value and upper and/or lower limit lines indicating upper and/or lower “safe” limits of the parameter being displayed. 6. The apparatus of claim 5, wherein the logic executed by the processor includes presenting on the display the at least one selectable parameters. 7. The apparatus of claim 5, wherein the visualization includes a graph of time on an x-axis versus parameter value on a y-axis. 8. Apparatus comprising: processor configured for controlling a display and for communicating with the Internet through a network interface;the processor configured for accessing instructions on a computer readable storage medium for executing the instructions for configuring the processor for:presenting a first user interface (UI) on the display enabling a person to log in to a health monitoring feature, such that the apparatus maintains health records by... | Sony Corporation | H04N007/173 G06F019/00 A61B005/024 A61B005/00 H04L029/08 G08B021/04 A61B005/02 H04N021/475 H04N021/478 H04N021/488 A61B019/00 A61B005/0205 A61B005/021 A61B005/145
US10376222 | B2 | 2019-08-13 | Method and system for functional balance assessment | 2015-11-06 | A system comprising: at least one platform configured for, in use, having a person on it, said platform comprising a plurality of sensors configured for capturing information of forces/pressures applied by the feet of the person or of the position of the person's feet; means for, from said captured information, respectively extracting either a center of pressure or the position of the feet; at least one reference signal disposed at a known location with respect to said platform; at least one sensor which, in use, is connected to the body of the person situated on said platform, the sensor being configured for capturing said at least one reference signal; processing means for extracting information from said at least one captured reference signal and for estimating the position and/or orientation of said at least one sensor; processing means for estimating the position and/or orientation of the body of said person; processing means for calculating a stability index from said position and/or orientation of the body of said person and from said extracted feet position or center of pressure. | 1. A system comprising: at least one platform configured for, in use, having a person on it, said platform comprising a plurality of sensors configured for capturing information of forces/pressures applied by the feet of the person or of the position of the person's feet;a computer processor configured for, from said captured information, respectively extracting either a center of pressure or the position of the feet;the system being characterized in that it further comprises:at least one reference module providing a reference signal, the at least one reference module being disposed at a fixed location with respect to said platform;at least one sensor which, in use, is connected to the body of the person situated on said platform, the at least one sensor being worn by said person, the sensor worn by said person capturing the reference signal provided by the at least one reference module;wherein the computer processor is further configured for:extracting information from said at least one reference signal captured by the at least one sensor worn by said person,estimating from said reference signal captured by the at least one sensor worn by said person the position and orientation of said at least one sensor worn by said person with respect to the at least one reference module providing said reference signal,estimating the position and orientation of the body of said person from said estimated position and orientation of the at least one sensor in turn obtained from said at least one reference signal, andcalculating a stability index from said position and orientation of the body of said person and from said extracted feet position or center of pressure. 2. The system of claim 1, wherein said at least one sensor connected to the body of the person situated on said platform comprises a camera and said at least one reference signal provided by said at least one reference module disposed at a fixed location with respect to said platform is a visual pattern. 3. The system of claim 2, wherein said at least one visual pattern is an infrared visual pattern and said at least one camera is configured to capture images in the infrared range of light. 4. The system of claim 1, wherein said platform is a plantar pressure distribution system or a force platform. 5. The system of claim 1, wherein said stability index is a function of the position and orientation of said at least one sensor connected to the body of the person situated on said platform and at least one of the following inputs: a base of support (BoS), a center of pressure (CoP), a ground reaction force (GRF) and one or more features derived from the base of support (BoS), the center of pressure (CoP) or the ground reaction force (GRF). 6. The system of claim 1, wherein said at least one sensor connected to the body of the person situated on said platform is adapted to be connected to the trunk of said person. 7. The system of claim 1, wherein said at least one sensor connected to the body o... | - | A61B005/00 A61B005/103 A61B005/11
US10405758 | B2 | 2019-09-10 | Carrier assembly for blood pressure module | 2015-11-06 | A carrier assembly includes a sub-assembly including a pump, at least one solenoid valve, and a check valve. The pump, the at least one solenoid valve, and the check valve are in fluid communication with each other and with at least one pressure transducer via a manifold. Also, the carrier assembly is configured to be mounted to a medical device's main assembly board. | 1. A blood pressure module, comprising: a blood pressure cuff connector;a main printed circuit assembly;a manifold positioned on the main printed circuit assembly and forming a fluid connection with the blood pressure cuff connector;at least one pressure transducer positioned on the main printed circuit assembly; anda carrier configured to be mounted to the main printed circuit assembly, wherein the carrier supports: a pump;at least one solenoid valve positioned adjacent to the pump;a check valve positioned adjacent to both the pump and the at least one solenoid valve, wherein the pump, the at least one solenoid valve and the check valve are in fluid communication with each other and with the at least one pressure transducer via the manifold;a speaker;a wireless radio; andan integrated wire harness holding at least one wire for each of: the pump, the at least one solenoid valve, the check valve, the speaker, and the wireless radio;wherein the main printed circuit assembly has a first surface area on a top surface of the main printed circuit assembly; andwherein the components supported by the carrier occupy an amount of surface area on the top surface that is at least less than half of the first surface area. 2. The blood pressure module of claim 1, wherein the carrier is carried on the main printed circuit assembly. 3. The blood pressure module of claim 2, wherein the carrier and the main printed circuit assembly form an integrated assembly. 4. The blood pressure module of claim 1, wherein the carrier comprises pump retention snaps for supporting the pump. 5. The blood pressure module of claim 1, wherein the carrier includes a speaker mount for the speaker and a wireless radio mount for the wireless radio. 6. The blood pressure module of claim 1, wherein the carrier includes an integrated pump/valve harness. | Welch Allyn, Inc. | A61B005/02 A61B005/022 A61B005/0235 A61B005/00 G02F001/1333 H02J001/00 H02J007/00 A61B005/0205 H05K001/02 A61B005/021 A61B050/13 H05K001/18
US20160120452 | A1 | 2016-05-05 | BLOOD GLUCOSE MANAGEMENT | 2016-01-14 | This disclosure provides equipment and processes for blood glucose management. Embodiments may comprise a lancet device that includes a base that supports a lancet and a cover. The cover can move controllably from an un-depressed position to a depressed position when a predetermined force is applied. The cover can provide access to the lancet in the depressed position. In another aspect, a biological test kit may include a reusable module and a disposable module. The disposable module may include one or more lancet devices such as that described above to collect a biological sample, one or more strip stations supporting biological strips and a disposable module connector. The reusable module may support measuring equipment to measure a property of the biological sample and a reusable module connector that engages removably with the disposable module connector. | 1. A lancet device comprising: (a) a base that includes a support surface;(b) a lancet supported by the base and extending away from the support surface at an angle of approximately 90°;(c) a cover that includes (i) a peripheral edge along which the cover is connected to the base, (ii) a cover region configured to cover the lancet when the cover is in an un-depressed position, the cover region having a cover region geometry, (iii) a lancet aperture through which the lancet is configured to protrude when the cover is in a depressed position, and (iv) a peripheral hinge near the peripheral edge, the peripheral hinge having a peripheral hinge geometry that differs from the cover region geometry to facilitate controlled movement of the cover from the un-depressed position to the depressed position upon application of a predetermined depression force. 2-26. (canceled) |  | A61B005/15 A61B005/151
US20160128618 | A1 | 2016-05-12 | DIAGNOSTIC APPARATUS USING HABIT, DIAGNOSIS MANAGEMENT APPARATUS, AND DIAGNOSTIC METHOD USING SAME | 2016-01-19 | A diagnostic apparatus and a diagnosis management apparatus that utilizes habit data, and a diagnostic method using the same are provided. The diagnostic apparatus includes a habit analyzer configured to generate habit data of a user by analyzing sensor data detected from at least one sensor, and a diagnoser configured to determine whether the user is at risk of a disease, by comparing the generated habit data with diagnostic data including habit data of healthy people. | 1. A diagnostic apparatus, comprising: a habit analyzer configured to generate habit data of a user by analyzing sensor data detected from at least one sensor; anda diagnoser configured to determine whether the user is at risk of a disease, by comparing the generated habit data with diagnostic data stored in a memory, wherein the diagnostic data comprises habit data of healthy people. 2. The diagnostic apparatus of claim 1, wherein the habit analyzer comprises a log analyzing module configured to generate the habit data by analyzing log data stored in a usage log. 3. The diagnostic apparatus of claim 1, wherein the habit analyzer comprises an input analyzing module configured to generate the habit data by analyzing data input by the user. 4. The apparatus diagnostic of claim 1, wherein the habit analyzer is configured to generate the habit data in a normalized form. 5. The diagnostic apparatus of claim 1, wherein the diagnoser comprises a search module configured to search the diagnostic data that matches profile information of the user. 6. The apparatus diagnostic of claim 1, wherein the diagnoser is configured to, in response to a differential value between the habit data and the diagnostic data being greater than a preset threshold, determine that the user is at risk of the disease. 7. The diagnostic apparatus of claim 1, further comprising: a memory storage configured to store the habit data that are generated at every preset cycle,wherein the diagnoser comprises a tendency analyzing module configured to determine whether a differential value between the stored habit data and the diagnostic data has a tendency to increase, and, in response to determining that the difference has a tendency to increase, determine that the user is at risk of the disease. 8. The diagnostic apparatus of claim 1, further comprising: a memory storage configured to store habit data that are generated at every preset cycle,wherein the diagnoser comprises a correlation analyzing module configured to transform a change in the habit data of the user stored in the storage into a sequence and analyze correlation between the transformed sequence with a sequence indicating a change in habit data of a patient suffering from a specific disease, and to, in response to the correlation being greater than a preset threshold, determine that the user is at risk of the specific disease. 9. The diagnostic apparatus of claim 1, further comprising: a preventive measure provider configured to, in response to a determination that the user is at risk of the disease, either provide the user with information on the disease or inform a doctor or a family member of a result of the determination. 10. A diagnostic management apparatus, comprising: a habit manager configured to generate habit data of a user by analyzing behavior data of the user received from a diagnostic apparatus; anda diagnosis manager configured to determine whether the user is at risk of a disease, by comparing the ... | Samsung Electronics Co., Ltd. | A61B005/16 A61B005/11 A61B005/00
US20160143582 | A1 | 2016-05-26 | Wearable Food Consumption Monitor | 2014-11-22 | This invention is a device or system to monitor a person's food consumption and/or caloric intake which is worn like a hearing aid, ear-mounted bluetooth device, ear bud, or ear ring. This invention comprises: a sensor-positioning member which is worn on, in, or around a person's ear; a sensor which detects when the person is eating; an imaging member which is triggered to take pictures of food when the person is eating; and an image-analyzing member which analyzes the pictures of food to estimate the person's food consumption and/or caloric intake. | 1. A device which automatically monitors a person's food consumption and/or caloric intake comprising: a sensor-positioning member which is configured to be worn on, in, and/or around a person's ear;at least one sensor which is positioned by the sensor-positioning member, wherein this sensor collects data which is used to detect when the person is eating;at least one imaging member which is automatically triggered to take pictures of food when data from the at least one sensor indicates that the person is eating; andan image-analyzing member which automatically analyzes the pictures of food. 2. The device in claim 1 wherein the device is worn like a hearing aid, ear bud, and/or ear-mounted bluetooth device. 3. The device in claim 1 wherein the device is worn like an ear ring. 4. The device in claim 1 wherein the device is inserted into the person's ear. 5. The device in claim 1 wherein the device loops around a portion of the perimeter of the person's ear. 6. The device in claim 1 wherein at least one of the sensors is a motion sensor. 7. The device in claim 1 wherein at least one of the sensors is an accelerometer. 8. The device in claim 1 wherein at least one of the sensors is a strain gauge. 9. The device in claim 1 wherein at least one of the sensors is a piezoelectric sensor. 10. The device in claim 1 wherein at least one of the sensors is a sound sensor. 11. The device in claim 1 wherein at least one of the sensors is a microphone. 12. The device in claim 1 wherein at least one of the sensors is an electromagnetic energy sensor. 13. The device in claim 1 wherein at least one of the sensors is an EEG or EMG sensor. 14. The device in claim 1 wherein at least one of the sensors is a chew sensor. 15. The device in claim 1 wherein at least one of the sensors is a swallow sensor. 16. The device in claim 1 wherein at least one of the sensors is a GPS and/or location sensor. 17. The device in claim 1 wherein the imaging-analyzing member automatically estimates the types and/or quantities of food consumed by the person. 18. The device in claim 1 wherein the image-analyzing member uses one or more methods selected from the group consisting of: pattern recognition or identification; human motion recognition or identification; face recognition or identification; gesture recognition or identification; food recognition or identification; word recognition or identification; logo recognition or identification; bar code recognition or identification; and 3D modeling. 19. The device in claim 1 wherein the image-analyzing member analyzes one or more factors selected from the group consisting of: number of reachable food sources; types of reachable food sources; changes in the volume of food at a reachable food source; number of times that the person brings food to their mouth; sizes of portions of food that the person brings to their mouth; number of chewing movements; frequency of speed of chewing movements; and number of swallowing movemen... | Medibotics LLC | A61B005/00 A61B005/11 A61B005/117 A61B005/0476 A61B005/0488 A61B007/04 A61B019/00 A61B007/00
US20160143558 | A1 | 2016-05-26 | APPARATUS AND METHOD FOR ANALYZING BODY TISSUE LAYER IN ELECTRONIC DEVICE | 2015-11-24 | An electronic device, including a receiver configured to receive signals reflected from an object; and a controller configured to generate information corresponding to at least one tissue layer of the object based on the signals and a plurality of positions of the electronic device, wherein the plurality of positions are determined while the electronic device moves. | 1. An electronic device comprising: a receiver configured to receive signals reflected from an object; anda controller configured to generate information corresponding to at least one tissue layer of the object based on the signals and a plurality of positions of the electronic device,wherein the plurality of positions are determined while the electronic device moves. 2. The electronic device of claim 1, wherein the object comprises a body, wherein the at least one tissue layer comprises at least one from among a muscle, a skin and a fat, andwherein the information comprises a thickness associated with the at least one tissue layer. 3. The electronic device of claim 1, further comprising: a transmitter configured to radiate the signals to the object while the electronic device moves along a surface of the object. 4. The electronic device of claim 1, further comprising: a sensor configured to determine the plurality of positions while the electronic device moves. 5. The electronic device of claim 1, further comprising: a display configured to display an image representing the information. 6. The electronic device of claim 1, further comprising: a communicator configured to transmit the information to another electronic device. 7. The electronic device of claim 1, further comprising: at least one antenna configured to radiate the signals and to detect the signals reflected from the object,wherein the at least one antenna comprises flexible materials. 8. The electronic device of claim 1, further comprising: a reference coupler configured to generate a marker signal for a calibration relating to a signal delay associated with the signals. 9. The electronic device of claim 1, wherein the controller is further configured to measure a magnitude attenuation and a phase delay of the signals. 10. The electronic device of claim 1, wherein the information is generated based on a magnitude attenuation and a phase delay of the signals, and an estimation of signal attenuation corresponding to a thickness of the at least one tissue layer. 11. A method for operating an electronic device, the method comprising: receiving signals reflected from an object; andgenerating information corresponding to at least one tissue layer of the object based on the signals and a plurality of positions of the electronic device,wherein the plurality of positions are determined while the electronic device moves. 12. The method of claim 11, wherein the object comprises a body, wherein the at least one tissue layer comprises at least one from among a muscle, a skin and a fat, andwherein the information comprises a thickness associated with the at least one tissue layer. 13. The method of claim 11, further comprising: radiating the signals to the object while the electronic device moves along a surface of the object. 14. The method of claim 11, further comprising: determining the plurality of positions while the electronic device moves. 15. The method of claim 11, further ... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/05 G01S013/02 G01S007/41 A61B005/107
US09329046 | B2 | 2016-05-03 | Methods and systems generating driver workload data | 2013-07-01 | A method of determining data indicative of an individual driver workload is disclosed. The method involves obtaining a generic driver workload parameter, the generic driver workload data being based on data indicative of the behavior of multiple drivers when performing a maneuver at a node or traversing a set of one or more segments of an electronic map. An individual modifier representative of the behavior of an individual driver is generated and used to determine an individual driver workload parameter indicative of the workload of the individual driver when performing the maneuver at a node or traversing the one or more segments of an electronic map based on the generic driver workload data for performing the maneuver at the node or traversing the one or more segments of the electronic map. | 1. A method of determining data indicative of an individual driver workload, the method comprising: obtaining generic driver workload data, the generic driver workload data being representative of a generic driver workload when performing a manoeuvre at a node or traversing a set of one or more segments of an electronic map;generating an individual modifier representative of the behaviour of an individual driver, the individual modifier being proportional to likely deviation of an individual workload of the driver from corresponding generic workloads of drivers when performing manoeuvres or traversing segments; andusing the individual modifier to determine individual driver workload data indicative of the workload of the individual driver when performing the manoeuvre at the node or traversing the one or more segments of the electronic map based on the generic driver workload data for the manoeuvre at the node or the traversal of the one or more segments. 2. The method of claim 1, wherein the generic driver workload data is a generic driver workload parameter, and the individual driver workload data is an individual driver workload parameter. 3. The method of claim 1, wherein the generic driver workload data is based upon data indicative of the behaviour of multiple drivers when performing the manoeuvre and/or traversing the one or more segments. 4. The method of claim 1, wherein the generic driver workload data forms part of a generic driver workload profile comprising data indicative of a generic workload of drivers when performing each of plurality of different manoeuvres at a node or nodes and/or traversing a plurality of sets of one or more segments of the electronic map, wherein the generic driver workload data for a given manoeuvre and/or traversal of a set of one or more segments is associated with electronic map data indicative of the manoeuvre or the set of one or more segments. 5. The method of claim 1, wherein the step of determining the individual modifier comprises generating the individual modifier using data representative of the real-time or historic behaviour of the individual driver, or combinations thereof. 6. The method of claim 5, further comprising obtaining the data representative of the behaviour of the individual driver, wherein the data collected automatically during driving by the individual driver. 7. The method of claim 1, wherein the individual modifier is generated using data indicative of any one of the acceleration, deceleration, curve aggression factor, speed or braking of the driver. 8. The method of claim 1, wherein the manoeuvre at the node or traversal of a set of one or more segments of the electronic map is a first manoeuvre at a node or traversal of a set of one or more segments of the electronic map, and the method further comprises obtaining data indicative of an individual driver workload for at least one further manoeuvre at a node and/or the traversal of at least one further set of one or more... | TomTom Navigation B.V. | B60W040/00 G01C021/00 G01C021/26 A61B005/18 G01C021/34 G01C021/36 B60W040/09 A61B005/053 A61B005/11
US09351999 | B2 | 2016-05-31 | Use of platelet rich plasma composition in the treatment of cardiac conduction abnormalities | 2014-08-29 | Methods and kits for treating a cardiac arrhythmia using a platelet rich plasma (PRP) composition are provided. Any type of arrhythmia may be treated using the PRP composition. The PRP composition may comprise PRP developed using blood collected from a patient suffering the cardiac arrhythmia. The PRP composition may be buffered to a physiological pH and may include one or more anti-arrhythmic agents, anti-coagulants, or other drugs. The PRP composition may be delivered using a nebulizer, minimally invasively, or surgically. | 1. A method for treating a cardiac conduction abnormality, comprising: identifying a cardiac conduction abnormality in a patient; anddelivering a platelet rich plasma composition to the patient, wherein the platelet rich plasma composition is in an inhalable form. 2. The method of claim 1, wherein the conduction abnormality is a cardiac arrhythmia. 3. The method of claim 1, wherein the patient has a negative cardiac enzyme elevation. 4. The method of claim 1, further comprising monitoring the conduction abnormality after delivering the platelet rich plasma composition. 5. The method of claim 1, wherein the conduction abnormality is selected from the group consisting of bradycardia, tachycardia, ventricular tachycardia, and ventricular fibrillation. 6. The method of claim 1, wherein the conduction abnormality is chronic, acute or episodic. 7. The method of claim 1, wherein determining a cardiac conduction abnormality exists comprises using an electrocardiogram, Holter monitor or a cardiac event monitor. 8. The method of claim 1, wherein the platelet rich plasma composition is prepared from whole blood of the patient. 9. The method of claim 1, wherein the platelet rich plasma composition is buffered to a physiological pH, and wherein the physiological pH is between about 7.3 and about 7.5. 10. The method of claim 1, wherein the inhalable form is an inhalable powder. 11. The method of claim 1, wherein the treatment further comprises administering an anti-arrhythmic agent selected from the group consisting of a sodium channel blocker, a beta blocker, a potassium channel blocker, and a calcium channel blocker. 12. The method of claim 1, wherein the inhalable form is delivered with a nebulizer or inhaler. 13. The method of claim 1, wherein the conduction abnormality is due at least in part to reperfusion therapy. |  | A61K035/16 A61N001/36 A61N001/362 A61B005/0402 A61B005/00 A61K045/06 A61M015/00 A61M016/14
US20160128628 | A1 | 2016-05-12 | PROJECTED CAPACITIVE DETECTING SYSTEM FOR DETECTING HUMAN ACTIVITIES | 2015-11-03 | A projected capacitive detecting system comprises a sensing pad device and a monitoring module. The sensing pad device comprises at least one sensing array unit and at least one control unit. The at least one sensing array unit comprises multiple sensing electrodes arranged in an array. The at least one control unit is electrically connected to the multiple sensing electrodes to receive capacitances of the multiple sensing electrodes. The at least one control unit performs a capacitance conversion to convert the capacitances of the multiple sensing electrodes to obtain multiple capacitive sensing values and digitize the multiple capacitive sensing values to obtain Boolean true/false results. The monitoring module is electrically connected to the at least one control unit to receive and display the multiple capacitive sensing values and the Boolean true/false results as references for determining human activities. | 1. A projected capacitive detecting system for detecting human activities, the projected capacitive detecting system comprising: a sensing pad device comprising: at least one sensing array unit comprising multiple sensing electrodes arranged in an array;at least one control unit electrically connected to the multiple sensing electrodes to receive capacitances of the multiple sensing electrodes; andthe at least one control unit performing a capacitance conversion to convert the capacitances of the multiple sensing electrodes to obtain multiple capacitive sensing values and digitize the multiple capacitive sensing values to obtain Boolean true/false results; anda monitoring module electrically connected to the at least one control unit to receive and display the multiple capacitive sensing values and the Boolean true/false results as references for determining human activities. 2. The projected capacitive detecting system as claimed in claim 1, wherein when the at least one control unit digitizes each one of the multiple capacitive sensing values, the at least one control unit determines each one of the capacitive sensing values higher than a reference value as a first logic level and determines each one of the capacitive sensing values lower than the reference value as a second logic level complementary to the first logic level. 3. The projected capacitive detecting system as claimed in claim 2, wherein the monitoring module calculates a total number of the sensing electrodes determined to be the first logic level and determines whether the total number is higher than a threshold value; when the total number is higher than the threshold value, the monitoring module determines that a human lies supine; when the total number is lower than the threshold value, the monitoring module determines that the human is in a side-lying position;the at least one sensing array unit of the sensing pad device includes multiple sensing array units respectively representing different body portions of the human, such that the monitoring module determines movements of the different body portions of the human according to the Boolean true/false results of the multiple sensing array units. 4. The projected capacitive detecting system as claimed in claim 3, wherein the monitoring module calculates a number of times of turning over and movement of the human as references to determine the human's activity, sleep quality, and sleep stage. 5. The projected capacitive detecting system as claimed in claim 3, wherein the monitoring module filters the capacitive sensing values of the sensing electrodes determined to be the first logic level by a band-pass filtering process to obtain multiple filtered results corresponding to a human breathing frequency. 6. The projected capacitive detecting system as claimed in claim 3, wherein the at least one control unit subtracts a base value from the capacitive sensing value of each one of the sensing electrodes to obtain a relative ca... |  | A61B005/00 A61B005/113 A61B005/08
US20160125766 | A1 | 2016-05-05 | PHOTOACOUSTIC BLOOD MODEL | 2014-05-20 | The photoacoustic blood model contains two or more kinds of light absorbing compounds in a blood model base material, in which the absorption coefficient ratios μ[λ2]/μ[λ1] at arbitrary two wavelengths λ1 and λ2 (λ12) of 600 nm or more and 1100 nm or less of the compounds are different from each other and the parameter S calculated from Equation (1) is 0 or more and 100 or less, in which HbO2[λ1] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ1, HbO2[λ2] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ2, Hb[λ1] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ1, Hb[λ2] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ2, and P′ indicates a ratio (Pλ2/Pλ1) of a photoacoustic signal intensity Pλ2 obtained by irradiation with light of the wavelength λ2 to a photoacoustic signal intensity Pλ1 obtained by irradiation with light of the wavelength λ1. | 1. A photoacoustic blood model, comprising: two or more kinds of light absorbing compounds in a blood model base material,wherein absorption coefficient ratios μ[λ2]/μ[λ1] at arbitrary two wavelengths λ1 and λ2 (λ1&lt;λ2) of 600 nm or more and 1100 nm or less of the light absorbing compounds are different from each other and a parameter S calculated from the following equation (1) is 0 or more and 100 or less, andwherein a coefficient of linear thermal expansion of the blood model base material is 100 ppm/K or more and 1000 ppm/K or less,[Math.1]S=P′·Hb[λ1]-Hb[λ2](HbO2[λ2]-Hb[λ2])-P′·(HbO2[λ1]-Hb[λ1])·100Expression(1)wherein HbO2[λ1] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ1, HbO2[λ2] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ2, Hb[λ1] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ1, Hb[λ2] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ2, and P′ indicates a ratio (Pλ2/Pλ1) of a photoacoustic signal intensity Pλ2 obtained by irradiation with light of the wavelength λ2 to a photoacoustic signal intensity Pλ1 obtained by irradiation with light of the wavelength λ1. 2. The photoacoustic blood model according to claim 1, comprising at least one light absorbing compound in which the absorption coefficient ratio μ[λ2]/μ[λ1] satisfies one of the following expression (2) or expression (3),[Math.2]Smin≥(μ[λ2]/μ[λ1])·Hb[λ1]-Hb[λ2](HbO2[λ2]-Hb[λ2])-(μ[λ2]/μ[λ1])·(HbO2[λ1]-Hb[λ1])·100Expression(2)Smax≤(μ[λ2]/μ[λ1])·Hb[λ1]-Hb[λ2](HbO2[λ2]-Hb[λ2])-(μ[λ2]/μ[λ1])·(HbO2[λ1]-Hb[λ1])·100Expression(3)wherein Smin indicates a lower limit of the parameter S, Smax indicates an upper limit of the parameter S, HbO2[λ1] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ1, HbO2[λ2] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ2, Hb[λ1] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ1, Hb[λ2] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ2, μ[λ1] indicates an absorption coefficient of the light absorbing compound at the wavelength λ1, and μ[λ2] indicates an absorption coefficient of the light absorbing compound at the wavelength λ2. 3. The photoacoustic blood model according to claim 2, further comprising at least one light absorbing compound in which the absorption coefficient ratio μ[λ2]/μ[λ1] satisfies the other one of the expression (2) or the expression (3). 4. (canceled) 5. The photoacoustic blood model according to claim 1, wherein an acoustic velocity of the blood model base material is 800 m/s or more and 2000 m/s or less. 6. The photoacoustic blood model according to claim 5, wherein the acoustic velocity of the blood model base material is 1300 m/s or more and 1700 m/s or less. 7. The photoacoustic blood model according to claim 1, wherein the blood model base material contains a polymer material. 8. The p... | CANON KABUSHIKI KAISHA | G09B023/28 A61B005/00 G01N029/30
US09339595 | B2 | 2016-05-17 | Disposable vacuum source and medical applications thereof | 2013-08-03 | An apparatus including: a vacuum container configured to provide and substantially maintain a built-in vacuum pressure; a contact member configured to contact and substantially enclose a surface area; a vacuum device configured to couple the built-in vacuum pressure to the surface area thereby applying a suction force over the surface area; and a conduit coupled to the vacuum container; wherein the vacuum container is also configured to retain a substance collected from adjacent the surface area, wherein the conduit is configured to direct the substance to the vacuum container, and to maintain substantially the same level of vacuum pressure as the vacuum container, wherein the vacuum container comprises a collapsible container, and wherein the collapsible container comprises a spring member configured to maintain a tension in a wall of the vacuum container. | 1. An apparatus, comprising: a vacuum container configured to provide and substantially maintain a built-in vacuum pressure;a contact member configured to contact and substantially enclose a surface area;a vacuum device configured to couple the built-in vacuum pressure to the surface area thereby applying a suction force over the surface area; anda conduit coupled to the vacuum container;wherein the vacuum container is also configured to retain a substance collected from adjacent the surface area,wherein the conduit is configured to direct the substance to the vacuum container, and to maintain substantially the same level of vacuum pressure as the vacuum container,wherein the vacuum container comprises a collapsible container, and wherein the collapsible container comprises a spring member configured to maintain a tension in a wall of the vacuum container. 2. The apparatus of claim 1, wherein the vacuum device comprises: a first valve configured to, if opened, couple the built-in vacuum pressure to the surface area; anda second valve configured to equalize a pressure over the surface area with an ambient atmospheric pressure thereby releasing the suction force. 3. The apparatus of claim 2, wherein the first valve, if closed, is also configured to maintain a pressure difference between the vacuum container and the surface area. 4. The apparatus of claim 1, wherein the vacuum device comprises a first vacuum valve configured to equalize the built-in vacuum pressure with the pressure on the surface area. 5. The apparatus of claim 4, wherein the vacuum device further comprises a second vacuum valve configured to release the suction force from the surface area. 6. The apparatus of claim 1, further comprising a needle and a catheter, wherein the suction force is configured for fixation of the needle and the catheter, and wherein the apparatus is configured as an anesthesia apparatus. 7. The apparatus of claim 1, further comprising a needle, wherein the apparatus is configured as a disposable lancing apparatus. 8. The apparatus of claim 1, further comprising a suction tube, and wherein the apparatus is configured as a uterine tissue extraction device. 9. The apparatus of claim 1, wherein the vacuum device comprises a suction cup, and wherein the apparatus is configured as a fetal attraction apparatus. 10. A vacuum apparatus, comprising: a vacuum container configured to provide and substantially maintain a built-in vacuum pressure;a contact member configured to contact and substantially enclose a surface area;a vacuum device configured to couple the built-in vacuum pressure to the surface area thereby applying a suction force over the surface area;wherein the vacuum container is also configured to retain a substance collected from adjacent the surface area,wherein the vacuum container comprises a collapsible container, andwherein the collapsible container comprises a spring member configured to maintain a tension in a wall of the vacuum container... |  | A61M027/00 A61M001/00 A61H009/00 A61B005/15 A61B010/02 A61B017/44 A61H007/00 A61H039/04 A61M019/00
US09326691 | B1 | 2016-05-03 | Dispensable single-use protective film for a sphygmomanometer | 2008-11-25 | A pad of disposable protective barrier films for adhesive attached to a blood pressure cuff prior to its placement on a limb of a subject. The barrier film is a thin, non-permeable polymer, typically having a thickness of between approximately 1.0 and 3.0 mil. A portion of the side of the barrier adapted for attachment of the blood pressure cuff is coated with a low tack adhesive. The low tack adhesive holds the barrier in place during installation and use of the blood pressure but allows easy removal and disposal thereof upon completion of the blood pressure measurement. A unique packaging system is provided whereby no adhesive backing strips are required, successive barrier sheets are merely peeled from a stack. In still other embodiments, the barrier or the adhesive may include a germicidal material to further help protect the blood pressure cuff to from germs or other contamination. | 1. A plurality of protective barrier films for covering a central region of a blood pressure cuff, comprising: a) a plurality of thin, single-layer films adapted and configured for placement against an inner surface of a central region along a major axis of a blood pressure cuff; andb) adhesive placed on a portion of a single surface of each of said thin films, said single surface being adapted for releasable contact with said inner surface of said central region of said blood pressure cuff;each of said plurality of thin films being placed against another in a padded dispensing configuration whereby an upper most one of said plurality of thin films is readily removable from said pad. 2. The plurality of protective barrier films for a blood pressure cuff as recited in claim 1, wherein each of said plurality of thin films is impermeable. 3. The plurality of protective barrier films for a blood pressure cuff as recited in claim 2, wherein each of said plurality of thin films comprises a polymer film. 4. The plurality of protective barrier films for a blood pressure cuff as recited in claim 3, wherein each of said thin films comprises a polymer film having a thickness in the range of approximately 0.001 inch to 0.003 inch. 5. The plurality of protective barrier films films for a blood pressure cuff as recited in claim 1, wherein said adhesive is disposed adjacent at least one edge of each of said thin films parallel to a major axis thereof. 6. The plurality of protective barrier films for a blood pressure cuff as recited in claim 5, wherein said adhesive is disposed adjacent both edges of said thin film parallel to a major axis thereof. 7. The plurality of protective barrier films for a blood pressure cuff as recited in claim 1, wherein said adhesive is disposed adjacent at least one edge of each of said thin films substantially perpendicular to a major axis thereof. 8. The plurality of protective barrier films films for a blood pressure cuff as recited in claim 1, wherein at least one of said thin films is absent a backing membrane in a region containing said adhesive. 9. The plurality of protective barrier films for a blood pressure cuff as recited in claim 1, wherein at least one of said plurality of thin films comprises a germicidal material. 10. The plurality of protective barrier films for a blood pressure cuff as recited in claim 1, wherein said adhesive comprises at least one material chosen from the group: an antibacterial material, a germicidal material, and an antifungal material. 11. The plurality of protective barrier films for a blood pressure cuff as recited in claim 1, wherein said packaging configuration comprises a fan-fold configuration. 12. A method of preventing contamination of a blood pressure cuff, the steps comprising: a) providing the plurality of protective barrier films of claim 1;b) removing one of said thin films from said pad;c) placing said removed thin film over an inner surface of a central region of a blo... |  | A61B005/021 B32B005/00
US09339192 | B2 | 2016-05-17 | Systems and methods for minimally-invasive optical-acoustic imaging | 2015-08-26 | An imaging guidewire can include one or more optical fibers communicating light along the guidewire. At or near its distal end, one or more blazed or other fiber Bragg gratings (FBGs) directs light to a photoacoustic transducer material that provides ultrasonic imaging energy. Returned ultrasound is sensed by an FBG sensor. A responsive signal is optically communicated to the proximal end of the guidewire, and processed to develop a 2D or 3D image. In one example, the guidewire outer diameter is small enough such that an intravascular catheter can be passed over the guidewire. Techniques for improving ultrasound reception include using a high compliance material, resonating the ultrasound sensing transducer, using an attenuation-reducing coating and/or thickness, and/or using optical wavelength discrimination. Techniques for improving the ultrasound generating transducer include using a blazed FBG, designing the photoacoustic material thickness to enhance optical absorption. Techniques for distinguishing plaque or vulnerable plaque may be used to enhance the displayed image. | 1. A method of imaging a blood vessel of a subject comprising: providing an elongated body sized and shaped to be inserted into the blood vessel, the elongated body having proximal and distal ends, the elongated body including a plurality of optical fibers extending lengthwise along the elongated body, each optical fiber including at least one imaging portion including an optical-to-acoustic transducer and an acoustic-to-optical transducer;sequentially transmitting light along the plurality of optical fibers;sequentially transducing the transmitted light into acoustic energy;sequentially providing the acoustic energy to a region of interest in the vasculature;sequentially receiving at the imaging portion responsive acoustic energy from the region of interest;sequentially transducing the received acoustic energy into responsive light;sequentially communicating the responsive light along the plurality of optical fibers; andforming a viewable image of at least a portion of the region of interest based at least in part on the responsive light. 2. The method of claim 1, in which the forming a viewable image of at least a portion of the region of interest includes forming an image from a plurality of the acoustic-to-optical transducers. 3. The method of claim 1, comprising: sequentially accessing selected optical fibers. 4. The method of claim 1, wherein the at least one imaging portion includes at least two imaging portions per each of the plurality of optical fibers, and in which the forming a viewable image of at least a portion of the region of interest includes forming a three-dimensional image of the region of interest. 5. The method of claim 4, in which the forming a viewable image of at least a portion of the region of interest includes: sequentially accessing selected optical fibers; andsequentially accessing selected imaging portions. 6. The method of claim 5, wherein sequentially accessing selected imaging portions includes selectively individually addressing the different imaging portions. 7. The method of claim 6, selectively individually addressing the different imaging portions includes individually addressing the different imaging portions using different wavelengths of light. 8. The method of claim 1, wherein sequentially transmitting light along the plurality of optical fibers includes sequentially transmitting a first optical signal along the plurality of optical fibers, and wherein sequentially communicating the responsive light along the plurality of optical fibers includes sequentially providing responsive modulation of a second optical signal. 9. The method of claim 1, comprising: identifying a plaque deposit,wherein forming a viewable image of at least a portion of the region of interest includes providing a distinctive display of the plaque deposit. 10. The method of claim 1, comprising: inserting a guidewire including the elongated body into the blood vessel of the subject. | Vascular Imaging Corporation | A61B005/00 A61B001/00 A61B008/12 G01H009/00 A61B005/02
US09345407 | B2 | 2016-05-24 | Method and apparatus for analyzing subsurfaces of a target material | 2015-03-10 | A system that incorporates teachings of the present disclosure may include, for example, a method for aligning first and second light signals on an optical path directed to a target, where the first light signal provides a visualization of the target, and a portion of the second light signal reflects from at least one subsurface of the target. The method also includes aligning a first focal point of the first light signal and a second focal point of the second light signal, where the first focal point is at least in a first proximate location of the second focal point, and adjusting a first position of the first and second focal points to be in at least a second proximate location of the target without adjusting the at least first proximate location of the first focal point relative to the second focal point. Other embodiments are disclosed. | 1. A device, comprising: a first light source for emitting a first light signal;a second light source for emitting a second light signal, wherein the first light signal and the second light signal differ; andan optical component that coaxially aligns the first light signal and the second light signal, wherein the optical component maintains a first focal point of the first light signal in at least a proximate location of a second focal point of the second light signal, and wherein the optical component further maintains the first focal point at the proximate location of the second focal point after a focal adjustment of the first focal point or the second focal point. 2. The device of claim 1, wherein the first light signal operates in a first region of a light spectrum, and wherein the second light signal operates in a second region of the light spectrum. 3. The device of claim 2, wherein the first region and the second region differ. 4. The device of claim 1, wherein the first light signal enables visualization of a target. 5. The device of claim 1, wherein the second light signal reflects a portion of the second light signal from a subsurface of a target. 6. The device of claim 5, wherein the target comprises a tissue. 7. The device of claim 5, wherein the portion of the second light signal reflected from the subsurface of the target when processed by a computing device identifies an aberration of the target. 8. The device of claim 5, wherein the portion of the second light signal reflected from the subsurface of the target is directed to a system that detects a signal from the subsurface of the target. 9. The device of claim 8, wherein the system comprises a spectrometer that processes the signal. 10. The device of claim 1, comprising a beam splitter incident with a coaxial path of the first light signal and the second light signal. 11. The device of claim 1, wherein the first light source comprises a light emitting diode. 12. The device of claim 1, wherein the second light source comprises a laser. 13. A method, comprising: receiving from a device a portion of a first light signal reflected from a subsurface of a target, wherein the device coaxially aligns the first light signal and a second light signal and maintains a first focal point of the first light signal in at least a proximate location of a second focal point of the second light signal;directing the device to perform a focal point adjustment of the first focal point or the second focal point, wherein the device maintains the first focal point at the proximate location of the second focal point after the focal point adjustment;receiving from the device an updated instance of the first light signal reflected from the subsurface of the target responsive to the focal point adjustment; andproducing from the updated instance of the first light signal an output for analyzing the subsurface. 14. The method of claim 13, further comprising determining from the output a presence ... | WELCH ALLYN, INC. | A61B005/00 G01B009/02
US20160135736 | A1 | 2016-05-19 | MONITORING TREATMENT COMPLIANCE USING SPEECH PATTERNS CAPTURED DURING USE OF A COMMUNICATION SYSTEM | 2015-06-03 | Methods and systems for monitoring compliance of a patient with a prescribed treatment regimen are described. Patient speech is detected during use of a communication system such as a mobile telephone and analyzed to determine compliance with a treatment for a brain-related disorder, for example. Speech data representing one or more patient speech pattern and an identity signal containing information used to determine presence/identity of the patient are transmitted from a circuitry-based system at the patient location to a monitoring location. Identity of the patient as user of the communication system is determined through, e.g., biometric or authentication techniques. Speech data is analyzed to determine whether a patient speech pattern matches one or more characteristic speech patterns. Outcome of the analysis is reported to a medical caregiver or other party, for example. | 1. A system comprising: at least one receiving device for use at a monitoring location for receiving a speech data signal transmitted to the monitoring location from a patient location, the speech data signal containing speech data, the speech data representing at least one speech pattern in speech sensed from a patient with at least one audio sensor in a communication system at the patient location during use of the communication system by the patient, and the patient having a brain-related disorder and a prescribed treatment regimen for treating at least one aspect of the brain-related disorder;patient identification circuitry configured to determine a presence of the patient from at least one identity signal received at the monitoring location from the patient location;signal processing circuitry configured to analyze the speech data signal to determine whether the speech data represents at least one speech pattern that matches at least one characteristic speech pattern;compliance determination circuitry configured to determine whether the patient has complied with the prescribed treatment regimen based upon whether the speech data represents the at least one speech pattern that matches the at least one characteristic speech pattern; andreporting circuitry configured to report a conclusion based on the determination of whether the patient has complied with the prescribed treatment regimen. 2. The system of claim 1, further comprising: speech identification circuitry configured to identify patient speech data in the speech data based at least in part on the determination of the presence of the patient by the patient identification circuitry. 3. The system of claim 2, wherein the identity signal includes at least a portion of the speech data signal, wherein the patient identification circuitry is configured to analyze the speech data signal to determine at least a portion of the speech data signal containing speech data representing a speech pattern that matches a known speech pattern of the patient, and wherein the speech identification circuitry is configured to identify patient speech data by identifying the at least a portion of the speech data representing the speech pattern that matches a known speech pattern of the patient. 4.-21. (canceled) 22. The system of claim 1, wherein the signal processing circuitry includes a speech processor. 23. The system of claim 1, wherein the compliance determination circuitry includes a speech analyzer for analyzing the speech data to determine the speech pattern and a comparator for comparing the speech pattern represented by the speech data with the at least one characteristic speech pattern. 24. (canceled) 25. The system of claim 1, wherein the compliance determination circuitry includes a comparator for comparing the speech data with characteristic speech data representing the characteristic speech pattern. 26. (canceled) 27. The system of claim 1, wherein the compliance determination circuitr... | Elwha LLC, a limited liability comany of the State of Delaware | A61B005/00 G06Q050/24 G10L017/00 G06F019/00
US10383521 | B2 | 2019-08-20 | Non-cylindrical hypotubes | 2015-11-04 | A non-cylindrical hypotube is disclosed, such as for use in OCT and endoscopy. The hypotube is defined by a non-cylindrical, rotationally symmetric tube and has an interior, a proximal-end section with an outer diameter D1, a distal-end section with an outer diameter D3, and a middle section between the proximal-end and distal-end sections and having an outer diameter D2, wherein D21, and D23. The distal-end section is sized to accommodate the optical probe and includes an outer surface with an aperture that allows for optical communication therethrough. | 1. An optical coherence tomography (OCT) system, comprising: a hypotube comprising a non-cylindrical, rotationally symmetric rigid tube having an interior, a proximal-end section with a length L1 and an outer diameter D1, a distal-end section with a length L3 and an outer diameter D3 and an outer surface with an aperture, and a middle section between the proximal-end section and the distal-end section and having a length B and an outer diameter D2, and wherein the length B is greater than the length L1 and/or the length L3, and wherein D2&lt;D3;an optical probe operably arranged in the distal-end section relative to the aperture to allow for optical communication therethrough;an optical fiber operably connected to the optical probe at a fiber-probe interface, with the fiber-probe interface residing within the distal-end section of the hypotube;a torque coil operably connected to a proximal end of the proximal-end section of the of the hypotube; andan inner lumen that contains the hypotube and the torque coil, wherein the inner lumen has a critical bend radius that is smaller as compared to using a cylindrical hypotube having an outer diameter equal to D3. 2. The OCT system according to claim 1, wherein D2&lt;D1. 3. The OCT system according to claim 1, wherein D1=D3. 4. The OCT system according to claim 1, wherein the outer diameter D2 of the middle section varies as a function of length along the hypotube. 5. The OCT system according to claim 4, wherein the middle section has an hourglass shape. 6. The OCT system according to claim 1, wherein at least a portion of the middle section has a constant outer diameter D2. 7. The OCT system according to claim 1, wherein the middle section transitions to the distal-end section over a distal-end transition region wherein the hypotube varies in diameter over a transition length LT3 in the range 0 mm≤LT3≤0.5 mm. 8. The OCT system according to claim 7, wherein the middle section transitions to the proximal-end section over a proximal-end transition region wherein the hypotube varies in diameter over a transition length LT1 in the range 0 mm≤LT1≤0.5 mm. 9. The OCT system according to claim 8, wherein the proximal-end transition region is either curved or straight, and wherein the distal-end transition region is either curved or straight. 10. The OCT system according to claim 1, wherein the hypotube includes at least one of a proximal-end bevel at a proximal end of the proximal-end section and a distal-end bevel at a distal end of the distal-end section. 11. A hypotube assembly, comprising: an optical probe operably connected to an optical fiber at a fiber-probe interface;a hypotube defined by a non-cylindrical, rotationally symmetric tube with an interior and a proximal-end section with an outer diameter D1 and a length L1, a middle section with an outer diameter D2 and a length B, and a distal-end section with an outer diameter D3 and a length L3, an outer surface, and an aperture formed in the out... | Corning Incorporated | A61B005/00 A61B001/00
US10667744 | B2 | 2020-06-02 | Skin function evaluation device and skin evaluation method | 2014-06-20 | An evaluation method for determining a skin sensory evaluation value from an image, including: acquiring an image of an area including skin of a subject; extracting a skin region from the image; calculating at least two characteristic indices representing characteristics of the image of the skin region; and determining a skin sensory evaluation value based on the calculated at least two characteristic indices. | 1. An evaluation method, performed using circuitry, for determining a skin sensory evaluation value from an image captured using a camera, comprising: (a) acquiring an image of an area including skin of a subject using the camera;(b) using the circuitry, extracting a skin region from the image comprising: detecting a face of the subject in the image, andremoving, from the detected face, areas containing certain parts of the face based on a position of the areas of the detected face;(c) calculating using the circuitry at least two characteristic indices representing characteristics of the extracted skin region based on polarized light values and/or color components of the imagined skin region;(d) determining a skin sensory evaluation value based on the calculated at least two characteristic indices using the circuitry; and(e) displaying on a display device the determined skin sensory evaluation value, wherein: the at least two characteristic indices include at least two of a blemish amount, a wrinkle amount, and a luminance level,the sensory evaluation value is an evaluation value of a skin clarity, andthe evaluation value of the skin clarity is determined based on a correlation between the at least two characteristic indices measured on a plurality of subjects in advance and an evaluation value of the skin clarity determined by an evaluation performed on the skin of the plurality of subjects. 2. The evaluation method according to claim 1, wherein: in the step (b), the skin region is divided by the circuitry into a plurality of unit blocks;in the steps (b) and (c), the at least two characteristic indices are calculated by the circuitry to determine the skin sensory evaluation value on a unit block-by-unit block basis; andin step (e), the skin sensory evaluation value, found on a unit block-by-unit block basis by the circuitry, is displayed on the display device in association with a position of each of the unit blocks. 3. The evaluation method according to claim 2, wherein the step (e) displays, on the display device, the skin sensory evaluation value found on a unit block-by-unit block basis with a color tone or a gray scale suitable to the skin sensory evaluation value in association with the position of the corresponding unit block. 4. The evaluation method according to claim 1, wherein the correlation is found by multiple regression analysis. 5. The evaluation method according to claim 1, wherein information on beauty equipment or cosmetics regarding each of the calculated characteristic indices or the determined sensory evaluation value is further displayed on a display device. 6. A sensory evaluation device, comprising: a camera configured to acquire an image of an area including skin of a subject;circuitry configured to: extract a skin region from the image;calculate at least two characteristic indices representing characteristics of the extracted skin region based on polarized light and/or color components of the imagined skin region;... | Panasonic Intellectual Property Corporation of America | A61B005/00
US09341629 | B2 | 2016-05-17 | Synthesis and composition of amino acid linking groups conjugated to compounds used for the targeted imaging of tumors | 2015-05-19 | The present disclosure relates to compounds that are useful as near-infrared fluorescence probes, wherein the compounds include i) a pteroyl ligand that binds to a target receptor protein, ii) a dye molecule, and iii) a linker molecule that comprises an amino acid or derivative thereof. The disclosure further describes methods and compositions for making and using the compounds, methods incorporating the compounds, and kits incorporating the compounds. | 1. A kit comprising a compound having the formula:or a pharmaceutically acceptable salt thereof,or isotopes thereof, wherein:X is an single amino acid or a single amino acid derivative thereof, wherein the single amino acid or single amino acid derivative contains an —OH, —NH2, or —SH functional group, andY is a dye that has a fluorescence excitation and emission spectra in the near infra red range, wherein Y is represented by the formula:wherein, X′ is independently selected from the group consisting of O, S, N and C, andR′ is independently selected from the group consisting of CH2 and CH2CH2,and the compound maintains or enhances the fluorescence of the dye, Y. 2. The kit of claim 1 wherein the amino acid is selected from the group consisting of tyrosine, cysteine, lysine, serine or a derivative thereof of tyrosine or cysteine. 3. The kit of claim 1 wherein the amino acid is tyrosine, wherein the aromatic ring of tyrosine comprises a carbon isotope and/or a hydrogen isotope. 4. The kit of claim 1 wherein the amino acid derivative is a derivative of tyrosine selected from the group consisting of:or racemic mixtures thereof. 5. The kit of claim 1 wherein the amino acid including a sulfur-containing side chain group is cysteine. 6. The kit of claim 1 wherein the compound has the formula:or is a potassium, sodium, or ammonium salt thereof. 7. The kit of claim 1 wherein the compound has the formula:or a pharmaceutically acceptable salt thereof. 8. The kit of claim 1 wherein the compound has an absorption and emission maxima between about 600 nm and 800 nm. 9. The kit of claim 1 wherein the compound is made to fluoresce after distribution thereof in the tissue cells. 10. The kit of claim 1 wherein the compound is highly selective for targeting to a tumor cell. 11. The kit of claim 1, wherein the kit further comprises a pharmaceutically acceptable carrier, excipient or diluent. | Purdue Research Foundation | A61B005/00 A61B008/00 A61B010/00 G01N033/574 C07D475/04 G01N033/58 A61K049/00
US09327116 | B2 | 2016-05-03 | Wireless cardioresonance stimulation | 2015-02-02 | An apparatus for the cardio-synchronized stimulation of skeletal or smooth muscle, but excluding the heart muscles, in a counterpulsation mode of a patient. The apparatus comprises an active and a passive electrode for attachment to said patient, a signal processor having a configuration input for varying a time delay associated with counterpulsation mode stimulation, and a sensing system for sensing information relating to the performance of the patient's heart and for transmission of information signals to said signal processor, said signal processor producing control signal information relating to stimulation signals to be applied to said active electrode in a counterpulsation mode, a stimulation signal generator associated with said active electrode for generating stimulation signals, wireless transmission means for transmitting said control signal information from said signal processor to said stimulation signal generator whereby said stimulation signal generator applies stimulation signals to said active electrode in accordance with said signal information. | 1. An apparatus for cardio-synchronised stimulation of skeletal or smooth muscle, but excluding heart muscles, in a counterpulsation mode of a patient having a cardiovascular system and a heart having a heart rate, the apparatus comprising: at least one active electrode and at least one passive electrode configured for attachment to said patient at positions at which the electrodes are in communication with said skeletal or smooth muscle when said skeletal or smooth muscle is in its original, natural position within the patient;a signal processor having a configuration input and configured to produce control signal information relating to stimulation signals to be applied to said at least one active electrode in the counterpulsation mode; a sensing system configured to sense information relating to performance of the patient's heart, the information comprising at least the heart rate, and to transmit information signals relating to the information to said signal processor;a medical evaluation unit configured to transmit signal configuration information to the configuration input and said signal processor being adapted to take account of said signal configuration information when generating said control signal information,wherein the connection between the configuration input and the signal processor, and/or the connection between the configuration input and the medical evaluation unit can be realized by any device with wireless transmitter and/or receiver capabilities. 2. The apparatus in accordance with claim 1, wherein the device is selected from the group of devices comprising a mobile phone, a personal digital assistant with phone function, and any standard piece of equipment having a transceiver, a microprocessor, a memory for storing software and data, a battery, or other source of power and a clock; and wherein software relevant to the operation of the signal processor and data relating to the patient's heart rhythm and the performance of the heart and data relating to the stimulation applied or to be applied is stored in the device. 3. The apparatus in accordance with claim 1, wherein said information relating to the performance of the heart is selected from the group consisting of heart rate information, electrocardiogram (ECG) information, ECG derived information, ECG information and information resulting from electrical stimulation, ECG derived trigger signals, R-R information, end of T-wave information, blood pressure information, and blood pressure derived information. 4. The apparatus in accordance with claim 1, wherein said sensing system comprises at least one of an invasive sensor, an intercavity sensor, a non-invasive sensor, a body surface sensor, and a remote sensing system detached from the patient's body. 5. The apparatus in accordance with claim 3, further comprising a remote sensing system, wherein said signal processor is integrated into said remote sensing system. 6. The apparatus in accordance with claim 1, said si... | CARDIOLA LTD. | A61N001/00 A61N001/36 A61B005/00 A61B005/0245 A61B005/0456 G06F019/00
US10617343 | B2 | 2020-04-14 | Methods and systems for quantitative brain assessment | 2015-11-03 | Systems and methods for measuring macromolecular proton fraction in a subject are provided. A nuclear magnetic resonance apparatus applies a magnetic field to a body region on the subject, and radiofrequency modes are applied to the body region as well. Each radiofrequency mode delivers a plurality of radiofrequency pulses separated by time delays, wherein at least one of the radiofrequency modes causes suppression of signal components from an unwanted tissue, and at least one of the radiofrequency modes causes magnetization exchange between water and macromolecules in tissues in the body region. Amplitudes corresponding to magnetic signals received from the body region are measured and macromolecular proton fraction based on the amplitudes can be calculated. | 1. A method for measuring macromolecular proton fraction (MPF) in a subject, the method comprising: applying, via a nuclear magnetic resonance apparatus, a magnetic field to a body region of the subject, the body region including a first tissue and a second tissue;applying a first radiofrequency (RF) mode and a second RF mode to the body region such that (i) the first RF mode causes suppression of signal components returned from the first tissue, and (ii) the second RF mode causes magnetization exchange between water and macromolecules in the second tissue;detecting a plurality of magnetic signals emitted from the body region in response to applying the first RF mode and the second RF mode to the body region; andcalculating an MPF in the body region based on the plurality of magnetic signals. 2. The method of claim 1, wherein the body region encompasses one or both of a brain or a spinal cord of the subject. 3. The method of claim 2, wherein the first tissue does not belong to the brain of the subject. 4. The method of claim 3, wherein the first tissue comprises one or more of: (i) a fat tissue; (ii) a muscle tissue; and (iii) cerebrospinal fluid. 5. The method of claim 2, wherein applying the first RF mode and the second RF mode causes one or more of the following: (i) a saturation of a nuclear magnetization of fat protons outside the brain; (ii) a saturation of a nuclear magnetization of muscle protons outside the brain; (iii) a saturation of a nuclear magnetization of cerebrospinal fluid protons inside the brain; (iv) a variable excitation of a nuclear magnetization of protons of the brain; or (v) a variable saturation of a nuclear magnetization of macromolecular protons in the brain parenchyma. 6. The method of claim 2, wherein detecting the plurality of magnetic signals is performed after varying one or more of the following: (i) excitation of nuclear magnetization of protons of the brain; (ii) saturation of nuclear magnetization of macromolecular protons in brain parenchyma; or (iii) a duration of a time delay separating pulses of the first RF mode or the second RF mode. 7. The method of claim 1, wherein the subject is a human subject. 8. The method of claim 1, further comprising: comparing the MPF with a control. 9. The method of claim 1, further comprising: determining a myelin content in the body region based on the MPF. 10. The method of claim 1, further comprising: using the plurality of magnetic signals to calculate a proton density of the body region and a longitudinal relaxation time of the body region, andwherein calculating the MPF comprises using the proton density and the longitudinal relaxation time to calculate the MPF. 11. A magnetic resonance apparatus comprising: a magnet;a radiofrequency pulse generator; anda processor and a non-transitory computer-readable medium storing program instructions thereon executable by the processor to cause the apparatus to perform functions comprising: applying, via the magnet, a ma... | University of Washington | A61B005/00 A61B005/055 A61N001/40
US09339206 | B2 | 2016-05-17 | Adaptor for endovascular electrocardiography | 2010-06-14 | An adaptor for endovascular electrocardiography includes a member having two ends and an inner lumen such that when one of the ends is connected to a catheter and the other end is connected to a syringe, fluid from the syringe is flushed into the catheter through the inner lumen. The adaptor also includes a metal insert, disposed along the inner lumen of the member, which connects through a sealed electrical connection to the outer side of the member. | 1. An adaptor for endovascular electrocardiography, comprising: a member having a first end, a second end, and an inner lumen, the first end designed to be connected to a luer of a catheter and the second end designed to be connected to a syringe, such that fluid from the syringe is flushed into the catheter through the inner lumen; anda metal ring, disposed along the inner lumen of the member and located fully inside the member, the metal ring connecting through a sealed electrical connection to an outer side of the member, the metal ring separated from the catheter by a section of the member when the first end of the member is connected to the luer of the catheter. 2. The adaptor according to claim 1, wherein the metal ring comprises a metal ring disposed along an inner surface of the inner lumen. 3. The adaptor according to claim 1, wherein the metal ring terminates prior to the two ends of the member. 4. The adaptor according to claim 1, further comprising a wire having a first end connected to the metal ring. 5. The adaptor according to claim 4, further comprising a connector coupled to an opposing end of the wire. 6. The adaptor according to claim 5, wherein the connector comprises external insulation. 7. The adaptor according to claim 5, wherein the connector is configured to plug into an electronic module. 8. The adaptor according to claim 5, wherein the connector comprises a nipple shape. 9. The adaptor according to claim 1, wherein an exterior most surface of the one of the ends of the adaptor is tapered. 10. The adaptor according to claim 9, wherein the other end comprises a luer connector. 11. The adaptor according to claim 1, wherein a body portion provides a connection between the metal ring along an inner surface of the inner lumen of the member and a wire external to the member through a perforation in a wall of the member. 12. The adaptor according to claim 11, wherein the connection through the wall is waterproof. 13. The adaptor according to claim 1, wherein the two ends of the member comprise different external and internal diameters. 14. An assembly for endovascular electrocardiography, comprising: the adaptor according to claim 1; anda connector separate from a body of the adaptor coupled to the metal ring. 15. The assembly according to claim 14, wherein the connector is coupled to the adaptor by a wire between the metal ring and the connector, the wire configured to freely orient the connector with respect to the adaptor. 16. The assembly according to claim 14, wherein the connector is configured to plug into an electronic module. 17. The assembly according to claim 14, wherein the connector comprises a nipple shape. | Bard Access Systems, Inc. | A61B005/042 A61B005/06 A61M025/00 A61B005/04 A61B005/0452 A61M039/10 A61B005/00 A61B019/00
US09326693 | B2 | 2016-05-03 | Placement of electrodes in proximity to the heart | 2014-04-15 | A method, consisting of receiving respective electrocardiograph (ECG) signals from body-surface electrodes at respective locations in proximity to a heart of a subject. The ECG signals may be processed to generate respective signal parameters characteristic of positions of the body-surface electrodes with respect to the heart, and the respective locations may be adjusted so as to achieve a specified geometrical relationship between the body-surface electrodes and the heart in response to the respective signal parameters. | 1. A method, comprising: receiving respective electrocardiograph (ECG) signals from body-surface electrodes at respective locations in proximity to a heart of a subject;processing the ECG signals to generate respective signal parameters thereof characteristic of positions of the body-surface electrodes with respect to the heart by finding baselines for the signals;generating the respective signal parameters by determining whether differences between maximum deviations from the baselines are positive or negative;enumerating as a first number the body-surface electrodes having a positive difference, and enumerating as a second number the body-surface electrodes having a negative difference, wherein achieving the specified geometrical relationship comprises the first and the second numbers differing by no more than a preset number; andadjusting the respective locations so as to achieve a specified geometrical relationship between the body-surface electrodes and the heart in response to the respective signal parameters. 2. The method according to claim 1, wherein the body-surface electrodes are configured to receive currents, from a catheter electrode within the subject, indicative of a location of the catheter electrode. 3. The method according to claim 2, wherein receiving the respective ECG signals comprises receiving the ECG signals while not receiving the currents. 4. The method according to claim 2, wherein receiving the respective ECG signals comprises receiving the ECG signals while receiving the currents. 5. The method according to claim 1, wherein the geometric relationship comprises the body-surface electrodes surrounding the heart. 6. The method according to claim 1 wherein the preset number comprises a positive whole number less than a total number of the body-surface electrodes. 7. The method according to claim 1, and comprising dividing the body-surface electrodes into a plurality of sub-groups, and comprising for a given sub-group enumerating as a first number the body-surface electrodes therein having a positive difference, and enumerating as a second number the body-surface electrodes therein having a negative difference, and wherein achieving the specified geometrical relationship comprises the first and the second numbers differing by no more than a preset number for the given sub-group. 8. The method according to claim 7, wherein the preset number comprises a positive whole number less than a total number of the body-surface electrodes in the given sub-group. 9. The method according to claim 1, wherein the signal parameters comprise binary parameters. 10. The method according to claim 1, wherein the signal parameters comprise non-binary parameters. 11. Apparatus, comprising: A plurality of body-surface electrodes, each body surface electrode positioned in a respective location in proximity to a heart of a subject, each body surface electrode configured to receive respective electrocardiograph (ECG) signals from the resp... | Biosense Webster (Israel) Ltd. | A61N001/00 A61B005/04 A61B005/0408 A61B005/06
US09901263 | B2 | 2018-02-27 | Radial artery blood pressure waveform measuring device | 2015-07-24 | A radial artery blood pressure waveform measuring device includes a stress sensor, a stress concentrating elastic structure, and a stress guiding elastic member. The stress sensor is in contact with a skin surface of a wrist corresponding to a position of a radial artery. The stress concentrating elastic structure is disposed on the stress sensor. The stress guiding elastic member has a top surface, a bottom surface, and at least one side surface. The bottom surface is disposed on the stress concentrating elastic structure, and an area of the bottom surface is greater than an area of a horizontal section of the stress concentrating elastic structure. | 1. A radial artery blood pressure waveform measuring device, comprising: a stress sensor configured for contact with a skin surface of a wrist corresponding to a position of a radial artery;a stress concentrating elastic structure disposed on the stress sensor;a stress guiding elastic member having a top surface, a bottom surface, and at least one side surface, wherein the bottom surface is disposed on the stress concentrating elastic structure, and an area of the bottom surface is greater than an area of a horizontal section of the stress concentrating elastic structure; andan elastic stress bugger disposed below the bottom surface and on one side of the stress concentrating elastic structure, for buffering a contractile force applied on the side surface by the elastic band, wherein the elastic stress buffer is formed of a sponge material. 2. The radial artery blood pressure waveform measuring device of claim 1, further comprising an elastic band surrounding the stress guiding elastic member and being configured to surround the wrist, and capable of applying the contractile force on the wrist, the top surface, and the side surface. 3. The radial artery blood pressure waveform measuring device of claim 1, wherein a Modulus of Elasticity of the elastic stress buffer is less than a Modulus of Elasticity of the stress guiding elastic member. 4. The radial artery blood pressure waveform measuring device of claim 1, wherein a side profile of the stress guiding elastic member is trapezoidal in shape. 5. The radial artery blood pressure waveform measuring device of claim 1, wherein the stress guiding elastic member is formed of a foam material. 6. The radial artery blood pressure waveform measuring device of claim 1, wherein the stress concentrating elastic structure comprises: a stress applying layer disposed on the stress sensor;a buffer layer disposed on the stress applying layer; anda stress receiving layer disposed on the buffer layer. 7. The radial artery blood pressure waveform measuring device of claim 6, wherein each of the stress applying layer and the stress receiving layer is formed of a foam material. 8. The radial artery blood pressure waveform measuring device of claim 6, wherein the buffer layer is formed of a sponge material or rubber. | NATIONAL TAIWAN UNIVERSITY OF SCIENCE AND TECHNOLOGY | A61B005/02 A61B005/021 A61B005/00
US09597035 | B2 | 2017-03-21 | Skin-mountable medical device | 2015-11-11 | The invention concerns a skin-mountable medical device comprising a patch portion with a contact surface facing the skin of a patient during use, a process unit attached to the patch portion, and a plaster kit for securing the patch portion to the skin, wherein the plaster kit has a primary pad adapted to releasably adhere to at least a section of a peripheral skin surface surrounding the contact surface. According to the invention, the plaster kit further comprises at least one auxiliary pad which is connectable or connected to the patch portion and adapted to functionally replace the primary pad after a wearing time. | 1. Skin-mountable medical device, comprising: a process unit;a patch portion housing the process unit;a primary pad extending from the patch portion, the primary pad including a primary carrier layer having a tear-off section configured to detach the primary pad from the patch portion, anda primary adhesive on the primary carrier layer to releasably adhere to skin; andan auxiliary pad layered at least partially over the primary pad, the auxiliary pad being bendably connected to the patch portion for pivoting to uncover the primary pad during removal, the auxiliary pad including an auxiliary carrier layer, andan auxiliary adhesive on the auxiliary carrier layer to releasably adhere to the skin after removal of the primary pad. 2. Medical device according to claim 1, further comprising: a base pad joined to the patch portion and the primary pad, the base pad including a base carrier layer, anda base adhesive on the base carrier layer to releasably adhere to the skin; andwherein the tear-off section of the primary carrier layer includes a perforation line between the base carrier layer and the primary carrier layer. 3. Skin-mountable medical device, comprising: a patch portion with a contact surface facing the skin of a patient during use;a process unit particularly formed by a sensor or medication delivery set attached to the patch portion;a adhesive kit for securing the patch portion to the skin, wherein the adhesive kit has a primary pad adapted to releasably adhere to at least a section of a peripheral skin surface surrounding the contact surface, characterized in that the adhesive kit further comprises at least one auxiliary pad which is connectable or connected to the patch portion and adapted to functionally replace the primary pad after a wearing time;wherein the primary pad and the at least one auxiliary pad at least partially overlap one another in a layered configuration;wherein the primary pad has a tear-off perforation for allowing removal of the primary pad; andwherein the auxiliary pad is bendably connected to the patch portion and forms a flap to uncover the primary pad for removal. 4. Medical device according to claim 3, wherein an adhesive coating of the auxiliary pad is separated by a liner from the primary pad. 5. Medical device according to claim 3, wherein the adhesive kit further comprises a base pad fixedly connected to the contact surface of the patch portion. 6. Medical device according to claim 3, wherein the auxiliary pad is formed as a separate ring which can be applied over the periphery of the patch portion. 7. Medical device according to claim 3, further comprising a ring-shaped frame which is affixed to the inner edge of an annular auxiliary pad, wherein the frame can be engaged to the patch portion. 8. Medical device according to claim 3, further comprising a stack of at least two auxiliary pads which are joined together in a parallel folding pattern particularly in a Leporello fashion for sequential unfoldin... | Roche Diabetes Care, Inc. | A61B005/00 A61M005/142 A61B005/145 A61N001/04
US10080624 | B2 | 2018-09-25 | Powering and reading implanted devices | 2014-05-12 | A method of providing RF energy to a wireless transponder device includes outputting primary RF energy in a form wherein the primary RF energy comprises a plurality of RF bursts, with each consecutive pair of the RF bursts being separated by an associated time interval, and outputting secondary continuous wave RF energy in between each of the RF bursts of the primary RF energy during at least a portion of each of the associated time intervals. Also, a reader device implementing this method. | 1. A radio frequency identification reader device, comprising: a primary radio frequency source structured to output primary radio frequency energy in a form wherein the primary radio frequency energy comprises a plurality of radio frequency bursts, each consecutive pair of the radio frequency bursts being separated by an associated time interval;a supplemental radio frequency source structured to output secondary continuous wave radio frequency energy in between each of the radio frequency bursts of the primary radio frequency energy during at least a portion of each of the associated time intervals; andcontrol circuitry structured to detect an end of each of the radio frequency bursts and responsive to each detection cause the supplemental radio frequency source to output the secondary continuous wave radio frequency energy during at least a part of the time interval associated with the radio frequency burst. 2. The radio frequency identification reader device according to claim 1, further comprising a touch probe coupled to the primary radio frequency source and the supplemental radio frequency source, the touch probe being structured to receive the primary radio frequency energy and the secondary continuous wave radio frequency energy and to provide the primary radio frequency energy and the secondary continuous wave radio frequency energy to a wireless transponder device through a direct electrical connection not over an air interface between the touch probe and the wireless transponder device. 3. The radio frequency identification reader device according to claim 2, wherein the direct electrical connection is through transcutaneous contact. 4. The radio frequency identification reader device according to claim 1, wherein the supplemental radio frequency source comprises an radio frequency oscillator. 5. The radio frequency identification reader device according to claim 4, wherein the primary radio frequency source is part of a commercial off the shelf radio frequency identification reader device and the supplemental radio frequency source and the control circuitry are separate from the commercial off the shelf radio frequency identification reader device. 6. A method of providing radio frequency energy to a wireless transponder device, comprising: outputting primary radio frequency energy in a form wherein the primary radio frequency energy comprises a plurality of radio frequency bursts, each consecutive pair of the radio frequency bursts being separated by an associated time interval;outputting secondary continuous wave radio frequency energy in between each of the radio frequency bursts of the primary radio frequency energy during at least a portion of each of the associated time intervals; anddetecting an end of each of the radio frequency bursts and responsive to each detection causing the secondary continuous wave radio frequency energy to be output during at least a part of the time interval associated with the radio frequency ... | University of Pittsburgh—Of the Commonwealth System of Higher Education | A61B090/98 A61B005/00 G01S003/02 H04B007/06 G01F023/296 A61F002/38 A61B005/07 A61F002/30
US10548500 | B2 | 2020-02-04 | Apparatus for measuring bioelectrical signals | 2015-11-20 | An apparatus for measuring bioelectrical signals is provided. The apparatus includes a sensor electrode, a sensor support, and a main body. The sensor electrode has a tapering portion that narrows toward one end and a protruding portion that extends from the one end of the tapering portion, contacts a body part, and senses bioelectrical signals. The sensor support maintains the contact between the sensor electrode and the body part. The main body is connected to the sensor support and is wearable on a living body. | 1. An apparatus comprising: a sensor electrode comprising a tapering portion that narrows toward one end and a flexible protruding portion that extends from the one end of the tapering portion and is configured to contact a body part, and senses bioelectrical signals;an electrode supporter to which an other end opposite from the one end of the tapering portion is attached and which is configured to support the sensor electrode;a sensor support configured to support the sensor electrode to maintain the contact between the sensor electrode and the body part; anda main body connected to the sensor support and being wearable on a living body,wherein, a gradient of a longitudinal section of an outer surface of the flexible protruding portion is greater than a gradient of a longitudinal section of an outer surface of the tapering portion with respect to a bottom surface of the other end of the tapering portion, andwherein the protruding portion of the sensor electrode is configured to bend and the tapering portion remains stationary with respect to the electrode supporter when a force is applied to the sensor electrode in the direction of the body part, wherein an outer circumferential surface of the protruding portion contacts the body part. 2. The apparatus of claim 1, wherein the flexible protruding portion comprises a flexible material, has a flexibility greater than a flexibility of the tapering portion, and is bent when contacting the body part so that a side portion of the flexible protruding portion contacts the body part, and the tapering portion supports the flexible protruding portion. 3. The apparatus of claim 2, wherein the flexible material is conductive silicone or conductive rubber. 4. The apparatus of claim 1, wherein an electrode pattern for electrically connecting to the sensor electrode is provided on one of a mount surface of the electrode supporter on which the sensor electrode is mounted or a rear surface of the electrode supporter. 5. The apparatus of claim 4, further comprising a sensor circuit disposed on the rear surface of the electrode supporter or disposed inside the electrode supporter and configured to process the bioelectrical signals detected by the sensor electrode into analog signals or digital signals. 6. The apparatus of claim 1, wherein the sensor support elastically supports a sensor module that includes the sensor electrode and the electrode supporter so that the sensor module moves in 3-axis directions. 7. The apparatus of claim 6, wherein the sensor support comprises a spiral spring, and the sensor module is coupled to one of two ends of the spiral spring, wherein a surface on which one end of the spiral spring is placed protrudes beyond a surface on which the other end of the spiral spring is placed. 8. The apparatus of claim 7, wherein an elasticity of the spiral spring along a central axis of the spiral spring is greater than an elasticity of the spiral spring in a direction perpendicular to the cent... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/0478 A61B005/00 A61B005/04
US10813388 | B2 | 2020-10-27 | Athletic band with removable module | 2015-11-19 | Methods of detecting and reducing error of athletic sensor readouts are disclosed. Methods relate to receiving properties, such as an athlete's body mass index and a physiological property, such as heart rate. Based on a body mass index or other parameter being outside a range, selecting a protocol, such as a heart rate measurement protocol. In further embodiments, a heart rate measurement protocol may be augmented. | 1. A computer-implemented method of selecting and modifying a heart rate measurement protocol: receiving at a processor a body mass index of an athlete and demographic information of the athlete;determining by the processor that the athlete's body mass index is less than a lower threshold of a predetermined index range having at least the lower threshold;receiving a first heart rate value from an optical heart rate sensor retained in a housing configured to be in contact with a body of the athlete, where the optical heart rate sensor is at a first orientation and at a first position with respect to the athlete's body;based upon the athlete's demographic information and the determination that the athlete's body mass index is less than the lower threshold of the predetermined index range, selecting a first heart rate measurement protocol instead of at least a second heart rate measurement protocol, wherein the first heart rate measurement protocol is different than the at least a second heart rate measurement protocol with respect to at least one parameter selected from the group consisting of: wavelength of light transmitted by the optical heart rate sensor, intensity of light transmitted from the optical heart rate sensor, algorithm utilized, weights, confidences, and combinations thereof;calculating, using the first heart rate measurement protocol, a first heart rate measurement from the first heart rate value;based upon both the athlete's body mass index being less than the lower threshold of the predetermined index range, and the first heart rate measurement being outside of a first heart rate range, determining with the processor that at least one of the first orientation and the first position is not optimal for obtaining athletic data from the athlete;based upon determining that the at least one of the first orientation and the first position is not optimal for obtaining athletic data from the athlete, outputting a human-detectable output to the athlete configured prompt the athlete to change the at least one of the first orientation and the first position of the optical heart rate sensor;storing the first heart rate measurement on a non-transitory computer-readable medium;receiving a second heart rate value from the optical heart rate sensor; andcalculating a second heart rate measurement calculated using the first heart rate measurement protocol from the second heart rate value. 2. The method of claim 1, wherein: the human-detectable output to the athlete is further configured to guide the athlete in changing the at least one of the first orientation and the first position of the optical heart rate sensor. 3. The method of claim 1, comprising: receiving at the processor athletic data from an external database from a plurality of athletes obtained while each of the plurality of athletes performed athletic activities known to cause muscle flexure in a first muscle group at the first position;wherein the human-detectable output is based up... | NIKE, Inc. | A61B005/00 A41D001/00 A63B024/00 A41D020/00 G08B006/00 G06F003/02 G06F003/01 G06F001/16 A41H043/04 G06F019/00 A61B005/0205 A61B005/024 A61B005/11 A41D027/20 B29C065/02 B29C065/52 B29C065/78 B29C065/00 A45F005/00 G01C021/36 G16H020/40 B29K623/00 B29L029/00 B29L031/48
US10470703 | B2 | 2019-11-12 | System and method for functional state and/or performance assessment and training program adjustment | 2015-10-21 | A system and method for training program generation and modification based on user input goal data and assessed functional state and/or workload performance. A user-interface preferably operating on a mobile device permits a user to record bio-signals indicative of functional state. Assessments are performed on the received bio-signal data to produce training session targets based on the current functional state of the user. Workload performance may be monitored and the training session targets modified based on measured past performance to improve future performance. Various embodiments are disclosed. | 1. A method for a portable electronic device to interact with at least one user borne sensor electrode to combine sensed data with user input goal data to generate a training session target, comprising the steps of: detecting, by the portable electronic device, input specifying user goal data, the goal data including a type of activity for the user, and storing the goal data in memory;detecting, by the portable electronic device, bio-signal data sensed from at least one sensor electrode that is positioned at an appropriate location on the user, for a heart rate variability (HRV) assessment of a current functional state of the user, and storing the detected HRV bio-signal data in memory;processing, by the portable electronic device or another device in communication with the portable electronic device, an HRV assessment of the detected HRV bio-signal data to generate a first index parameter indicative of the state of a first body system of the user that corresponds to the HRV assessment, and storing the first index parameter in memory;generating, by the portable electronic device or another device in communication with the portable electronic device, a session target for a given user training session that is based on retrieved user goal data and the first index parameter, the session target including a first session target value and a second session target value; andgenerating, by the portable electronic device, and displaying on a screen thereof, a visual indication of the first and the second session target values;wherein the first body system is the autonomic nervous system and the first and second session target values are based, at least in part, on the relative levels of the sympathetic nervous system and the parasympathetic nervous system of that user;wherein the first session target value is an intensity target and the second session target value is a volume target; andwherein the session target generating step further includes decreasing the volume target more than the intensity target when the parasympathetic level is high relative to the sympathetic level and decreasing the intensity target more than the volume target when the sympathetic level is high relative to the parasympathetic level. 2. The method of claim 1, further comprising the steps of: detecting, by the portable electronic device, sensed bio-signal data for an electrocardiogram (ECG) assessment of the current functional state of that user, and storing the detected ECG bio-signal data in memory;processing, by the portable electronic device or another device in communication with the portable electronic device, an ECG assessment of the ECG bio-signal data to generate a second index parameter indicative of the state of a second body system that corresponds to the ECG assessment, and storing the second index parameter in memory;adjusting, by the portable electronic device or another device in communication with the portable electronic device, the session target based on the se... | - | A61B005/00 A61B005/0205 A61B005/0402 A61B005/0476 A61B005/0488 A61B005/11 G06F019/00 A61B005/04 A61B005/024
US09351712 | B1 | 2016-05-31 | Simultaneous multiple method out-patient uterus biopsy device and method | 2014-12-24 | A biopsy device, comprising a flexible coaxial structure, comprising an obturator within a sheath, the obturator being adapted to be displaced with respect to the sheath along the coaxial axis by a force applied at a proximal end; a disruptor, at a distal end of the obturator, adapted to disrupt a tissue surface to free cells therefrom, having a first position covered within the sheath and a second position freely extending beyond the sheath; an element, having a fixed position on the obturator, and a close clearance with respect to an inner surface of the sheath, such that a retraction of the obturator within the sheath causes an associated change in pressure in a space between the obturator distal to the element and the sheath, to thereby draw fluid proximate to an end of the sheath within the sheath. | 1. A biopsy device, comprising: a flexible coaxial structure, comprising a flexible impermeable sheath and a displaceable flexible tensile and compressive force transmitting structure having a distal end within the sheath, the sheath being configured for advancement in direct external contact with tissues into an organ due to a pressure exerted at a proximal end;a tissue disruptive sampling structure attached to the displaceable flexible tensile and compressive force transmitting structure at the distal end, configured to preserve three-dimensional structures among cell aggregates of a disrupted tissue; anda piston-like structure, disposed around the displaceable flexible tensile and compressive force transmitting structure, disposed proximally with respect to the tissue disruptive sampling structure, configured to slide within the sheath and to draw a vacuum within the sheath and thereby to draw a liquid sample containing the three-dimensional structures among cell aggregates of a disrupted tissue into the sheath from a space surrounding the tissue disruptive sampling structure when the displaceable flexible tensile and compressive force transmitting structure is supplied with a tensile force and the piston-like structure withdrawn into the sheath. 2. The biopsy device according to claim 1, wherein the tissue disruptive sampling structure comprises a plurality of bristles. 3. The biopsy device according to claim 2, wherein the plurality of bristles extend from the displaceable flexible tensile and compressive force transmitting structure, having an unconstrained diameter greater than a diameter of the sheath, and adapted to be withdrawn into the sheath. 4. The biopsy device according to claim 1, further comprising a blunt cap at a distal tip of the tissue disruptive sampling structure. 5. The biopsy device according to claim 1, further comprising an atraumatic bulb at a distal tip of the tissue disruptive sampling structure. 6. The biopsy device according to claim 1, wherein the displaceable flexible tensile and compressive force transmitting structure comprises a guidewire. 7. The biopsy device according to claim 6, wherein the guidewire comprises a spirally twisted flexible wire. 8. The biopsy device according to claim 7, wherein the sampling end comprises a plurality of bristles fixed within the spirally twisted wire at a distal end, which form a tapered brush. 9. The biopsy device according to claim 1, wherein the sheath has an outer diameter of between 1 and 3 mm. 10. The biopsy device according to claim 1, wherein the sheath comprises at least two lumens, a first lumen containing the displaceable flexible tensile and compressive force transmitting structure, and the second lumen being configured to inject a liquid at a distal end of the sheath. 11. A biopsy method, comprising: providing:a flexible coaxial structure, comprising a flexible impermeable sheath and a displaceable flexible tensile and compressive force transmitting str... |  | A61B005/00 A61B010/02 A61M025/09
US09341632 | B1 | 2016-05-17 | Ketone measurement system capable of detecting correlations between measurements and user behaviors | 2015-07-23 | Systems and related methods are provided for measuring a ketone in the fluid of a patient or other user. The systems include a portable measurement device for analyzing a plurality of fluid samples and for generating a corresponding plurality of ketone measurement results. The systems also include an electronic or communications device, for example, such as a cell phone or smartphone, configured to communicate with the measurement device. The system further includes a software application. The systems and related methods according to one aspect provide reminders, preferably interactive reminders that facilitate user compliance with a weight management or diabetic ketoacidosis (DKA) monitoring or prevention program. The systems and related methods according to another aspect incorporate ketone tags and/or trigger points to improve accuracy of results and to advance user compliance. | 1. A system for measuring and monitoring ketone levels in a human, comprising: a portable ketone measurement device configured to measure ketone levels in breath samples, the portable ketone measurement device comprising a breath input port that fluidly couples a breath input device to a ketone sensor, and a wireless transceiver that enables the portable ketone measurement device to communicate with a mobile communications device of a user of the portable ketone measurement device; andnon-transitory computer storage having stored thereon program instructions that direct a computing system comprising one or more computing devices to perform a process that comprises: storing data representing a trigger point associated with the user, the trigger point representing a user behavior associated with non-compliance with a behavioral program;receiving and storing a plurality of ketone measurements generated from respective breath samples, each ketone measurement generated at a different point in time over a time period by the portable ketone measurement device;receiving, via a user interface, indications from the user of a plurality of trigger point occurrences in which the user engaged in the behavior over the time period;analyzing the plurality of ketone measurements in conjunction with the indications of the trigger point occurrences to determine whether the trigger point is associated with a ketone level pattern;recording a correlation detected based on the analysis, the correlation identifying a particular ketone level pattern associated with the trigger point;subsequent to the time period, receiving and storing a sequence of two or more ketone measurements generated by the portable ketone measurement device based on additional respective breath samples;determining whether the sequence corresponds to the trigger point at least partly by comparing the sequence to the particular ketone level pattern; andin response to determining that the sequence corresponds to the trigger point, causing the mobile communications device to display a user interface that identifies a possible occurrence of the trigger point, the user interface including a user option to indicate whether the trigger point has occurred. 2. The system of claim 1, wherein the behavioral program is a weight loss program. 3. The system of claim 1, wherein the program instructions further direct the computing system to: prompt the user to select, from a master set of trigger points, a subset of trigger points that are applicable to the user, the subset having less than all of the trigger points in the master set; andgenerate, on the mobile communications device, a user interface that displays the subset of trigger points and enables the user to use the subset to record occurrences of trigger points. 4. The system of claim 1, wherein the process further comprises: recording, in association with at least some of the trigger point occurrences, trigger point state information specifying at lea... | Invoy Technologies, Inc. | G01N033/48 G01N033/64 A61B005/00
US09339201 | B2 | 2016-05-17 | Database and algorithm for evaluating efficacy of an electrophysiology procedure | 2013-10-08 | The present invention provides an improved, Internet-based system that seamlessly collects cardiovascular data from a patient before, during, and after a procedure for EP or an ID. During an EP procedure, the system collects information describing the patient's response to PES and the ablation process, ECG waveforms and their various features, HR and other vital signs, HR variability, cardiac arrhythmias, patient demographics, and patient outcomes. Once these data are collected, the system stores them on an Internet-accessible computer system that can deploy a collection of user-selected and custom-developed algorithms. Before and after the procedure, the system also integrates with body-worn and/or programmers that interrogate implanted devices to collect similar data while the patient is either ambulatory, or in a clinic associated with the hospital. A data-collection/storage module, featuring database interface, stores physiological and procedural information measured from the patient. | 1. A system for evaluating an efficacy of an electrophysiology procedure, comprising: a software interface configured to receive ablation locations indicating where a patient's heart has been ablated during a cardiac electrophysiology (EP) procedure, and store numerical representations of the ablation locations in a database;an ECG-measuring system configured to measure a time-dependent ECG waveform from the patient after the EP procedure when the patient is ambulatory, and then transmit numerical representation of the ECG waveform to the database;a first processing system configured to processes the numerical representations of the time-dependent ECG waveform to determine a heart rate and heart rate variability; anda second processing system configured to compare a first set of numerical values indicating the ablation locations to the heart rate and heart rate variability to determine a parameter estimating the efficacy of the EP procedure. | TOSENSE, INC. | A61B005/04 A61B005/00 A61B005/024 A61B005/044
US09345436 | B2 | 2016-05-24 | Apparatus and methods for computing cardiac output of a living subject | 2013-03-14 | Apparatus and methods for calculating cardiac output (CO) of a living subject. In one embodiment, the apparatus and methods build a nonlinear mathematical model to correlate physiologic source data vectors to target CO values. The source data vectors include one or more measurable or derivable parameters such as: systolic and diastolic pressure, pulse pressure, beat-to-beat interval, mean arterial pressure, maximal slope of the pressure rise during systole, the area under systolic part of the pulse pressure wave, gender (male or female), age, height and weight. The target CO values are acquired using various methods, across a plurality of individuals. Multidimensional nonlinear optimization is then used to find a mathematical model which transforms the source data to the target CO data. The model is then applied to an individual by acquiring physiologic data for the individual and applying the model to the collected data. | 1. A method of computing cardiac output (CO) of a living subject, comprising: obtaining a plurality of target CO values;generating an mathematical model based at least in part on said target CO values;obtaining a plurality of physiologic input source data from said living subject;transforming said plurality of physiologic input source data using at least said mathematical model; andexpressing a result of said transformation as a CO value in physiologic units. 2. The method of claim 1, wherein said target CO values are obtained by performing at least one of: waveform-based methods;Doppler echocardiography-based methods; orthermodilution-based methods. 3. The method of claim 1, wherein said optimization module comprises a multi-dimensional non-linear optimization algorithm. 4. The method of claim 1, wherein said plurality of physiologic input source data comprises at least one of: systolic pressure of said living subject;diastolic pressure of said living subject; andpulse pressure of said living subject. 5. The method of claim 1, wherein said plurality of physiologic input source data comprises at least one of: beat-to-beat interval;mean arterial pressure;maximal slope of pressure rise during systole; andarea under a systolic portion of a pulse pressure wave. 6. The method of claim 1, wherein said plurality of physiologic input source data comprises at least one of: age of said living subject;gender of said living subject;height of said living subject; andweight of said living subject. 7. The method of claim 1, wherein said act of transforming comprises fitting said physiologic input data into said target CO data in a least-square optimal fashion. 8. A non-transitory computer readable apparatus configured to store at least one computer program thereon, said at least one computer program comprising a plurality of instructions which are configured to, when executed by a processing apparatus: receive a plurality of first data, individual ones of said plurality of first data each comprising cardiac output (CO) values representative of living subjects having a respective plurality of physiologic parameters;generate a mathematical model to correlate said plurality of CO values to said respective plurality of physiologic parameters;receive second data, said second data comprising one or more measured physiologic parameters of a particular living subject;link, in a mathematically optimal way, said second data to one of said plurality of CO values via at least:detection of one or more specific data points within continuous biosignals received from said particular living subject;selection of a data vector representative of said continuous bio-signals based at least in part on said detected one or more specific data points; andapplication of a mathematical transformation to said selected data vector to arrive at said one of said plurality of CO values; and cause a display of said one of said plurality of CO values. 9. The apparatus of claim 8, wherein... | HighDim GmbH | A61B005/02 A61B005/00 A61B005/029 A61B005/021 A61B005/0452 A61B008/02 A61B008/06 A61B005/0205 A61B003/16 A61B005/024 A61B008/08
US09333312 | B2 | 2016-05-10 | Peep regulation for a breathing apparatus | 2009-03-27 | A breathing apparatus has an expiratory flow sensor unit, an expiratory pressure sensor unit, an expiratory pressure regulator, and an expiratory control unit operatively connected to the expiratory pressure sensor unit for a feedback control of an expiratory pressure during an expiratory phase of a breathing cycle. The expiratory control unit is adapted to provide a real time target pressure value during the expiratory phase to the expiratory pressure regulator for adjusting a desired positive end expiratory pressure (PEEP) level (PEEPset) based on an inspiratory tidal volume of an inspiratory phase of the same breathing cycle and a currently accumulated expiratory tidal volume measured by the expiratory flow sensor unit. In this manner advantageous regulation of PEEP with minimized work of breathing is obtained. | 1. A breathing apparatus comprising: a breathing circuit adapted to interact with a patient respiratory in successive breathing cycles each comprising an inspiratory phase and expiratory phase, to provide respiratory assistance to the patient, wherein the breathing circuit comprises an inspiratory gas flow section and an expiratory gas flow section, wherein the expiratory gas flow section includes an expiratory pressure regulator;an expiratory control unit configured to provide a real time variable target pressure (Ptarget) during an expiratory phase of a breathing cycle to said expiratory pressure regulator for regulating a desired positive end expiratory pressure (PEEP) level (PEEPset) of said expiratory pressure regulator, said expiratory control unit being configured to provide said target pressure (Ptarget) based on an inspiratory tidal volume (VTi) provided by said breathing circuit of an inspiratory phase in the same breathing cycle and a currently accumulated expiratory tidal volume; andsaid expiratory control unit being configured to provide said target pressure (Ptarget) below said desired PEEP level (PEEPset) until an expiratory flow is zero or equal to a bias flow in said expiratory gas flow section of said breathing circuit. 2. The breathing apparatus of claim 1, wherein said expiratory control unit is configured to adapt to said patient and to changes in expiratory control conditions during or within said same breathing cycle in said respiratory assistance by said breathing circuit. 3. The breathing apparatus of claim 1, comprising an inspiratory flow sensor unit arranged to measure an inspiratory gas flow during an inspiratory phase of said same breathing cycle, and configured to provide said inspiratory tidal volume. 4. The breathing apparatus of claim 1, wherein said expiratory control unit is configured to control said expiratory pressure regulator so as to remain partly closed at a beginning of said expiratory phase in accordance with a condition that said expiratory pressure regulator is controlled to open not more than to provide said target pressure (Ptarget) with a maximum, pre-determined gradient. 5. The breathing apparatus of claim 1, wherein said expiratory control unit is configured to provide said target pressure (Ptarget) as Ptarget(t)=f(Vexp(t),Vinsp,Vleakage)·PEEPset whereinf(Vexp(t),Vinsp,Vleakage)≦1;Vinsp=∫0Tinsp⁢V.⁢⁢ⅆtis the inspiratory tidal volume;Vexp⁡(t)=-∫Tinspt⁢V.⁢⁢ⅆtis the accumulated, currently expired expiratory volume; and Vleakage is a leakage volume. 6. The breathing apparatus of claim 5, whereinf⁡(Vexp⁡(t),Vinsp,Vleakage)=(Vexp⁡(t)(Vinsp-Vleakage)·k)n,whereink is a positive numeral, and k≦1; andn is a positive numeral, and n>1. 7. The breathing apparatus of claim 1, wherein said expiratory control unit is configured to provide said target pressure (Ptarget) asPtarget⁡(t)=Vexp⁡(t)×PEEPset×marginVinsp+Paccumulated⁡(t)whereinVexp(t) is the accumulated, currently expired expiratory volume;Vinsp is ... | Maquet Critical Care AB | A61M016/00 A61M016/10 A61B005/091 A61M016/20
US09972081 | B2 | 2018-05-15 | Method and system for estimating a deflated lung shape for video assisted thoracic surgery | 2015-10-29 | The present teaching relates to surgical procedure assistance. In one example, a first volume of air inside a lung is obtained based on a first image of the lung captured prior to a surgical procedure. The lung has a first shape on the first image. A second volume of air deflated from the lung is determined based on a second image of the lung captured during the surgical procedure. A second shape of the lung is estimated based on the first shape of the lung and the first air volume inside the lung and second volume of air deflated from the lung. A surgical plan is updated based on the estimated second shape of the lung. | 1. A method, implemented on a computing device having at least one processor, storage, and a communication platform capable of connecting to a network for surgical procedure assistance, comprising: obtaining a first volume of air inside a lung based on a first image of the lung captured prior to a surgical procedure, wherein the lung has a first shape on the first image;capturing a video image of the lung during the surgical procedure;computing, via registering the captured video image with the first image, a second volume of air deflated from the lung;estimating a second shape of the lung based on the first shape of the lung, the first air volume inside the lung, and the second volume of air deflated from the lung, the second shape being estimated by condensing a plurality of voxels in the first image along a gravity direction; andupdating a surgical plan based on the estimated second shape of the lung. 2. The method of claim 1, wherein the surgical procedure includes a video assisted thoracic surgery (VATS), and the first image includes a computed tomography (CT) image. 3. The method of claim 1, further comprising: obtaining an angle of rotation of the lung from a first position at which the first image was captured to a second position at which the video image was captured; andadjusting the first shape of the lung on the first image based on the angle of rotation, wherein the second shape of the lung is estimated based on the adjusted first shape of the lung, the first volume, and the second volume. 4. The method of claim 1, wherein computing the second volume of the air comprises: obtaining at least the video image having the lung and surrounding structures;extracting features from at least the video image;reconstructing three dimensional (3D) coordinates of the lung and the surrounding structures based on the features; andestimating a third volume of space between the lung and the surrounding structures based on the 3D coordinates. 5. The method of claim 1, wherein computing the second volume of the air comprises: estimating a plurality of simulated second shapes of the lung as a function of a deflation control parameter;determining one of the plurality of simulated second shapes of the lung based on the video image of the lung; anddetermining the deflation control parameter associated with the determined simulated second shape of the lung. 6. The method of claim 1, wherein estimating the second shape of the lung comprises: calculating a third volume of air inside the lung during the surgical procedure based on the second volume; anddetermining a ratio between the first volume and the third volume. 7. The method of claim 6, wherein estimating the second shape of the lung further comprises: estimating, for each voxel of the plurality of voxels included in the first shape of the lung on the first image, the first volume of air inside each voxel based on an intensity of each voxel;calculating the second volume of air inside each voxel bas... | EDDA TECHNOLOGY, INC. | G06T007/00 A61B019/00 A61B005/091 G06T017/20 G06T019/00 A61B006/03
US09326688 | B2 | 2016-05-03 | Object information acquiring apparatus | 2013-02-26 | The present invention employs an object information acquiring apparatus that includes a light source, a holding plate configured to hold an object, a supporting unit configured to support the holding plate, and a probe configured to receive, via the holding plate, an acoustic wave generated from the object on which light is irradiated from the light source. A light reflecting member is provided between the supporting unit and the holding plate. | 1. An object information acquiring apparatus comprising: a light source;a holding plate configured to hold an object, said holding plate being transparent to light from said light source;a supporting unit configured to support said holding plate;a probe configured to receive, via said holding plate, an acoustic wave generated from the object on which light is irradiated from said light source; anda light reflecting member provided between said supporting unit and said holding plate, wherein said light reflecting member is in contact with both said supporting unit and said holding plate. 2. The object information acquiring apparatus according to claim 1, wherein said light reflecting member is provided on said holding plate. 3. The object information acquiring apparatus according to claim 1, wherein said light reflecting member is provided on said supporting unit. 4. The object information acquiring apparatus according to claim 1, wherein said light reflecting member is a member having reflectance higher than that of said supporting unit. 5. The object information acquiring apparatus according to claim 1, further comprising a processing unit configured to generate property information of an inside of the object on the basis of the acoustic wave. 6. An object information acquiring apparatus according to claim 1, wherein said supporting unit is further configured to support said holding plate via a joining region. 7. An object information acquiring apparatus according to claim 1, further comprising an optical member disposed between said light source and said holding plate, said optical member being at least one of a lens, a mirror, a diffuser, and an optical fiber. 8. An object information acquiring apparatus according to claim 1, wherein said holding plate includes polymethylpentene. 9. An object information acquiring apparatus according to claim 1, wherein said holding plate includes at least one of plastic and glass. 10. An object information acquiring apparatus according to claim 1, wherein said supporting unit is made from tungsten carbide. 11. An object information acquiring apparatus according to claim 1, wherein said light reflecting member includes at least one of a dielectric multilayer film and a metal film. 12. An object information acquiring apparatus according to claim 1, wherein said light reflecting member has a reflectivity of 80% or more at the wavelength of the light emitted from said light source. 13. An object information acquiring apparatus comprising: a light source;a holding plate configured to hold an object;a supporting unit configured to support said holding plate;a probe configured to receive, via said holding plate, an acoustic wave generated from the object; anda light reflecting member provided between said supporting unit and said holding plate,wherein said light reflecting member is in contact with both said supporting unit and said holding plate, andwherein said probe is configured to be moved to receiv... | CANON KABUSHIKI KAISHA | A61B008/00 A61B006/00 A61B005/00
US09345881 | B2 | 2016-05-24 | Genetic algorithm based auditory training | 2015-04-13 | Embodiments of the present invention are generally directed to the use of a genetic algorithm for the purpose of providing progressive and adaptive auditory training (rehabilitation) to a recipient of a hearing prosthesis. In general, the genetic algorithm is used to adapt the training process to automatically increase the difficulty of the training based on recipient feedback and performance. That is, the genetic algorithm progressively removes perceivable sounds from the training process so as to generate groups of sounds that are difficult for a recipient to perceive. | 1. An auditory training system, comprising: a user interface; anda processor configured to: present an initial set of auditory training sound tokens to a recipient of a hearing prosthesis;receive recipient feedback in response to presented sound tokens from the initial set of auditory training sound tokens, wherein the feedback indicates the recipient's perception of the presented sound tokens; andexecute an interactive augmented genetic algorithm to automatically increase the difficulty of the auditory training based on the recipient's perception of the presented sound tokens to create a second set of auditory sound tokens for presentation to the recipient. 2. The system of claim 1, further comprising: a memory comprising an auditory training library,wherein the processor is configured to select the initial set of auditory training sound tokens from the auditory training library. 3. The method of claim 1, wherein the auditory training library comprises recipient-specific sound tokens added to the auditory training library by a user. 4. The system of claim 1, wherein the processor is configured to: sequentially present the initial set of auditory training sound tokens to the recipient via the hearing prosthesis; andreceive recipient feedback after each sequentially presented sound token. 5. The system of claim 4, wherein each sound token in the initial set of auditory training sound tokens corresponds to a known sound, and wherein to receive recipient feedback in response to a first presented sound token the processor is configured to: receive an indication of whether the recipient correctly perceived the known sound corresponding to the first presented sound token. 6. The system of claim 5, wherein if the recipient correctly perceived the known sound corresponding to the first presented sound token, the processor is configured to add the first presented sound token to a tabu list. 7. The system of claim 6, wherein the processor is configured to: present a visual representation of a plurality of sounds to the recipient in relation to presentation of a selected sound token;present a request for the recipient to select one of the visual representations that corresponds to the selected sound token; andreceive a selection from the recipient of one of the visual representations. 8. One or more non-transitory computer readable storage media encoded with software comprising computer executable instructions and when the software is executed operable to: control a hearing prosthesis to sequentially present, during auditory training, an initial set of auditory training sound tokens to a recipient of the hearing prosthesis;receive, from a user interface, recipient feedback in response to presented sound tokens from the initiate set of auditory training sound tokens, wherein the feedback indicates whether the recipient correctly perceived known sounds corresponding to each the presented sound tokens; andexecute a genetic algorithm to automatically inc... | Cochlear Limited | A61N001/36 G06F019/00 G06N003/12 H04R025/00 A61B005/12 A61N001/05 A61N001/372
US09578903 | B2 | 2017-02-28 | Athletic band with removable module | 2015-11-19 | A device for monitoring athletic activity of a user. In one example, the device has a sensor, and executes a calibration process using data received from the sensor to determine a correct positioning of the device on the user. | 1. A device configured to be worn on an arm of a user comprising: a first component configured to be worn at any one of a plurality of locations along an arm of a user and further configured to removably receive at least a portion of a second component;a second component comprising a housing, the housing including: a sensor, such that upon proper reception of at least a portion of the housing of the second component to the first component, the sensor is aligned at a first orientation that is directly on or over the arm of the user;a processor; anda non-transitory computer-readable medium that comprises computer-executable instructions, that when executed by the processor, perform a calibration comprising at least: receive first sensor data from the sensor at a first time period in which: (1) the first component is located at a first location of the plurality of locations, and (2) the second component is properly received with respect to the first component;calculate, using the first sensor data, a first value of a physiological parameter for the user;instruct the user to perform a specific activity configured to trigger a known tension within a specific muscle within the arm at the first location;receive second sensor data from the sensor during the specific activity;calculate, using the second sensor data, a second value of the physiological parameter for the user; andcompare the first value for the physiological parameter to the second value for the physiological parameter,wherein if the first value for the physiological parameter is within a threshold range of the second value for the physiological parameter, communicate a notification to the user that a calibration is complete, andwherein if the first value for the physiological parameter is outside of a threshold range of the second value for the physiological parameter, instruct the user to reposition the device on the arm to a second location within the plurality of locations. 2. The device of claim 1, wherein the specific activity is configured to cause movement to trigger a known tension within a specific muscle of a forearm of the user. 3. The device of claim 1, wherein the first sensor data is associated with a period of rest for the user. 4. The device of claim 1, wherein when at the first location, the device is worn on the forearm of the user in a position having a first distance to the user's elbow, and wherein the second position is located on the forearm at a second distance to the user's elbow, wherein the second distance is longer than the first distance. 5. The device of claim 1, wherein the physiological parameter comprises a heart rate. 6. The device of claim 1, wherein the sensor comprises an optical heart rate sensor. 7. The device of claim 1, wherein the sensor comprises an accelerometer. 8. The device of claim 1, wherein the comparison of the first value for the physiological parameter to the second value for the physiological parameter accounts for demographic in... | NIKE, Inc. | A61B005/024 A61B005/11 A61B005/00 A61B005/0205 A41D001/00 A63B024/00 A41D020/00 G08B006/00 G06F003/02 G06F003/01 G06F001/16 A41D027/20 B29C065/02 B29C065/52 B29C065/78 B29C065/00 A41H043/04 G06F019/00 B29L029/00 B29L031/48 A45F005/00
US09330462 | B2 | 2016-05-03 | Object information acquiring apparatus and control method of object information acquiring apparatus | 2014-04-11 | An object information acquiring apparatus that acquires information inside an object by receiving an acoustic wave that has arrived from inside the object through a layer having an acoustic impedance that is different from that of the object, and analyzing the acoustic wave, the object information acquiring apparatus comprises an acoustic wave probe that receives an acoustic wave and converts the acoustic wave into an electric signal; a whole image generation unit that generates a whole image, which is an image indicating information inside the object, based on the electric signal after the conversion; a partial image generation unit that extracts a partial image, which is a part of the whole image, from the whole image; and a similar image search unit that searches an area similar to the partial image, from the whole image. | 1. An object information acquiring apparatus that acquires information inside an object by receiving an acoustic wave that has arrived from inside the object through a layer having an acoustic impedance that is different from that of the object, and analyzing the acoustic wave, the object information acquiring apparatus comprising: an acoustic wave probe that receives an acoustic wave and converts the acoustic wave into an electric signal;a whole image generation unit that generates a whole image, which is an image indicating information inside the object, based on the electric signal after the conversion;a partial image generation unit that extracts a partial image, which is a part of the whole image, from the whole image; anda similar image search unit that searches an area similar to the partial image, from the whole image,wherein the similar image search unit searches an area similar to the partial image from a start position, which is a position where the partial image has been extracted, in a direction where the acoustic wave that has arrived from inside the object is reflected by the layer having acoustic impedance that is different from that of the object. 2. The object information acquiring apparatus according to claim 1, wherein the similar image search unit specifies a false image that has appeared on the whole image, based on the search result. 3. The object information acquiring apparatus according to claim 2, further comprising: a display unit that displays the whole image and highlights a false image specified by the similar image search unit. 4. The object information acquiring apparatus according to claim 2, further comprising: a removal unit that removes a false image specified by the similar image search unit, from the whole image; anda display unit that displays the whole image after the removing unit removes the false image. 5. The object information acquiring apparatus according to claim 1, wherein the partial image generation unit determines an area to extract a partial image, based on pixel values of pixels of the whole image. 6. The object information acquiring apparatus according to claim 5, wherein the partial image generation unit extracts, as a partial image, an area constituted by pixels having a brightness of a predetermined value or more, out of the whole image. 7. The object information acquiring apparatus according to claim 1, wherein the partial image generation unit determines an area to extract a partial image, based on the location of the area in the whole image. 8. The object information acquiring apparatus according to claim 7, wherein the partial image generation unit extracts a partial image from an area of which distance from a surface of the object is shorter than a predetermined value, out of the whole image. 9. The object information acquiring apparatus according to claim 2, wherein the whole image is an image that indicates an intensity of the received acoustic wave, and the object informatio... | CANON KABUSHIKI KAISHA | G06K009/68 G06T007/00 A61B005/00 G06T005/00 A61B008/08
US10342475 | B2 | 2019-07-09 | Threshold estimation apparatus, threshold estimation method, and non-transitory computer-readable information recording medium | 2015-08-07 | In a threshold estimation apparatus, a setting unit sets an initial intensity value for each electrode. A selection unit randomly selects an electrode. A stimulation unit, via the selected electrode, stimulates a living subject at an intensity set for the electrode. A detection unit detects the presence or absence of a response evoked by the stimulation. An update unit reduces or increases the intensity associated with the selected electrode, based on the presence or absence of the response evoked by the stimulation. An estimation unit, if the thus updated intensity associated with each electrode satisfies a convergence condition, estimates a convergence value as a threshold associated with the electrode. The threshold estimation apparatus repeats these processes until thresholds associated with all electrodes are estimated. If the same electrode was consecutively selected, the time interval between a stimulation and a subsequent stimulation is adjusted. | 1. A threshold estimation apparatus including a plurality of stimulators being electrodes or coils applying stimuli respectively to a living subject, the threshold estimation apparatus including a processor and a memory which stores a program executable by the processor, the program comprising: a setting code that causes the processor to set an initial value of an intensity associated with each of the plurality of stimulators;a selecting code that causes the processor to select randomly, from the plurality of stimulators, a selected stimulator for which a threshold associated therewith has not been estimated;a stimulating code that causes the processor to stimulate the living subject via the selected stimulator at an intensity associated with the selected stimulator;a detecting code that causes the processor to detect whether or not the living subject generated a response evoked by the applied stimulation;an updating code that causes the processor to update the intensity by reducing the intensity associated with the selected stimulator if the response was generated, orincreasing the intensity associated with the selected stimulator if the response was not generated; andan estimating code that causes the processor to estimate a convergence value of the updated intensity associated with each stimulator as a threshold that satisfies a convergence condition,wherein,the selecting code, the stimulating code, the detecting code, the updating code and the estimating code are repeatedly executed by the processor until all thresholds associated with all of the plurality of stimulators are estimated;when stimulators consecutively selected by the selecting code are:(a) the same stimulator, a first stimulus is applied via the stimulator and then, after a first stand-by period, a subsequent stimulus is applied via the same stimulator,(b) when stimulators consecutively selected by the selecting code are different stimulators, a first stimulus is applied via one of the consecutively selected stimulators that was selected first and then, after a second stand-by period, a subsequent stimulus is applied via the other stimulator of the consecutively selected stimulators that was selected later; andthe second stand-by period is shorter than the first stand-by period. 2. The threshold estimation apparatus according to claim 1, further comprising a generating code that causes the processor to generate a threshold map representing a distribution of the thresholds by depicting the thresholds on relative positions of the plurality of stimulators respectively, by concentration or contour. 3. The threshold estimation apparatus according to claim 2, wherein the threshold estimation apparatus, after generating the threshold map, deletes an intensity update history associated with each of the plurality of stimulators, and again performs the repeating process as well as the generation of a threshold map using the estimated thresholds associated with the respective stimulators... | RIKEN | A61B005/00 A61B005/0488 A61B005/04 A61N001/37 A61N001/36
US10265139 | B2 | 2019-04-23 | Medicine ingestion state management method, medicine and medicine ingestion state management device | 2016-01-26 | Disclosed is a medicine ingestion state management method capable of objectively managing medicine ingestion states of patients. The medicine ingestion state management method involves an operation of prescribing medicine (10) encapsulating, together with a medicament, a medicine information transmitting unit (20) having a function of transmitting medicine information capable of specifying a type and a quantity of the medicament to each individual patient, and an operation of grasping the medicine ingestion state of each patient by collecting the medicine information from each medicine information transmitting unit (20) in each patient. | 1. A swallowable medicament, comprising: a combination of a medicine and an IC chip, the combination being made using a material soluble by gastric juice,wherein the IC chip includes an electrode and starts a given operation using power generated by exposure of the electrode to the gastric juice,wherein the exposure of the electrode to the gastric juice is caused by dissolution of the material by the gastric juice. 2. The swallowable medicament according to claim 1, wherein the IC chip includes a memory, and the given operation includes an access to the memory. 3. A medicament, comprising: a medicine; andan IC chip exists in the medicament in a status mixed with the medicine,wherein the IC chip has an electrode that forms a battery by contact with gastric juice and the IC chip operates in response to power supply from the battery,wherein the contact between the electrode and the gastric juice is caused by dissolution of a material soluble by the gastric juice covering the electrode. 4. A medicament, comprising: an IC chip that has an electrode to form a battery within a body and starts a given operation using power generated by exposure of the electrode to gastric juice; anda medicine that exists on a contact position with the electrode of the IC chip and is dissolved within the body,wherein the exposure of the electrode to the gastric juice is caused by dissolution of the material by the gastric juice. | FUJITSU LIMITED | A61B090/98 A61B005/00 A61B005/07 G06F019/00 G06Q010/10 G16H010/60 G16H020/13 A61J003/00
US20160128577 | A1 | 2016-05-12 | System for Measuring Skin Hydration | 2016-01-18 | A system for measuring skin hydration at a measurement site with a independent measurement probe and a portable data processing system. Wireless communications enable the transfer of commands and data between the probe and portable data processing system. The portable data processing system can comprise a wireless tablet or smartphone. The portable data processing system provides input functions and output functions on its display and includes a camera for imaging the measurement site. | 1. A system for measuring skin hydration comprising: A) a probe for measuring skin impedance at a measurement site including an independently powered measurement control for obtaining measurements in response to measurement commands and for generating information including the outcome of the skin impedance measurement,B) a portable data processing system for generating the measurement commands and for receiving measurement information generated said probe, said data processing system having an operator interface and visual display, andC) a wireless communications network interconnecting said probe and said portable data processing system whereby skin impedance measurements can be displayed at said portable data processing system. 2. The system as recited in claim 1 wherein said portable data processing system additionally includes imaging means for imaging and displaying the measurement site. |  | A61B005/00 A61B005/053
US09339211 | B2 | 2016-05-17 | Body-worn monitor for measuring respiration rate | 2014-05-31 | The invention provides a multi-sensor system that uses an algorithm based on adaptive filtering to monitor a patient's respiratory rate. The system features an impedance pneumography sensor and a motion sensor (e.g., an accelerometer) configured to attach to the patient's torso and measure therefrom a motion signal. The system further comprises a processing system, configured to operably connect to the impedance pneumography sensor and motion sensor, and to determine a respiration rate value by applying filter parameters obtained from the impedance pneumography sensor signals to the motion sensor signals. | 1. A system for continuously monitoring a patient's respiration rate comprising: a combined impedance pneumography, electrocardiography (ECG), and motion sensor unit comprising a housing configured to be mounted on the patient's torso and comprising: an impedance pneumography circuit within the housing configured to generate an impedance pneumography waveform, an ECG circuit within the housing configured to generate an ECG waveform, and an accelerometer within the housing configured to generate a motion waveform;a photoplethysmography sensor configured to be mounted on a finger of the patient and to generate a photoplethysmography waveform; anda processing system configured to, on a continuous basis;(i) receive the impedance pneumography waveform, the ECG waveform, the motion waveform, and the photoplethysmography waveform,(ii) select using a metric of signal quality one of the received waveforms to provide an initial respiration rate,(iii) determine the initial respiration rate from the waveform selected in (ii), and(iv) process one or more of the received waveforms not selected in (ii) using an adaptive filtering algorithm that uses parameters determined from the initial respiration rate to obtain and report the respiration rate for the patient. | SOTERA WIRELESS, INC. | A61B005/08 A61B005/0205 A61B005/0402 A61B005/11 A61B005/113 A61B005/00 A61B005/024 A61B005/053
US09351658 | B2 | 2016-05-31 | Device and method for sensing electrical activity in tissue | 2006-08-08 | An exemplary embodiment providing one or more improvements includes apparatus and methods for sensing electrical activity in tissue of a person in a manner which is substantially limits or eliminates interference from noise in a surrounding environment. | 1. A method comprising: detecting electrical activity from a tissue of a user via a sensor of a headpiece worn by the user, the sensor including: a sensor electrode to contact a first point on skin of the user;a reference electrode to contact a second point on the skin of the user, the contact between the reference electrode and the skin of the user to cause a contact resistance between a surface of the reference electrode and the skin of the user; andan electronics module positioned proximate the reference electrode to reduce noise, a local reference potential in the electronics module and a potential of the second point on the skin being approximately equal to allow the surface of the reference electrode to be in direct contact with the skin of the user at the second point without having to prepare the skin or use a material that decreases the contact resistance, the electrical activity represented by a voltage signal including a signal of interest, undesired signals, and noise, the electronics module including: an amplifier to (1) amplify the signal of interest and the undesired signals within a first frequency range while (2) filtering the noise occurring in a second frequency range, which includes frequencies higher than the first frequency range, prior to transmitting the amplified signal of interest and undesired signals to one or more filters; andtransforming, with a processor and a Fourier transform, the signal of interest to form an output signal. 2. A method as defined in claim 1, wherein the amplifier is to amplify the signal of interest and the undesired signals by using the local reference potential of the electronics module, and connecting the second point directly to the local reference potential. 3. A method as defined in claim 1, wherein the amplifier is to amplify the signal of interest at the electronics module by using the local reference potential and source voltages, and positioning the electronics module in a region of the headpiece adjacent to the second point so as to electrically couple the local reference potential and source voltages to the second point. 4. A method as defined in claim 1, wherein the amplifier is to amplify the signal of interest at the electronics module and detecting the electrical activity includes using conductors extending from the electronics module to the sensor and reference electrodes at the first and second points, and electrically isolating the electronics module by keeping the electronics module free of wires, optical fibers or other extensions other than the conductors extending to the sensor and reference electrodes. 5. A method as defined in claim 1 further including filtering an output signal of the amplifier with the electronics module to attenuate the undesired signals and to isolate the signal of interest from the undesired signals. 6. A method as defined in claim 5 further including communicating the signal of interest to a computing device. 7. A method as defined in claim 6,... | The Nielsen Company (US), LLC | A61B005/0476 A61B005/04 A61B005/048 A61B005/00
US09351662 | B2 | 2016-05-31 | MRI scanner that outputs bone strength indicators | 2011-06-17 | Described herein are various technologies pertaining to diagnosing and/or prescribing treatment for osteoporosis. A bone of a patient is subjected to an MRI scan, and the resultant signal is subjected to a likelihood function. The output of the likelihood function are values that are parameters that are employed to ascertain connectivity of trabeculae in the bone of the patient and volumetric trabecular density of the bone. | 1. A method, comprising: using a magnetic resonance imaging (MRI) system, performing one-dimensional magnetic resonance scans of a bone of a patient to generate signals, the magnetic resonance scans performed at multiple respective angles relative to the bone of the patient;evaluating, by a processor, a likelihood function over the signals;generating, by the processor, values that are indicative of size, aspect ratio, and orientation of marrow phase pockets of the bone of the patient based upon the evaluating of the likelihood function over the signals; andproviding data to an operator of the MRI system that is indicative of strength of the bone of the patient, the data based upon the values that are indicative of the size, aspect ratio, and orientation of the marrow phase pockets of the bone of the patient. 2. The method of claim 1, further comprising generating a value that is indicative of at least one of connectivity of the bone of the patient or volumetric trabecular density of the bone of the patient based upon the values that are indicative of the size, aspect ratio, and orientation of the marrow phase pockets of the bone of the patient, the data provided to the operator is based upon the value that is indicative of the at least one of connectivity of the bone of the patient or the volumetric trabecular density of the bone of the patient. 3. The method of claim 1, wherein the likelihood function is based upon a model of trabecular bone microstructure that comprises a model of marrow phase pockets of the trabecular bone microstructure. 4. The method of claim 3, further comprising modeling the marrow phase pockets of the trabecular bone microstructure as at least one of an ellipsoid model, a spherical model, a cuboid model, or a cylindrical model. 5. The method of claim 3, further comprising generating the model of the trabecular bone microstructure based upon ex vivo experiments on trabecular bones, the likelihood function based upon the model of the trabecular bone microstructure. 6. The method of claim 3, further comprising generating the model of trabecular bone microstructure based upon a plurality of numerical simulations, the likelihood function based upon the model of the trabecular bone microstructure. 7. The method of claim 1, the data provided to the operator is further indicative of a probability that the patient has osteoporosis. 8. The method of claim 1, wherein memory that is accessible to the processor comprises a computer-implemented model of trabecular bone microstructure, and wherein the likelihood function is evaluated over the signals by the processor based upon the computer-implemented model of the trabecular bone microstructure. 9. A system comprising: a magnetic resonance imaging (MRI) scanner that is configured to generate magnetic resonance signals based upon respective magnetic resonance scans performed over a portion of a trabecular bone of a patient, the magnetic resonance scans performed at different ang... | Microsoft Technology Licensing, LLC | A61B005/055 A61B005/00
US09326686 | B2 | 2016-05-03 | System and method for mitigating the effects of tissue blood volume changes to aid in diagnosing infiltration or extravasation in animalia tissue | 2013-07-30 | A system including a sensor and a device coupled to the sensor. The sensor is configured to detect in Animalia tissue (i) a first electromagnetic radiation extinction dominated by absorption of a first wavelength and (ii) a second electromagnetic radiation extinction dominated by scattering of a second wavelength. The device is configured to aid in diagnosing at least one of infiltration and extravasation in the Animalia tissue based on the first and second electromagnetic radiation extinctions detected by the sensor. | 1. A system to aid in diagnosing at least one of infiltration and extravasation in Animalia tissue, the system comprising: a sensor configured to emit first and second signals entering the Animalia tissue and to detect third and fourth signals exiting the Animalia tissue, the first signal having a peak wavelength between 800 nanometers and 1,050 nanometers, the second signal having a peak wavelength between 560 nanometers and 660 nanometers, the third signal including a portion of the first signal that is at least one of reflected, scattered and redirected from the Animalia tissue, and the fourth signal including a portion of the second signal that is at least one of reflected, scattered and redirected from the Animalia tissue; anda device coupled to the sensor and configured to output a notice based on the third and fourth signals, the device is configured to detect infusate accumulation over time in the Animalia tissue based on the third signal, and the device is configured to detect tissue blood volume changes in the Animalia tissue based on the fourth signal, the device includes: an analog-to-digital converter configured to (i) represent the third signal with a first sequence of values and (ii) represent the fourth signal with a second sequence of values; anda processor coupled to the analog-to-digital converter, the processor being configured to (i) compute a predicted sequence of values based on the second sequence of values and (ii) compare the first and predicted sequences of values. 2. The system of claim 1 wherein the device comprises an indicator coupled to the processor, the indicator being configured to output the notice when (i) the predicted sequence of values is less than a first threshold value and (ii) the predicted sequence of values and the first sequence of values diverge less than a second threshold value. 3. The system of claim 2 wherein the first threshold value is 0.90 and the second threshold value is 0.05. | ivWatch, LLC | A61B005/00 A61M005/168 A61M005/50 A61F013/02 A61B005/1455 A61B005/05 A61M005/172
US09351688 | B2 | 2016-05-31 | Low power monitoring systems and method | 2013-01-29 | The present disclosure relates to systems and methods for collecting patient data via a monitoring system, with reduced power consumption. In one embodiment, the monitoring system is configured to emit pulses of light, and detect the light after passing through patient tissue. The light data is emitted sporadically, and a waveform is reconstructed from the sporadically sampled light data. Physiological parameters from the patient may be calculated from the reconstructed waveform. The sporadic sampling may reduce the power consumption by the monitoring system. | 1. A system comprising: light drive circuitry configured to drive an emitter of a pulse oximetry sensor at an average frequency and an average pulse width to cause the emitter to emit light into a tissue of the patient at the average frequency and the average pulse width;a processor configured to: receive a set of data samples generated by a detector of the pulse oximetry sensor based on the light that passed through the tissue of the patient, wherein the set of data samples are indicative of a physiological signal of the patient, and the average frequency and the average pulse width cause the set of data samples to comprise an average sampling frequency less than twice the Nyquist frequency of the physiological signal;construct a waveform representative of the physiological signal based on the set of data samples; andcalculate a physiological parameter based on the waveform. 2. The system of claim 1, comprising the pulse oximetry sensor, wherein the pulse oximetry sensor and the processor are coupled wirelessly. 3. The system of claim 1, further comprising a display coupled to the processor configured to display the waveform and/or the physiological parameter. 4. The system of claim 3, wherein the display is configured to display a portion of the waveform spanning a first time duration, and wherein the processor is configured to update the waveform based on an updated set of data samples spanning a second time duration shorter than the first time duration. 5. The system of claim 1, wherein the set of data samples comprise an average sampling frequency less than the Nyquist frequency of the physiological signal. 6. The system of claim 1, wherein the set of data samples comprise an average sampling frequency less than a fully sampled frequency of the physiological signal. 7. The system of claim 1, wherein the set of data samples comprise an average sampling frequency less than one quarter of a fully sampled frequency of the physiological signal. 8. The system of claim 1, wherein the processor is configured to construct the waveform by matching pursuit. 9. The system of claim 8, wherein the processor is configured to access a dictionary of components. 10. The system of claim 9, wherein the dictionary is customized based on the physiological signal. 11. The system of claim 1, comprising the pulse oximetry sensor, the emitter, and the detector, wherein the emitter comprises a first light emitting diode (LED) configured to emit red light and a second LED configured to emit infrared light, the detector is configured to generate the set of data samples corresponding to the red light and a second set of data samples corresponding to the infrared light, the processor is configured to construct the waveform by matching pursuit based on the set of data samples, and the processor is configured to construct a second waveform by shifting and/or scaling the waveform based on the second set of data samples. 12. The system of claim 11, wherein the proc... | COVIDIEN LP | A61B005/1455 A61B005/00
US09345532 | B2 | 2016-05-24 | Methods and devices for ablation of tissue | 2012-05-11 | Various apparatus and methods for ablating tissue are described therein. In certain variations, a method for ablating tissue may include advancing one or more electrodes, e.g., a first electrode and a second electrode, to or near a target tissue in a subject. The electrodes may be energized such that the electrodes create a plasma arc. The plasma arc generates heat which desiccates or ablates the target tissue. An apparatus for ablating tissue may include an elongate instrument, e.g., a tube or rod. One or more electrodes, e.g., a first electrode and a second electrode, may be provided in the elongate instrument. The electrodes may be energized such that the electrodes create a plasma arc for generating heat to desiccate or ablate a target tissue. | 1. A method for ablating tissue, the method comprising: advancing an instrument having a first electrode and a second electrode to or near a target tissue in a subject; andenergizing the electrodes such that the electrodes create a plasma arc encapsulated within the instrument, wherein the plasma arc generates heat; and ablating the target tissue using the heat generated by the plasma arc. 2. The method of claim 1, wherein the plasma arc is positioned next to a tumor such that the heat desiccates or ablates blood vessels supplying blood to the tumor. 3. The method of claim 1, wherein the plasma arc generates a heated steam which desiccates and ablates the target tissue. 4. The method of claim 1, wherein the first and second electrodes are positioned within an elongate instrument, wherein the elongate instrument comprises an electrical cutting distal tip for cutting tissue to advance the electrodes to the target tissue. 5. The method of claim 1, wherein light generated by the plasma arc provides feedback regarding the localization of the plasma arc in the patient. 6. The method of claim 5, wherein the light is visible through the subject's skin, e.g., allowing for localization of the apparatus. 7. The method of claim 1, further comprising, providing imaging to monitor target tissue shrinkage. 8. The method of claim wherein the first and second electrodes are advanced through an opening through an airway wall to position the plasma arc next to a target tissue or tumor in the lung. 9. The method of claim 1, wherein the heat or plasma arc may create temperatures ranging from about 1000 to about 6000 degrees C. 10. An apparatus for ablating tissue comprising: an elongate instrument;a first electrode and a second electrode enclosed by a wall of the elongate instrument, wherein the wall of the elongate instrument enclosing the electrodes is glass, and wherein the electrodes are configured to create a plasma arc within the instrument for generating heat which is thermally conducted through the wall at the elongate instrument to the tissue to be ablated, and wherein said wall is shaped and positioned around the electrodes such that the plasma arc is captured by said elongate instrument. 11. The apparatus of claim 10, wherein the elongate instrument comprises a tube having a first channel and a second channel, wherein at least a portion of the first electrode is positioned within the first channel and at least a portion of the second electrode is positioned in the second channel. 12. The apparatus of claim 10, wherein the elongate instrument is configured to encapsulate vapor released from the electrodes. 13. The apparatus of claim 10, wherein the electrodes are configured to burn off after a period of time. 14. The apparatus of claim 10, wherein the electrodes are made from tungsten or molybdenum. 15. The apparatus of claim 10, wherein the elongate instrument comprises a conductive cap positioned at a distal end of the elongate instrument, w... | Broncus Medical Inc. | A61B018/04 A61B018/14 A61B018/08 A61B018/00 A61B018/12 A61B019/00 A61B005/06
US09345387 | B2 | 2016-05-24 | Preventing instrument/tissue collisions | 2007-06-13 | Collisions between a minimally invasive surgical instrument and patient tissue are prevented in various ways. In one aspect a body cavity is mapped by recording positions of a surgical instrument, and the map is used to insert another surgical instrument. In another aspect, a surgical instrument is inserted into a space vacated by a guide tube. In yet another aspect, a volume is defined in which a control system permits an instrument to move, and the volume is a sum of one volume that is defined by previous movements of the instrument and other volume that is defined by the boundaries of an image capture component. In yet another aspect, a real time image of a surgical site is mosaiced over a previously recorded more distal image of the site. Surgical instruments are visible in the real time image, and representations of the surgical instruments as they would appear in the previously recorded image are generated and displayed on the previously recorded image. Consequently, a person moving the surgical instruments sees a representation of the instruments outside the field of view of an imaging system taking the real time images. | 1. A minimally invasive surgical system comprising: a first surgical device comprising a distal end;an actuator control system coupled to the first surgical device to move the distal end to a plurality of positions within a body cavity as the first surgical device is advanced within the body cavity;a memory coupled to the actuator control system, wherein the actuator control system stores, in the memory, data that defines the plurality of positions; anda second surgical device coupled to the actuator control system, wherein the actuator control system advances the second surgical device by using the stored data that defines the plurality of positions. 2. The system of claim 1, further comprising: a position sensor associated with the first surgical device;wherein the position sensor determines the plurality of positions of the distal end of the first surgical device. 3. The system of claim 2: wherein the position sensor comprises a bend sensor. 4. The system of claim 1, further comprising: a sensor positioned outside the patient;wherein the sensor senses the plurality of positions of the distal end of the first surgical device. 5. A method of advancing a minimally invasive surgical device, comprising: storing data by an actuator control system in a memory, wherein the data defines a plurality of positions of a part of a first surgical device within a body cavity as the first surgical device is advanced within the body cavity; andadvancing a second surgical device within the body cavity by the actuator control system using the stored data that defines the plurality of positions. 6. The method of claim 5: wherein the plurality of positions of the part of the first surgical device is determined from a position sensor associated with the first surgical device. 7. The system of claim 6: wherein the position sensor comprises a bend sensor. 8. The method of claim 5: wherein the plurality of positions of the part of the first surgical instrument is determined from information sensed by a sensor positioned outside of a patient. 9. The minimally invasive surgical system of claim 1: the first surgical device comprising a first guide tube having at least one port proximal to the distal end;the second surgical device comprising a second guide tube extending from the port; anda third surgical device extending from the second guide tube. | INTUITIVE SURGICAL OPERATIONS, INC. | A61B001/00 A61B019/00 A61B001/002 A61B001/04 A61B001/06 A61B005/00 A61B008/12 A61B001/05 A61B017/00 A61B001/005 A61B001/018 A61B017/34
US10362970 | B2 | 2019-07-30 | Phase-to-amplitude/slope mapping | 2016-01-20 | A method includes obtaining a signal that includes a plurality of cycles and generating a map that maps motion phases to the signal based on both an amplitude and a slope of the signal. A system includes a processor that identifies a set of motion signal timestamps, for a plurality of motion cycles in a motion signal indicative of cyclic motion of a moving object, based on a predetermined motion phase of interest and a phase-to-amplitude/slope mapping, wherein the set of motion signal timestamps correspond to a common signal amplitude. A method include identifying a peak of a plurality of peaks in a motion cycle of a noisy cyclic signal having irregular periodicity, wherein the peak corresponds to a point lying between two points with amplitudes below a predetermined threshold, comparing points before and after the peak with the peak, and identifying the peak as a local maximum when the peak is greater than the points. | 1. A system, comprising: a processor configured to:identify a set of motion signal timestamps, for a plurality of motion cycles in a continuous cyclical motion signal indicative of cyclic motion of a moving object, based on a predetermined motion phase of interest and a phase-to-amplitude/slope mapping, wherein the set of motion signal timestamps correspond to a common signal amplitude, wherein each descending amplitude region and each ascending amplitude region of the continuous cyclical motion signal is mapped using a lesser of a mean maximum and a local maximum, and a greater of a mean minimum and a local minimum, and divided for at least one of the local maximum less than the mean maximum or the local minimum greater than the mean minimum; andgenerate volumetric image data based on data acquired by a scanner and the identified set of motion signal timestamps. 2. The system of claim 1, further comprising: a signal characteristic identifier that identifies signal maxima and minima for one or more cycles of the motion signal; anda signal characteristic processor that determines the mean maximum and the mean minimum based on the identified maxima and minima. 3. The system of claim 2, wherein the signal characteristic identifier identifies a maximum for an evaluation range within a motion cycle as a point lying between two points with amplitudes below a maximum determining threshold, compares other points in the evaluation range before and after the maxima with the maximum; and identifies the maxima as a local maximum for the evaluation range when the maximum is greater than the compared other points. 4. The system of claim 2, further comprising: a signal slope identifier that identifies a descending amplitude region and an ascending amplitude region for each cycle; anda phase-to-amplitude/slope mapper that maps motion phases to the motion signal based on the mean maximum and the mean minimum and the descending and ascending amplitude regions. 5. The system of claim 4, further comprising: a timestamp identifier that identifies timestamps of the motion signal that correspond to predetermined phase of interest and a common signal amplitude. 6. The system of claim 5, wherein the timestamp identifier selects a phase based on a linearly divided amplitude in at least one of the ascending amplitude region or the descending amplitude region. 7. The system of claim 4, wherein the phase-to-amplitude/slope mapper maps motion phases the ascending amplitude regions separately from the descending amplitude regions. 8. The system of claim 4, wherein the phase-to-amplitude/slope mapper maps the motion phases dividing linearly each descending amplitude region and each ascending amplitude region between the mean maximum and the mean minimum by amplitude. 9. The system of claim 2, wherein the signal characteristic identifier removes at least one of the identified maxima based on a time distance between a preceding identified maxima and a subsequent local max... | KONINKLIJKE PHILIPS N.V | A61B005/00 A61B006/00 G06K009/00 A61B005/113 G06F017/00
US09336672 | B2 | 2016-05-10 | Healthcare communication system for programming bed alarms | 2015-04-14 | A system that monitors various conditions of a plurality of hospital beds located in different rooms of a healthcare facility is provided. Alternatively or additionally, other types of equipment may be monitored by the system. Various configurations of network interface units that are coupleable to or integrated into a hospital bed are also disclosed. The system receives data from the hospital beds and/or other equipment and initiates a communication to a wireless communication device of at least one designated caregiver in response to the received data being indicative of an alarm condition. | 1. A system for alerting caregivers carrying wireless communication devices of alarm conditions associated with patient beds in a healthcare facility, the system comprising at least one computer device that is remote from the patient beds and that receives data from the patient beds, a display screen associated with the computer device, the computer device being operable to display at least one alarm selection screen on the display screen, the at least one alarm selection screen permitting a user to select at least one alarm condition for at least one of the patient beds that must be violated before an alarm condition is considered to exist, the computer device being operable to initiate a communication to the wireless communication device of at least one of the designated caregivers in response to the alarm condition. 2. The system of claim 1, wherein the at least one alarm condition comprises one or more of the following: an angle at which a head section of a patient bed is raised, a height at which an upper frame of the patient bed is raised, a position of a siderail of the patient bed, a patient position with respect to the patient bed, and a patient weight as measured by a weigh scale system of the patient bed. 3. The system of claim 1, wherein the at least one alarm condition comprises an angle below which a head section of a patient bed should not be lowered. 4. The system of claim 1, wherein the at least one alarm condition comprises an angle above which a head section of a patient bed should not be raised. 5. The system of claim 1, wherein the at least one alarm condition is transmitted to the patient bed from the at least one computer device and the patient bed monitors itself for the occurrence of the at least one alarm condition. 6. The system of claim 1, wherein the at least one alarm condition is assigned automatically based on patient acuity level. 7. The system of claim 1, wherein the at least one alarm condition changes from a first value to a second value after an elapsed amount of time. 8. The system of claim 7, wherein the elapsed amount of time is selectable using the computer device. 9. The system of claim 1, wherein the at least one alarm condition changes from a first value to a second value at a preprogrammed time. 10. The system of claim 1, wherein the communication initiated by the computer device comprises an alphanumeric message destined for receipt by the wireless communication device of the at least one caregiver. 11. The system of claim 1, wherein the communication initiated by the computer device comprises an audio message destined for receipt by the wireless communication device of the at least one caregiver. 12. The system of claim 1, wherein the computer device is operable to initiate communications with pagers, wireless badges, and wireless handsets. 13. The system of claim 1, wherein the computer device is operable to communicate via a network of the healthcare facility with any at least one of th... | Hill-Rom Services, Inc. | G08B001/08 G08B021/02 A61B005/11 A61B005/00 G06F019/00 G07C003/00 G08B021/04 G08B025/01 G08B005/22 G08C019/00 G08B025/00
US09526618 | B2 | 2016-12-27 | Method and device for percutaneous left ventricular reconstruction | 2015-08-28 | A method for reducing left ventricular volume, which comprises identifying infarcted tissue during open chest surgery; reducing left ventricle volume while preserving the ventricular apex; and realigning the ventricular apex, such that the realigning step comprises closing the lower or apical portion of said ventricle to achieve appropriate functional contractile geometry of said ventricle in a dyskinetic ventricle of a heart. | 1. A method for reducing a chamber volume in a heart comprising: applying suction to an epicardium of said chamber to bring opposing walls of said chamber adjacent one another and thereby reduce a size of said chamber, said opposing walls including a first wall and a second wall; andpenetrating a needle across said first and second walls of said chamber with said first wall adjacent said second wall;wherein said suction is applied by engaging the epicardium with a suction device so that the epicardium is drawn into the suction device. 2. The method of claim 1, wherein the epicardium is drawn into the suction device so that inner surfaces of said first wall and second wall are adjacent one another. 3. The method of claim 1, further comprising sensing infarcted tissue of the first wall or second wall. 4. The method of claim 1, further comprising: clamping said first and second walls; andsecuring a clamped portion of said first and second walls. 5. The method of claim 4, further comprising applying a patch on the clamped portion. 6. The method of claim 1, further comprising: identifying infarcted tissue;reducing a volume of said chamber while preserving an apex of the chamber; andrealigning the apex of the chamber;wherein said realigning step comprises closing of a lower or apical portion of said chamber to achieve appropriate functional contractile geometry of said chamber. 7. The method of claim 6, wherein the step of realigning the apex of the chamber further comprises the steps of altering the geometry of the chamber and decreasing wall tension. 8. The method of claim 1, further comprising applying stitching to said first and second walls of said chamber to affix said first and second walls together. 9. The method of claim 3, wherein the step of sensing the infarcted tissue is performed by a procedure selected from the group consisting of: pacing, angiography, echocardiography, and tissue Doppler. 10. The method of claim 1, wherein said method is a minimally invasive surgery. | BioVentrix, Inc. | A61F002/24 A61B017/04 A61B005/042 A61B008/12 A61B017/00 A61B018/00
US09327138 | B2 | 2016-05-03 | Hair treatment system and method | 2014-07-05 | A hand held device generates a predetermined number of pulses of electromagnetic radiation having a predetermined electromagnetic spectrum, a predetermined duration, a predetermined inter-pulse interval, and a predetermined total energy. The pulse sequence is delivered to a skin surface to reduce or eliminate Xray or ultraviolet radiation damage to the skin surface. | 1. A method for treating a skin area of a human subject, the skin area including a surface portion and hair wholly below the surface portion, comprising: administering to the subject at the skin area light to retard growth of the hair,the light including a first light application during a first treatment session and a second light application during a second treatment session, andthe first light application and the second light application separated by a time period during which no light treatment occurs, andthe time period extending longer than a time-delay between treatment pulses but for a duration less than the amount of time required for protruding growth of the hair following the first light application absent any light application following the first light application. 2. The method of claim 1, wherein the first light application includes light to retard growth of the hair. 3. The method of claim 1, wherein the second light application includes light to retard hair growth of the hair. 4. The method of claim 1, wherein the first light application includes light to retard growth of the hair and the second light application includes light to retard growth of the hair. 5. A hair treatment method comprising: applying light to a hair before the hair protrudes above the skin by delivering light to the hair below the skin surface,the light delivered in a first application during a first treatment session and a second application during a second treatment session,wherein no light treatment occurs between the first treatment session and the second treatment session,wherein the time between the first application and the second application is longer than a time-delay between treatment pulses, andwherein the light increases the time until the hair protrudes above the skin as compared to the time until the hair protrudes above the skin absent any applying of light. 6. The hair treatment method of claim 5, wherein one of the two separate applications includes more than one light pulse. 7. The hair treatment method of claim 5, wherein each of the two separate applications includes more than one light pulse. 8. The hair treatment method of claim 5, wherein one of the two separate applications increases the time until the hair protrudes above the skin as compared to the time until the hair protrudes above the skin absent any applying of light. 9. The hair treatment method of claim 5, wherein each of the two separate applications increases the time until the hair protrudes above the skin as compared to the time until the hair protrudes above the skin absent any applying of light. 10. The hair treatment method of claim 5, wherein the two separate applications together increase the time until the hair protrudes above the skin as compared to the time until the hair protrudes above the skin absent any applying of light. |  | A61N005/06 A61B018/18 A61B005/00 A61B005/103 A61B018/20 A61B017/00 A61B018/00 A61N005/067
US10159410 | B2 | 2018-12-25 | Method and apparatus for biometric data gathering and dissemination | 2014-11-24 | A system includes a processor configured to determine that a vehicle accident has occurred. The processor is also configured to connect to a detected wearable device. The processor is additionally configured to request measured present biometric information from the wearable device, in response to the accident determination. Also, the processor is configured to receive measured present biometric information from the wearable device and relay the biometric information to a connected emergency operator. | 1. A system comprising: a processor configured to:connect to a detected wearable device;retrieve a user-profile corresponding to a received wearable device identification;determine a user-location corresponding to the wearable device location; andadjust a vehicle safety system based on information contained in the user profile. 2. The system of claim 1, wherein the safety system includes a seat-belt restraint tensioner. 3. The system of claim 1, wherein the safety system includes an airbag. 4. The system of claim 1, wherein the processor is further configured to: request present biometric information from the wearable device, measurable by the wearable device;receive the present biometric information;compare the present biometric information to saved biometric information previously received from the wearable device; andverify the identity of a user based on a correlation between the received present biometric information and the saved biometric information. 5. The system of claim 1, wherein the processor is further configured to: request input of information lacking in the user profile, needed to perform one or more vehicle safety system adjustments; andupdate the user-profile based on the received information. | Ford Global Technologies, LLC | G08C019/22 H04Q009/00 A61B005/00 B60K028/02 G06F009/445 G06F019/00 G16H040/67 B60R021/00 A61B005/024 G08B025/01
US09332936 | B2 | 2016-05-10 | Concentration determination apparatus and concentration determination method for detecting an absorbance of living body tissue based on a light intensity and measuring a concentration of a measured component contained in living body tissue | 2015-03-24 | A concentration determination apparatus in accordance with exemplary embodiments has a storage body and a temperature control plate. The concentration determination apparatus may also have a first and second input light guide and a first and second light receiving light guide, and an operational unit that contains an absorbance calculation device for calculating an absorbance of liquid and living body tissue and a concentration calculation device for comparing the absorbance of the liquid and the living body tissue. | 1. A concentration determination apparatus comprising: a storage body that holds liquid;a temperature control plate having one surface configured for close contact to a living body tissue and another surface contacting the liquid;a first input light guide configured for causing a first incident light to be incident on the storage body;a first light receiving light guide configured for receiving a reference light transmitted through the liquid from the first input light guide;a second input light guide configured for causing a second incident light to be incident on the living body tissue;a second light receiving light guide configured for receiving a measurement light from the second input light guide via the living body tissue;a photodiode configured for detecting a light intensity of the reference light and a light intensity of the measurement light; andan operator comprising:an absorbance calculator configured for calculating an absorbance of the liquid and an absorbance of the living body tissue based on the light intensity of the reference light and the light intensity of the measurement light, anda concentration calculator configured for comparing the absorbance of the liquid with the absorbance of the living body tissue and calculating a concentration of a measured component in the living body tissue. 2. The concentration determination apparatus according to claim 1, wherein the liquid is a liquid having a scattering characteristic. 3. The concentration determination apparatus according to claim 1, wherein the liquid is water or liquid corresponding to body fluid. 4. The concentration determination apparatus according to claim 1, wherein the temperature control plate includes a thermal conductor configured for causing the body temperature of the living body tissue and the liquid temperature of the liquid to be heat-exchanged with each other. 5. The concentration determination apparatus according to claim 1, wherein the storage body further includes a reflector configured for reflecting the first incident light incident from the first input light guide to the first light receiving light guide. 6. The concentration determination apparatus according to claim 1, further comprising a plurality of light sources. 7. The concentration determination apparatus according to claim 1, wherein the storage body further comprises a temperature controller that heats or cools the liquid. 8. The concentration determination apparatus according to claim 1, further comprising: a light source; anda spectroscopic device configured for dividing light irradiated from the light source toward the first input light guide and the second input light guide. 9. The concentration determination apparatus according to claim 1, wherein the absorbance calculator is configured for respectively detecting the absorbance of the liquid and the absorbance of the living body tissue if a difference between the temperature of the living body tissue and the temperature of the li... | SEIKO EPSON CORPORATION | A61B005/1455 A61B005/01 A61B005/145 G01N021/17
US09351646 | B2 | 2016-05-31 | Photoacoustic measurement apparatus and probe unit for use therewith, and endoscope | 2013-10-25 | A photoacoustic measurement apparatus which includes a diffuser plate which diffuses laser light inputted from the upstream side of an optical system, a lens system which focuses the laser light diffused by the diffuser plate, and an optical fiber cable which includes an optical fiber having a core and cladding structure and is disposed such that the laser light focused by the lens system is inputted from one end, in which the light projection unit projects the laser light guided by the optical fiber cable as the measuring light. In a photoacoustic measurement performed by guiding the measuring light using an optical fiber, high energy measuring light can be guided while inhibiting end face damage of the optical fiber. | 1. A photoacoustic measurement apparatus, comprising: a diffuser plate which diffuses measuring light inputted from the upstream side of an optical system;a lens system which focuses the measuring light diffused by the diffuser plate and is disposed such that the measuring light diffused by the diffuser is inputted from the upstream side of the lens system;an optical fiber cable which includes an optical fiber having a core and cladding structure and is disposed such that the measuring light focused by the lens system is inputted from one end;a light projection unit which projects the measuring light guided by the optical fiber cable into a subject;an electroacoustic transducer unit which detects a photoacoustic wave generated in the subject by the projection of the measuring light and converts the photoacoustic wave into an electrical signal; anda signal measurement unit which performs a signal measurement based on the electrical signal. 2. The photoacoustic measurement apparatus of claim 1, wherein: the lens system focuses the measuring light such that a minimum beam diameter of the measuring light is ⅓ to ⅔ of a core diameter of the optical fiber; andthe optical fiber cable is disposed such that the measuring light is inputted with a beam diameter of the measuring light being ⅓ to ⅔ of the core diameter of the optical fiber. 3. The photoacoustic measurement apparatus of claim 2, wherein a divergence angle φ of the measuring light when incident on the diffuser plate, a diffusion angle θ of the diffuser plate, a focal length f of the lens system, and a core diameter d of the optical fiber satisfy a formula (1) given below:d3≤2.4·f·tan((ϕ2)2+(θ2)2)≤2⁢d3.(1) 4. The photoacoustic measurement apparatus of claim 1, wherein the lens system has a focal length of 4 to 40 mm. 5. The photoacoustic measurement apparatus of claim 1, wherein: the diffuser plate is a holographic diffuser; andthe holographic diffuser has a diffusion angle of 0.2 to 2.0°. 6. The photoacoustic measurement apparatus of claim 1, wherein: the diffuser plate is a light shaping diffuser in which tiny concave lenses are arranged randomly on one surface of a substrate; andthe light shaping diffuser has a diffusion angle of 0.2 to 2.0°. 7. The photoacoustic measurement apparatus of claim 1, wherein the optical fiber has a core diameter of 100 to 600 μm. 8. The photoacoustic measurement apparatus of claim 1, wherein the apparatus comprises a holding section which integrally holds the lens system and the optical fiber cable. 9. The photoacoustic measurement apparatus of claim 8, wherein the holding section also integrally holds the diffuser plate. 10. The photoacoustic measurement apparatus of claim 1, wherein: the apparatus comprises a branch section which branches the measuring light; andthe diffuser plate, the lens system, and the optical fiber cable are disposed with respect to each optical path of the branched measuring light portions. 11. The photoacoustic measurement appa... | FUJIFILM Corporation | A61B008/08 A61B005/00 A61B001/07
US10485486 | B2 | 2019-11-26 | Clinical metric for predicting onset of cardiorespiratory deterioration in patients | 2015-11-16 | A system generates and displays a clinical metric based on continuously collected patient physiological data, wherein the clinical metric provides a predictive measure of the likelihood of the onset of a cardiorespiratory deterioration event in the patient in a predetermined time period in the future. If the clinical metric has a configured relationship with a predetermined threshold value, embodiments may generate an alarm in addition to or instead of displaying the clinical metric. The clinical metric thus allows clinical staff to take medically indicated actions to prevent or reduce the effects of the predicted deterioration. | 1. A method of predicting an onset of cardiorespiratory deterioration of a patient, comprising: collecting real time patient physiological data from a plurality of patient physiological sensors;generating a clinical metric associated with a likelihood of cardiorespiratory deterioration of the patient in a predetermined future time period in real time, based on the real time patient physiological data; andgenerating an alarm if the clinical metric has a predetermined relationship with a threshold value,wherein the patient is an infant after Stage 1 surgical palliation for hypoplastic left heart syndrome. 2. The method of claim 1, wherein the predetermined relationship is greater than or equal. 3. The method of claim 1, wherein generating a clinical metric comprises: generating the clinical metric based on the real time patient physiological data and a patient-specific baseline physiological data. 4. The method of claim 3, wherein generating the clinical metric comprises: normalizing the clinical metric to an average risk of deterioration in a patient population; andcalculating a risk index value in real time based on the real time patient physiological data and the patient-specific baseline physiological data. 5. The method of claim 1, further comprising: reporting the clinical metric in real time. 6. The method of claim 1, wherein the predetermined future time period is one to two hours in the future. 7. The method of claim 1, wherein generating a clinical metric comprises: optimizing a model based upon one or more of: heart rate;respiration rate variability;SpO2,3D ST segment variability;ST segment; andheart rate variability. 8. The method of claim 1, further comprising: collecting a patient-specific baseline data from the patient physiological sensors, andwherein generating a clinical metric comprises: generating the clinical metric at time t based upon real time patient physiological data collected at time t and patient-specific baseline data collected at a predetermined time before time t. 9. The method of claim 1, wherein the clinical metric is a clinical metric associated with a likelihood of cardiac arrest in an infant after Stage 1 surgical palliation. 10. The method of claim 1, wherein generating a clinical metric comprises: optimizing a model using logistic regression; andemploying the model to analyze the real time patient physiological data. 11. The method of claim 1, wherein the cardiorespiratory deterioration comprises one or more of: an intubation event;a shock;a respiratory arrest;a cardiac arrest; andhypoxemia. 12. A non-transitory computer readable medium, on which are stored instructions, comprising instructions that when executed cause a patient data monitoring system to: collect real time patient physiological data from a plurality of patient physiological sensors associated with a patient, wherein the patient is an infant after Stage 1 surgical palliation for hypoplastic left heart syndrome;generate a real time cl... | Baylor College of Medicine | A61B005/00 A61B005/0205 A61B005/021 A61B005/024 A61B005/0402 A61B005/1455 A61B005/08
US20160120467 | A1 | 2016-05-05 | PAIN MANAGEMENT FOR PATIENT CONTROLLED ANALGESIA AND EPIDURALS | 2014-12-31 | Methods, computer systems and computer readable media for receiving data and information from medical devices in a healthcare setting are provided. In embodiments, information is received from an infusion device. After a pain score or sedation score is received for a patient, a graphical representation that indicates a correlation between the information and the pain score or sedation score is provided. | 1. One or more computer storage media storing computer-useable instructions that, when used by one or more computing devices, cause the one or more computing devices to perform operations comprising: receiving information from an infusion device;receiving a pain score or sedation score for a patient, the pain score or sedation score being correlated with the infusion; andproviding, via a user interface, a graphical representation that indicates a correlation between the information and the pain score or sedation score. 2. The media of claim 1, wherein the information includes duration, rate, dose, or volume of an infusion. 3. The media of claim 1, wherein the information includes whether the infusion was patient administered, provider administered, or automated. 4. The media of claim 1, wherein the information includes an indication of a number of doses attempted. 5. The media of claim 1, wherein the information includes an indication of a number of doses delivered. 6. The media of claim 1, further comprising, providing, via the user interface, a type of medication associated with the infusion. 7. The media of claim 1, further comprising, receiving vital signs associated with each dose delivered via the infusion. 8. The media of claim 7, wherein the vital signs are automatically taken each time a dose is attempted or delivered via the infusion. 9. The media of claim 7, further comprising, providing, via the user interface, a graphical representation that indicates a correlation between the information and the vital signs. 10. A computer-implemented method in a clinical computing environment comprising: receiving information from an infusion device, via a first computing process, the information including duration, rate, dose, or volume of an infusion and whether the infusion was patient administered, provider administered, or automated; andreceiving, via a second computing process, a pain score or sedation score for a patient;providing, via a third computing process, a graphical representation that indicates a correlation between the information and the pain score or sedation score;wherein each computing process is performed by one or more computing devices. 11. The media of claim 10, wherein the information includes an indication of a number of doses attempted. 12. The media of claim 10, wherein the information includes an indication of a number of doses delivered. 13. The media of claim 10, further comprising, providing, via the user interface, a type of medication associated with the infusion. 14. The media of claim 10, further comprising, receiving vital signs associated with each dose delivered via the infusion. 15. The media of claim 14, wherein the vital signs are automatically taken each time a dose is attempted or delivered via the infusion. 16. The media of claim 14, further comprising, providing, via the user interface, a graphical representation that indicates a correlation between the information and the vital signs. ... |  | A61B005/00 A61M005/172
US20160128584 | A1 | 2016-05-12 | MEASURING DEVICE, METHOD OF MEASURING BLOOD PRESSURE, AND PROGRAM | 2015-10-30 | A measuring device includes a measuring part that performs an invasive blood pressure measurement using a transducer and detects that a blood pressure measured by the invasive blood pressure measurement becomes a predetermined abnormal state during the invasive blood pressure measurement, a detecting part that detects whether or not a zero-point calibration of the transducer is being carried out, and a notification control part that controls whether or not to output an alarm based on a detection state of the detecting part when the blood pressure becomes the predetermined abnormal state. | 1. A measuring device, comprising: a measuring part that performs an invasive blood pressure measurement using a transducer and detects that a blood pressure measured by the invasive blood pressure measurement becomes a predetermined abnormal state during the invasive blood pressure measurement;a detecting part that detects whether or not a zero-point calibration of the transducer is being carried out; anda notification control part that controls whether or not to output an alarm based on a detection state of the detecting part when the blood pressure becomes the predetermined abnormal state. 2. The measuring device according to claim 1, wherein the measuring part detects, as the predetermined abnormal state, a state where, during the invasive blood pressure measurement, there is no detected pulse and an averaged blood pressure value is a predetermined value or less. 3. The measuring device according to claim 1, wherein the measuring part detects, as the predetermined abnormal state, case where a state in which, during the invasive blood pressure measurement, there is no detected pulse and an averaged blood pressure value is a predetermined value or less is continued during a predetermined period of time or more. 4. The measuring device according to claim 1, wherein the notification control part suppresses outputting an alarm if the blood pressure measured by the measuring part is in the predetermined abnormal state and the zero-point calibration is being carried out, and outputs an alarm if the blood pressure is still in the predetermined abnormal state after a predetermined period of time from a predetermined timing. 5. The measuring device according to claim 4, wherein the predetermined timing is any one of an operation timing of a interface for the zero-point calibration, a timing of starting displaying a message related to that the zero-point calibration is being carried out, an ending timing of the zero-point calibration, a timing of starting displaying a message related to ending of the zero-point calibration, and a timing of ending displaying of the message related to ending of the zero-point calibration. 6. The measuring device according to claim 1, wherein the notification control part suppresses outputting an alarm if the blood pressure measured by the measuring part is in the predetermined abnormal state and the zero-point calibration is being carried out, and thereafter outputs an alarm if the blood pressure measured by the measuring part is decided as becoming again the predetermined abnormal state. 7. A method of measuring a blood pressure, comprising: performing an invasive blood pressure measurement using a transducer;detecting that a blood pressure becomes a predetermined abnormal state during the invasive blood pressure measurement;detecting whether or not a zero-point calibration of the transducer is being carried out; andcontrolling whether or not to output an alarm based on a detection state of the zero-point calibrati... |  | A61B005/0215 A61B005/00
US20160120479 | A1 | 2016-05-05 | Respiration Monitoring Method and Device with Context-Aware Event Classification | 2014-10-31 | A respiration monitoring method and device with context-aware noise detection use context supplied by neighboring events when deciding how to classify a target event. More particularly, the present invention considers similarities between a target event and neighboring events relative to event attributes and timing to inform the decision whether to classify the target event as a respiration event or a noise event. Target events classified as noise events are removed from the respiration signal or otherwise ignored when estimating respiration parameters. | 1. A respiration monitoring device, comprising: a body sensor configured to sense respiration activity and generate a respiration signal representing the respiration activity;an event identification engine configured to identify candidate respiration events in the respiration signal;an event classification engine configured to select individual target events from the candidate respiration events, compute an attribute similarity metric for each of the target events by comparing a target value of an attribute of the target event with a reference value for the attribute determined from one or more events neighboring the target event, and classify the target event as one of a noise event or a respiration event using the attribute similarity metric;a respiration parameter estimator configured to compute a respiration parameter estimate using the target events classified as respiration events; anda respiration data output interface configured to output respiration data using the respiration parameter estimate. 2. The device of claim 1, wherein the attribute similarity metric is computed as a quotient of the target value and the reference value. 3. The device of claim 1, wherein the target event is classified using a result of a comparison of the attribute similarity metric with a threshold value. 4. The device of claim 1, wherein the target event is classified using a result of a comparison of the attribute similarity metric and a timing similarity metric with a threshold value. 5. The device of claim 1, wherein the reference value is selected as a value of the attribute for one of the neighboring events which is most similar to the values of the attribute for the other neighboring events. 6. The device of claim 1, wherein the attribute is mean event amplitude. 7. The device of claim 1, wherein the attribute is event width. 8. The device of claim 1, wherein the attribute is event area. 9. The device of claim 1, wherein the timing similarity metric is computed using a period and phase interval associated with the target event. 10. The device of claim 1, wherein the body sensor is a sound sensor. 11. The device of claim 1, wherein the respiration parameter is respiration rate. 12. The device of claim 1, wherein the event classification engine is further configured to be able to classify at least one of the candidate respiration events as a noise event using a result of a comparison of a value of an attribute of the candidate respiration event with a threshold value of the attribute. 13. A respiration monitoring method for a respiration monitoring device having a body sensor, comprising: sensing, by the body sensor, respiration activity;generating, by the device, a respiration signal representing the respiration activity;identifying, by the device, candidate respiration events in the respiration signal;selecting, by the device, individual target events from the candidate respiration events;computing, by the device, an attribute similarity metr... | Sharp Laboratories of America, Inc. | A61B005/00 A61B007/00 A61B005/08
US20160128617 | A1 | 2016-05-12 | SOCIAL CUING BASED ON IN-CONTEXT OBSERVATION | 2014-11-10 | System and techniques for social cuing based on in-context observation are described herein. A plurality of context attributes about a user and a subject may be received. A period in which an interpersonal interaction between the user and the subject occurs may be determined. A social-interaction-emotional-state may be determined during the period from the plurality of context attributes using a pre-defined classification methodology. A feedback response may be generated based on the determined social-interaction-emotional-state. The feedback may be presented to the user. | 1. A system for social cuing based on in-context observation, the system comprising: a context circuit-set to receive a plurality of context attributes about a user and a subject;an interpersonal interaction circuit-set to determine a period in which an interpersonal interaction between the user and the subject occurs;an emotional assessment circuit-set to determine a social-interaction-emotional-state, during the period, from the plurality of context attributes using a pre-defined classification methodology;a feedback circuit-set to generate a feedback response based on the determined social-interaction-emotional-state; anda presentation circuit-set to present the feedback response to the user. 2. The system of claim 1, wherein the plurality of context attributes are derived from sensor data from a sensor of a device, the sensor positioned so as to observe a physical area of the interpersonal interaction. 3. The system of claim 1, including an access control mechanism to control access to a set of at least one device-present context attribute, the set being a subset of the plurality of context attributes, based on an access control list controlled by a provider and identification of a potential recipient, the set having been granted access to the user as the potential recipient. 4. The system of claim 3, wherein the access control list includes role based access for a portion of the set, roles in the role based access corresponding to positions in interpersonal communications. 5. The system of claim 4, wherein the positions in interpersonal communications include a social status of communicating persons. 6. The system of claim 1, wherein to determine a social-interaction-emotional-state, at least one of the plurality of context attributes is compared to a normalized value corresponding to the at least one of the plurality of context attributes to determine a variance, the variance used to determine the social-interaction-emotional-state from the pre-defined classification methodology. 7. The system of claim 1, including a recommendation circuit-set to: receive observations of a second interpersonal interaction including attributes of parties to the second interpersonal interaction and a second social-interaction-emotional-state, the attributes including a second plurality of context attributes including a set of interest attributes;assess suitability of the second interpersonal interaction for the user based on both the second social-interaction-emotional-state and the set of interest attributes; andprovide the user with a recommendation as to whether to join the second interpersonal interaction based on the assessed suitability of the second interpersonal interaction. 8. The system of claim 7, wherein the recommendation includes directions to bring the user to the second interpersonal interaction. 9. A method for social cuing based on in-context observation comprising: receiving, via a transceiver, a plurality of context attributes abou... |  | A61B005/16 G09B019/00
US20160128674 | A1 | 2016-05-12 | ULTRASOUND IMAGING APPARATUS AND CONTROL METHOD THEREOF | 2015-04-15 | Disclosed herein is an ultrasound imaging apparatus including: an image creator configured to create an ultrasound elastic image that represents a degree of elasticity for a region of interest including a predetermined point of interest of an object; and a display configured to display the ultrasound elastic image together with an ultrasound image for the object and a high-resolution medical image matching with the ultrasound image. | 1. An ultrasound imaging apparatus comprising: an image creator configured to create an ultrasound elastic image that represents a degree of elasticity for a region of interest including a predetermined point of interest of an object; anda display configured to display the ultrasound elastic image together with an ultrasound image for the object and a high-resolution medical image matching with the ultrasound image. 2. The ultrasound imaging apparatus according to claim 1, further comprising: a connector connected to an ultrasound probe; anda controller configured to determine an irradiation focal point of the ultrasound probe in order to acquire the ultrasound elastic image. 3. The ultrasound imaging apparatus according to claim 2, wherein the controller decides an offset between the irradiation focal point and the point of interest in order to determine the irradiation focal point. 4. The ultrasound imaging apparatus according to claim 3, wherein the controller decides the offset between the irradiation focal point and the point of interest, based on an amount of change in axis direction of shear waves irradiated from the irradiation focal point. 5. The ultrasound imaging apparatus according to claim 3, wherein the controller determines, as the irradiation focal point, a point located at the same depth as the point of interest and spaced by the offset in a lateral direction from the point of interest. 6. The ultrasound imaging apparatus according to claim 2, wherein the controller determines an irradiation width of the ultrasound probe according to a predetermined number of irradiation focal points. 7. The ultrasound imaging apparatus according to claim 2, further comprising an ultrasound probe configured to irradiate a focus beam to the irradiation focal point so as to transfer shear waves from the irradiation focal point to the point of interest. 8. The ultrasound imaging apparatus according to claim 1, wherein the high-resolution medical image is at least one image of a magnetic resonance image and an X-ray image. 9. The ultrasound imaging apparatus according to claim 1, further comprising: a connector connected to an ultrasound probe;a storage in which at least one high-resolution medical image is stored; anda controller configured to match the ultrasound image with the high-resolution medical image. 10. The ultrasound imaging apparatus according to claim 9, further comprising an input configured to receive a users input of inputting at least one of coordinate information and a feature point of the ultrasound image and at least one of coordinate information and a feature point of the high-resolution medical image, wherein the controller matches the ultrasound image with the high-resolution medical image, based on the users input. 11. The ultrasound imaging apparatus according to claim 9, wherein the connector receives an ultrasound signal and a location signal of the ultrasound probe, from the ultrasound probe, the image creator c... |  | A61B008/08 A61B006/00 A61B005/055
US09351088 | B2 | 2016-05-24 | Evaluation of sound quality and speech intelligibility from neurograms | 2014-06-16 | An approach is described for evaluating a selected hearing stimulation strategy in a hearing impaired patient. A healthy neurogram represents a normal auditory neural response to a selected sound input by a healthy auditory system. An impaired neurogram represents an impaired auditory neural response to the selected sound input by a hearing impaired patient auditory system using the hearing stimulation strategy. The impaired neurogram is compared to the healthy neurogram based on: i. partitioning the neurograms into corresponding temporal and frequency domain tiles, ii. determining local variability between corresponding individual tiles by comparison of selected tile features, iii. determining global variability between the neurograms by weighted averaging of the individual tile comparisons, and iv. developing from the local and global variabilities a perception estimate representing effectiveness of the selected hearing stimulation strategy in approaching normal healthy hearing for the implanted patient. | 1. A method for evaluating a selected hearing stimulation strategy in a hearing impaired patient, the method comprising: accessing a healthy neurogram representation of a normal auditory neural response to a selected sound input by a healthy auditory system;accessing an impaired neurogram representation of an impaired auditory neural response to the selected sound input by a hearing impaired patient auditory system using the hearing stimulation strategy; andcomparing the impaired neurogram to the healthy neurogram based on: i. partitioning the neurograms into corresponding temporal and frequency domain tiles,ii. determining local variability between corresponding individual tiles by comparison of selected tile features,iii. determining global variability between the neurograms by weighted averaging of the comparison of the selected tile features, andiv. developing from the local and global variabilities a perception estimate representing effectiveness of the selected hearing stimulation strategy in approaching normal healthy hearing for the implanted patient. 2. A method according to claim 1, wherein the perception estimate further is based on a computational model of a hearing prosthesis using the selected hearing stimulation strategy. 3. A method according to claim 1, wherein the local variability is determined based on a Levenshtein distance or an approximation thereof between the selected tile features. 4. A method according to claim 3, wherein a Needleman-Wunsch algorithm or an approximation thereof is used for evaluating or estimating the Levenshtein distance. 5. A method according to claim 4, wherein the Needleman-Wunsch algorithm is applied to a matrix containing neurogram instantaneous firing rate. 6. A method according to claim 1, wherein a weighted partition sum is used for determining global variability. 7. A method according to claim 1, wherein a rate-place code with temporal resolution greater than 300 milliseconds is used for comparing the neurograms. 8. A method according to claim 1, wherein temporal fine structure with temporal resolution higher than 1 millisecond is used for comparing the neurograms. 9. A non-transitory computer readable storage medium having stored therein a plurality of instructions configured for causing one or more processors to execute the steps of: accessing a healthy neurogram representation of a normal auditory neural response to a selected sound input by a healthy auditory system;accessing an impaired neurogram representation of an impaired auditory neural response to the selected sound input by a hearing impaired patient auditory system using the hearing stimulation strategy; andcomparing the impaired neurogram to the healthy neurogram including: i. partitioning the neurograms into corresponding temporal and frequency domain tiles,ii. determining local variability between corresponding individual tiles by comparison of selected tile features,iii. determining global variability between the neuro... | MED-EL Elektromedizinische Geraete GmbH Technische Universität München | H04R029/00 H04R025/00 A61N001/36 A61B005/048 A61B005/0484 A61B005/12 A61N001/05 A61B005/00
US09867566 | B2 | 2018-01-16 | Technologies for white matter hyperintensity quantification | 2015-11-20 | A method comprises selecting, via a computer, an intensity threshold value. The method also comprises defining, via the computer, a plurality of hyperintensities on imaging data based on the intensity threshold value. The method further comprises extracting, via the computer, a plurality of voxels from the imaging data based on the defining. The method additionally comprises determining, via the computer, a total volume of the voxels based on the extracting. The method also comprises determining, via the computer, a regional distribution of WMH based on an anatomical atlas and the total volume. | 1. A method comprising: selecting, via a computer, an intensity threshold value;defining, via the computer, a plurality of hyperintensities on imaging data based on the intensity threshold value;extracting, via the computer, a plurality of voxels from the imaging data based on the defining;determining, via the computer, a total volume of the voxels based on the extracting; anddetermining, via the computer, a regional distribution of WMH based on an anatomical atlas and the total volume, wherein the intensity threshold value is selected from a range of specificity-to-sensitivity intensity threshold values. 2. The method of claim 1, wherein the selecting is based on an image analysis based on at least one of image brightness information, image historic information, image degradation information, or imaging machine information. 3. The method of claim 1, further comprising: analyzing, via the computer, the imaging data for at least one of hyperintensity under-labeling or hyperintensity over-labeling based on at least one of a voxel location or a voxel light intensity. 4. The method of claim 3, further comprising: editing, via the computer, the imaging data based on the analyzing. 5. A system comprising: a hardware processor;a memory operably coupled to the hardware processor, wherein the memory storing a set of instructions for execution via the hardware processor, wherein the set of instructions instructs the hardware processor to perform a method comprising: selecting an intensity threshold value;defining a plurality of hyperintensities on imaging data based on the intensity threshold value;extracting a plurality of voxels from the imaging data based on the defining;determining a total volume of the voxels based on the extracting; anddetermining a regional distribution of WMH based on an anatomical atlas and the total volume, wherein the intensity threshold value is selected from a range of specificity-to-sensitivity intensity threshold values. 6. The system of claim 5, wherein the selecting is based on image analysis based on at least one of image brightness information, image historic information, image degradation information, or imaging machine information. 7. The system of claim 5, wherein the method further comprising: analyzing the imaging data for at least one of hyperintensity under-labeling or hyperintensity over-labeling based on at least one of a voxel location or a voxel light intensity. 8. The system of claim 7, wherein the method further comprising: editing the imaging data based on the analyzing. 9. A non-transitory computer-readable storage medium storing a set of instructions for execution via a processing circuit to implement a method comprising: selecting, via a computer, an intensity threshold value;defining, via the computer, a plurality of hyperintensities on imaging data based on the intensity threshold value;extracting, via the computer, a plurality of voxels from the imaging data based on the defining;determining, ... | The Trustees of Columbia University in the City of New York | A61B005/00 A61B005/055 G06K009/46 G06T015/08 G06T007/00 G06T007/11 G06T007/136 A61B005/16
US10842411 | B2 | 2020-11-24 | Non-invasive estimation of intra-pleural pressure and/or computation of work of breathing based on a non-invasive estimation of intra-pleural pressure | 2014-06-19 | A method includes obtaining a first physiological parameter indicative of a non-invasively measured airway pressure of a subject, obtaining a second physiological parameter indicative of a non-invasively measured air flow into the lungs of the subject, and estimating a third physiological parameter indicative of an intra-pleural pressure of the subject based on the first and second physiological parameters and generating a signal indicative thereof. An other method includes obtaining a first physiological parameter indicative of a non-invasively estimated intra-pleural pressure of a subject, determining a second physiological parameter indicative of a lung volume of the subject that is based on a third physiological parameter indicative of a non-invasively measured air flow into the lungs of the subject, and determining a work of breathing based on the first and second physiological parameters and generating a signal indicative thereof. | 1. A method of estimating work of breathing in a subject receiving mechanical ventilation via a ventilator that includes (a) a sensor to monitor a physiological parameter and (b) a controller, comprising: ventilating adapted to provide mechanical ventilation to the subject;determining, in the controller, a first physiological parameter indicative of a lung volume of a subject that is based on a second physiological parameter indicative of a non-invasively measured air flow into the lungs of the subject monitored via the sensor;obtaining a non-invasively measured airway pressure of the subject;obtaining a third physiological parameter indicative of a non-invasively estimated intra-pleural pressure of the subject by non-invasively estimating the intra-pleural pressure of the subject based on the non-invasively measured airway pressure of the subject and the first physiological parameter indicative of the lung volume;determining, in the controller, a work of breathing based on the first and third physiological parameters and generating a signal indicative thereof, wherein estimating the intra-pleural pressure is done based on a lung mechanics model, which is a function of the non-invasively measured airway pressure, the air flow into the lungs, the determined lung volume, a lung resistance, a lung compliance, and a predetermined constant; andcontrolling operation of the ventilator to ventilate the subject based on the work of breathing. 2. The method of claim 1, further comprising: generating a Campbell diagram based on the first and third physiological parameters; anddetermining the work of breathing based on the Campbell diagram. 3. The method of claim 2, further comprising: generating a chest wall compliance line;determining an area within a pressure-volume loop of the Campbell diagram and the chest wall compliance line; anddetermining the work of breathing based on the area. 4. The method of claim 1, further comprising: determining the work of breathing for each breath. 5. The method of claim 1, further comprising: determining the lung volume by integrating the air flow into the lungs. 6. The method of claim 1, further comprising: estimating the intra-pleural pressure by fitting the lung mechanics model to the non-invasively measured airway pressure, the air flow into the lungs, the determined lung volume, and the predetermined constant. 7. The method of claim 6, further comprising: estimating the lung resistance and the lung compliance by fitting the lung mechanics model to the non-invasively measured airway pressure, the air flow into the lungs, the determined lung volume, and the predetermined constant. 8. The method of claim 7, further comprising: estimating the airway pressure based on the lung mechanics model using the estimated intra-pleural pressure, the measured air flow into the lungs, the determined lung volume, the estimated lung resistance, the estimated lung compliance, and the predetermined constant;determining a difference... | Koninklijke Philips N.V. | A61B005/087 A61B005/085 A61M016/00 A61B005/03 A61B005/091 A61B005/00
US10617327 | B2 | 2020-04-14 | Exit detection system with compensation | 2015-10-20 | A person support apparatus, such as a bed, stretcher, cot, recliner, or the like, includes an exit detection system having a plurality of force sensors that support the weight of an occupant positioned on a support surface. The force sensors are part of an exit detection system that issues an alarm when the occupant exits, or is about to exit, the person support apparatus. The distribution of weight applied to the force sensors is used to determine if the occupant is about to exit the person support apparatus. Compensation is made to the exit detection system for changes in the weight distribution that are not caused by movement of the occupant. Such changes may be due to siderail movement, support surface pivoting, and/or movement of other components of the person support apparatus. | 1. A person support apparatus comprising: a support deck adapted to support thereon an occupant of the person support apparatus, the support deck comprising a first section and a second section, the first section being pivotable with respect to the second section about a generally horizontal pivot axis;a plurality of force sensors, each of the force sensors adapted to detect a force, wherein a sum of the forces detected by the plurality of force sensors includes a weight of the support deck and a weight of the occupant when the occupant is supported on the support deck;an angle sensor adapted to determine an angle of the first section with respect to a reference; anda controller adapted to issue an exit alert if a distribution of the forces among the plurality of force sensors changes by more than a threshold, the controller taking into account the angle when determining if the distribution of the forces among the plurality of force sensors changes by more than the threshold. 2. The person support apparatus of claim 1 wherein the controller is further adapted to determine a center of gravity of the forces and the threshold corresponds to a distance of the center of gravity of the forces from a boundary of a defined area. 3. The person support apparatus of claim 2 wherein the controller detects a change in the center of gravity of the forces and estimates, using the angle, what portion of the change is due to movement of the occupant relative to the support deck and what portion of the change is due to pivoting of the first section. 4. The person support apparatus of claim 2 further comprising: a litter frame adapted to tilt with respect to a horizontal plane, the litter frame supporting the support deck;a litter frame angle sensor adapted to determine an angle of the litter frame with respect to horizontal; andwherein the controller is further adapted to detect a change in the center of gravity of the forces and estimate, using the litter frame angle, what portion of the change is due to movement of the occupant relative to the support deck and what portion of the change is due to tilting of the litter frame. 5. The person support apparatus of claim 2 further comprising: a siderail movable between an up position and a down position;a siderail sensor adapted to detect whether the siderail is in the up position or the down position; andwherein the controller is further adapted to detect a change in the center of gravity of the forces and estimate, using the siderail position, what portion of the change is due to movement of the occupant relative to the support deck and what portion of the change is due to movement of the siderail. 6. The person support apparatus of claim 5 wherein the controller takes into account, when estimating what portion of the change is due to movement of the siderail, a weight of the siderail. 7. The person support apparatus of claim 2 further comprising: a first lift adapted to change a height of a first end of the s... | Stryker Corporation | A61B005/11 A61G007/015 A61G007/012 A61G007/018 A61G007/005 A61B005/00
US10799143 | B2 | 2020-10-13 | Split bridge for PET-MR scanner | 2015-12-22 | A PET-MRI apparatus includes a patient support assembly having a cradle to accommodate a subject, a carriage assembly to attach to the cradle and cause the cradle to selectively travel, and a rear bridge to receive the cradle and carriage. The patient support assembly also includes a front bridge to receive the cradle and carriage assembly and provide for translation therealong, with the front bridge including a first section and a second section that is spaced apart from the first section to form a gap in the front bridge between the first section and the second section in an area adjacent a PET detector array of the PET-MRI apparatus. The front and rear bridges include an anti-tip rail assembly to interact with the cradle and the carriage assembly so as to prevent tipping thereof when traversing over the gap between the first and second sections of the front bridge. | 1. A patient support assembly for use in a combination PET-MRI imaging system having a single bore, the patient support assembly comprising: a cradle configured to accommodate a subject of interest;a carriage assembly attachable to the cradle and configured to drive the cradle through an imaging volume of the PET-MRI imaging system; a bridge assembly configured to accommodate the cradle and the carriage assembly and provide for translation therealong, the bridge assembly including: a rear bridge;a front bridge comprising a first section and a second section, wherein the first and second sections are distinct from one another and define a gap, the gap positioned within a PET imaging volume defined by a PET detector array of the PET-MRI imaging system;wherein the gap is sized to reduce degradation of PET image quality; andan anti-tip assembly extending from the first section to the second section and extending over the gap, wherein the anti-tip assembly is configured to interact with the cradle and the carriage assembly so as to prevent tipping thereof when traversing over the gap between the first and second sections of the front bridge. 2. The patient support assembly of claim 1 wherein the cradle and the carriage assembly each include slide blades formed thereon configured to interact with the anti-tip rail assembly to prevent tipping. 3. The patient support assembly of claim 2 wherein the anti-tip rail assembly comprises a pair of anti-tip guide rails formed on opposing sides of the front bridge, each of the pair of anti-tip guide rails comprising a groove and upper flange, with the upper flange configured to apply a downward pressure onto the slide blades as the carriage assembly and cradle traverse over the gap between the first and second sections of the front bridge. 4. The patient support assembly of claim 3 wherein the pair of anti-tip guide rails extend along a length of the rear bridge and the front bridge, with the pair of anti-tip guide rails extending over the gap between the first and second sections of the front bridge so as to prevent tipping of the carriage assembly as it traverses over the gap. 5. The patient support assembly of claim 1 wherein the bridge assembly further comprise a plurality of roller bearings positioned along a length thereof and configured to support the cradle as it is translated through the imaging volume of the PET-MRI imaging system. 6. The patient support assembly of claim 5 wherein the front bridge includes a coil positioning area formed therein configured to house a coil matrix array (CMA) of the PET-MRI imaging system, the coil positioning area including a portion of the plurality of roller bearings therein to support the cradle as it passes over the coil positioning area. 7. The patient support assembly of claim 1 further comprising a set of mounting blocks configured to provide for mounting of the first and second sections of the front bridge to the a PET tube of the PET-MRI imaging system, wh... | GE Precision Healtcare LLC | A61B005/055 G01R033/48 A61B005/00 A61B006/03 A61B006/04 A61B006/00
US09984277 | B2 | 2018-05-29 | Systems, apparatus, and methods for analyzing blood cell dynamics | 2015-11-24 | Systems, apparatus, and methods for non-invasively detecting a white blood cell (WBC) event from in vivo capillary data may include obtaining a plurality of images, followed by specifying two-dimensional (2D) coordinates corresponding to internal contour and external contour of a capillary visible in the images. The two sets of 2D coordinates may be resampled to generate two sets of resampled coordinates, between which intermediate curves, including a middle curve, are generated. Curvilinear distances may be defined based on the middle curve. Intensity values, each of which corresponds to one of the images, one of the intermediate curves, and one of the curvilinear distances, may be extracted and transformed to the Radon domain. A plurality of maxima locations in the Radon domain corresponding to a flow trajectory inside the capillary may be identified. Detection of a visual gap in the flow trajectory inside the capillary indicates a WBC event. | 1. A method for detecting a white blood cell (WBC) event from in vivo capillary data associated with a subject, the method comprising: non-invasively obtaining the capillary data, the capillary data comprising a plurality of images captured over a first time period;specifying a first set of two-dimensional (2D) coordinates associated with the plurality of images, the first set of 2D coordinates corresponding to a plurality of internal contour points of a capillary visible in the plurality of images;specifying a second set of 2D coordinates associated with the plurality of images, the second set of 2D coordinates corresponding to a plurality of external contour points of the capillary;interpolating the first set of 2D coordinates to generate a first set of resampled coordinates;interpolating the second set of 2D coordinates to generate a second set of resampled coordinates;generating a plurality of intermediate curves based on the first set of resampled coordinates and the second set of resampled coordinates, the plurality of intermediate curves including a middle curve;defining a plurality of curvilinear distances based on the middle curve;extracting a plurality of intensity values from the plurality of images, each extracted intensity value corresponding to one of the plurality of images, one of the plurality of intermediate curves, and one of the plurality of curvilinear distances;transforming the plurality of extracted intensity values to a Radon domain;identifying a plurality of maxima locations corresponding to a flow trajectory inside the capillary; anddetecting at least one visual gap in the flow trajectory inside the capillary, the at least one visual gap indicating the WBC event. 2. The method of claim 1, wherein the plurality of images are obtained from nailfold capillaroscopy video. 3. The method of claim 1, wherein the plurality of images are obtained from video with a frame rate of 10 frames per second to 25 frames per second. 4. The method of claim 1, wherein the plurality of images are obtained from video with a camera pixel size of 0.5 μm to 5 μm. 5. The method of claim 1, wherein the plurality of images are a three-dimensional image stack, the stack comprising: a horizontal frame size of 640 pixels to 2560 pixels;a vertical frame size of 480 pixels to 1280 pixels; anda plurality of frames from 50 frames to 1000 frames. 6. The method of claim 1, wherein the plurality of images are single-channel images. 7. The method of claim 1, wherein: the plurality of images includes at least one multi-channel image; andobtaining the capillary data comprises converting the at least one multi-channel image to a single-channel image. 8. The method of claim 7, wherein converting the at least one multi-channel image to the single-channel image comprises: averaging each channel pointwise to generate a plurality of average intensity values corresponding to each two-dimensional location in the at least one multi-channel image; andnormalizing th... | Massachusetts Institute of Technology | G06K009/00 G06T001/20 G06T007/00 G06T005/50 A61B005/00 G06K009/52 G06K009/62 A61B005/145 G06T007/33 G06T007/12 G06T007/136 G01N015/14
US10080872 | B2 | 2018-09-25 | System and method for FFR guidewire recovery | 2014-11-04 | A kit for coronary treatment comprising a pressure sensing guidewire that includes a pressure sensor for measuring pressure within a patient's vasculature. A sheath comprising a tube having a proximal end and a distal end and defining an internal lumen sized for receiving the pressure sensing guidewire; a hemostasis valve positioned over the proximal end of the tube; a cylindrical element attached to the distal end of the tube, the cylindrical element having an internal lumen sized for receiving a conventional guidewire; wherein the pressure sensing guidewire is configured to be slideably insertable into the tube from the proximal end of the tube; and further wherein, the tube is configured to permit the pressure sensing guidewire, including the pressure sensor, to freely slide distally and proximally within the tube at the election of a user. | 1. A method for assessing a lesion in a patient's vasculature, comprising: inserting into the vasculature of the patient a pressure sensing guidewire that includes a pressure sensor;withdrawing the pressure sensing guidewire from the patient's vasculature;inserting the pressure sensing guidewire into a proximal end of a tube via a hemostasis valve attached to the tube;inserting a conventional guidewire having a proximal end and a distal end into the vasculature of the patient until the distal end of the conventional guidewire is located at a desired location in relation to the lesion;inserting the proximal end of the conventional guidewire into a distal end of a lumen of a cylindrical element that is attached to the tube;advancing the cylindrical element distally over the conventional guidewire, along with the pressure sensing guidewire inside the tube, through the patient's vasculature until the pressure sensor reaches a desired location with respect to the lesion;measuring pressure in the vasculature using the pressure sensor. | Abbott Cardiovascular Systems Inc. | A61B005/0215 A61M025/01 A61M025/09 A61B005/00 A61B005/026
US09662026 | B2 | 2017-05-30 | Three-dimensional electrode and a biological probe comprising the same | 2015-02-13 | The present invention provides a three-dimensional electrode having high cell affinity and capacitive coupling, comprising a pillar portion and a spherical portion, wherein the diameter of the spherical portion is larger than that of the pillar portion, and the carbon nanotubes are coated on the spherical portion, and pillar portion and the spherical portion are made of material selected from metal materials. The present invention may be used for developing biological probes having high cell affinity and capacitive coupling so as to provide high accuracy for measurement of neural cells or electrocardiograms and prevent from distortion. | 1. A three-dimensional electrode with cell affinity and capacitive coupling, comprising: a pillar portion and a spherical portion connected to each other, wherein a radius of the spherical portion is greater than a radius of the pillar portion, carbon nanotubes are formed on the the spherical portion and the pillar portion and the spherical portion are made of a metal material. 2. The three-dimensional electrode with cell affinity and capacitive coupling according to claim 1, wherein the pillar portion and the spherical portion are made of gold, platinum or titanium. 3. The three-dimensional electrode with cell affinity and capacitive coupling according to claim 1, wherein the radius of the spherical portion ranges from about 0.1 μm to about 100 μm. 4. The three-dimensional electrode with cell affinity and capacitive coupling according to claim 1, wherein a height of the pillar portion ranges from about 0.1 μm to about 100 μm. 5. The three-dimensional electrode with cell affinity and capacitive coupling according to claim 1, having an impedance less than 10 Ω/mm2. 6. The three-dimensional electrode with cell affinity and capacitive coupling according to claim 1, having a capacitance more than 10 mF/cm2. 7. The three-dimensional electrode with cell affinity and capacitive coupling according to claim 1, wherein the pillar portion further comprises an insulating layer configured to at least cover partially the surface of the pillar portion. 8. The three-dimensional electrode with cell affinity and capacitive coupling according to claim 7, wherein the carbon nanotubes are formed on a surface of the pillar portion which is not covered by the insulating layer. 9. The three-dimensional electrode with cell affinity and capacitive coupling according to claim 1, wherein surfaces of the carbon nanotubes are modified with hydroxyl, carboxyl or amino group. 10. A biological probe, comprising: a base;an output contact disposed on the base;a three-dimensional electrode array disposed on the base, the three-dimensional electrode array being composed of a three-dimensional electrode with cell affinity and capacitive coupling; andan interconnect conductive layer disposed on the base and electrically connecting the output contact and the three-dimensional electrode array;wherein the three-dimensional electrode with cell affinity and capacitive coupling comprising:a pillar portion and a spherical portion connected to each other, wherein a radius of the spherical portion is greater than a radius of the pillar portion, carbon nanotubes are formed on the spherical portion and the pillar portion and the spherical portion are made of a metal material. 11. The biological probe according to claim 10, wherein the pillar portion and the spherical portion are made of gold, platimum, or titanium. 12. The biological probe according to claim 10, wherein the radius of the spherical portion ranges from about 0.1 μm to about 100 μm. 13. The biological probe according to... | National Tsing Hua University | A61B005/04 A61B005/0408 B82Y030/00 B82Y015/00
US09326729 | B2 | 2016-05-03 | Electrode catheter and method for manufacturing the same | 2013-03-26 | An electrode catheter is provided with a catheter shaft (10), an operating handle (20) provided with a connector (70), a tip electrode (31), ring-shaped electrodes (32 to 34), lead wires (41 to 44) connected to the tip electrode (31) and the ring-shaped electrodes (32 to 34), respectively, and a pull wire (50) fixed to the tip electrode (31). The catheter shaft (10) is constituted of a shaft proximal end portion (11) formed of a metal tube having a spiral slit (115) formed in a tip portion, a shaft distal end portion (12) formed of a resin tube of multi-lumen structure, and a resin covering layer (13) covering outer peripheries of the shaft proximal end portion (11) and a rear end portion of the shaft distal end portion (12). The lead wires (41 to 44) and the pull wire (50) extend through different lumens of the resin tube constituting the shaft distal end portion (12). This electrode catheter can exhibit good kink resistance, torque transmissibility and pushability in the entire shaft. | 1. An electrode catheter, comprising: a catheter shaft;an operating handle connected to a base end side of the catheter shaft;a connector provided in the operating handle;a tip electrode attached to a tip of the catheter shaft;a lead wire having its tip connected to the tip electrode, extending along an axial direction inside the catheter shaft, and having its rear end connected to the connector; anda wire having its tip fixed to the tip electrode or a tip portion of the catheter shaft, extending along the axial direction inside the catheter shaft, and having its rear end fixed to the operating handle or a base end portion of the catheter shaft,wherein the catheter shaft comprises: a shaft proximal end portion formed of a metal tube having a spiral slit formed at least in a tip portion;a shaft distal end portion formed of a resin tube having a multi-lumen structure coupled to the shaft proximal end portion by inserting a rear end region of its rear end portion into a tip region of the tip portion of the shaft proximal end portion; anda resin covering layer covering outer peripheries of the shaft proximal end portion and the rear end portion of the shaft distal end portion, andwherein the lead wire and the wire extend through different lumens of the shaft distal end portion,wherein a constituent resin of the shaft distal end portion flows into the slit of the shaft proximal end portion in a coupling portion to the shaft distal end portion,wherein the resin covering layer is formed by shrinking a heat-shrinkable resin tube into which the shaft proximal end portion and the rear end portion of the shaft distal end portion are inserted, and a melting point of a heat-shrinkable resin constituting the heat-shrinkable resin tube is higher than a melting point of the constituent resin of the shaft distal end portion, andwherein a width ratio (W1/W0) is 1.3 or more, where (W1) represents a width of the slit in the shaft proximal end portion in the coupling portion between the shaft proximal end portion and the shaft distal end portion, and (W0) represents a width of the slit of the shaft proximal end portion in a portion other than the coupling portion. 2. The electrode catheter according to claim 1, wherein a pitch of the slit formed in the shaft proximal end portion narrows sequentially or stepwise in a direction toward a tip. 3. The electrode catheter according to claim 1, wherein an outside diameter of the catheter shaft is 1.4 mm or less. 4. The electrode catheter according to claim 1, wherein an area ratio of resin constituting the tube in the resin tube forming the shaft distal end portion is 60% or more in a transverse sectional view. 5. The electrode catheter according to claim 1, wherein the rear end of the wire is capable of a pulling operation, and the tip of the catheter shaft is deflectable by the pulling operation of the rear end of the wire. 6. A method for manufacturing the electrode catheter according to claim 1, the method comprisin... | Japan Lifeline Co., Ltd. | A61B005/04 A61N001/05 A61B005/00 A61B005/042 B32B037/14 B32B038/00 A61M025/00
US09726603 | B2 | 2017-08-08 | Grating pattern element, target material measuring apparatus, and target material measuring method | 2015-07-27 | A grating pattern element includes peaks which are periodically arranged and valleys provided between the peaks, and at least one of the peaks and at least one of the valleys includes an adsorbing material on a surface thereof which adsorbs a target material. | 1. An apparatus for measuring an amount of a target material provided on a surface of an object, the apparatus comprising: a grating pattern element including a concavo-convex cross-section in which peaks and valleys are periodically arranged and at least one of the peaks and at least one of the valleys comprises an adsorbing material which adsorbs the target material provided on a surface thereof;a light source configured to emit light and irradiate the grating pattern element with the light;a measuring unit configured to measure intensities of diffraction patterns of light diffracted by the grating pattern element; andan information outputter configured to output information indicating the amount of the target material provided on the surface of the object based on the intensities of the diffraction patterns of the diffracted lightwherein the information outputter comprises a processor and a software program instructing the processor to calculate a height difference between the peaks and the valleys based on the intensities of the diffraction patterns. 2. The target material measuring apparatus of claim 1, wherein one of the peaks and of the valleys comprises a hydrophobic material provided on a surface thereof and the other one of the peaks and of the valleys comprises a hydrophilic material provided on a surface thereof. 3. The target material measuring apparatus of claim 1, wherein the measuring unit measures at least one of intensities of a plurality of constructive interference patterns included in the diffraction patterns. 4. The target material measuring apparatus of claim 3, wherein the information outputter calculates an intensity ratio between two different constructive interference patterns among the plurality of constructive interference patterns. 5. The target material measuring apparatus of claim 4, wherein one of the two different constructive interference patterns is a (2m+1)th constructive interference pattern and the other is a 2nth constructive interference pattern, wherein m and n are integers. 6. The target material measuring apparatus of claim 1, wherein a height difference between the peaks and the valleys is smaller than ¼ of a wavelength of the light emitted by the light source. 7. A target material measurement method comprising: contacting a surface of an object with a grating pattern element including a concavo-convex cross-section in which peaks and valleys are periodically arranged and at least one of each of the peaks and each of the valleys includes an adsorbing material which adsorbs the target material provided on a surface thereof;irradiating the grating pattern element with light;measuring intensities of diffraction patterns of light diffracted by the grating pattern element; andoutputting information indicating the amount of the target material provided on the surface of the object based on the intensities of the diffraction patterns of the diffracted lightwherein the outputting of the information indic... | SAMSUNG ELECTRONICS CO., LTD. | G01N021/47 A61B010/00 A61B005/00 G01N033/543
US09348005 | B2 | 2016-05-24 | Interventional or non-interventional instrument for use in an MRI apparatus | 2012-08-29 | An interventional or a non-interventional instrument like a catheter, a surgical device, a biopsy needle, a pointer, a stent or another invasive or non-invasive device, like a position marker, or a surface or local coil like a head coil is disclosed, wherein these instruments are provided with an MR-safe RF transmission line or cable (2, 3) for connecting the instrument with related RF transmit/MR receive units or other signal processing units for operating the instrument during an MR imaging or MR examination of an examination object. Basically, MR safety is obtained or increased by means of a plurality of fuses (6) which are serially connected into the transmission line or cable (2, 3). | 1. An instrument for use during magnetic resonance imaging, the instrument comprising: cable for connecting the instrument with a signal processing unit,wherein:the cable comprises a plurality of fuses that are connected in series to the cable,adjacent fuses have a distance from one another that is equal to or shorter than a quarter of an effective wavelength of an radio frequency common mode current that is induced in the cable when exposed to radio frequency excitation field of a magnetic resonance imaging apparatus, andeach fuse is selected such that it turns into an at least substantially non-conducting state when the current flowing through the fuse exceeds a threshold or when the temperature of the fuse due to the current flowing through the fuse reaches or exceeds a threshold value. 2. The instrument of claim 1, wherein the fuses are melt fuses. 3. The instrument of claim 1, wherein the fuses are integrated into the cable in the form of a local reduction of the size or area of the cross section of the cable. 4. The instrument of claim 1, wherein the fuses are realized in the form of discrete fuse elements. 5. The instrument of claim 1, wherein the fuses are realized in the form of a time-lag fuse or a slow acting fuse. 6. The instrument of claim 1, wherein the cable is a coaxial line and the fuses are realized in the form of a coaxial fuse comprising a cylindrical outer conductor and an inner central conductor. 7. The instrument of claim 6, wherein the cylindrical outer conductor is a cylindrical foil. 8. The instrument of claim 6, wherein the cylindrical outer conductor is provided by a plurality of wires which are arranged parallel to each other on a surface of a cylindrical and parallel to the longitudinal axis of the cylinder around the inner central conductor. 9. The instrument of claim 1, wherein the cable comprises at least one electrically conducting wire into which the fuses are serially connected. 10. The instrument of claim 1 in the form of a catheter, the catheter comprising a hand-piece, a casing and at least one of a tracking coil, a temperature sensor and electrocardiography electrodes within the casing, wherein the cable is guided within the casing for electrically connecting at least one of the tracking coil, temperature sensor and electrocardiography electrodes with the hand-piece. 11. The instrument of claim 10, wherein the cable is provided for connecting the hand-piece with a remote signal receiving and/or a remote signal processing unit. 12. The instrument of claim 1 is selected from the group consisting of a catheter, a surgical device, a biopsy needle, a pointer, a local coil, a surface coil, a head coil, wherein the cable is provided for connecting the instrument with a remote signal receiving and/or remote signal processing unit. 13. A magnetic resonance imaging system comprising the instrument of claim 1. 14. The instrument of claim 1, wherein the fuses are resettable fuses. 15. The instrument of cl... | Koninklijke Philips N.V. | A61B005/05 G01R033/28 G01R033/36 A61B005/01 A61B005/042 A61B005/00 A61B010/02 A61M025/01 G01R033/54 A61B005/055
US09332939 | B2 | 2016-05-10 | Detecting, quantifying, and/or classifying seizures using multimodal data | 2014-11-17 | Methods, systems, and apparatus for detecting an epileptic event, for example, a seizure in a patient using a medical device. The determination is performed by providing an autonomic signal indicative of the patient's autonomic activity; providing a neurologic signal indicative of the patient's neurological activity; and detecting an epileptic event based upon the autonomic signal and the neurologic signal. | 1. A medical device system for detecting an epileptic event based upon a patient's cardiac signal and kinetic activity, comprising: a first sensor for detecting a cardiac signal indicative of a patient's heart activity;a second sensor for detecting a kinetic signal indicative of body movement of the patient; anda medical device (MD) coupled to the patient's body and operatively coupled to the first sensor and the second sensor; the MD comprising: a kinetic signal module for receiving the kinetic signal;a cardiac signal module for receiving the cardiac signal;a cardiac index module for determining at least one cardiac index from the cardiac signal, wherein said at least one cardiac index is selected from a heart rate, a heart rhythm, an EKG morphology feature, or an EKG wave interval;a detection module for determining whether an epileptic event has occurred based on a change in the at least one cardiac index, wherein said change in said at least one cardiac index is selected from an increase in the patient's heart rate; a decrease in the patient's heart rate; or a change in said at least one heart rhythm, EKG morphology feature, or EKG wave interval; andan action module for taking a further action based on said kinetic signal, wherein the further action is selected from confirming the detecting of the epileptic event, or not confirming the detecting of the epileptic event. 2. The medical device system of claim 1, further comprising: a kinetic index module for determining at least one kinetic index from the kinetic signal, wherein the kinetic index is selected from:a body posture,a body position,a fall,a cessation of movement,a parameter derived from an accelerometer signal selected from a movement, a frequency per unit time, a time interval between movements, a movement direction in at least one plane, an acceleration, a velocity, a force,a parameter derived from an inclinometer signal, ora change in one or more of the foregoing,wherein said action module takes said further action based on the at least one kinetic index. 3. The medical device system of claim 1, wherein said action module takes said further action based upon a temporal relationship between a change in the at least one cardiac index and the kinetic signal. 4. The medical device system of claim 1, further comprising a classification module for classifying said epileptic event based upon at least one of said cardiac index, a kinetic index, a respiratory index, a skin index, a blood gases signal, a blood pH, a force index, vocal sounds, ocular activity, an EEG signal, an ECoG signal, an EMG signal, a responsiveness of the patient, an awareness of the patient, or a stress marker. 5. The medical device system of claim 4, wherein said classification module is configured to classify the event as one or more of: an unconfirmed seizure;a confirmed seizure;a non-epileptic seizure;a primarily generalized tonic-clonic seizure,a secondarily generalized tonic-clonic seizure;an absence seizure... |  | A61B005/00 A61B005/025 A61B005/11 A61B005/0205 A61B005/0476 A61B005/0245 A61B005/145 A61B003/113 A61B005/0468 A61B005/0472 A61B005/0488 A61B005/16 A61B007/00 A61N001/36 A61N001/365 A61B005/024 A61B005/053 A61B005/08
US09341587 | B2 | 2016-05-17 | Electric potential measuring apparatus | 2013-06-10 | An electric potential measuring apparatus includes a detecting unit, a generating unit, and a measuring unit. The detecting unit is provided on a surface of a measurement subject and detects a second signal including a first signal within a first frequency range from the measurement subject. The generating unit generates a fourth signal by averaging a third signal within a second frequency range outside the first frequency range. The measuring unit measures electric potential of the first signal included in the second signal detected by the detecting unit with reference to electric potential of the fourth signal generated by the generating unit. | 1. An electric potential measuring apparatus comprising: a first module; anda second module which is a module separate from the first module; whereinthe first module includes: a detecting unit configured to detect a second signal including a first signal from a measurement subject by contacting a surface of the measurement subject, the first signal being an electric potential measurement target signal within a first frequency range, which propagates through the measurement subject;a generating unit configured to generate a fourth signal by averaging a third signal within a second frequency range which is different from the first frequency range; anda measuring unit for measuring an electric potential of the first signal with reference to an electric potential of the fourth signal;the second module includes a producing unit configured to produce the third signal and supplies the third signal to the measurement subject;the first module has the detecting unit detect the second signal including the first signal and the third signal and has the generating unit generate the fourth signal by averaging the third signal included in the detected second signal;the detecting unit includes a first electrode and a second electrode that detect the second signal;the generating unit generates the fourth signal by averaging the third signal included in the second signal detected by the first electrode; andthe measuring unit measures the electric potential of the first signal included in the second signal detected by the second electrode. 2. The electric potential measuring apparatus according to claim 1, wherein the first module is provided in plurality. 3. The electric potential measuring apparatus according to claim 1, wherein the first module includes a first housing, which accommodates the generating unit and the measuring unit, and the detecting unit is provided on a surface of the first housing; and wherein the second module includes a second housing, which accommodates the producing unit, and a supply unit configured to supply the third signal produced by the producing unit to the measurement subject is provided on a surface of the second housing. 4. The electric potential measuring apparatus according to claim 3, wherein an adhesive part that is adhesively attachable to the measurement subject is provided on the surface of the first housing where the detecting unit is placed, and on the surface of the second housing where the supply unit is placed. 5. An electric potential measuring apparatus comprising: a first module including: a first detecting unit configured to detect a second signal including a first signal from a measurement subject by contacting a surface of the measurement subject, the first signal being an electric potential measurement target signal within a first frequency range, which propagates through the measurement subject;a first generating unit configured to generate a fourth signal by averaging a third signal within a second frequen... | FUJITSU LIMITED | G01R031/02 G01N027/00 G01R025/00 G01R019/00 G01R023/02 G01R029/12 A61B005/053 G01R021/133 G01R031/40
US11193480 | B2 | 2021-12-07 | Exhalation measurement device, and control method | 2014-06-16 | The exhalation measurement device of certain implementations comprise a chamber, a measurement component, a piezoelectric pump, a first learning controller, and a second learning controller. The chamber may temporarily hold exhalation. The measurement component may measure a specific component in the exhalation. The piezoelectric pump may supply the measurement component with the exhalation held in the chamber. The first learning controller may perform operational setting on the piezoelectric pump before the piezoelectric pump supplies the exhalation in the chamber to the measurement component. The second learning controller may perform operational setting on the piezoelectric pump after the piezoelectric pump has started supplying the exhalation in the chamber to the measurement component, but before the measurement component performs its measurement. | 1. An exhalation measurement device, comprising: a chamber that temporarily holds exhalation;a measurement component that measures a specific component in the exhalation;a piezoelectric pump that supplies the measurement component with the exhalation held in the chamber;a first learning controller that performs a first operational setting on the piezoelectric pump before the piezoelectric pump supplies the exhalation in the chamber to the measurement component;a second learning controller that performs a second operational setting on the piezoelectric pump after the piezoelectric pump has started supplying the exhalation in the chamber to the measurement component, but before the measurement component performs its measurement, anda flow sensor that senses a flow of gas supplied by the operation of the piezoelectric pump,wherein the first learning controller selects a first drive frequency for driving the piezoelectric pump, andwherein the second learning controller selects a second drive frequency from within a specific range that includes the first drive frequency selected by the first learning controller, and selects a drive duty ratio that is a duty ratio of a drive application voltage;the second learning controller has:a frequency change component that changes the first drive frequency at a specific frequency interval while changing the duty ratio of the drive application voltage to maintain the flow of gas at a second target flow;a duty ratio increase/decrease determination component that determines an increase or decrease in the duty ratio when the frequency is changed by the frequency change component; anda duty ratio selector that selects a smallest duty ratio representing a minimum within a range of the determined increase or decrease in the duty ratio when the frequency has been changed,the second drive frequency is the frequency at the smallest duty ratio, andthe drive duty ratio is the smallest duty ratio selected by the duty ratio selector; anda voltage duty ratio adjuster that uses the second drive frequency selected by the second learning controller to control the duty ratio of the drive application voltage that is applied to the piezoelectric pump so that a flow rate is constant. 2. The exhalation measurement device according to claim 1, wherein the first learning controller also selects the drive application voltage that is applied to the piezoelectric pump. 3. The exhalation measurement device according to claim 2, wherein the first learning controller has a resonance frequency detector that detects the first drive frequency at which a piezoelectric element of the piezoelectric pump resonates by varying the first drive frequency in a state in which a specific drive application voltage has been applied, andthe first drive frequency is a frequency of resonance. 4. The exhalation measurement device according to claim 2, comprising: a zero gas generator that generates zero gas, the zero gas being obtained by eliminating the speci... | PHC HOLDINGS CORPORATION | F04B045/047 A61B005/097 G01N033/497 A61B005/091 A61B005/087 A61B005/00
US09355573 | B2 | 2016-05-31 | Performance monitoring, apparatuses, systems, and methods | 2013-03-08 | We have disclosed systems and methods for monitoring an individual's performance during a physical activity. The systems may include a mount adapted to releasably secure the portable performance monitoring device to the individual's body during the physical activity and a main body coupled to the mount. The main body may include a satellite positioning system receiver adapted to obtain position points during the physical activity, a processor adapted to determine performance information for the individual based on position points during the physical activity, an output adapted to convey at least some of the performance information to the individual during the physical activity, and a user input adapted to be activated in response to the individual tapping. | 1. A portable performance monitoring device for monitoring an individual's performance during a physical activity, the portable performance monitoring device comprising: a mount adapted to releasably secure the portable performance monitoring device to the individual's body during the physical activity;a main body coupled to the mount, the main body comprising: a satellite positioning system receiver adapted to obtain position points during the physical activity;a processor adapted to determine performance information for the individual based on the position points during the physical activity;a display screen adapted to convey at least some of the performance information to the individual during the physical activity;a housing; anda user input adapted to be activated in response to the individual tapping the housing. 2. The portable performance monitoring device of claim 1, wherein the mount is adapted to releasably secure the portable performance monitoring device to the individual's wrist. 3. The portable performance monitoring device of claim 1, wherein the main body is releasably coupled to the mount. 4. The portable performance monitoring device of claim 1, wherein the satellite positioning system receiver comprises a Global Positioning System-compatible receiver. 5. The portable performance monitoring device of claim 1, wherein the performance information that the processor is adapted to determine comprises speed or pace information. 6. The portable performance monitoring device of claim 1, wherein the main body further comprises: an accelerometer module adapted to obtain acceleration data during the physical activity,wherein the processor is further adapted to determine the performance information further based on the acceleration data. 7. The portable performance monitoring device of claim 1, wherein the main body further comprises: a wireless personal-area network. transceiver adapted to communicate with another device that is coupled to the individual's body during the physical activity. 8. The portable performance monitoring device of claim 1, wherein tapping comprises the individual's finger striking a solid surface. 9. The portable performance monitoring device of claim 1, wherein tapping comprises a. single detected tap. 10. The portable performance monitoring device of claim 1, wherein tapping comprises a series of detected taps. 11. The portable performance monitoring device of claim 1 , wherein the user input is adapted to affect a stopwatch function of the processor in response to the tapping. 12. The portable performance monitoring device of claim 1, wherein the user input is adapted to mark the individual's completion of a lap in response to the tapping. 13. The portable performance monitoring device of claim 1, wherein the performance information conveyed to the individual during the physical activity comprises speed or pace information. 14. The portable performance monitoring device of claim 1, wherein the user... | adidas AG | A63B024/00 G09B019/00 A61B005/103 A61B005/00 A63B071/06 H04L029/06 H04W012/06 G06F003/16 G06F017/40 G01C021/34 G01C021/36 A61B005/024 A61B005/11 H04L029/08 G06F011/32 G09B005/06 H04W004/02 G09B005/02 A45F005/00 A63B022/00 A63B022/02 A63B069/00 A63B069/16 G06F019/00 H04B001/3827 H04M001/60 H04W012/02 H04W084/18
US09345548 | B2 | 2016-05-24 | Patient-specific pre-operative planning | 2010-12-20 | A method for pre-operative orthopedic planning includes obtaining only a high-resolution knee-joint scan of a patient, determining hip rotation center and ankle rotation center from anthropometric data based on personal data of the patient, and determining a mechanical axis of the knee joint based on the anthropometric data. The method also includes preparing at least a two-dimensional image model of the knee joint using the knee-joint scan and the determined mechanical axis, and preparing a pre-operative surgical plan based on the image of the knee joint. | 1. A method for pre-operative orthopedic planning comprising: creating a diseased bone database;obtaining image-free personal data of a patient including age, gender, weight and height of the patient;selecting a best-fit bone model from the diseased bone database based on the obtained image-free personal data of the patient;transforming the best-fit bone model to a customized bone model matching the obtained image-free personal data of the patient using morphing software;preparing a pre-operative surgical plan based on the customized bone model without obtaining image data of the patient; andphysically manufacturing a customized physical alignment guide based on the customized bone model. 2. The method of claim 1, wherein transforming the best-fit bone model to a customized bone model matching the personal data of the patient includes transforming the best-fit bone model to match a hip height of the patient. 3. The method of claim 1, wherein transforming the best-fit bone model to a bone model matching the personal data of the patient includes transforming the best-fit bone model to match a tibial height of the patient. 4. The method of claim 1, wherein transforming the best-fit bone model to a bone model matching the personal data of the patient includes transforming the best-fit bone model to match a hip angle of the patient. 5. The method of claim 1, wherein transforming the best-fit bone model to a bone model matching the personal data of the patient includes transforming the best-fit bone model to match a tibial angle of the patient. 6. The method of claim 1, wherein preparing a pre-operative surgical plan further comprises selecting an implant for the patient. 7. The method of claim 6, wherein the implant is selected from a customized or semi-customized or non-customized implant based on the customized bone model. 8. The method of claim 1, further comprising selecting a best-fit bone model from the diseased bone database including daily activity of the patient. 9. The method of claim 1, further comprising selecting a best-fit bone model from the diseased bone database including body mass index of the patient. 10. The method of claim 1, wherein the diseased bone database includes images and measurement of diseased and deformed bones and bone joints. 11. A method for pre-operative orthopedic planning comprising: creating a diseased bone database;obtaining image-free personal data of a patient including one or more of age, gender, disability, race, ethnicity and weight of the patient;selecting a best-fit bone model from the diseased bone database based on the personal data of the patient;transforming the best-fit bone model to a customized bone model by matching at least one of hip height, tibial height, hip angle and tibial angle of the patient using morphing software;preparing a pre-operative surgical plan based on the customized bone model without obtaining image data of the patient; andmanufacturing a physical customized alignmen... | BIOMET MANUFACTURING, LLC | A61B005/05 A61B019/00 A61B017/15 A61F002/30 A61F002/46
US09345424 | B2 | 2016-05-24 | Clinical force sensing glove | 2009-04-03 | A clinical sensing glove system to quantify force, shear, hardness, etc., measured in manual therapies is disclosed. A sensor is disposed in a clinical glove. The sensor undergoes micro-bending, macro-bending, evanescent coupling, a change in resonance, a change in polarization, a change in phase modulation, in response to pressure/force applied. The amount of micro-bending, macro-bending, evanescent coupling, change in resonance, change in polarization, and/or change in phase modulation is proportional to the intensity of the pressure/force. A clinician can quantitatively determine the amount of pressure, force, shear, hardness, rotation, etc., applied. | 1. A sensing apparatus, comprising: a glove; anda sensor disposed in the glove, the sensor comprising: a first flexible applicator;a second flexible applicator having a groove disposed therein;an optical fiber having a first end, a second end, and a mirror formed on a tip of the second end, wherein at least part of the second end is disposed in the groove of the second flexible applicator; andan elastomeric polymer disposed on the optical fiber and between the first flexible applicator and the second flexible applicator;an optical circulator having an input port, an output port, and a port optically coupled to the first end of the optical fiber, wherein the optical circulator receives an input optical signal at the input port and directs the input optical signal into the optical fiber and wherein the optical circulator receives a reflected optical signal from the optical fiber and directs the reflected optical signal to the output port;wherein the optical fiber is further adapted and configured to experience micro-bend loss in response to a force applied through the first flexible applicator such that intensities of the input optical signal and the reflected optical signal in the optical fiber are attenuated in response to the applied force, wherein the attenuation is proportional to the applied force. 2. The apparatus of claim 1, further comprising: a light source operationally coupled to the input port of the optical circulator, the light source being adapted and configured to emit the optical signal; anda light detector operationally coupled to the output port of the optical circulator, the light detector being adapted and configured to receive the reflected optical signal from the output port. 3. The apparatus of claim 2, further comprising a control channel including the light source, a second optical fiber, and a second light detector, the second optical fiber having a first end coupled to the light source and a second end coupled to the second light detector, wherein the control channel is adapted and configured to provide a reference intensity of the optical signal to measure against the attenuated intensity of the optical signal in the optical fiber. 4. The apparatus of claim 2, further comprising: a data acquisition module operationally coupled to the light detector, the data acquisition module being adapted and configured to receive the optical signal from the light detector and determine the applied force based on the received optical signal; anda display module operationally coupled to the data acquisition module, the display module being adapted and configured to graphically display a representation of the force applied to the optical fiber. 5. The apparatus of claim 1, wherein the glove is selected from at least one of a surgical glove and a clinical glove, wherein the sensor is disposed in the glove using an adhesive, and wherein the sensor is disposed in at least one of a finger portion of the glove and a palm portion of the ... | University of Washington | A61B005/103 A61B005/00 A61B005/02
US09347089 | B2 | 2016-05-24 | Therapeutic and diagnostic strategies | 2009-09-18 | The present invention encompasses the finding that microRNAs (miRNAs) regulate certain key proteins involved in DNA repair. In some embodiments, a miRNA suppresses levels and/or activity of one or more DNA repair proteins. In some such embodiments, such suppression renders cells hypersensitive to certain DNA damage agents (e.g., γ-irradiation and genotoxic drugs, among others). The present invention provides various reagents and methods associated with these findings including, among other things, strategies for treating cell proliferative disorders, certain diagnostic systems, etc. | 1. A method comprising steps of: administering to an individual who is suffering from or susceptible to a cell proliferative disorder a composition comprising an miR-24 agent in an amount sufficient to inhibit cell proliferation in combination with a DNA damage agent and/or cell proliferation disorder treatment method that produces DNA double-strand breaks,isolating cells from the individual administered the miR-24 agent and preparing Wright-stained metaphase chromosome spreads from the isolated cells or staining the isolated cells from the individual with SYBR Green and performing single cell gel electrophoresis on the isolated cells, thereby measuring DNA damage and detecting DNA damage repair to be reduced in the isolated cells, as compared to control cells, thereby identifying the individual administered the miR-24 agent to possess reduced DNA damage repair,further comprising administering to the individual so identified to possess reduced DNA damage repair the miR-24 agent and γ-irradiation or a genotoxic drug. 2. A method comprising steps of:administering to an individual who is suffering from or susceptible to a cell proliferative disorder a composition comprising an miR-24 agent in an amount sufficient to render cells of the individual hypersensitive to a DNA damage agent and/or cell proliferation disorder treatment method that produces DNA double-strand breaks, isolating cells from the individual administered the miR-24 agent and preparing Wright-stained metaphase chromosome spreads from the isolated cells or staining the isolated cells from the individual with SYBR Green and performing single cell gel electrophoresis on the isolated cells, thereby measuring DNA damage and detecting DNA damage repair to be reduced in the isolated cells, as compared to control cells, thereby identifying the individual administered the miR-24 agent to possess reduced DNA damage repair,further comprising administering to the individual so identified to possess reduced DNA damage repair the miR-24 agent and γ-irradiation or a genotoxic drug. 3. The method of claim 1, wherein the amount is an amount sufficient to suppress levels and/or activity of one or more DNA repair proteins. 4. The method of claim 2, wherein the DNA repair proteins are selected from the group consisting of photolyases, methyltransferases, AP endonucleases, exonucleases, histone variants, transcription factors that regulate expression of DNA repair genes, and DNA glycosylases, and combinations thereof. 5. The method of claim 1, wherein the step of administering comprises administering in combination with other cell proliferation disorder treatment methods and/or therapeutic agents. 6. The method of claim 2, wherein the amount is an amount sufficient to suppress levels and/or activity of one or more DNA repair proteins. 7. The method of claim 2, wherein the step of administering comprises administering in combination with other cell proliferation disorder treatment methods and/or the... | Children's Medical Center Corporation | C12Q001/68 A61B005/00 A61P035/00 C12N015/11 C12N015/113
US10506928 | B2 | 2019-12-17 | Systems, devices, and methods for identifying portions of a wound filler left at a tissue site | 2015-11-04 | Methods for treating a wound include positioning a wound filler having a coating agent containing a fluorescent material proximate the wound. The wound filler is removed after a time period. The wound is then scanned using a fluorescence scanner to determine whether a portion of the wound filler remains at the wound. In response to the portion of the wound filler remaining at the wound, the portion of the wound filler is removed. Other methods and systems are presented. | 1. A method for treating a wound, the method comprising: positioning a wound filler treated with a liquid coating agent containing fluorescent markers proximate the wound;removing the wound filler from the wound after a time period;scanning the wound using a fluorescence scanner to determine whether a residual portion of the wound filler remains at the wound by detecting the fluorescent markers on to the residual portion with the fluorescence scanner; andresponsive to detecting the residual portion of the wound filler remaining at the wound, removing the residual portion of the wound filler from the wound. 2. The method of claim 1, wherein the fluorescent markers are an ultraviolet dye and the fluorescence scanner is an ultraviolet scanner. 3. The method of claim 1, wherein the fluorescent markers are an infrared dye and the fluorescence scanner is an infrared scanner. 4. The method of claim 3, wherein the fluorescence scanner emits infrared light with wavelengths in a range of about 700 nm to 900 nm. 5. The method of claim 1, wherein the coating agent is applied to a surface of the wound filler. 6. The method of claim 1, wherein the coating agent is at least partially embedded within the wound filler. 7. The method of claim 1, wherein the wound filler comprises a polyvinyl alcohol adapted to distribute reduced pressure to the wound. 8. The method of claim 1, wherein the wound filler is a reticulated foam adapted to distribute reduced pressure to the wound. 9. The method of claim 1, further comprising: sizing the wound filler prior to positioning the wound filler proximate the wound. 10. The method of claim 1, further comprising: applying reduced pressure to the wound filler prior to removing the wound filler. 11. The method of claim 1, wherein the wound filler is a first wound filler, the method further comprising: responsive to scanning the wound and finding that the first wound filler has been completely removed, positioning a second wound filler having a coating agent proximate the wound. 12. The method of claim 1, wherein the time period is greater than forty hours. 13. The method of claim 1, wherein the step of scanning the wound comprises using a receiver unit to detect light emitted by the fluorescent markers. 14. A method of treating a tissue site on a patient with reduced pressure, the method comprising: deploying a manifold treated with a liquid coating agent containing a fluorescent dye proximate the tissue site;covering the manifold and a portion of intact epidermis of the patient with a sealing member to form a sealed space in which the manifold is disposed, wherein the sealing member has a first side and a second, tissue-facing side;fluidly coupling a reduced-pressure interface proximate to the first side of the sealing member, wherein the reduced-pressure interface is in fluid communication with the manifold and the tissue site;fluidly coupling a reduced-pressure source to the reduced-pressure interface;applying the r... | KCI Licensing, Inc. | A61B005/00 A61B005/06 A61M001/00 A61F013/00 A61L015/42 A61B006/00 A61B006/12
US11437133 | B2 | 2022-09-06 | Athletic band with removable module | 2015-11-19 | A performance monitoring system includes an athletic band or other article of apparel worn by a user, with a pocket supported by the article and defining a cavity, where the pocket has a sensor opening extending through an inner wall of the pocket and an access opening providing access to the cavity. The system also includes an electronic module received in the cavity such that the module can be inserted and removed from the cavity through the access opening. The module has a projection on an underside of the electronic module and a sensor mounted on the projection, where the projection extends through the sensor opening when the electronic module is received in the cavity and is configured to place the sensor in close proximity to skin of the user, such that the sensor is configured to sense a physiological parameter of the user. | 1. A performance monitoring system comprising: an article of apparel configured to be worn by a user and at least partially formed of a first layer and a second layer of a flexible material, the second layer joined to the first layer in a superimposed, layered configuration, the article of apparel having a pocket supported by the article of apparel and defined by the flexible material between the first and second layers, wherein the pocket has an opening providing access to the pocket, the article of apparel further comprising a housing connected to the article of apparel within the pocket and being formed of a polymer material having increased rigidity relative to the flexible material defining the pocket, wherein the housing has an inner wall, an outer wall, and side walls connecting the inner wall and the outer wall, wherein the inner wall, the outer wall, and the side walls define a cavity within the pocket, and wherein the housing has an access opening through the inner wall of the housing that provides access to the cavity and is accessible through the opening in the pocket, such that the cavity is accessible through the opening and the access opening, wherein the housing comprises a flange extending around at least a portion of the housing, wherein the flange is positioned between the first and second layers of the flexible material, and the first and second layers are bonded to the flange to secure the housing within the cavity; andan electronic module received in the cavity and engaged with the housing such that the inner wall and the outer wall each engage at least a portion of an underside and a top side of the electronic module, respectively, and the side walls engage at least a portion of a side of the electronic module, wherein the electronic module is configured to be inserted and removed from the cavity through the access opening, the electronic module comprising a projection on the underside of the electronic module and a sensor mounted on the projection, wherein the projection extends through the housing when the electronic module is received in the cavity and is configured to place the sensor in close proximity to skin of the user, such that the sensor is configured to sense a physiological parameter of the user. 2. The performance monitoring system of claim 1, wherein the pocket has a sensor opening that is separate from the opening, such that the projection extends through the sensor opening when the electronic module is received in the cavity. 3. The performance monitoring system of claim 1, wherein the electronic module comprises a casing defining the projection, a computer device contained within the casing and in communication with the sensor, and an interface defined on a top side of the casing and configured to receive user input for operation of the computer device, and wherein the inner wall, the outer wall, and the side walls of the housing engage the casing of the electronic module. 4. The performance monitoring ... | NIKE, Inc. | A41D020/00 G16H020/30 A41D001/00 A63B024/00 G08B006/00 G06F003/02 G06F003/01 G06F001/16 A61B005/00 A41H043/04 G16H040/67 G16Z099/00 A61B005/0205 A61B005/024 A61B005/11 A41D027/20 B29C065/02 B29C065/52 B29C065/78 B29C065/00 A45F005/00 G01C021/36 G16H020/40 B29K623/00 B29L029/00 B29L031/48
US20160128631 | A1 | 2016-05-12 | VAPORIZATION HEAT LOSS MEASURING APPARATUS, METABOLIC EXPENDITURE MEASURING APPARATUS AND METHOD OF MEASURING VAPORIZATION HEAT LOSS | 2015-11-02 | A vaporization heat loss measuring apparatus measures humidity and a temperature of an atmosphere of a surrounding environment using an environmental humidity measuring part and an environmental temperature measuring part. Further, the apparatus measures humidity and a temperature of an atmosphere near a surface of a measuring object using an object humidity measuring part and an object temperature measuring part. Further, the apparatus calculates a vaporization heat loss from the surface of the measuring object using the environmental humidity, the environmental temperature, the object humidity, and the object temperature. Furthermore, the apparatus calculates a non-evaporative heat loss based on a measurement result of a heat flow measuring part, and calculates a metabolic expenditure using the vaporization heat loss and the non-evaporative heat loss. | 1. A vaporization heat loss measuring apparatus comprising: a first temperature sensor that measures a first temperature as a temperature near a surface of a measuring object;a first humidity sensor that measures first humidity as humidity near the surface of the measuring object;a second temperature sensor that measures a second temperature as a temperature of an environment;a second humidity sensor that measures second humidity as humidity of the environment; anda vaporization heat loss calculation part that calculates a vaporization heat loss of the measuring object using the first temperature, the first humidity, the second temperature, and the second humidity. 2. The vaporization heat loss measuring apparatus according to claim 1, wherein the vaporization heat loss calculation part calculates first saturated water vapor pressure as saturated water vapor pressure near the surface of the measuring object based on the first temperature, calculates second saturated water vapor pressure as saturated water vapor pressure of the environment based on the second temperature, and calculates the vaporization heat loss based on a ratio between the first saturated water vapor pressure and the first humidity and a ratio between the second saturated water vapor pressure and the second humidity. 3. A vaporization heat loss measuring apparatus comprising: a first saturated water vapor pressure measuring part that measures first saturated water vapor pressure as saturated water vapor pressure near a surface of a measuring object;a second saturated water vapor pressure measuring part that measures second saturated water vapor pressure as saturated water vapor pressure of an environment; anda vaporization heat loss calculation part that calculates a vaporization heat loss of the measuring object using the first saturated water vapor pressure and the second saturated water vapor pressure. 4. A vaporization heat loss measuring apparatus comprising: a first humidity sensor that measures first humidity as humidity near a surface of a measuring object;a second humidity sensor that measures second humidity as humidity of an environment; anda vaporization heat loss calculation part that calculates a vaporization heat loss of the measuring object using the first humidity and the second humidity. 5. A metabolic expenditure measuring apparatus comprising: the vaporization heat loss measuring apparatus according to claim 1;a heat flow measuring part that measures a heat flow from the surface of the measuring object to the environment;a non-evaporative heat loss calculation part that calculates a non-evaporative heat loss from the measuring object based on the heat flow; anda metabolic expenditure calculation part that calculates a metabolic expenditure of the measuring object using the vaporization heat loss measured by the vaporization heat loss measuring apparatus and the non-evaporative heat loss. 6. A metabolic expenditure measuring apparatus comprising: the vapo... |  | A61B005/00
US20160143540 | A1 | 2016-05-26 | MULTIFREQUENCY ELECTRICAL IMPEDANCE IMAGING USING LORENTZ FIELDS | 2013-07-17 | In this invention, an approach is proposed to image the electrical impedance (conductivity and permittivity) properties of biological tissues. This approach is based on electrical current induction using ultrasound together with an applied static magnetic field. Acoustic vibrations are generated via piezoelectric transducers located on the surface of a biological body. A linear phased array piezoelectric transducer is used to form pressure distribution in human body/tissue. In the existence of a static magnetic field, the resultant (velocity) current density is sensed by a receiver coil encircling the tissue or placed near the tissue and used for reconstructing the conductivity and permittivity distribution. | 1-19. (canceled) 20. A high resolution, non-invasive imaging method to generate images of the conductivity and permittivity distribution of a specimen comprising the steps of: generating a static magnetic field with a static magnetic field generator;applying an acoustic signal wave to a specimen via a plurality of ultrasonic transducers;sensing the induced current due to the acoustic signal and the static magnetic field inside the specimen by a plurality of receiver coils;filtering, amplifying, transmitting; and converting an analog signal to a digital signal; andconstructing a plurality of reconstructed images of conductivity and permittivity distribution of the specimen. 21. The method according to claim 1, wherein the plurality of ultrasonic transducers is a single element or a linear phased array transducer. 22. The method according to claim 1, wherein the plurality of receiver coils comprises coils which are sensitive to electrical current density, at least in the acoustic propagation direction and perpendicular to the acoustic propagation direction. 23. The method according to claim 1, wherein the plurality of reconstructed images are single or multi-frequency conductivity or permittivity images. 24. An apparatus to generate images of the conductivity and permittivity distribution of a specimen comprising: a static magnetic field generator configured to generate a uniform static magnetic field distribution;an ultrasonic transducer configured to generate acoustic waves in predetermined directions; anda plurality of receiver coils configured to sense the induced currents due to acoustic waves and static magnetic field inside the specimen;wherein the apparatus means for data acquisition for filtering, amplifying, transmitting, and converting analog signals to digital signals; andmeans for construction of a plurality of reconstructed images of conductivity and permittivity. |  | A61B005/00 A61B008/00 A61B005/053
US09327070 | B2 | 2016-05-03 | Medical device therapy based on posture and timing | 2009-04-30 | This disclosure describes techniques implemented by a medical device, such as an implantable medical device (IMD). The IMD may be configured to detect a posture state of a patient, and deliver posture-responsive therapy. In particular, the IMD not only detects the posture state of a patient, but also detects timing associated with the detected posture state, such as the time of day, the day of the week, or a specific time of day associated with a specific day. In this way, the posture-responsive therapy delivered by the IMD may be dependent not only on the posture state of the patient, but also on the timing associated with the posture state. The same posture state, therefore, may result in different types of therapy to the patient if the same posture occurs at different times of the day. | 1. A medical device system comprising: an implantable medical device that includes: a therapy delivery unit that delivers therapy to a patient,a posture state module that detects a posture state of the patient,a timing unit that detects timing associated with the detected posture state, anda processor; anda programmer device, wherein the programmer device: receives input from a user defining a time interval, a particular posture state, and a new set of therapy parameters for the time interval and the particular posture state, andprograms the implantable medical device based on the inputs,wherein, in response to the implantable medical device being programmed based on the inputs and by the programmer device, the processor causes the therapy delivery unit to deliver therapy to the patient according to the new set of therapy parameters upon subsequently detecting the particular posture state during the time interval. 2. The medical device system of claim 1, wherein the posture state module detects a posture transition from a first posture state to a second posture state. 3. The medical device system of claim 1, wherein the posture state module detects the posture state from a plurality of pre-defined posture states. 4. The medical device system of claim 3, wherein the plurality of posture states includes a first posture state defining a first posture and a second posture state defining the first posture in conjunction with an activity. 5. The medical device system of claim 3, wherein the plurality of posture states includes a first posture state defining a first posture and a second posture state defining a second posture. 6. The medical device system of claim 1, wherein the timing unit detects a time of day. 7. The medical device system of claim 1, wherein the timing unit detects a day of the week and a time of the day and the processor defines different therapies for similar times of different days for at least one posture state of a plurality of posture states. 8. The medical device system of claim 1, wherein the timing unit detects a time interval and the processor defines different therapies for different time intervals of one or more days. 9. The medical device system of claim 1, wherein the medical device system comprises one or more of: an implantable stimulation device, andan implantable drug pump. 10. The medical device system of claim 1, wherein the medical device system includes a memory that stores a lookup table that maps a set of posture states and timings of the postures states with different sets of therapy parameters, wherein upon detecting the particular posture state during the time interval, the medical device system performs a lookup in the lookup table. 11. The medical device system of claim 1, wherein the posture state module, the timing unit and the processor all reside within the implantable medical device. 12. The medical device system of claim 1, wherein the posture state module and the timing unit reside in th... | Medtronic, Inc. | A61M005/142 A61B005/00 A61B005/11 A61N001/36 A61N001/372
US09345431 | B2 | 2016-05-24 | Method and optical system for evaluating concentrations of components in tissue | 2013-07-17 | A method and an optical system for evaluating the spatial distribution of the concentrations of components in tissue are disclosed. The novel detecting probe included in the optical system comprises plural optical fiber sets, each optical fiber set respectively comprises at least one source optical fiber and at least one detector optical fiber, the source optical fiber connects with the multi-wavelength light source, the source optical fiber delivers light from the multi-wavelength light source onto a tested tissue; and the angle between one optical fiber set and another optical fiber set is greater than 0° and less than and not equal to 180°. Through the optical system of the present invention, the spatial distribution of the concentrations of components such as water, hemoglobin, melanin, lipid, and collagen in the tested tissue can be derived by an equation (I) defined in the present specification. | 1. A method for evaluating concentrations of components in a tissue, comprising the following steps: (A) providing an optical system for evaluating concentrations of components in a tissue, comprising:a multi-wavelength light source;a detecting probe comprising at least one optical fiber set and a diffuser, wherein each optical fiber set respectively comprises at least two source optical fibers and at least one detector optical fiber, and source-to-detector distances in each optical fiber set are different, the source optical fibers connect with the multi-wavelength light source, the source optical fibers deliver light from the multi-wavelength light source onto a tested tissue via the diffuser, and the detector optical fiber penetrates the diffuser to receive the diffused light from the source optical fibers as reflected by the tested tissue;an aligner disposed on the detecting probe for positioning a light emitting from the multi-wavelength light source on the tested tissue;a detecting device coupled to the detector optical fiber to receive reflected light from the tested tissue to obtain diffuse reflectance spectra; andan output device connected to the detecting device;(B) illuminating a tested tissue with light from the multi-wavelength light source, in which the light passes into the tested tissue;(C) detecting reflected light from the tested tissue with the detecting device coupled to the detector optical fiber to receive reflected light from the tested tissue to obtain diffuse reflectance spectra;(D) converting the diffuse reflectance spectra into absorption spectra;(E) fitting the absorption spectra with known chromophore absorption spectra to derive concentrations of components in the tested tissue by an equation (I) represented as follows: μa(λ)=CHbO2μaHbO2(λ)+CHbμaHb(λ)+CMelaninμaMelanin(λ)+CWaterμaWater(λ) +CFatμaFat(λ)+CCollagenμaCollagen(λ) (I)wherein μa(λ) is a total absorption coefficient, μaHbO2(λ) is an absorption coefficient of oxygenated hemoglobin, μaHb(λ) is an absorption coefficient of deoxygenated hemoglobin, μaMelanin(λ)is a absorption coefficient of melanin, μaWater(λ)is an absorption coefficient of water, μaFat(λ) is an absorption coefficient of lipid, μaCollagen(λ) is an absorption coefficient of collagen, CHbO2 is a concentration of oxygenated hemoglobin, CHb is a concentration of deoxygenated hemoglobin, CMelanin is a concentration of melanin, CWater is a concentration of water, CFat is a concentration of lipid, and Ccollagen is a concentration of collagen;(F) obtaining a ratio of the concentration of one of the components measured at a first direction to that measured at a second direction by an equation (II) represented as follows: R =(Cv/Cp)−1 (II)wherein R is the ratio of the concentration of the component measured at the first direction to that measured at the second direction, Cv is the concentration of the component measured at the first direction, Cp is the concentration of the component measured at the se... | National Cheng Kung University | A61B005/00 A61B005/1455
US09351713 | B2 | 2016-05-31 | Needle structure and method of performing needle biopsies | 2014-12-30 | A biopsy needle having a longitudinal channel formed within an inner conductor of a coaxial antenna is disclosed. The coaxial antenna terminates in a rigid insertion tip e.g. a ceramic cone that is insertable into biological tissue. Microwave energy (e.g. having a frequency of 1 to 100 GHz) delivered to the coaxial antenna is emitted at the insertion tip. The insertion tip may be arranged to match the impedance of the coaxial antenna to a predetermined tissue impedance. The emitted radiation can be used to measure properties of or treat (e.g., ablate) tissue at the insertion tip. Needle biopsy apparatus is also disclosed, in which a microwave energy is controllably delivered to a needle from a microwave generator. The apparatus may include an impedance tuner to dynamically match the impedance of the needle with tissue at the insertion tip. | 1. Surgical apparatus for introducing material through a needle into biological tissue, the apparatus comprising: a power source for delivering microwave energy,a needle that is insertable into biological tissue, the needle having:an elongate body terminating at a distal end with an insertion tip and at a proximal end with a microwave connector;a coaxial cable connected between the power source and the microwave connector;a longitudinal channel formed within the elongate body for transporting material to or from the tissue, wherein the longitudinal channel extends through the insertion tip and terminates at an opening in the insertion tip; and a coaxial antenna comprising an inner conductor along an inside wall of the longitudinal channel and an outer conductor coaxial with the inner conductor and separated from it by a first dielectric material,wherein the insertion tip is made from a second dielectric material and is arranged to create a matched condition between the first dielectric material and a representative tissue impedance for the microwave energy;wherein the needle is connected to the power source via a first signal path and a second signal path to receive energy from the power source, and the coaxial antenna is connected to couple the received energy into biological tissue at the insertion tip, andwherein the energy delivered on the first signal path has a first frequency and the energy delivered on the second signal path has a second frequency greater than the first frequency. 2. Surgical apparatus according to claim 1, wherein the inner conductor is a conductive tube having a diameter of about 0.5 mm and a wall thickness of about 0.01 mm. 3. Surgical apparatus according to claim 1, wherein the insertion tip comprises a rigid element attached to the first dielectric material of the coaxial antenna, the rigid element extending away from the outer conductor at a distal end of the elongate body. 4. Surgical apparatus according to claim 1, wherein the needle is insertable through any one of an endoscope, a bronchoscope and a cystoscope. 5. Surgical apparatus according to claim 1 including a controller arranged to control the amount of energy delivered into the biological tissue. 6. Surgical apparatus according to claim 1, wherein the needle is arranged to convey a material through the longitudinal channel to react with the energy delivered from the coaxial antenna. 7. Surgical apparatus according to claim 1, wherein the needle is arranged to convey a radioactive implant through the longitudinal channel. | CREO MEDICAL LIMITED | A61B010/04 A61B010/02 A61B001/018 A61B001/267 A61B001/307 A61B005/05 A61B005/06 A61M005/158 A61N005/02 A61N005/10 A61B018/18 A61B018/00 A61B018/14
US20160143541 | A1 | 2016-05-26 | System and Method For Acousto-Electromagnetic Neuroimaging | 2015-11-20 | A system and method for determining an electrical activity in a subject using ultrasound and electromagnetics are provided. In some aspects, the method includes directing ultrasound energy to a portion of a subject's anatomy using an ultrasound system, wherein the ultrasound energy inducing a perturbation to locations in the subject's anatomy, and sensing, using sensors arranged about the subject, a plurality of electromagnetic signals representing an electrical activity of the subject. The method also includes identifying electromagnetic signals that are modulated by the perturbation, and generating a report of the electrical activity using the identified electromagnetic signals. | 1. A method for determining electrical activity in a subject using ultrasound and electromagnetics, the method comprising: directing ultrasound energy to a portion of a subject's anatomy using an ultrasound system, the ultrasound energy inducing a perturbation to locations in the subject's anatomy;sensing, using sensors arranged about the subject, a plurality of electromagnetic signals representing an electrical activity of the subject;identifying electromagnetic signals that are modulated by the perturbation; andgenerating a report of the electrical activity using the identified electromagnetic signals. 2. The method of claim 1, wherein the electrical activity is neural activity. 3. The method of claim 1, wherein the ultrasound energy is a focused ultrasound energy. 4. The method of claim 1, wherein the ultrasound energy is transcranial focused ultrasound (tFUS) and wherein identifying the electromagnetic signals that are modulated by the perturbation includes demodulating from known mechanical carriers to decode intrinsic neural information. 5. The method of claim 1, wherein the electromagnetic signals include electroencephalography (“EEG”) signals. 6. The method of claim 1, wherein the electromagnetic signals include magnetoencephalography (“MEG”) signals. 7. The method of claim 6 wherein the MEG signals are acquired at room temperature. 8. The method of claim 1, the method further comprising generating a map indicative of the electrical activity in the subject. 9. The method of claim 1 further comprising frequency shifting the ultrasound energy with frequency up to at least 1 kHz to reduce noise in the electromagnetic signals. 10. The method of claim 1 further comprising performing an anatomical imaging of the subject's anatomy and correlating the report of the electrical activity with an anatomical image of the subject's anatomy. 11. The method of claim 10 wherein the anatomical imaging includes magnetic resonance imaging (MRI). 12. The method of claim 1 wherein the electrical activity of the subject includes electrical activity from the subject's brain, muscles, heart, or organs. 13. A system for determining neural activity in a subject using ultrasound and electromagnetics, the system comprising: a plurality of sensors capable of detecting electromagnetic signals associated with an electrical activity of a subject;an ultrasound system configured to direct ultrasound energy to a portion of a subject's anatomy; anda computer programmed to: i. control the ultrasound system to induce a perturbation to locations in the subject's anatomy;ii. receive electromagnetic signal data from sensors arranged about the subject;iii. identify, from the electromagnetic signal data, signals modulated by the perturbation;iv. determine spatial information related to modulated signals;v. generate a report of the electrical activity in the subject using the electromagnetic signal data and determined spatial information. 14. The system of claim 13, whe... |  | A61B005/00 A61B005/04 A61B005/0476
US20160128571 | A1 | 2016-05-12 | SYSTEMS AND METHODS FOR TRANSMITTING GLUCOSE METER READINGS | 2015-11-10 | Devices, methods, and systems for transmitting, via USSD messages, blood glucose readings from a glucometer over a GSM network to a database storing blood glucose readings. The database is configured to transmit electronic messages including the blood glucose readings received from the glucometer to one or more personal electronic devices of one or more predetermined recipients. | 1. An electronic blood glucose measurement device comprising: a blood glucose module including a microprocessor;a global system for mobile (GSM) module in communication with the blood glucose module; anda subscriber identification module (SIM) in communication with the GSM module;wherein the GSM module is configured to: receive a blood glucose reading from the blood glucose module;connect to a network via the SIM; andtransmit an Unstructured Supplementary Service Data (USSD) message including the blood glucose reading over a network to a server in communication with a database configured to store the blood glucose reading. 2. The device of claim 1, further comprising an input configured to receive a test strip, wherein the blood glucose module is configured to generate a blood glucose reading based on a blood sample applied to the test strip and to send the generated blood glucose reading to the GSM module in response to removal of the test strip from the input. 3. The device of claim 1 wherein at least one of the microprocessor of the blood glucose module and the GSM module are configured to format the at least on blood glucose reading in order to generate the USSD message. 4. The device of claim 1, wherein the SIM is one of a physical SIM card and a virtual SIM 5. The device of claim 1, wherein the USSD message is up to 80 characters. 6. The device of claim 1, wherein the at least one blood glucose reading in the USSD message includes a reading representing a last blood glucose reading successfully transmitted to the server and the blood glucose reading. 7. The device of claim 1, further comprising a non-volatile memory and an indication pointer configured to point to an address in the non-volatile memory representing a last blood glucose reading successfully transmitted to the server. 8. The device of claim 7, wherein the device is configured to receive a confirmation from the server indicating that the USSD message transmitted from the device was received by the server, and to update the indication pointer to point to an address in the non-volatile memory representing the at least one blood glucose reading successfully transmitted to the server. 9. The device of claim 7, wherein the device is configured, upon receiving an indication that the USSD message was not successfully transmitted from the device to the server, to store the blood glucose reading in the non-volatile memory, and to transmit another USSD message including the stored blood glucose reading after another blood glucose reading is generated by the blood glucose module when a connection to the server becomes available. 10. The device of claim 8, wherein the device is configured to receive the confirmation from the server in a form of an USSD message. 11. A system for transmitting blood glucose readings, the system comprising: an electronic blood glucose measurement device; anda server in communication with the electronic blood glucose measurement device over a global s... |  | A61B005/00 A61B005/15 A61B005/145
US20160140986 | A1 | 2016-05-19 | MONITORING TREATMENT COMPLIANCE USING COMBINED PERFORMANCE INDICATORS | 2015-11-12 | Methods and systems for monitoring compliance of a patient with a prescribed treatment regimen are described. Two or more patient activities (including e.g., speech activity) are analyzed to determine compliance with a treatment for a brain-related disorder. Activity is detected unobtrusively during performance of routine activities with activity sensor(s) at the patient location, patient speech is detected during use of a communication system such as a mobile telephone, and activity data is sent to a monitoring system at a monitoring location. Activity and/or speech data is processed at the patient location or monitoring location to identify activity parameters or patterns that indicate whether the patient has complied with the treatment regimen. Patient identity may be determined through biometric identification or other authentication techniques. The system may provide a report to an interested party, for example a medical care provider or insurance company, regarding patient compliance. | 1. A system comprising: at least one receiving device for use at a monitoring location for receiving at least one activity data signal and at least one audio data signal from a communication system, the at least one audio data signal including audio data representing speech from a patient at a patient location sensed with at least one audio sensor at the patient location during use of the communication system and transmitted to the monitoring location, the patient having a brain-related disorder and a prescribed treatment regimen for treating at least one aspect of the brain-related disorder, the at least one activity data signal including activity data indicative of whether the patient has complied with the prescribed treatment regimen, the activity data representing at least one first activity of the patient;signal processing circuitry configured to process the at least one activity data signal to determine, based upon the at least one first activity of the patient and at least one second activity of the patient, whether the patient has complied with the prescribed treatment regimen; andreporting circuitry configured to report a conclusion based on the determination of whether the patient has complied with the prescribed treatment regimen. 2.-4. (canceled) 5. The system of claim 1, including patient identification circuitry configured to determine a presence of the patient from at least one identity signal received by the at least one receiving device at the monitoring location from the patient location; wherein the signal processing circuitry is configured to identify patient activity data corresponding to an activity of the patient based at least in part on the determination of the presence of the patient by the patient identification circuitry. 6. The system of claim 5, wherein the identity signal includes at least a portion of the audio data signal, wherein the patient identification circuitry includes speech analysis circuitry for identifying at least a portion of the audio signal data that resembles known speech of the patient, and wherein the signal processing circuitry is configured to identify the patient activity data by identifying activity data corresponding to a portion of the audio signal that resembles known speech of the patient. 7.-24. (canceled) 25. The system of claim 1, wherein the at least one activity data signal includes non-speech activity data, and wherein the signal processing circuitry is configured to process the at least one activity data signal to identify at least one non-speech activity pattern that corresponds to unprompted performance of a non-speech activity by the patient. 26. The system of claim 1, wherein the at least one activity data signal includes non-speech activity data, and wherein the signal processing circuitry includes an activity analyzer for analyzing the activity data to determine a non-speech activity pattern; and a comparator for comparing the non-speech activity pattern represented by ... | Elwha LLC | G10L025/66 G10L017/00 A61B005/16 A61B005/0205 A61B003/11 A61B003/113 G06F019/00 A61B005/00
US09355534 | B2 | 2016-05-31 | Causing display of a notification on a wrist worn apparatus | 2013-12-12 | A method comprising determining of a notification based, at least in part, on occurrence of an event, the notification comprising a notification type associated with the event, causing display of a representation of the notification at a position on a wrist adherence portion of a wrist worn apparatus, the wrist adherence portion being distinct from a primary display, and the representation of the notification being indicative of the notification type, receiving information indicative of a notification selection input associated with the notification, and causing display of a different representation of the notification on the primary display of the wrist worn apparatus based, at least in part, on the notification selection input is disclosed. | 1. An apparatus, comprising: at least one processor;at least one memory including computer program code, the memory and the computer program code configured to, working with the processor, cause the apparatus to perform at least the following: determine a notification based on occurrence of an event, the notification comprising a notification type associated with the event;cause display of a representation of the notification at a position on a wrist adherence portion of a wrist worn apparatus, the wrist adherence portion being distinct from a primary display, and the representation of the notification being indicative of the notification type, wherein the position of the notification on the wrist adherence portion is determined to be closer or further from the primary display based on a notification status;receive information indicative of a notification selection input associated with the notification; andcause display of a different representation of the notification on the primary display of the wrist worn apparatus in response to the notification selection input, wherein the different representation comprises additional details relative to detail of the representation of the notification on the wrist adherence portion. 2. The apparatus of claim 1, wherein the position on the wrist adherence portion on which the notification is displayed is determined based on a notification status associated with the notification. 3. The apparatus of claim 1, wherein the memory includes computer program code configured to, working with the processor, cause the apparatus to perform identification of a notification attribute associated with the notification. 4. The apparatus of claim 1, wherein a color of the representation of the notification is indicative of at least one of a level of urgency, a time, an application, an importance, a count of associated notifications, a notification status, or the notification type. 5. The apparatus of claim 1, wherein the notification comprises a notification status, and wherein the memory includes computer program code configured to, working with the processor, cause the apparatus to perform determination of the position of the representation of the notification on the wrist adherence portion of the wrist worn apparatus relative to other notifications represented on the wrist adherence portion, based on the notification status. 6. The apparatus of claim 5, wherein the memory includes computer program code configured to, working with the processor, cause the apparatus to perform: determine that the notification status of the notification has changed to a different notification status;determine a different position of the representation of the notification on the wrist adherence portion of the wrist worn apparatus based on the different notification status; andcause display of the representation of the notification at the different position on the wrist adherence portion of the wrist worn apparatus. 7. The apparatus of... | Nokia Technologies Oy | G08B005/22 G06F001/16 G06F003/14 A61B005/11
US09345417 | B2 | 2016-05-24 | Method, implantable medical device, and system for determining the condition of a heart valve | 2015-07-15 | An implantable medical device has an impedance processor for determining atrial impedance data reflective of the cardiogenic impedance of an atrium of a heart during diastole and/or systole of heart cycle. Ventricular impedance data reflective of the cardiogenic impedance of a ventricle during diastole and/or systole are also determined. The determined impedance data are processed by a representation processor for estimating a diastolic and/or a systolic atrial impedance representation and a diastolic and/or a systolic ventricular impedance representation. A condition processor determines the presence of any heart valve malfunction, such as valve regurgitation and/or stenosis, of at least one heart valve based on the estimated atrial and ventricular impedance representations. | 1. A method of determining a condition of a valve of a heart in a subject, the method comprising: in computerized processing circuitry, determining, based on a first electric signal applied over at least a portion of an atrium of a first side of the heart and a first resulting electric signal measured over at least a portion of the atrium, atrial impedance data reflective of a cardiogenic impedance of the atrium during at least a sub-phase of at least one heart cycle;in computerized processing circuitry, determining, based on a second electric signal applied over at least a portion of a ventricle of the first side of the heart and a second resulting electric signal measured over at least a portion of the ventricle, ventricular impedance data reflective of a cardiogenic impedance of the ventricle during the at least a sub-phase of the at least one heart cycle;in computerized processing circuitry, estimating an atrial impedance representation based on the atrial impedance data;in computerized processing circuitry, estimating a ventricular impedance representation based on the ventricular impedance data; andin computerized processing circuitry, determining the condition of the valve based on the atrial impedance representation and the ventricular impedance representation, and emitting in electronic form, an indication of the condition of the valve from the computerized processing circuitry. 2. The method according to claim 1, wherein: the determining the atrial impedance data comprises determining, based on the first electric signal and the first resulting electric signal, atrial impedance data reflective of a cardiogenic impedance of the atrium during a diastolic phase and a systolic phase of the at least one heart cycle;determining the ventricular impedance data comprises determining, based on the second electric signal and the second resulting electric signal, ventricular impedance data reflective of a cardiogenic impedance of the ventricle during the diastolic phase and the systolic phase of the at least one heart cycle;the estimating the atrial impedance representation comprises estimating i) a diastolic, atrial impedance representation based on the atrial impedance data reflective of the cardiogenic impedance of the atrium during the diastolic phase and ii) a systolic, atrial impedance representation based on the atrial impedance data reflective of the cardiogenic impedance of the atrium during the systolic phase;the estimating the ventricular impedance representation comprises estimating i) a diastolic, ventricular impedance representation based on the ventricular impedance data reflective of the cardiogenic impedance of the ventricle during the diastolic phase and ii) a systolic, ventricular impedance representation based on the ventricular impedance data reflective of the cardiogenic impedance of the ventricle during the systolic phase; andthe determining comprises determining the condition of the valve based on i) the diastolic, atrial im... | ST. JUDE MEDICAL AB | A61B005/042 A61B005/02 A61B005/04 A61B005/0432 A61B005/00 A61B005/07 A61N001/362 A61N001/365 A61N001/368 A61N001/37
US09347817 | B2 | 2016-05-24 | Surgical object and comprehensive fluid monitoring system having capability of mobile monitoring and having highly sensitive and reliable detection of objects being placed in a container | 2015-02-24 | A surgical sponge and comprehensive fluid monitoring system and method are provided. The system includes a support ring for securely supporting a container with at least one port opening for receiving surgical sponges placed into the container. A hollow spine structure includes a channel and a rod that can move in the channel. The rod transfers weight force from the support ring to a weight force sensing device. A flexure maintains the rod vertically aligned in the channel with a gap separating the length of the rod from the channel walls. A processor receives an information signal indicative of at least one of urine outflow, suction fluid outflow, irrigation fluid intake; and intravenous fluid intake, of a surgical patient. The processor accumulates a total amount of estimated net fluid volume loss for the patient from the surgical sponge, and from other measured fluid outflow and intake. | 1. A surgical object and comprehensive fluid monitoring system for a surgical patient, comprising: a support ring for securely supporting a container with at least one port opening of the container being disposed for receiving surgical objects placed into the container through the at least one port opening with the aid of gravity;a hollow spine structure vertically oriented and comprising a channel with channel walls along the length of the hollow spine structure;a rod movably disposed in the channel within the hollow spine structure, the rod moveable along the channel;a weight force sensing device mechanically coupled to the rod, the support ring mechanically coupled to the rod, and the rod being in a moving arrangement that transfers weight force from the support ring to the weight force sensing device;a flexure mechanically coupled with the rod and the hollow spine structure, the flexure adding very little if any downward force onto the rod and thereby onto the weight force sensing device while maintaining the rod vertically aligned in the channel with a gap separating the length of the rod from the channel walls, the gap allowing the rod to move vertically without friction from contact with the channel walls while the rod in the channel is protected from impact by the channel walls; anda processor communicatively coupled with the weight force sensing device, wherein the processor, responsive to executing computer instructions, performs operations comprising: receiving an information signal indicative of fluid outflow and/or fluid intake amount associated with a surgical patient, the information signal comprising an indication of at least one of urine fluid outflow amount from a surgical patient;suction fluid outflow amount from a surgical cavity of the surgical patient;irrigation fluid intake amount into a surgical cavity of the surgical patient; andintravenous fluid intake amount into the surgical patient;determining an incremental weight of the contents of the container from a surgical object placed into the container through one of the at least one port opening, based at least on an information signal received from the weight force sensing device indicating a change in weight force sensed by the weight force sensing device contemporaneous with the placement of the surgical object into the container through the at least one port opening with the aid of gravity;identifying the type of surgical object placed into the container through the one of the at least one port opening;determining a fluid weight of the surgical object placed into the container through the at least one port opening by at least comparing the determined incremental weight of the contents of the container from the surgical object placed into the container to an expected “dry” (non-fluid-filled) weight value for the identified type of the surgical object; andaccumulating a total amount of estimated net fluid volume loss for a surgical patient, based at least on the determine... | Surgitrac Corporation | A61B005/02 G01G017/04 G01G019/414 G01G021/23 G01G021/28 G01G019/387 G06F019/00 A61B019/02 A61B019/00 B65F001/14 G01G019/40 A61M001/00 A61B005/20 B62B003/10 A61B005/00
US09934697 | B2 | 2018-04-03 | Modular wearable device for conveying affective state | 2014-11-06 | A wearable device conveys information to a user. The device includes a master soft circuit cell and a plurality of actuation soft circuit cells. These cells are physically interconnected to form a garment that is worn by a user and each of these cells includes an electrically non-conductive fabric covering. Each of the actuation cells is electrically connected to and operates under the control of the master cell. The master cell is configured to wirelessly receive actuation instructions and activate a combination of the actuation cells based on the received actuation instructions. Each of the actuation cells is configured to generate a particular actuation that is perceived by one or more senses of the user whenever the actuation cell is activated by the master cell. A system also conveys affective state information to a user. | 1. A wearable device for conveying information to a user, comprising: a master soft circuit cell which is a pliable electronic circuit; anda plurality of actuation soft circuit cells each of which is another pliable electronic circuit,the master soft circuit cell and each of the actuation soft circuit cells comprising an electrically non-conductive fabric covering,a physical interconnection of the master and actuation soft circuit cells forming a garment that is worn by the user,each of the actuation soft circuit cells being electrically connected to and operating under the control of the master soft circuit cell,the master soft circuit cell being configured to wirelessly receive actuation instructions and activate a combination of the actuation soft circuit cells based on the received actuation instructions, andeach of the actuation soft circuit cells being configured to generate a particular actuation that is perceived by one or more senses of the user whenever the actuation soft circuit cell is activated by the master soft circuit cell. 2. The wearable device of claim 1, wherein, the master soft circuit cell further comprises a battery, a microcontroller, and a wireless receiver,the electrical connection between each of the actuation soft circuit cells and the master soft circuit cell comprises a power distribution bus and a communication bus that serially pass through the master and actuation soft circuit cells,the received actuation instructions specify a set of actuations,the power distribution bus is configured to supply power from the battery to each of the actuation soft circuit cells, andthe microcontroller is configured to interpret the received actuation instructions and send commands over the communication bus to each of the actuation soft circuit cells, said commands causing the actuation soft circuit cells whose particular actuation is in said set to be activated, and causing the actuation soft circuit cells whose particular actuation is not in said set to be deactivated. 3. The wearable device of claim 2, wherein the wireless receiver comprises one of: a Bluetooth personal area network receiver; ora Wi-Fi local area network receiver. 4. The wearable device of claim 2, wherein the communication bus comprises an Inter-Integrated Circuit (I2C) bus and said commands follow the I2C message protocol. 5. The wearable device of claim 1, wherein, each of the actuation soft circuit cells further comprises a microcontroller and an actuator unit,the electrical connection between each of the actuation soft circuit cells and the master soft circuit cell comprises a power distribution bus and a communication bus that serially pass through the master and actuation soft circuit cells,each of the actuation soft circuit cells is configured to receive power from the master soft circuit cell over the power distribution bus, andthe microcontroller on each of the actuation soft circuit cells is configured to receive commands from the master soft cir... | MICROSOFT TECHNOLOGY LICENSING, LLC | G09B019/00 G06F001/16 G06F003/01 A41D001/04 A41D023/00 A41F009/00 A61B005/16 G08B007/02 G08B007/06 G09B005/00 A61B005/00
US20160143561 | A1 | 2016-05-26 | SELF-CONTAINED, PORTABLE H2/CO2 (AIR) RATIO APPARATUS | 2015-09-23 | A self-contained, portable H2/CO2 (air) ratio apparatus has an exhaled breath system, an air analyzing chamber, a hydrogen concentration analyzer, a carbon dioxide (air) concentration analyzer, an air outlet, a computational analyzer and a display unit. The display unit can be a part of the self-contained, portable H2/CO2 (air) ratio apparatus, remotely a part of the self-contained, portable H2/CO2 (air) ratio apparatus, or combinations thereof. The apparatus can assist in the diagnosis of a neonatal patient developing necrotizing enterocolitis and/or lactose intolerance. | 1. A hydrogen/carbon dioxide ratio apparatus comprising: a self-contained, portable hydrogen/carbon dioxide ratio unit having: an exhaled breath system having a pacifier, an exhaled delivery tube having a length that allows the distal orifice of the exhaled delivery tube to be positioned in the patient's pharynx area, and an aspirator to create a negative pressure in the exhaled deliver tube to collect the patient's exhaled breath into the exhaled delivery tube;an air analyzing chamber interconnected to a hydrogen concentration analyzer, a carbon dioxide concentration analyzer, and an air outlet, the air analyzing chamber receives the collected patient's exhaled breath so (a) the hydrogen concentration analyzer can measure the amount of hydrogen in the collected air breath and (b) the carbon dioxide concentration analyzer can measure the amount of carbon dioxide in the collected air breath; anda computational analyzer (a) receives the hydrogen measurement and the carbon dioxide measurement, (b) determines whether the collected air breath is an adequate specimen to retain the hydrogen measurement and the carbon dioxide measurement, (c) only when the collected air breath is adequate, calculates a hydrogen/carbon dioxide ratio, and records the hydrogen measurement, the carbon dioxide measurement, the calculated hydrogen/carbon dioxide ratio and time the measurements occurred; anda display unit (a) receives the hydrogen measurement, the carbon dioxide measurement, the calculated hydrogen/carbon dioxide ratio and time the measurements occurred, and (b) presents the hydrogen measurement, the carbon dioxide measurement, the calculated hydrogen/carbon dioxide ratio and time the measurements occurred on a screen. 2. The apparatus of claim 1, further comprising a hydrogen auto-calibration unit for the hydrogen concentration analyzer. 3. The apparatus of claim 1, further comprising a carbon dioxide auto-calibration unit for the carbon dioxide concentration analyzer. 4. The apparatus of claim 1, wherein the aspirator is mechanically controlled by the computational analyzer. 5. The apparatus of claim 1, wherein the exhaled breath system has a flow sensor that indicates an exhaled breath. 6. The apparatus of claim 5, wherein the flow sensor interconnects to the computational analyzer and computational analyzer initiates the aspirator to create negative pressure in the exhaled breath system when the flow sensor indicates an exhaled breath. 7. The apparatus of claim 1, wherein the air outlet has a valve. 8. The apparatus of claim 1, wherein the valve is mechanically controlled by the computational analyzer. 9. The apparatus of claim 1, wherein the self-contained, portable hydrogen/carbon dioxide ratio apparatus is used to assist in diagnosing necrotizing enterocolitis. 10. The apparatus of claim 1, wherein the self-contained, portable hydrogen/carbon dioxide ratio apparatus is used to assist in diagnosing lactose intolerance. 11. The apparatus of claim... | ONY, Inc. | A61B005/08 A61B005/00 A61B005/097 A61B005/087
US10512416 | B2 | 2019-12-24 | Estimation of blood flow rates | 2015-11-11 | Disclosed are various embodiments for estimating the flow of blood through a blood vessel in a region of interest. Passage of a bolus of fluid through an imaged blood vessel over a period of time is tracked by a computing device. The computing device fits the tracked passage of the bolus of fluid to a modeled passage of the bolus of fluid. The computing device then estimates a volume of blood flow through the imaged blood vessel based at least in part on a fit of the tracked passage of the bolus of fluid to the modeled passage of the bolus of fluid. | 1. A system, comprising: at least one computing device;an imaging device in data communication with the at least one computing device, the imaging device configured to acquire an image of a region of interest of a blood vessel; andan application executed in the at least one computing device, the application comprising non-transitory computer-readable media comprising machine readable instructions that, when executed, cause the at least one computing device to at least: identify a passage of a bolus of fluid through the region of interest of the imaged blood vessel over a period of time; andestimate, with a mass transit model, blood flow through the region of interest of the imaged blood vessel based at least in part on the passage of the bolus of fluid through region of interest of the imaged blood vessel during the period of time. 2. The system of claim 1, wherein causing the at least one computing device to estimate the blood flow through the imaged blood vessel further comprises causing the at least one computing device to at least estimate the blood flow based at least in part on an anatomical data of the blood vessel. 3. The system of claim 1, wherein the bolus of fluid comprises an arterial spin labeling (ASL) tagged bolus of fluid. 4. The system of claim 3, wherein the ASL tagged bolus of fluid corresponds to a pulsed ASL tagged bolus of fluid. 5. The system of claim 3, wherein the ASL tagged bolus of fluid corresponds to a pseudo-continuous ASL tagged bolus of fluid. 6. The system of claim 3, wherein the arterial spin labeling (ASL) tagged bolus of fluid has a box-car signal profile. 7. The system of claim 1, wherein the mass transit model includes an axial velocity profile of the bolus of fluid. 8. The system of claim 1, wherein causing the at least one computing device to at least estimate, with a mass transit model, blood flow through the region of interest of the imaged blood vessel estimates variation in axial velocity with radius of the bolus of fluid passing through the region of interest of the imaged blood vessel. 9. The system of claim 1, wherein the application causes the at least one computing device to at least acquire dynamic angiographic data which covers the region of interest of the imaged blood vessel in a subject. 10. The system of claim 9, wherein causing the at least one computing device to acquire the dynamic angiographic data comprises causing the at least one computing device to acquire the angiographic data from a medical imaging device, and wherein the subject is positioned in association with the medical imaging device and the medical imaging device is used to acquire the dynamic angiographic data. 11. A method, comprising: tracking, via a computing device, a passage of a bolus of fluid through an imaged blood vessel over a period of time;fitting, via the computing device, the tracked passage of the bolus of fluid to a modeled passage of the bolus of fluid, wherein the modeled passage of the bolus of f... | Oxford University Innovation Limited | A61B005/055 A61B005/026 G01R033/563
US09346515 | B2 | 2016-05-24 | Pedal assisted bicycle and method of controlling the pedal assisted bicycle | 2013-02-13 | A pedal assisted bicycle a first and a second wheel; a pedaling assembly mechanically decoupled from the first and second wheels by which an user can supply a pedaling power; an electric motor mechanically coupled to at least one of the wheels capable of taking a motor power; a generator device adapted to generate a generator device electric power from the pedaling power, arranged in an energy exchange relationship with the pedaling assembly and the electric motor; an energy storage device arranged in an energy exchange relationship with the electric motor and with the generator device; a control system including a module for controlling the power required to the generator device to be supplied to the electric motor and/or to the storage device; a heartbeat sensor adapted to generate a signal representing heartbeat. | 1. A pedal assisted bicycle comprising: first and second wheels,a pedaling assembly mechanically decoupled from said first and second wheels, by which a user can supply a pedaling power,an electric motor mechanically coupled to at least one of said first and second wheels capable of taking a motor power,a generator device adapted to generate a generator device electric power from said pedaling power, arranged in an energy exchange relationship with the pedaling assembly and the electric motor,an energy storage device arranged in an energy exchange relationship with said electric motor and generator device,a control system comprising a module for controlling the power required to the generator device to be supplied to the electric motor and/or to the storage device, anda heartbeat sensor adapted to generate a signal representing the heartbeat,wherein said module for controlling the electric power required to the generator device is operatively connected to the heartbeat sensor and is configured in order to determine the electric power required to the generator device as a function of said signal representing the heartbeat. 2. The pedal assisted bicycle according to claim 1, wherein said control system comprises a module for controlling the electric motor such that the bicycle moves according to predetermined motion laws, said control system being further configured for comparing the electric motor required power and the generator device required power and for driving the generator device and storage device so that: if the motor required power is equal to the generator device required power, the generator device required power is totally supplied to the electric motor;if the motor required power is greater than the generator device required power, the generator device required power is totally supplied to the electric motor and the portion of the lacking power is supplied to the electric motor by the storage device;if the motor required power is less than the generator device required power, the portion of the generator device required power equal to the motor required power is supplied to the electric motor, and the portion of the excessive generator device required power is supplied to the storage device. 3. The pedal assisted bicycle according to claim 1, wherein said module for controlling the generator device required power is configured for determining the generator device required electric power by performing a comparison between a predetermined heartbeat pattern and the heartbeat pattern detected by said heartbeat sensor, and by performing a closed-loop control of the error between said predetermined heartbeat pattern and said heartbeat pattern detected by said heartbeat sensor. 4. The pedal assisted bicycle according to claim 1, wherein said heartbeat sensor is connected to the control system according to a wireless mode. 5. The pedal assisted bicycle according to claim 1, wherein said module for controlling the generator device requir... | Politecnico di Milano Università degli Studi di Bergamo | B62M006/40 B62M006/45 B62M006/50 B62M006/60 B60L011/00 B60L007/12 B60L011/18 B60L015/20 A61B005/024 B62K011/00 B62M006/90 G01P003/44 B62J099/00
US20160120471 | A1 | 2016-05-05 | PIECE OF CLOTHING FOR A HUMAN BEING | 2015-12-02 | A piece of clothing is described and which detects external effects, and which further generates a signal which can be relayed to an analysis unit, and a multiplicity of sensors are incorporated into the piece of clothing, and which are arranged in zones of the piece of clothing, and which further adjoins a vital organ of a human body wearing the piece of clothing. | 1. A piece of clothing for a human being, comprising: a first sensor made integral with a piece of clothing, and which is worn by a human body, and wherein the piece of clothing has a zone which overlies a vital organ of the human body when the piece of clothing is worn correctly, and wherein the first sensor is made integral with the zone, and wherein the first sensor is formed, at least in part, of a polyvinylidene fluoride (PVDF) structure which has piezoelectric properties, and which further is capable of generating a first electrical signal which corresponds to a pressure which is applied to the first sensor by an exterior force, and wherein the first electrical signal has a given first frequency;a second sensor made integral with the zone of the clothing, and which further overlies the vital organ of the human being, and wherein the second sensor is formed, at least in part, of an electrically conductive polymer which defines an electrical circuit, and wherein the electrically conductive polymer is applied in a manner where the second sensor is oriented in at least partial, covering relation relative to the first sensor, and wherein the first sensor is electrically coupled to the electrical circuit, and the first electrical signal is transmitted by the electrical circuit;a source of electrical power coupled to the electrical circuit, and wherein a second electrical signal having a given frequency is generated by the second sensor when the electrical circuit is partially destroyed by a breach which is formed through the piece of clothing by the exterior force, and wherein the frequency of the second electrical signal is different than the frequency of the first electrical signal, and which is generated by the first sensor; andan analysis unit coupled in signal receiving relation relative to the first and second electrical signals which are transmitted by the electrical circuit, and wherein the analysis unit distinguishes, based upon the frequencies of the first and second electrical signals, the amount of the exterior force applied to the piece of clothing; the type of destruction experienced by the piece of clothing; and the location of the breach which is formed through the zone of the piece of clothing, and which overlies the vital organ of the human being wearing the piece of clothing. 2. A piece of clothing as claimed in claim 1, and wherein the first and second sensors include multiple first and second sensors which are positioned in multiple discrete zones which are individually assigned to specific vital organs of the human body. 3. A piece of clothing as claimed in claim 1, and wherein the electrically conductive polymer is fabricated from polyethylene dioxythiophene (PEDOT). 4. A piece of clothing as claimed in claim 1, and further comprising a radio unit which is coupled in first and second electrical signal receiving relation relative to the piece of clothing. 5. A piece of clothing as claimed in claim 1, and wherein the firs... | HEXONIA GMBH | A61B005/00 A61B005/103 A61B005/0205 F41H001/02 A41D001/00
US09730604 | B2 | 2017-08-15 | Methods for detecting atrial tachyarrhythmia in implantable devices without dedicated atrial sensing | 2016-01-27 | An apparatus comprises an implantable cardiac signal sensing circuit configured to provide a sensed depolarization signal from a ventricle and a processor. The processor includes a signal analyzer module and a tachyarrhythmia discrimination module. The signal analyzer module is configured to determine a measure of stability of ventricular (V-V) depolarization intervals using the depolarization signal, and determine a rate of change of the measure of stability. The tachyarrhythmia discrimination module is configured to detect an episode of tachyarrhythmia using the depolarization signal, determine whether the detected tachyarrhythmia is indicative of atrial tachyarrhythmia using the determined rate of change, and provide the determination to a user or process. | 1. An apparatus comprising: an implantable cardiac signal sensing circuit, configured to provide a sensed depolarization signal from a ventricle;a processor communicatively coupled to the cardiac signal sensing circuit, the processor including: a signal analyzer module configured to measure V-V depolarization intervals; anda tachyarrhythmia discrimination module configured to: detect an episode of tachyarrhythmia using the measured V-V depolarization intervals;detect a sudden presence of a mix of both long and short V-V depolarization intervals;classify the detected tachyarrhythmia as atrial tachyarrhythmia or ventricular tachyarrhythmia using a detected pattern of long and short V-V depolarization intervals included in the mix of both long and short V-V depolarization intervals; andprovide the tachyarrhythmia classification to a user or process. 2. The apparatus of claim 1, including: a memory communicatively coupled to the processor and configured to a store a signal morphology template; anda far-field sensing channel configured to provide a far-field atrial and ventricular signal using an electrode location that is outside of an atrium or ventricle, andwherein the tachyarrhythmia discrimination module is configured to: upon classifying the detected tachyarrhythmia as atrial tachyarrhythmia, initiate sensing of the far-field signal; andinitiate a comparison a morphology of the far-field signal to a stored morphology template by the signal analyzer module; andconfirm the classification as atrial tachyarrhythmia using the morphology comparison. 3. The apparatus of claim 2, wherein the signal analyzer module is configured to: calculate a measure of similarity between a segment of the far-field signal and the morphology template, wherein the stored morphology template is representative of normal sinus rhythm;calculate a central tendency of the measure of similarity and a variance of the measure of similarity, andwherein the tachyarrhythmia discrimination module is configured to confirm the classification as atrial tachyarrhythmia when the central tendency of the measure of similarity is above a central tendency threshold, the variance of the measure of similarity is within a variance range, and the depolarization rate or interval satisfies a rate or interval detection threshold. 4. The apparatus of claim 1, including: a memory communicatively coupled to the processor; anda communication circuit, communicatively coupled to the processor, configured to communicate information wirelessly with a second separate device, andwherein the tachyarrhythmia discrimination module is configured to: upon classifying the detected tachyarrhythmia as atrial tachyarrhythmia, initiate storage of a segment of the depolarization signal when a duration of the detected arrhythmia exceeds a duration threshold;initiate storage of a histogram of V-V depolarization intervals, wherein the histogram indicates regular and irregular V-V depolarization intervals; andcommunicate... | Cardiac Pacemakers, Inc. | A61B005/0464 A61B005/00 A61B005/04 A61N001/362 A61N001/365
US10095837 | B2 | 2018-10-09 | Real-time phase detection of frequency band | 2015-11-18 | Techniques are described for real-time phase detection. For the phase detection, a signal is correlated with a frequency component of a frequency band whose phase is being detected, and the correlation includes predominantly decreasing weighting of past portions of the signals. | 1. A method comprising: receiving, with a processor of a medical device, a first signal generated from a second signal;correlating, with the processor of the medical device, the first signal with a frequency component of a frequency band of the second signal, wherein correlating includes predominantly decreasing weighting of past portions of the first signal, and wherein predominantly decreasing weighting of past portions of the first signal comprises predominantly weighting present and more recent portions of the first signal more heavily than earlier portions of the first signal, and wherein the frequency band includes one or more frequency components;determining, with the processor of the medical device, a phase of the frequency component based on a filtered signal outputted from the correlation;determining, with the processor of the medical device, a phase of the frequency band based on the determined phase of the frequency component;instructing, with the processor of the medical device, a therapy module to deliver therapy based on the determined phase of the frequency band; anddelivering, with the therapy module, the therapy in response to the instruction to deliver the therapy. 2. The method of claim 1, wherein the first signal is one of: the second signal pre-filtered to generate the first signal; orthe same as the second signal. 3. The method of claim 1, further comprising: determining a frequency weighting for the frequency component based on the filtered signal,wherein determining the phase of the frequency band comprises determining the phase of the frequency band based on the determined phase of the frequency component and the frequency weighting for the frequency component. 4. The method of claim 1, wherein determining the phase of the frequency band comprises: adding or subtracting a phase delay caused from generating the first signal from the second signal to the determined phase of the frequency component; anddetermining the phase of the frequency band based on the result of the adding or subtracting. 5. The method of claim 1, wherein correlating the first signal with the frequency component of the frequency band of the second signal comprises: correlating the first signal with one or more waveforms of a periodic basis set, wherein the frequency of the waveforms of the periodic basis set is the frequency component. 6. The method of claim 5, wherein the waveforms of the periodic basis set include a sine wave and a cosine wave. 7. The method of claim 1, wherein correlating the first signal with the frequency component of the frequency band of the second signal comprises: Fourier transforming the first signal with an exponentially decreasing weighting of past portions of the first signal. 8. The method of claim 1, further comprising: determining a phase of the filtered signal, wherein the phase of the filtered signal includes the phase of the frequency component and a phase error;determining the phase error;subtracting the det... | Medtronic, Inc. | G16H040/63 G06F019/00 A61B005/048 A61B005/04 A61B005/00
US20160135745 | A1 | 2016-05-19 | MONITORING GLOVE AND SYSTEM | 2014-11-19 | A measuring glove including a first sensing unit and a second sensing unit is provided. The first sensing unit measures a first physiological signal. The second sensing unit measures a second physiological signal. When the first sensing unit operates, the second sensing unit stops operating. When the second sensing unit operates, the first sensing unit stops operating. | 1. A measuring glove, comprising: a first sensing unit measuring a first physiological signal; anda second sensing unit measuring a second physiological signal, wherein when the first sensing unit operates, the second sensing unit stops operating, and when the second sensing unit operates, the first sensing unit stops operating. 2. The measuring glove as claimed in claim 1, wherein the first sensing unit is one of a light sensor, a voltage sensor, a vibration sensor, a pressure sensor, and an audio sensor, and the second sensing unit is another of the light sensor, the voltage sensor, the vibration sensor, the pressure sensor, and the audio sensor. 3. The measuring glove as claimed in claim 1, wherein the first physiological signal is one of a heart sound signal, an ECG signal, a SpO2 signal, a temperature signal and a pulse signal, and the second physiological signal is another of the heart sound signal, the ECG signal, the SpO2 signal, the temperature signal and the pulse signal. 4. The measuring glove as claimed in claim 1, wherein the first sensing unit comprises a light emitter and a light receiver. 5. The measuring glove as claimed in claim 4, wherein the light emitter is disposed in a first side of a finger or a ear of a patient, the light receiver is disposed in a second side of the finger or the ear of the patient, and the first side is opposite to the second side. 6. The measuring glove as claimed in claim 5, wherein the light emitter emits light emitting through the finger or the ear, the light receiver generates the first physiological signal according to the light emitting through the finger or the ear. 7. The measuring glove as claimed in claim 4, further comprising: a thumb portion configured to dispose the light receiver; anda first finger portion configured to disposed the light emitter. 8. The measuring glove as claimed in claim 1, further comprising: a palm portion configured to dispose the first sensing unit. 9. The measuring glove as claimed in claim 1, further comprising: a thumb portion configured to dispose the first sensing unit; anda middle finger portion, wherein the first sensing unit is an ECG sensor, and the ECG sensor comprises a first electrode disposed in the thumb portion and a second electrode disposed in the middle finger portion. 10. The measuring glove as claimed in claim 1, further comprising: a ring finger portion configured to dispose the first sensing unit, wherein the first sensing unit is a temperature sensor. 11. The measuring glove as claimed in claim 1, further comprising: a first finger portion;a middle finger portion; anda ring finger portion, wherein the first sensing unit is a pressure sensor disposed in the first finger portion, the middle finger portion and the ring finger portion. 12. The measuring glove as claimed in claim 1, further comprising: a film covering the first sensing unit. 13. The measuring glove as claimed in claim 1, further comprising: a third sensing unit measuring... |  | A61B005/00 A61B007/04 A61B005/1455 A61B005/0205 A61B005/0408
US09326742 | B2 | 2016-05-03 | Systems for integrated radiopharmaceutical generation, preparation, transportation and administration | 2007-12-28 | A hazardous fluid handling system which includes a housing having a radiation shielded internal chamber disposed therein which may accommodate a container holding a radioactive fluid, one or more radioactivity detectors positioned within the shielded internal chamber in operational proximity to the radioactive fluid in the container, and a dosimeter control device electronically coupled to the radioactivity detector(s) is disclosed. The dosimeter control device is operational for determining information regarding the radioactive fluid in the container based on individual measurements received from the one or more radioactivity detectors. The hazardous fluid handling system may be integrated into a patient support platform which includes a patient stimulus apparatus, an imager proximate the patient support platform, a radiopharmaceutical fluid delivery system for infusing a radiopharmaceutical fluid into a patient, a patient monitor to be associated with the patient, and an integrated system controller. | 1. A hazardous fluid handling system comprising: a housing;a radiation shielded internal chamber disposed within the housing and configured to hold a container holding a radioactive fluid and having a known internal geometry;at least two radioactivity detectors positioned within the shielded internal chamber in operational proximity to the radioactive fluid in the container; anda dosimeter control device electronically coupled to each of the at least two radioactivity detectors,wherein the dosimeter control device is operational for determining information on a fill volume of the radioactive fluid in the container based on individual measurements received from the at least two radioactivity detectors, andwherein the dosimeter control device is configured to use the individual measurements received from the at least two radioactivity detectors for determining error information comprising at least one of a presence of bubbles in the container and a presence of spills or leaks of the radioactive fluid from the container. 2. The hazardous fluid handling system of claim 1, further comprising: a pump;a user interface; anda pump control device electronically linked to the pump and the user interface. 3. The hazardous fluid handling system of claim 2, wherein the dosimeter control device is electronically linked to the pump control device. 4. The hazardous fluid handling system of claim 2, further comprising a communications interface electronically linked to the pump control device. 5. The hazardous fluid handling system of claim 4, wherein the communications interface is configured to electronically link to an information network. 6. The hazardous fluid handling system of claim 2, further comprising a fluid path in fluid communication with the pump. 7. The hazardous fluid handling system of claim 6, wherein the fluid path is further adapted to communicate with an external hazardous fluid source. 8. The hazardous fluid handling system of claim 1, further comprising a data recording device electronically linked to the dosimeter control device. 9. The hazardous fluid handling system of claim 1, wherein the hazardous fluid handling system is integrated into a patient support platform. 10. The hazardous fluid handling system of claim 9, wherein the patient support platform comprises: a patient stimulus apparatus;an imager proximate the patient support platform;a radiopharmaceutical fluid delivery system for infusing a radiopharmaceutical fluid into a patient;a patient monitor to be associated with the patient; andan integrated system controller operably associated with the patient stimulus apparatus, the imager, the radiopharmaceutical fluid delivery system, and the patient monitor. 11. The hazardous fluid handling system of claim 10, wherein the radiopharmaceutical fluid delivery system comprises a plurality of fluid injectors each individually controlled by the integrated system controller. 12. The hazardous fluid handling system of claim 1, wh... | Bayer Healthcare LLC | A61M003/00 A61B006/00 A61B006/03 A61B006/04 A61M005/172 G06F019/00 G21G004/08 A61B006/10 A61B005/0205 A61B005/00 A61M005/00 A61M005/145 A61M005/168 A61M005/178
US09345419 | B2 | 2016-05-24 | Method and apparatus for the objective detection of auditive disorders | 2009-07-24 | This invention relates to the apparatus for the objective detection of auditive disorders, by recording arising potentials of a steady state with simultaneous acoustic stimulation via the bones or the airway to be used in newborns. The apparatus comprises the necessary means for the independent generation of two different sound stimuli presented to the subject simultaneously via osseous, and airway electroacoustic transducers, the synchronous recording of the cerebral electrical activity of the subject, the continuous evaluation of the contact of the electrodes, the interactive measurement of the levels of ambient noise, and the wireless digital transmission of the bioelectric activity towards a computer. | 1. An apparatus for the objective evaluation of hearing through brain responses or steady state auditive potentials, comprising: means for simultaneously providing a patient with different stimuli through electro-acoustic transducers for both air and bone conduction modes;means for synchronous recording of the neuroelectric activity of the patient, further comprising first and second audio channels, said stimuli comprise a first stimulus comprising one or several mixed low frequencies, amplitude modulated, carrier tones with a M1 modulation frequency through the bone transducer and first audio channel, and the second stimulus comprises one or several mixed high frequencies, amplitude modulated, carrier tones with a M2 modulation frequency through the airway transducer and second audio channel, with synchronous recording of the neuroelectric activity caused by the simultaneous bone and airway stimulation;andmeans for processing and analyzing the neuroelectric activity to obtain a result of the primary hearing test for classifying, the hearing disorder as a conductive or sensorineural disorder. 2. The apparatus according to claim 1, further comprising means for collecting and measuring the ambient noise level. 3. The apparatus of claim 2, further comprising means for evaluating the ambient noise value and the impedance values for determining whether a pre-established rejection criterion is met for interruption of the hearing evaluation. 4. An apparatus for objectively evaluating the hearing of a subject comprising: (a) a first electro-acoustical transducer comprising means for generating a first stimuli comprising first frequencies specific for bone stimulation and providing a resultant first impedance value;(b) a second electro-acoustical transducer comprising means for generating a second stimuli comprising second frequencies specific for airway stimulation providing a resultant second impedance value, said first frequencies being lower than said second frequencies;(c) said (a) and (b) further comprise means for continuously and simultaneously stimulating of the respective bone and airway; and(d) means for calculating the respective impedance values of (a) and (b) with the simultaneous stimulation to determine whether there is a conductive or sensorineural disorder in the subject. 5. The apparatus of claim 4, further comprising means for measuring the ambient noise level, wherein the ambient noise level and impedance value are calculated according to a pre-established rejection criterion as to whether or not to interrupt the evaluation of the hearing of the subject. 6. The apparatus of claim 4, said first transducer comprises a bone conductor and said second transducer comprises an airway conductor, and further comprising means for synchronous recording of the neuroelectric activity from the respective bone and airway conductors. 7. The apparatus of claim 5, said first transducer comprises a bone conductor and said second transducer compris... | CENTRO DE NEUROSCIENCIAS DE CUBA | A61B005/0484 A61B005/12
US09339205 | B2 | 2016-05-17 | Waveform analyzer and waveform analysis method | 2012-01-26 | A waveform analyzer includes a converter which converts a logical function, where a pair of data including a time and a value at the time is variable, created according to data sets of a time and a value of a signal waveform at the time into a second function expressed by a binary decision diagram, an acquisition unit which obtains for each of characteristic points of a reference waveform a condition representative of constraints on a relationship between time information specified by the points and a value corresponding to the time information in the signal waveform according to a value of the reference waveform at the points and a specified tolerance given to a value of the reference waveform, and a searching unit which applies the condition for each of the points to the second function to obtain a time range which meets the entirety of the conditions. | 1. A waveform analyzer which searches a signal waveform of an input signal for a portion similar to a reference waveform, the waveform analyzer comprising a computer that includes: an electrocardiogram sensor which measures an action potential of cardiac muscle and which outputs an electrocardiogram signal; anda processor which performs a process including: a signal waveform obtaining process which obtains a pair of data sets of a time and a value of the signal waveform of the electrocardiogram signal, which is outputted from the electrocardiogram sensor, at the time obtained by sampling the signal waveform;a convert process which creates according to the pair of data sets a characteristic function which is a logical function where a pair of data including a time and a value at the time is variable, and which converts the created characteristic function into a second function which is an expression on a binary decision diagram;a constraint condition acquisition process which obtains for each of characteristic points of the reference waveform a constraint condition representative of constraints on a relationship between time information specified by the characteristic points and a value corresponding to the time information in the signal waveform according to a value of the reference waveform at the characteristic points and a specified tolerance given to a value of the reference waveform by creating a binary decision diagram representing the constraint condition according to the value of the reference waveform at the characteristic points and the specified tolerance; anda searching process which applies the constraint condition for each of the plurality of characteristic points to the second function to obtain a time range which meets the constraint condition for all the plurality of characteristic points, and which outputs information of the time range as a result of a search, the constraint condition being applied by performing a logical operation between the binary decision diagram representing the second function and the binary decision diagram representing the constraint condition. 2. The waveform analyzer according to claim 1, wherein the searching process firstly performs on the second function for each of the characteristic points a variable replacement operation, which corresponds to a time difference between a time specified by the characteristic points and a reference time preliminarily set for the reference waveform, and subsequently applies the constraint condition obtained for the characteristic points to the characteristic function on which the variable replacement operation is performed to calculate the time range which meets the constraint condition for the characteristic points. 3. The waveform analyzer according to claim 1, wherein the searching process applies to the second function for each of the characteristic points the constraint condition obtained for each of the characteristic points to calculate the time range which ... | FUJITSU LIMITED | A61B005/0452 A61B005/0464 G06K009/00 G01R029/02
US09326728 | B2 | 2016-05-03 | BioMEMS sensor and apparatuses and methods therefor | 2011-02-25 | Electronic devices, apparatus, systems, and methods of operating and constructing the devices, apparatus, and/or systems include a wireless sensor configured to measure strain of hardware implanted in a subject. In various embodiments, temporal measurement of the hardware strain includes monitoring changes of the resonant frequency of the sensor. The sensor can be realized as an inductively powered device that operates as an all-on-chip resonator, where the components of the sensor are biocompatible. Additional apparatus, systems, and methods are disclosed. | 1. A method of monitoring changes in a hardware implant device using a sensor device mounted to or integrated with the hardware device, which is implanted in a subject, wherein the sensor device comprises:a substrate; anda resonator configured as a strain gauge composed essentially of: a solid dielectric material disposed over the substrate as a dielectric capacitor;a conductive layer between the substrate and the solid dielectric material and in contact with the substrate and the solid dielectric material; andat least one conductive coil disposed on the solid dielectric material,wherein the substrate, the conductive layer, the solid dielectric material, and the conductive coil are configured to undergo strain as the hardware implant device mounted to or integral with the sensor device undergoes strain,wherein the substrate, the conductive layer, the solid dielectric material, and the conductive coil are configured as a biocompatible sensor device for implantation in a biological subject,wherein the resonator is configured to be inductively powered without an incorporated power supply,wherein the resonator is configured to reflect electromagnetic waves corresponding to a resonant frequency of the resonator in response to an alternating magnetic or electromagnetic field directed toward the sensor device, andwherein the resonant frequency of the resonator shifts in correspondence with strain applied to the resonator by the hardware implant device,wherein the method comprises the steps of:inductively powering the resonator by directing an alternating magnetic or electromagnetic field toward the sensor device at different times over a period;receiving electromagnetic waves reflected from the resonator, which undergoes strain as the hardware implant device undergoes strain, at different times over the period;determining a shift in a resonant frequency of the resonator based on the received reflected electromagnetic waves at different times over the period; anddetermining a temporal change in strain of the hardware implant device based on the determined shift in the resonant frequency of the resonator over the period. 2. The method of claim 1, wherein the hardware implant device is implanted in a biological subject. 3. The method of claim 2, wherein: the resonant frequency determining step determines the resonant frequency of the resonator by inputting the received signals to a spectrum analyzer, andthe resonator is configured to have a resonant frequency ranging from 50 MHz to 7 GHz at no load. 4. The method of claim 2, further including the step of determining a change in the biological subject based on the temporal change in strain of the hardware implant device. 5. The method of claim 2, wherein the temporal change in strain of the hardware implant device is determined using a strain-frequency calibration of the hardware implant device. 6. The method of claim 2, further including the step of monitoring surface bending strains of the hardware i... | INNOVATIVE IN VIVO SENSING, LLC | A61B005/103 A61B005/117 A61B005/00 H01F041/04 H01F017/00
US09332942 | B2 | 2016-05-10 | Systems, processes and computer-accessible medium for providing hybrid flourescence and optical coherence tomography imaging | 2008-01-28 | Apparatus and method can be provided for obtaining information regarding at least one portion of a biological structure. For example, using a first arrangement, it is possible to generate a first electro-magnetic radiation to be forwarded to the portion and receive, from the portion, a second fluorescent radiation associated with the first electro-magnetic radiation Further, using a second arrangement, it is possible to obtain information associated with the second fluorescent radiation, and generate at least one image of the portion as a function of the second fluorescent radiation. According to an exemplary embodiment of the present invention, the first arrangement can be inserted into at least one lumen, and the image of the portion may be associated with the at least one lumen. | 1. An apparatus for obtaining information regarding at least one portion of a chosen biological target, comprising: a catheter having proximal and distal ends and an axis and including a first channel and a second channel, wherein the first and second channels together are configured to transmit at least first, second, and third radiations,wherein said first channel includes a first optical waveguide and is configured to transmit through the distal ends towards the proximal end, the first radiation that includes one of a mechanical radiation and an optical radiation including fluorescent radiation, the first radiation generated within said biological target in response to absorption by said target of third radiation transmitted through catheter from the proximal end;wherein said second channel includes a second optical waveguide and is configured to transmit, from said biological target through the distal end towards the proximal end, the second radiation representing anatomical characteristics of said target and containing optical-computed-tomography (OCT) information;wherein each of the first and second channels is structured to deflect a radiation transmitted through such channel to cause said radiation to traverse a first portion of such channel along the axis and a second portion of such channel along a line that is transverse to the axis;wherein said first optical waveguide dimensioned to channel only the first radiation and the second optical waveguide is dimensioned to channel only the second radiation;wherein the catheter is configured to outcouple said first and second radiations from the first and second optical waveguides at respectively corresponding first and second locations,wherein the first location is defined at an outer surface of the first optical waveguide away from an output facet of the first optical waveguide at a proximal end thereof, the second location defined at an output facet of the second optical waveguide at a proximal end thereof,and a detector operably connected with the proximal end to acquire radiation energy and to produce output data representing said target, said output data including a first data portion representing the first radiation and a second data portion representing the second radiation. 2. The apparatus according to claim 1, wherein the first and second optical waveguides extend beyond a sheath of the catheter,wherein the first optical waveguide is configured to emit said first radiation from an output surface of the first optical waveguide at a location prior to a proximal end of the first optical waveguide. 3. The apparatus according to claim 1, further comprising a programmable processor operably connected with said detector, wherein said first radiation includes incoherent radiation, said second radiation includes coherent radiation,wherein said processor is enabled by program code stored on a tangible non-transitory computer-readable medium to extract, from said output data, data components... | THE GENERAL HOSPITAL CORPORATION | A61B005/00 G01N021/47 G01N021/64
US09336354 | B1 | 2016-05-10 | Systems and methods for simulation of hemodialysis access and optimization | 2015-01-13 | Systems and methods are disclosed for simulating or optimizing hemodialysis access. One method includes receiving a patient-specific anatomic model of a patient's vasculature; computing a pre-treatment hemodynamic characteristic of a pre-treatment geometry of a portion of the anatomic model; simulating a post-treatment geometry of a vascular access in the portion of the anatomic model; computing a post-treatment hemodynamic characteristic of the post-treatment geometry of the portion of the anatomic model having the vascular access; and generating a representation of the pre-treatment hemodynamic characteristic or the post-treatment hemodynamic characteristic. | 1. A computer-implemented method of simulating or optimizing hemodialysis access, the method comprising: receiving a patient-specific, three-dimensional anatomic model of a patient's vasculature;computing a pre-treatment hemodynamic characteristic of a pre-treatment geometry of a portion of the received three-dimensional anatomic model;modifying the geometry of the received three-dimensional anatomic model such that the geometry of the received three-dimensional anatomic model defines a post-treatment geometry of a vascular access in the portion of the received three-dimensional anatomic model, the post-treatment geometry being different from the pre-treatment geometry of the portion of the received three-dimensional anatomic model;determining a computational model of a hemodynamic characteristic of the modified three-dimensional anatomic model defining the post-treatment geometry of the vascular access;computing a post-treatment hemodynamic characteristic using the computational model associated with the modified three-dimensional anatomic model; andgenerating a representation of the post-treatment hemodynamic characteristic. 2. The computer-implemented method of claim 1, wherein the hemodynamic characteristic includes blood pressure, blood velocity, or cardiac output. 3. The computer-implemented method of claim 1, further comprising: receiving or determining a geometry of a planned treatment defining the vascular access; andmodifying the received three-dimensional anatomic model such that the post-treatment geometry is based on the geometry of the planned treatment. 4. The computer-implemented method of claim 3, wherein the planned treatment includes a graft. 5. The computer-implemented method of claim 1, further comprising: receiving or measuring one or more candidate locations of the received three-dimensional anatomic model, wherein the portion of the received three-dimensional anatomic model is selected from one of the one or more candidate locations. 6. The computer-implemented method of claim 1, further comprising: defining a cost function for optimizing the vascular access; andsolving the cost function using the pre-treatment hemodynamic characteristic and the post-treatment hemodynamic characteristic. 7. The computer-implemented method of claim 1, further comprising: receiving one or more infeasible surgical geometries for the hemodialysis access, andcomputing the post-treatment hemodynamic characteristic using the one or more infeasible geometries as constraints for the simulating of the post-treatment geometry. 8. The computer-implemented method of claim 1, further comprising: selecting a treatment or treatment location of the vascular access based on the post-treatment hemodynamic characteristic. 9. A system for simulating or optimizing hemodialysis access, the system comprising: a data storage device storing instructions for simulating or optimizing hemodialysis access; anda processor configured to execute the instructions t... | HeartFlow, Inc. | G06F019/10 A61B005/02 G06K009/00 G06K009/36 G06F019/00
US10499823 | B2 | 2019-12-10 | Selection of electroencephalography (EEG) channels valid for determining cognitive load of a subject | 2015-03-23 | Disclosed is a method and system for selection of Electroencephalography (EEG) channels valid for determining cognitive load of subject. According to one embodiment, EEG signals are obtained from EEG channels associated with subject performing cognitive tasks are received. Time-frequency features of EEG signals are extracted for a frequency band comprise maximum energy value, minimum energy value, average energy value, maximum frequency value, minimum frequency value, and average frequency value. Weight of an EEG channel associated with time-frequency feature is derived using statistical learning technique. Binary values for EEG channels corresponding to time-frequency feature are assigned using weight of EEG channel associated with time-frequency feature. Intersections of binary values of EEG channels corresponding to maximum energy value and average energy value, minimum energy value and average energy value, maximum frequency value and average frequency value, and minimum frequency value and average frequency value are computed. Unions of intersections are computed, wherein the unions represent EEG channels valid to determine cognitive load of subject. | 1. A method for selecting a set of Electroencephalography (EEG) channels valid for determining a cognitive load of a subject, the method comprising: receiving, by a processor, EEG signals obtained from a plurality of EEG channels associated with a subject performing one or more cognitive tasks, wherein the plurality of EEG channels are connected to a low resolution EEG device;decomposing, by the processor, the EEG signals into a time-frequency domain;extracting, by the processor, time-frequency features of the decomposed EEG signals for each of multiple frequency bands of each EEG channel of the plurality of EEG channels, wherein the time-frequency features include at least one of a maximum energy value, a minimum energy value, an average energy value, a maximum frequency value, a minimum frequency value, or an average frequency value, and wherein the multiple frequency bands of each of the EEG channels of the plurality of EEG channels include alpha, beta, theta, and delta;deriving, by the processor, a weight for each time-frequency feature associated with each EEG channel of the plurality of EEG channels using a statistical learning technique;assigning, by the processor, a binary value for each EEG channel of the plurality of EEG channels using the weight of each time-frequency feature associated with each EEG channel of the plurality of EEG channels, wherein the binary value of each EEG channel of the plurality of EEG channels is equal to 1 when the weight associated with the maximum energy value, the minimum energy value, the average energy value, the maximum frequency value, the minimum frequency value, and the average frequency value of each EEG channel of the plurality of EEG channels is greater than a threshold value and wherein the binary value of each EEG channel of the plurality of EEG channels is equal to 0 when the weight associated with the maximum energy value, the minimum energy value, the average energy value, the maximum frequency value, the minimum frequency value, and the average frequency value of each EEG channel of the plurality of EEG channels is less than or equal to the threshold value;computing, by the processor, a first intersection of the binary value of each EEG channel of the plurality of EEG channels corresponding to the maximum energy value and the average energy value, a second intersection of the binary value of each EEG channel of the plurality of EEG channels corresponding to the minimum energy value and the average energy value, a third intersection of the binary value of each EEG channel of the plurality of EEG channels corresponding to the maximum frequency value and the average frequency value, and a fourth intersection of the binary value of each EEG channel of the plurality of EEG channels corresponding to the minimum frequency value and the average frequency value, wherein computation of the first intersection, the second intersection, the third intersection, and the fourth intersection is performed for ... | Tata Consultancy Services Limited | A61B005/04 A61B005/0476 A61B005/00 A61B005/16
US09554749 | B2 | 2017-01-31 | Evaluation of the quality of electrode contact with a skin surface | 2016-01-15 | An instrument that utilizes body contact electrodes evaluates the quality of the connections made between the electrodes and the body. An electrode contact quality evaluation circuit performs the quality evaluation, such as by determining contact impedances for the electrodes. The corresponding contact quality of each electrode is conveyed to the user. | 1. An electrocardiograph comprising: a base unit comprising: a housing;a computing device at least partially contained within the housing; andan output device at least partially contained within the housing and configured to output information associated with electrical signals from the heart of a patient detected by the electrocardiograph;a lead interface unit having a lead interface housing separate from the housing of the base unit, the lead interface unit comprising: a lead interface arranged and configured to physically connect the lead interface unit to electrically conductive leads, wherein the leads are configured to be physically connected to at least three surface contact electrodes comprising an injection electrode, a sense electrode, and a return electrode;electrode contact detection circuitry within the lead interface housing that detects contact between the surface contact electrodes and the patient;electrode contact quality evaluation circuitry within the lead interface housing, wherein the electrode contact quality evaluation circuitry determines a quality of electrical connections between the surface contact electrodes and the patient, the electrode contact quality evaluation circuitry including driven return circuitry coupled between the sense electrode and the return electrode, the driven return circuitry configured to generate a return current having opposite polarity to a test current that is supplied to the injection electrode; andinstrument electronics within the lead interface housing that detect electrical signals from the heart of the patient; anda communication interface between the lead interface unit and the base unit. 2. The electrocardiograph of claim 1, wherein the lead interface unit further comprises a display, the display providing an indication of the qualities of the electrical connections between the surface contact electrodes and the patient. 3. The electrocardiograph of claim 2, wherein the indication of the qualities of the electrical connections are displayed at locations of a representation of the patient using color coded graphical elements, wherein the color coded graphical elements are displayed with a first color if determined that the quality is good, a second color if the quality is poor, and a third color if no connection is detected. 4. The electrocardiograph of claim 1, wherein the output device of the base unit is a display, the display providing an indication of the qualities of the electrical connections between the surface contact electrodes and the patient. | Welch Allyn, Inc. | A61B005/04 A61B005/00 A61B005/0424 A61B005/0478 A61B005/0492 A61B005/0408 A61B005/0428 A61B005/053
US10292656 | B2 | 2019-05-21 | Fabrication for ultra-compliant probes for neural and other tissues | 2015-12-08 | Methods of fabricating ultra-miniature, ultra-compliant probe arrays through spin coating, wherein a dissolvable material in hydrogel form is dispensed onto an assembled mold with wires. Once the dissolvable material is dispensed onto the mold, centrifuging spin casts the material by evaporating the solvent, forming a dried dissolvable polymer. Finally, a device is used with water to remove excess dissolvable material to obtain a dissolvable needle with wires. | 1. A method of spin coating fabrication of needle for a neural probe comprising: providing a top mold and a bottom mold to form an assembled mold with wires;dispensing a dissolvable material in a hydrogel form onto the assembled mold with the wires;centrifuging the dissolvable material to spin-cast the dissolvable material by evaporating solvent to form a dried dissolvable polymer; andusing a wiping device with intermittent applications of water to remove excess dissolvable material to obtain needle shape only in the assembled mold with the wires to form a final dissolvable needle with wires embedded. 2. The method according to claim 1, further comprising the step of using a precision assembly approach to assemble a top mold onto a wired bottom Si mold to form the assembled with the wires. 3. The method according to claim 1, further comprising the step of using a device bonder to assemble a top mold onto a wired bottom mold to form the assembled mold with the wires. 4. The method according to claim 1, wherein the dissolvable material is carboxymethyl cellulose (CMC) or a combination of CMC and at least one of sucrose, glucose, maltodextrin, poly(vinylpyrrolidone), polyvinyl alcohol, maltose, polyactic acid. 5. The method according to claim 1, wherein the hydrogel is prepared by mixing powdered carboxymethyl cellulose (CMC) with water. 6. The method according to claim 1, wherein the solvent is water. 7. The method according to claim 1, wherein the step of centrifuging is performed at 20 deg. C at 4700 rpm spinning speed, and for 5 hours. | CARNEGIE MELLON UNIVERSITY | A01B005/00 H01L021/50 A61B005/00 A61B005/0478 A61N001/05 B29C041/04 B29C035/02 B29C035/08 B29C041/20 B29C041/42 G03F007/00 B29K001/00 B29L031/00
US09724154 | B2 | 2017-08-08 | Irrigated ablation catheter with multiple sensors | 2014-11-24 | Systems and methods are disclosed for providing and using an irrigated ablation catheter. The catheter may include a distal shell electrode having irrigation apertures. An insert disposed within the electrode has protrusions that mate with orifices in the shell of the electrode. Each protrusion has a port communicating with at least one interior lumen in the insert and a sensor is disposed in each port. A support seals the proximal end of the electrode and engages the insert. The plurality of sensors may be used to measure electrical and thermal characteristics surrounding the electrode and may help assess contact between the electrode and tissue and/or determine movement of the electrode during ablation. | 1. A catheter, comprising: an elongated body;an electrode mounted at a distal end of the elongated body, wherein the electrode is configured as a shell defining an interior space;a plurality of irrigation apertures formed in the shell and communicating with the interior space;an insert disposed within the interior space having a plurality of protrusions configured to mate with a corresponding plurality of orifices in the shell of the electrode, wherein each protrusion extends at least flush with an exterior surface of the electrode and has a port communicating with at least one interior lumen in the insert;a plurality of sensors, wherein each sensor is disposed within one of the ports of the protrusions; anda support which forms a flight tight seal with a proximal-most end of the electrode and engages a proximal end of the insert to stabilize the insert against rotational motion. 2. The catheter of claim 1, wherein the insert comprises at least one longitudinally extending arm with at least one protrusion. 3. The catheter of claim 2, wherein the at least one arm has an interior lumen in communication the port of the at least one protrusion. 4. The catheter of claim 3, wherein the at least one arm has a plurality of protrusions, such that the interior lumen of the at least one arm is in communication with a plurality of ports. 5. The catheter of claim 3, further comprising at least one guide tube extending from a through-hole in the support to the interior lumen of the at least one arm. 6. The catheter of claim 2, wherein each protrusion has a shoulder positioned radially outwards from a surface of the arm, such that the shoulder engages an interior surface of the electrode surrounding the orifice. 7. The catheter of claim 6, further comprising a minimum separation between the insert and an interior surface of the electrode, wherein the minimum separation is defined by a distance from the surface of the arm and the shoulder. 8. The catheter of claim 2, further comprising a plurality of arms. 9. The catheter of claim 8, further comprising at least one passageway between the plurality of arms to allow circulation of irrigation fluid within the interior space. 10. The catheter of claim 3, wherein the insert comprises an outer portion and an inner portion and wherein the outer portion and the inner portion mate to form the at least one interior lumen. 11. The catheter of claim 10, wherein the inner portion supports the outer portion against inward deflection. 12. The catheter of claim 1, wherein at least some of the plurality of sensors are temperature sensors. 13. The catheter of claim 1, wherein at least some of the plurality of sensors are electrical sensors. 14. The catheter of claim 1, wherein at least one of the plurality of sensors is a combined temperature and electrical sensor. 15. A method for the ablation of a portion of tissue of a patient by an operator comprising: inserting a catheter into the patient, wherein the catheter ... | BIOSENSE WEBSTER (ISRAEL) LTD. | A61B018/12 A61B018/14 A61B005/01 A61B005/053 A61B005/042 A61B018/00 A61B017/00 A61B090/00
US10028695 | B2 | 2018-07-24 | Evaluation of movement disorders | 2015-11-12 | This disclosure relates to a system and method to evaluate movement disorders. A movement data aggregator can combine data from a plurality of sensors into an aggregate movement data describing multi-dimensional movement of a handheld device. A calculator to compute an indication of a movement disorder based on the movement vector data and user input data, the user input data being generated in response to physical interaction between the handheld device and a human machine interface of a computing device/machine that is separate from the handheld device. In some examples, the handheld device can communicate with computing device via a wireless interface. | 1. A system comprising: a handheld device comprising a plurality of sensors to record motion and position data of the handheld device in three dimensions during a predefined task; anda computing device comprising: a touchscreen interface to display a visualization related to the predefined task;a non-transitory computer-readable medium to store instructions; anda processor to facilitate execution of the instructions to at least: receive the motion and position data recorded by the plurality of sensors on the handheld device during the predefined task;receive pressure data from the touchscreen interface related to an interaction between the handheld device and the touchscreen interface during the predefined task;convert the motion and position data into relative movement data corresponding to a reference frame for the touchscreen interface during the predefined task; anddetermine an indication of a movement disorder based on the relative movement data and the pressure data for the predefined task. 2. The system of claim 1, wherein the motion and position data is acquired by the plurality of sensors during user interactions in response to movement of the handheld device with respect to the touchscreen interface. 3. The system of claim 2, wherein the touchscreen interface displays a given target related to the predefined task, and wherein the touchscreen interface records user interaction with the touchscreen interface based on movement of the handheld device with respect to the given target. 4. The system of claim 1, wherein the computing device computes the indication the movement disorder based on the relative movement data, the pressure data, and user input data, corresponding to the user interaction with the touchscreen interface, during the predefined task. 5. The system of claim 2, wherein the touchscreen interface generates user input data in response to the movement of the handheld device against the touchscreen interface, and wherein the indication of the movement disorder is determined further based on the user input data. 6. The system of claim 1, wherein the indication of the movement disorder is determined based on the pressure data and multiple sets of relative movement data determined by the computing device based on the motion and position data recorded by the plurality of sensors of the handheld device for a plurality of predefined tasks. 7. The system of claim 1, wherein the predefined task is associated with a test interval. 8. The system of claim 1, wherein the processor further facilitates execution of the instructions to send the indication of the movement disorder to one or more providers via a network connection. 9. The system of claim 1, wherein the indication of the movement disorder corresponds to stimulation parameters set for a stimulation therapy for a patient completing the predefined task. 10. A method comprising: displaying, by a touchscreen interface associated with a computing device comprising a processo... | The Cleveland Clinic Foundation | A61B005/00 A61B005/11
US10182766 | B2 | 2019-01-22 | Intelligent apparatus for patient guidance and data capture during physical therapy and wheelchair usage | 2016-01-29 | A system for guiding and evaluating physical positioning, orientation and motion of the human body, comprising: a cloud computing-based subsystem including an artificial neural network and spatial position analyzer said cloud computing-based subsystem adapted for data storage, management and analysis; at least one motion sensing device wearable on the human body, said at least one motion sensing device adapted to detect changes in at least one of spatial position, orientation, and rate of motion; a mobile subsystem running an application program (app) that controls said at least one motion sensing device, said mobile subsystem adapted to capture activity data quantifying said changes in at least one of spatial position, orientation, and rate of motion, said mobile subsystem further adapted to transfer said activity data to said cloud computing-based subsystem, wherein said cloud computing-based subsystem processes, stores, and analyzes said activity data. | 1. A system for guiding and evaluating physical positioning, orientation and motion of a human patient during physical therapy, comprising: a cloud computing-based subsystem, said subsystem including an artificial neural network and spatial position analyzer, said artificial neural network and said spatial position analyzer determining personalized boundary values for at least one of said physical positioning, orientation, and motion of said patient, achieved independent of personal mobility devices aiding said positioning, orientation, and motion, and outputting said values;at least one of a wearable garment, head gear, or body attachment instrumented with at least one gyroscopic motion sensing device and at least one processing module, said at least one motion sensing device detecting changes and rate of changes in motion, including at least one of said physical positioning, orientation, and movement when worn on said patient, and said at least one processing module receiving from said motion sensing device unprocessed motion change indicators and angle changes between at least two vectors relative to any reference physical orientation of said patient, and outputting said indicators and changes;a mobile subsystem receiving personalized boundary values from said cloud computing-based subsystem and receiving said angle changes between at least two vectors relative to any reference physical orientation from said wearable garment, said mobile subsystem comparing said angle changes with said boundary values to quantify degree of deviation, and, thereafter, creating a patient record comprising said changes and said rate of changes in at least one of said physical positioning, orientation, and movement, and transmitting said patient record to said cloud computing-based subsystem,said mobile subsystem further providing actionable aural guidance to said patient in substantially real-time to at least one of limiting or actuating changes and said rate of changes in at least one of physical positioning, orientation, and motion relative to said boundary values, andsaid actionable aural guidance provided to said patient including corrective indicators directed to patient performance adjustments to achieve therapeutic motion technique and preclude injury, said indicators responsive to said degree of deviation from said personalized boundary values and gaged patient response to said guidance during prescribed therapeutic exercise. 2. The system of claim 1, wherein said mobile subsystem further includes at least an artificial neural network and spatial position analyzer to operating independently without interacting with said cloud computing-based subsystem. 3. The system of claim 1, wherein said at least one motion sensing device is adapted to mount in a position on said at least one of said garment or body attachment so that arm movements can be measured during physical therapy. 4. The system of claim 1, wherein said at least one motion sensing device moun... | University of Central Oklahoma | A61B005/00 A61B005/11 A61G005/10 G06N003/02 A61G007/057 G06F019/00 G16H050/20
US09724046 | B2 | 2017-08-08 | Detecting apparatus for arrhythmia and detecting method of the detecting apparatus | 2015-10-29 | An apparatus (1) and method for detecting pulse-related parameters, such as pulse arrhythmia is presented. The apparatus (1) detects a series of pulses from a user, e.g. through a cuff-related measurement. Time differences, amplitude and pattern differences between multitudes of N pulses are determined. The apparatus investigates the degree of similarity of multiple pulse periods and/or pulse amplitudes and/or pulse pattern and subsequently generates a statistical set of similarity values based on a plurality of compared results. Basing on this, the apparatus generates a decision value based on the statistical sets of similarity values, and uses the decision value to determine whether or not the user discloses a normal pulse rhythm, atrial fibrillation, premature atrial or ventricular contractions, tachycardia, bradycardia and/or unspecified pulse arrhythmia. Further an artefact index is generated, informing the user whether a measurement was taken under sufficient artefact-free measurement conditions. | 1. An arrhythmia detecting method adopted in an arrhythmia detection apparatus, comprising: a) obtaining a pulse period sequence and a pulse amplitude sequence through a cuff of the arrhythmia detection apparatus, and in which to obtain a plurality of pulse periods and a plurality of pulse amplitudes through a pulse period and a pulse amplitude measurement module of the arrhythmia detection apparatus;b) comparing similarities of any two of the pulse periods and any two of the pulse amplitudes for respectively generating similarity statistical comparison values through a comparison generating module of the arrhythmia detection apparatus, wherein each of the generated similarity statistical comparison values is 1 if the two compared pulse periods or the two compared pulse amplitudes are similar, and is 0 if the two compared pulse periods or the two compared pulse amplitudes are dissimilar;c) calculating a statistic value comprising a set of the similarity statistical comparison values corresponding to the similarities of the pulse periods and another statistic value comprising a set of the similarity statistical comparison values corresponding to the similarities of the pulse amplitudes through a judgment generating module of the arrhythmia detection apparatus;d) establishing an arrhythmia analysis matrix according to the statistic values, the pulse period sequence and the pulse amplitude sequence through the judgment generating module, wherein the arrhythmia analysis matrix(AAM) is[⁢ZP⁢⁢ZAPPiAi], where ZP is the statistic value comprising the set of the similarity statistical comparison values corresponding to the similarities of the pulse periods, ZA is the statistic value comprising the set of the similarity statistical comparison values corresponding to the similarities of the pulse amplitudes, PPi, is the pulse period sequence, and Ai , is the pulse amplitude sequence;e) determining an atrial fibrillation, through a type determination module of the arrhythmia detection apparatus, when a coefficient of variation Var(ZP) is greater than a first atrial fibrillation threshold value AF−1, a pulse period average Mean (PPi,) is greater than a second atrial fibrillation threshold value PPmin, AF or a coefficient of variation Var(ZA) is between minimum and maximum atrial fibrillation threshold values (AF−2min, AF−2 max);f) determining, through the type determination module, if the coefficient of variation Var(ZP) is between first minimum and maximum premature contraction pulse period threshold values (APC−1min, APC−1max), or the coefficient of variation Var(ZA) is between second minimum and maximum premature contraction pulse amplitude threshold values (APC−2min, APC−2max);g) determining, through the type determination module, a premature contraction when multitude continuous dissimilar pulse periods or pulse amplitudes are detected if the coefficient of variation Var(ZP) is determined between the first minimum and maximum premature contraction pulse p... | ROSSMAX INTERNATIONAL LTD. | A61B005/0464 A61B005/00 A61B005/024 A61B005/022 A61B005/11 A61B005/021 A61B005/0255 G06F019/00
US09833190 | B2 | 2017-12-05 | Methods of detecting parameters of a lumen | 2016-01-21 | A system, device and method are presented for utilizing stretchable active integrated circuits with inflatable bodies. The invention allows for such operative features to come into direct contact with body structures, such as the inner wall of a lumen. Such direct contact increases accuracy of measurement and delivery of therapy. | 1. A method to detect parameters inside of a lumen comprising: inserting an inflatable body into said lumen, said inflatable body having a stretchable surface and a stretchable electronic circuit affixed thereto, said stretchable surface having soft or elastic properties enabling the stretchable surface to be made longer or wider without tearing or breaking, said stretchable electronic circuit being embedded in or affixed to the stretchable surface, the stretchable electronic circuit comprising: a plurality of discrete operative devices arranged as device islands on the stretchable surface, the plurality of discrete operative devices including at least one discrete operative device and at least one other discrete operative device; anda plurality of stretchable interconnects embedded in or affixed to the stretchable surface to electrically interconnect at least some of the plurality of discrete operative devices, wherein at least one stretchable interconnect of the plurality of stretchable interconnects is directly attached to the at least one discrete operative device and to the at least one other discrete operative device to directly electrically interconnect the at least one discrete operative device to the at least one other discrete operative device;inflating the inserted inflatable body to contact the at least one discrete operative device, the at least one other discrete operative device, or a combination thereof to said lumen; andmeasuring a parameter of said lumen in response to the contact. 2. The method of claim 1, further comprising providing a therapy to the lumen based on said measured parameter. 3. The method of claim 1, further comprising generating a graphical output based on said measured parameter. 4. The method of claim 3, wherein the graphical output is a map of the lumen. 5. The method of claim 1, wherein the at least one discrete operative device and the at least one other discrete operative device detect data related to the lumen. 6. The method of claim 5, wherein said data is related to electrical conductivity of plaque inside of the lumen. 7. The method of claim 5, wherein said data is related to pH. 8. The method of claim 5, wherein said data is related to enzymatic activity. 9. The method of claim 5, wherein said data is related to chemical activity. 10. The method of claim 5, further comprising continuously generating said data from the at least one discrete operative device and the at least one other discrete operative device. 11. The method of claim 10, further comprising continuously recording said data generated by the at least one discrete operative device and the at least one other discrete operative device. 12. The method of claim 1, wherein each stretchable interconnect of the plurality of stretchable interconnects is directly attached to respective discrete operative devices of the plurality of discrete operative devices to directly electrically interconnect the plurality of discrete operative devic... | MC10, Inc. | A61B005/01 A61B001/00 A61B005/02 A61B005/00 A61B005/053 A61B001/313 A61M025/10 A61F002/958
US09566033 | B2 | 2017-02-14 | Garment system with electronic components and associated methods | 2015-11-03 | The disclosure provides a wireless biometric monitoring system that may be capable of acquiring, compiling, analyzing, and transmitting biometric data in near real time/real time. The system may utilize either the most up-to-date Bluetooth protocol (currently Bluetooth Smart) or a similar wireless protocol. The system may further integrate through this wireless protocol with peripheral devices to expand the measurement capacity of the system.;In embodiments, the system may be capable of monitoring multiple biometric responses including, but not limited to: skin temperature, core temperature, respirations, heart rate, predicted tidal volume, chest wall movement, abdominal movement in conjunction with inspiration, abdominal movement in conjunction with expiration, HRR (heart rate reserve), HRV (heart rate variability), body position relevant to perpendicular, shoulder position relevant to hip position, general body posture, up time, down time, and malfunctions. | 1. A garment system for providing data comprising: a garment body comprising a form-fitting fabric having an open interior defining a torso, wherein the form-fitting fabric comprises a first layer, a second layer, and a third layer coupled together, the first layer comprising a conductive elastic fabric, the second layer comprising an insulative fabric, and the third layer comprising an elastic fabric;wherein the garment body comprises an anterior portion and a posterior portion, the garment body having a plurality of components integrated in the garment body, the components comprising: a processor;a memory;a battery unit;a respiratory monitor sub-system comprising a plurality of respiratory monitoring sites distributed in the garment body, wherein the plurality of respiratory monitoring sites comprises at least a first respiratory monitoring site, a second respiratory monitoring site, a third respirators monitoring site, and a fourth respiratory monitoring site; wherein the first respiratory monitoring site is coupled to the second respiratory monitoring site with a first portion of the conductive elastic fabric, and the third respiratory monitoring site is coupled to the fourth respiratory monitoring site with a second portion of the conductive elastic fabric, the first portion of the conductive elastic fabric being oriented in a first direction in the garment body and the second portion of the conductive elastic fabric being oriented in a second direction in the garment body, the first direction being different than the second direction:wherein the respiratory monitor sub-system is configured to assess a first resistance of the first portion of the conductive elastic fabric and assess a second resistance of the second portion of the conductive elastic fabric, the first resistance being assessed independently of the second resistance;a plurality of kinetic generators wired to a power modulator connected to the battery;a GPS monitor; anda plurality of accelerometers, at least one of the plurality of accelerometers wired to the processor;the processor, the GPS monitor, and the plurality of accelerometers each comprising at least one of a wireless receiver and a wireless transmitter; anda data analysis module comprising a second wireless receiver and a wireless transmitter, the module being capable of receiving data procured from at least one of the processor, the respiratory monitor sub-system, the GPS monitor, and the plurality of accelerometers. 2. The system of claim 1, wherein the components are arranged in a skeletal substructure, the processor and memory being housed as a unit and coupled to the posterior portion of the garment body, the battery unit being coupled to the posterior portion of the garment body,the plurality of respiratory monitoring sites being coupled to the anterior portion of the garment body and the posterior portion of the garment body,at least one of the plurality of kinetic generators being coupled to the anterior por... |  | G08C019/22 H04Q009/00 G08B001/08 A61B005/00 A61B005/0205 A61B005/08 A61B005/11
US09955874 | B2 | 2018-05-01 | Phantom | 2014-05-20 | This invention provides a phantom whose light propagation characteristics and acoustic propagation characteristics are similar to those of human tissues and which allows detection of an absorber also with an ultrasonic wave diagnostic apparatus and allows adjustment of the oxygen saturation.;A phantom has light absorbers 12a to 12d in a phantom base material 11, in which the light absorbers 12a to 12d approximate the light absorption coefficient ratio at two or more wavelengths to that of biological tissues, and the acoustic propagation characteristics of the phantom base material 11 and the light absorbers 12a to 12d are different from each other. | 1. A phantom, comprising: a light absorber in a phantom base material, whereinthe light absorber approximates a light absorption coefficient ratio at two or more wavelengths to that of biological materials, andacoustic propagation characteristics of the phantom base material and the light absorber are different from each other,wherein the light absorber contains two or more kinds of light absorbing compounds in a light absorber base material, absorption coefficient ratios μ[λ2]/μ[λ1] of the light absorbing compounds at arbitrary two wavelengths λ1 and λ2 (λ1&lt;λ2) of 600 nm or more and 1100 nm or less are different from each other, and a parameter S calculated from the following equation (1) is 0 or more and 100 or less, andwherein a coefficient of linear expansion of the light absorber base material is 100 ppm/K or more and 1000 ppm/K or less,S=P′·Hb⁡[λ1]-Hb⁡[λ2](HbO2⁡[λ2]-Hb⁡[λ2])-P′·(HbO2⁡[λ1]-Hb⁡[λ1])·100equation⁢⁢(1)wherein HbO2[λ1] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ1, HbO2[λ2] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ2, Hb[λ1] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ1, Hb[λ2] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ2, and P′ indicates a ratio (Pλ2/Pλ1) of a photoacoustic signal intensity Pλ2 obtained by irradiation with light of the wavelength λ2 to a photoacoustic signal intensity Pλ1 obtained by irradiation with light of the wavelength λ1. 2. A phantom, comprising: a light absorber in a phantom base material, whereinthe light absorber approximates a light absorption coefficient ratio at two or more wavelengths to that of biological materials, andacoustic propagation characteristics of the phantom base material and the light absorber are different from each other,wherein the light absorber contains two or more kinds of light absorbing compounds in a light absorber base material, absorption coefficient ratios μ[λ2]/μ[λ1] of the light absorbing compounds at arbitrary two wavelengths λ1 and λ2 (λ1&lt;λ2) of 600 nm or more and 1100 nm or less are different from each other. 3. The phantom according to claim 2, wherein the light absorber contains at least one light absorbing compound in which the absorption coefficient ratio μ[λ2]/μ[λ1] satisfies one of the following equation (2) or equation (3),Smin≥(μ⁡[λ2]/μ⁡[λ1])·Hb⁡[λ1]-Hb⁡[λ2](HbO2⁡[λ2]-Hb⁡[λ2])-(μ⁡[λ2]/μ⁡[λ1])·(HbO2⁡[λ1]-Hb⁡[λ1])·100Equation⁢⁢(2)Smax≤(μ⁡[λ2]/μ⁡[λ1])·Hb⁡[λ1]-Hb⁡[λ2](HbO2⁡[λ2]-Hb⁡[λ2])-(μ⁡[λ2]/μ⁡[λ1])·(HbO2⁡[λ1]-Hb⁡[λ1])·100Equation⁢⁢(3)wherein Smin indicates a lower limit of the parameter S, Smax indicates an upper limit of the parameter S, HbO2[λ1] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ1, HbO2[λ2] indicates an absorption coefficient of oxyhemoglobin at the wavelength λ2, Hb[λ1] indicates an absorption coefficient of deoxyhemoglobin at the wavelength λ1, Hb[λ2] indicates an absorption coefficient of deo... | Canon Kabushiki Kaisha | A61B005/00 A61B008/00
US09349235 | B2 | 2016-05-24 | Preauthorized wearable biometric device, system and method for use thereof | 2015-11-16 | Embodiments are directed towards authenticating users using biometric devices. The biometric device may be arranged to capture one or more biometric feature of a user that may be wearing the biometric device such as biometric features that correspond to an electrocardiogram of the user. The user of the biometric device may be authenticated based on one or more biometric features, or a combination thereof. Authenticating the user of the biometric device, may include communicating information that includes biometric features to an authorized authentication device (AAD). When the user is authenticated, the biometric device may be preauthorized for the user. When the preauthorized biometric device senses at least one access point, an authorization signal may be provided to the access point. If the preauthorized biometric device is removed from the user, the biometric device is deauthorized, disabling access to access points by the user. | 1. A method for authenticating a user using a device that performs actions including: capturing information that identifies the user of the device;authenticating the user of the device based on the information, wherein the authentication includes: communicating the information to an authorized authentication device (AAD) that is separated from the device; andcomparing the information to a profile that corresponds to the user and is stored on the AAD, wherein the AAD preauthorizes the device for the user when the profile stored on the AAD corresponds to the information; andproviding the preauthorization to one or more access points, wherein the user is provided access to the one or more access points based on the preauthorization of the device. 2. The method of claim 1, wherein capturing the information, further comprises, capturing at least one or more biometric features that correspond to an electrocardiogram of the user. 3. The method of claim 1, wherein the device, further comprises: an electronic continuity detector based on a circuit within the device that encircles at least a portion of a user's wrist; andwhen a clasp of the device is opened, the device is deauthorized for the user. 4. The method of claim 1, further comprises, providing a user interface that includes, one or more of a keyboard, a liquid crystal display (LCD), a speaker, a microphone, mouse, stylus, one or more physical, or one or more electronic buttons that a user employs to communicate with the device. 5. The method of claim 1, wherein providing access to the one or more access points, further comprises, authorizing the user to access the access point based on one or more of a password, a Personal Identification Number (PIN), a gesture, a voice command, a finger tap, a distance between the preauthorized device and the one or more access points, or one or more additional information that identifies the user. 6. The method of claim 1, wherein preauthorizing the device for the user, further comprises, providing a list of one or more access points that the user is permitted to access based at least on profile information that is associated with the user. 7. The method of claim 1, further comprising, when two or more access points are sensed by the preauthorized device, providing the user access to one or more of the two or more access points based on at least one secondary condition. 8. The method of claim 1, further comprising: enabling an administrative user to provide profile information for a plurality of users of the device;when one of the plurality of users is wearing the device, authenticating that user based on captured information that identifies that user; andproviding access to the one or more access points based at least on the provided profile information for the authenticated one of the plurality of users. 9. A device for authenticating a user, comprising: a transceiver that communicates over a network;a memory that stores at least instructions; anda pro... | Nymi Inc. | A61B008/00 G05B019/00 A61B005/02 G06F021/00 G07C009/00
US10252063 | B2 | 2019-04-09 | Leadless intra-cardiac medical device with built-in telemetry system | 2016-01-27 | A leadless intra-cardiac medical device is configured to be implanted entirely within a heart of a patient. The device includes an intra-cardiac extension and a housing. The intra-cardiac extension includes a loop body having at least one loop segment retaining at least one coil group that is configured to one or both of receive and transmit radio frequency (RF) energy, wherein the loop body is configured to extend into a first chamber of the heart. The housing is in electrical communication within the loop body, and includes a transceiver, control logic and an energy source. The housing is configured to be securely attached to an interior wall portion of a second chamber of the heart, wherein the transceiver is configured to communicate with an external device through the RF energy. | 1. A method of operating a leadless intra-cardiac medical device (LIMD) implanted within a heart of a patient, wherein the LIMD includes a housing,a loop body that is external to and connected to the housing, the loop body having at least one loop segment retaining one or more windings of insulated conductive wire within at least one of the superior vena cava (SVC) or the right atrium (RA) of the heart, andan insulated conductor that extends between the loop body and the housing,the housing secured within the right ventricle (RV) of the heart, andthe housing comprising a telemetry circuit electrically coupled to the loop body via the insulated conductor,wherein the method comprises: the loop body of the LIMD, which is external to the housing of the LIMD and has the at least one loop segment retaining the one or more windings of insulated conductive wire within the at least one of the SVC or RA of the heart, receiving radio frequency (RF) energy from a non-implanted external device;the one or more windings of insulated conductive wire of the loop body retained within the at least one of the SVC or RA of the heart generating an induced current therein in response to the RF energy;the insulated conductor that extends between the loop body and the housing passing the induced current from the loop body retained within the at least one of the SVC or RA of the heart to the housing secured within the RV of the heart; andthe telemetry circuit within the housing demodulating a signal from the induced current, that is passed by the insulated conductor from the loop body retained within the at least one of the SVC or RA of the heart to the housing secured within the RV of the heart, to thereby enable communicating between the LIMD and the external device. 2. The method of claim 1, wherein the housing secured within the RV of the heart includes a rechargeable energy source of the LIMD, and the method further comprising recharging the rechargeable energy source of the LIMD using the induced current that is passed by the insulated conductor from the loop body retained within the at least one of the SVC or RA of the heart to the housing secured within the RV of the heart. 3. The method of claim 1, further comprising: contacting an internal wall of at least one of the SVC or RA of the heart with at least one electrode on a portion of the loop body; andusing electronics within the housing, which is secured within the RV of the heart, to provide one or both of sensing or stimulus through the at least one electrode. 4. The method of claim 1, further comprising calibrating the LIMD after implantation into the heart of the patient. 5. The method of claim 4, wherein the calibrating comprises adjusting drive frequencies of the at least one loop segment. 6. The method of claim 1, wherein the at least one loop segment has a resonant frequency, and the method further comprising tuning the at least one loop segment to thereby adjust the resonant frequency of the at lea... | PACESETTER, INC. | A61B005/0408 A61N001/05 A61B005/042 A61N001/375 A61N001/372 A61B005/00 A61N001/362 A61N001/378 A61N001/368 A61N001/39
US09327866 | B2 | 2016-05-03 | Carton holder with dispensing configurations, system, and method of use | 2013-01-30 | Systems and apparatuses directed to cartons with user selectable dispensing configurations. An example carton for dispensing a plurality of products stored therein includes a first removable feature configured to be selectively removed from a top surface to enable removal of a portion of a first end; a second removable feature configured to be selectively removed from the top surface to enable removal of a portion of a second end; and a third removable feature configured to be selectively removed from the top surface to enable removal of a portion of the top surface. The carton defines a closed configuration and a dispensing configuration that is defined by at least one of the portion of the first end being removed, the portion of the second end being removed, or the portion of the top surface being removed. Methods for manufacturing a carton with user selectable dispensing configurations are also contemplated. | 1. A carton for dispensing a plurality of products stored therein, the carton defining a top surface, a bottom surface, a front surface, a back surface, a first end, and a second end, the carton comprising: a first removable feature associated with the top surface and the front surface and positioned proximate the first end of the carton, wherein the first removable feature is configured to be selectively removed from the top surface and the front surface to enable removal of a first portion of the first end, a first portion of the top surface, and a first portion of the front surface, wherein a first tab, a first portion of the bottom surface and a first portion of the back surface remains after removal of the first portion of the first end, the first portion of the top surface, and the first portion of the front surface;a second removable feature associated with the top surface and the front surface and positioned proximate the second end of the carton, wherein the second removable feature is configured to be selectively removed from the top surface and the front surface to enable removal of a first portion of the second end, a second portion of the top surface, and a second portion of the front surface, wherein a second tab, a second portion of the bottom surface, and a second portion of the back surface remains after removal of the first portion of the second end, the second portion of the top surface, and the second portion of the front surface;a third removable feature associated with the top surface and positioned between the first removable feature and the second removable feature, wherein the third removable feature is configured to be selectively removed from the top surface to enable removal of a third portion of the top surface, wherein at least a portion of the first removable feature and at least a portion of the second removable feature are each positioned adjacently to the third removable feature; andwherein the carton defines a closed configuration and a dispensing configuration, wherein the dispensing configuration is defined by at least one of: the first portion of the first end, the first portion of the top surface, and the first portion of the front surface being removed;the first portion of the second end, the second portion of the top surface, and the second portion of the front surface being removed; orthe third portion of the top surface being removed,wherein, when in the dispensing configuration, the entire back surface and bottom surface of the carton remain to define a support for the plurality of products when the carton is in the dispensing configuration. 2. The carton according to claim 1, wherein the dispensing configuration is defined by the first portion of the first end being removed to reveal a first portion of each product for dispensing. 3. The carton according to claim 2, wherein the first tab extends upwardly from the first portion of the bottom surface, wherein the tab is configured to encourage the prod... | WELCH ALLYN, INC. | B65D005/54 B65D005/63 B65D083/08 A61B005/022
US09339436 | B2 | 2016-05-17 | Automatic cardiac therapy advisor with hidden markov model processing | 2012-04-23 | A method of automatically determining which type of treatment is most appropriate for (or the physiological state of) a patient. The method comprises transforming one or more time domain measurements from the patient into frequency domain data representative of the frequency content of the time domain measurements; processing the frequency domain data to form a plurality of spectral bands, the content of a spectral band representing the frequency content of the measurements within a frequency band; forming a weighted sum of the content of the spectral bands, with different weighting coefficients applied to at least some of the spectral bands; determining the type of treatment (or physiological state) based on the weighted sum. | 1. A method of automatically determining which of a plurality of possible types of cardiac interventions should be performed in treatment of a patient, the method comprising: storing first prior information representative of types of prior cardiac interventions performed on the patient;storing second prior information representative of the patient's reactions to the prior cardiac interventions;wherein the first and second prior information are other than heart rhythms or ECG waveforms; andprocessing the first and second prior information using a hidden Markov model to determine which of a plurality of possible further cardiac interventions should be performed. 2. The method of claim 1 further comprising; sensing the patient's reaction to the further cardiac intervention,storing first further information representative of the type of the further cardiac intervention;storing second further information representative of the patient's reaction to the further cardiac intervention,processing the first and second prior and first and second further information using a hidden Markov model to determine which of still further cardiac interventions should be performed in further treatment of the patient. 3. The method of claim 1 wherein the patient is a cardiac arrest victim. 4. The method of claim 2 wherein the patient is a cardiac arrest victim. 5. Apparatus for automatically determining which of a plurality of possible types of cardiac interventions should be performed in treatment of a patient, the apparatus comprising: a processor and associated memory for storing first prior information representative of types of prior cardiac interventions performed on the patient;memory for storing second prior information representative of the patient's reactions to the prior cardiac interventions;wherein the first and second prior information are other than heart rhythms or ECG waveforms; andwherein the processor is configured to process the first and second prior information using a hidden Markov model to determine which of a plurality of possible further cardiac interventions should be performed. 6. The apparatus of claim 5 further comprising; components for sensing the patient's reaction to the further cardiac intervention,memory for storing first further information representative of the type of the further cardiac intervention;memory for storing second further information representative of the patient's reaction to the further cardiac intervention, andwherein the processor is further configured to process the first and second prior and first and second further information using a hidden Markov model to determine which of still further cardiac interventions should be performed in further treatment of the patient. 7. The apparatus of claim 5 wherein the patient is a cardiac arrest victim. 8. The apparatus of claim 6 wherein the patient is a cardiac arrest victim. | ZOLL Medical Corporation | A61B005/02 A61H031/00 A61B005/0464 A61N001/39 G09B023/28 A61B005/00
US09357288 | B2 | 2016-05-31 | Earhealth monitoring system and method IV | 2013-05-02 | A system for monitoring sound pressure levels at the ear includes an ambient sound microphone (ASM) for receiving ambient sounds and an ear canal microphone (ECM) for producing audio signals as a function of ambient sound received at the ambient sound microphone and a sound signal received from an associated personal audio device. A logic circuit is operatively associated with the ASM and ECM calculates a total SPL_Dose experienced by the ear at a time t. The total SPL_Dose is calculated by determining SPL_Dose for periods Δt as measured at the ECM. The logic circuit may select an action parameter in response to the Total SPL_Dose. | 1. A system configured to monitor a sound pressure level dose at an ear comprising: an ear canal microphone configured to measure a first acoustic sound pressure level (SPL1) within an ear canal of the ear;an ambient microphone configured to measure a second sound pressure level (SPL2); anda logic circuit operatively coupled to the ear canal microphone, where the logic circuit is configured to: calculate a noise reduction rating (NRR) using the SPL1 and the SPL2;compare the NPR to a threshold value of the NRR; andsend a signal to indicate a sealing integrity compromise if the NPR is less than the threshold value of the NRR. 2. The system of claim 1, wherein the logic circuit calculates a Total SPL_Dose as a function of an estimated SPL_Dose during a time period Δt. 3. The system of claim 2, wherein the estimated SPL_Dose is calculated as: SPL_DoseECM+ASM=SPL_DoseECM+ASM−1+Time of Sound Exposure/Time100%, where Time 100%=24 hrs/2^((LECM−T)/3), where LECM is a measured Ear Canal d/BA SPL by the ear canal microphone, and where T is a threshold value related to noise level. 4. The system of claim 2, further comprising a readable memory operatively connected to the logic circuit, at least one action parameter stored in the readable memory, the logic circuit comparing the Total SPL_Dose to a threshold value, and if the Total SPL_Dose is greater than the threshold value, then the logic circuit reads the action parameter from the readable memory. 5. The system of claim 4, wherein the logic circuit performs an action associated with the action parameter, where the action is at least one of modifying an operation of an audio device, modifying at least one acoustic signal directed to an ear canal receiver (ECR), and sending an acoustic notification signal to a user. 6. The system of claim 3, wherein the logic circuit determines whether the estimated Total SPL Dose exceeds a minimum threshold and reduces the Total SPL_Dose by selection of an action parameter. 7. The system of claim 6, wherein the minimum threshold is an effective quiet level. 8. The system of claim 1, wherein the ambient sound microphone measures an ambient SPL. 9. The system of claim 1, wherein the logic circuit compares the SPL1, to a threshold value and determines whether the SPL1 corresponds to a voice of a user of the system. 10. The system of claim wherein the logic circuit is configured to compare the SPL1 to the SPL2 and determine whether the SPL1 includes an acoustic signal corresponding to a voice of user. 11. The system of claim 1, wherein the logic circuit further estimates a third sound pressure level (SPL3) emitted by an ear canal receiver (ECR. 12. The system according to claim 11, wherein the logic circuit further calculates the total sound pressure level (SPL_Total), using the SPL1 and at least one of the SPL2 or the SPL3. 13. The system according to claim 12, wherein the logic circuit further: calculates a Time_100% for a time increment Δt during which the SPL_T... | Personics Holdings, LLC | H04R029/00 G01H003/14 A61B005/12 H04R001/10 A61B005/00 A61F011/14
US09339196 | B2 | 2016-05-17 | Non-invasive method and device of measuring the real-time continuous pressure of fluid in elastic tube and the dynamic compliance of elastic tube | 2013-03-14 | The invention presents a non-invasive method and device of measuring the real-time continuous pressure of fluid fluctuating in an elastic tube and the dynamic compliance of the elastic tube, in which the theory of VLDT (Vascular Loading Decoupling Technique) is used. After searching the initial critical depth and determining the decoupled ratio, a DC controller generates a DC control gain to maintain the elastic tube at critical depth, and an AC controller employs the self-adaptive and Step-Hold control rules to create the pulsation of elastic tube without effect of surrounding tissues, and be capable of measuring the real-time continuous fluid pressure and dynamic compliance of elastic tube. | 1. A non-invasive method of measuring a real-time continuous pressure of fluid in an elastic tube and a dynamic compliance of the elastic tube, wherein elastic surrounding tissues are wrapped outside the elastic tube, and wherein a fluctuating fluid flows in the elastic tube, the fluid having a fluid pressure which can be divided into a DC part Pb and an AC part ΔPb, the method comprising: using an oscillometric method, pressing a detection head by a DC-driven actuator on the surface of the surrounding tissues to locate an initial critical depth λs;holding at the initial critical depth to determine a decoupled ratio K for partial or full tracking of the AC part of fluid pressure ΔPb;actuating a DC-driven actuator to maintain a DC part of a sensor's pressure Ps as equal to the DC part of the fluid pressure Pb, and setting a resulting new depth as the location of an updated critical depth λs′;setting an AC-driven actuator, separate from the DC-driven actuator, in an idling (hold) stage to obtain three previous reference pressures Δ{circumflex over (P)}b′(n−3), Δ{circumflex over (P)}b′(n−2), and Δ{circumflex over (P)}b′(n−1), and estimating a reference pressure Δ{circumflex over (P)}b′(n) for an actuated (step) stage;using a constant AC open loop gain Ka and a self-adaptive control rule to estimate a parallel impedance H1 and compute an impedance of the elastic tube Zv2(n);using the constant AC open loop gain Ka and the parallel impedance H1 to calculate an AC control gain Ga for a next step stage;sending an AC control gain Ka from AC controller, and moving the AC-driven actuator to track the AC part of the fluid pressure ΔPb and obtain the AC part of a displacement Δλs′(n);computing the AC part of the fluid pressure ΔPb′(n) according to the impedance of the elastic tube Zv2(n) and the AC part of the displacement Δλs′(n);adding the DC part of the fluid pressure Pb′(n), which is equal to the DC part of the sensor's pressure Ps, to the AC part of the fluid pressure ΔPb′(n), to compute the real-time continuous fluid pressure Pb(n);extracting a series of equivalent mechanical properties including mass Mv(n), damping Dv(n), and stiffness Kv(n) from the dynamic impedance of elastic tube Zv2(n) by using a parameter identification technique, and computing the dynamic compliance of elastic tube Cv(n) using the reciprocal of the stiffness Kv(n). 2. The method of claim 1, wherein the prediction measure is a spline curve fitting technique. 3. The method of claim 1, wherein the elastic tube to be measured is an arterial blood vessel, the surrounding tissues are muscle tissues, and the fluid pressure is a blood pressure in the arterial blood vessel. 4. The method of claim 1, wherein the detection head includes the DC-driven actuator, the AC-driven actuator, a DC displacement sensor, an AC displacement sensor, and a pressure sensor, the AC-driven actuator disposed in the DC-driven actuator and independently movable up and down relative to the DC-driven actuator... | Gong Bu Design Company | A61B005/02 A61B005/021 A61B005/00
US09351677 | B2 | 2016-05-31 | Analyte sensor with increased reference capacity | 2013-03-04 | Systems and methods of use for continuous analyte measurement of a host's vascular system are provided. In some embodiments, a continuous glucose measurement system includes an electrochemical sensor incorporating a silver/silver chloride reference electrode, wherein a capacity of the reference electrode is controlled. | 1. An implantable electrochemical sensor for measuring an analyte concentration in a host, comprising: a working electrode configured to measure a concentration of an analyte; anda silver/silver chloride reference electrode, wherein at least a portion of the silver/silver chloride reference electrode is covered with an enzyme layer, wherein the enzyme layer is configured, in vivo, to generate hydrogen peroxide upon exposure to a substrate, whereby the hydrogen peroxide regenerates silver chloride of the reference electrode such that a reference capacity of the reference electrode is increased. 2. The implantable electrochemical sensor of claim 1, wherein the enzyme is an oxidase enzyme. 3. The implantable electrochemical sensor of claim 2, wherein the substrate is selected from the group consisting of glucose, urate, ascorbate, citrate, L-lactate, succinate, D-glucose, and ethanol. 4. The implantable electrochemical sensor of claim 1, wherein the analyte is glucose. 5. The implantable electrochemical sensor of claim 1, wherein the reference electrode comprises a chloridized elongated silver body. 6. The implantable electrochemical sensor of claim 1, further comprising a diffusion barrier configured to substantially block diffusion of hydrogen peroxide between the silver/silver chloride reference electrode and the working electrode. 7. The implantable electrochemical sensor of claim 1, wherein the enzyme layer has a thickness from about 0.01 microns to about 12 microns thick. 8. The implantable electrochemical sensor of claim 1, wherein a percentage of a surface area of the reference electrode covered with the enzyme layer is from about 10% to about 100%. 9. An electrochemical sensor for measuring an analyte concentration in a host, comprising: a working electrode configured to measure a concentration of an analyte; anda reference electrode, wherein at least a portion of the reference electrode is covered with an enzyme layer, wherein the enzyme layer is configured, in vivo, to generate a reference electrode regenerating species upon exposure to a substrate, whereby the reference electrode regenerating species regenerates a component of the reference electrode such that a reference capacity of the reference electrode is increased. 10. The electrochemical sensor of claim 9, wherein the enzyme is an oxidase enzyme. 11. The electrochemical sensor of claim 9, wherein the substrate is selected from the group consisting of glucose, urate, ascorbate, citrate, L-lactate, succinate, D-glucose, and ethanol. 12. The electrochemical sensor of claim 9, wherein the analyte is glucose. 13. The electrochemical sensor of claim 9, wherein the reference electrode comprises a chloridized elongated silver body. 14. The electrochemical sensor of claim 9, wherein the reference electrode regenerating species is hydrogen peroxide. 15. The electrochemical sensor of claim 9, wherein the enzyme layer has a thickness from about 0.01 microns to about 12 microns t... | DexCom, Inc. | G01N033/483 G01N033/50 A61B005/1486 A61B005/145
US20160143572 | A1 | 2016-05-26 | INCONTINENCE MONITORING DEVICE | 2015-06-05 | An incontinence monitoring device has a face separated into substantially parallel upper and lower surfaces having a flexible material held therebetween. The device includes a means of registering excessive moisture levels with respect to at least one of the surfaces and a system for conveying information regarding the registering of excessive moisture levels to remote users. | 1. An incontinence monitoring device comprising a face separated into substantially parallel upper and lower surfaces having a flexible material held therebetween, a means of registering excessive moisture levels with respect to at least one of the surfaces and a system for conveying information regarding the registering of excessive moisture levels to remote users. 2. A device according to claim 1 comprised in a padded cushion. 3. A device according to claim 1 comprised in an insert for a padded cushion. 4-12. (canceled) 13. A device, comprising: a face separated into substantially parallel upper and lower surfaces;means of detecting a presence of fluid with respect to at least one of the surfaces; andmeans for conveying information regarding the registering of excessive moisture levels to remote users. 14. The device according to claim 13, wherein the means of detecting a presence of fluid with respect to at least one of the surfaces comprises a foil circuit mounted on a flexible sheet, the foil circuit being in communication with the least one of the surfaces such that when the fluid contacts the least one of the surfaces the foil circuit is completed or breached to provide an electrical output that is different from an output of the foil circuit when the fluid is not present. 15. The device according to claim 14, further comprising a microprocessor in electrical communication with the foil circuit, the microprocessor sensing a change in the electrical conductivity of the foil circuit so as to detect a completion or breach of the foil circuit caused by the presence of the fluid. 16. The device according to claim 13, further comprising a tab that is in fluid communication with the upper surface, the tab providing a visual indication when the fluid contacts the upper surface. 17. The device according to claim 16, wherein the tab comprises a hydrochromic, opacifier, or solvent that provide the visual indication. 18. The device according to claim 15, further comprising means for activating the microprocessor. 19. The device according to claim 18, wherein the means for activating the microprocessor are based on presence, or lack thereof, of a user in contact with the device. 20. The device according to claim 13, wherein the means for conveying information regarding the registering of excessive moisture levels to remote users comprises a wired or wireless communication module that transmits a signal generated by the microprocessor that indicates that the fluid has contacted or permeated the upper surface. 21. The device according to claim 13, further comprising a padding material that supports the means of detecting a presence of fluid with respect to at least one of the surfaces. 22. The device according to claim 21, further comprising a module disposed below the padding material, the module comprising a microprocessor and a communication module. 23. The device according to claim 22, further comprising a pocket formed on an inside sur... |  | A61B005/20 G08C019/00 G08B005/22 A61B005/00 G08B021/18
US20160128639 | A1 | 2016-05-12 | GRAPHICAL TECHNIQUE FOR DETECTING CONGESTIVE HEART FAILURE | 2015-10-07 | The invention provides a system to better detect the onset of congestive heart failure (CHF). The system consists of two primary components—the CHF Index Graph (CHFIG) and a body-worn sensor for measuring the necessary parameters—that correctly measure and analyze a set of parameters related to CHF. The CHFIG typically consists of 7 or 8 unique axes, each corresponding to a different CHF-related parameter. Each axis is normalized and scaled so that the plotted values trend outwards as a patient's trajectory towards CHF worsens. Titus, the area of a region constructed from the plotted values of the parameters increases as the values move further away from the center of the graph; this indicates the patient is trending towards CHF. Conversely, a healthy patient's parameters will be clustered relatively close to the center of the graph, creating a region with a much smaller area. | 1. A method for characterizing heart failure in a patient, comprising the following steps: generating the following first set of data parameters with a group of one or more sensors: fluid level, stroke volume, blood pressure, cardiac output, and SpO2, the first set of data parameters measured from the patient when the patient is in a relatively healthy state;transmitting the first set of data parameters from the group of one or more sensors to a display device;generating the following second set of data parameters with a group of one or more sensors: fluid level, stroke volume, blood pressure, cardiac output, and SpO2, the second set of data parameters measured from the patient after the first set of data parameters;transmitting the second set of data parameters from the group of one or more sensors to a display device;calculating a third set of data parameters from the first and second set of data parameters, the third set of data parameters indicating a change in a data parameter and calculated by subtracting a value from the first set of data parameters from a corresponding value in the second set of data parameters;determining a maximum value for the third set of data parameters for each of the data parameters;normalizing the third set of data parameters for each of the data parameters to fit on plotting axes by normalizing each data parameter in the third set of data parameters;generating a graphical image of a plot comprising: i) multiple axes, with each axis corresponding to a single, normalized change in a data parameter, and organized so that a value corresponding to the normalized change in the data parameter is further separated from the plot's origin as the patient's state of heart failure worsens; and ii) a region calculated from values plotted on each axis that increases in area as the patient's state of heart failure worsens;calculating the area of the region created from the values plotted on each axis using a mathematical algorithm;generating a graphical image on a display device, the graphical image showing a plot comprising: i) two axes; and ii) a line created from the calculated area and organized such that the slope of the line increases as the patients state of heart failure worsens. 2. The method of claim 1, wherein the horizontal axis corresponds to the date that that the measurements were recorded from the one or more sensors. 3. The method of claim 1, wherein the vertical axis corresponds to calculated congestive heart failure index number. 4. The method of claim 1, wherein the formula for calculating the congestive heart failure index number comprises the following steps: breaking the figure down into a series of triangles wherein two adjacent axes create two sides, side A and side B, and the line connecting the points on those axes creates the third side, side C;determining the angle at the origin, angle α, by dividing 360 by the total number of axes;using the function: ½AB×sin(α), to determine the area of each of t... |  | A61B005/00 A61B005/0205
US10568569 | B2 | 2020-02-25 | Oral appliance for ventilation flow measurement | 2015-10-30 | A system and method for managing ventilation of a user, through the use of an oral appliance, is disclosed. The appliance is configured to fit between the user's teeth and lips or cheeks to inhibit ventilation by the user around the appliance, and includes a passageway. An airway resistor within the passageway includes an aperture that is sized to allow a metered amount of air flow by the ventilation of the user. A pressure sensor measures an air pressure in the aperture, which measurement is transmitted to an external computer processor to generate feedback data for the user. | 1. An apparatus for managing ventilation of a user, the apparatus comprising: a front plate having a curved peripheral portion with an upper edge and a lower edge, the curved peripheral portion being configured to fit between teeth of the user and lips or cheeks of the user to inhibit ventilation by the user around the curved peripheral portion;a post region that is substantially tubular and extends out of a central region of the front plate and in a direction toward inside a mouth of the user when worn, the post region positioned between the upper and lower edges of the front plate, bounded by the curved peripheral portion at one end, and extending substantially orthogonally to the central region of the front plate;a passageway extended through the front plate and the post region;an airway resistor slidingly and removably inserted into a post within the passageway, the post being an elongate member extended downward vertically and connected between opposing surfaces in the post region, the post being configured to engage with a slot in the airway resistor and configured as a stop against the slot to limit a distance of insertion of the airway resistor into the passageway in the direction toward inside the mouth of the user when worn, the airway resistor having an aperture that is sized to allow a metered amount of air flow by the ventilation of the user;one or more sensors associated with the airway resistor, the one or more sensors including a pressure sensor to measure an air pressure in the aperture of the metered amount of air flow by the ventilation of the user, the pressure sensor being located at a distal end of the post region;a micro controller connected with the pressure sensor to generate measurement data based on the air pressure measured by the pressure sensor and the metered amount of air flow, the micro controller further being configured to format the measurement data for transmission via a wireless transmission medium; anda transmitter connected with the micro controller to transmit the measurement data via the wireless transmission medium to an external computer processor. 2. The apparatus in accordance with claim 1, wherein the curved peripheral portion further comprises a flange configured to form a seal with the user's lips and be positioned adjacent the user's lips when the apparatus is inserted in the user's mouth, wherein the airway resistor is removably attached in the passageway. 3. The apparatus in accordance with claim 1, wherein the size of the aperture of the airway resistor is reconfigurable to provide different metered amounts of air flow; and the front plate further comprises an on-off switch to selectively provide power to the one or more sensors. 4. The apparatus in accordance with claim 1, wherein the one or more sensors further include a carbon dioxide sensor to measure carbon dioxide in the metered amount of air flow by the ventilation of the user. 5. The apparatus in accordance with claim 1, further com... | Connected Rock, Inc. | A61B005/08 A61B005/00 A63B023/18 A61B005/091
US20160143569 | A1 | 2016-05-26 | SYSTEM TO IMPROVE THE QUALITY OF THE BLOOD SAMPLE OBTAINED FROM INFANTS FEET AND DSIPOSABLE HEATABLE BOOTY THEREOF | 2016-02-04 | The present invention relates to methods used to improve a quantity and a quality of a blood sample obtained from a heel of a foot of a newborn baby, particularly as sample for the neonatal screening. | 1. A method to improve a quantity and a quality of a blood sample obtained from a heel of a foot of a newborn baby comprising: providing a disposable heatable booty, the heatable booty including an insert located on a bottom side of the booty and heatable seeds located over the inset, and a cover completely surrounding the insert and heatable seeds;warming the entire disposable heatable booty including the insert, heatable seeds, and cover in a warming device;placing the heated disposable booty in the foot of the newborn baby;removing the booty from the foot of the newborn baby;puncturing the heel of the newborn baby foot with a lancet;pressing the heel of the foot of the newborn baby;letting newborn blood sample drops from the heel of the newborn baby and be deposited on a filter paper;attaching the filter paper to a screening reader. 2. A disposable heatable booty designed to fit on a foot of a newborn baby comprising: a body, the body including an upper section, a lower section, a bottom section, and an interior section, the lower section of the body is designed to entirely cover the foot of the newborn baby, the booty is made of a disposable, heatable, and soft material selected from the group consisting of paper fabric, micro fleece, and polar fleece fabric, cotton fabric, vinyon fabric, flannel fabric,a cuff secured to the upper section of the body, the cuff is designed to cover an ankle of the newborn baby, the cuff is made of a stretchable and heatable fabric selected from the group consisting of an elastic fabric, an aramid fabric, a thermoplastic polyurethane fabric, or a vinyon fabric, the cuff adjusts to the ankle of the newborn baby;an insert made of disposable heatable fabric selected from the group consisting of a paper fabric, a micro fleece fabric, a polar fleece fabric, a cotton fabric, a vinyon fabric, or a flannel fabric, the insert is placed on the interior bottom section of the body;heatable seeds placed over the insert, the heatable seeds are selected from the group consisting of wheat seeds, lavender seeds, and linen seeds; anda cover completely surrounding the insert and the heatable, the cover being made of polar fleece fabric, cotton fabric, or flannel fabric. 3. The disposable heating booty according to claim 2, wherein the disposable heatable booty is made of a one piece fabric. 4. The disposable heating booty according to claim 2, wherein the body and the cuff are made of two different pieces of fabric sewn together. 5. The disposable heating booty according to claim 2, wherein the insert and the heatable seeds are replaced by a thermal gel capsule having heatable seeds inside. 6. A method to improve a quality and a quantity of blood as sample for the neonatal screening, obtained from a foot of a newborn baby comprising: a) providing a disposable heatable booty designed to fit the foot of the newborn baby and including: body, the body including an upper section, a lower section, a bottom section, and an interior... |  | A61B005/15 A43B003/30 A43B007/02 A61B005/151
US20160128629 | A1 | 2016-05-12 | APNEA DETECTION SYSTEM AND METHOD | 2015-11-11 | Provided is a method and system to predict a sleep apnea event based upon a sequence of sleep events and mitigate the deleterious effects such an event may have on a person. Some embodiments receive a sequence of sleep events from sensors taking physiological measurements of a person sleeping. These sleep events are compared against grouped sleep apnea events previously gathered from other sleepers with a similar sleep architecture to predict if the person is likely to have a sleep apnea event within an apnea prediction window. If a similar pattern is found, the method sends a sleep apnea warning event responsive to the requested comparison indicating that the person sleeping is likely to experience an imminent sleep apnea event. Various types of sleep entrainment operations may be performed to rouse the person from the sleep to sufficiently to mitigate potential harm from the imminent sleep apnea event. | 1. A computer-implemented method of mitigating effects of a sleep apnea event for a person based upon a sequence of sleep events gathered during the person's sleep, comprising: receiving a sequence of sleep events from one or more sensor devices taking physiological measurements of a person during sleep;requesting a comparison of the sequence of sleep events from the person against grouped sleep apnea events gathered from other sleepers with a similar sleep architecture to predict if person is likely to have a sleep apnea event within an apnea prediction window;receiving a sleep apnea warning event responsive to the requested comparison indicating that the person associated with the sequence of sleep events is likely to experience an imminent sleep apnea event;performing one or more sleep entrainment operations to rouse the person from the sleep to sufficiently to mitigate potential harm from the imminent sleep apnea event. 2. The computer-implemented method of claim 1 wherein the sequence of sleep events are received continuously at regular time intervals. 3. The computer-implemented method of claim 1 further comprising: determining if at least one sleep event gathered from the person during sleep indicates that the person is entering a sleep apnea event;adding a sleep apnea marker into the sequence of sleep events gathered from the person during sleep if it is determined that the person is entering into the sleep apnea event; andsubmitting the sequence of sleep events from the person to be included with grouped sleep apnea events gathered from other sleepers with a similar sleep architecture useful in predicting a sleep apnea event. 4. The computer-implemented method of claim 3 wherein the sleep apnea marker identifies the sequence of sleep events as a precursor to a sleep apnea event. 5. The computer-implemented method of claim 3 wherein submitting the sequence of sleep events increases an apnea prediction confidence associated with predicting a sleep apnea event within an apnea prediction window. 6. The computer-implemented method of claim 1 wherein the one or more sensors includes: a headband strap configured to be positioned around the forehead region of the person during sleep embedded with one or more position sensors to indicate positioning of the headband on the forehead, at least one electroencephalogram (EEG) sensor to measure brainwaves of the person;a first downward extension from the headband strap that positions an electrooculagraphy (EOG) sensor near an eye to detect eye movement and an electromyography (EMG) sensor near muscles of the person's face to detect muscle activity;a second downward extension from the headband strap that positions a carbon dioxide (CO2) sensor near the person's mouth to measure exhaled gasses; anda wireless communications module powered by a portable power supply that transmits the sequence of sleep events wirelessly to a remote computer device. 7. The method of claim 1 wherein requesting the comp... |  | A61B005/00 A61B005/097 A61B005/0488 A61B005/08 A61B005/0476 A61B005/0496
US20160128638 | A1 | 2016-05-12 | SYSTEM AND METHOD FOR DETECTING AND QUANTIFYING DEVIATIONS FROM PHYSIOLOGICAL SIGNALS NORMALITY | 2014-11-10 | A system for detecting and quantifying deviations from physiological signals normality and methods for making and using same. Each subject physiology follows unique patterns. The physiological signals can be affected by one or more factors such as circadian rhythm, disease and/or external stressors. Deviations of physiological signals from the normality of a subject can be indicative of external events that might require proper lifestyle management or just in time interventions, such as being exposed to high stress or the progress/onset of specific disease conditions. The disclosed system advantageously can quantify such deviations. | 1. A method for detecting a deviation from physiological normality, comprising: acquiring physiological data and at least one of physical data and environmental data from a sensor associated with a subject device;extracting contextual data from the acquired physical data and environmental data;analyzing the physiological data with relevant contextual data to form an individualized physiological model;comparing current physiological data and relevant current contextual data with the physiological model;identifying the deviation based upon said comparing; andreporting the deviation via the subject device. 2. The method of claim 1, wherein said analyzing the physiological data with the relevant contextual data includes generating the physiological model as a probability density function based upon historical physiological data and relevant historical contextual data, the probability density function establishing a normal physiological baseline for each of a plurality of selected contexts. 3. The method of claim 2, wherein said acquiring physiological data includes acquiring a plurality of different physiological data types; andwherein said generating the physiological model as the probability density function comprises generating a composite probability density function for a selected combination of the physiological data types. 4. The method of claim 3, wherein said extracting contextual data includes extracting a plurality of different contextual data types; andwherein said generating the composite probability density function comprises generating the composite probability density function for each of the contextual data types. 5. The method of claim 2, wherein said acquiring physiological data includes acquiring a plurality of different physiological data types; andwherein said generating the physiological model as the probability density function comprises generating a probability density function for each of the physiological data types. 6. The method of claim 2, wherein said extracting contextual data includes extracting a plurality of different contextual data types; andwherein said generating the physiological model as the probability density function comprises generating a probability density function for each of the contextual data types. 7. The method of claim 2, wherein said generating the physiological model includes establishing an average value of the physiological data for each of the selected contexts; andwherein said comparing comprises determining whether the current physiological data is within a predetermined range of the average value for the selected context corresponding to the relevant current contextual data. 8. The method of claim 7, wherein said generating the physiological model includes establishing a standard deviation value of the physiological data for each of the selected contexts; andwherein said determining whether the current physiological data is within the predetermined range comprises determining whethe... |  | A61B005/00 A61B005/08 A61B005/053 A61B005/0205
US10194821 | B2 | 2019-02-05 | Medical device having automated ECG feature extraction | 2015-10-29 | A medical device having automated electrocardiogram (ECG) feature extraction is disclosed. The medical device includes input circuitry configured to receive an ECG signal. Processing circuitry coupled to the input circuitry is configured to identify at least one fiducial point of heartbeat signature of the ECG signal. The processing circuitry is further configured to perform substantially simultaneously both a discrete wavelet transform (DWT) and a curve length transform (CLT) to identify the at least one fiducial point. | 1. A method for a medical device having automated electrocardiogram (ECG) feature extraction comprising: receiving by way of input circuitry an ECG signal generated by heartbeats;processing the ECG signal by substantially simultaneously performing in parallel both a discrete wavelet transform (DWT) process that generates DWT data associated with a T wave and a P wave of the ECG signal and a curve length transform (CLT) process that generates CLT data associated with a QRS complex of the ECG signal by way of processing circuitry coupled to the input circuitry;detecting an R peak of the QRS complex from the CLT data by identifying an amplitude within the CLT data that is greater than a QRS threshold updated as a portion of a sum of a previous QRS threshold associated with a previous one of the heartbeats and an average of CLT data generated from the previous one of the heartbeats; anddetecting a T wave peak and a P wave peak of the ECG signal from the DWT data by identifying within a forward search window following the QRS complex an amplitude that is greater than a T threshold updated as an average of the ECG signal within the forward search window and by identifying within a backward search window before the QRS complex an amplitude that is greater than a P threshold updated as an average of the ECG signal within the backward search window. 2. The method for the medical device having automated ECG feature extraction of claim 1 further including adjusting locations of boundaries of the forward search window and the backward search window relative to the QRS complex and an RR interval associated with the previous heartbeat. 3. The method for the medical device having automated ECG feature extraction of claim 1 further including extracting parameters that include PQ interval variability, QP interval variability, RT interval variability, TR interval variability, PS interval variability, and SP interval variability automatically by way of feature extraction circuitry coupled to the processing circuitry. 4. The method for the medical device having automated ECG feature extraction of claim 1 further including filtering the ECG signal by way of the input circuitry to substantially reduce noise having frequencies lower and higher than at least one wave within the ECG signal that includes the QRS complex. 5. A medical device having automated electrocardiogram (ECG) feature extraction comprising: input circuitry configured to receive an ECG signal generated from heartbeats; andprocessing circuitry configured to: execute a curve length transform (CLT) that generates CLT data associated with a QRS complex of the ECG signal;execute a discrete wavelet transform (DWT) that generates DWT data associated with a T wave and a P wave of the ECG signal, wherein the CLT and DWT are executed in parallel at substantially the same time;detect an R peak of the QRS complex from the CLT data by identifying an amplitude within the CLT data that is greater than a QRS thres... | Khalifa University of Science and Technology | A61B005/0456 A61B005/0468 A61B005/0428 A61B005/0245 A61B005/00 A61B005/04 A61B005/0464 A61B005/024 A61B005/0402
US09355448 | B2 | 2016-05-31 | Method and apparatus for image registration | 2013-12-20 | An image registering method includes extracting a first cross-section image from first volume data about a target object, which is acquired via a first image acquisition modality; displaying the first cross-section image; acquiring second volume data about the target object, the second volume data including a second cross-section image corresponding to the first cross-section image and position information of a probe, via a second image acquisition modality using the probe; and registering the first volume data with the second volume data, based on the first cross-section image, the second cross-section image, and the position information of the probe. | 1. An image registering method comprising: extracting a first cross-section image from first volume data about a target object based on information about an anatomical structure stored in correspondence to the target object, wherein the first volume data is acquired via a first image acquisition modality;displaying the first cross-section image;acquiring second volume data about the target object, the second volume data including a second cross-section image corresponding to the first cross-section image and position information of a probe, via a second image acquisition modality using the probe; andregistering the first volume data with the second volume data, based on the first cross-section image, the second cross-section image, and the position information of the probe. 2. The image registering method of claim 1, further comprising mapping information about an anatomical structure of a certain region in the target object with the target object and storing a result of the mapping. 3. The image registering method of claim 1, wherein the target object is selected by receiving an input of selecting the target object from a user. 4. The image registering method of claim 1, wherein the acquiring of the second volume data comprises: displaying the second cross-section image; anddisplaying a degree of correlation between the first cross-section image and the second cross-section image. 5. The image registering method of claim 4, wherein the acquiring of the second volume data comprises, when the degree of correlation is a predetermined value or more, acquiring the second volume data based on the second cross-section image. 6. The image registering method of claim 1, wherein the acquiring of the second volume data comprises: acquiring a plurality of pieces of cross-section image data from the target object via the probe;acquiring information about a plurality of positions of the probe that correspond to the plurality of pieces of cross-section image data; andproducing the second volume data based on the plurality of pieces of cross-section image data and the information about the plurality of positions of the probe. 7. The image registering method of claim 1, wherein the first image acquisition modality comprises a computed tomography (CT) image acquisition modality, and the second image acquisition modality comprises an ultrasound image acquisition modality. 8. The image registering method of claim 1, wherein the registering of the first volume data with the second volume data comprises producing third volume data by registering the first volume data with the second volume data. 9. The image registering method of claim 8, further comprising: acquiring information about a current position of the probe;selecting a cross-section image corresponding to the information about the current position of the probe from the third volume data; anddisplaying the selected cross-section image. 10. An image registering method comprising: receiving an input o... | SAMSUNG MEDISON CO., LTD. | G06K009/00 G06T007/00 A61B008/08 A61B006/12 A61B006/00 A61B006/03 A61B008/00 A61B005/055 A61B019/00 A61B005/00
US09351675 | B2 | 2016-05-31 | Noninvasive multi-parameter patient monitor | 2014-12-02 | Embodiments of the present disclosure include a handheld multi-parameter patient monitor capable of determining multiple physiological parameters from the output of a light sensitive detector capable of detecting light attenuated by body tissue. For example, in an embodiment, the monitor is capable of advantageously and accurately displaying one or more of pulse rate, plethysmograph data, perfusion quality, signal confidence, and values of blood constituents in body tissue, including for example, arterial carbon monoxide saturation (“HbCO”), methemoglobin saturation (“HbMet”), total hemoglobin (“Hbt”), arterial oxygen saturation (“SpO2”), fractional arterial oxygen saturation (“SpaO2”), or the like. In an embodiment, the monitor displays a line associated with a patient wellness level. | 1. A method comprising: processing one or more signals with one or more signal processors, said signals responsive to attenuation of light by body tissue of a monitored patient;noninvasively, electronically determining measured values of oxygen saturation and arterial carbon monoxide saturation, based at least in part on the processed one or more signals;providing a first oxygen saturation threshold for the oxygen saturation based at least in part on at least one measured value of the arterial carbon monoxide saturation satisfying an arterial carbon monoxide saturation threshold;providing a second oxygen saturation threshold for the oxygen saturation based at least in part on the at least one measured value of the arterial carbon monoxide saturation not satisfying the arterial carbon monoxide saturation threshold; andactivating a caregiver alert associated with a display based at least in part on at least one measured value of the oxygen saturation satisfying at least one of the first oxygen saturation threshold or the second oxygen saturation threshold. 2. The method of claim 1, wherein the caregiver alert comprises at least one of an LED activation, an alarm message, and an audible alarm. 3. The method of claim 1, wherein the caregiver alert changes as measured values of the oxygen saturation move further from the at least one of the first oxygen saturation threshold or the second oxygen saturation threshold. 4. The method of claim 1, wherein activating the caregiver alert comprises emphasizing on the display the measured values of the oxygen saturation. 5. The method of claim 1, wherein activating the caregiver alert further comprises at least one of altering a size of measured values of the oxygen saturation on the display and alternately displaying and removing from the display the measured values of the oxygen saturation. 6. The method of claim 1, further comprising: displaying on the display the at least one measured value of the oxygen saturation; anddisplaying on the display a prediction indicator corresponding to the oxygen saturation, wherein the prediction indicator indicates a predicted trend of measured values of the oxygen saturation. 7. A patient monitor, comprising: a sensor port configured to receive one or more signals from a sensor, wherein the signals are responsive to attenuation of light by body tissue of a monitored patient; anda processor in communication with the sensor port and configured to: noninvasively, electronically determine measured values of oxygen saturation and arterial carbon monoxide saturation using the one or more signals received from the sensor;provide a first oxygen saturation threshold for the oxygen saturation based at least in part on at least one measured value of the arterial carbon monoxide saturation satisfying an arterial carbon monoxide saturation threshold;provide a second oxygen saturation threshold for the oxygen saturation based at least in part on the at least one measured value of ... | Cercacor Laboratories, Inc. | G08B023/00 A61B005/1455 G06F019/00 A61B005/0205 A61B005/145 A61B005/00 A61B005/026 A61B005/0295 A61B001/00 A61B005/024
US10413241 | B2 | 2019-09-17 | Apparatus, method and computer program for detecting physiological parameters | 2014-06-25 | An apparatus, method and computer program where the apparatus comprises: a plurality of sensors (3, 5) configured to detect a physiological parameter: wherein at least one first (5) sensor is configured to have a first sensitivity to the physiological parameter and at least one second sensor (3) is configured to have a second sensitivity to the physiological parameter; such that a parameter profile, comprising a plurality of measurements of the physiological parameter at different is provided by the apparatus. | 1. An apparatus comprising: a plurality of sensors configured to detect a physiological parameter:wherein at least one first sensor is configured to have a first sensitivity to the physiological parameter and at least one second sensor is configured to have a second sensitivity to the physiological parameter;a permeable layer located between the at least one first sensor and a user's skin, the permeable layer being configured to enable moisture exchange between the at least one first sensor and the user's skin;a parameter profile sensor, wherein the parameter profile sensor, with the at least one first sensor and the at least one second sensor, provides a parameter profile comprising a plurality of measurements of the physiological parameter at different sensitivities;wherein the at least one first sensor and the at least one second sensor each provide a different response for a production of the physiological parameter; andwherein the at least one first sensor and the at least one second sensor are arranged in a stack and overlay each other. 2. An apparatus as claimed in claim 1 wherein the parameter profile provided using the apparatus is configured to be compared to one or more known parameter profiles of the parameter to enable the physiological parameter to be quantified. 3. An apparatus as claimed in claim 1 wherein the apparatus is configured so that the at least one first sensor and the at least one second sensor are positioned at different distances from the user's skin. 4. An apparatus as claimed in claim wherein the apparatus comprises at least one proximity sensor configured to detect the distance between the user's skin and at least one of the plurality of sensors. 5. An apparatus as claimed in claim 1 wherein the apparatus comprises a temperature sensor. 6. An apparatus as claimed in claim 1, wherein at least one third sensor is configured to have a third sensitivity to the physiological parameter. 7. An apparatus as claimed in claim wherein the at least one first sensor comprises a first sensing material and the at least one second sensor comprises a second different sensing material. 8. An apparatus as claimed in claim 1 wherein at least one of the plurality of sensors comprises a two dimensional material. 9. An apparatus as claimed in claim 1 wherein at least one of the plurality of sensors comprises at least one of graphene oxide, graphene, functionalised graphene, boron nitride, and molybdenite. 10. An apparatus as claimed in claim 1 wherein the physiological parameter gives an indication of sweat gland activity of the user. 11. An apparatus as claimed in claim 1 wherein the apparatus is configured to be worn by a user. 12. A method comprising: detecting a physiological parameter using a plurality of sensors, wherein at least one first sensor is configured to have a first sensitivity to the physiological parameter and at least one second sensor is configured to have a second sensitivity to the physiological parameter... | Nokia Technologies Oy | A61B005/01 A61B005/00
US10166071 | B2 | 2019-01-01 | Multi frequency and multi polarity complex impedance measurements to assess ablation lesions | 2016-01-15 | A method of assessing a tissue ablation treatment, including positioning a medical device adjacent a target tissue; measuring a first impedance magnitude a first frequency with the medical device; measuring a first impedance phase at a second frequency with the medical device; ablating at least a portion of the target tissue with the medical device; measuring at second impedance magnitude at a third frequency with the medical device; measuring a second impedance phase at a fourth frequency with the medical device; comparing at least one of (i) the first and second impedance magnitudes and (ii) the first and second impedance phases; and providing an indication of the efficacy of the ablation treatment based at least in part on the comparison. | 1. A medical system, comprising: a medical device positionable proximate a tissue site, the medical device comprising a plurality of electrodes and a patient return electrode operable to treat the tissue site; anda control unit in communication with the medical device, the control unit programmed to: deliver non-ablative radiofrequency energy at a first plurality of frequencies to the target tissue in a unipolar mode and a bipolar mode, wherein in the unipolar mode, the non-ablative radiofrequency energy is delivered between one electrode from the plurality of electrodes and the patient return electrode and in the bipolar mode, the non-ablative radiofrequency energy is delivered between two electrodes of the plurality of electrodes;receive baseline impedance characteristics at a first plurality of frequencies in both the unipolar and bipolar modes;deliver radiofrequency energy to the plurality of electrodes to ablate the target tissue;deliver non-ablative radiofrequency energy at a second plurality of frequencies to the target tissue in both the unipolar and bipolar modes;receive post-treatment impedance characteristics at the second plurality of frequencies in both the unipolar and bipolar modes;perform a comparison between the post-treatment impedance characteristics and the baseline impedance characteristics; andgenerate an indication of the efficacy of the ablation of the target tissue based at least in part on the comparison between the post-treatment and the baseline impedance characteristics. 2. The system of claim 1, wherein the baseline and the post-treatment impedance characteristics each includes at least one impedance magnitude value, and the comparison includes calculating a percentage difference between the at least one impedance magnitude value of the baseline and post-treatment impedance characteristics. 3. The system of claim 1, wherein the baseline and the post-treatment impedance characteristics each includes at least one impedance phase value, and the control unit is programmed to calculate a percentage difference between the at least one impedance phase value of the baseline and post-treatment impedance characteristics. 4. The system of claim 1, wherein the first plurality of frequencies are substantially the same as the second plurality of frequencies. 5. A medical system for assessing a tissue ablation treatment, comprising: a medical device positionable adjacent a target tissue, the medical device comprising a plurality of electrodes and a patient return electrode operable to treat the tissue site; anda control unit in communication with the medical device, the control unit programmed to: deliver non-ablative radiofrequency energy at a first frequency and a second frequency to the target tissue in both unipolar and bipolar modes, wherein in the unipolar mode, the non-ablative radiofrequency energy is delivered between one electrode from the plurality of electrodes and the patient return electrode and in the bipolar mod... | Medtronic Ablation Frontiers LLC | A61B018/14 A61B018/02 A61B018/12 A61B005/00 A61B017/00 A61B018/00
US09345427 | B2 | 2016-05-24 | Method of using a combination vein contrast enhancer and bar code scanning device | 2013-06-05 | It is known in the art to use an apparatus to enhance the visual appearance of the veins and arteries in a patient to facilitate insertion of needles into those veins and arteries. This application discloses a number of inventions that add additional data collection and presentation capabilities to a handheld vein enhancement apparatus and a set of processes for the collection of blood and the delivery of IV medicines that use the handheld device to mediate the process. | 1. A method of facilitating a venipuncture procedure for a patient, using a handheld vein enhancing device, said method comprising: storing in a memory, a list of identification data for one or more patients, and storing corresponding required venipuncture procedure data; and associating said data for each of the one or more patients with a corresponding bar code on a wrist band;identifying the patient to receive the venipuncture procedure by using a scanner of said device for scanning of the bar coded patient wrist band using a first wavelength of light, and comparing, by a microprocessor in said device, said scanned bar code to each of the corresponding patient bar codes stored in said memory until said scanned bar code matches one of said stored patient bar codes;displaying, on a display means of said device, said data for said required venipuncture procedure for said imaged bar code;gathering equipment needed to perform the required venipuncture procedure, in accordance with said required venipuncture procedure data;locating one or more subcutaneous veins in a target area on the patient, by scanning said first wavelength of light in a pattern onto the target area, thereby creating a contrasted image of the subcutaneous veins of the target area, resulting from differential absorption and reflection between the subcutaneous veins and the surrounding tissue therein, and by an image capture means receiving said contrasted image and outputting a signal representing said contrasted image to said second laser;receiving, by said second laser, said signal output by said image capture means; andemitting, by said second laser, said contrasted image using a second wavelength of light, and scanning said contrasted image onto the target area to overlay the subcutaneous veins therein, for locating a suitable vein for the required venipuncture procedure. 2. The method of performing a venipuncture according to claim 1, further comprising: storing a list, in said memory, of one or more valid users having required training and privileges to operate said device; and verifying access privileges to said device for a current user, by using an identification means therein to compare said current user to said list of valid users; and permitting use of said device when said current user is a valid user. 3. The method of performing a venipuncture procedure according to claim 2, wherein said required venipuncture procedure data comprises a quantity of blood to be removed; and wherein said required venipuncture procedure comprises removing blood by filling one or more vials to equal said quantity. 4. The method of performing a venipuncture procedure according to claim 2, wherein said required venipuncture procedure comprises injecting a medicine; and wherein said required venipuncture procedure data comprises a type of medicine to be injected. | AccuVein, Inc. | A61B006/00 A61B005/15 A61B005/00 A61B005/153 A61M005/42 A61B005/117 A61B005/024 A61B019/00
US09351653 | B1 | 2016-05-31 | Multi-channel reconfigurable systems and methods for sensing biopotential signals | 2012-11-29 | A biopotential monitoring device includes a configurable receiver circuit having a plurality of channels for receiving a plurality of biopotential signals from a biological tissue via a plurality of inputs coupled with the electrodes, and each channel substantially removes a DC (direct current) offset from a corresponding one of the biopotential signals and then band-pass amplifies such corresponding biopotential signal at a configurable gain and particular frequency range based on frequency control signals. The device further includes a controller circuit for receiving commands for configuring frequency characteristics of each biopotential signal. The controller automatically generates the frequency control signals based on such commands and outputs such frequency control signals to the configurable receiver circuit. The controller outputs a representation of each biopotential signal to an analyzer device that is configured to analyze such biopotential signal. | 1. A device for monitoring biopotentials of biological tissue through a plurality of electrodes, the device comprising: a configurable receiver circuit having a plurality of channels for receiving a plurality of biopotential signals from a biological tissue via a plurality of inputs coupled with the electrodes, each channel being configured to substantially remove a DC (direct current) offset from a corresponding one of the plurality of biopotential signals and then band-pass amplify the corresponding one of the biopotential signals at a configurable particular frequency range based on one or more frequency control signals;a controller circuit operable to receive commands for configuring or reconfiguring one or more frequency characteristics pertaining to receiving each biopotential signal;an impedance measurement module for generating a capacitively-coupled AC (alternating current) current waveform, having an amplitude and frequency, that is applied to a particular input coupled to a particular channel to measure a voltage signal received at the particular channel in response to the AC current waveform applied to the particular input,wherein the controller is further configurable to program the impedance measurement module to generate the capacitively-coupled AC current waveform to be applied to the particular selected input so as to monitor the impedance of the corresponding one of the electrodes by calculating the impedance based on the measured voltage signal received at the particular channel,wherein the controller is further configurable to automatically generate the one or more frequency control signals based on the commands and output the one or more frequency control signals to the configurable receiver circuit,wherein the controller is further configurable to output a representation of each biopotential signal and a representation of the calculated impedance to an analyzer device that is configured to analyze the representation of each biopotential signal and the representation of the calculated impedance. 2. The device of claim 1, further comprising: a multiplexer (MUX) coupled to the channels of the configurable receiver circuit so as to output a selected one of the amplified biopotential signal based on a received channel select signal; andan analog to digital converter (ADC) for receiving the selected amplified biopotential signal from the MUX and converting the selected amplified biopotential signal to a digitized representation of the selected amplified biopotential signal, wherein the controller is further operable to output a channel selection signal to the MUX indicative of which channel the MUX is to output one of the biopotential signals to the ADC. 3. The device of claim 2, wherein the configurable receiver circuit comprises: a configurable band-pass amplifier in each channel for amplifying the corresponding biopotential signal, wherein the one or more frequency characteristics comprise a lower cutoff frequency and/or an u... | Intan Technologies, LLC | A61B005/04
US09345437 | B2 | 2016-05-24 | Method for obtaining a set of measured data relating to a breathing object of interest by using magnetic resonance technology, magnetic resonance system, computer program, and electronically readable data storage medium | 2013-04-19 | In a method for determining a respiratory phase from a series of measurements of a respiratory position, the respiratory phase is determined by a finite state machine, which processes a current measured respiratory position and at least one previously measured respiratory position, and assigns a current respiratory phase to the current respiratory position. For this purpose, the temporal spacing between two successive measurements of the respiratory position is small in relation to a respiratory cycle. This method allows for a continuous measurement of imaging measurement data to be influenced in a prospective manner in a determined respiratory phase. | 1. A method for magnetic resonance (MR) imaging, comprising: using a control computer to operate an MR image data scanner to obtain MR raw data from a subject situated in the MR image data scanner;using said control computer to operate said MR image data scanner, during the acquisition of said MR raw data, to detect respiratory movement of the subject by obtaining a plurality of navigator spirals that comprise MR navigator data representing a series of measurements of respiratory position and phase, including measurement of a current respiratory position, in a respiratory cycle of the subject, with a temporal spacing between successive measurements in said series being small in relation to said respiratory cycle;using said control computer to process said MR navigator data representing said measurements of said respiratory position in said series to assign a status value, among multiple status values respectively designating said phases, to each respiratory position that precedes said current respiratory position;using said control computer to retrieve a momentary status value for the phase for a respiratory position in said series that precedes said current respiratory position;when said series comprises more than two measurements, using said control computer to determine, as of a last change to said momentary status value, a momentarily valid maximum respiratory position or a momentarily valid minimum respiratory position in said series;also when said series comprises more than two measurements, using said control computer to make a comparison, dependent on the momentary status value of the phase of the current respiratory position, with said maximum respiratory position or with said minimum respiratory position, to obtain a comparison result;also when said series comprises more than two measurements, using said control computer to assign a current status value to the respiratory phase of said current respiratory position as occurring during inhalation or exhalation, dependent on said comparison result;when said series comprises two or fewer measurements, using said control computer to assign a current status value of “unknown” to said current respiratory position;using said control computer to apply a prospective gating algorithm to said MR raw data, with an acceptance window in said prospective gating algorithm being determined from said current status value of said current respiratory position; andusing said control computer to reconstruct MR image data from said raw data after applying said prospective gating algorithm thereto, and making the MR image data available in electronic form as a data file from said control computer. 2. The method according to claim 1, wherein the respiratory phase transitions from the momentary status value “unknown” to a current status value “inhalation” when the current measured respiratory position is less than or equal to the determined momentary maximum respiratory position, minus a threshold value provided... | Siemens Aktiengesellschaft | A61B005/055 A61B005/00 A61B005/08 G01R033/567 G01R033/48
US10197645 | B2 | 2019-02-05 | Multi-channel endorectal coils and interface devices therefor | 2015-10-05 | An intracavity probe for use with a magnetic resonance system includes: a pair of coil loops arranged in a phased array configuration; a pair of decoupling circuits; a pair of output cables; and a spacer material positioned adjacent to an anterior surface of the coil loops. Each coil loop has a drive capacitor and a tuning capacitor. Each decoupling circuit is connected across the tuning capacitor of one of the coil loops. Each output cable is connected at a first end thereof across the drive capacitor of one of the coil loops such that each of the drive capacitors is provided with a separate ground. The spacer material assures a predetermined distance between the pair of coil loops and the region of interest, which thereby reduces intensity of the magnetic resonance signals in proximity to the coil loops, maintains SNR at depth within the region of interest and reduces artifacts. | 1. An interface device for interfacing a coil with a magnetic resonance system, the coil comprising a pair of coil loops arranged in a phased array configuration each of which for receiving magnetic resonance signals from a region of interest corresponding thereto, the interface device comprising: (a) a first preamplifier for receiving a signal from a first coil loop of the pair of coil loops to produce a first amplified signal;(b) a second preamplifier for receiving a signal from a second coil loop of the pair of coil loops to produce a second amplified signal;(c) a first splitter operatively connected to the first preamplifier for dividing the first amplified signal into a right loop signal that is provided to a first channel output and a first composite signal;(d) a second splitter operatively connected to the second preamplifier for dividing the first amplified signal into a left loop signal that is provided to a second channel output and a second composite signal;(e) a third splitter operatively connected to the first splitter for dividing the first composite signal;(f) a fourth splitter operatively connected to the second splitter for dividing the second composite signal;(g) a zero degree combiner operatively connected to the third splitter and the fourth splitter for combining signals received therefrom to produce a saddle signal that is provided to a third channel output; and(h) a 180 degree combiner operatively connected to the third splitter and the fourth splitter for combining signals received therefrom to produce a whole loop signal that is provided to a fourth channel output,wherein the interface device is configured to selectively recognize each of the first, second, third, and fourth channel outputs, thereby allowing the magnetic resonance system coupled to the interface device to produce images in a plurality of different modes. 2. The interface device of claim 1, wherein the first preamplifier and the second preamplifier are provided with a predetermined reduced supply voltage as compared to a rated supply voltage of the first preamplifier and the second preamplifier. 3. The interface device of claim 1, further comprising at least one attenuator providing an attenuation nominally in the range of 3 dB to 6 dB, the at least one attenuator positioned at at least one of (a) between the first preamplifier and the first splitter; (b) between the second preamplifier and the second splitter; (c) after the first splitter; and (d) after the second splitter. 4. The interface device of claim 1, wherein the plurality of modes include: Left Loop; Right Loop; Whole Loop; Whole Saddle; Right Loop and Left Loop (LL); Whole Loop and Whole Saddle; and Right Loop, Left Loop, Whole Loop, and Whole Saddle (LLLS). 5. A system for obtaining images of a region of interest, the system comprising: (a) an intracavity probe comprising:(i) a pair of coil loops arranged in a phased array configuration each of which receives magnetic resonance signals from... | Bayer HealthCare LLC | G01R033/3415 G01R033/34 G01R033/36 A61B005/055
US09339198 | B2 | 2016-05-17 | Headphones measuring heart-rate | 2012-11-06 | The invention relates to headphones and a method for listening to music or some other audio signal, which headphones comprise at least one sound source such as a an earphone to be placed in the vicinity of the ear, for producing sound, and a galvanic conductor for feeding an electrical signal from a signal source, such as a music player or similar, to the sound source. In accordance with the invention at least one sound source is equipped with means for measuring heart rate, as well as means for forwarding heart-rate information wirelessly. | 1. Headphones for listening to music or some other audio signal, which comprise at least one sound source configured to be placed in the vicinity of the ear, for producing sound, anda galvanic conductor for feeding an electrical signal from a signal source to the sound source, said sound source being equipped with means for measuring heart-rate, as well as means for forwarding heart-rate information wirelessly, andmeans for starting heart-rate measurememt once an audio signal fed to the headphones is detected. 2. Headphones according to claim 1, wherein the heart-rate measurement is implemented using infrared technology. 3. Headphones according to claim 1, wherein the heart-rate measurement is implemented from the outer auditory canal using infrared technology. 4. Headphones according to claim 1, wherein power supply to the measurement and transmission means is implemented using a dry-cell battery or rechargeable battery. 5. Method in headphones, comprising the steps of: feeding a signal from a signal source to a sound source, with the aid of galvanic conductors,equipping the sound source with means for measuring heart-rate, as well as means for forwarding heart-rate information wirelessly, andstarting heart-rate measurememt once an audio signal fed to the headphones is detected. 6. Method according to claim 5, wherein the heart-rate measurement is implemented using infrared technology. 7. Method according to claim 5, wherein the heart-rate measurement is implemented from the outer auditory canal using infrared technology. 8. Method according to claim 5, wherein power supply to the measurement and transmission means is implemented using a dry-cell battery or rechargeable battery. | Suunto Oy | A61B005/024 H04R001/10 A61B005/00
US09339236 | B2 | 2016-05-17 | Continuous transdermal monitoring system and method | 2014-12-09 | Various embodiments of methods and systems for continuous transdermal monitoring (“CTM”) are disclosed. One exemplary embodiment of a continuous transdermal monitoring system comprises a sensor package. The sensor package may include a pulse oximetry sensor having a plurality of light detectors arranged as an array. One exemplary method for continuous transdermal monitoring begins by positioning a pulse oximetry sensor system, similar to the system described immediately above, adjacent to a target tissue segment. Then, the method continues by detecting a light reflected by the target tissue segment. Then, the method continues by transmitting a pulse oximetry reading(s), based at least in part on the light reflected by the target tissue segment, of the target tissue segment. Then, the method continues by analyzing the pulse oximetry reading(s). Then, the method continues by assessing the accuracy of the pulse oximetry reading from the first light detector relative to the pulse oximetry reading from the second light detector. | 1. A system for continuous transdermal monitoring of a target tissue segment, the system comprising a sensor package comprising: a pulse oximetry sensor comprising a plurality of light detectors arranged as an array, each of the plurality of light detectors configured to: detect a light reflected by the target tissue segment; andtransmit a pulse oximetry reading of the target tissue segment, the pulse oximetry reading based, at least in part, on the light reflected by the target tissue segment;a shift sensor configured to detect movement related information about the plurality of light detectors;wherein the sensor package is configured to: designate, as a preferred light detector, a first light detector from the plurality of light detectors;output a pulse oximetry reading from the first light detector, based on the first light detector being designated as the preferred light detector;determine, based on the movement related information about detected by the shift sensor, that a second light detector from the plurality of light detectors is better positioned, relative to the target tissue segment, than the first light detector;designate, as the preferred light detector, the second light detector from the plurality of light detectors; andoutput a pulse oximetry reading from the second light detector, based on the second light detector being designated as the preferred light detector. 2. The system of claim 1, wherein the pulse oximetry sensor comprises a light source configured to emit light towards the target tissue segment. 3. The system of claim 1, wherein the shift sensor comprises an accelerometer. 4. The system of claim 3, wherein the accelerometer is one of the group consisting of a 3-axis accelerometer and a 6-axis accelerometer. 5. The system of claim 1, wherein the shift sensor is a displacement sensor configured to measure distance of each of the plurality of light detectors from the target tissue segment. 6. The system of claim 5, wherein the displacement sensor comprises an infrared sensor configured to visualize the location of a target blood vessel of the target tissue segment. 7. A method of continuous transdermal monitoring of a target tissue segment, the method comprising: designating, as a preferred light detector within a pulse oximetry sensor, a first light detector from a plurality of light detectors arranged as an array, each of the plurality of light detectors configured to detect a light reflected by the target tissue segment and transmit a pulse oximetry reading of the target tissue segment, the pulse oximetry reading based, at least in part, on the light reflected by the target tissue segment;outputting avulse oximetry reading from the first light detector, based on the first light detector being designated as the preferred light detector;determining, based on movement related information detected by a shift sensor configured to detect movement related information of the plurality of light detectors, that a second l... |  | A61B005/1455 A61B005/00 A61B005/0205 A61B005/024 A61B005/01 A61B005/021 A61B005/053 A61B005/11 A61B005/06 A61B005/145
US09336592 | B2 | 2016-05-10 | Method and apparatus for determining tumor shift during surgery using a stereo-optical three-dimensional surface-mapping system | 2013-02-01 | A system and method for determining intraoperative locations of a lesion in tissue from lesion locations determined in preoperative imaging includes determining three dimensional locations of surface features of the organ in the preoperative images. A preoperative surface map is extracted from stereo images annotated with surface features from preoperative images. An intraoperative surface map of the organ is extracted from stereo images, and surface features are identified in the stereo images corresponding to surface features annotated into the preoperative surface map. Three dimensional displacements of the surface features are determined and used to constrain a computer model of deformation of the organ. In embodiments, the model of deformation is adapted or constrained to model locations and dimensions of surgical cavities using an optical flow method and/or locations of surgical instruments in the organ. The model of deformation is used to determine intraoperative locations for the lesion. | 1. A method for determining a 3D model of a surface comprising: calibrating 3D reconstruction parameters for at least one reference setting of an optical system;calibrating image warping parameters for at least one secondary calibration settings, the image warping parameters adapted to control an image warping routine to warp images taken at that secondary calibration setting into warped images corresponding to images taken at the reference setting;taking an observed stereo image through the optical system with the optical system at a current setting;determining warping parameters from the image warping parameters for at least one secondary calibration setting of the at least one secondary calibration settings, the warping parameters for warping the observed stereo image taken through the optical system with the optical system at the current setting into a warped stereo image corresponding to a stereo image taken at the reference setting;warping the observed stereo image at the current setting into the warped stereo image corresponding to a stereo image taken at the reference setting;determining three-dimensional (3D) warping parameters for warping a first image of the warped stereo image corresponding to a stereo image taken at the reference setting into a second image of the warped stereo image corresponding to a stereo image taken at the reference setting; andusing the 3D warping parameters to determine the 3D model of the surface. 2. The method of claim 1 wherein there are image warping parameters for more than one secondary calibration setting stored in a table, and wherein the step of determining warping parameters from the image warping parameters for at least one secondary calibration setting includes interpolating between image warping parameters stored in the table. 3. The method of claim 2 wherein the step of determining 3D warping parameters uses vertically unconstrained fitting, and wherein pixels initially having vertical warp parameters exceeding a threshold are excluded from fitting the 3D warping parameters. 4. The method of claim 3 further comprising using the 3D model of the surface to constrain a mechanical model of tissue to determine a displacement of a particular structure in the tissue. 5. The method of claim 4 further comprising extracting a 3 dimensional representation of a surgical cavity, and wherein the mechanical model of tissue is adapted with the 3-dimensional model of the surgical cavity such that the cavity is taken into account in determining displacement of the particular structure in the tissue; wherein the 3-dimensional representation of a surgical cavity is determined by projecting a pre-resection 3D model of the surface onto a plane as a first projection, projecting a post-resection model of the surface onto the same plane as a second projection, determining an image-warping that maps the first projection into the second projection or the second projection into the first projection, this mapping is used... | THE TRUSTEES OF DARTMOUTH COLLEGE | G06K009/00 G06T007/00 G06T017/00 A61B006/00 A61B005/00 A61B005/107 G06K009/60 G06T019/20 A61B005/055 A61B019/00 A61B008/08
US10154506 | B2 | 2018-12-11 | MBAN channel management scheme using patient acuity information | 2014-06-19 | manage medical body area network (MBAN) channels to provide differentiated A system (120) and method (150) quality of service (QoS) based on patient acuity level. An available set of channels to be used by a plurality of MBAN systems (52) is identified (154). Channel access rules describing a set of channels for each of a plurality of MBAN priorities are generated (156). The plurality of MBAN priorities are based on patient acuity level. The channel access rules are communicated (158) to the plurality of MBAN systems (52). An MBAN system (52) of the plurality of MBAN systems (52) uses a channel from the set of channels corresponding to an MBAN priority assigned to the MBAN system (52). A hub device (60) configured to use the system (120) and method (150) is also provided. | 1. A medical body area network (MBAN) channel management system comprising: at least one controller comprising an electronic processor executing stored instructions to: identify an available set of channels to be used by a plurality of MBAN systems;generate channel access rules describing a set of channels for each of a plurality of MBAN priorities, the channels of the set of channels being channels of the available set of channels, the plurality of MBAN priorities corresponding to MBAN priorities assigned to the plurality of MBAN systems, and the MBAN priorities assigned to the plurality of MBAN systems based on patient acuity level;communicate the channel access rules to the plurality of MBAN systems, wherein an MBAN system of the plurality of MBAN systems uses a channel from the set of channels corresponding to the MBAN priority assigned to the MBAN system;receive reports from the plurality of MBAN systems describing changes in MBAN priority and changes in service area or service zone; andupdate at least one of the channels from a first MBAN priority to a second MBAN priority based on a change in the patient acuity level by, during the updating, increasing the number of channels in the sets of channels for the most prominent MBAN priorities, the most prominent MBAN priorities being those MBAN priorities most prominent in MBAN priority distributions of the plurality of MBAN systems. 2. The MBAN channel management system according to claim 1, wherein the sets of channels for the plurality of MBAN priorities differ based upon the plurality of MBAN priorities. 3. The MBAN channel management system according to claim 1, wherein the at least one controller is further configured to: generate the channel access rules for each of a plurality of service areas, or each of a plurality of service zones, of a medical institution within which the MBAN channel management system is used, the plurality of service areas corresponding to service areas of access points through which the plurality of MBAN systems communicate with the MBAN channel management system, and the plurality of service zones corresponding to collections of service areas;wherein the MBAN system uses a channel from the set of channels corresponding to the priority of the MBAN system and the service area or service zone within which the MBAN system is located. 4. The MBAN channel management system according to claim 1, wherein each of the plurality of MBAN systems is located in one of a plurality of service areas, or one of a plurality of service zones, of a medical institution within which the MBAN channel management system is used, the plurality of service areas corresponding to service areas of access points through which the plurality of MBAN systems communicate with the MBAN channel management system, and the plurality of service zones corresponding to collections of service areas sharing common role. 5. The MBAN channel management system according to claim 1, wherein the set of chann... | KONINKLIJKE PHILIPS N.V. | H04W072/10 H04W004/70 A61B005/00 H04W084/18 H04W074/00
US09877662 | B2 | 2018-01-30 | Vessel sensing device with automatic amendment function | 2015-06-22 | A vessel sensing device with automatic amendment function includes an analog processing circuit, a signal generating circuit and a DC voltage detecting circuit. The analog processing circuit includes an optical sensing module. An optical sensor of the optical sensing module generates an analog signal with a skin characteristic signal and a vessel characteristic signal according to an optical reflecting signal. The skin characteristic signal and the vessel characteristic signal respectively correspond to a skin feature and a vessel feature of the user. The signal generating circuit is coupled to the analog processing circuit to transform the analog signal into a compensation signal. The DC voltage detecting circuit is coupled to the signal generating circuit and adapted to compensate the analog processing circuit according to the compensation signal, so as to decrease the skin characteristic signal within the analog signal. | 1. A vessel sensing device with automatic amendment function, comprising an analog processing circuit, the analog processing circuit comprising an optical sensing module, an optical sensor of the optical sensing module generating an analog signal with a skin characteristic signal and a vessel characteristic signal according to an optical reflecting signal, wherein the skin characteristic signal corresponds to a skin feature of an user, and the vessel characteristic signal corresponds to a vessel feature of the user, the optical sensing module further comprising a switch electrically connected to the optical sensor;a signal generating circuit electrically connected to the analog processing circuit and the switch and adapted to transform the analog signal into a compensation signal; anda direct current (DC) voltage detecting circuit electrically connected to the signal generating circuit and adapted to compensate the analog processing circuit according to the compensation signal, so as to decrease the skin characteristic signal within the analog signal, the DC voltage detecting circuit controlling signal quantity through the switch by varying gate voltage of the switch so as to compensate the analog processing circuit. 2. The vessel sensing device of claim 1, wherein the DC voltage detecting circuit reduces the gate voltage to deduct a DC voltage value, so as to decrease the skin characteristic signal within the analog signal. 3. The vessel sensing device of claim 1, wherein the optical sensing module further comprises a gain module electrically connected between the optical sensing module and the signal generating circuit, the DC voltage detecting circuit controls gain result of the gain module applied to the analog signal by varying a compensating voltage value of a compensation capacity of the gain module, so as to compensate the analog processing circuit. 4. The vessel sensing device of claim 3, wherein the DC voltage detecting circuit utilizes the compensation capacity to reduce the gain result of the gain module applied to the analog signal, so as to decrease the skin characteristic signal within the analog signal. 5. The vessel sensing device of claim 1, wherein the switch is an N-type Metal oxide semiconductor field effect transistor, a P-type Metal oxide semiconductor field effect transistor or a transmission gate. 6. A vessel sensing device with automatic amendment function, comprising an analog processing circuit, the analog processing circuit comprising an optical sensing module, an optical sensor of the optical sensing module generating an analog signal with a skin characteristic signal and a vessel characteristic signal according to an optical reflecting signal, wherein the skin characteristic signal corresponds to a skin feature of an user, and the vessel characteristic signal corresponds to a vessel feature of the user, the optical sensing module further comprising a switch and a charge capacity;a signal generating circuit elec... | PixArt Imaging Inc. | A61B005/02 A61B005/024 A61B005/08 A61B005/00
US09532831 | B2 | 2017-01-03 | Systems and methods for activating transducers | 2016-01-19 | Transducer-based systems and methods may be configured to display a graphical representation of a transducer-based device, the graphical representation including graphical elements corresponding to transducers of the transducer-based device, and also including between graphical elements respectively associated with a set of the transducers and respectively associated with a region of space between the transducers of the transducer-based device. Selection of graphical elements and/or between graphical elements can cause activation of the set of transducers associated with the selected elements. Transducer activation characteristics, such as initiation time, activation duration, activation sequence, and energy delivery characteristics, can vary based on numerous factors. Visual characteristics of graphical elements and between graphical elements can change based on an activation-status of the corresponding transducers. Activation requests for a set of transducers can be denied if it is determined that a transducer in the set of transducers is unacceptable for activation. | 1. A method of ablating tissue with a catheter device system, the catheter device system comprising an energy source device system and a catheter device, the catheter device including a plurality of transducers, the plurality of transducers arranged in a distribution, the distribution positionable in a bodily cavity, the bodily cavity defined by at least one tissue wall, and the energy source device system connected to at least some of the plurality of transducers, the method comprising: activating each group of a plurality of groups of adjacent transducers in the distribution to cause energy from the energy source device system to be delivered to each of the transducers in each group of the plurality of groups of adjacent transducers in the distribution to form a series of ablated regions in the tissue wall, each of the ablated regions in the series positioned one after the other in spatial succession and corresponding to one of the plurality of groups of adjacent transducers, wherein each group of the plurality of groups of adjacent transducers in the distribution has a same transducer as another group of the plurality of groups of adjacent transducers in the distribution, and each group of the plurality of groups of adjacent transducers in the distribution is activated in a sequence that causes at least one of the ablated regions in the series to be formed in a region of the tissue wall that has not been previously ablated, the region in the tissue wall that has not been ablated being positioned between at least two previously formed ones of the ablated regions in the series. 2. The method of claim 1 wherein the at least one of the ablated regions in the series is spatially separated from at least one of the at least two previously formed ones of the ablated regions in the series. 3. The method of claim 1 wherein each of the ablated regions in the series is positioned one after the other in spatial succession to form a continuous ablated region. 4. The method of claim 3 wherein the continuous ablated region circumferentially encloses a region of the tissue wall. 5. The method of claim 3 wherein the continuous ablated region circumferentially encloses a port of a bodily opening that interrupts the tissue wall. 6. The method of claim 1 wherein each of the ablated regions in the series is adjacently positioned in the series between a respective pair of ablated regions in the series. 7. The method of claim 1 wherein the energy from the energy source device system delivered to each of the transducers in each group of the plurality of groups of adjacent transducers in the distribution is configured to cause bipolar tissue ablation. 8. The method of claim 1 wherein the energy from the energy source device system delivered to each of the transducers in each group of the plurality of groups of adjacent transducers in the distribution is configured to cause monopolar tissue ablation. 9. A catheter device system comprising: an energy source devic... | KARDIUM INC. | A61B018/00 A61B018/12 A61B018/14 A61B005/053 A61B005/00 A61B005/042 A61N001/372 G06F003/0482 G06F003/0484 A61B005/026 A61B005/01 A61N001/362
US10045738 | B2 | 2018-08-14 | Tissue resistance measurement | 2015-11-25 | The disclosure describes tissue resistance measurement techniques. To avoid errors caused by bandwidth limitations of the amplifier that amplifies a signal used to determine the tissue resistance, the disclosure describes determining values at different frequencies, where the values include respective error values. The respective error values are proportional to the respective frequencies, and based on this relationship the error value can be removed from the tissue resistance measurement. | 1. A method for tissue resistance measurement, the method comprising: generating, at an electrode, a first electrical signal of a first type through outputting, via the electrode, a first electrical signal of a second type at a first frequency;processing, at least in part with a bandwidth-limited amplifier, the first electrical signal of the first type at the first frequency to generate a first output signal, wherein processing the first electrical signal comprises amplifying, with the amplifier, the first electrical signal of the first type to generate a first intermediate electrical signal, and wherein a slew rate of the first intermediate electrical signal generated by the amplifier is less than a slew rate of a rectangular wave;generating, at the electrode, a second electrical signal of the first type through outputting, via the electrode, a second electrical signal of the second type at a second, different frequency;processing, at least in part with the bandwidth-limited amplifier, the second electrical signal of the first type at the second frequency to generate a second output signal, wherein processing the second electrical signal comprises amplifying, with the amplifier, the second electrical signal of the first type to generate a second intermediate electrical signal, and wherein a slew rate of the second intermediate electrical signal generated by the amplifier is less than the slew rate of the rectangular wave; anddetermining a tissue resistance at the electrode based on the first output signal, the second output signal, and a ratio between the first frequency and the second frequency, wherein the ratio includes a value based on a division of a value based on the first frequency and a value based on the second frequency. 2. The method of claim 1, wherein the first electrical signal of the first type comprises a first voltage signal, wherein the first electrical signal of the second type comprises a first current signal, wherein the second electrical signal of the first type comprises a second voltage signal, and wherein the second electrical signal of the second type comprises a second current signal. 3. The method of claim 1, wherein processing the first electrical signal of the first type comprises: chopping the first intermediate electrical signal at the first frequency to generate a first chopped signal; andaveraging the first chopped signal to generate the first output signal, wherein the first output signal is a first value, andwherein processing the second electrical signal of the first type comprises: chopping the second intermediate electrical signal at the second frequency to generate a second chopped signal; andaveraging the second chopped signal to generate the second output signal, wherein the second output signal is a second value. 4. The method of claim 3, wherein the amplifier comprises a transconductance amplifier, wherein amplifying the first electrical signal of the first type comprises amplifying the first elect... | Medtronic Bakken Research Center B.V. | A61B005/05 A61B005/00 A61B005/053 A61B005/04 A61N001/04 A61N001/36 A61B005/0478 A61N001/05
US10231667 | B2 | 2019-03-19 | Non-invasive dehydration monitoring | 2015-10-26 | The present invention relates to a device (10) and method for non-invasively measuring a hydration state of a living being (62). To improve with respect to the convenience-to-use, the cost-effectiveness and the accuracy of the delivered results, the proposed device comprises a first light source (14) for emitting light of a first wavelength into a tissue portion (12) of the living being (62), polarization means (18) for polarizing at least part of the emitted light prior to an interaction of the emitted light with the tissue portion (12), a first light detector (26) for detecting polarized light of the first wavelength reflected at the tissue portion (12) and processing means (32) configured to derive a tissue hydration parameter from the detected reflected polarized light and to determine a body hydration index based on said tissue hydration parameter. | 1. A device for measuring a hydration state of a living being, comprising: a first light source for emitting light of a first wavelength into a tissue portion of the living being;polarization means for polarizing at least part of the emitted light prior to an interaction of the emitted light with the tissue portion;a first light detector for detecting polarized light of the first wavelength reflected at the surface of the tissue portion, which has retained its polarization in the reflection;a second light detector for detecting non-polarized light of the first wavelength after an interaction of said light of the first wavelength with the tissue portion; andprocessing means configured to: compare an intensity of the detected reflected polarized light with a reference intensity,derive a tissue hydration parameter based on the comparison,derive a tissue color parameter and at least one vital sign parameter, wherein the at least one vital sign parameter includes a heart rate and/or a breathing rate, from the detected non-polarized light, anddetermine a body hydration index based on said tissue hydration parameter, said tissue color parameter and said vital sign parameter. 2. The device as claimed in claim 1, further comprising: a second light source for emitting light of a second wavelength into the tissue portion of the living being;the second light detector configured to detect non-polarized light of the first wavelength and non-polarized light of the second wavelength after an interaction of said light of the first wavelength and said light of the second wavelength with the tissue portion;wherein the processing means is configured to further derive the tissue color parameter and the at least one vital sign parameter, wherein the at least one vital sign parameter includes the heart rate, the breathing rate and/or a blood oxygen saturation, from the detected non-polarized light and to determine the body hydration index further based on said tissue color parameter and based on said at least one vital sign parameter. 3. The device as claimed in claim 2, wherein the processing means is configured to further determine a pulmonary parameter being indicative of a pulmonary problem of the living being based on the at least one vital sign parameter and a temperature parameter, wherein the temperature parameter is determined by a temperature sensor and is being indicative of a body temperature of the living being. 4. The device as claimed in claim 1, further comprising: a contact impedance sensor for determining an impedance parameter of the tissue portion;wherein the processing means is configured to derive the tissue hydration parameter further based on said impedance parameter. 5. The device as claimed in claim 1, further comprising: a temperature sensor for determining a temperature parameter being indicative of a body temperature of the living being;wherein the processing means is configured to determine the body hydration index further based on sai... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/024 A61B005/0205 A61B005/01 A61B005/08 A61B005/053 A61B005/087 A61B005/1455
US09597004 | B2 | 2017-03-21 | Wearable monitor | 2015-10-30 | The present disclosure relates to a wearable monitor device and methods and systems for using such a device. In certain embodiments, the wearable monitor records cardiac data from a mammal and extracts particular features of interest. These features are then transmitted and used to provide health-related information about the mammal. | 1. A wearable device for monitoring data, the device comprising: a sensor and circuit configured to detect or derive cardiac signals from a user, the circuit configured to estimate a heartbeat time series from the cardiac signals;a transmitter, the transmitter configured to transmit the heartbeat time series to a computing device when the user triggers the device based on a perceived cardiac event, the computing device configured to infer a likelihood of an occurrence of past cardiac arrhythmia from the heartbeat time series; andwherein the computing device is configured to provide a report, the report comprising the likelihood of the occurrence of past cardiac arrhythmia. 2. The device of claim 1, wherein the computing device is a server. 3. The device of claim 1, wherein the computing device is a smartphone. 4. The device of claim 2, wherein the computing device communicates with the transmitter through a smartphone intermediary. 5. The device of claim 1, wherein the report comprises indications for the presence of atrial fibrillation. 6. The device of claim 5, wherein the indications for the presence of atrial fibrillation comprise detecting at least 1000 beats. 7. The device of claim 1, wherein the computing device infers a plurality of most probable cardiac rhythms by filtering the heartbeat time series according to a predetermined threshold. 8. The device of claim 1, wherein the circuit is further configured to estimate a signal amplitude from the cardiac signals and the transmitter is further configured to transmit the signal amplitude to the computing device. 9. The device of claim 1, wherein the circuit is further configured to estimate a noise level from the cardiac signals and the transmitter is further configured to transmit the noise level to the computing device. 10. The device of claim 1, wherein the circuit is configured to collect a secondary signal and transmit the secondary signal to the computing device. 11. The device of claim 10, wherein the secondary signal is accelerometer data. 12. The device of claim 10, wherein the secondary signal is electrode contact quality data. 13. A wearable device for monitoring cardiac data, the device comprising: a sensor and circuit configured to record a plurality of heart beat signals from a user;a transmitter, the transmitter configured to transmit a segment of the heart beat signals to a server when the user uses a trigger, the server configured to analyze the segment and infer a likelihood of an occurrence of past cardiac arrhythmia; andwherein the server is configured to prepare a report from the likelihood of cardiac arrhythmia and transmit the report to the user, the report comprising the likelihood of the occurrence of past cardiac arrhythmia. | iRhythm Technologies, Inc. | A61B005/0408 A61B005/046 A61B005/00 A61B005/0205 A61B005/04 A61B005/0456 A61B007/04 A61B005/0432 A61B005/0468 A61B005/0472 G06F019/00 A61B005/024 A61B005/0404
US10561366 | B2 | 2020-02-18 | Measuring system and method for evaluating condition of patient | 2015-10-30 | A measuring system includes a first sensor, a second sensor, a first measurement value calculator that, based on a change of an amount of received light that is measured by the first sensor, calculates a first measurement value indicating an index related to a blood oxygen saturation or a blood refill time, a second measurement value calculator that, based on a change of an amount of received light that is measured by the second sensor, calculates a second measurement value indicating an index related to a blood oxygen saturation or a blood refill time, and a pressure controller that controls the first sensor and the second sensor so that a time difference between a start of compression release by the first sensor and a start of compression release by the second sensor is within a predetermined time period. | 1. A measuring system comprising: a first sensor configured to emit light to a first portion of a subject, to compress the first portion, and to measure reflected light or transmitted light from the first portion;a second sensor configured to emit light to a second portion of the subject, to compress the second portion, and to measure reflected light or transmitted light from the second portion;a first measurement value calculator configured to calculate, based on a change of an amount of received light that is measured by the first sensor, a first measurement value indicating a first blood refill time;a second measurement value calculator configured to calculate, based on a change of an amount of received light that is measured by the second sensor, a second measurement value indicating a second blood refill time; anda pressure controller configured to control the first sensor and the second sensor so that a time difference between a start of compression release by the first sensor and a start of compression release by the second sensor is within a predetermined time period. 2. The measuring system according to claim 1, wherein the pressure controller is further configured to control a first time period from the start of compression release by the first sensor to an end of compression, and a second time period from the start of compression release by the second sensor to an end of compression so that the first time period and the second time period are substantially equal to each other. 3. The measuring system according to claim 1 further comprising an evaluator configured to evaluate a cardiovascular system of the subject based on a comparison process using the first measurement value and the second measurement value. 4. The measuring system according to claim 3, wherein the evaluator is further configured to calculate: a difference value between the first measurement value and a threshold, and a difference value between the second measurement value and a threshold; ora difference value between the first measurement value and the second measurement value. 5. The measuring system according to claim 3, wherein the evaluator is further configured to calculate a first difference value between the first measurement value and a threshold, to calculate a second difference value between the second measurement value and the threshold, and to evaluate the cardiovascular system of the subject based on the first difference value and the second difference value. 6. The measuring system according to claim 5, wherein, in a case where both the first measurement value and the second measurement are smaller than the threshold, the evaluator is further configured to determine that the cardiovascular system of the subject is normal, and, in the other case, the evaluator is further configured to determine that the cardiovascular system of the subject is abnormal. 7. The measuring system according to claim 3, wherein the evaluator is further configured to calc... | NIHON KOHDEN CORPORATION | A61B005/1455 A61B005/026 A61B005/00
US20160128556 | A1 | 2016-05-12 | VIDEO LARYNGOSCOPE | 2015-11-11 | A video laryngoscope includes a handle, a blade, a lens, a photo sensor, and an output terminal. The blade is connected to the handle. The lens is disposed on the blade. The photo sensor is connected to the lens to transform an image obtained by the lens into an electrical signal. The output terminal is disposed on the handle and is electrically connected to the photo sensor. The output terminal is configured to allow a smart communication device to be connected thereto, such that the electrical signal can be output to the smart communication device through the output terminal. | 1. A video laryngoscope, comprising: a handle;a blade connected to the handle;a lens disposed on the blade;a photo sensor connected to the lens for transforming an image obtained by the lens into an electrical signal; andan output terminal disposed on the handle and being electrically connected to the photo sensor, wherein the output terminal is configured to allow a smart communication device to be connected thereto, such that the electrical signal can be output to the smart communication device through the output terminal. 2. The video laryngoscope of claim 1, wherein the handle comprises: a main body; andat least a position fixing feature disposed on the main body for temporarily fixing the smart communication device on the main body. 3. The video laryngoscope of claim 2, wherein the position fixing feature is a tool-free position fixing feature. 4. The video laryngoscope of claim 2, wherein the output terminal is connected to the smart communication device when the smart communication device is temporarily fixed on the main body by the position fixing feature. 5. The video laryngoscope of claim 2, wherein the position fixing feature is disposed on a terminal of the main body away from the blade. 6. The video laryngoscope of claim 1, wherein the output terminal is disposed on a terminal of the main body away from the blade. 7. The video laryngoscope of claim 1, wherein the photo sensor is accommodated in the handle. 8. The video laryngoscope of claim 1, further comprising: an optical fiber connecting the lens with the photo sensor, wherein the optical fiber is partially accommodated in the blade. 9. The video laryngoscope of claim 1, wherein the output terminal can be directly connected to the smart communication device. 10. The video laryngoscope of claim 1, further comprising: a connecting wire being capable of coupling the output terminal with the smart communication device. |  | A61B001/267 A61B001/00 A61B005/00 A61B001/05
US20160135981 | A1 | 2016-05-19 | FOOT ORTHOTIC FABRICATION SYSTEM AND METHODS OF OPERATING | 2015-11-16 | Method and system is disclosed for manufacturing an orthopedic insert. The process includes, initializing a plurality of sensors as pixel-based information, making one or more pressure measurement from each of the plurality of sensors, analyzing and conditioning the pressure measurements to derive a pressure value for each pressure sensor, converting the pressure values into a pixel intensity value, mapping the pixel intensity values to a graphic image, determining pixel intensity values between each mapped pixel intensity value along a first axis, determining pixel intensity values along a second axis, and machining an orthotic insert based upon the graphical image. | 1. A method of manufacturing an orthotic shoe insert, comprising: receiving foot pressure information from a plurality of pressure sensors;generating a graphical-based representation based upon the received foot pressure information;determining machining coordinates and machining depth based upon the graphical-based representation; andmachining the orthotic shoe insert based upon the determining. 2. The method of claim 1, where the machining is executed on only a single side of the orthotic shoe insert. 3. The method of claim 1, wherein the generating the graphical-based representation based upon the received foot pressure information comprises converting the received foot pressure information into pixel-based information. 4. The method of claim 3, wherein the pixel-based information includes color information and location information associated with the graphical-based representation. 5. The method of claim 4, wherein the graphical-based representation includes pixel cells comprising an area of pixels, wherein each pressure sensor of the plurality of pressure sensors has a unique pixel cell on the graphical-based representation. 6. The method of claim 5, further comprising: determining color information for pixels not associated with a location on the graphical-based representation that is associated with one of the plurality of pressure sensors based upon: (1) proximity to the location on the graphical-based representation that is associated with one of the plurality of pressure sensors and (2) the color that is associated with the unique location of the one of the plurality of pressure sensors. 7. The method of claim 6, wherein the proximity is determined based upon a single, predefined axis. 8. The method of claim 5, further comprising: iteratively, identifying a first pixel not associated with a location on the graphical-based representation that is associated with one of the plurality of pressure sensors along a first axis;identifying a second pixel associated with one of the plurality of pressure sensors along the first axis on a first side of the first pixel;identifying a third pixel associated with one of the plurality of pressure sensors along the first axis on a second side of the first pixel; anddetermining color information for the first pixel based upon proximity to the second pixel, proximity to the third pixel, color information associated with the second pixel, and color information associated with the third pixel. 9. The method of claim 8, further comprising: iteratively, identifying a pixel not having color information;identifying a closest upper pixel having color information and a closest lower pixel having color information, each upper and lower pixel along a second axis; anddetermining color information for the identified pixel not having color information based upon proximity to the upper pixel, proximity to the lower pixel, color information associated with the upper pixel, and color information associated with th... |  | A61F005/14 A61B005/103
US20160120472 | A1 | 2016-05-05 | Low Dissolution Rate Device and Method | 2015-10-30 | An implantable device includes a circuit protected with a low dissolution rate layer, wherein the circuit is either (a) fully encapsulated by the low dissolution rate layer and configured for non-electrical conduction contact sensing (e.g., capacitive sensing) or (b) partially encapsulated by the low dissolution rate layer with an electrode at least partially exposed outside the layer; wherein the implantable device is suitable for implantation inside the body of a living animal; and wherein the low dissolution rate layer comprises an element selected from the group consisting of gallium, boron, nitrogen, oxygen, zirconium, aluminum, and titanium. Such devices can be made by lithographic and other means, with coating layers applied by atomic layer deposition. | 1. An implantable device comprising: a circuit protected with a low dissolution rate layer, wherein the circuit is either (a) fully encapsulated by the low dissolution rate layer and configured to perform non-electrical conduction contact sensing, or (b) partially encapsulated by the low dissolution rate layer with an electrode at least partially exposed outside the layer;wherein the implantable device is suitable for implantation inside the body of a living animal; andwherein the low dissolution rate layer comprises at least one element selected from the group consisting of gallium, boron, nitrogen, oxygen, zirconium, aluminum, and titanium. 2. The device of claim 1, wherein the low dissolution rate layer comprises a material selected from the group consisting of gallium, boron, nitride, oxide, zirconium, aluminum, titanium, gallium nitride, boron nitride, zirconium oxide, zirconia oxide, diamond, aluminum oxide, titanium nitride, titanium carbide, titanium dioxide, and combinations thereof. 3. The device of claim 1, wherein the device has shear modulus of 200 to 1500 Pa. 4. The device of claim 1, wherein said low dissolution rate layer has a dissolution rate of less than 3 nm/day at 96.4° C. while in an in vivo environment. 5. The device of claim 1, wherein said low dissolution rate layer is in a state of having been deposited by atomic layer deposition. 6. The device of claim 1, wherein said device is less than 25 microns thick. 7. The device of claim 1, further comprising an electrically insulating layer beneath said low dissolution rate layer, and wherein said low dissolution rate layer is electrically conductive. 8. The device of claim 7, wherein said low dissolution rate layer comprises GaN, TiN, and/or TiO2 and is surrounded by an outer layer of AlN, each layer in a state of having been deposited by atomic layer deposition. 9. The device of claim 1, comprising circuit elements configured to perform wireless power transfer. 10. The device of claim 1, wherein said circuit includes at least one bond wire protected by said low dissolution rate layer. 11. The device of claim 1, wherein said low dissolution rate layer is surrounded by a polymeric coating. 12. The device of claim 1, wherein said polymeric coating is parylene. 13. The device of claim 1, having a rough outer surface. 14. An implantable device comprising: a circuit protected with a coating material comprising two layers in a state of having been deposited by atomic layer deposition, namely an inner insulation layer and an outer low dissolution rate layer in intimate contact therewith,wherein the circuit is either (a) fully encapsulated by both layers and configured to perform non-electrical conduction contact sensing, or (b) partially encapsulated by both layers with an electrode at least partially exposed outside the layers;wherein the implantable device is suitable for implantation inside the body of a living animal; andwherein the low dissolution rate layer is in a ... |  | A61B005/00 C23C016/455
US09589107 | B2 | 2017-03-07 | Monitoring treatment compliance using speech patterns passively captured from a patient environment | 2014-11-17 | Methods and systems are described for monitoring patient speech to determine compliance of the patient with a prescribed regimen for treating for a brain-related disorder. Patient speech is detected with an audio sensor at the patient location, and speech data is transmitted to a monitoring location. Patient speech is processed at the patient location and/or monitoring location to identify speech parameters and/or patterns that indicate whether the patient has complied with the prescribed treatment regimen. Patient identity may be determined through biometric identification or other authentication techniques. The system may provide a report to an interested party, for example a medical care provider, based on whether (and/or the extent to which) the patient has complied with the prescribed treatment regimen. The monitoring system may transmit a report to a wireless device such as a pager or cell phone, generate an alarm or notification, and/or store information for later use. | 1. A system comprising: at least one receiving device for use at a monitoring location for receiving a speech data signal transmitted to the monitoring location from a patient location, the speech data signal containing speech data, the speech data including patient speech data representing spontaneous speech sensed from a patient with at least one audio sensor at the patient location, and the patient having a brain-related disorder and a prescribed treatment regimen for treating at least one aspect of the brain-related disorder;speech identification circuitry configured to identify the patient speech data corresponding to speech from the patient in the speech data, the patient speech data including data indicative of at least one patient speech pattern;compliance determination circuitry configured to determine compliance of the patient with the prescribed treatment regimen based on whether the patient speech data includes data indicative of the at least one patient speech pattern matching at least one characteristic speech pattern; andreporting circuitry configured to report a conclusion based on the determination of whether the patient has complied with the prescribed treatment regimen. 2. The system of claim 1, further comprising control circuitry for controlling at least one of the at least one receiving device, the speech identification circuitry, the compliance determination circuitry, and the reporting circuitry. 3. The system of claim 1, wherein the speech identification circuitry includes: patient identification circuitry configured to determine a presence of the patient from at least one identity signal received at the monitoring location from the patient location;wherein the speech identification circuitry is configured to identify patient speech data corresponding to speech from the patient in the speech data based at least in part on the determination of the presence of the patient by the patient identification circuitry. 4. The system of claim 3, wherein the identity signal includes at least a portion of the speech data signal, wherein the patient identification circuitry is configured to analyze the speech data signal to determine the presence of the patient by identifying at least a portion of the speech data signal that resembles a known speech data signal of the patient, and wherein the speech identification circuitry is configured to identify patient speech data by identifying speech data corresponding to the presence of the patient. 5. The system of claim 3, wherein the identity signal includes an image signal received from an imaging device at the patient location, wherein the patient identification circuitry is configured to analyze the image signal to determine the presence of the patient, and wherein the speech identification circuitry is configured to identify patient speech data by identifying speech data corresponding to the presence of the patient detected from the image signal. 6. The system of claim 5, wherein t... | Elwha LLC | G10L015/00 G06F019/00 A61B005/00 G06Q050/24 A61B005/16
US09332544 | B2 | 2016-05-03 | Channel switching method in a medical body area network | 2012-03-29 | The disclosure of the present invention provides a method for a Medical Body Area Network (MBAN) to control, through a master, the channel switching of an MBAN terminal that uses a channel of a first frequency band in an MBAN system. The method includes: receiving information related to the use of each channel of a second frequency band from an MBAN controller; and transmitting information on a channel of the second frequency band that the MBAN terminal is to switch on the basis of the received information related to the use of each channel of the second frequency band. The first frequency band and the second frequency band are distinguished from each other based on whether the MBAN terminal uses a corresponding frequency band first. | 1. A method for channel switching in a wireless communication system that supports medical body area network (MBAN) services and operates in the 2360 MHz-2400 MHz frequency band, the method comprising: determining, by an MBAN coordinator, whether a primary user is scheduled to use an operating channel that is currently used by an MBAN master device and an MBAN device for communication between the MBAN master device and the MBAN device; andwhen the first MBAN coordinator determines that the primary user is scheduled to use the operating channel, transmitting, by the MBAN coordinator, first channel switching information to protect the primary user,wherein the first channel switching information is transmitted to the MBAN master device,wherein the first channel switching information includes a first remaining time field and a first channel number field,wherein the first remaining time field includes information on a first time when the channel switching occurs by the MBAN master device,wherein the first time is determined based on a transmission time of the first channel switching information,wherein the first channel number field includes information on a channel to be switched,wherein the channel switching information further includes a first channel switching method field,wherein the first channel switching method field includes information on a frequency band to be switched,wherein the frequency band includes the channel to be switched,wherein the MBAN master device transmits second channel switching information to the MBAN device,wherein the second channel switching information is generated based on the first channel switching information by the MBAN master device,wherein the second channel switching information includes a second remaining time field, a second channel number field, and a second channel switching method field,wherein the second remaining time field includes information on a second time when the channel switching occurs by the MBAN device,wherein the second time reaches 0 based on a transmission time of the second channel switching information,wherein the channel number field includes the information on the channel to be switched,wherein the channel switching method field includes information on the frequency band to be switched, andwherein the second channel switching information is transmitted to the MBAN device based on a broadcast method via a beacon frame by the MBAN master device. 2. A medical body area network (MBAN) coordinator for channel switching in a wireless communication system that supports MBAN services and operates in the 2360 MHz-2400 MHz frequency band, the MBAN coordinator comprising: a controller configured to determine whether a primary user is scheduled to use an operating channel that is currently used by an MBAN master device and an MBAN device for communication between the MBAN master device and the MBAN device; anda wireless communication unit configured to transmit first channel switching information ... | LG Electronics Inc. | H04W072/04 H04W036/06 H04B013/00 H04W028/04 H04W048/16 H04W072/00 H04W084/10 A61B005/00
US09336535 | B2 | 2016-05-10 | Neuro-response data synchronization | 2014-02-11 | An example system includes a headset to gather first data comprising first neuro-response data and second neuro-response data from a user while the user is exposed to stimulus material. In the example system, the headset comprises a first sensor to gather the first neuro-response data, the first neuro-response data comprising at least one of electroencephalographic data or magnetoencephalographic data, and a second sensor to gather the second neuro-response data, the second neuro-response data comprising facial emotion encoding data. The headset also comprises a processor to synchronize the first neuro-response data, the second neuro-response data and the stimulus material to generate synchronized data and determine an effectiveness of a portion of the stimulus material based on the synchronized data. | 1. A system comprising: a headset to gather data including first neuro-response data and second neuro-response data from a user while the user is exposed to stimulus material, the headset including: a band;a first sensor extending from the band to be located on a first position on a head of the user when the user is wearing the headset, the first sensor to gather the first neuro-response data, the first neuro-response data including at least one of electroencephalographic data or magnetoencephalographic data;a second sensor extending from the band to be located on a second position on the head of the user when the user is wearing the headset, the second sensor to gather second neuro-response data, the second neuro-response data including at least one of electroencephalographic data or magnetoencephalographic data, wherein when coupled to the band, the first sensor and the second sensor are rotatably positionable relative to the band, wherein rotation of the first sensor from the first position to a third position on the head of the user rotates the second sensor relative to the band from the second position to a fourth position on the head of the user when the user is wearing the headset; anda processor programmed to: synchronize the first neuro-response data, the second neuro-response data and the stimulus material to generate synchronized data; anddetermine an effectiveness of a portion of the stimulus material based on the synchronized data. 2. The system of claim 1, wherein the headset further includes an identifier to identify the stimulus material. 3. The system of claim 1, wherein the first sensor and the second sensor are electrodes and the processor is programmed to process signals from active ones of the electrodes for analysis. 4. The system of claim 1, wherein the first sensor and the second sensor are electrodes and the processor is programmed to process signals from a subset of the plurality of electrodes for analysis that are active channels. 5. The system of claim 1, wherein the headset is a first headset, the data is first data, and the user is a first user, further including a second headset to gather second data from a second user while the second user is exposed to the stimulus material. 6. The system of claim 5, further including a receiver to receive the first data from the first user and the second data from the second user, the effectiveness based on the first data and the second data. 7. The system of claim 1, further including a third sensor to gather eye-tracking data, wherein the processor is programmed to analyze the eye-tracking data to determine a direction of the attention of the user. 8. The system of claim 1, wherein the headset further includes a first rotatable hub and the first sensor includes a first electrode coupled to the first hub at a first distance. 9. The system of claim 8, wherein the second sensor includes a second electrode coupled to the first hub at a second distance, the second distance d... | The Nielsen Company (US), LLC | A61B005/04 G06Q030/02 A61B005/0476 A61B005/00 A61B005/0408 A61B005/0488 A61B005/0496
US20160143610 | A1 | 2016-05-26 | MEDICAL IMAGING DEVICE AND METHOD FOR OPERATING A MEDICAL IMAGING DEVICE | 2015-11-25 | A medical imaging device with a contactless data transmission device that registers interference fields acting thereon is provided. The device includes a bit error registration unit of the data transmission apparatus. The bit error registration unit is embodied and programmed to establish the bit error rate of the data transmission apparatus caused by the interference fields, and to output a warning signal when a predeterminable threshold of the bit error rate is exceeded. | 1. A medical imaging device comprising: a contactless data transmission apparatus configured to register acting interference fields, the contactless data transmission apparatus comprising: a bit error registration unit configured to: establish a bit error rate of the contactless data transmission apparatus caused by the acting interference fields; andoutput a warning signal when a predeterminable threshold of the bit error rate is exceeded. 2. The medical imaging device of claim 1, further comprising an indication unit configured to represent the warning signal optically, acoustically, or optically and acoustically. 3. The medical imaging device of claim 1, further comprising a switching-off unit configured to interrupt operation of the medical imaging device when the warning signal occurs. 4. The medical imaging device of claim 1, wherein the contactless data transmission device comprises a stationary unit and a rotating unit, between which data is transmitted contactlessly. 5. The medical imaging device of claim 1, wherein the medical imaging device is a computed tomography scanner or a magnetic resonance imaging scanner. 6. The medical imaging device of claim 2, further comprising a switching-off unit configured to interrupt operation of the medical imaging device when the warning signal occurs. 7. The medical imaging device of claim 2, wherein the contactless data transmission device comprises a stationary unit and a rotating unit, between which data is transmitted contactlessly. 8. The medical imaging device of claim 3, wherein the contactless data transmission device comprises a stationary unit and a rotating unit, between which data is transmitted contactlessly. 9. The medical imaging device of claim 2, wherein the medical imaging device is a computed tomography scanner or a magnetic resonance imaging scanner. 10. The medical imaging device of claim 3, wherein the medical imaging device is a computed tomography scanner or a magnetic resonance imaging scanner. 11. The medical imaging device of claim 4, wherein the medical imaging device is a computed tomography scanner or a magnetic resonance imaging scanner. 12. A method for operating a medical imaging device comprising a contactless data transmission, the method comprising: registering bit errors in the contactless data transmission caused by acting interference fields;establishing a bit error rate; andemitting a warning signal when a predeterminable threshold of the bit error rate is exceeded. 13. The method of claim 12, further comprising optically, acoustically, or optically and acoustically indicating the warning signal. 14. The method of claim 12, further comprising interrupting the operation of the medical imaging device when the warning signal occurs. 15. The method of claim 13, further comprising interrupting the operation of the medical imaging device when the warning signal occurs. 16. The method of claim 12, wherein the data transmission is performed between unit... |  | A61B006/00 A61B005/055 H04L029/08 A61B006/03 H04N001/00
US09339231 | B2 | 2016-05-17 | Heart failure management to avoid rehospitalization | 2014-03-04 | Systems and methods are described for subject rehospitalization management. In an example, multiple physiologic signals can be obtained from a subject using multiple sensors. In response to a hospitalization event, pre-hospitalization characteristics of the multiple physiologic signals can be identified. Post-hospitalization characteristics of the multiple physiologic signals can be identified, including characteristics that differ from their corresponding pre-hospitalization characteristics. Later subsequent physiologic signals can be further monitored after the hospitalization event, such as using the same multiple sensors, and subsequent physiologic signal characteristics can be identified. In an example, a heart failure diagnostic indication can be determined using information about the pre-hospitalization characteristics, the post-hospitalization characteristics, and the subsequent characteristics. Information about relative changes in signal characteristics from multiple sensors can be used to identify particular subject physiologic signals to monitor during subsequent periods. | 1. A system comprising: multiple physiologic sensors configured to sense respective physiologic signals from a subject; anda heart failure analysis module, coupled to the multiple physiologic sensors, the heart failure analysis module including a processor circuit configured to:identify a hospitalization event; andin response to the hospitalization event, identify pre-hospitalization physiologic signal characteristics corresponding to respective physiologic signals obtained using one or more of the physiologic sensors;identify one or more post-hospitalization physiologic signal characteristics that are different than their corresponding one or more pre-hospitalization physiologic signal characteristics, the one or more post-hospitalization physiologic signal characteristics obtained using the same respective one or more physiologic sensors as the pre-hospitalization physiologic signal characteristics;monitor a subsequent physiologic signal, after the hospitalization event, the subsequent physiologic signal corresponding to one of the one or more post-hospitalization physiologic signal characteristics that are different than their corresponding one or more pre-hospitalization physiologic signal characteristics, the subsequent physiologic signal obtained using the same one or more physiologic sensors as the pre-hospitalization and post-hospitalization physiologic signal characteristics;identify a subsequent physiologic signal characteristic using the subsequent physiologic signal; anddetermine a heart failure parameter for the subject using information about the subsequent physiologic signal characteristic relative to its corresponding pre-hospitalization physiologic signal characteristic. 2. The system of claim 1, wherein the processor circuit is configured to update a therapy parameter for the subject using the determined heart failure parameter. 3. The system of claim 1, wherein the processor circuit is configured to determine a parameter indicative of worsening heart failure when the subsequent physiologic signal characteristic trends toward its corresponding pre-hospitalization physiologic signal characteristic. 4. The system of claim 1, wherein the processor circuit is configured to identify the hospitalization event in response to a subject episode, and wherein the processor circuit is configured to identify pre-hospitalization physiologic signal characteristics including pre-episode physiologic signal characteristics corresponding respectively to the physiologic signals obtained using one or more of the physiologic sensors. 5. The system of claim 1, wherein the processor circuit is configured to identify the hospitalization event in response to a subject heart failure decompensation episode. 6. The system of claim 1, wherein the processor circuit is configured to provide a rehospitalization indication when the subsequent physiologic signal characteristic trends toward its corresponding pre-hospitalization physiologic signal characteris... | Cardiac Pacemakers, Inc. | A61N001/36 A61B005/00 A61B005/02 A61B007/04 A61B005/0205 A61N001/362 A61N001/365
US20160128488 | A1 | 2016-05-12 | APPARATUS AND METHODS FOR HEATING OR COOLING A BED BASED ON HUMAN BIOLOGICAL SIGNALS | 2015-12-15 | Introduced are methods and systems for an adjustable bed device configured to: gather biological signals associated with multiple users, such as heart rate, respiration rate, or temperature; analyze the gathered human biological signals; and heat or cool a bed based on the analysis. | 1. A method for automatically heating or cooling a mattress based on a biological signal associated with a user: obtaining, by a processor, said biological signal associated with said user, wherein said biological signal comprises a respiration rate associated with said user, a heart rate associated with said user, a motion associated with said user, or a temperature associated with said user;identifying, by said processor, said user based on at least one of: said heart rate associated with said user, said respiration rate associated with said user, said motion associated with said user, or said temperature associated with said user;based on said user identification, obtaining, by said processor, from a database a normal biological signal range associated with a sleep phase in a plurality of sleep phases associated with said user, wherein said normal biological signal range comprises a normal temperature range associated with said user;based on said normal biological signal range and said biological signal, identifying, by said processor, said sleep phase in said plurality of sleep phases associated with said user; andwhen said temperature associated with said sleep phase is outside of said normal temperature range associated with said sleep phase, sending, by said processor, a control signal to a temperature control device coupled to said mattress, said control signal comprising an instruction to heat or cool said mattress to a temperature within said normal temperature range. 2. A method of claim 1, wherein said obtaining said biological signal comprises: measuring, by a sensor strip coupled to said mattress, said biological signal associated with said user, wherein said sensor strip comprises a piezo sensor, and a temperature sensor. 3. A method of claim 1, wherein said obtaining said biological signal comprises: receiving, by said processor, said biological signal associated with said user from a wearable sensor coupled to said user. 4. A method of claim 1, wherein said sending said control signal to said temperature control device comprises: determining, by said processor, a current time;based on said user identification, obtaining, by said processor, a wake-up time associated with said user; andwhen said current time is at most 3 hours before said wake-up time, sending by said processor, said control signal to said temperature control device coupled to said mattress, said control signal comprising an instruction to turn off. 5. A method of claim 1, wherein said identifying said sleep phase comprises detecting, by said processor, a slowdown in said heart rate, a drop in said temperature, and a regular respiration rate. 6. A method of claim 1, wherein said identifying said sleep phase comprises detecting, by said processor, an end of a preceding sleep phase in said plurality of sleep phases. 7. A method of claim 1, further comprising: obtaining, by said processor, perspiration associated with said user; andwhen said user is perspiring, ... |  | A47C021/04 A61B005/024 A61B005/11 G05B019/042 A61B005/0205 A61B005/00 A47G009/02 G05D023/13 A61B005/08 A61B005/01
US09875581 | B2 | 2018-01-23 | Automated 3D reconstruction of the cardiac chambers from MRI or ultrasound | 2015-11-02 | The disclosure relates to a method of automatically producing a three-dimensional (3D) segmentation of a heart chamber, the method comprising: obtaining data sets from cardiac magnetic resonance imaging (MRI) or ultrasound, generating a 3D segmentation of the heart chamber from the data sets using an active contour method, modifying the 3D segmentation by adding a plurality of intra-chamber structures; and identifying an enclosing myocardium using the 3D segmentation generated by the method. | 1. A method of automatically producing a three-dimensional (3D) segmentation of a heart chamber, the method comprising: (a) obtaining data sets from cardiac magnetic resonance imaging (MRI) or ultrasound,(b) generating a 3D segmentation of the heart chamber from the data sets using an active contour method comprising minimizing an energy function, E(Φ), when a contour lies on a boundary of the heart chamber, wherein E(Φ) is defined as E(Φ)=Eint(Φ)+Eext(Φ),wherein Eint is the internal energy function and Eext is the external energy function of the heart chamber in a 3D domain,(c) modifying the 3D segmentation by adding a plurality of intra-chamber structures; and(d) identifying an enclosing myocardium using the 3D segmentation generated in step (b). 2. The method of claim 1, where minimizing the energy function, E(Φ) comprises using an external energy function, Eext(Φ), defined as Eext(Φ)=w2Ereg+w3Eedge+w4Egeom wherein Ereg is a region-based term, Eedge is an edge-based term, Egeom is a geometric term, and where w2, w3, and w4 are a plurality of weighting parameters. 3. The method of claim 2, further comprising normalizing the MRI or ultrasound data sets and reusing the same weighting parameters across the entire MRI or ultrasound data set. 4. The method of claim 1, where the MRI or ultrasound data sets comprise short-axis cardiac magnetic resonance images, long-axis cardiac magnetic resonance images, sagittal MRI images, coronal MRI images, axial MRI images, or any combination thereof. 5. The method of claim 1, where modifying the 3D segmentation by adding a plurality of cardiac substructures comprises: identifying a plurality of points on a convex hull of the 3D segmentation;computing a centroid for the plurality of points;calculating the radius and angle of the plurality of points on the convex hull with respect to the centroid to produce cylindrical coordinates for the plurality of points on the convex hull; andinterpolating the cylindrical coordinates to produce a closed convex curve which includes the plurality of cardiac substructures. 6. The method of claim 5, where identifying an enclosing myocardium using the 3D segmentation comprises removing a portion of endocardium of the cardiac structure from the 3D segmentation and refilling the 3D segmentation with a pattern representing the myocardium of the cardiac structure in its place as the distance from the centroid is increased. 7. The method of claim 1, where generating a 3D segmentation of the cardiac structure from the MRI or ultrasound data sets comprises simultaneously segmenting the MRI or ultrasound data sets and reconstructing 3D images therefrom. 8. The method of claim 1, wherein said heart chamber is selected from the group consisting of the left ventricle, the right ventricle, the left atrium and the right atrium. 9. The method of claim 1, wherein said modifying the 3D segmentation by adding a plurality of intra-chamber structures comprises adding papillary muscles to a ... | THE REGENTS OF THE UNIVERSITY OF CALIFORNIA | G06K009/00 G06K009/40 G06T019/20 A61B005/055 A61B005/00 A61B008/08 A61B008/00 G06T007/12 G06T007/149
US09345406 | B2 | 2016-05-24 | Occlusion-crossing devices, atherectomy devices, and imaging | 2012-11-13 | Described herein are methods for producing and identifying characteristic (“crescent shaped”) regions indicative of an atherectomy plaque within a vessel, and systems and devices adapted to take advantage of this characteristic region. | 1. A method of identifying a plaque mass within a vessel, the method comprising: inserting a device into the vessel, the device applying a force on the vessel to cause the vessel wall to stretch away from the device;imaging the vessel to create an image;identifying a location and orientation of a plurality of pointed crescent-shaped structures extending radially from a substantially semi-circular central portion in the image, the substantially semi-circular central portion having a substantially rounded portion and a substantially linear portion, each pointed crescent-shaped structure extending from a junction of the substantially rounded portion and the substantially linear portion; anddetermining a position of the plaque mass within the vessel relative to the device based upon the identified location and orientation of the pointed crescent-shaped structures in the image. 2. The method of claim 1, wherein imaging the vessel comprises imaging the vessel with optical coherence tomography. 3. The method of claim 1, wherein imaging the vessel comprises imaging with an imaging sensor attached to the device. 4. The method of claim 3, further comprising orienting the device within the vessel based on the location and orientation of the crescent-shaped structures. 5. The method of claim 4, wherein orienting the device comprises pointing a directional cutter of the device at the position of the plaque mass. 6. The method of claim 4, wherein orienting the device comprises directing the device based upon the relationship between markers in the image created by a portion of the device and the location and orientation of the crescent-shaped structures. 7. The method of claim 3, further comprising rotating the imaging sensor to obtain the image. 8. The method of claim 1, wherein identifying a location and orientation of a plurality of pointed crescent-shaped structures comprises determining which direction a tip of each crescent-shaped structure points. 9. The method of claim 8, wherein determining a position of the plaque mass comprises identifying a location between two pointed crescent-shaped structures in which the tips of the two pointed crescent-shaped structures point towards one another. | Avinger, Inc. | A61B005/05 A61B005/00 A61B017/3207 A61B005/02 A61B019/00 A61B006/12 A61B006/00 A61B017/22
US09330820 | B2 | 2016-05-03 | Method for making electrically conductive three-dimensional structures | 2013-04-30 | Methods are provided for fabricating three-dimensional electrically conductive structures. Three-dimensional electrically conductive microstructures are also provided. The method may include providing a mold having at least one microdepression which defines a three-dimensional structure; filling the microdepression of the mold with at least one substrate material; molding the at least one substrate material to form a substrate; and depositing and patterning of at least one electrically conductive layer either during the molding process or subsequent to the molding process to form an electrically conductive structure. In one embodiment, the three-dimensional electrically conductive microstructure comprises an electrically functional microneedle array comprising two or more microneedles, each including a high aspect ratio, polymeric three dimensional substrate structure which is at least substantially coated by an electrically conductive layer. | 1. A method for fabricating a three-dimensional, electrically conductive structure comprising: providing a mold having at least one microdepression which defines a three-dimensional structure;depositing and patterning at least one electrically conductive layer within the microdepression of the mold;filling the microdepression of the mold with at least one substantially electrically non-conductive substrate material;molding the at least one substrate material to form a substrate on top of the at least one electrically conductive layer while the at least one non-conductive substrate material abuts an outer surface of the at least one electrically conductive layer opposite the mold; andtransferring the at least one electrically conductive layer to the substrate, thereby forming an electrically conductive structure. 2. The method of claim 1, further comprising removing the electrically conductive structure from the mold. 3. The method of claim 1, further comprising electroplating onto the electrically conductive layer. 4. The method of claim 1, wherein the mold comprises a photoresist or polydimethlysiloxane. 5. The method of claim 1, wherein the substrate material comprises a polymer. 6. The method of claim 1, wherein the electrically conductive layer comprises nickel, iron, gold, titanium, copper, platinum, palladium, a stainless steel, or an alloy thereof. 7. The method of claim 1, wherein the patterning comprises selectively removing at least a portion of the electrically conductive layer from an upper surface of at least one protrusion of the mold. 8. The method of claim 1, wherein the patterning comprises depositing the electrically conductive layer onto protruding and/or recessed surfaces of the mold through a shadow mask or a stencil mask. 9. The method of claim 1, wherein the at least one microdepression defines a high-aspect ratio three-dimensional structure. 10. The method of claim 1, wherein the mold further comprises at least one upper protrusion that defines a channel within the three-dimensional structure. 11. A method for fabricating a three-dimensional, electrically conductive structure comprising: providing a mold having at least one microdepression which defines a three-dimensional structure and at least one upper protrusion which defines a channel within the three-dimensional structure;depositing and patterning at least one electrically conductive layer within the microdepression of the mold, the patterning comprising selectively removing at least a portion of the electrically conductive layer from an upper surface of the upper protrusion of the mold;filling the microdepression of the mold with at least one substrate material; andmolding the substrate material to form a substrate structure. 12. The method of claim 11, further comprising electroplating onto the electrically conductive layer. 13. The method of claim 11, wherein the mold comprises a photoresist or polydimethlysiloxane. 14. The method of claim 11, wherein... | GEORGIA TECH RESEARCH CORPORATION | H01B013/00 A61B005/00 A61M037/00 H05K001/02 C25D001/00 C25D001/10 C25D005/02 A61N001/32 H05K003/40 H05K003/20
US09332935 | B2 | 2016-05-10 | Device having embedded antenna | 2013-09-24 | Body-mountable devices and methods for embedding a structure in a body-mountable device are described. A body-mountable device includes a transparent polymer and a structure embedded in the transparent polymer. The transparent polymer defines a posterior side and an anterior side of the body-mountable device. The structure has an outer diameter and an inner diameter and includes a sensor configured to detect an analyte and an antenna. The antenna includes a plurality of conductive loops spaced apart from each other between the outer diameter and the inner diameter. | 1. A body-mountable device comprising: a transparent polymer, wherein the transparent polymer defines a posterior side and an anterior side of the body-mountable device; anda structure embedded in the transparent polymer, wherein the structure has an outer circumference defined by an outer diameter and an inner circumference defined by an inner diameter and includes a sensor configured to detect an analyte, electronics, and an antenna, wherein the antenna includes a plurality of conductive loops spaced apart from each other between the outer circumference and the inner circumference, wherein, with respect to a vertical axis between the anterior side of the body-mountable device and the posterior side of the body-mountable device, a particular conductive loop in the plurality of conductive loops is spaced vertically apart and horizontally apart from an adjacent conductive loop in the plurality of conductive loops,wherein each conductive loop in the plurality of conductive loops comprises a respective metal layer disposed between a first respective polymer layer and a second respective polymer layer, and wherein the respective metal layer, the first respective polymer layer, and the second respective polymer layer are each embedded in the transparent polymer,wherein the transparent polymer extends from the anterior side of the body-mountable device to the posterior side of the body-mountable device between adjacent conductive loops in the plurality of conductive loops, andwherein each conductive loop in the plurality of conductive loops is connected to a first interconnect and a second interconnect, wherein the electronics and sensor are disposed between the first and second interconnects, and wherein each conductive loop in the plurality of conductive loops is electrically connected to the electronics and sensor via the first and second interconnects. 2. The body-mountable device of claim 1, wherein the body-mountable device comprises an eye-mountable device. 3. The body-mountable device of claim 1, wherein the body-mountable device comprises a tooth-mountable device. 4. The body-mountable device of claim 1, wherein the body-mountable device comprises a skin-mountable device. 5. The body-mountable device of claim 1, wherein the conductive loops are substantially concentric. 6. The body-mountable device of claim 1, wherein the plurality of conductive loops comprises at least three conductive loops. 7. The body-mountable device of claim 1, wherein the metal and first and second polymer layers in each conductive loop are spaced apart from the metal and first and second polymer layers in each adjacent conductive loop in the plurality of conductive loops. 8. The body-mountable device of claim 7, wherein the metal and first and second polymer layers of a particular conductive loop in the plurality of conductive loops are spaced apart from the metal and first and second polymer layers of a first adjacent conductive loop in the plurality of conduct... | Verily Life Sciences LLC | A61B005/00 A61B005/145 G02C007/04 A61B005/1486
US09339203 | B2 | 2016-05-17 | Internet-based system for evaluating ECG waveforms to estimate the degree of coronary blockage | 2013-10-08 | The present invention provides an improved, Internet-based system that seamlessly collects cardiovascular data from a patient before, during, and after a procedure for EP or an ID. During an EP procedure, the system collects information describing the patient's response to PES and the ablation process, ECG waveforms and their various features, HR and other vital signs, HR variability, cardiac arrhythmias, patient demographics, and patient outcomes. Once these data are collected, the system stores them on an Internet-accessible computer system that can deploy a collection of user-selected and custom-developed algorithms. Before and after the procedure, the system also integrates with body-worn and/or programmers that interrogate implanted devices to collect similar data while the patient is either ambulatory, or in a clinic associated with the hospital. A data-collection/storage module, featuring database interface, stores physiological and procedural information measured from the patient. | 1. A system for evaluating a degree of blockage of a test patient's heart, comprising: an electrocardiogram (ECG)-measuring system configured to measure an ECG waveform from the test patient; anda computer system comprising a database stored on a non-transitory storage medium operably linked to the computer system, the database comprising a set of data fields wherein each data field in the set corresponds to an individual patient in a plurality of comparison patients and wherein each data field comprises an ECG waveform and a parameter indicating a degree of blockage of an artery feeding the individual comparison patient's heart, wherein the database is operably connected to the ECG-measuring system to receive the test patient's ECG waveform and store the test patient's ECG waveform in a data field corresponding to the test patient,an ECG-analysis software system configured to (i) receive the ECG waveforms from the database corresponding to each individual comparison patient in the plurality of comparison patients and to process the ECG waveform in each data field to extract a set of parameters from an individual comparison patient's ECG waveform,(ii) correlate the set of parameters to the degree of blockage of an artery feeding the individual comparison patient's heart,(iii) processes multiple data fields to determine an average correlation factor mapping the set of parameters to the parameter indicating the degree of blockage of an artery feeding each individual comparison patient's heart, and(iv) evaluate the degree of blockage of the test patient's heart by extracting a new set of parameters from the test patient's ECG waveform, and then comparing the new set of parameters to the average correlation factor to estimate a new parameter indicating the degree of blockage of an artery feeding the test patient's heart. 2. The system of claim 1, wherein the ECG-analysis software system is further configured to process the test patient's ECG waveform to determine a temporal location of at least one of the following parameters: P wave, Q wave, R wave, S wave, T wave, and U wave. 3. The system of claim 2, wherein the ECG-analysis software system is further configured to determine at least one of the following parameters from the P wave, Q wave, R wave, S wave, T wave, and U wave: PR interval, ST segment, PR segment, QRS interval, and QT interval. 4. The system of claim 1, wherein the ECG-analysis software system is further configured to process the test patient's ECG waveform to determine a magnitude of at least one of the following parameters: P wave, Q wave, R wave, S wave, T wave, and U wave. 5. The system of claim 3, wherein the average correlation factor correlates an average value of the PR interval to the degree of blockage of an artery feeding the test patient's heart. 6. The system of claim 5, wherein the ECG-analysis software system is further configured to determine the presence of blockage of an artery feeding the test patient's heart ... | TOSENSE, INC. | A61B005/04 A61B005/0402 A61B005/02 A61B005/00 G06F019/00 A61B005/0205 A61B005/044 A61B005/0452
US09820664 | B2 | 2017-11-21 | Catheter with high density electrode spine array | 2014-11-20 | A catheter adapted or high density mapping and/or ablation of tissue surface has a distal electrode matrix having a plurality of spines arranged in parallel configuration on which a multitude of electrodes are carried in a grid formation for providing uniformity and predictability in electrode placement on the tissue surface. The matrix can be dragged against the tissue surface upon deflection (and/or release of the deflection) of the catheter. The spines generally maintain their parallel configuration and the multitude of electrodes generally maintain their predetermined relative spacing in the grid formation as the matrix is dragged across the tissue surface in providing very high density mapping signals. The spines may have free distal ends, or distal ends that are joined to form loops for maintaining the spines in parallel configuration. | 1. A catheter comprising: an elongated catheter body; anda distal electrode assembly, the distal electrode assembly comprising at least three spines extending in a common plane, each spine having at least one electrode, a proximal spine portion, a distal spine portion, and a free distal end,wherein the catheter body defines a longitudinal axis, and each proximal spine portion is configured to extend at an angle from the longitudinal axis of the catheter body and each distal spine portion is configured to extend generally parallel to the longitudinal axis of the catheter body and generally parallel to each other distal spine Portion. 2. The catheter of claim 1, further comprising an intermediate deflection section between the catheter body and the distal electrode assembly. 3. The catheter of claim 1, wherein the distal electrode assembly has an electrode density ranging between about 12 and 16 electrodes per square centimeter. 4. The catheter of claim 1, wherein the distal electrode assembly has an electrode density of about 32 electrodes per 2.2 cm2. 5. The catheter of claim 1, wherein each spine comprises a shape memory member and a nonconductive covering. 6. The catheter of claim 1, wherein the at least three spines ranges between about three and ten. 7. The catheter of claim 1, wherein the distal electrode assembly is configured to maintain the distal portions of the at least three spines in parallel arrangement upon exertion of a contact force. 8. The catheter of claim 1, wherein inner proximal spine portions extend at lesser angles from the longitudinal axis of the catheter body, and outer proximal spine portions extend at greater angles from the longitudinal axis of the catheter body. 9. A catheter comprising: an elongated catheter body; anda distal electrode assembly, the distal electrode assembly comprising a plurality of spines extending in a common plane, each spine having one or more electrodes, a proximal spine portion, a distal spine portion, and a free distal end,wherein the one or more electrodes are arranged in a predetermined pattern within the distal electrode assembly;wherein the catheter body defines a longitudinal axis, each proximal spine portion is configured to extend at an angle from the longitudinal axis of the catheter body, and each distal spine portion is configured to extend generally parallel to the longitudinal axis of the catheter body, andwherein inner proximal spine portions extend at lesser angles from the longitudinal axis of the catheter body, and outer proximal spine portions extend at greater angles from the longitudinal axis of the catheter body. 10. The catheter of claim 9, further comprising an intermediate deflection section between the catheter body and the distal electrode assembly. 11. The catheter of claim 9, wherein the distal electrode assembly has an electrode density ranging between about 12 and 16 electrodes per square centimeter. 12. The catheter of claim 9, wherein the distal elec... | Biosense Webster (Israel) Ltd. | A61B005/042 A61B018/14 A61B005/00 A61B018/00
US10130287 | B2 | 2018-11-20 | Hearing test system | 2014-11-18 | Techniques are described for improving results of hearing tests. For example, the hearing tests described herein may be divided into multiple hearing test segments which may be provided to the test subject individually. Additionally, the hearing test may include multiple cognitive test segments that may be interleaved between the hearing test segments to gauge the test subject's awareness level. | 1. A device comprising: one or more user interfaces for receiving user input;an input interface for releasably coupling to at least one speaker;a display;one or more processors; andnon-transitory computer-readable storage media storing computer-executable instructions, which when executed by the one or more processors cause the one or more processors to:receive image data associated with an ear of the test subject after sound damping foam has been placed in an ear canal of the ear;present on the display, image data associated with an ear of the test subject after sound damping foam has been placed in an ear canal of the ear and at least one image illustrating a correct placement of the sound damping foam within an ear canal;provide audio associated with a hearing test segment to the speaker;receive user inputs associated with the audio output by the speaker;present a cognitive test segment on the display;receive user inputs associated with the cognitive test segment;determine an awareness level of the test subject based at least in part on the user inputs associated with the cognitive test segment;determine the awareness level is above a threshold;mark the hearing test segment as complete; anddisplay test results on the display. 2. The device as recited in claim 1, wherein the computer-readable storage media further stores computer-executable instructions, which when executed by the one or more processors cause the one or more processors to select the hearing test segment from a plurality of hearing test segments based at least in part on information about the test subject. 3. The device as recited in claim 1, wherein the computer-readable storage media further stores computer-executable instructions, which when executed by the one or more processors cause the one or more processors to select the cognitive test segment from a plurality of cognitive test segments based at least in part on demographic information associated with a plurality of prior test subjects. 4. The device as recited in claim 1, wherein the computer-readable storage media further stores computer-executable instructions, which when executed by the one or more processors cause the one or more processors to determine that each of a plurality of hearing test segments are marked as completed before outputting the test results. 5. The device as recited in claim 1, wherein the hearing test segment includes outputting at the speaker at least one of: a test tone;a part of speech;a phrase;a melody;a pattern of tones; orenvironmental noises. 6. A device comprising: one or more user interfaces for receiving user input;an input interface for releasably coupling to at least one speaker;a display;one or more processors; andnon-transitory computer-readable storage media storing computer-executable instructions, which when executed by the one or more processors cause the one or more processors to: process image data associated with an ear of the test subject after sound damping foam has b... | Audicus, Inc. | A61B005/12 A61B019/00 A61B005/16 A61B005/00
US09345392 | B2 | 2016-05-24 | Endoscope system and method for imaging exposure control | 2013-02-27 | In a special observation mode, an oxygen saturation frame period, a normal frame period, and a vessel pattern frame period are repeatedly performed. A brightness detector detects the brightness of a latest frame image of an oxygen saturation video image, being a key video image. The intensity of light to be applied in the next oxygen saturation frame period is determined from the detected brightness. From the determined light intensity and a light intensity ratio among frames, the intensity of light to be applied in the next normal frame period and the next vessel pattern frame period is calculated. The exposure time of the next oxygen saturation frame period is determined from the detected brightness. From the determined exposure time and an exposure time ratio among frames, the exposure time of the next normal frame period and the next vessel pattern frame period is determined. | 1. An endoscope system comprising: a lighting section for sequentially and repeatedly applying a plurality of types of illumination light having different wavelength bands to an internal body portion;an image sensor for sequentially capturing reflected light from said internal body portion in each frame period in which said illumination light is applied, and producing a plurality of types of frame images corresponding to said types of illumination light;a video image processing unit for independently producing a plurality of modes of video images corresponding to said plurality of types of frame images, based on said plurality of types of frame images;a display section for displaying said plurality of modes of video images;a brightness detector for detecting brightness of a specific mode of video image out of said plurality of modes of video images;an exposure condition determiner having a memory for storing correlation in exposure conditions among said plurality of modes of video images, calculating said exposure condition of all of said plurality of modes of video images including said specific mode of video image based on said brightness detected by said brightness detector and said correlation, and determining exposure conditions of said plurality of modes of video images to said calculated conditions in a transition among said plurality of modes; and an exposure controller for regulating exposure of said image sensor in accordance with said determined exposure conditions. 2. The endoscope system according to claim 1, wherein said correlation in said exposure conditions is a light intensity ratio among said plurality of types of illumination light. 3. The endoscope system according to claim 1, wherein said correlation in said exposure conditions is an exposure time ratio of said image sensor to obtain each of said plurality of types of frame images. 4. The endoscope system according to claim 1, wherein said memory stores a first correlation for making visibility of a specific mode of video image produced from said plurality of modes of video images higher than that of the other modes of video images, and a second correlation for making visibility of two or more of said modes of video images including said specific mode of video image higher than that of the other modes of video images; andone of said first and second correlation is chosen manually. 5. The endoscope system according to claim 4, wherein said plurality of modes of video images produced with use of said first or second correlation have different brightness from each other. 6. The endoscope system according to claim 4, wherein said plurality of modes of video images produced with use of said first or second correlation have different frame rates from each other. 7. The endoscope system according to claim 1, wherein said plurality of modes of video images include: a normal video image obtained under white light;an oxygen saturation video image that images an oxygen saturation... | FUJIFILM Corporation | A61B001/06 A61B001/045 A61B001/00 A61B001/005 A61B001/04 A61B005/1455 A61B005/1459
US09345441 | B2 | 2016-05-24 | Apparatus and method for nuclear imaging | 2012-12-20 | A system for endoscopic nuclear imaging, a nuclear probe and related methods are provided. In one embodiment, a system includes a movable endoscopic nuclear probe which comprises a collimator and a detector, the endoscopic nuclear probe having an elongated body with a longitudinal axis (X), wherein the detector and the collimator are arranged on the elongated body, and the collimator having a field of view with an axis (Y) and the axis (Y) is angled with respect to the axis (X) of the elongated body of the endoscopic nuclear probe; means for determining the position and orientation of the elongated body of the endoscopic nuclear probe; an evaluation unit including means for computationally determining an actual position and actual orientation of the elongated body of the endoscopic nuclear probe; and a program storage section and means for computing at least one quality parameter of the image. | 1. A system for endoscopic nuclear imaging, which comprises: a movable endoscopic nuclear probe which includes a collimator, a crystal for detecting radioactive radiation, and a photomultiplier coupled to the crystal, wherein the endoscopic nuclear probe has an elongated body with a longitudinal axis (X), which is designed to be inserted into the body of a patient, wherein the crystal and the collimator are arranged on the elongated body of the endoscopic nuclear probe, wherein the collimator surrounds the crystal and the photomultiplier, wherein the collimator has a laterally directed opening which delimits a detection angle through a field of view of the collimator, and wherein the field of view has an axis (Y), wherein the axis (Y) of the field of view of the collimator is inclined with respect to the axis (X) of the elongated body of the endoscopic nuclear probe, and wherein the field of view of the collimator is limited anisotropically with respect to the longitudinal axis X of the elongated body;means for determining the position and orientation of the elongated body of the endoscopic nuclear probe; andan evaluation unit, which includes: i) an interface system for transmitting data with information about the detected radioactive radiation, and information about the position and orientation of the elongated body of the endoscopic nuclear probe to the evaluation unit;ii) a data storage section for storing the information;iii) means for synchronising the information about the position and orientation of the elongated body of the endoscopic nuclear probe with the information about the radiation detected by the endoscopic nuclear probe;iv) means for computationally determining an actual position and actual orientation of the elongated body of the endoscopic nuclear probe;v) a program storage section with a program for repeated determination of at least one quality value Q with regards to the imaging from the information mentioned; andvi) an output system for outputting a three-dimensional reproduction of the information about the actual position and actual orientation of the elongated body of the endoscopic nuclear probe and information about the detected radiation. 2. The system according to claim 1, wherein the evaluation unit further comprises: means for computing a desired position and desired orientation of the elongated body of the endoscopic nuclear probe, wherein these data are determined under the additional condition that a corresponding placing of the elongated body of the endoscopic nuclear probe leads to an improvement of the at least one quality value Q; andan output system for outputting an instruction to a user or robot for moving the endoscopic nuclear probe to the desired position and desired orientation. 3. The system according to claim 1, further comprising: a reference which is fixed in relation to the patient; andmeans for determining the position and orientation of the reference, wherein the interface system of the evalu... | SURGICEYE GMBH | A61B006/00 A61B001/317 A61B001/267 A61B001/233 A61B001/313 A61B001/273 A61B001/31 A61B001/307 A61B001/303 A61B001/227 A61B001/06 A61B001/04 A61B001/00 A61B006/03 A61B005/06 A61B008/08 G01T001/161 A61B008/00 A61B008/12 A61B006/12 A61B019/00
US09345907 | B2 | 2016-05-24 | Bladder and/or rectum extender with exchangeable and/or slideable tungsten shield | 2012-07-27 | A method of using a shielding device with an applicator includes inserting the distal portion of the shielding device to a treatment site of a patient, the shield being made of plastic, imaging the treatment site, planning treatment of the treatment site based on the imaging, setting up the patient for treatment, imaging the treatment site, removing the distal portion of the shielding device from the treatment site, exchanging the shield made of plastic with a shield made of tungsten, inserting the distal portion of the shielding device to the treatment site, the shield being made of tungsten, imaging the treatment site, and performing the treatment. | 1. A method of using at least one shielding device with an applicator, the shielding device including: two elongated, parallel tubes joined at a distal portion of the tubes, and a shield at least one of removably and slidably situated at the distal portion, the distal portion including a separate shield retainer attached to distal ends of the tubes, the shield retainer including a retaining lip and a key slot configured to retain the shield, the shield including a retaining ring and a key portion configured to be received by the retaining lip and the key slot, respectively, comprising: inserting the distal portion of the shielding device to a treatment site of a patient, the shield being made of plastic;imaging the treatment site a first time;planning treatment of the treatment site based on the imaging;setting up the patient for treatment;imaging the treatment site a second time after imaging the treatment site the first time;removing the distal portion of the shielding device from the treatment site;exchanging the shield made of plastic with a shield made of tungsten;inserting the distal portion of the shielding device to the treatment site, the shield being made of tungsten;imaging the treatment site a third time after imaging the treatment site the first and second times; andperforming the treatment. 2. The method according to claim 1, further comprising: after the inserting of the distal portion, adjusting a position of the distal portion of the plastic shield relative to a distal end of the applicator. 3. The method according to claim 2, wherein the adjusting includes setting a longitudinal depth of insertion of the distal portion. 4. The method according to claim 2, wherein the adjusting includes setting a rotational position of the distal portion relative to the distal end of the applicator. 5. The method according to claim 2, further comprising: after the adjusting, and before the removing of the distal portion, returning the distal portion to an unadjusted position. 6. The method according to claim 5, further comprising: after the returning to the unadjusted position, and after the removing the distal portion of the plastic shield and the inserting of the distal portion of the tungsten shield, reproducibly adjusting the distal portion of the tungsten shield to a previously adjusted position. 7. The method according to claim 1, further comprising: at least one of moving and removing at least one of the plastic shield and the tungsten shield. 8. The method according to claim 1, further comprising: sliding at least one of the plastic shield and the tungsten shield relative to the tubes while the distal portion of the shielding device is inserted to the treatment site. 9. The method according to claim 1, wherein the imaging includes at least one of fluoroscopy, computed tomography scanning, and magnetic resonance imaging. | MICK RADIO-NUCLEAR INSTRUMENTS, INC. | A61B005/05 A61N005/10 A61B005/055 A61B005/00 A61B019/00 A61B006/03 A61B006/00
US09332907 | B2 | 2016-05-10 | Extracting application dependent extra modal information from an anatomical imaging modality for use in reconstruction of functional imaging data | 2013-03-28 | Segmentation is provided in multi-modal reconstruction of functional information. Rather than using HU values for segmentation, the HU values are converted into linear attenuation coefficients, such as in a μ-map (Mu map). The conversion uses different functions for different ranges of the HU values. Linear attenuation coefficients are less likely subject to variation by patient, protocol, or scanner. The resulting segmentation may be more consistent across various clinical settings, providing for more accurate multi-modal reconstruction. | 1. A method for segmentation in multi-modal reconstruction of functional information, the method comprising: obtaining computed tomography (CT) data representing a region of a patient;converting the CT data to linear attenuation coefficients;segmenting, with the linear attenuation coefficients, locations in the region into at least five types of anatomical tissue, the locations comprising zones segmented between respective types of the anatomical tissue;generating separate zonal image objects for each of the zones from the functional information;reconstructing, with a processor, the functional information into an image object, the reconstructing including separately forward projecting the functional information for each zone into data space, wherein a contribution of each zone to a data model is weighted according to a zonal scaling factor for the respective zones; andgenerating an image as a function of the image object. 2. The method of claim 1 wherein obtaining the CT data comprises obtaining the CT data in Hounsfield measurement units and wherein converting comprises converting to a μ-map. 3. The method of claim 1 wherein obtaining the CT data comprises obtaining with a CT scanner for a patient using a protocol, and wherein converting the CT data to linear attenuation coefficients comprises adapting to the CT scanner, the patient, and the protocol. 4. The method of claim 1 wherein converting comprises: comparing a value of the CT data to a predetermined range;when the value is within a predetermined range, calculating a first linear attenuation coefficient from the value using a first function;when the value is outside the predetermined range, calculating the first linear attenuation coefficient from the value using a second function different than the first function. 5. The method of claim 4 wherein the predetermined range indicates air-soft tissue and wherein calculating using the first function comprises converting using a linear function calculated from data points of air and water, for which the energy dependence is obtained from a fit to a double-power law, and wherein calculating using the second function comprises converting by a linear fit from soft tissue and bone tissue with a constraint that the second function is continuous over an entire range of Hounsfield units of the CT data. 6. The method of claim 1 wherein segmenting comprises comparing the linear attenuation coefficients for each location with thresholds, the thresholds separating the anatomical tissue by type, the at least five types including dense bone, soft bone, soft tissue, adipose, and lung. 7. The method of claim 1 wherein generating separate zonal image objects comprises separately grouping the functional information for locations into the respective zones. 8. The method of claim 1 further comprising: smoothing a transition between zones. 9. The method of claim 1 further comprising: correcting the functional information for attenuation with the linear atte... | Siemens Medical Solutions USA, Inc. | G06K009/00 A61B005/00 G06T011/00 G06T007/00
US09332923 | B2 | 2016-05-10 | Sensor element with an insulation layer | 2011-06-16 | A sensor element comprises a sensor section comprising a sensor unit configured to measure a physiological variable or any other signal in a living body and to generate a sensor signal in response to the variable or other signal, and a bond section comprising contact members configured to electrically connect at least one signal transmitting microcable. The bond section is coated with an electrically insulating material and the sensor unit is left uncoated. The sensor element may further comprise an intermediate section between the sensor section and the bond section. The intermediate section includes electric connection lines configured to connect the contact members to the sensor unit. The intermediate section is also coated with the electrically insulating material. | 1. A sensor element, comprising: a sensor support body;a sensor unit disposed on the sensor support body, the sensor unit being configured to measure a variable in a living body and to generate a sensor signal in response to said measurement, and the sensor unit comprising a cavity covered by a membrane;a plurality of contact members disposed on the sensor support body, the contact members being configured to be electrically connected to at least one signal transmitting microcable; anda plurality of electric connection lines that connect the respective contact members to the sensor unit,wherein all exposed surfaces of the contact members and at least a portion of each electric connection line are coated with an electrically insulating layer, andwherein (i) an entirety of an exposed surface of the membrane, and (ii) a surface of the sensor support body that peripherally surrounds the sensor unit on at least three sides of the sensor unit, are left uncoated by said electrically insulating layer. 2. The sensor element according to claim 1, wherein the sensor element has an elongated rectangular shape. 3. The sensor element according to claim 1, wherein the sensor element has a length in a range of 1.0 to 2.2 mm, a width in a range of 0.1 to 0.6 mm, and a thickness in a range of 0.05 to 0.3 mm. 4. The sensor element according to claim 1, wherein a material of the electrically insulating layer is biocompatible and chemically stable in the living body. 5. The sensor element according to claim 1, wherein a material of the electrically insulating layer is selected from a group consisting of silicone and epoxy. 6. The sensor element according to claim 1, wherein said sensor unit is sensitive to one or more physiological variables comprising pressure, temperature, and flow velocity. 7. The sensor element according to claim 1, wherein the sensor unit comprises at least one of a piezoresistive element, a capacitor, and a mechanically resonating sensor. 8. The sensor element according to claim 1, wherein the sensor unit comprises at least one piezoresistive element arranged in connection with the membrane. 9. A sensor wire for an intravascular measurement of a variable in a living body, wherein the sensor wire has a proximal region, a distal sensor region and a distal tip region, the sensor wire comprising: a sensor element according to claim 1 and arranged in the distal sensor region;a jacket accommodating at least a part of the sensor element;a core wire; andthe at least one signal transmitting microcable connected to the contact members and running along the sensor wire. 10. The sensor element according to claim 1, wherein a surface of the sensor support body that peripherally surrounds the sensor unit on four sides of the sensor unit is uncoated by said electrically insulating layer. 11. A method of producing a sensor element, the method comprising: forming a sensor unit on a sensor support body, the step of forming the sensor unit comprising: f... | ST. JUDE MEDICAL COORDINATION CENTER BVBA | A61B005/05 A61B005/01 A61B005/0215 A61B005/00
US09327100 | B2 | 2016-05-03 | Selective drug delivery in a lumen | 2013-03-12 | Methods and systems are disclosed for selective drug or fluid delivery in a lumen through a coating or fluid delivery channels. One system includes an elongate catheter having a proximal end and a distal end with an axis therebetween, the catheter having a radially expandable balloon near the distal end and an energy delivery portion proximate the balloon for transmission of energy, a thermally changeable coating having a releasable drug coupled to the balloon, the thermally changeable coating being oriented to be urged against the body tissue when the expandable balloon expands and an energy source operatively coupled to the energy delivery portion configured to energize the energy delivery portion to heat and liquefy the thermally changeable coating to release the drug to the body tissue. | 1. A medical device, comprising: a catheter shaft having a distal region;an expandable member coupled to the distal region;a coating disposed along the expandable member, the coating including a releasable drug;a plurality of electrodes coupled to the expandable member;wherein the plurality of electrodes includes one or more pairs of bipolar electrodes; andwherein the electrodes are configured to induce a phase change in the coating and release the releasable drug from the coating. 2. The medical device of claim 1, wherein the expandable member includes a balloon. 3. The medical device of claim 2, wherein the electrodes are mounted to an outer surface of the balloon. 4. The medical device of claim 2, wherein the balloon is configured to be fully inflated with a pressure of 6 atmospheres or less. 5. The medical device of claim 2, wherein the balloon is configured to be fully inflated with a pressure of 1 to 2 atmospheres. 6. The medical device of claim 1, further comprising a generator coupled to the catheter shaft for selectively energizing the plurality of electrodes. 7. The medical device of claim 6, wherein the generator is configured to energize the electrodes for 1 to 180 seconds. 8. The medical device of claim 6, wherein the generator is configured to energize the electrodes with 0.25 to 5 Watts average power. 9. The medical device of claim 6, wherein the generator includes an RF energy source. 10. The medical device of claim 1, wherein at least some of the plurality of electrodes are configured to sense impedance. 11. The medical device of claim 1, wherein at least some of the plurality of electrodes are configured to characterize target tissue based on sensed impedance. 12. The medical device of claim 1, wherein the coating has coating melting temperature above 37° Celsius. 13. The medical device of claim 1, wherein energizing the electrodes liquefies at least a portion of the coating. 14. A medical device, comprising: a catheter shaft having a distal region;a balloon coupled to the distal region;a coating disposed on the balloon, the coating including a releasable drug;one or more pairs of bipolar electrodes coupled to the balloon;wherein the electrodes are configured to induce a phase change in the coating and release the releasable drug from the coating, wherein the phase change of the coating from a solid to a liquid releases the drug; anda generator coupled to the catheter shaft, the generator being configured to selectively energize the electrodes. 15. The medical device of claim 14, wherein the generator includes an RF energy source. 16. The medical device of claim 14, wherein at least some of the one or more pairs of bipolar electrodes are configured to sense impedance and characterize target tissue based on sensed impedance. 17. The medical device of claim 14, wherein the coating has coating melting temperature above 37° Celsius. 18. The medical device of claim 14, wherein energizing the electrodes liquefies at l... | VESSIX VASCULAR, INC. | A61M025/10 A61B005/053 A61B005/00 A61B018/14 A61L029/08 A61L029/14 A61L029/16 A61B018/24 A61B018/00 A61B018/18 A61M037/00 A61N007/02
US09339221 | B1 | 2016-05-17 | Diagnosing intestinal ischemia based on oxygen saturation measurements | 2009-03-24 | Devices and systems have a sensor probe configured to measure tissue oxygen saturation in the intestine or mesentery. The devices and systems can determine the oxygenation state of the entire thickness of the intestine or mesentery. Thus, embodiments of the invention can be applied in diagnosing intestinal ischemia in a patient, as well as in monitoring tissue oxygen saturation of the intestine or mesentery during or after a surgical procedure. | 1. A system comprising: a catheter device comprising:an outer sheath, anda sensor probe, disposed inside the outer sheath, comprising a first optical fiber and a second optical fiber wherein distal ends of the first and second optical fibers are coupled to a tip of the sensor probe, where an oximeter sensor is formed,the tip of the sensor probe comprises a planar surface in which a first source structure and first detector structure are planar with the planar surface, wherein the first source structure is coupled to a cross-sectional end of the first optical fiber and the first detector structure is coupled to a cross-sectional end of the second optical fiber,at the tip, the first source structure and first detector structure are arranged in a line and there is a fixed distance between the first source structure and the first detector structure,the sensor probe comprises a second source structure and a second detector structure, a first distance is between the first source structure and the first detector structure, a second distance is between the first source structure and the second detector structure, a third distance is between the second source structure and the first detector structure, a fourth distance is between the second source structure and the second detector structure,the first distance is not equal to the second, third, and fourth distances, the second distance is not equal to the third and fourth distances, and the third distance is not equal to the fourth distance,wherein a third optical fiber couples to the second source structure, and a fourth optical fiber couples to the second detector structure,the tip comprises only the first source structure, second source structure, first detector structure, and second detector structure, wherein there are no source or detector structures between the first source structure and the first detector structure, second source structure and the second detector structure, the first source structure and the second source structure, the first source structure and the second source structure, andthere are no source or detector structures between the outer sheath and the first source structure, the outer sheath and the second source structure, the outer sheath and the first detector structure, and the outer sheath and the second detector structure,at the tip, a filler material is between the first source structure and first detector structure, and between the first source structure and first detector structure and the sheath to hold the first and second two optical fibers at the fixed distance,the oximeter sensor is adapted to measure oxygen saturation of tissue having a thickness greater than 0, where the fixed distance between the first source structure and first detector structure is greater than 0, the second distance is greater than 0, the third distance is greater than 0, and the fourth distance is greater than 0, thereby allowing the oximeter sensor to measure oxygen saturation for the thickn... | ViOptix, Inc. | A61B005/00 A61B005/1459
US09333270 | B2 | 2016-05-10 | Synthesis and composition of amino acid linking groups conjugated to compounds used for the targeted imaging of tumors | 2013-08-26 | The present disclosure relates to compounds that are useful as near-infrared fluorescence probes, wherein the compounds include i) a pteroyl ligand that binds to a target receptor protein, ii) a dye molecule, and iii) a linker molecule that comprises an amino acid or derivative thereof. The disclosure further describes methods and compositions for making and using the compounds, methods incorporating the compounds, and kits incorporating the compounds. | 1. A method of optical imaging of a biological tissue that expresses a folate receptor, the method comprising: (a) contacting the biological tissue with a composition comprising a compound having the formula:or a pharmaceutically acceptable salt thereof,or isotopes thereof, wherein:X is an single amino acid or a single amino acid derivative thereof, wherein the single amino acid or single amino acid derivative contains an —OH, —NH2 or —SH functional group, andY is a dye that has a fluorescence excitation and emission spectra in the near infra red range, wherein Y is represented by the formula:wherein, X′ is independently selected from the group consisting of O, S, N and C, andR′ is independently selected from the group consisting of CH2 and CH2CH2;and the compound maintains or enhances the fluorescence of the dye, Y;(b) allowing time for the compound in the composition to distribute within the biological target;(c) illuminating the tissue with an excitation light of a wavelength absorbable by the compound; and(d) detecting the optical signal emitted by the compound. 2. The method of claim 1, wherein the signal emitted by the compound is used to construct an image. 3. The method of claim 1, wherein the tissue is in a subject and the subject is an animal or human. 4. The method of claim 1, wherein in step (a) two or more fluorescent compounds whose signal properties are distinguishable are contacted with the tissue, and optionally the tissue is in a subject. 5. The method of claim 1, wherein the illuminating and detecting steps are performed using an endoscope, catheter, tomographic system, hand-held optical imaging system, surgical goggles, or intra-operative microscope. 6. The method of claim 1 wherein the biological tissue that expresses a folate receptor has a disease that overexpresses the receptor. 7. The method of claim 6, wherein the disease is selected from the group consisting of cancer, cardiovascular diseases, neurodegenerative diseases, immunologic diseases, autoimmune diseases, respiratory diseases, metabolic diseases, inherited diseases, infectious diseases, bone diseases, and environmental diseases. 8. The method of claim 1 wherein the amino acid is selected from the group consisting of tyrosine, cysteine, or a derivative of tyrosine or cysteine. 9. The method of claim 1, wherein the amino acid is tyrosine. 10. The method of claim 1 wherein a carbon isotope is on the aromatic ring of tyrosine. 11. The method of claim 1 wherein a hydrogen isotope is a substituent of the aromatic ring of tyrosine. 12. The method of claim 1 wherein the amino acid derivative is a derivative of tyrosine selected from the group consisting of:or racemic mixtures thereof. 13. The method of claim 1 wherein the amino acid includes a sulfur-containing side chain group. 14. The method of claim 13 wherein the amino acid including a sulfur-containing side chain group is cysteine. 15. The method of claim 13 wherein the amino acid including a sulfur-c... | Purdue Research Foundation | A61B005/00 A61B008/00 A61B010/00 A61K049/00 G01N033/569 G01N033/566 G01N033/50 G01N033/58 G01N033/68
US10405799 | B2 | 2019-09-10 | Device features and design elements for long-term adhesion | 2016-01-25 | An electronic device for long-term adhesion to a mammal includes a housing with an electronic component. The electronic device may include a first wing and a second wing, each being integrally formed with the housing. An electrode is positioned on a bottom surface of each of the wings, the electrodes electrically connected to the electronic component. An adhesive layer is provided for adhesion to a surface of the mammal. The adhesive layer may cover a portion of the bottom surfaces of the wings but generally does not cover the electrode or a bottom surface of the housing. A method of applying an electronic device to a mammal includes removing first and second adhesive covers from first and second wings of the electronic device to expose an electrode and an adhesive coated on a bottom surface of each wing. | 1. An electronic device for long-term adhesion to a mammal, the device comprising: a housing;a physiologic data collection circuit contained in the housing;a plurality of wings extending from the housing;an electrode positioned on a bottom surface of each of the wings, the electrodes electrically connected to the physiologic data collection circuit;each wing comprising a first adhesive layer positioned on a bottom surface of the wing, wherein the first adhesive layer provides adhesion to the skin of the mammal; andeach wing comprising a second adhesive layer functioning as an adhesion guard protecting the first adhesive layer; the second adhesive layer positioned over a top of the wing and extending horizontally outward beyond a boundary of each wing. 2. The device of claim 1, wherein the first adhesive layer is not on a bottom surface of the housing. 3. The electronic device of claim 1, wherein the wings are more flexible than the housing. 4. The electronic device of claim 1, wherein the wings and the housing are made from the same material. 5. The electronic device of claim 1, wherein the wings and the housing are made from different materials. 6. The electronic device of claim 1, further comprising a hinge portion adjacent the housing. 7. The electronic device of claim 1, further comprising a flap connected to each wing, the flaps extending below the housing but not directly attached to a bottom surface of the housing, wherein the first adhesive layer extends onto a bottom surface of each flap. 8. The electronic device of claim 7, further comprising a connector segment configured to connect the flaps together without directly attaching to the bottom surface of the housing. 9. The electronic device of claim 1, wherein the first adhesive layer comprises an adhesive that can absorb fluids. 10. The electronic device of claim 1, wherein the first adhesive layer comprises a hydrocolloid adhesive. 11. The electronic device of claim 1, wherein the first adhesive layer comprises a pressure-sensitive adhesive. 12. The electronic device of claim 11, wherein the pressure sensitive adhesive is selected from the group consisting of a polyacrylate, a polyisobutylene and a polysiloxane. 13. The electronic device of claim 1, further comprising; an additional adhesive layer between each wing and the first adhesive layer; anda synthetic material layer between the wing and the first adhesive layer. 14. The electronic device of claim 1, wherein the physiologic data collection circuit is configured to collect cardiac rhythm data from the mammal. 15. The electronic device of claim 1, wherein each wing comprises a fabric or synthetic fiber layer. 16. The electronic device of claim 15, wherein the fabric or synthetic fiber layer is positioned over the first adhesive layer. 17. The electronic device of claim 1, further comprising a third adhesive layer positioned between the first adhesive layer and the second adhesive layer. 18. The electronic device ... | iRhythm Technologies, Inc. | A61B005/0408 A61B005/04 A61B005/00 A61B005/0478 A61B005/0488
US09764191 | B2 | 2017-09-19 | Training apparatus | 2015-10-12 | A training apparatus includes an operating rod, a strength detector, a motion position detector, a strength speed calculator, a boundary line arrival speed calculator, and a motion speed calculator. The operating rod moves a held limb. The strength detector outputs a strength component signal based on a magnitude of a strength component. The motion position detector detects a motion position of the operating rod. The strength speed calculator calculates a strength speed. The boundary line arrival speed calculator calculates a boundary line arrival speed whose absolute value is smaller as a boundary line distance is shorter. The motion speed calculator calculates a lower one of the strength speed and the boundary line arrival speed as the motion speed at which the operating rod should move. | 1. A training apparatus for training a user's four limbs including upper limbs and/or lower limbs according to a predetermined training program, the apparatus comprising: an operating rod configured for moving a held limb, the operating rod being supported to a stationary frame placed on or near a floor surface so as to be movable at one or more degrees of freedom;a strength detector configured for detecting a strength component that is a component of a strength applied to the operating rod in each freedom degree direction where the operating rod is movable, and outputting a strength component signal based on a magnitude of the detected strength component;a motion position configured detector for detecting a motion position of the operating rod in the corresponding freedom degree direction where the operating rod is movable;a strength speed calculator configured for calculating a strength speed of the operating rod based on the strength component signal output from the strength detector;a boundary line arrival speed calculator configured for calculating a boundary line arrival speed whose absolute value is smaller as a boundary line distance is shorter, the boundary line distance being from a current motion position of the operating rod to a mobile region boundary line for defining a boundary of an operating rod mobile region for setting a range where the operating rod is movable; anda motion speed calculator configured for calculating a lower one of the strength speed and the boundary line arrival speed as a motion speed at which the operating rod is to be moved. 2. The training apparatus according to claim 1, wherein the motion speed is limited to a maximum motion speed or less. 3. The training apparatus according to claim 1, wherein when a determination is made that the current motion position of the operating rod is present outside the operating rod mobile region, the motion speed calculator calculates a motion speed including a speed component toward a motion position reference point as a reference point of the motion position of the operating rod. 4. The training apparatus according to claim 2, wherein when a determination is made that the current motion position of the operating rod is present outside the operating rod mobile region, the motion speed calculator calculates a motion speed including a speed component toward a motion position reference point as a reference point of the motion position of the operating rod. | MURATA MACHINERY, LTD. | A63B024/00 A63B023/12 A61H001/02 A63B021/00 G06F019/00 A61B005/22 A63B021/005 A63B023/035 A61B005/11 A63B071/06
US09642545 | B2 | 2017-05-09 | Neural probe with optical stimulation capability | 2016-02-01 | A neural probe is disclosed for optically stimulating or silencing neurons and recording electrical responses to the stimulus. Using patterning techniques, an integral optical waveguide may be fabricated on the probe for transmitting neuron-affecting light from a light source to a probe tip. The probe tip may include one or more electrodes to receive electrical responses from stimulated neurons for recording or further processing. According to various embodiments, the disclosed neural probes may utilize multiple light sources simultaneously, switch between multiple light sources, or utilize a single light source to stimulate or silence multiple neuron locations simultaneously via multiple probe tips or via multiple light-emitting sites located along the length of the probe. Neural probes are thereby provided that have sufficient spatial resolution to accurately target, stimulate, and record the reaction of neurons, or as few as a single neuron, utilizing a slim, compact structure. | 1. A neural probe, comprising: a probe body;a shank extending from the probe body to a tip;a light-emitting diode (LED) light source attached to the tip for providing neuron-affecting light at said tip, wherein the shank is configured to locate the LED light source at a probing depth to target one or more neurons; andone or more recording electrodes attached to said tip for receiving electrical responses to the neuron-affecting light. 2. A neural probe as defined in claim 1, wherein the LED light source is located near the one or more recording electrodes. 3. A neural probe as defined in claim 2, wherein the one or more recording electrodes comprise a plurality of recording electrodes, and wherein the LED light source is located between at least some of the plurality of recording electrodes. 4. A neural probe as defined in claim 1, comprising a plurality of LED light sources attached to the tip. 5. A neural probe as defined in claim 4, wherein at least two of the plurality of LED light sources are in a staggered configuration. 6. A neural probe as defined in claim 4, wherein one of the LED light sources is located closer to the tip than at least one other of the LED light sources. 7. A neural probe as defined in claim 4, wherein at least two of the plurality of LED light sources each emit a different wavelength of light than the other. 8. A neural probe as defined in claim 7, wherein one wavelength of light stimulates a neuron and another wavelength of light silences a neuron. 9. A neural probe as defined in claim 4, wherein the plurality of LED light sources have separate electrical connections to control the plurality of LED light sources. 10. A neural probe as defined in claim 9, comprising a plurality of recording electrodes having separate electrical connections, and the electrical connections for the plurality of recording electrodes are located on either side of the electrical connections for the LED light sources. 11. A neural probe as defined in claim 1, wherein the LED light source is an edge-emitting LED for providing neuron-affecting light from a side of the LED light source. 12. A neural probe as defined in claim 1, further comprising an optical waveguide extending from a source end at the probe body to an emitting end near said tip, wherein a second light source is in communication with the source end and the neuron-affecting light is provided at the tip by the emitting end and the LED light source. 13. A neural probe as defined in claim 12, wherein the second light source comprises a fiber optic having a first end in communication with an external light source and a second end abutting the source end of the optical waveguide. 14. A neural probe as defined in claim 12, wherein the second light source comprises an LED light source mounted adjacent the source end of the optical waveguide. 15. A neural probe as defined in claim 1, wherein the LED light source is an organic light-emitting diode (OLED) light source. 16. A ... | The Regents of the University of Michigan | A61B005/04 A61N005/06 A61B005/00 A61N001/05
US09330206 | B2 | 2016-05-03 | Producing a three dimensional model of an implant | 2014-05-14 | Determining a shape of a medical device to be implanted into a subject produces an image including a defective portion and a non-defective portion of a surface of a tissue of interest included in the subject. The tissue of interest is segmented within the image. A template, representing a normative shape of an external anatomical surface of the tissue of interest, is superimposed to span the defective portion. An external shape of an implant, is determined as a function of respective shapes of the defective portion as seen in the template, for repairing the defective portion. | 1. A computer implemented method of obtaining data for determining a 3-dimensional shape of a medical device, the method comprising: (a) obtaining computer readable image data of a target tissue wherein the target tissue comprises two portions, a portion with a defect and a portion without a defect;(b) rendering from the image data a computer-generated 3-dimensional representation of the target tissue;(c) superimposing a three-dimensional template onto the 3-dimensional representation, wherein the three-dimensional template represents a normative shape of an anatomical surface of the target tissue; and(d) deforming the three-dimensional template to the computer-generated 3-dimensional representation to determine the 3-dimensional shape of the medical device. 2. The method as set forth in claim 1, wherein the 3-dimensional image data is obtained from image slices of the target tissue. 3. The method as set forth in claim 1, wherein the 3-dimensional image data is obtained as one or more voxels of the target tissue. 4. The method of claim 1, wherein the medical device is an implant to be implanted in a subject. 5. The method of claim 1, further comprising generating the computer-readable image data of the target tissue. 6. The method of claim 5, further comprising fabricating the 3-dimensional shape of the medical device. 7. The method of claim 1, further comprising fabricating the 3-dimensional shape of the medical device. | Osteoplastics LLC | A61B005/00 G06F017/50 A61B005/107 G06T019/00 A61B005/103 A61F002/30 G06T015/00 A61B019/00 A61F002/28
US09330453 | B2 | 2016-05-03 | Apparatus and method for determining a skin inflammation value | 2012-03-20 | The invention relates to an apparatus and to a method for determining a skin inflammation value. The apparatus comprises an optoelectronic measuring device, preferably a 3D scanner, for recording a three-dimensional image of an inflammation region on human or animal skin, wherein area-related, spatial, and color values of the three-dimensional image can be detected by the optoelectronic measuring device, a computing unit for calculating the skin inflammation value from the area-related, spatial, and color values detected by the measuring device, and a display unit for displaying the calculated skin inflammation value. | 1. A device for determining a skin inflammation score or value, comprising an optoelectronic measuring instrument or device for directly taking a three-dimensional image of an inflammation area on human or animal skin, wherein area-related, space and color values of the three-dimensional image are directly detected by the optoelectronic measuring instrument,a processing or computing unit for calculating the skin inflammation score from the area-related, space and color values detected by the measuring instrument, anda display unit for displaying the calculated skin inflammation score,wherein the processing unit distinguishes the area-related values of the scanned three-dimensional image into an inflammation focus and a focus-surrounding area adjoining and surrounding the inflammation focus by delimiting the color values of the individual pixels and by delimiting the space values of the individual pixels and in that each color value corresponds to a magenta value in the CMYK color model;wherein the scanned three-dimensional image of the inflammation area is comprised of a multitude of pixels arranged in a three-dimensional coordinate system in grid-like fashion, wherein each area-related value corresponds to a single pixel that is unique in the coordinate system;wherein both a single color value and a single space value are assigned to each area-related value of a three-dimensional image scanned by the optoelectronic measuring instrument;wherein each space value corresponds to a height value of the respective pixel in the three-dimensional coordinate system;wherein a relative overall color value of the inflammation focus is determined by comparing the averaged color values in the inflammation focus and the averaged color values in the focus-surrounding area;wherein the optoelectronic measuring instrument or device is a 3D sensor;wherein a relative overall volume value of the inflammation focus is determinable by comparing averaged space values in the inflammation focus to averaged space values in the focus-surrounding area;wherein the relative overall volume value is a comparative value of the surface roughness in the inflammation focus to the surface roughness in the focus-surrounding area. 2. The device according to claim 1, wherein an absolute overall volume value of the inflammation focus is determinable from the space values in the inflammation focus. 3. The device according to claim 1, wherein an area-related value corresponds to a peripheral value corresponding to the periphery of the inflammation focus, and/or an area-related value corresponds to an area value representing the surface area of the inflammation focus. 4. The device according to claim 3, wherein an area-related value is formed as a function of the area value and of the peripheral value and corresponds to a compactness value representing the ratio of the peripheral value to the area value. 5. The device according to claim 1, wherein a relative overall volume value correspo... | RED. SOFT IT-SERVICE GMBH | G06T007/00 A61B005/00
US09737262 | B2 | 2017-08-22 | Wearable medical monitoring device | 2016-01-29 | A wearable monitoring device includes a plurality of cardiac sensing electrodes, a monitor, at least one motion sensor, and a controller. The plurality of cardiac sensing electrodes are positioned outside a body of a subject and to detect cardiac information of the subject. The monitor administers a predetermined test to the subject, and has a user interface configured to receive quality of life information from the subject. The at least one motion sensor is positioned outside the body of the subject and to detect subject motion during the predetermined test. The controller is communicatively coupled to the plurality of cardiac sensing electrodes, the monitor, and the at least one motion sensor, and receives and stores the detected cardiac information, the quality of life information, and the detected subject motion. The controller further communicates the stored detected subject motion and the quality of life information to a remote computer. | 1. A wearable monitoring device, comprising: a plurality of cardiac sensing electrodes configured to be positioned outside a body of a subject and to detect cardiac information of the subject;a monitor configured to administer a predetermined test to the subject and having a user interface configured to receive quality of life information from the subject;at least one motion sensor configured to be positioned outside the body of the subject and to detect subject motion during the predetermined test; anda controller communicatively coupled to the plurality of cardiac sensing electrodes, the monitor, and the at least one motion sensor, the controller configured to receive and store the detected cardiac information, the quality of life information, and the detected subject motion, and to communicate the stored detected subject motion and the quality of life information to a remote computer. 2. The wearable monitoring device of claim 1, wherein the quality of life information identifies at least one of a fatigue level and a breathing level of the subject. 3. The wearable monitoring device of claim 1, wherein the quality of life information identifies information about at least one of the subject's lifestyle, body weight, blood pressure, eating and exercise habits, symptom information, fatigue information, breathing level, and how the subject currently feels. 4. The wearable monitoring device of claim 3, wherein the predetermined test is a six minute walk test. 5. The wearable monitoring device of claim 4, wherein the at least one motion sensor includes at least one accelerometer. 6. The wearable monitoring device of claim 5, wherein the controller determines a distance travelled by the subject during the six minute walk test based upon the detected subject motion from the at least one accelerometer. 7. The wearable monitoring device of claim 6, further comprising at least one additional sensor communicatively coupled to the controller to detect at least one of the subject's blood pressure, blood oxygen level, respiration, and temperature. 8. The wearable monitoring device of claim 7, wherein the controller wirelessly communicates the stored detected subject motion and the quality of life information to the remote computer. 9. The wearable monitoring device of claim 4, wherein the at least one motion sensor includes a pedometer, and wherein the controller determines a distance travelled by the subject during the six minute walk test based upon the detected subject motion from the pedometer. 10. The wearable monitoring device of claim 4, further comprising at least one additional sensor communicatively coupled to the controller to detect at least one of the subject's blood pressure, blood oxygen level, respiration, hypovolemia, and temperature. 11. The wearable monitoring device of claim 4, wherein the quality of life information further includes information about at least one of the subject's lifestyle, body weight, blood pressure, eating an... | ZOLL MEDICAL CORPORATION | A61N001/00 A61B005/00 A61B005/0205 A61B005/11 A61N001/04 A61N001/39 A61B005/021 A61B005/024 A61B005/0408 A61B005/08 A61B005/1455 A61N001/365 A61N001/372
US09687307 | B2 | 2017-06-27 | Navigation system and method for tracking objects using optical and non-optical sensors | 2016-01-13 | Systems and methods that utilize optical sensors and non-optical sensors to determine the position and/or orientation of objects. A navigation system includes an optical sensor for receiving optical signals from markers and a non-optical sensor, such as a gyroscope, for generating non-optical data. A navigation computer determines positions and/or orientations of objects based on optical and non-optical data. | 1. A navigation system comprising: a localizer;a tracker for communicating with said localizer and including three markers and a non-optical sensor; anda computing system having at least one processor configured to: determine positions of each of said markers in a first coordinate system based on optical and non-optical signals,determine a position of a virtual point in said first coordinate system, andmatch said positions of said markers and said virtual point in said first coordinate system with positions of said markers and said virtual point in a model of said tracker established relative to a second coordinate system to obtain a transformation matrix to transform said second coordinate system to said first coordinate system. 2. The navigation system of claim 1, wherein said at least one processor is configured to determine positions of a plurality of said virtual points and match said positions of said markers and said plurality of said virtual points in said model of said tracker to obtain said transformation matrix. 3. The navigation system of claim 2, wherein said plurality of said virtual points comprises a first set of virtual points positioned along each of x, y, and z axes of said second coordinate system. 4. The navigation system of claim 3, wherein said plurality of said virtual points comprises a second set of virtual points positioned along each of said x, y, and z axes of said second coordinate system wherein said second set of virtual points are positioned along said x, y, and z axes on an opposite side of an origin of said second coordinate system from said first set of virtual points. 5. The navigation system of claim 2, wherein said plurality of said virtual points are located radially outwardly from said markers in said model of said tracker. 6. The navigation system of claim 1, wherein said at least one processor is configured to match said positions of said markers and said virtual point in said first coordinate system with positions of said markers and said virtual point in a model of said tracker established relative to said second coordinate system using a rigid body matching algorithm or point matching algorithm. 7. The navigation system of claim 1, wherein said at least one processor is configured to change said position of said virtual point in said model based on movement of said tracker. 8. The navigation system of claim 1, wherein said at least one processor is configured to determine said positions of said markers and said virtual point at a frequency of at least 500 Hz. 9. The navigation system of claim 1, wherein said first coordinate system is further defined as a coordinate system of said localizer and said second coordinate system is further defined as a coordinate system of said tracker. 10. The navigation system of claim 1, wherein said at least one processor is configured to determine a position of a first of said markers at a first time based on a first optical signal from said first marker. 11... | STRYKER CORPORATION | A61B005/05 A61B034/20 A61B005/06 A61B034/30 A61B090/00
US09352166 | B2 | 2016-05-31 | Wearable cardiac defibrillator system sounding to bystanders in patient's own voice | 2014-10-27 | A wearable cardiac defibrillator (“WCD”) system may include a support structure that a patient can wear, an energy storage module that can store an electrical charge, and a discharge circuit that can discharge the electrical charge through the patient so as to shock him or her, while the patient is wearing the support structure. Embodiments may actively take into account bystanders, both to protect them from an inadvertent shock, and also to enlist their help. In some embodiments, the WCD system includes a speaker system and a memory. Prompts have been saved in advance in the patient's own voice, and stored in the memory. In case of an emergency, the prompts may be played by the speaker system in the patient's own voice, and heard by a bystander. | 1. A wearable cardiac defibrillator (WCD) system, comprising: a support structure configured to be worn by a patient;an energy storage module configured to be coupled to the support structure and to store an electrical charge;a discharge circuit configured to be coupled to the energy storage module and configured to discharge the electrical charge through the patient while the support structure is worn by the patient;a memory configured to store at least one data file encoding a sound recorded in the voice of the patient; anda speaker system coupled to the support structure and configured to output a transmitted sound that includes the sound recorded in the voice of the patient. 2. The WCD system of claim 1, in which the sound recorded in the voice of the patient further communicates a request to say: “CLEAR”, “YES”, “ALL CLEAR”, “NO”, or “WAIT”. 3. The WCD system of claim 1, further comprising: a microphone configured to sense the voice of the patient, andin which the sensed voice becomes the recorded sound. 4. The WCD system of claim 1, further comprising: a user interface configured to input the data file, andin which the inputted data file becomes stored in the memory. 5. The WCD system of claim 1, in which the memory is further configured to store another data file encoding another sound in a voice other than the voice of the patient, andthe transmitted sound further includes the other sound. 6. A non-transitory computer-readable storage medium storing one or more programs which, when executed by at least one processor a wearable cardiac defibrillator (WCD) system, the WCD system including a support structure configured to be worn by a patient, a power source, an energy storage module configured to be coupled to the support structure, to be charged from the power source and to store an electrical charge, a discharge circuit configured to be coupled to the energy storage module, a measurement circuit, a speaker system coupled to the support structure and a memory, they result in operations comprising: storing in the memory at least one data file encoding a sound recorded in the voice of the patient;rendering, by the measurement circuit, a physiological input from a physiological signal of the patient while the patient is wearing the support structure;determining from the physiological input whether or not a shock condition is met and, if met, controlling the discharge circuit to discharge the electrical charge through the patient while the support structure is worn by the patient; andoutputting, via the speaker system, a transmitted sound that includes the sound recorded in the voice of the patient after it is determined that the shock condition is met and before the discharge circuit is controlled to discharge the electrical charge through the patient. 7. The non-transitory computer-readable storage medium of claim 6, in which the sound recorded in the voice of the patient further communicates a request to not touch the patient. 8. Th... | WEST AFFUM HOLDINGS CORP. | A61N001/00 A61N001/39 A61B005/00 A61B005/11 A61B005/053
US10922874 | B2 | 2021-02-16 | Medical imaging apparatus and method of displaying medical image | 2015-11-02 | Provided is a medical imaging apparatus. The medical imaging apparatus includes: a user interface configured to receive an input for setting a region of interest (ROI) in a first medical image and an input for setting first volume rendering properties for the ROI and second volume rendering properties for a remaining region of the first medical image other than the ROI; an image generator configured to generate a second medical image by performing volume rendering on the ROI and the remaining region other than the ROI based on the first and second volume rendering properties, respectively; and a display configured to display the second medical image. | 1. A medical imaging apparatus comprising: a user interface configured to receive an input for setting a region of interest (ROI) in a first medical image and an input for setting first volume rendering properties for the ROI and second volume rendering properties for a remaining region of the first medical image, wherein the remaining region is identified as a region other than the ROI set by a user in the first medical image;a display; andone or more processors configured to generate a second medical image by performing volume rendering on the ROI and the remaining region other than the ROI based on the first and second volume rendering properties, respectively, and control the display to display the second medical image,wherein the one or more processors are further configured to: display the ROI in the first medical image and at least one item related to the first volume rendering properties via the display, and receive the input for setting the first volume rendering properties via the user interface while displaying the first medical image and the at least one item related to the first volume rendering properties, anddisplay the remaining region of the first medical image and at least one item related to the second volume rendering properties via the display, and receive the input for setting the second volume rendering properties via the user interface while displaying the second medical image and the at least one item related to the second volume rendering properties. 2. The medical imaging apparatus of claim 1, wherein the one or more processors perform the volume rendering on the ROI and the remaining region other than the ROI based on the first and second volume rendering properties, respectively, and synthesize a volume rendered medical image of the ROI with a volume rendered medical image of the remaining region other than the ROI, thereby generating the second medical image. 3. The medical imaging apparatus of claim 1, wherein the second volume rendering properties comprise transparency. 4. The medical imaging apparatus of claim 1, wherein the second volume rendering properties comprise a color map. 5. The medical imaging apparatus of claim 1, wherein the second volume rendering properties comprise properties for obtaining a contour of the remaining region. 6. The medical imaging apparatus of claim 1, wherein at least one of the first and second volume rendering properties is reset. 7. The medical imaging apparatus of claim 1, wherein the second volume rendering properties comprise at least one selected from the group consisting of a threshold, a gamma curve, a post gain, and an image filter. 8. The medical imaging apparatus of claim 1, wherein the first volume rendering properties comprise at least one selected from the group consisting of transparency, a color map, a threshold, a gamma curve, a post gain, and an image filter. 9. The medical imaging apparatus of claim 1, wherein the first and second medical images comprise ... | SAMSUNG MEDISON CO.. LTD. | G06T015/08 A61B008/00 A61B005/00 A61B008/08 G06T015/04 G06T007/00 G06T019/00
US11484274 | B2 | 2022-11-01 | Systems and methods for monitoring a patient | 2016-01-08 | A system for monitoring a patient includes an inflatable cuff configured to at least partially occlude an artery of the patient, and a sensor configured to determine a first parameter associated with the at least partially occluded artery and to generate an output signal indicative of the first parameter. The system also includes a processor configured to receive the output signal and information indicative of an occlusion efficiency of the cuff. The processor is configured to determine a hemodynamic parameter of the patient based on the output signal and the information. | 1. One or more non-transitory computer-readable storage media storing instructions that, when executed, cause one or more processors to perform a method of monitoring a patient, the method comprising: receiving information comprising an occlusion efficiency particular to a respective inflatable cuff, the respective cuff being of a first type;causing inflation of the respective cuff while the respective cuff is disposed about a limb of the patient, wherein causing inflation of the respective cuff causes the respective cuff to at least partially occlude an artery disposed within the limb;causing a sensor in communication with the one or more processors to determine a parameter associated with the artery of the patient, wherein the parameter is determined while the artery is at least partially occluded by the cuff; anddetermining a hemodynamic parameter of the patient based on the parameter and the occlusion efficiency, wherein the occlusion efficiency: quantifies an effectiveness, relative to other inflatable cuffs of the first type, with which air pressure within the respective cuff is transferred to force on the artery; andis used by the one or more processors as an input value in determining the hemodynamic parameter. 2. The one or more non-transitory computer-readable storage media of claim 1, wherein the method further comprises causing inflation of the cuff to a reference pressure, using an inflation device, while the cuff is disposed about a test structure, andcausing a sensor to determine a pressure applied to the test structure, by the cuff, while the cuff is inflated to the reference pressure, wherein receiving information comprising the occlusion efficiency comprises determining the occlusion efficiency based on the pressure applied to the test structure. 3. The one or more non-transitory computer-readable storage media of claim 1, wherein the method further comprises: receiving information indicative of an identification parameter particular to the cuff, the identification parameter being indicative of at least one of a type, size, design, model, style, part number, model number, or serial number of the cuff, whereinreceiving information comprising the occlusion efficiency comprises retrieving the occlusion efficiency from the one or more computer-readable storage media at least partially based on the identification parameter. 4. The one or more non-transitory computer-readable storage media of claim 1, wherein the method further comprises determining receipt of an incorrect occlusion efficiency, the instructions causing the one or more processors to prohibit a hemodynamic parameter calculation at least partly in response to determining receipt of the incorrect occlusion efficiency. 5. The one or more non-transitory computer-readable storage media of claim 1, wherein the method further comprises: receiving additional information indicative of at least one of an empirically derived variable, a cumulative time associated with inflatin... | Welch Allyn, Inc. | A61B005/00 A61B005/0225 A61B005/022 A61B005/02 A61B005/025
US09339212 | B2 | 2016-05-17 | Bone cutting system for alignment relative to a mechanical axis | 2013-09-13 | A bone cutting system is disclosed that supports one or more bone cuts that are aligned relative to a mechanical axis. The system comprises a first bone cutting jig, a second bone cutting jig, a sensored insert, a bone jig adapter shim, and a device having at least two reference surfaces. The sensored insert includes a three-axis accelerometer to measure position, rotation, and tilt and includes a plurality of sensors to measure a parameter of the muscular-skeletal system. The reference surface device can be an operating table having a first reference surface and a second reference surface that is perpendicular to the first reference surface for referencing the three-axis accelerometer. The bone jig adapter shim can include a tab that fits into a slot of the first or second bone cutting jigs. A remote system receives accelerometer data to calculate offset relative to a mechanical axis. | 1. A bone cutting system for muscular-skeletal joint installation comprising: a bone cutting jig wherein the bone cutting jig is coupled to a bone;a sensored insert configured to measure load and position of load when installed in a prosthetic joint wherein the sensored insert is a prosthetic component, wherein the sensored insert includes a position measurement system, and wherein measurement data is transmitted from the sensored insert;a bone jig adapter shim configured to couple to the sensored insert where the bone jig adapter shim is configured to couple the sensing insert to the bone cutting jig; anda remote system configured to receive the measurement data from the sensored insert wherein the measurement data from the sensored insert is used to align the bone cutting jig. 2. The system of claim 1 where the bone jig adapter shim includes a tab wherein the tab is configured to temporarily retain and align the sensored insert to the bone jig adapter shim. 3. The system of claim 2 where the tab of the bone jig adapter shim couples to a cutting slot of the bone cutting jig. 4. The system of claim 1 wherein the sensored insert is configured to be removed, wherein the sensored insert is configured to be installed in a prosthetic joint, and wherein the sensored insert is configured to measure load and position of load in the prosthetic joint. 5. The system of claim 1 further including an operating table where the sensored insert is configured to be referenced to a first surface of the operating table and where the sensored insert is configured to be referenced to a second surface of the operating table where the second surface is perpendicular to the first surface. 6. The system of claim 1 wherein the position measurement system comprises a three-axis accelerometer and wherein the position measurement system measures position, rotation, and slope. 7. The system of claim 1 wherein the position measurement system comprises a three-axis gyroscope and wherein the position measurement system measures position, rotation, and slope. 8. The system of claim 1 wherein a display coupled to the remote system is configured to display at least one angle measured by the sensored insert that corresponds to a final position of the bone cutting jig, wherein the sensored insert is removed from the bone cutting jig, and wherein the bone is cut to the at least one angle measured by the sensored insert. 9. A bone cutting system for muscular-skeletal joint installation comprising: a bone cutting jig having a slot wherein the bone cutting jig is configured to couple to a bone;a sensored insert configured to measure load and position of load when installed in a prosthetic joint wherein the sensored insert is a prosthetic component, wherein the sensored insert includes a position measurement system configured to measure position, rotation, or tilt and wherein measurement data is transmitted from the sensored insert;a bone jig adapter shim configured to couple to th... | ORTHOSENSOR INC | A61B017/15 A61B005/103 A61B005/00 A61B005/11 A61B017/17 A61F002/46 A61F002/38 A61B005/107 A61B019/00
US20160120441 | A1 | 2016-05-05 | Portable Device for Direct Nasal Respiration Measurement | 2014-11-02 | A portable and wireless connected device, comprising a breathing channel part, a sensor box, a nose piece and a cover, allows measuring air flow of respiration through each nostril. All components are in conjunction with other parts through quick assembly mechanisms. Multiple parameters including air flow speed, volume and respiration rate are determined by analyzing sensors data in real time, thereby allowing diagnosis, monitoring and interactive breathing training in patients with pulmonary diseases. | 1. A portable device for direct nasal respiration measurement, the said device comprising: a breathing channel part, wherein respiration air flow is guided through and portioned;a sensor box to house air flow sensors, a power source, electronic circuits and a wireless communication unit;a nosepiece with one end in communication with nostrils, the other end in clip-in conjunction with said breathing channel part; anda cover for protection. 2. The device of claim 1, wherein the said device measures respiration information of each individual nostril separately. 3. The device of claim 1, wherein the said device wirelessly sends respiration information to a computer, tablet, smart phone or other electronic device for data analysis and graph display. 4. The device of claim 3, where said graph is used for guiding breathing rehabilitation exercise. 5. The device of claim 1, wherein said breathing channel part is made of metal, plastics or other essentially rigid materials. 6. The device of claim 1, wherein said sensor box is made of metal, plastics or other essentially rigid materials. 7. The device of claim 1, wherein said nosepiece is made of silicone, rubber or other substantially soft materials. 8. The device of claim 1, wherein said cover is made of metal, plastics or other essentially rigid materials. 9. The device of claim 1, wherein said breathing channel part comprises: two breathing channels;two protruded connectors at the end of each breathing channel in conjunction with said nosepiece;an orifice structure in each of the channels; andan air flow bypass in each of the channels. 10. The breathing channel part of claim 9, wherein said air flow bypass comprises two slits from the inner wall, a narrow channel and an opening through bottom surface. 11. The device of claim 1, wherein said breathing channel part is washable for disinfection and sanitizing. 12. The device of claim 1, wherein said sensor box comprises: two protrusions from the top surface with cavities for sensors mounting;two MEMS air flow sensors to convert the breathing air flow into electrical signal;two air flow sensor amplifier circuit s to amplify said MEMS air flow sensors signal;a wireless data acquisition module to collect and send said MEMS air flow sensors data to a computer or a portable electronic device;a low battery indicator circuit;a device power on LED;a low battery status LED;a battery; anda switch to turn on and off the said device. 13. The sensor box in claim 12, wherein said MEMS air flow sensors are mounted in cavities of said protrusions and form flat surfaces. 14. The sensor box in claim 12, wherein shape of said protrusion is a cuboid or a cylinder or any prism. 15. The sensor box in claim 12, wherein said MEMS air flow sensors have no electrical contact pad on surface and can be sanitized by wiping with an alcohol wipe. 16. The sensor box in claim 12, wherein said wireless data acquisition module is a Bluetooth module with a function of data ac... |  | A61B005/097 A61B005/00 A61B005/087
US09808194 | B2 | 2017-11-07 | Cushioning support for monitoring user activity | 2015-09-09 | Multi-layer cushion supports are described which may generally comprise a first support having a first contact surface for contacting a portion of a body and a second surface opposite to the first surface, the first support defining a central chamber and a peripheral chamber surrounding the central chamber, wherein the first support is filled with a first gas or liquid and a second support attached to the first support along the first contact surface. The second support may be filled with a second gas or liquid which is relatively more viscous than the first gas or liquid. In particular, the first support may be filled with a volume of air and the second support may be filled with oil which is less than the volume of air. | 1. A monitoring system comprising: a support for supporting at least a portion of a user's body, wherein the support comprises a first layer and a second layer opposite the first layer, wherein the first and second layers define a central chamber and a peripheral chamber surrounding the central chamber, wherein the central chamber and the peripheral chamber are filled with a first gas or liquid, wherein the support further comprises a third layer, wherein the third layer is opposite both the first and second layers, wherein the second and third layers define an additional chamber, and wherein the additional chamber is filled with a second gas or liquid which is more viscous than the first gas or liquid;at least one pressure sensor in communication with the support; andan electronics unit in communication with the at least one pressure sensor,wherein the electronics unit is configured to transmit information relating to pressure sensed from the at least one pressure sensor to a remote device, and wherein the electronics unit and/or remote device is programmed to provide a warning or alert to inform the user that their posture is proper or improper when seated. 2. The system of claim 1, wherein the additional chamber is sized to contain a maximum volume which is less than a maximum volume of the central and peripheral chambers combined. 3. The system of claim 2 wherein the first gas or liquid comprises a volume of air. 4. The system of claim 2 wherein the second gas or liquid comprises a volume of oil. 5. The system of claim 2 further comprising a foam layer positioned adjacent to the first layer. 6. The system of claim 5 further comprising a cover configured to envelope the support and the foam layer. 7. The system of claim 1 wherein the electronics unit comprises a microprocessor. 8. The system of claim 1 wherein the electronics unit is further configured for wireless communication with the remote device. 9. The system of claim 1 wherein the remote device comprises a smart phone, tablet, PDA, laptop, or computer. 10. The system of claim 1 wherein the electronics unit and/or remote device is programmed to track a duration of time upon loading of the support. 11. The system of claim 10 wherein the electronics unit and/or remote device is programmed to provide a warning or alert to the user when the duration of time reaches or exceeds a predetermined level. 12. The system of claim 1 wherein the support comprises at least two fluid chambers separated from one another, each fluid chamber in communication with a corresponding pressure sensor which are in communication with the electronics unit. 13. The system of claim 12 wherein the electronics unit and/or remote device is programmed to determine a pressure differential between the fluid chambers and to provide a warning or alert if the pressure differential exceeds a predetermined level. 14. The system of claim 1 wherein the pressure sensor is further configured to detect and/or monitor o... | PRS Medical Technologies, Inc. | A61B005/00 A61B005/0205 A61B005/103 G01G019/44
US10004419 | B2 | 2018-06-26 | Electrode for measuring living body signal | 2015-09-15 | An electrode device for measuring a living body is herein disclosed. The electrode device for measuring a living body signal includes a body comprising a handle part and a tongs part, the body elastically moving by an external force, an electrode part comprising needles disposed on inner surfaces of the tongs part to face each other, and a wire part connected to one end of the electrode part for a connection with a measuring device. | 1. An electrode device for measuring a living body signal, comprising: a body comprising a handle part and a tongs part, the body elastically moving by an external force;an electrode part comprising needles disposed on inner surfaces of the tongs part to face each other, wherein the needles of the electrode part are disposed on inner surfaces at opposite ends of the tongs part to face each other, such that the needles are fixed to a living body as the tongs part is closed, wherein root portions of the needles protrude outwardly from the opposite ends; anda wire part connected to one end of the electrode part for a connection with a measuring device, wherein the wire part is connected to the one end of the electrode part via a connection of the wire to a first electrode member which connects the root portions of the needles at a first opposite end of the opposite ends. 2. The electrode device of claim 1, wherein the electrode part is arranged to affix to a scalp using a restoring force obtainable by opening the tongs part of the body by the external force applied to the handle part. 3. The electrode device of claim 1, wherein the electrode part is arranged to affix to a living body via a restoring force obtainable by opening the tongs part of the body by the external force applied to the handle part. 4. The electrode device of claim 3, wherein the electrode part is arranged to affix to an outer skin layer of the living body by the restoring force. 5. The electrode device of claim 1, wherein opposite ends of the tongs part are not touched with each other even though the external force is not applied to the handle part. 6. The electrode device of claim 1, wherein the needles of the electrode part are arranged to contact the living body in an angle of less than 90 degrees as the tongs part is closed. 7. The electrode device of claim 1, wherein the wire part is further connected to a second end of the electrode part via a connection of the wire to a second electrode member connecting the root portions of the needles at a second opposite end of the opposite ends. 8. The electrode device of claim 7, wherein the wire part connects the electrode members to obtain one electrode signal, or the electrode members are respectively connected to a plurality of wire parts to obtain a plurality of electrode signals. 9. The electrode device of claim 7, wherein the needles, the electrode members, and the wire part are integrated into one and are detached from or attached to the body. 10. The electrode device of claim 1, wherein the electrode part comprises a plurality of needles formed sharply. 11. The electrode device of claim 1, wherein the electrode part comprises a plurality of needles sharply formed, and wherein root portions of the needles are connected to each other. 12. The electrode device of claim 1, wherein the electrode part is formed of sheet metal, and a plurality of needles are formed at one side of the electrode part so as to be fixed to a... | InBody Co., Ltd. | A61B005/0478 A61B005/00 A61B005/04 A61B005/01 A61B005/0402 A61B005/0476 A61B005/0488 A61B005/053
US10413253 | B2 | 2019-09-17 | Method and apparatus for processing medical image | 2015-11-23 | Provided is a medical image processing apparatus including: a data acquisition unit configured to respectively acquire a plurality of medical images representing an object including at least one target at a plurality of different time points; and an image processor configured to generate, based on the acquired plurality of medical images, a diagnostic image showing a degree of change that has occurred in the at least one target over the plurality of different time points. | 1. A medical image processing apparatus comprising: a data acquisition unit configured to acquire a plurality of medical images representing an object, including at least one bone in the object, corresponding to a plurality of different time points, wherein at least one of the plurality of medical images correspond to one of the plurality of different time points;an image processor configured to: obtain a plurality of bone-extracted images respectively corresponding to the plurality of different time points by extracting the at least one bone from each of the plurality of medical images,generate a plurality of registered images by registering the plurality of bone-extracted images,calculate, based on variation in pixel intensity between corresponding pixels in the plurality of registered images, at least one degree of change that has occurred in the at least one bone over the plurality of different time points,classify the at least one degree of change into one of a plurality of stages based on the variation in pixel intensity, the plurality of stages comprising: a bone formation stage if the variation has a first type value, a bone destruction stage if the variation has a second type value, and a no change stage if the variation is zero, andgenerate, based on the plurality of registered images, a diagnostic image that shows the at least one degree of change in the at least one bone over the plurality of different time points; anda display configured to display a screen including the diagnostic image, in which a portion corresponding to the no change stage in the bone, a portion corresponding to the bone formation stage, and a portion corresponding to the bone destruction stage, are respectively represented by a first color, a second color and a third color, for representing a progression of metastatic bone cancer over time including the plurality of different time points, andwherein the first color, the second color and the third color are different from each from other. 2. The medical image processing apparatus of claim 1, wherein the image processor is configured to distinguish each of the plurality of stages from others of the plurality of stages. 3. The medical image processing apparatus of claim 1, wherein the image processor is configured to register the plurality of medical images to generate the plurality of registered images. 4. The medical image processing apparatus of claim 1, wherein the image processor is configured to distinguish each of the plurality of stages in the diagnostic image by use of at least one of shapes, marks, and lines. 5. The medical image processing apparatus of claim 4, further comprising a user interface unit configured to receive input to select the at least one of the shapes, the marks, and the lines to distinguish each of the plurality of stages. 6. The medical image processing apparatus of claim 1, wherein the at least one degree of change comprises at least one of a physiological change, a position cha... | Samsung Electronics Co., Ltd. | G06K009/00 A61B005/00 A61B005/055 A61B006/03 A61B006/00 G06T011/00 H04N005/202 G16H050/30
US09351833 | B2 | 2016-05-31 | Dynamic annuloplasty ring sizer | 2015-09-28 | A dynamic, adjustable annuloplasty ring sizer can include an adjustable ring replica, which can be adjusted through a range of sizes corresponding to available prosthetic annuloplasty repair ring sizes. Actuation of an adjustment trigger on a handle portion of the ring sizer can displace tension wires that extend through a malleable shaft and through a plurality of articulating segments that form the ring replica. Displacement of the tension wires causes flexion of the joints between adjacent articulating segments, thereby reducing the overall size of the ring replica. Releasing the tension wires can allow an elastic extension wire to act on the ring replica, enlarging the ring replica to its maximum, at-rest size. In this manner, the appropriate size of annuloplasty ring prosthesis can be determined with a single device, without requiring a plurality of static ring sizers that require individual insertion and placement for the conventional trial-and-error sizing methods. | 1. A device for sizing a heart valve annulus, comprising: a proximal handle having an actuator;a shaft coupled to and extending in a distal direction from the handle;an annuloplasty ring replica coupled to a distal end of the shaft configured to be adjusted to a plurality of different sizes, the annuloplasty ring replica being elongated and defining a shape of the annulus being sized and being biased to a maximum size; andat least one elongated member coupled to the actuator in the handle and extending along the shaft to the annuloplasty ring replica, the elongated member being configured, when actuated, to reduce the size of the annuloplasty ring replica to a plurality of different sizes smaller than the maximum size. 2. The device according to claim 1, wherein the annuloplasty ring replica comprises a generally C-shaped ring replica having a first free end and a second free end. 3. The device according to claim 2, wherein the handle couples to a mid-point along the C-shaped ring replica. 4. The device according to claim 1, wherein the handle is malleable and couples to the C-shaped ring replica with a swivel joint. 5. The device according to claim 1, wherein the shaft has a hollow bore and the at least one elongated member comprises a flexible tension member that connects to the actuator, passes through the hollow bore, and extends through aligned channels in a plurality of articulating segments defining an arc of the annuloplasty ring replica, wherein tension on the flexible tension member pivots the articulating segments with respect to one another to reduce the size of the annuloplasty ring replica. 6. The device according to claim 1, wherein the annuloplasty ring replica includes a flexible extension wire extending along its length that biases the annuloplasty ring replica to its maximum size. 7. The device according to claim 1, further comprising a display on the handle configured to provide the size of annuloplasty ring that corresponds to incremental sizes of the annuloplasty ring replica. 8. A device for sizing a heart valve annulus, comprising: a proximal handle having an actuator;a shaft coupled to and extending in a distal direction from the handle;an annuloplasty ring replica coupled to a distal end of the shaft configured to be adjusted to a plurality of different sizes, the annuloplasty ring replica comprising a plurality of articulating segments defining an arc of the annuloplasty ring replica and each segment being pivotable with respect to one or more respective adjacent articulating segment; andat least one elongated member coupled to the actuator in the handle and extending along the shaft to the annuloplasty ring replica, the elongated member being configured, when actuated, to pivot the articulating segments so as to adjust the size of the annuloplasty ring replica to a plurality of different sizes. 9. The device according to claim 8, wherein the annuloplasty ring replica comprises a generally C-shaped ring replica ... | Edwards Lifesciences Corporation | A61B005/00 A61F002/24 A61B005/107 A61B019/00
US20160128158 | A1 | 2016-05-05 | LED ENVIRONMENT ENGINE | 2015-10-21 | An environment engine with LED lighting is disclosed herein. According to one embodiment, a lighting system is disclosed having a sensor that detects a physical condition of a user. The physical condition includes one or more of a heart rate, a body temperature, and a rate of motion. The lighting system further includes a processor that receives an identification of the user via a radio module that communicates with a mobile device. The processor controls a lighting module based on the physical conditions detected by the sensor and the identification of the user. | 1. A system, comprising: a sensor that detects a physical condition of a user, wherein the physical condition includes one or more of a heart rate, a body temperature, and a rate of motion; anda processor that receives an identification of the user via a radio module that communicates with a mobile device, wherein the processor controls a lighting module based on the physical conditions detected by the sensor and the identification of the user. 2. The system of claim 1, wherein the sensor comprises at least one of a microphone, a temperature sensor, a camera, and an infrared sensor. 3. The system of claim 1, wherein the radio module comprises at least one of an RF transceiver, a cellular modem, a Bluetooth transceiver and a WiFi transceiver. 4. The system of claim 1, wherein the processor is operable to transmit information received by the radio module to an external device for storage. 5. The system of claim 1, wherein the lighting module comprises at least one LED controllable by the processor to change an output color of the LED. 6. The system of claim 5, wherein the lighting module comprises color rendering index lighting. 7. The system of claim 1, wherein the lighting module comprises an array of light sources coupled to an LED output module that controllably emits light. 8. A method, comprising: detecting a physical condition of a user using a sensor, wherein the physical condition includes one or more of a heart rate, a body temperature, and a rate of motion;receiving an identification of the user via a radio module that communicates with a mobile device; andcontrolling a lighting module based on the physical conditions detected by the sensor and the identification of the user. 9. The method of claim 8, wherein the sensor further comprises at least one of a microphone, a temperature sensor, a camera, and an infrared sensor. 10. The method of claim 8, wherein the radio module comprises at least one of an RF transceiver, a cellular modem, a Bluetooth transceiver and a WiFi transceiver. 11. The method of claim 8, wherein the lighting module is controlled using an Internet-of-things platform. 12. The method of claim 8, wherein the lighting module is controlled using at least one of a smartphone application and a web portal. 13. The method of claim 8, further comprising transmitting information received by the radio module to an external device for storage 14. The method of claim 8, wherein the lighting module comprises at least one LED having an adjustable output color. 15. The method of claim 14, wherein the lighting module comprises color rendering index lighting. 16. The method of claim 8, wherein the lighting module comprises an array of light sources coupled to an LED output module that controllably emits light. 17. A non-transitory computer readable medium having stored thereon computer-readable instructions, and a processor coupled to the non-transitory computer readable medium, wherein the processor executes the compute... |  | H05B033/08 A61B005/0205 H05B037/02
US10105070 | B2 | 2018-10-23 | Intrauterine access catheter for delivering and facilitating operation of a medical apparatus for assisting parturition | 2015-11-16 | A catheter to access and navigate the uterine cavity through the cervical canal for delivering, deploying, and facilitating operation of a peripheral intrauterine medical apparatus during parturition includes an elongated, flexible tubular catheter body with a handle at the proximal end controlling a steering mechanism and deflecting the distal tip. The catheter includes a lumen that extends the length of catheter body to a distal end region of the catheter. The peripheral medical apparatus is located in the distal end region and is configured to deploy from the distal tip of the catheter into the uterine cavity. A soft catheter tip reduces the likelihood of injury during intrauterine navigation and placement, and sensors associated with the catheter body provide signals that allow for the monitoring of infant, umbilical cord, placenta and uterine muscle activity during labor to assist operation of the peripheral medical apparatus. | 1. A birthing system for augmenting expulsive uterine forces towards a cervical canal during delivery of a fetus from a uterus, the birthing system comprising: a medical apparatus having a balloon body configured to be placed in an intrauterine cavity at a location between an infant and uterine walls, the balloon body configured to transition between a compacted state and an expanded state, wherein when in the expanded state the balloon body applies a force to a base of the fetus in a direction of the cervical canal;a catheter configured to receive the medical apparatus, circumnavigate the fetus and one or more structures within the uterus to place the balloon body in the intrauterine cavity at the location between the infant and the uterine walls, and obtain electrical signals indicative of intrinsic uterine contractions; anda controller coupled to the catheter to receive the electrical signals, and to a source of an agent, the controller configured to: process the electrical signals to detect one of an onset of an intrinsic uterine contraction, an increase in uterine contraction forces, and a decrease in uterine contraction forces,in response to either of a detection of an onset of an intrinsic uterine contraction or an increase in uterine contraction forces, deliver agent from the agent source to the balloon body to at least partially expand the balloon body toward the cervical canal and into contact with the fetus, andin response to a detection of a decrease in uterine contraction forces, discharge agent from the balloon body to at least partially collapse the balloon body away from the fetus and the cervical canal. 2. The birthing system of claim 1, wherein the medical apparatus further comprises a conduit body having a proximal region, a distal region, and an internal lumen, wherein the balloon body is coupled to the distal region of the conduit body and has an internal chamber in fluid communication with the internal lumen of the conduit body. 3. The birthing system of claim 1, wherein the catheter comprises: a handle having a proximal end, a distal end, and a port at the proximal end;an elongated body having a proximal section, a distal section, a port at a distal tip of the distal section; andan inner lumen extending from the distal tip of the elongated body to the port at the proximal end of the handle, the inner lumen configured to slidably receive the medical apparatus through the port at the proximal end of the handle. 4. The birthing system of claim 3, wherein the catheter further comprises: a plurality of sensors at the distal section configured to provide the electrical signals indicative of intrinsic uterine contractions;a connector associated with the handle; anda plurality of electrical wires electrically connecting the plurality of sensors to the controller through the connector associated with the handle. 5. The birthing system of claim 1, wherein: the medical apparatus comprises a pressure sensor configured to obtain an ... | 3VO Medical, Inc. | A61B005/00 A61B005/04 A61B005/06 A61B017/42 A61B005/03 A61B017/00
US09332944 | B2 | 2016-05-10 | Method and system for providing data management in data monitoring system | 2014-01-31 | Method and system for providing a fault tolerant data receiver unit configured with a partitioned or separate processing units, each configured to perform a predetermined and/or specific processing associated with the one or more substantially non-overlapping functions of the data monitoring and management system is provided. | 1. A method, comprising: generating, by an analyte data receiver, a quiet host signal prior to a data reception period; andin response to the generated quiet host signal, completing, by the analyte data receiver, a current processing cycle, if any, and entering, by the analyte data receiver, an inactive mode during the data reception period. 2. The method of claim 1, further including receiving, by the analyte data receiver, one or more analyte related signals during the data reception period. 3. The method of claim 2, wherein the one or more analyte related signals are received from a transmitter communicatively coupled with an analyte sensor, wherein at least a portion of the analyte sensor is in contact with a fluid in a user's body. 4. The method of claim 3, wherein the analyte sensor is a glucose sensor. 5. The method of claim 1, further including receiving, by the analyte data receiver, data associated with an analyte sample. 6. The method of claim 5, wherein the analyte sample is received from a glucose test strip. 7. The method of claim 2, further including processing the one or more analyte related signals by the analyte data receiver. 8. The method of claim 5, wherein the analyte sample includes a blood glucose measurement. 9. The method of claim 1, further including outputting one or more signals associated with a physiological condition including one or more of a directional arrow indicator, a color indicator, or a monitored condition level indicator. 10. The method of claim 9, wherein outputting the one or more signals associated with the physiological condition includes progressively increasing or decreasing an associated outputted signal over a predetermined time period. 11. A method, comprising: completing a first processing cycle and subsequently entering an inactive mode if a predetermined signal is received during the first processing cycle, or entering an active mode and determining whether a second processing cycle is to be performed during a data reception period if the predetermined signal is received other than during the first processing cycle;expediting the second processing cycle, and entering the inactive mode if it is determined that the first processing cycle is to be performed during the data reception period; andperforming the second processing cycle and subsequently entering the inactive mode if it is determined that the first processing cycle is to be performed during the data reception period. 12. The method of claim 11, wherein the second processing cycle includes a housekeeping routine. 13. A method, comprising: performing, using a first processing unit, a first predetermined processing routine, and performing, using a second processing unit, a second predetermined processing routine, wherein the first and the second predetermined processing routines are non-overlapping; andprior to a start of a data reception cycle, placing one of the first or the second processing units in an inactive mode, where... | ABBOTT DIABETES CARE INC. | G08B023/00 A61B005/00 A61B005/145 G06F019/00 G01N033/487
US09345492 | B2 | 2016-05-24 | Shielded capacitor sensor system for medical applications and method | 2014-03-06 | A measurement system for measuring a parameter of the muscular-skeletal system is disclosed. The measurement system comprises a capacitor, a signal generator, a digital counter, counter register, a digital clock, a digital timer, and a data register. The sensor of the measurement system is the capacitor. The measurement system generates a repeating signal having a measurement cycle that corresponds to the capacitance of the capacitor. The capacitor comprises more than one capacitor mechanically in series. Electrically, the capacitor comprises more than one capacitor in parallel. In one embodiment, the capacitor includes a dielectric layer comprising polyimide. A force, pressure, or load is applied to the capacitor that elastically compresses the device. The capacitor is shielded from parasitic coupling and parasitic capacitance. | 1. An orthopedic surgical system comprising: a reamer having a non-planar surface where the reamer is configured to cut bone of a muscular-skeletal system and where the reamer comprises: a sensor coupled to and conforming to the non-planar surface of the reamer;electronic circuitry coupled to the sensor where the electronic circuitry includes a transmitter and where the transmitter is configured to transmit measurement data from the sensor; anda housing where the electronic circuitry is within the housing of the reamer. 2. The orthopedic surgical system of claim 1 further including a remote system having a processor where the remote system receives and processes measurement data from the reamer. 3. The orthopedic surgical system of claim 2 where the reamer is configured to transmit position and trajectory data. 4. The orthopedic surgical system of claim 3 where the remote system is configured to use the position and trajectory data of the reamer to direct a position and trajectory of a tool. 5. The orthopedic surgical system of claim 1 where the sensor is configured to measure a force applied by the reamer to bone. 6. The orthopedic surgical system of claim 5 where the sensor couples to a cutting head of the reamer. 7. The orthopedic surgical system of claim 6 where the sensor is elastically compressible within a predetermined range. 8. The orthopedic surgical system of claim 7 where the non-planar surface is configured to be compliant to transfer the force to the sensor. 9. The surgical system of claim 8 where the sensor is an elastically compressible capacitor. 10. The surgical system of claim 8 further including a shield configured to reduce parasitic coupling to the sensor. 11. An orthopedic surgical system comprising: a reamer configured to cut bone of a muscular-skeletal system where the reamer comprises: a cutting head having a non-planar surface;a sensor coupled to and conforming to the non-planar surface of the cutting head;electronic circuitry coupled to the sensor where the electronic circuitry includes a transmitter to transmit measurement data from the sensor;a housing where the electronic circuitry is in the housing of the reamer; anda remote system having a processor configured to receive and process the measurement data. 12. The orthopedic surgical system of claim 11 where the reamer is configured to transmit position and trajectory data. 13. The orthopedic surgical system of claim 12 where the remote system is configured to use the position and trajectory data of the reamer to direct a position and trajectory of a tool. 14. The orthopedic surgical system of claim 11 where the sensor is configured to measure a force applied by the reamer to bone and where the sensor is elastically compressible within a predetermined range. 15. The orthopedic surgical system of claim 14 where the non-planar surface is configured to be compliant to transfer the force to the sensor. 16. The orthopedic surgical system of claim 15 where ... | ORTHOSENSOR INC. | A61B005/22 A61B017/16 A61B005/00 A61B005/07
US09345401 | B2 | 2016-05-24 | Handheld vision tester and calibration thereof | 2015-04-17 | In one aspect, there is provided an embodiment of a calibration system for use with a vision tester having a display. In this particular embodiment, the calibration system comprises a calibration stand and a reflective surface. The calibration stand is configured to hold the vision tester. The reflective surface is coupled to the calibration stand and is oriented substantially parallel to the display. The calibration stand is configured to hold the display of the vision tester at a fixed distance from the reflective surface. | 1. A calibration system for use with a vision tester having a display, comprising: a calibration stand configured to hold said vision tester; anda reflective surface coupled to said calibration stand and oriented substantially parallel to said display, said stand configured to hold said display at a fixed distance from said reflective surface. 2. The calibration system as recited in claim 1 wherein said reflective surface is configured to reflect images displayed on said display to a camera of said vision tester to allow said vision tester to employ said reflected images to calibrate and verify functionality of said vision tester. 3. The calibration system as recited in claim 2 wherein said images include varying colors, shapes and brightness. 4. The calibration system as recited in claim 2 wherein said images include calibration images. 5. The calibration system as recited in claim 2 wherein said images move. 6. The calibration system as recited in claim 1 wherein said vision tester includes calendar functions configured to remind a user to employ said calibration system to calibrate and verify functionality of said vision tester. 7. The calibration system as recited in claim 1 further comprising at least one security feature configured to protect information and privacy. 8. The calibration system as recited in claim 1 wherein said vision tester includes damage assessment functions configured to prompt a user to employ said calibration system to calibrate and verify functionality of said vision tester. 9. The calibration system as recited in claim 1 further comprising an auxiliary camera configured to capture an image. 10. The calibration system as recited in claim 9 wherein said auxiliary camera is configured to communicate with said vision tester. 11. The calibration system as recited in claim 9 wherein said auxiliary camera is configured to communicate with an external device. 12. A calibration system for use with a vision tester having a display, comprising: a reflective surface oriented and configured to reflect images displayed on said display; anda camera associated with said vision tester and configured to allow said vision tester to employ said reflected images to calibrate and verify functionality of said vision tester. | VITAL ART AND SCIENCE, LLC | A61B003/02 A61B003/032 A61B003/00 A61B005/145 H04N005/351
US09326698 | B2 | 2016-05-03 | Method for automatic, unsupervised classification of high-frequency oscillations in physiological recordings | 2012-02-17 | Oscillatory signals may be used to determine a region of a patient's body that is associated with a medical condition. As described herein, oscillatory signals may be detected using a high sensitivity, low specificity detector. The oscillatory signals may be representative of discrete events in a patient's body. The detected signals may be tested in the context of surrounding background activity to identify anomalous discrete physiologic events that are sufficiently different from the surrounding background activity. The anomalous discrete physiologic events having correlative morphological, time, or location characteristics may be automatically clustered and clusters of anomalous physiologic events may be determined that are indicative of at least one region of the patient's body that is associated with a medical condition. | 1. A method for automatically detecting, classifying, and processing oscillatory signals for use as a biomarker for a physiologic physiological event in a patient, the method comprising: detecting and classifying oscillatory signals representative of discrete physiologic events in a patient's body using a high sensitivity, low specificity detector;processing, using a processor, the detected signals by: testing the detected signals in the context of surrounding background activity to identify anomalous discrete physiologic events that are sufficiently different from the surrounding background activity; andautomatically clustering said anomalous discrete physiologic events into clusters of anomalous physiologic events having correlative morphological, time, or location characteristics;determining at least one cluster of anomalous physiologic events as a biomarker indicative of at least one region of the patient's body that is associated with a medical condition; andproviding treatment, using a treatment device responsive to said biomarker, of said medical condition to the at least one region of the patient's body identified by said biomarker. 2. The method of claim 1, wherein the discrete physiologic events are associated with the patient's brain, wherein the at least one region of the patient's body comprises at least one region of the brain, and wherein the medical condition comprises epileptic seizure. 3. The method of claim 2, comprising the further steps of implanting said high sensitivity, low specificity detector in the patient's body and processing outputs of said high sensitivity, low specificity detector to localize seizure onset channels of said patient. 4. The method of claim 1, wherein each discrete physiologic event is a high frequency oscillation event. 5. The method of claim 4, wherein each discrete physiologic event is detected by bandpass filtering and thresholding electroencephalographic recordings in at least one of a time domain, a frequency domain, or a domain of another set of applicable basis functions. 6. The method of claim 5, comprising the further step of determining a power ratio in the frequency domain by estimating power in a fast ripple band to that in a ripple band. 7. The method of claim 5, comprising the further step of determining a spectral centroid in the frequency domain by distinguishing physiologic events having energy concentrated in a ripple sub-band or a fast ripple sub-band alone from those with components in the ripple and fast ripple band, with relatively broadband spectra, or with predominant energy in an intermediate band. 8. The method of claim 5, comprising the further step of determining a line length in the time domain by detrending and normalizing detections. 9. The method of claim 5, comprising the further step of determining a spectral peak in the frequency domain by identifying a frequency corresponding to a peak of an estimated power spectral density. 10. The method of claim 5, comp... | The Trustees of the University of Pennsylvania | A61B005/04 A61B005/0476 A61B005/048 A61B005/00
US09351642 | B2 | 2016-05-31 | Non-contact optical system, computer-accessible medium and method for measurement at least one mechanical property of tissue using coherent speckle technique(s) | 2010-03-12 | Exemplary embodiments of apparatus and method for determining at least one material property of an anatomical structure can be provided. According to one exemplary embodiment, it is possible to apply at least one first coherent radiation to at least one portion of the anatomical structure, and receive at least one second coherent radiation from such portion(s). The first and second coherent radiations can be associated with one another. In addition, it is possible to determine the material property based on the second coherent radiation(s). Such determination can be performed without (i) any portion of an apparatus performing the procedure causing an induction of at least one mechanical deformation on or in the anatomical structure, and/or (ii) any mechanical deformation on or in the anatomical structure. | 1. An apparatus, comprising: at least one first source arrangement which is configured to apply at least one first coherent radiation to at least one portion of an anatomical structure; andat least one second arrangement includes a detector which is configured to receive at least one second coherent radiation from the at least one portion, the first and second coherent radiations being (i) associated with one another and (ii) time varying,wherein the at least one second arrangement is further includes a computer which is configured to determine at least one material property of the anatomical structure which is an viscoelastic modulus, an elastic modulus or a viscous modulus of the anatomical structure based on the at least one second coherent radiation, andwherein the determination by the computer is performed without at least one of:i. any portion of the apparatus causing an induction of at least one mechanical deformation on or in the anatomical structure, orii. a requirement for the mechanical deformation on or in the anatomical structure. 2. The apparatus according to claim 1, wherein at least one of the at least one first coherent radiation or the at least one second coherent radiation is an electro-magnetic radiation. 3. The apparatus according to claim 1, wherein the at least one first arrangement is configured to scan the anatomical structure at multiple locations. 4. The apparatus according to claim 3, wherein the at least one first arrangement scans the anatomical structure at the multiple locations simultaneously. 5. The apparatus according to claim 3, wherein the at least one first arrangement scans the anatomical structure at the multiple locations sequentially. 6. The apparatus according to claim 3, wherein the detector is further configured to detect a scan of the anatomical structure at the multiple locations simultaneously. 7. The apparatus according to claim 3, wherein the detector is further configured to detect a scan of the anatomical structure at the multiple locations sequentially. 8. The apparatus according to claim 1, wherein the computer determines the viscoelastic modulus or the viscosity viscous modulus that is spatially-varying or depth-varying. 9. The apparatus according to claim 1, wherein the at least one second arrangement is configured to determine the at least one material property that is a macroscopic property of the anatomical structure. 10. The apparatus according to claim 1, wherein the at least one second arrangement is configured to determine the at least one material property that is a microscopic property of the anatomical structure. 11. The apparatus according to claim 1, wherein the at least one second arrangement is configured to determine the at least one material property that is a mesoscopic property of the anatomical structure. 12. The apparatus according to claim 1, wherein the at least one second arrangement is configured to determine the at least one material property as a function... | The General Hospital Corporation | A61B006/00 A61B005/00 A61B005/02
US09352136 | B2 | 2016-05-31 | Drug delivery and substance transfer facilitated by nano-enhanced device having aligned carbon nanotubes protruding from device surface | 2015-05-11 | The present invention relates to a nano-enhanced device for substance transfer between the device and a tissue. The device comprises a substrate with substantially aligned carbon nanotubes anchored within the substrate, and with at least one end of the carbon nanotubes protruding from the substrate. The protruding nanotube ends may be coated with a drug for delivery of the drug into body tissue. The present invention may be incorporated into an angioplasty catheter balloon or into a patch that is worn on the skin. The carbon nanotubes can be grouped in clusters to effectively form nano-needles which can transfer fluid to or from the subdermal tissue. The nano-needles can be used in conjunction with a sensor to ascertain body fluid information such as pH, glucose level, etc. | 1. A method of transferring a substance, comprising: (a) applying a device to a body of a patient, the device comprising a substrate and an array of carbon nanotubes with each carbon nanotube in the array having two ends, wherein the array of carbon nanotubes is anchored within the substrate with at least one end of each carbon nanotube protruding from the substrate,wherein the carbon nanotubes within the array are grouped in clusters such that each cluster effectively functions as a needle,wherein the clusters are arranged as a plurality of rings with all of the carbon nanotubes of the clusters being in a ring periphery, andwherein application of the device to the body causes the clusters to enter tissue of the patient's body; and(b) allowing the transfer of the substance through the substrate and into the tissue. 2. The method of claim 1, wherein the substrate comprises a porous material and the substance is transferred into the tissue via inherent capillary action of the carbon nanotubes. 3. The method of claim 1, wherein the substrate is loaded with the substance to be transferred to the tissue. 4. The method of claim 1, wherein the carbon nanotubes are coated with the substance transferred to the tissue. 5. The method of claim 1, wherein the device is a patch comprising an adhesive, and wherein the patch is applied to the body of the patient such that the adhesive adheres to the body. 6. The method of claim 5, wherein the adhesive coats only a perimeter portion of the patch with the protruding nanotube clusters located on a portion interior to the perimeter portion. 7. The method of claim 5, wherein the tissue is epidermal tissue. | CALIFORNIA INSTITUTE OF TECHNOLOGY | A61M037/00 A61K047/48 A61L027/44 A61L027/54 A61L027/56 B29C070/64 B82Y005/00 B82Y030/00 B82Y040/00 C01B031/02 A61B005/145 A61M005/44 A61M025/10 A61N001/30 A61N005/06 B29K105/16
US09339200 | B2 | 2016-05-17 | Systems and methods for determining blood flow characteristics using flow ratio | 2015-07-20 | Embodiments include a system for determining cardiovascular information for a patient which may include at least one computer system configured to receive patient-specific data regarding a geometry of an anatomical structure of a patient; create a model representing at least a portion of the anatomical structure; create a physics-based model relating to a blood flow characteristic within the anatomical structure; determine a first blood flow rate at at least one point of interest in the model; modify the model; determine a second blood flow rate at a point in the modified model corresponding to the at least one point of interest in the model; and determine a fractional flow reserve value as a ratio of the second blood flow rate to the first blood flow rate. | 1. A system for determining cardiovascular information for a patient, the system comprising: at least one computer system configured to: receive patient-specific image data regarding a geometry of an anatomical structure of a patient;create an anatomical model representing at least a portion of the anatomical structure of the patient based on the patient-specific image data;determine at least one point of interest of vasculature of the anatomical model;determine a first blood flow rate through the at least one point of interest of the vasculature;dilate vessel geometry of the vasculature at the at least one point of interest;determine a second blood flow rate resulting from the dilating of the vessel geometry at the at least one point of interest; anddetermine a flow rate ratio of the second blood flow rate to the first blood flow rate. 2. The system of claim 1, wherein the anatomical model is a three-dimensional model. 3. The system of claim 1, wherein dilating the vessel geometry comprises removing one or more anatomic restrictions or plaque proximal to the at least one point of interest. 4. The system of claim 1, wherein the anatomical model representing at least the portion of the patient's anatomical structure includes at least a portion of an aorta and at least a portion of a plurality of coronary arteries emanating from the portion of the aorta. 5. The system of claim 4, wherein the at least one computer system is configured to determine the flow rate ratio at a plurality of locations in the plurality of coronary arteries. 6. The system of claim 1, wherein the at least one computer system is configured to create the anatomical model based on the imaging data by locating boundaries of lumens of coronary arteries of the patient's heart using the imaging data. 7. The system of claim 1, wherein the at least one computer system is further configured to: create a physics-based model representing a blood flow through the anatomical model, wherein the first blood flow rate or the second blood flow rate are determined using the created physics-based model. 8. The system of claim 1, wherein the at least one computer system is configured to determine the first blood flow rate at the at least one point of interest using a parameter associated with at least one of a level of hyperemia, a level of exercise, or a medication. 9. The system of claim 8, wherein the at least one computer system is configured to determine the first blood flow rate at the at least one point of interest using a parameter associated with the level of hyperemia, and the parameter relates to a coronary artery resistance of the patient, an aortic blood pressure of the patient, or a heart rate of the patient. 10. A method for determining patient-specific cardiovascular information using at least one computer system, the method comprising: receiving, at least one computer system, patient-specific image data regarding a geometry of an anatomical structure of a patient;creatin... | HeartFlow, Inc. | G06F019/10 G06F019/00 A61B005/02 G06K009/00 G06K009/36 A61B005/029 G06F017/50 G06F017/10 A61B005/107 G06T019/00 A61B005/00 A61B005/021 A61B005/026
US09333350 | B2 | 2016-05-10 | Psychiatric disorder therapy control | 2009-04-17 | A therapy system for managing a psychiatric disorder of the patient may be controlled based on a patient mood state. Therapy may be delivered to a patient according to a therapy program, and a physiological parameter of the patient may be monitored during or after therapy delivery. The patient mood state may be determined based on the monitored physiological parameter, and the therapy delivery may be controlled based on the determined mood state. In some embodiments, the therapy delivery is stopped prior to determining the patient mood state and the therapy delivery is restarted upon detecting a negative mood state. In other embodiments, therapy delivery is delivered until a positive mood state is detected, at which point the therapy delivery may be stopped. | 1. A method comprising: delivering psychiatric disorder therapy to a patient via a medical device according to a therapy program;monitoring a physiological parameter of the patient in response to the psychiatric disorder therapy, wherein the physiological parameter comprises at least one of a respiratory rate, electrodermal activity, thermal activity or muscle activity;determining a patient mood state based on the physiological parameter; andcontrolling the delivery of the psychiatric disorder therapy based on the determined patient mood state, wherein controlling the delivery of the psychiatric disorder therapy comprises stopping the delivery of therapy by the medical device based on the determined patient mood state. 2. The method of claim 1, wherein delivering psychiatric disorder therapy to the patient comprises delivering electrical stimulation to a brain of the patient. 3. The method of claim 1, wherein delivering psychiatric disorder therapy to the patient comprises delivering the psychiatric disorder therapy to the patient during a period of time and determining the patient mood state comprises determining the patient mood state after the period of time. 4. The method of claim 3, wherein controlling the delivery of the psychiatric disorder therapy further comprises initiating the delivery of the psychiatric disorder therapy after the period of time. 5. The method of claim 4, further comprising modifying a value of at least one therapy parameter defined by the therapy program based on the patient mood state, wherein initiating the delivery of the psychiatric disorder therapy comprises initiating the delivery of the psychiatric disorder therapy according to the modified value of the at least one therapy parameter. 6. The method of claim 1, wherein monitoring the physiological parameter of the patient comprises monitoring the physiological parameter of the patient while the psychiatric disorder therapy is delivered to the patient. 7. The method of claim 1, further comprising modifying a value of at least one therapy parameter defined by the therapy program based on the patient mood state. 8. The method of claim 1, further comprising receiving patient feedback to the delivery of the psychiatric disorder therapy according to the therapy program, wherein determining the patient mood state comprises determining the patient mood state based on the physiological parameter and the patient feedback. 9. The method of claim 1, further comprising: receiving patient feedback to the delivery of the psychiatric disorder therapy according to the therapy program;determining whether the patient feedback is consistent with the mood state determined based on the monitored physiological parameter; andgenerating an indication if the patient feedback is not consistent with the mood state determined based on the monitored physiological parameter. 10. The method of claim 1, wherein determining the patient mood state comprises: receiving a signal from a sen... | Medtronic, Inc. | A61N001/00 A61N001/36 A61B005/16 A61B005/01 A61B005/0488 A61B005/08 A61B005/00
US09327110 | B2 | 2016-05-03 | Devices, systems and methods for the targeted treatment of movement disorders | 2012-02-02 | Devices, systems and methods are provided for the targeted treatment of movement disorders. Typically, the systems and devices are used to stimulate one or more dorsal root ganglia while minimizing or excluding undesired stimulation of other tissues, such as surrounding or nearby tissues, ventral root and portions of the anatomy associated with body regions which are not targeted for treatment. The dorsal root ganglia are utilized in particular due to their specialized role in movement. It is in these areas that sensory fibers are isolated from motor fibers. Sensory fibers are involved in a variety of reflexes that are involved in movement control, and these reflexes can be utilized in the treatment of various movement disorders. Thus, by stimulating sensory fibers in these areas, fundamental reflexes can be affected to lessen the symptoms of movement disorders. | 1. A method of treating a movement disorder of a patient using selective stimulation of a soma within a dorsal root ganglion associated with an involved sensory neuron of the movement disorder, comprising: positioning an introducing needle into an epidural space of the patient;advancing a sheath through the introducing needle and along the epidural space of the patient to a spinal level having the dorsal root ganglion related to the involved sensory neuron of the movement disorder;advancing a lead through the needle and sheath toward the dorsal root ganglion so that the at least one electrode is disposed near the dorsal root ganglion; andproviding stimulation energy to the at least one electrode so as to stimulate at least a portion of the dorsal root ganglion to access a lower motor neuron via a reflex arc in a manner which reduces a symptom of the movement disorder at least partially caused by damage or disease affecting upper neurons. 2. A method as in claim 1, wherein the movement disorder includes Parkinson's Disease, Multiple Sclerosis or a Demylenating Movement Disorder. 3. A method as in claim 1, wherein the movement disorder includes Cerebral Palsy, Chorea, Dystonia, Spasm, Tic disorder or Tremor. 4. A method as in claim 1, wherein providing stimulation energy comprises adjusting at least one signal parameter to reduce the symptom of the movement disorder. 5. A method as in claim 4, wherein adjusting the at least one signal parameter comprises adjusting a frequency of the stimulation energy. 6. A method as in claim 5, wherein adjusting a frequency of the stimulation energy comprises selecting a frequency less than or equal to approximately 100 Hz. 7. A method as in claim 6, wherein adjusting a frequency of the stimulation energy comprises selecting a frequency less than or equal to approximately 50 Hz. 8. A method as in claim 1 wherein providing stimulation energy comprises adjusting at least one signal parameter provided to the at least one electrode to inhibit improper action potentials directed to the lower motor neuron. 9. A method as in claim 8 wherein the improper action potentials directed to the lower motor neuron causes a synergistic muscle pair involved in the movement disorder to contract. 10. A method as in claim 8 wherein the improper action potentials directed to the lower motor neuron causes an antagonistic muscle involved in the movement disorder to relax. 11. A method as in claim 1 wherein providing stimulation energy comprises adjusting at least one signal parameter provided to the at least one electrode based on an input from one or more sensors configured to sense a status of a symptom of the movement disorder of the patient. 12. A method as in claim 11 wherein the symptom is the onset of the symptoms of the movement disorder. 13. A method as in claim 12, wherein the stimulation signal is provided to reduce or avoid the onset of the symptom. 14. A method as in claim 12, wherein the stimulation signal is p... | St. Jude Medical Luxembourg Holdings SMI S.A.R.L. (“SJM LUX SMI”) | A61N001/36 A61N001/05 A61B005/00 A61B005/0488 A61B005/11
US09330452 | B2 | 2016-05-03 | Statistical mapping in an optoacoustic imaging system | 2013-03-11 | Electromagnetic energy is deposited into a volume, an acoustic return signal from energy deposited in the volume is measured, and a parametric map that estimates values of at least one parameter as spatially represented in the volume is computed. A reference level of a region of interest is determined, and upper and lower color map limits are specified, with at least one of them being determined in relation to the reference level. The parametric map is then rendered in the palette of a color map by mapping the estimated values of the parametric map onto the color map according to the color map limits. Two wavelengths of energy can be applied to the volume, and the parametric map computation can be adapted by applying an implicit or explicit model of, or theoretical basis for, distribution of electromagnetic energy fluence within the volume pertaining to the two wavelengths. The actual electromagnetic energy fluence caused by each wavelength has a propensity, due to variability within the volume, to differ from the modeled or theoretical electromagnetic energy fluence. | 1. A method to produce a qualitative image, comprising: depositing electromagnetic energy into a volume, where a spatial profile of the deposited energy is subject to variability;measuring acoustic return signals from energy deposited in the volume;computing a parametric map comprising estimated values of at least one parameter as spatially represented in the volume;producing a mapping configuration to adjust for effects resulting from at least some of the variability in the spatial profile of the deposited energy by: i) determining at least one reference level within at least one region of interest of the parametric map, the step of determining at least one reference level comprising using a particular mathematical characteristic of the values of the parametric map within the at least one region of interest; andii) specifying an upper color map limit and a lower color map limit, the upper color map limit and the lower color map limit being determined in relation to the at least one reference level, wherein the upper color limit is the reference level plus a first constant times a scalar result of a statistical function applied to the values of the parametric map within the at least one region of interest, and the lower color limit is the reference level minus a second constant times the scalar result of the statistical function applied to the at least one region of interest; and,rendering the parametric map in the palette of a color map by mapping the estimated values of the parametric map onto the color map according to the upper and lower color map limits of the mapping configuration to produce a rendered parametric map. 2. The method of claim 1, in which the color map consists of at least one channel and the mapping step comprises mapping numerical values of the parametric map onto each channel of the color map according to interpolating between the upper and lower color map limits. 3. The method of claim 1, in which the particular mathematical characteristic for determining the reference level is the mean of numerical values of pixels within the region of interest. 4. The method of claim 1, in which the particular mathematical characteristic for determining the reference level is a statistical function. 5. The method of claim 1, further comprising a step of masking the rendered map with a transparency. 6. The method of claim 5, further comprising a step of co-registering the rendered map with an ultrasound image. 7. The method of claim 1, further comprising a step of defining the at least one region of interest of the parametric map. 8. The method of claim 7, in which the step of defining at least one region of interest comprises receiving user input specifying the region of interest. 9. The method of claim 8, in which the step of receiving user input comprises receiving, via a graphical user interface, input defining a top and bottom depth. 10. The method of claim 9, in which the input defining a top and bottom depth comprises inpu... | Seno Medical Instruments, Inc. | G06T007/00 A61B008/00 A61B005/00
US09345413 | B2 | 2016-05-24 | Heart rate extraction using neural wavelet adaptive gain control and neural pattern processing | 2013-08-12 | An improved heart rate monitor is provided that can detect and distinguish a heartbeat from an otherwise contaminated system with noise components potentially larger than the signal of interest. Embodiments of the inventive monitor have an amplification system that eliminates large noise components so as not to saturate the system during detection of a desired low amplitude signal. In embodiments the elimination of noise components is accomplished through wavelet decomposition, and the removal of undesired components including interference components during adaptive gain control (AGC), in addition to hunting algorithms which minimize the error with techniques such as neural network least mean squares type back propagation algorithms. | 1. A heart rate measurement device comprising: a wavelet template;an instrumentation amplifier;a set of two electrodes configured to be coupled to a subject's skin, and in electrical communication with and providing a set of signals to said instrumentation amplifier;an adaptive gain control (AGC) algorithm;an analog or digital filter in a feedback path of said instrumentation amplifier for at least one of removal of saturating signals and unwanted waveform components responsive to said AGC algorithm;wherein said AGC algorithm utilizes said wavelet template and responds to the correlation of said wavelet template to a measured signal obtained from said set of signals by said instrumentation amplifier; andwherein said device is configured to capture said set of signals utilizing a high common-mode resection ratio (CMRR) termination, which does not include a DC path to ground or overload said electrodes. 2. The device of claim 1 wherein said instrumentation amplifier further comprises: two input amplifiers, where a level of gain is set by at least one of a variable resistor or impedance between the two input amplifiers, as well by feedback components around the two input amplifiers; anda summing amplifier. 3. The device of claim 1 wherein said instrumentation amplifier includes a common mode driver, a common mode accelerator, and an active feedback network. 4. The device of claim 1 wherein said instrumentation amplifier is configured for baseline capture by changing a set of feedback components such that saturation recovery is accelerated. 5. The device of claim 4 wherein the baseline capture is obtained with a third electrode driven to create a baseline between said set of signals from said two electrodes placed a minimum impedance away from each other. 6. The device of claim 1 wherein said measured signal is sent wirelessly to the cloud or a remote server for digital processing. 7. The device of claim 1 wherein said instrumentation amplifier precedes a hardware bessel filter. 8. The device of claim 1 wherein said measured signal is analyzed by a wavelet decomposition that builds up a three or more dimensional image of said measured signal. 9. The device of claim 8 wherein said three or more dimensional image is analyzed with pattern recognition methods to determine heart beat components responsive to a training algorithm. 10. The device of claim 9 wherein said pattern recognition methods are a neural network trained to recognize heart beat components from said measured signal. 11. The device of claim 10 wherein a periodic loss of signal is recovered through a prediction based on said neural network, where a rate or waveform may be reconstructed from discontinuous snapshots of heart beat trains. 12. The device of claim 1 wherein said wavelet is in resonance with one or more of the P, QRS and/or T reference components of a heart beat signal. 13. The device of claim 12 wherein said wavelet is configured to meet the following: transform an... |  | A61B005/0452 A61B005/04 A61B005/00 A61B005/0245
US09326714 | B2 | 2016-05-03 | Analyte monitoring device and methods of use | 2010-06-29 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. An analyte sensor unit comprising: a housing adapted for placement on the surface of skin having a bottom surface for contacting with the skin and wherein the housing comprises: an electrochemical sensor comprising a plurality of electrodes including a working electrode, wherein the sensor comprises a portion within the housing and a portion exterior to the housing to permit insertion of the electrochemical sensor to a depth of about 2 to about 7 mm beneath the surface of the skin, wherein the electrochemical sensor comprises a substrate comprising all of the plurality of electrodes;sensor electronics comprising an RF communication unit for communicating data obtained by the electrochemical sensor, the sensor electronics coupled to the electrochemical sensor through the electrical contacts, wherein the sensor electronics comprise electrical contacts that are positioned between at least a portion of the electrochemical sensor and the bottom surface of the housing; anda readable or readable/writable memory for storing factory determined calibration information for calibrating the electrochemical sensor, wherein the electrochemical sensor does not require user-implemented calibration following insertion of a portion of the sensor beneath the surface of the skin; andan adhesive disposed on the bottom surface of the housing to attach the housing to the surface of the skin. 2. The analyte sensor unit of claim 1, wherein the electrochemical sensor is inserted beneath the surface of the skin at a position located substantially beneath the housing when the housing is placed on the skin. 3. The analyte sensor unit of claim 1, wherein the electrochemical sensor has a length of about 0.3 to about 5 cm and comprises a proximal end that is in contact with electrical contacts within the housing and a distal end that extends through an opening at the bottom of the housing. 4. The analyte sensor unit of claim 3, wherein at least a portion of the distal end of the electrochemical sensor is bent so that a portion of the sensor is in a plane substantially perpendicular to the electrical contacts in the housing. 5. The analyte sensor unit of claim 3, wherein the distal end of the electrochemical sensor has a length of up to about 6 mm. 6. The analyte sensor unit of claim 3, wherein the electrochemical sensor has a length of up to about 2 cm. 7. The analyte sensor unit of claim 1, wherein the housing comprises a data storage unit to store analyte data obtained by the electrochemical sensor. 8. The analyte sensor unit of claim 1, wherein the analyte sensor unit automatically detects use following insertion by transitioning from a sleep mode to an active mode. 9. The analyte sensor unit of claim 8, wherein automatically detecting use comprises detecting a change in resistance in the electrochemical sensor. 10. The analyte sensor unit of claim 9, wherein automatically detecting use further comprises activating the analyte sensor unit. 11. The analyte sensor ... | Abbott Diabetes Care Inc. | A61B005/1486 A61B005/1473 A61B005/1495 A61B005/01 C12Q001/54 A61M005/172 A61B005/145 H01L023/31 H01L023/495 G06F019/00 A61B005/00
US09339235 | B2 | 2016-05-17 | Methods and systems for determining signal-to-noise information from a physiological signal | 2012-09-11 | A physiological monitoring system may determine physiological information, such as physiological rate information, and other information, such as signal-to-noise information, from a physiological signal. The system may generate at least one difference signal based on the physiological signal and sort the at least one difference signal to generate at least one sorted difference signal. The system may analyze the at least one sorted difference signal to determine at least two values indicative of noise. The system may determine a value indicative of a signal-to-noise ratio based on the two or more values indicative of noise. | 1. A subject monitoring pulse oximetry system for determining physiological information of a subject, comprising: a pulse oximetry sensor configured to detect light attenuated by the subject and generate a photoplethysmographic (PPG) signal based on the detected attenuated light; anda pulse oximeter coupled to the pulse oximetry sensor, wherein the pulse oximeter is configured to: store a rate algorithm;generate at least one difference signal based on the PPG signal;sort the at least one difference signal in an order based on the values of the difference signal to generate at least one sorted difference signal;analyze the at least one sorted difference signal to determine at least two values indicative of noise;determine a value indicative of a signal-to-noise ratio based on the two or more values indicative of noise;determine an algorithm setting based on the value indicative of a signal-to-noise ratio;apply the algorithm setting to the rate algorithm stored by the pulse oximeter; anddetermine the physiological information of the subject using the rate algorithm and the PPG signal, wherein the physiological information comprises pulse rate of the subject. 2. The system of claim 1, wherein the pulse oximeter is configured to determine the value indicative of a signal-to-noise ratio using a reference of associated values indicative of noise and values indicative of signal-to-noise ratios. 3. The system of claim 1, wherein the pulse oximeter is further configured to: analyze the at least one sorted difference signal to determine three values indicative of noise;select the maximum value of the three values; anddetermine the value indicative of a signal-to-noise ratio based on the maximum value. 4. The system of claim 1, wherein the pulse oximeter is further configured to: determine a best fit line based on the at least one sorted difference signal;generate at least one threshold based on the best fit line; anddetermine a number of points of the at least one sorted difference signal that exceed the at least one threshold, wherein at least one value of the at least two values indicative of noise is based on the number of points. 5. The system of claim 1, wherein the pulse oximeter is further configured to: generate at least one second difference signal based on the PPG signal;sort the at least one second difference signal in an order based on the values of the second difference signal to generate at least one second sorted difference signal;determine a first best fit line based on the sorted difference signal;determine a second best fit line based on the second sorted difference signal; andcompare the first best fit line and the second best fit line to determine at least one value of the at least two values indicative of noise. 6. A pulse oximetry processing module for determining physiological information of a subject, wherein the pulse oximetry processing module is configured to: store a rate algorithm;receive a photoplethysmographic (PPG) sign... | Covidien LP | G06F007/00 A61B005/00 A61B005/024 A61B005/1455
US09351682 | B2 | 2016-05-31 | Sensor interface system | 2011-11-09 | A sensor interface system for providing a connection between at least one sensor and a maternal-fetal monitor, wherein the interface system converts electrical muscle activity captured by the sensor(s) into uterine activity data signals for use by the maternal-fetal monitor. The sensor interface system of the invention preferably includes a conversion means for converting the signals from the sensor(s) into signals similar to those produced by a tocodynamometer. | 1. An interface system for interfacing a plurality of electrodes with a maternal-fetal monitor, wherein the maternal-fetal monitor has a port configured to operably connect with, receive output from, and display output from at least one of a tocodynamometer, an intrauterine pressure catheter, or a fetal scalp electrode, wherein the interface system comprises: a sensor array dimensioned to attach to a maternal abdomen, wherein the sensor array comprises the plurality of electrodes and wherein each of the plurality of electrodes is configured to detect a physiological signal;a first component comprising: an electrode interface configured to operably and removably connect with the sensor array and to receive the physiological signals from the plurality of electrodes; anda wireless signal transmitter configured to receive the physiological signals from the electrode interface and to wirelessly transmit the received physiological signals to a wireless signal receiver located remote from the first component; anda second component comprising: the wireless signal receiver configured to receive the physiological signals transmitted by the wireless signal transmitter;a signal converter configured to receive the physiological signals from the wireless signal receiver and to process the physiological signals into output data that mimics electrical output from the at least one of the tocodynamometer, the intrauterine pressure catheter, or the fetal scalp electrode; anda maternal-fetal monitor port interface configured to operably and removably connect to the port of the maternal-fetal monitor, wherein the maternal-fetal monitor port interface is configured to receive the output data from the signal converter and to communicate to the maternal-fetal monitor the output data when the maternal-fetal monitor port interface is connected to the port of the maternal-fetal monitor. 2. The interface system of claim 1, wherein the sensor array includes a curved arm with one of the plurality of electrodes located at a distal region of the curved arm. 3. The interface system of claim 2, wherein the sensor array includes a plurality of curved arms, each having one of the plurality of electrodes located at a distal region of the curved arm. 4. The interface system of claim 1, wherein the sensor array includes a portion having a serpentine shape that is configured to allow the sensor array to conform to a shape of the maternal abdomen. 5. The interface system of claim 4, wherein the sensor array includes a plurality of portions having a serpentine shape. 6. The interface system of claim 1, wherein the sensor array comprises between 2 and 8 electrodes. 7. The interface system of claim 1, wherein the sensor array is disposable and wherein the first component and the second component are both reusable. 8. The interface system of claim 1, wherein the sensor array is formed of a mesh. 9. The interface system of claim 1, wherein the sensor array further comprises an alignm... | Convergent Engineering, Inc. | A61B005/0444 A61B005/00 A61B005/03 A61B005/024 A61B005/0448
US10537277 | B2 | 2020-01-21 | Methods, systems, and computer readable media for visualization of resection target during epilepsy surgery and for real time spatiotemporal visualization of neurophysiologic biomarkers | 2014-05-28 | The subject matter described herein relates to methods, systems, and computer readable media for visualization of a resection target during epilepsy surgery and for real time spatiotemporal visualization of neurophysiologic biomarkers. One exemplary method includes a real time neurophysiologic biomarker visualization system implemented by at least one computer, receiving, as input, a pre-electrode-implantation MRI of an epilepsy patient's brain. | 1. A method for visualization of a resection target during epilepsy surgery, the method comprising: at a real time neurophysiologic biomarker visualization system stored in a non-transitory computer readable medium and implemented by at least one computer: receiving, as input, a pre-electrode-implantation magnetic resonance image (MRI) of an epilepsy patient's brain;receiving, as input, a post-electrode-implantation computed tomography (CT) image of the patient's brain;receiving, as input, a post-electrode-implantation MRI of the patient's brain;receiving, as input, raw EEG data produced by an implanted electrode array, processing the raw EEG data, and generating, in real time, intraoperatively, and from the raw EEG data, spatiotemporal visualizations of occurrences of a neurophysiologic biomarker mapped to locations of the occurrences of the neurophysiologic biomarker, wherein the spatiotemporal visualizations of the occurrences change over time with changes in the raw EEG data and the spatiotemporal visualizations comprise a spatiotemporal color map of results of band-limited, shifting-time-windowed sampling of high frequency oscillations of EEG signals rendered on top of a 3D model of the patient's brain;receiving, as input, data identifying a resection target; andreceiving, as input, coordinates in real space produced by a neuronavigation system;registering the MRIs, the CT image, the spatiotemporal visualizations of the occurrences of the neurophysiologic biomarker, the data identifying the resection target and the coordinates in real space produced by the neuronavigation system to produce a registered image. 2. The method of claim 1 wherein processing the raw EEG data includes processing the raw EEG data to identify a particular biomarker in the EEG data. 3. The method of claim 1 wherein generating the spatiotemporal visualizations of the occurrences of the neurophysiologic biomarker includes generating a heat map illustrating intensities of the occurrences of the neurophysiologic biomarker mapped to the locations of the occurrences. 4. The method of claim 3 wherein generating the spatiotemporal visualizations includes utilizing color to illustrate relative intensities of the occurrences of the neurophysiologic biomarker. 5. The method of claim 1 wherein the registered image includes an image of the patient's brain, a visual indication of the resection target mapped to the image of the patient's brain, a visualization of occurrence of the neurophysiologic biomarker mapped to the image of the patient's brain, and an image of at least one object tracked by the neuronavigation system mapped to the image of the patient's brain. 6. The method of claim 1 comprising continually receiving the raw EEG data as input and updating the spatiotemporal visualizations of occurrences of the neurophysiologic biomarker mapped to locations of the occurrences. 7. The method of claim 1 wherein outputting the registered image includes outputting the registe... | The Trustees of the University of Pennsylvania | A61B005/00 A61B005/04 A61B005/0478 G01R033/48
US09351686 | B1 | 2016-05-31 | Tissue oximeter intraoperative sensor | 2015-06-22 | An oximeter probe includes sensor head with a probe face having one or more sensor structures to make measurements, a handle, and an elastic member connected between the handle and the base. A user can hold the handle while measurements are made and the elastic member permits the handle to flex relative to the sensor head with one or more sensor structures. | 1. An oximeter device comprising: a sensor head comprising a probe face that faces a tissue to be measured;first, second, and third sensor structures formed on the probe face, wherein a first distance is between the first and second sensor structures, and a second distance is between the first and third sensor structures, and the first distance is different from the second distance;a handle; andan elastic member comprising a spring constant, coupled between the sensor head and handle. 2. The device of claim 1 wherein the first sensor structure comprises a source structure and the second and the third sensor structures comprise detector structures. 3. The device of claim 1 wherein the first sensor structure comprises a detector structure and the second and the third sensor structures comprise source structures. 4. The device of claim 1 wherein the first and second sensor structures form a first line, the first and third sensor openings form a second line, and the first line is not parallel to the second line. 5. The device of claim 1 wherein the first, second, and third sensor structures are arranged in a line. 6. The device of claim 1 comprising: a pressure sensor, coupled to the sensor head. 7. The device of claim 1 comprising: a light emitting diode, coupled to the first sensor structure;a photodetector, coupled to the second sensor structure; anda battery, electrically coupled to the light emitting diode and the photodetector. 8. The device of claim 1 wherein a first line passes through the base portion, in a first direction that is transverse to the bottom surface of the base portion, a second line extends in the first direction and intersects a surface of the handle,the first line passes through the first sensor structure,the first line does not pass through the handle, anda first distance is between the first line and the second line in a second direction, transverse to the first direction, wherein the first distance is greater than about 1 millimeter. 9. The device of claim 1 comprising: a light emitting diode, coupled to the first sensor structure;a photodetector, coupled to the second sensor structure; anda battery, electrically coupled to the light emitting diode and the photodetector, whereina first line passes through the base portion, in a first direction that is transverse to the bottom surface of the base portion,a second line extends in the first direction and intersects a surface of the handle,the first line passes through the first sensor structure,the first line does not pass through the handle, anda first distance is between the first line and the second line in a second direction, transverse to the first direction, wherein the first distance is at least about 1 millimeter. 10. The device of claim 1 wherein the oximeter device is a tissue oximeter probe. 11. An oximeter device comprising: a probe face that faces a tissue to be measured;at least one sensor structure formed on the probe face;a pressure sensor, coupled ... | ViOptix, Inc. | A61B005/00 A61B005/1455
US10028701 | B2 | 2018-07-24 | Disposable insert having sensor and RFID | 2015-08-03 | A diagnostic sheet for use in conjunction with a disposable absorbent garment is disclosed. The diagnostic sheet may be operable to wick fluid vertically, and includes one or more diagnostic elements and a radio frequency identification device (RFID). The one or more diagnostic elements may be integrated with the diagnostic sheet and positioned at a first region of the diagnostic sheet, and the RFID may be integrated with the diagnostic sheet and positioned proximate the first region. In the presence of the fluid, each of the one or more diagnostic elements may be configured to indicate a result of a particular diagnostic test. In the presence of the fluid, the RFID may be configured to provide an indication that the fluid wicked by the diagnostic sheet has reached the one or more diagnostic elements. | 1. An article for use in conjunction with a disposable absorbent garment, the article comprising: a diagnostic sheet operable to wick fluid vertically, the diagnostic sheet including: one or more diagnostic elements integrated with the diagnostic sheet and positioned at a first region of the diagnostic sheet; anda radio frequency identification device (RFID) integrated with the diagnostic sheet and positioned proximate the first region of the diagnostic sheet, wherein the RFID includes: an antenna configured to receive and transmit signals;signal generation circuitry configured to generate a signal for transmission by the antenna based on a wireless signal received at the antenna from an interrogator device;power generation circuitry configured to provide operational power to the signal generation circuitry based on the wireless signal received at the antenna from the interrogator device; andpower control circuitry configured to disable the provisioning of power, by the power generation circuitry, to at least the signal generation circuitry until at least a portion of the power control circuitry is in the presence of the fluid,wherein, in the presence of the fluid wicked to one or more diagnostic elements by the diagnostic sheet, each of the one or more diagnostic elements is configured to indicate a result of a particular diagnostic test, andwherein the signal generated by the signal generation circuitry is configured to provide an indication to the interrogator device that the fluid has been wicked to the one or more diagnostic elements by the diagnostic sheet, the indication configured to cause the interrogator device, in response to receiving the signal, to generate a notification that indicates the fluid has been wicked to the one or more diagnostic elements by the diagnostic sheet. 2. The article of claim 1, wherein each of the one or more diagnostic elements is selected from the list consisting of: a pH test, a specific gravity test, a protein test, a blood test, a glucose test, a ketone test, a bilirubin test, a urobilinogen test, a nitrites test, and a leukocytes test. 3. The article of claim 1, wherein the article is integrated into a disposable absorbent garment as a top sheet. 4. The article of claim 1, wherein the diagnostic sheet includes a front edge, wherein first region is proximate a front edge of the diagnostic sheet, and wherein the RFID is positioned between the one or more diagnostic elements and the front edge of the diagnostic sheet. 5. A disposable absorbent garment comprising: a backsheet;an absorbent core; anda top sheet configured to wick fluid, the top sheet including:one or more diagnostic elements integrated with the top sheet and positioned at a first region of the top sheet, wherein each of the one or more diagnostic elements is configured to, in the presence of the fluid wicked to one or more diagnostic elements by the top sheet, indicate a result of a particular diagnostic test; anda radio frequency identifi... | ATTENDS HEALTHCARE PRODUCTS, INC. | G08B023/00 A61B005/00 A61F013/49 A61F013/42 A61B005/145 A61B005/20
US10321890 | B2 | 2019-06-18 | Apparatus and method for endovascular device guiding and positioning using physiological parameters | 2016-01-28 | An endovascular navigation system and method are disclosed. The endovascular navigation system includes an elongate flexible member, a endovascular electrogram lead disposed at a distal end of the elongate flexible member and configured to sense an endovascular electrogram signal, a processor and an output device. The processor is configured to receive the endovascular electrogram signal, to determine a peak amplitude of a P-wave in the endovascular electrogram signal, and to determine that the position of the distal end of the elongate flexible member is within a predetermined structure within the venous vasculature of the patient. The output device is configured to display a visual indication that the distal end of the elongate flexible member is within the predetermined structure within the venous vasculature of the patient, where the visual indication is different from the endovascular electrogram signal of the venous vasculature of the patient. | 1. An endovascular navigation system, comprising: an elongate flexible member configured to access the venous vasculature of a patient;an endovascular electrogram lead disposed at a distal end of the elongate flexible member and configured to sense an endovascular electrogram signal of the venous vasculature of the patient;a processor configured to: receive the endovascular electrogram signal of the venous vasculature of the patient provided by the endovascular electrogram lead;determine a peak amplitude of a P-wave in the endovascular electrogram signal;determine an amplitude within the QRS complex of the endovascular electrogram signal;compare the peak amplitude of the P-wave to the amplitude within the QRS complex to determine a P-wave relative amplitude; anddetermine, based at least in part on the P-wave relative amplitude, that the position of the distal end of the elongate flexible member is within a predetermined structure within the venous vasculature of the patient; andan output device configured to display, in response to the determination that the position of the distal end of the elongate flexible member is within the predetermined structure within the venous vasculature of the patient, a visual indication that the distal end of the elongate flexible member is within the predetermined structure within the venous vasculature of the patient,wherein the visual indication is different from the endovascular electrogram signal of the venous vasculature of the patient. 2. The system of claim 1, wherein the predetermined structure within the venous vasculature of the patient is the superior vena cava of the patient. 3. The system of claim 2, wherein the predetermined structure within the venous vasculature of the patient is the lower one-third of the superior vena cava of the patient. 4. The system of claim 1, wherein the processor is configured to determine the peak amplitude of the P-wave in the endovascular electrogram signal by: determining a first peak amplitude of a first P-wave in a first heartbeat of the endovascular electrogram signal; anddetermining a second peak amplitude of a second P-wave in a second heartbeat of the endovascular electrogram signal. 5. The system of claim 4, wherein the processor is configured to determine that the position of the distal end of the elongate flexible member is within the predetermined structure within the venous vasculature of the patient by determining that the second peak amplitude of the second P-wave is greater than the first peak amplitude of the first P-wave. 6. The system of claim 4, wherein the processor is configured to: determine the amplitude within the QRS complex of the endovascular electrogram signal by: determine a first amplitude within a first QRS complex in the first heartbeat of the endovascular electrogram signal; anddetermine a second amplitude within a second QRS complex in the second heartbeat of the endovascular electrogram signal; andcompare the peak amplitude of the ... | ARROW INTERNATIONAL, INC. | A61B008/00 A61B005/026 A61B005/042 A61B005/0452 A61B005/0472 A61B005/0476 A61B005/06 A61B008/06 A61B008/12 G09B023/28 A61B005/0205 A61B007/04 A61B005/00 A61B008/02 A61B005/0456 A61B005/0488 A61B034/20 A61B034/00 A61B090/00 A61B005/0215 A61B005/029 A61B005/1459 A61B008/08 A61B017/00 A61B017/22
US09326743 | B2 | 2016-05-03 | Multi-energy CT imaging system and imaging method | 2013-12-19 | The present invention relates to a Computed Tomography (CT) imaging system, in particular to a multi-energy CT imaging system and imaging method. | 1. A multi-energy CT imaging system comprising: a stage for carrying an object to be inspected;a voltage-regulatable X-ray generation device for emitting X-rays for radiating said object to be inspected;a detector for receiving X-rays emitted from said X-ray generation device and penetrating said object to be inspected, and for outputting detection data;a rack having said X-ray generation device and said detector mounted thereon; anda data processing and control device for controlling said stage, said X-ray generation device, said detector and said rack, processing detected data from said detector, and during one rotation of scan of said X-ray generation device, evenly dividing a circular track of said X-ray generation device into angle intervals of a predetermined number according to a predetermined number of energies and setting a different high voltage of said X-ray generation device for each angle interval, and when said rack rotates from the current angle interval into a next angle interval, controlling said X-ray generation device according to an instruction from said data processing and control device to switch said X-ray generation device to a voltage set in the next angle interval. 2. A method for imaging with a multi-energy CT imaging system, comprising: (a) providing a multi-energy CT imaging system of claim 1;(b) evenly dividing the circular track of the X-ray generation device of said multi-energy CT imaging system into angle intervals of a predetermined number according to the predetermined number of energy and setting a different high voltage of said X-ray generation device for each angle interval;(c) rotating said X-ray generation device along the circular track and working in each angle interval at a high voltage set for said angle interval to emit X-rays to irradiate the object to be inspected;(d) acquiring projection data of different energies using the detector of said multi-energy CT imaging system;(e) obtaining line integral data of the projection data of said different energies using the data processing and control device of said multi-energy CT imaging system;(f) splicing said line integral data of different energies one by one to obtain a complete line integral data covering 360°;(g) reconstructing according to said complete line integral data to obtain a hybrid reconstruction result;(h) using said hybrid reconstruction result as a priori image to obtain attenuation coefficient reconstruction results in different energies;(i) obtaining decomposition coefficients from said attenuation coefficient reconstruction results and obtaining an atomic number distribution image and electron density distribution image using said decomposition coefficient. 3. The method of claim 2, wherein the circular track of step (b) is divided into N angle intervals and the line integral data in different energies obtained in step (e) are made to be P1, P2, . . . Pn, and in step (g), reconstruction is performed using a common circular track filt... | Nuctech Company Limited Tsinghua University | A61B006/00 G01N023/04 G01V005/00 A61B006/02 A61B005/00
US09333368 | B2 | 2016-05-10 | Treatment of biological tissues using subnanosecond electric pulses | 2014-02-03 | A system for treatment of biological tissues is provided. The system includes a lens having a hollow, substantially hemispherical shape with an outer surface and an inner surface, the inner surface defining a substantially hemispherical cavity for inserting the biological tissues. The system further includes an antenna assembly for generating and directing electromagnetic radiation towards the outer surface. In the system, the lens is configured to direct the electromagnetic energy to an area in the cavity, a dielectric constant of the lens at the inner surface substantially matches a dielectric constant of the biological tissues, the dielectric constant monotonically increases from the outer surface to the inner surface, and the electromagnetic energy is generated via a series of pulses having a transient of less than about 1 nanosecond. | 1. A system for treatment of biological tissues, comprising: a lens having a hollow, substantially hemispherical shape with an outer surface, an inner surface, one or more lossy portions, and a dielectric constant increasing monotonically from the outer surface to the inner surface, the inner surface defining a substantially hemispherical cavity for inserting the biological tissues; andan antenna assembly for generating electromagnetic energy via a series of pulses having a transient time of less than about 1 nanosecond and directing the electromagnetic radiation towards the outer surface;wherein the lens is configured to direct the electromagnetic radiation to an area in the cavity, wherein the dielectric constant of the lens at the inner surface substantially matches a dielectric constant of the biological tissues, and wherein the lossy portions of the lens are configured to attenuate a portion of the electromagnetic radiation being directed to the cavity from at least one direction. 2. The system of claim 1, wherein the lens comprises a plurality of layers, and wherein at least a portion of the plurality of layers comprise lossy dielectric materials in an azimuthal direction of the electromagnetic radiation defining the lossy portions of the lens. 3. The system of claim 1, wherein the lens comprises one or more lossy elements along radial directions defining the lossy portions of the lens, the lossy elements configured to attenuate the electromagnetic radiation in at least one pre-defined radial path to control the angle of the power incidence. 4. The system of claim 1, wherein the dielectric constant of the lens increases exponentially from the outer surface to the inner surface. 5. The system of claim 1, wherein the antenna assembly comprises: an antenna for receiving the series of pulses and generating the electromagnetic radiation; anda reflector for directing the electromagnetic radiation to the lens. 6. The system of claim 5, wherein the antenna comprises one of a discone antenna, resistively-loaded dipole antenna, or a resistance terminated transmission line antenna. 7. The system of claim 5, wherein the reflector comprises a prolate spheroid reflector with the antenna at a first focal point of the prolate spheroid reflector, and wherein the lens and the antenna assembly are arranged such that a second focal point of the reflector is in the cavity. 8. The system of claim 1, further comprising a pulse generator coupled to the antenna assembly and configured for generating the series of pulses. 9. The system of claim 8, wherein the pulse generator is configured for generating each of the series of pulses to have a pulse with transient time between about 1 ps and 1 ns. 10. The system of claim 8, wherein the pulse generator is configured for generating each of the series of pulses to have a magnitude between about 1 kV and 1 MV. 11. The system of claim 8, wherein the pulse generator is configured for generating each of the series ... | OLD DOMINION UNIVERSITY RESEARCH FOUNDATION | A61N001/00 A61B005/04 A61N002/00 A61N005/02 A61N001/40
US20160143586 | A1 | 2016-05-26 | A STRAP FOR A PORTABLE PULSE MEASURING DEVICE AND A PORTABLE PULSE MEASURING DEVICE | 2014-06-11 | A strap for a pulse measuring device, wherein the strap includes an indicator configured to indicate tightness of the strap. A portable pulse measuring device having a strap configured to fasten the portable pulse measuring device on a human; and a mechanical indicator configured to indicate tightness of the strap are also disclosed. | 1. A strap for a portable pulse measuring device, wherein the strap comprises: an indicator configured to indicate tightness of the strap, wherein the strap is a stretchable strap comprising a non-stretchable sliding part built-in in the stretchable strap and wherein the stretchable strap comprises a section through which the indicator on the non-stretchable sliding part is visible. 2. The strap for a portable pulse measuring device, wherein the strap comprises: an indicator configured to indicate tightness of the strap, wherein the strap is a stretchable strap comprising a stretchable sliding part built-in in the stretchable strap, the stretchable sliding part being less stretchable than the stretchable strap and wherein the stretchable strap comprises a section through which the indicator on the stretchable sliding part is visible. 3. The strap for a portable pulse measuring device, wherein the strap comprises: an indicator configured to indicate tightness of the strap, wherein the indicator is provided by knitting the strap such that when the strap is stretched, the knitting enables a meter to be seen on the strap, wherein the meter alone indicates the tightness of the strap. 4. A portable pulse measuring device comprising a strap as claimed in claim 1, to fasten the portable pulse measuring device on a human. 5. A portable pulse measuring device comprising: a main body comprising a link part;a strap configured to fasten the portable pulse measuring device on a human; andwherein the strap is connected to the link part and wherein the link part comprises a mechanical indicator configured to indicate tightness of the strap. 6. The portable pulse measuring device as claimed in claim 5, wherein the link part comprises a window in which the mechanical indicator is movable to indicate the tightness of the strap. 7. The portable pulse measuring device as claimed in claim 6, wherein the link part comprises a rotating part to which the strap is attached, and wherein the indicator is attached to the rotating part such that when the rotating part rotates, the indicator moves in the window. 8. The portable pulse measuring device as claimed in claim 6, wherein the link part comprises a sliding part to which the strap is attached, and wherein the indicator is attached to the sliding part such that when the sliding part moves, the indicator moves in the window. 9. The portable pulse measuring device comprising a strap as claimed in claim 2, to fasten the portable pulse measuring device on a human. 10. The portable pulse measuring device comprising a strap as claimed in claim 3, to fasten the portable pulse measuring device on a human. |  | A61B005/00 A61B005/024
US09333349 | B2 | 2016-05-10 | Bio-feedback type device for physical pain relief at childbirth | 2013-11-13 | Disclosed is a bio-feedback type device for physical pain relief at childbirth, comprising a maternal-fetal monitoring module, a hand-grip pressure transducer, an electrical-stimulation wave group generator, an electrical-stimulation wave group output module and a childbirth pain-relief workstation. The maternal-fetal monitoring module is used for monitoring maternal-fetal vital-sign signals; the hand-grip pressure transducer is used to convert the degree of pain experienced by the mother into a pressure value; the childbirth pain-relief workstation combines and analyzes uterine contraction stress values from the maternal-fetal vital-sign signals with the hand-grip pressure values to produce electrical-stimulation wave group output parameters; on the basis of said output parameters, the electrical-stimulation wave group generator regulates the electrical-stimulation wave group and outputs same to the maternal body surface by means of the electrical-stimulation wave group output module and electrodes; the present invention is also provided with input and display modules to enable interaction between operators and equipment. The present invention is able to provide timely adjustment of electrical-stimulation wave group output parameters, the electrical-stimulation wave group signals taking into account the felt physical reception of different frequencies and waveforms of electrical stimulation pulses. | 1. A bio-feedback type device for physical pain relief at childbirth, comprising a childbirth pain-relief station and an electrical-stimulation wave group output module, wherein the apparatus further comprising a maternal-fetal monitoring module, a hand-grip pressure transducer, an electrical-stimulation wave group generator and an electrical-stimulation wave group output module; the maternal-fetal monitoring module, the hand-grip pressure transducer and the electrical-stimulation wave group generator are coupled respectively to the childbirth pain-relief station, and the electrical-stimulation wave group generator is coupled to the electrical-stimulation wave group output module; the maternal-fetal monitoring module is configured to detect a maternal-fetal physical signal and transmit the detected maternal-fetal physical signal to the childbirth pain-relief station;the hand-grip pressure transducer is configured to convert a degree of a pain born by a parturient to a hand-grip pressure value and transmit the converted hand-grip pressure value to the childbirth pain-relief station;the childbirth pain-relief station is configured to regulate an output parameter of electrical-stimulation wave group by analyzing a pressure value of uterine contraction in the maternal-fetal physical signal and the hand-grip pressure value from the hand-grip pressure transducer, and calculating the output parameter of electrical-stimulation wave group according to the analyzing; the output parameter of electrical-stimulation wave group including type, frequency and amplitude of the electrical-stimulation wave group;the electrical-stimulation wave group generator is configured to generate a corresponding electrical-stimulation wave group according to the output parameter of electrical-stimulation wave group regulated by the childbirth pain-relief station, and the electrical-stimulation wave group output module is configured to output the electrical-stimulation wave group to the parturient's body. 2. The bio-feedback type device for physical pain relief at childbirth of claim 1, wherein regulating the output parameter of electrical-stimulation wave group by the childbirth pain-relief station comprises: step 1, detecting a maximum hand-grip pressure value of the parturient, Fmax, by the hand-grip pressure transducer;step 2, detecting a real time pressure value of uterine contraction, TOCO, by the maternal-fetal monitoring module;step 3, detecting a real time hand-grip pressure value of the parturient, F, by the hand-grip pressure transducer; andstep 4, originally calculating a type XTOCO and an amplitude ATOCO of the electrical-stimulation wave group according to the TOCO detected in the step 2, andwhen TOCO>TOCOthreshold, the type XTOCO of the electrical-stimulation wave group is set to an analgesic wave group, and when TOCO&lt;TOCOthreshold, the type XTOCO of the electrical-stimulation wave group is a massage wave group, wherein, the TOCOthreshold is a pain threshold ... | Guangzhou Sunray Medical Apparatus Co., Ltd. | A61N001/00 A61N001/36 A61B005/00
US09355297 | B2 | 2016-05-31 | Biometric information input apparatus and biometric information input method | 2013-11-14 | A biometric information input apparatus includes an image capturing section configured to obtain a captured image of a biological object; a pliable part detecting section configured to detect whether there is a pliable part on a surface of the biological object by obtaining a distance to the surface of the biological object from the captured image and comparing the obtained distance with a predetermined distance to be compared set beforehand; and an extracting section configured to extract biometric information from the captured image if the pliable part is not detected by the pliable part detecting section. | 1. A biometric information input apparatus comprising: an image capturing camera device configured to obtain a captured image of a biological object;a memory device configured to store a program; anda central processing unit configured to execute the program to perform:obtaining distances between a position of an image capturing camera device and a plurality of points on a surface of a biological object from an image captured by the image capturing camera device to calculate an average value of the obtained distances;comparing each of the obtained distances with the calculated average value of the obtained distances to detect whether there is a pliable part on the surface of the biological object; andextracting biometric information from the captured image if the pliable part is not detected. 2. The biometric information input apparatus as claimed in claim 1, wherein the central processing unit is configured to execute the program to calculate the average value of the obtained distances by smoothing the obtained distances, and detects the pliable part based on an amount of shift between the calculated average value and each of the obtained distances. 3. The biometric information input apparatus as claimed in claim 1, wherein the central processing unit is configured to execute the program to display a message indicating that the pliable part is detected if an area of the pliable part on the surface of the biological object is greater than a predetermined threshold value. 4. The biometric information input apparatus as claimed in claim 3, wherein the displayed message includes positional information of the detected pliable part. 5. The biometric information input apparatus as claimed in claim 1, wherein the central processing unit is configured to execute the program to determine whether to detect whether there is a pliable part, based on a posture of the surface of the biological object obtained from the captured image obtained by the image capturing camera device. 6. The biometric information input apparatus as claimed in claim 1, wherein the central processing unit is configured to execute the program to set a candidate pliable part from a form of the surface of the biological object, and detect whether there is a pliable part on the set candidate pliable part. 7. The biometric information input apparatus as claimed in claim 1, wherein the central processing unit is configured to execute the program to execute authentication using the extracted biometric information. 8. A non-transitory computer-readable recording medium having a program stored therein for causing a computer to execute a biometric information input method, the method comprising: capturing a captured image of a biological object with an image capturing camera device;obtaining distances between a position of the image capturing camera device and a plurality of points on a surface of the biological object from the captured image to calculate an average value of the obtained... | FUJITSU LIMITED | G06K009/00 A61B005/117 A61B005/00
US10362973 | B2 | 2019-07-30 | Systems and methods for monitoring and managing life of a battery in an analyte sensor system worn by a user | 2015-12-18 | Systems and methods for analyte monitoring, particularly systems and methods for monitoring and managing life of a battery in an analyte sensor system worn by a user, are provided. | 1. A method for prolonging life of a battery installed in an analyte sensor system configured to be worn by a user comprising: measuring a first analyte value at a first time;causing a transmission of the measured first analyte value along with a predicted second analyte value;measuring a second analyte value at a second time;calculating a difference between the measured second analyte value and the predicted second analyte value;determining whether the calculated difference between the measured second analyte value and the predicted second analyte value is within a predefined range;determining that a remaining power level of the battery is less than a predetermined power level; andresponsive to the determining that the remaining power level of the battery is less than the predetermined power level and determining that the calculated difference is within the predefined range,skipping a transmission of the measured second analyte value,wherein the predefined range is a variable predefined range that is determined based at least partly on a clinical error grid. 2. The method of claim 1, wherein the analyte sensor system is a continuous glucose sensor system. 3. The method of claim 1, wherein at least one additional predicted future analyte value is transmitted along with the measured first analyte value and the predicted second analyte value, and skipping transmission of at least one additional measured future analyte value if a difference between the at least one additional measured future analyte value and the at least one additional predicted future analyte value is within the predetermined range. 4. An analyte sensor system configured to be worn by a user, comprising: an analyte sensor;a sensor measurement module coupled to the analyte sensor and configured to provide measured analyte values;a transceiver configured to transmit and receive wireless signals to and from a display device configured to display measured analyte values; a battery; anda control module coupled to the sensor measurement module and the transceiver and configured to: receive a measured first analyte value at a first time,cause a transmission of the measured first analyte value to the display device along with a predicted second analyte value,receive a measured second analyte value from the sensor measurement module at a second time,calculate a difference between the measured second analyte value and the predicted second analyte value,determine whether the calculated difference between the measured second analyte value and the predicted second analyte value is within a predefined range; anddetermine that a remaining power level of the battery is less than a predetermined power level, andresponsive to the determining that the remaining power level of the battery is less than the predetermined power level and determining that the calculated difference is within the predefined range, skip a transmission of the measured second analyte value,wherein the predefined range is a... | DexCom, Inc. | A61B005/00 A61B005/145 H04W052/02
US09331265 | B2 | 2016-05-03 | Non-volatile memory with linear hot-electron injection technique and strain gauge using the same | 2013-01-28 | A linear hot-electron injection technique is provided for a non-volatile memory arrangement. The non-volatile memory is comprised of: a floating gate transistor; a capacitor with a first terminal electrically coupled to the gate node of the floating gate transistor; a current reference circuit electrically coupled to the source node of the floating gate transistor; and a feedback circuit electrically coupled between the source node of the floating gate transistor and a second terminal of the capacitor. The feedback circuit operates to adjust a voltage at the gate node of the floating gate transistor in accordance with a source-to-drain voltage across the floating gate transistor. | 1. A self-powered static-strain sensor, comprising: a piezoelectric member that generates a voltage in response to as mechanical strain; anda non-volatile memory circuit powered by the voltage received from the piezoelectric member and includes a floating gate transistor having a gate node and a source node;a capacitor having a first terminal and a second terminal, the first terminal electrically coupled to the gate node of the floating gate transistor;a current reference circuit electrically coupled to the source node of the floating gate transistor and operable to source a current therein; anda feedback circuit electrically coupled between the source node of the floating gate transistor and the second terminal of the capacitor, and operable to adjust a voltage at the gate node of the floating gate transistor in accordance with a source-to-drain voltage across the floating gate transistor. 2. The self-powered static-strain sensor of claim 1 wherein the current reference circuit operates to hold the source current constant, thereby ensuring that the source-to-gate voltage of the floating gate transistor remains constant. 3. The self-powered static-strain sensor of claim 1 wherein the current reference circuit is implemented by two transistors in a cascading arrangement. 4. The self-powered static-strain sensor of claim 1 wherein the feedback circuit is further defined as an operational amplifier having an inverting terminal electrically coupled to the source node of the floating gate transistor, a non-inverting terminal electrically coupled to a reference voltage and an output electrically coupled to the second terminal of the capacitor. 5. The self-powered static-strain sensor of claim 1 wherein the feedback circuit is further defined as a single-stage, single ended amplifier. 6. The self-powered static-strain sensor of claim 1 wherein the non-volatile memory circuit further comprises a tunnel capacitor electrically coupled to the gate node of the floating gate transistor for removing electrons from the floating gate transistor. 7. The self-powered static-strain sensor of claim 1 wherein the non-volatile memory circuit further comprises: a switch electrically coupled between the second terminal of the capacitor and ground; anda controller interfaced with the switch to control operation thereof, such that the switch is opened to inject current into the floating gate transistor and is closed to remove electrons from the floating gate transistor. 8. The self-powered static-strain sensor of claim 1 further comprises a rectifier interposed between the piezoelectric member and the non-volatile memory circuit. 9. A self-powered static-strain sensor, comprising: a piezoelectric member that generates a voltage in response to as mechanical strain;a non-volatile memory arrangement having an array of memory circuits powered by voltage received from the piezoelectric member, each memory circuit includes a floating gate transistor having a gate node a... | Board of Trustees of Michigan State University | G01L001/16 H01L041/113 G11C016/34 G11C027/00 G01B007/16 A61B005/00 H01L027/20 H01L029/788
US20160132656 | A1 | 2016-05-12 | PROCEDURE FOR THE MODELLING OF THE DIRECTION OF INCIDENT ABDOMINAL PRESSURES TOWARDS THE FEMALE PELVIC CAVITY AND THE DIRECTION OF THE REFLECTED PRESSURE IN THE PELVIC SPACE | 2015-10-27 | A method for modelling the direction of incident abdominal pressures through a strait and directed towards the female pelvic cavity and for modelling the direction of pressures reflected by the pelvic paraboloid of an individual so as to allow correlating with occurrence of prolapses and incontinence in the said individual. The method includes compiling morphological data of a plane of a pelvic cavity strait by 3D cube MRI; modelling a center of gravity CG1 of the strait at a level at which the incident pressure vector penetrates in the true pelvis; modelling by 3D cube MRI a plane of a pelvic paraboloid, as well as an axis of the paraboloid; determining a CG2 reflection point of the incident vector on the plane; and determining the orientation of a reflected pressure vector. | 1-2. (canceled) 3. A method for modelling the direction of incident abdominal pressures towards a female pelvic cavity and a direction of reflected pressures in a pelvic space of an individual for use in correlating with an occurrence of prolapses and incontinence in the individual, said method comprising the following steps: compiling morphological data of a plane of a pelvic cavity strait by 3D cube MRI;modelling a center of gravity CG1 of the strait at a level at which an incident pressure vector penetrates in the true pelvis,modelling by 3D cube MRI a plane of a pelvic paraboloid and an axis of the paraboloid;determining a CG2 reflection point of an incident vector on the plane; anddetermining the orientation of a reflected pressure vector. 4. The method of claim 3, comprising the additional step, prior to said step of determining the orientation of a reflected pressure vector, of refining the determination of the center of gravity CG1 of the upper strait using variations in a density of intestinal content to define a weighted center of gravity (CG1′). |  | G06F019/00 A61B005/055 A61B005/20
US20160135738 | A1 | 2016-05-19 | MONITORING TREATMENT COMPLIANCE USING PASSIVELY CAPTURED TASK PERFORMANCE PATTERNS | 2015-11-12 | Methods and systems are described for unobtrusively monitoring patient activity to determine compliance of the patient with a prescribed regimen for treating for a brain-related disorder. Patient activity associated with routine patient activities is detected with an activity sensor at the patient location, and activity data is transmitted to a monitoring location. Patient activity is processed at the patient location and/or monitoring location to identify activity parameters and/or patterns that indicate whether the patient has complied with the prescribed treatment regimen. Patient identity may be determined through biometric identification or other authentication techniques. The system may provide a report to an interested party, for example a medical care provider or insurance company, regarding compliance of the patient with the prescribed treatment regimen. | 1. A system comprising: at least one receiving device for use at a monitoring location for receiving an activity data signal transmitted to the monitoring location from a patient location, the activity data signal containing activity data representing at least one non-speech activity pattern in activity sensed from a patient with at least one activity sensor in an unobtrusive activity-detection system at the patient location during performance of the non-speech activity by the patient, the patient having a brain-related disorder and a prescribed treatment regimen for treating at least one aspect of the brain-related disorder;signal processing circuitry configured to analyze the activity data signal to determine whether the activity data represents at least one non-speech activity pattern that matches at least one characteristic activity pattern;compliance determination circuitry configured to determine whether the patient has complied with the prescribed treatment regimen based upon whether the activity data represents the at least one non-speech activity pattern that matches the at least one characteristic activity pattern; andreporting circuitry configured to report a conclusion based on the determination of whether the patient has complied with the prescribed treatment regimen. 2. The system of claim 1, wherein the signal processing circuitry is configured to analyze the activity data signal to identify at least one non-speech activity pattern that corresponds to unprompted performance of the non-speech activity by the patient. 3.-34. (canceled) 35. The system of claim 1, including patient identification circuitry configured to determine a presence of the patient from at least one identity signal received by a receiving device at the monitoring location from the patient location; wherein the signal processing circuitry is configured to identify patient activity data corresponding to an activity of the patient based at least in part on the determination of the presence of the patient by the patient identification circuitry. 36.-64. (canceled) 65. The system of claim 1, wherein the at least one receiving device is adapted to receive a physiological activity data signal indicative of at least one physiological signal sensed with at least one physiological sensor operatively connected to the unobtrusive activity-detection system. 66.-71. (canceled) 72. A method of monitoring compliance of a patient with a prescribed treatment regimen, comprising: receiving an activity data signal with a receiving device at a monitoring location, the activity data signal transmitted to the monitoring location from a patient location, the activity data signal containing activity data representing at least one non-speech activity pattern in activity sensed from a patient with at least one activity sensor in an unobtrusive activity-detection system at the patient location during performance of the non-speech activity by the patient, the patient having a brain-r... |  | A61B005/00 A61B005/117 A61B003/11 A61B003/113 A61B005/0205 A61B005/0476
US20160135755 | A1 | 2016-05-19 | Remote home healthcare system | 2013-08-19 | A remote home health care system is described. The system includes: a fusion sorting subsystem configured to receive physical sign data parameters collected by a sensor in real time, perform fusion sorting processing on the physical sign data parameters, pre-diagnose a physical condition of a user in real time according to physiological data and a physiological model in a physiological model library and feed back the pre-diagnosed physical condition; a resource optimization subsystem configured to optimize the physiological data in a physiological database periodically, generate a personalized physiological model for the user according to historical physiological data in the physiological database, store the generated physiological model in the physiological model library, and update the physiological model in the physiological model library according to the latest physiological data in the physiological database; and a comprehensive evaluation subsystem configured to predict a changing trend and a dynamic change range of the physical signs of the user according to the physiological data in the physiological database and the physiological model in the physiological model library and evaluate a health condition of the user according to the physiological data and the result of the prediction. | 1. A remote home health care system, comprising: a fusion sorting subsystem configured to receive physical sign data parameters collected by a sensor in real time, perform fusion sorting processing on the physical sign data parameters and pre-diagnose a physical condition of a user in real time according to the physical sign data parameters and a physiological model in a physiological model library while discovering and filtering error data in the physical sign data parameters, and store data resulting from the fusion sorting processing in a physiological database as physiological data;a resource optimization subsystem configured to recover and optimize the physiological data in the physiological database, generate a personalized physiological model for the user according to historical physiological data in the physiological database, store the generated physiological model in the physiological model library, and update the physiological model in the physiological model library according to the latest physiological data in the physiological database;a comprehensive evaluation subsystem configured to predict a changing trend and a dynamic change range of the physical signs of the user according to the physiological data in the physiological database and the physiological model in the physiological model library, and evaluate a health condition of the user according to the physiological data and the changing trend and the dynamic change range of the physical signs;a physiological database configured to store the physiological data of the user; anda physiological model library configured to store the physiological model of the user. 2. The system according to claim 1, wherein the physiological data in the physiological database includes: physical sign data, electronic medical record and health record. 3. The system according to claim 2, wherein the fusion sorting subsystem is further configured to delete the error data in the physical sign data parameters through the fusion sorting processing before storing the physical sign data parameters in the physiological database. 4. The system according to claim 2, wherein the fusion sorting subsystem comprises: a motion state detection module configured to detect whether the user falls or is in motion according to the physiological data collected by the sensor in real time, give a fall alarm or an abnormal body position alarm and send the fall alarm or the abnormal body position alarm to an alarming module after detecting that the user falls, or send motion information to a health detection module after detecting that the user is in motion;a health detection module configured to perform data fusion correlation processing and historical data correlation processing according to the acquired physiological data and motion information, diagnose a disease and discover a physiological data error according to corresponding physiological data and a corresponding physiological model, output a corresponding disease... |  | A61B005/00 A61B005/0205
US20160135715 | A1 | 2016-05-19 | METHOD FOR RESPIRATORY MEASUREMENT | 2014-07-02 | The invention is directed method for measuring respiration using impedance pneumography over a duration of at least several minutes, several hours or over the duration of night sleep, the method comprising: using at least one electrode (11) configured to be in contact with an arm (2) of a human body (1) and at least one electrode (22) configured to be in skin contact with the thorax of a human body (1); defining impedance signal changes which relate to the respiratory volume changes or time-differentiated impedance signal changes which relate to the respiratory flow; and analysing variation over time in flow-time, volume-time, flow-volume curves, or derived numerical indices, or plain respiratory impedance signal or its time derivate over the duration of at least several minutes, several hours or over the duration of night sleep. | 1. A method for measuring respiration using impedance pneumography over a duration of at least several minutes, several hours or over the duration of night sleep, the method comprising: using at least one electrode configured to be in contact with an arm of a human body and at least one electrode configured to be in skin contact with the thorax of the human body;defining impedance signal changes which relate to the respiratory volume changes or time-differentiated impedance signal changes which relate to the respiratory flow; andanalyzing variation over time in flow-time, volume-time, flow-volume curves, or derived numerical indices, or plain respiratory impedance signal or its time derivate over the duration of at least several minutes, several hours or over the duration of night sleep. 2. The method of claim 1, further comprising suppressing the cardiogenic part of the measured impedance signal. 3. The method of claim 1, further comprising deriving flow-time, volume-time, or flow-volume curves from individual or averaged multiple breaths from the impedance signal or the time-differentiated impedance signal. 4. The method of claim 1, wherein the numerical indices describe the flow-time, volume-time, or flow-volume curves. 5. The method of claim 4, wherein the indices relate to the shape of the curves, timing of peak values or other points of interest, or ratio of values or timings of two points of interest in the said curves. 6. The method of claim 1, further comprising analyzing the curves, indices, or signals regarding their variation over time using methods such as nonlinear dynamics, entropy, detrended fluctuation analysis, Lyapunov exponent, correlation dimension, recurrence plot, noise limit, frequency spectrum or using descriptive statistics such as mean, variance, or distribution analysis. 7. The method of claim 1, comprising using two electrodes configured to be in contact with one arm of the human body and two electrodes configured to be in skin contact with the thorax of the human body on the side opposite to the arm. 8. The method of claim 1, comprising using two electrodes configured to be in contact with opposite arms of the human body and two electrodes configured to be in skin contact with the thorax on opposite sides of the human body. 9. The method of claim 1, comprising preventing the skin contact between the arm and the torso by an insulation material configured to be positioned between the arm and the torso. 10. The method of claim 9, comprising the insulation material being configured to be a sleeve preventing the skin contact. 11. The method of claim 9, comprising the insulation material being configured to be a shirt or a vest preventing the skin contact. 12. An apparatus for measuring respiration using impedance pneumography over a duration of at least several minutes, several hours or over the duration of night sleep, the apparatus comprising: using at least one electrode configured to be in contact with an a... |  | A61B005/08 A61B005/087 A61B005/091 A61B005/00 A61B005/053
US10296965 | B2 | 2019-05-21 | Device configuration | 2014-11-07 | A system and method for configuring a device include determining a plurality of configuration settings for the device. The configuration settings are saved in a configuration file, and a sales order for the device is generated. The configuration file is assigned to the sales order. Based on the sales order, the configuration file is retrieved and applied to the configuration settings to the device. | 1. A method for configuring a vital signs monitor device, comprising: determining a plurality of configuration settings for the vital signs monitor device, the plurality of configuration settings including: location settings defining a clinical location at which the vital signs monitor device will be used; andone or more clinical workflows for the vital signs monitor device associated with the clinical location, wherein the one or more clinical workflows provide different functionalities of the vital signs monitor device;saving the configuration settings in a configuration file;generating a sales order for the vital signs monitor device;assigning the configuration file to the sales order; andapplying the configuration settings to the vital signs monitor device. 2. The method of claim 1, further comprising retrieving the configuration file based on a sales order number of the sales order. 3. The method of claim 1, further comprising assigning the configuration file to a line item of the sales order. 4. The method of claim 1, further comprising saving the configuration settings in a plurality of configuration files. 5. The method of claim 1, further comprising: saving the configuration settings in a plurality of configuration files; andassigning the plurality of configuration files to a plurality of line items of the sales order. 6. The method of claim 1, wherein determining the plurality of configuration settings includes receiving responses to a plurality of predetermined questions. 7. The method of claim 6, wherein receiving responses to a plurality of predetermined questions includes receiving responses to a first number of predetermined questions, wherein a second number of configuration settings are determined in response to the first number of responses, and wherein the second number is greater than the first number. 8. The method of claim 1, wherein the plurality of configuration settings are determined before generating a sales order. 9. The method of claim 1, wherein the plurality of configuration settings are determined after generating the sales order. 10. The method of claim 1, further comprising modifying the configuration settings after the configuration settings have been saved in the configuration file. 11. The method of claim 1, further comprising flagging the sales order in response to assigning the configuration file to the sales order. 12. A computer storage medium including instructions that when executed by a computing system perform a method for configuring a device, comprising: determining a plurality of configuration settings for a vital signs monitor device, the configuration settings including: clinical settings identifying a location for the vital signs monitor device, the location being one of hospital, office, or long term care; andone or more clinical workflows for the vital signs monitor device associated with the selected location, wherein the one or more clinical workflows provide different functionalit... | WELCH ALLYN, INC. | G06Q030/06 G06F009/445 A61B005/00 A61B005/01 A61B005/021 A61B005/08 A61B005/145
US09326723 | B2 | 2016-05-03 | Method and apparatus of monitoring foot inflammation | 2015-03-19 | A method and apparatus for evaluating foot inflammation each uses at least one temperature detection modality to form a first thermogram and a second thermogram of the sole of at least one foot. Each thermogram forms a substantially continuous set of two-dimensional temperature values across the sole of the (at least one) foot. The thermograms have features; namely, the first thermogram has first features and the second thermogram has second features. The method and apparatus thus control a device to apply at least one transformation to the first and second thermograms to align the first features with corresponding second features, and determine, at any thermogram location, if at least one of the thermograms presents one of a plurality of patterns indicative of inflammation. Finally, the method and apparatus each produce output information indicating the result of the determination of whether the thermograms present one of the plurality of patterns. | 1. A method of evaluating foot inflammation of at least one foot of a patient, each foot having a sole, the method comprising: providing one or more processors;generating, using at least one of the one or more processors and temperature data produced by at least one temperature detection modality, a first thermogram and a separate second thermogram of a sole of each of the at least one foot, each thermogram forming a spatially continuous data set of two-dimensional temperature values across a sole of each of the at least one foot, the first thermogram having first features, the second thermogram having second features;controlling at least one of the one or more processors to apply at least one transformation to one or both of the first and second thermograms to align the first features of the first thermogram with corresponding second features of the second thermogram;determining, by at least one of the one or more processors, at any thermogram location, if at least one of the thermograms presents one of a plurality of patterns indicative of foot inflammation; andproducing, by at least one of the one or more processors, output information indicating a result of the determination of whether the thermograms present one of the plurality of patterns. 2. The method as defined by claim 1 wherein the at least one transformation comprises a non-affine transformation. 3. The method as defined by claim 1 wherein the at least one transformation comprises an affine transformation. 4. The method as defined by claim 3 wherein the at least one affine transformation comprises at least one of reflection, rotation, scaling and translation. 5. The method as defined by claim 3 wherein the at least one transformation further comprises a non-affine transformation to at least one of the first and second thermograms. 6. The method as defined by claim 1 wherein the first thermogram represents a sole a left foot of a given person and the second thermogram represents a sole of a right foot of the given person. 7. The method as defined by claim 6 further comprising: using the modality to obtain temperature data across the sole of the left foot at a first time;using the modality to obtain temperature data across the sole of the right foot at a second time,the first time and second time being different times. 8. The method as defined by claim 1 wherein both the first thermogram and the second thermogram represent a sole of the same foot of a given person, the temperature data used to form the first and second thermograms being obtained at different times. 9. The method as defined by claim 1 wherein the at least one temperature detection modality includes a thermal camera. 10. The method as defined by claim 9 wherein using said generating comprises: a person holding the thermal camera in an unconstrained manner in at least three degrees of freedom in free space when the thermal camera obtains temperature data of a sole of each of the at least one foot, the thermal cam... | Podimetrics, Inc. | A61B005/01 A61B005/00
US09339230 | B2 | 2016-05-17 | Valve aperture sizing devices and methods | 2013-03-26 | Valve aperture sizing devices and methods. In an exemplary method to size a valve aperture, the method comprises introducing at least part of a first device into a luminal organ at a valve aperture, the first device having a balloon positioned thereon, inflating the balloon within the valve aperture until a point of apposition is achieved, and obtaining a first valve aperture measurement based upon the point of apposition. | 1. A method to size a valve annulus, comprising: introducing at least part of a first device into a luminal organ at a valve annulus, the first device having a balloon positioned thereon;inflating the balloon within the valve annulus until a point of apposition is achieved; andobtaining a first valve annulus measurement based upon the point of apposition. 2. The method of claim 1, further comprising: selecting a stent valve based upon the first valve annulus measurement. 3. The method of claim 2, further comprising: positioning the stent valve upon the balloon;re-introducing the at least part of the first device into the luminal organ; andre-inflating the balloon to place the stent valve within the luminal organ. 4. The method of claim 3, wherein the step of re-introducing is performed to re-introduce the at least part of the first device into the luminal organ at the valve annulus, and wherein the step of re-inflating is performed to place the stent valve within the luminal organ at or near the valve annulus. 5. The method of claim 2, further comprising: positioning the stent valve around at least part of a second device;introducing the at least part of the second device into the luminal organ; andpositioning the stent valve within the luminal organ using the second device. 6. The method of claim 1, wherein the first device is configured as a catheter having a lumen therethrough and defining a suction/infusion port so that the lumen is in communication with the balloon, and wherein the step of inflating is performed by introducing a fluid through the lumen, through the suction/infusion port, and into the balloon. 7. The method of claim 1, wherein the first device further comprises at least one electrode positioned within the balloon, and wherein the step of obtaining the first valve annulus measurement is performed by obtaining a size measurement within the balloon using the electrode. 8. The method of claim 1, further comprising: introducing a second device into the luminal organ, the second device selected from a plurality of devices based upon the first valve annulus measurement. 9. A method to size a percutaneous valve opening, comprising: introducing at least part of a first device into a luminal organ at a percutaneous valve opening, the first device having a balloon positioned thereon;inflating the balloon within the percutaneous valve opening until point of apposition is detected; andobtaining a first percutaneous valve opening measurement based upon the point of apposition. 10. The method of claim 9, further comprising: selecting a stent valve based upon the first percutaneous valve opening measurement. 11. The method of claim 10, further comprising: positioning the stent valve upon the balloon;re-introducing the at least part of the first device into the luminal organ; andre-inflating the balloon to place the stent valve within the luminal organ. 12. The method of claim 11, wherein the step of re-introducing is performed to ... | 3DT Holdings, LLC | A61B005/05 A61B005/117 A61B005/103 A61B005/00 A61B005/053 A61B005/107 A61B017/00 A61F002/24
US20160135721 | A1 | 2016-05-19 | Catheter Comprising a Detection Device for Supplying Real-Time Detection of a Sample Material | 2014-06-23 | The present invention relates to a catheter (1). To provide a device permitting the real-time tracing and detection of sample material (2), the catheter (1) according to the invention comprises at least one detection device (4, 4a) for supplying real-time detection of the sample material (2), wherein the at least one detection device (4, 4a) has a functionalized surface (5, 5a) for accumulating the sample material (2), a signal converter (6), which converts the accumulation of sample material (2) on the functionalized surface (5, 5a) into a binding signal (7), and a signal line (8) for transmitting the binding signal (7). | 1. A catheter comprising at least one detection device for the real-time detection of a sample material wherein the at least one detection device comprises a functionalized surface for the accumulation of the sample material, a signal converter which converts the accumulation of the sample material on the functionalized surface into a binding signal and a signal line for transmitting the binding signal. 2. The catheter according to claim 1 wherein the functionalized surface is arranged on the external side and/or the inner side of the catheter. 3. The catheter according to claim 1 wherein the at least one detection device is undetachably connected to the catheter and further wherein the at least one detection device is formed integrally with the catheter. 4. The catheter according to claim 1 wherein the functionalized surface is at least sectionally loaded with detection molecules. 5. The catheter according to claim 4 wherein the detection molecules comprise antibodies, binding fragments of antibodies, antigens, peptides, proteins, nucleic acids, ligands, receptors, chelates, haptens, enzymes, enzyme inhibitors, enzyme substrates, cofactors of enzymes, endotoxins or other molecules which specifically bind with the sample material arc used as the detection molecules. 6. The catheter according to claim 1 wherein the detection molecules bind with pathogen specific and/or pathogen associated sample material. 7. The catheter according to claim 1 wherein at least one of the detection molecules and the functionalized surface structurally changes upon binding of the sample material. 8. The catheter according to claim 1 wherein the signal converter and/or the catheter is at least sectionally coated with a polymer. 9. The catheter according to claim 8 wherein the polymer has functional groups. 10. The catheter according to claim 1 wherein the detection molecules are couple directly via a covalent bond or indirectly via a coupling molecule to the signal converter or a polymer coating the signal converter. 11. The catheter according to claim 1 wherein the at least one detection device further comprises an electrochemical, an optical, an acoustical, an electrical, a thermal and/or a piezo-electric signal converter. 12. The catheter according claim 11 wherein the signal converter further comprises at least one electrode or an optical fiber section. 13. The catheter according to claim 1 wherein the signal line is an electrical conductor or an optical fiber. 14. The catheter according to claim 1 wherein the signal line is arranged at the catheter and/or is enclosed in the catheter. 15. The catheter according to claim 1 wherein the at least one detection device further comprises a shield against disruptive factors and/or at least a reference measuring device. 16. The catheter according to claim 3 wherein the at least one detection device is undetachably connected to the catheter, and further wherein the at least one detection device is integrated i... |  | A61B005/145 A61B005/1459 A61B005/00 A61B005/1473
US09649062 | B2 | 2017-05-16 | Swallowing estimation device, information terminal device, and storage medium | 2016-02-01 | A swallowing estimation device includes: a sound detection part configured to detect sound of a larynx portion; a respiration detection part configured to detect respiration; and a swallowing estimation part configured to estimate swallowing based on sound information outputted from the sound detection part and based on respiration information outputted from the respiration detection part. The swallowing estimation part obtains a value of a parameter for swallowing estimation with respect to a biological sound generation interval that corresponds to a respiratory cessation interval longer than or equal to 400 msec, and estimates whether swallowing has occurred in the biological sound generation interval based on whether the obtained value of the parameter satisfies a swallowing determination condition. | 1. A swallowing estimation device comprising: a sound detector that detects sound of a larynx portion;a respiration detector that detects respiration;a computer and a non-transitory storage medium containing a program for causing the computer to estimate swallowing based on sound information outputted from the sound detector and based on respiration information outputted from the respiration detector, wherein estimating swallowing comprises: obtaining a value of a parameter for swallowing estimation with respect to a biological sound generation interval that corresponds to a respiratory cessation (apnea) interval longer than or equal to 400 msec; andestimating whether swallowing has occurred in the biological sound generation interval based on whether the obtained value of the parameter satisfies a swallowing determination condition;an output part configured to output information based on a result of the swallowing estimation performed, wherein the output part displays a screen in which a sound waveform based on the sound information outputted from the sound detector is superposed on a time axis, and a swallowing sound generation interval at which swallowing had been estimated to have occurred is indicated along the sound waveform;a storage part in which the sound information outputted from the sound detector is stored; andan input part capable of selecting, from the sound waveform displayed by the output part, the swallowing sound generation interval,wherein the output part obtains from the storage part the sound information having a time width including the swallowing sound generation interval selected via the input part, and outputs sound that is obtained by reproducing the obtained sound information, andwhen the swallowing sound generation interval is selected via the input part, the output part displays a zoomed-in state of the selected swallowing sound generation interval. 2. The swallowing estimation device according to claim 1, wherein the program further causes the computer to: detect, from the respiration information, respiratory phases before and after a timing at which swallowing had been estimated to have occurred, and evaluate whether there has been a possibility of aspiration at the timing based on the detected respiratory phases, and the output part outputs information based on a result of the evaluation. 3. The swallowing estimation device according to claim 1, wherein the program further causes the computer to: detect, from the sound information, inspiration sound and expiration sound before and after a timing at which swallowing had been estimated to have occurred, and evaluate whether there has been a possibility of aspiration at the timing based on the detected inspiration sound and expiration sound, and the output part outputs information based on a result of the evaluation. 4. The swallowing estimation device according to claim 1, wherein when the swallowing sound generation interval is selected via the input part, the o... | HYOGO COLLEGE OF MEDICINE EUSENSE MEDICAL CO., LTD. | A61B005/00 A61B005/08 A61B007/00 A61B005/087 A61B005/11
US09545226 | B2 | 2017-01-17 | Detecting, quantifying, and/or classifying seizures using multimodal data | 2015-10-20 | A method, comprising receiving at least one of a signal relating to a first cardiac activity and a signal relating to a first body movement from a patient; triggering at least one of a test of the patient's responsiveness, awareness, a second cardiac activity, a second body movement, a spectral analysis test of the second cardiac activity, and a spectral analysis test of the second body movement, based on at least one of the signal relating to the first cardiac activity and the signal relating to the first body movement; determining an occurrence of an epileptic event based at least in part on said one or more triggered tests; and performing a further action in response to said determination of said occurrence of said epileptic event. Further methods allow classification of epileptic events. Apparatus and systems capable of implementing the method. | 1. A method, comprising: receiving at least one of a kinetic signal indicative of the patient's kinetic activity and a cardiac signal indicative of the patient's cardiac activity;determining at least one of at least one kinetic feature from the kinetic signal and at least one cardiac feature from the cardiac signal;detecting an epileptic seizure based on the at least one of the at least one kinetic feature and the at least one cardiac feature;receiving a signal indicative of the patient's dermal activity;determining a dermal activity feature from the signal indicative of the patient's dermal activity; andclassifying the epileptic seizure, based at least in part on the dermal activity feature. 2. The method of claim 1, further comprising: performing a further action in response to said detecting, wherein said further action comprises one or more of:logging the occurrence of the epileptic seizure; logging the time of occurrence of the epileptic seizure; logging the date of occurrence of the epileptic seizure; providing at least one of a warning, AIArm or alert to the patient, a caregiver or a health care provider; providing a therapy to prevent, abort, reduce the severity, or reduce the duration of the epileptic seizure; assessing at least one of the awareness or responsiveness of the patient during the epileptic seizure; assessing the severity of the epileptic seizure; determining the end of the epileptic seizure; determining the beginning of a post-ictal period; determining the end of a post-ictal period; and assessing the patient's post-ictal impairment or recovery from the epileptic seizure. 3. The method of claim 1, further comprising: performing a further action in response to said classifying, wherein said further action comprises one or more of: logging the occurrence of the classified epileptic seizure; logging the time of occurrence of the classified epileptic seizure; logging the date of occurrence of the classified epileptic seizure; providing at least one of a warning, AIArm or alert to the patient, a caregiver or a health care provider; providing a therapy to prevent, abort, reduce the severity, or reduce the duration of the classified epileptic seizure; assessing at least one of the awareness or responsiveness of the patient during the classified epileptic seizure; assessing the severity of the classified epileptic seizure; determining the end of the classified epileptic seizure; determining the beginning of a post-ictal period; determining the end of a post-ictal period; and assessing the patient's post-ictal impairment or recovery from the classified epileptic seizure. 4. The method of claim 1, wherein the cardiac activity is heart rate, heart rate variability, or heart beat morphology. 5. The method of claim 1, wherein the signal indicative of the patient's cardiac activity is provided by at least one of an electrocardiogram (EKG), phonocardiogram (PKG), apexcardiography, blood pressure, and echocardiography. 6. The method of... | FLINT HILLS SCIENTIFIC LLC | A61B005/00 A61N001/36 A61B005/0205 A61B005/11 A61B005/0402 A61B005/08 A61B005/113 A61B005/0245 A61N001/365 A61B005/021 A61B005/024 A61B005/0488 A61B005/053 A61B007/04 A61B008/08
US09888853 | B2 | 2018-02-13 | Method for monitoring a physiological characteristic of human body based on WBAN | 2013-08-15 | The provided are a method and system for monitoring one or more physiological characteristic based on WBAN. The method includes: an node of a micro-network for monitoring the human body acquires physiological characteristic information of the human body; and a hub of the micro-network for monitoring the human body recognizes at least one identifier for identifying a physiological status of the human body according to the acquired physiological characteristic information of the human body, and selects, according to the at least one identifier for identifying the physiological status of the human body, whether to send the physiological characteristic information of the human body to a medical monitoring system terminal via a wireless network or not. According to the method, sign information of the monitored human body may be mastered in first time to provide a strong basis for subsequent treatment. | 1. A method for monitoring one or more physiological characteristics of a human body based on a Wireless Body Area Network (WBAN), applied in a micro-network for monitoring the human body based on a WBAN-technology, wherein the micro-network comprises a hub and at least one node arranged on a body surface of a monitored human body or inside the monitored human body, and the method comprises: acquiring, by the node, physiological characteristic information of the human body; andrecognizing, by the hub, at least one identifier for identifying a physiological status of the human body according to the acquired physiological characteristic information of the human body, and selecting, by the hub, according to the at least one identifier for identifying the physiological status of the human body, whether to send the physiological characteristic information of the human body to a medical monitoring system terminal via a wireless network or not;wherein the node is activated to acquire the physiological characteristic information of the human body after receiving an instruction of the hub;wherein, when the node is not activated by the hub, the following processing is performed respectively:when the node is in an inactive state in a non-wakeup beacon period, the node enters a hibernating state; when the node is in a wakeup beacon period, and is not required to receive a beacon, the node is in the inactive state within transmission time of the beacon; and when the node is not required to transmit a management type frame or a data type frame, the node is in the inactive state in a dedicated access stage and random access stage of a superframe. 2. The method as claimed in claim 1, wherein, before the hub and the node exchange a data type frame or a management type frame, mutual discovery is performed between the hub and the node, wherein the mutual discovery comprises mutual discovery between the hub and a disconnected node and mutual discovery between the hub and a connected node, wherein the mutual discovery between the hub and the disconnected node is as follows: before frame transmission between the hub and a disconnected node is performed, the hub ensures that the disconnected node is waken up and is in a same channel with the hub, and then provides a polling allocation interval for the disconnected node; andthe mutual discovery between the hub and the connected node is as follows: before frame transmission between the hub and a connected node is performed, the hub sends a wakeup frame and a polling frame to the connected node and provides a polling allocation interval for the connected node. 3. The method as claimed in claim 1, wherein, in a process of service interaction between the hub and the node, the hub or the node performs a power control through power adjustment information carried in an interacted frame, to increase or reduce a current power level of an opposite party. 4. The method as claimed in claim 3, wherein the power control is activel... | ZTE CORPORATION | G08C017/00 H04L012/16 H04Q011/00 A61B005/00 G06F019/00 G08B021/02 G08B021/04 A61B005/024
US10714208 | B2 | 2020-07-14 | Computer systems for treating diseases | 2016-01-12 | The present invention relates to computer-implemented methods and system for analysing a biomarker which cycles in a subject. In some other aspects, the present invention relates to analysing a biomarker which at least initially increases or decreases in amount in a subject following a treatment for a disease. In further aspects, the present invention relates to computer-implemented methods and systems for determining a preferred time to administer a therapy to treat a disease in a subject. The present invention also relates to computer program product to implement the methods. Further, the present invention relates to methods of determining the timing of treating a disease in a subject in which the immune system is cycling. | 1. A non-transitory computer-readable information storage medium having stored thereon instructions, that if executed by one or more processors, perform a method comprising: i) obtaining from a first memory measurements of an immune system marker in samples from a subject having a disease characterized by cycling immune system markers;ii) obtaining from a second memory a model of cycling of the immune system marker;iii) fitting the measurements to the model;iv) obtaining a best-fit curve of the measurements to the model to estimate a periodicity of the cycling of the immune system marker;v) projecting the obtained best-fit curve into the future; andvi) outputting a preferred time in the future, based on the projected best-fit curve, to administer a therapy to treat the disease. 2. The computer-readable information storage medium of claim 1, wherein the best-fit curve is a fit to a periodic model of the cycling of the immune system marker, and obtaining the best-fit curve further comprises fitting the periodic model to the measurements multiple times using different initial conditions for each fit. 3. The computer-readable information storage medium of claim 1, wherein the best-fit curve is a fit to a harmonic model of the cycling of the immune system marker, and obtaining the best-fit curve further comprises fitting the harmonic model to the measurements multiple times using different initial conditions for each fit. 4. The computer-readable information storage medium of claim 1, wherein the method further comprises generating a confidence measure by comparing a similarity in a fit result produced by each fitting. 5. The computer-readable information storage medium of claim 1, wherein obtaining the best-fit curve further comprises imposing a box constraint on at least some parameters of the model to guide optimization to biologically realistic regions. 6. The computer-readable information storage medium of claim 1, wherein the model has a heavy tail to allow for noisy nature of the measurements. 7. The computer-readable information storage medium of claim 1, wherein the method further comprises determining a variance-stabilizing transformation of the measurements prior to determining the periodicity based on the transformed measurements. 8. The computer-readable information storage medium of claim 7, wherein the variance-stabilizing transformation of measurements is logarithm of the measurements. 9. The computer-readable information storage medium of claim 1, wherein the measurements are for two or more different immune system markers which are cycling in the subject. 10. The computer-readable information storage medium of claim 1, wherein the disease is cancer, an infection, a degenerative disease, an autoimmune disease or transplant rejection. 11. The computer-readable information storage medium of claim 10, wherein the disease is cancer. 12. The computer-readable information storage medium of claim 1, wherein the immune system marke... | BIOTEMPUS PTY LTD | G16B005/00 G16B040/00 G01N033/564 G01N033/574 G16H050/50 A61B005/00 G06F019/00 A61K031/495 G01N033/68
US09826938 | B2 | 2017-11-28 | Motion compensation for optical heart rate sensors | 2015-09-14 | A method of optical heart rate sensing includes receiving a motion frequency from a motion sensor and receiving an optical signal from an optical sensor. The motion frequency is then filtered from the optical signal, and an estimated heart rate frequency is determined based on the filtered optical signal. A heart rate is reported that is based on the motion frequency, but not based on the estimated heart rate frequency, when a local neighbor magnitude ratio of the estimated heart rate frequency is below a confidence threshold. | 1. A method of optical heart rate sensing, comprising: receiving a motion frequency via a motion sensor;receiving an optical signal from an optical sensor;filtering the motion frequency from the optical signal;determining an estimated heart rate frequency based on the filtered optical signal;reporting a heart rate based on the motion frequency, but not based on the estimated heart rate frequency responsive to a local neighbor magnitude ratio of the estimated heart rate frequency being less than a confidence threshold; andreporting a heart rate based on the estimated heart rate frequency responsive to the local neighbor magnitude ratio of the estimated heart rate frequency being greater than the confidence threshold. 2. The method of claim 1, further comprising: dynamically adjusting the confidence threshold based on the estimated heart rate frequency. 3. The method of claim 1, wherein the local neighbor magnitude ratio of the estimated heart rate frequency comprises a ratio of a magnitude of the estimated heart rate frequency to a sum of magnitudes of heart rate frequencies within a threshold of the estimated heart rate frequency. 4. The method of claim 3, wherein the magnitudes of heart rate frequencies within a threshold of the estimated heart rate frequency include magnitudes of heart rate frequencies within 30 beats per minute of the estimated heart rate frequency. 5. The method of claim 3, further comprising: transforming a recent window of the filtered optical signal from a time domain to a frequency domain;determining the magnitude of the estimated heart rate frequency based on the transformed filtered optical signal; anddetermining magnitudes of heart rate frequencies within a threshold of the estimated heart rate frequency based on the transformed filtered optical signal. 6. The method of claim 5, wherein the recent window of the filtered optical signal is a recent five second window. 7. The method of claim 1, wherein an amount of motion signal filtering remains constant with changing motion frequency. 8. The method of claim 7, wherein filtering the optical signal based on the motion frequency includes applying a comb filter to the optical signal. 9. The method of claim 8, wherein filtering the optical signal based on the motion frequency further includes: adjusting one or more parameters of the comb filter based on a recent motion frequency while maintaining the amount of motion signal filtering. 10. An optical heart rate sensor, comprising: an optical source configured to illuminate one or more blood vessels through a user's skin;an optical sensor configured to measure illumination from the one or more blood vessels;one or more motion sensors; anda storage machine holding instructions executable by a logic machine to: receive an optical signal from the optical sensor;receive a motion signal from the one or more motion sensors;determine a motion frequency based on the motion signal;filter the motion frequency from the optical s... | MICROSOFT TECHNOLOGY LICENSING, LLC | A61B005/02 A61B005/00 A61B005/024 A61B005/01
US09345449 | B2 | 2016-05-24 | Prosthetic component for monitoring joint health | 2014-03-11 | A prosthetic component suitable for long-term implantation is provided. The prosthetic component includes electronic circuitry and sensors to measure a parameter of the muscular-skeletal system. The prosthetic component comprises a first structure having at least one support surface, a second structure having at least one feature configured to couple to bone, and at least one sensor. The electronic circuitry and sensors are hermetically sealed within the prosthetic component. The prosthetic component includes at least on transmissive region. The transmissive region can be located in a region that has exposure to a region outside the joint. The transmissive region can comprise glass. One or more sensors can be used to monitor synovial fluid in proximity to the joint to determine joint health. The transmissive region can be used to support communication between the electronic circuitry and remote system. | 1. An orthopedic system comprising: a sensor;electronic circuitry coupled to the sensor where the electronic circuitry includes a transmitter and where the transmitter is configured to transmit measurement data; anda light source configured to illuminate a region in proximity to the muscular-skeletal system where the electronic circuitry is in a housing and where the housing is coupled to the muscular-skeletal system. 2. The orthopedic system of claim 1 further including a remote system configured to receive and process the measurement data. 3. The orthopedic system of claim 1 where the housing has a transmissive region. 4. The orthopedic system of claim 3 where the transmissive region comprises glass. 5. The orthopedic system of claim 3 where the transmissive region comprises plastic. 6. The orthopedic system of claim 3 where the transmitter is configured to transmit through the transmissive region. 7. The orthopedic system of claim 3 where the sensor is an optical sensor and where the optical sensor is configured to detect through the transmissive region. 8. The orthopedic system of claim 1 where the sensor is an optical sensor and where the optical sensor measures color or turbidity. 9. The orthopedic system of claim 1 where the sensor is a temperature sensor or a pH sensor. 10. An orthopedic system comprising: a sensor;electronic circuitry coupled to the sensor where the electronic circuitry includes a transmitter and where the transmitter is configured to transmit measurement data; anda light emitting diode configured to illuminate a region in proximity to the muscular-skeletal system where the electronic circuitry is in a housing and where the housing is coupled to the muscular-skeletal system. 11. The orthopedic system of claim 10 where the housing has a transmissive region. 12. The orthopedic system of claim 11 where the transmissive region comprises glass. 13. The orthopedic system of claim 11 where the transmissive region comprises plastic. 14. The orthopedic system of claim 11 where the transmitter is configured to transmit through the transmissive region. 15. The orthopedic system of claim 11 where the sensor is an optical sensor and where the optical sensor is configured to detect through the tranmissive region. 16. The orthopedic system of claim 15 further including a pH or temperature sensor. 17. The orthopedic system of claim 15 where the optical sensor is configured to detect color or turbidity. 18. An orthopedic system comprising: a sensor;electronic circuitry coupled to the sensor where the electronic circuitry includes a transmitter and where the transmitter is configured to transmit measurement data;a light source configured to illuminate a region in proximity to the muscular-skeletal system where the electronic circuitry is in a prosthetic component configured to couple to the muscular-skeletal system. 19. The orthopedic system of claim 18 further including a remote system configured to receive and process ... | ORTHOSENSOR INC | A61B005/00 A61B008/08 A61B005/01 A61B005/103 A61B005/145 A61B008/00
US09782092 | B2 | 2017-10-10 | Patient signal analysis based on vector analysis | 2014-11-14 | Disclosed herein is a framework for facilitating patient signal analysis based on vector analysis. In accordance with one aspect, a set of vectors is generated from a patient signal data waveform. The vectors may be directed from a common center to points of interest on the patient signal data waveform. The framework may further extract one or more vector parameters from the set of vectors, and determine one or more vector ratios based on the vector parameters to monitor changes in the patient signal data waveform. | 1. A computer-implemented method of patient signal analysis, comprising: receiving, from a patient monitor, a patient signal data waveform;generating, by a processor, a set of vectors directed from a common center to points of interest on the patient signal data waveform;extracting, by the processor, one or more vector parameters from the set of vectors; anddetermining, by the processor, one or more vector ratios based on the one or more vector parameters to monitor changes in the patient signal data waveform, wherein determining the one or more vector ratios comprises determining one or more unilateral ratios from the vectors or determining one or more bilateral ratios from the vectors. 2. The method of claim 1 wherein the patient signal data waveform comprises an electrocardiogram waveform. 3. The method of claim 1 wherein generating the vectors comprises generating the vectors directed from an R wave peak center point on the patient signal data waveform. 4. The method of claim 1 wherein generating the vectors comprises generating the vectors directed from the common center point to at least one of P, Q, S, T or U wave peak points on the patient signal data waveform. 5. The method of claim 1 wherein generating the vectors comprises generating the vectors directed from the common center point to an H point defined at a half cycle time point. 6. The method of claim 1 wherein extracting the one or more vector parameters comprises extracting one or more time durations measured from a half cycle time stamp to time stamps of the points of interest. 7. The method of claim 6 wherein determining the one or more vector ratios comprises determining at least one ratio of distances associated with two vectors selected from the set of vectors. 8. The method of claim 6 wherein determining the one or more vector ratios comprises determining at least one ratio of time durations associated with two vectors selected from the set of vectors. 9. The method of claim 6 wherein determining the one or more vector ratios comprises determining at least one ratio comparing a vector amplitude associated with a first vector to a time duration associated with a second vector, wherein the first and second vectors are selected from the set of vectors. 10. The method of claim 9 wherein the first and second vectors are different. 11. The method of claim 9 wherein the first and second vectors are same. 12. The method of claim 1 wherein extracting the one or more vector parameters comprises extracting vector circles, wherein each of the vector circles is centered at the common center and has a radius defined by one of the vectors. 13. The method of claim 12 wherein determining the one or more vector ratios comprises determining at least one ratio of areas associated with two of the vector circles. 14. The method of claim 12 wherein determining the one or more vector ratios comprises determining at least one ratio that compares an area associated with a vector circle o... | SIEMENS MEDICAL SOLUTIONS USA, INC. | A61N001/00 A61B005/04 A61B005/00 A61B005/0456 A61B005/046 A61B005/0464 A61N001/365
US09332931 | B2 | 2016-05-10 | Pulsating expression cap | 2008-10-03 | A body fluid sampling device includes an expression member placed around an incision in the skin of a user, a compression member to contact and traverse the expression member, and a housing configured to contact the skin near the incision. The expression member is configured to repeatedly move from a relaxed position over the incision to a contracted position over the incision. The expression member is at least partially resilient in order to repeatedly squeeze and/or compress the skin around the incision. | 1. A method of expressing a body fluid, comprising: positioning a sampling device having an expression member and a compression member over an incision in skin, the expression member having a resilient band for contacting the skin surrounding the incision, the expression member including a plurality of expression lobes and the compression member including a plurality of compression lobes; andrepeatedly rotating the compression member about the expression member between a first position and a second position, wherein the first position includes the plurality of compression lobes contacting the plurality of expression lobes to deform the resilient band of the expression member, the deformed resilient band squeezing and compressing the skin that surrounds the incision to express bodily fluid from the incision, and wherein the second position includes the plurality of compression lobes contacting the expression member and releasing pressure on the skin that surrounds the incision, the resilient band returning to an undeformed shape. 2. The method of claim 1, further comprising: compressing the skin that surrounds the incision with a housing to urge the bodily fluid to the incision, the housing contacting the skin surrounding the incision. 3. The method of claim 1, further comprising: collecting the bodily fluid with a test strip of the sampling device. 4. The method of claim 1, further comprising: releasing pressure on the skin that surrounds the incision as the plurality of compression lobes traverse the plurality of expression lobes. 5. The method of claim 4, further comprising: varying the deformation of the expression member to squeeze and compress the skin that surrounds the incision in an asymmetrical pattern. 6. A method of expressing a body fluid, comprising: positioning a sampling device having an expression member and a compression member over an incision in skin, the expression member having a deformable end portion for contacting the skin surrounding the incision, the expression member includes a plurality of expression lobes and the compression member includes a plurality of compression lobes; andrepeatedly contacting and traversing the plurality of expression lobes with the plurality of compression lobes as the compression member is rotated about the expression member from a first position to a second position, wherein the first position includes deforming the deformable end portion, the deformed end portion pulsating the skin surrounding the incision to pump bodily fluid from the incision, and wherein the second position includes the deformed end portion returning to an undeformed shape, the undeformed end portion releasing pressure on the skin surrounding the incision. 7. The method of claim 6, wherein the deformable end portion includes a plurality of resilient fingers configured in a semi-circular shape; and wherein pulsating the skin surrounding the incision includes dilating the plurality of resilient fingers as the pluralit... | Roche Diabetes Care, Inc. | A61B005/15
US09351668 | B2 | 2016-05-31 | Signal processing for continuous analyte sensor | 2009-10-12 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A glucose sensor system for monitoring glucose concentration in a biological sample of a host, the system comprising: a continuous glucose sensor configured to produce sensor data indicative of a host's glucose concentration, wherein the continuous glucose sensor comprises a sensing region and a sensing membrane located over the sensing region, and wherein at least a portion of the glucose sensor is configured for transcutaneous implantation; andan electronics module configured to receive sensor data from the continuous glucose sensor, the electronics module comprising a processor configured to process sensor data received from the continuous glucose sensor for a first time period so as to generate real-time glucose readings from the sensor data, to detect an approaching clinical risk based at least in part on an acceleration of the host's glucose concentration, and to determine a possibility of avoiding the approaching clinical risk, wherein the electronics module delays activation of an alarm when the possibility of avoiding the approaching clinical risk is determined. 2. The system of claim 1, wherein detecting an approaching clinical risk further comprises determining that a predicted potential glucose concentration value is greater than a high threshold value or less than a low threshold value. 3. The system of claim 1, wherein the alarm comprises an audible signal. 4. The system of claim 1, wherein the alarm comprises a tactile signal. 5. The system of claim 2, wherein the high threshold value and the low threshold value are user configurable. 6. The system of claim 1, wherein the processor module is further configured to calibrate the sensor data using a glucose concentration measured by a reference glucose monitor operatively linked to the electronics module, wherein the reference glucose monitor is integral with the electronics module. 7. The system of claim 1, wherein detecting an approaching clinical risk further comprises predicting a potential glucose concentration for a second time period based at least in part on both an extrapolation of real time glucose readings and on a variation analysis, and wherein the variation analysis comprises a statistical analysis. 8. The system of claim 1, wherein detecting an approaching clinical risk further comprises predicting a potential glucose concentration for a second time period based at least in part on both an extrapolation of real time glucose readings and on a variation analysis, and wherein the variation analysis comprises a clinical analysis. 9. The system of claim 1, wherein detecting an approaching clinical risk further comprises predicting a potential glucose concentration for a second time period based at least in part on both an extrapolation of real time glucose readings and on a variation analysis, and wherein the variation analysis comprises a physiological analysis. 10. The system of claim 1, wherein delay is for a predetermined amount of time. 11. A device for pro... | DexCom, Inc. | A61B005/145 A61B005/1468 A61B005/1486 A61B005/1495 A61B005/00
US09339177 | B2 | 2016-05-17 | Full-eye illumination ocular surface imaging of an ocular tear film for determining tear film thickness and/or providing ocular topography | 2014-11-18 | Ocular surface interferometry (OSI) devices, systems, and methods are disclosed for measuring a tear film layer thickness (TFLT) of the ocular tear film, including the lipid layer thickness (LLT) and/or the aqueous layer thickness (ALT). The TFLT can be used to diagnose dry eye syndrome (DES). Certain embodiments also include ocular topography devices, systems and methods for deducing corneal shape by capturing an image of a target reflecting from the surface of the cornea. The image of the target contains topography information that is reviewable by a clinician to diagnose the health of the patient's eye by detecting corneal aberrations and/or abnormalities in corneal shape. Certain embodiments also include a combination of the OSI and ocular topography devices, systems and methods to provide imaging that can be used to yield a combined diagnosis of the patient's tear film and corneal shape. | 1. An apparatus for imaging an eye, comprising: a multi-wavelength light source configured to emit light to an eye; anda control system configured to: (a) spatially modulate light from the multi-wavelength light source to project a first circular pattern onto the eye, such that at least one first portion of the eye receives emitted light from the multi-wavelength light source and at least one second portion of the eye does not receive the emitted light from the multi-wavelength light source;(b) receive at an imaging device, at least one first image containing at least one first signal associated with an ocular property of the eye comprising the emitted light reflected from the at least one first portion of the eye;(c) spatially modulate light from the multi-wavelength light source to project a second circular pattern onto the eye, such that the at least one first portion of the eye does not receive the emitted light from the multi-wavelength light source and the at least one second portion of the eye receives the emitted light from the multi-wavelength light source;(d) receive at the imaging device, at least one second image containing at least one second signal associated with the ocular property of the eye comprising the emitted light reflected from the at least one second portion of the eye; and(e) subtract the at least one second image from the at least one first image to generate at least one resulting image containing a resultant signal that represents the ocular property of the eye;wherein: the at least one first signal is optical wave interference of specularly reflected light from an ocular tear film combined with at least one background signal; andthe control system is configured to receive the at least one first image by being configured to: capture a first circular tiled pattern of the specularly reflected light including the at least one background signal from the at least one first portion of the ocular tear film in the at least one first image by the imaging device; andcapture a second circular tiled pattern of the specularly reflected light including the at least one background signal from the at least one second portion of the ocular tear film in the at least one second image by the imaging device. 2. The apparatus of claim 1, further comprising a light modulator system communicatively coupled to the control system to spatially modulate light from the multi-wavelength light source to alternately project the first circular pattern and the second circular pattern onto the eye thereby generating the at least one first signal and the at least one second signal of the ocular property. 3. The apparatus of claim 2, wherein the light modulator system includes a disk disposed between the imaging device and eye for sequentially projecting the first circular pattern and the second circular pattern onto the eye when the disk is illuminated, and wherein the disk includes a central aperture for the imaging device to receive light reflected f... | TearScience, Inc. | A61B003/14 A61B003/10 A61B003/02 A61B005/05 A61B003/00
US09357339 | B2 | 2016-05-31 | Proxy communication system and control method thereof in ban environment | 2012-11-08 | Disclosed is a control method of a communication system including at least one sensor that a user wears, a wireless communication apparatus and a coordinator. The control method of the communication system includes: in the coordinator, receiving a proxy authority request for proxy with respect to the at least one sensor and the coordinator from the wireless communication apparatus, and granting the proxy authority to the wireless communication apparatus; in the wireless communication apparatus, creating a user identifier corresponding to the user; in the wireless communication apparatus, searching for the at least one sensor that the user wears, and forming a pairing with the at least one found sensor; in the wireless communication apparatus, making a request to the at least one found sensor for association information for the association with the coordinator and the sensor, and receiving the association information; in the wireless communication apparatus, making a request to the coordinator for association proxy for an association between the at least one found sensor and the coordinator, and receiving the association proxy in response to the request; and in the coordinator, requesting data by forming an association with the at least one found sensor, and receiving data corresponding to the data request. | 1. A control method of a communication system including at least one sensor that a user wears, a wireless communication apparatus and a coordinator, the method comprising: transmitting, by the wireless communication apparatus, a proxy authority request for proxy with respect to the at least one sensor and the coordinator to the coordinator, and receiving a proxy authority from the coordinator;creating, by the wireless communication apparatus, a user identifier corresponding to the user who wears the at least one sensor;searching, by the wireless communication apparatus, for the at least one sensor that the user wears, and forming a pairing with the at least one searched sensor;transmitting, by the wireless communication apparatus, a request to the at least one searched sensor for association information for the association with the coordinator and the sensor, and receiving the association information, the request for association information comprising information on the coordinator and the association information comprising an identifier of the at least one searched sensor; andtransmitting, by the wireless communication apparatus, a request to the coordinator for an association proxy for the association between the at least one searched sensor and the coordinator, and receiving the association proxy comprising the identifier of the at least one searched sensor in response to the request for direct communication between the coordinator and the at least one searched sensor. 2. The method as claimed in claim 1, further comprising, after the receiving the association information, transmitting, by the wireless communication apparatus, the user identifier to the at least one searched sensor. 3. The method as claimed in claim 1, wherein, in the creating the user identifier, the user identifier is created on the basis of an external input for the user identifier. 4. The method as claimed in claim 1, wherein, in creating the user identifier, the user identifier is created on the basis of a RF signal received from an RF signal generating means located around the user. 5. The method as claimed in claim 1, wherein, in the forming of the pairing, the pairing is formed on the basis of the near-field communication (NFC) technique. 6. The method as claimed in claim 1, wherein, in the forming of the pairing, when the type of communication technique that the at least one searched sensor uses is identical, the pairing with respect to each of the at least one searched sensor is formed sequentially. 7. The method as claimed in claim 1, wherein the association information further comprises information on the sensor. 8. The method as claimed in claim 7, wherein the information on the sensor is at least one of a serial number, a form, and a battery type of the sensor. 9. The method as claimed in claim 7, wherein, if the coordinator includes the database with the information on the sensor, the association proxy comprises the identifier of the sensor. 10. The met... | Samsung Electronics Co., Ltd. | H04B007/00 H04W004/00 A61B005/00 G08C017/02 H04L029/08 H04W012/06
US09999372 | B2 | 2018-06-19 | Combination of inert gas rebreathing and multiple-breath wash-out techniques for determination of indices of ventilation inhomogeneity | 2015-12-29 | The present invention discloses a method to determine the lung clearance index (LCI) or other indices of ventilation inhomogeneity of the lungs by combining two pulmonary gas exchange techniques; Inert gas rebreathing (IGR) is used for rapid wash-in of the inert tracer gas and this is followed by multiple-breath wash-out (MBW). The functional residual capacity (FRC) can either be determined from the tracer gas concentration and the gas flow inhaled and exhaled during multiple-breath wash-out or by gas analysis alone from the inert gas rebreathing. The cumulative expired volume (VCE) required to clear the inert tracer gas from the lungs is determined from the multiple-breath wash-out, and LCI is calculated as the ratio between VCE and FRC. The advantages of the method are i) significant reduction of required test time, ii) significant reduction of consumed gas mixture for wash-in of tracer gas, iii) potential for further reduction of the use of tracer gas, and iv) potential for more accurate determination of the FRC by gas dilution alone. Furthermore the present invention relates to a corresponding system and computer-readable medium. | 1. A method of determining a lung clearance index (LCI) of a test subject, said method comprising the steps of: in a closed rebreathing assembly, rebreathing an inert tracer gas mixture having a starting concentration of an inert tracer gas until a constant concentration of the inert tracer gas is reached to wash-in the inert tracer gas, wherein the closed rebreathing assembly comprises a rebreathing bag filled with a volume of the inert tracer gas mixture corresponding to the resting tidal volume of the test subject;performing multiple-breath wash-out until end-tidal tracer gas concentration has fallen below a predetermined fraction of the starting concentration, where the concentrations are measured by a gas analysis; anddetermining the lung clearance index (LCI) by: determining, via a processor, functional residual capacity (FRC);determining, via the processor, cumulative expired volume (VCE) required to clear the inert tracer gas concentration from the lungs below the predetermined fraction of the starting concentration; andcalculating, via the processor, the LCI as a ratio between VCE and FRC. 2. The method according to claim 1, wherein the inert tracer gas is SF6. 3. The method according to claim 1, wherein the inert tracer gas mixture is an oxygen enriched gas mixture. 4. The method according to claim 1, wherein the gas analysis is performed by a photoacoustic spectroscopy (PAS) sensitive gas analyzer. 5. The method according to claim 1, wherein the FRC is determined from the tracer gas concentration and a gas flow inhaled and exhaled during multiple-breath wash-out. 6. A system adapted to determine a lung clearance index (LCI) of the lungs of a test subject, said system comprising: a closed rebreathing setup configured to allow a test subject to rebreathe an inert tracer gas mixture having a starting concentration of an inert tracer gas until a constant concentration of the inert tracer gas is reached to wash-in the inert tracer gas during a rebreathing period, wherein the closed rebreathing setup comprises at least a rebreathing bag prefilled with a volume of the inert tracer gas mixture corresponding to the resting tidal volume of said test subject to rebreathe to and from during the rebreathing period;at least one gas analyzer for obtaining fractional concentration of said inert tracer gas inhaled and exhaled by said test subject;a flowmeter for monitoring gas flow inhaled and exhaled by said test subject; and a non-transitory computer readable medium having stored therein instructions for causing a processing unit of the system to execute said instructions, comprising:determining, via said processing unit, the LCI of the lungs of the test subject by determining functional residual capacity (FRC) and using said gas flow to determine, via said processing unit, cumulative expired volume (VCE) required to clear said inert tracer gas concentration from the lungs below a predetermined fraction of the starting concentration, wherein LCI... | IPR Holding APS | A61B005/08 A61B005/087 A61B005/093 A61B005/00 A61B005/091
US09351637 | B2 | 2016-05-31 | Cornea shape measurement apparatus | 2011-12-20 | A cornea shape measurement apparatus includes: a first projecting optical system configured to project a first target on the cornea; a first imaging optical system configured to capture a reflection image of the first target projected on the cornea; a second projecting optical system configured to project a second target different from the first target on the cornea to measure a corneal shape; a second imaging optical system configured to capture a reflection image of the second target projected on the cornea; and a calculation controller configured to adjust the amount of projection light of the second projecting optical system based on the reflection image of the first target, the reflection image having been acquired by the first imaging optical system, and measuring the corneal shape of an examinee's eye based on the reflection image of the second target, the reflection image having been captured by the imaging optical system. | 1. A cornea shape measurement apparatus for measuring a corneal shape of an examinee's eye, comprising: a first projecting optical system configured to project a first target on the cornea;a first imaging optical system configured to capture a reflection image of the first target projected on the cornea;a second projecting optical system configured to project a second target different from the first target on the cornea to measure the corneal shape;a second imaging optical system configured to capture a reflection image of the second target projected on the cornea; anda calculation controller configured to measure a first corneal shape of the examinee's eye based on the reflection image of the first target, the reflection image having been captured by the first imaging optical system, adjust the amount of projection light of the second projecting optical system based on the first corneal shape, and measure a second corneal shape of the examinee's eye based on the reflection image of the second target, the reflection image having been captured by the second imaging optical system,wherein the calculation controller measures a radius of curvature of the cornea based on the reflection image of the first target, the reflection image having been captured by the first imaging optical system, sets the amount of projection light of the second projecting optical system smaller as the radius of curvature is larger, and sets the amount of projection light of the second projecting optical system larger as the radius of curvature is smaller. 2. The cornea shape measurement apparatus according to claim 1, wherein the first projecting optical system includes an infrared light source, and projects the first target on the cornea with infrared light, andthe second projecting optical system includes a visible light source, and projects the second target on the cornea with visible light. 3. The cornea shape measurement apparatus according to claim 1, wherein the second projecting optical system projects a target having a wider projection area on the cornea than the first target, as the second target on the cornea. 4. The cornea shape measurement apparatus according to claim 2, wherein the second projecting optical system projects a target having a wider projection area on the cornea than the first target, as the second target on the cornea. 5. The cornea shape measurement apparatus according to claim 1, wherein the first imaging optical system serves also as the second imaging optical system. 6. The cornea shape measurement apparatus according to claim 2, wherein the first imaging optical system serves also as the second imaging optical system. 7. The cornea shape measurement apparatus according to claim 3, wherein the first imaging optical system serves also as the second imaging optical system. 8. The cornea shape measurement apparatus according to claim 4, wherein the first imaging optical system serves also as the second imaging optical system. 9. The cor... | NIDEK CO., LTD. | A61B003/10 A61B003/107 A61B005/107 A61B005/103
US09642571 | B2 | 2017-05-09 | System and method for sensorized user interface | 2015-12-02 | A system and method for is provided for operation of an orthopedic system. The system includes a load sensor for converting an applied pressure associated with a force load on an anatomical joint, and an inertial sensing device configured to measure alignment. The inertial sensing device provides alignment measurement data to measure alignment. An ultrasonic transducer, MEMs microphone, electromagnets, optical elements, metallic objects or other transducers can be configured to convert or convey a physical movement to an electrical signal and support measurement of muscular-skeletal alignment. The load sensor can be used to measure load magnitude and position of load of an applied load by the muscular-skeletal system. | 1. A medical system for generating quantitative measurements on the muscular-skeletal system, comprising: a load sensing device configured to couple to the muscular-skeletal system to measure a load magnitude and a position of load applied thereto;an inertial sensing device configured to measure muscular-skeletal alignment; anda remote system having a display configured to wirelessly couple to the inertial sensing device and the load sensing device wherein the remote system is configured to receive quantitative measurement data from the load sensing device and the inertial sensing device and where the remote system is configured to display load magnitude, position of load, and alignment data. 2. The system of claim 1, further comprising a detector unit coupled to the remote system for receiving visual, audio, or motion queues from within a surgical field. 3. The system of claim 1, wherein the inertial sensing device comprises a three-axis accelerometer. 4. The system of claim 1 wherein the inertial sensing device is configured to provide position measurement data and wherein the remote system comprises a processor, display, and software. 5. The system of claim 4 wherein the inertial sensing device is configured to measure 3 and 6 axis tilt during motion and while stationary. 6. The system of claim 1 wherein at least one of an ultrasonic transducer, MEMs microphone, electromagnets, optical elements, metallic objects or other transducers are configured to convert or convey a physical movement to an electrical signal and support measurement of muscular-skeletal alignment. 7. The system of claim 6 wherein a time of flight or a differential time of flight signal is configured to be measured by at least one of an ultrasonic transducer, MEMs microphone, electromagnets, optical elements, metallic objects or other transducers corresponding to muscular-skeletal alignment. 8. The system of claim 1 further including a detector unit coupled to the remote system for receiving visual, audio, or motion queues from within a surgical field where an action is generated upon receiving one of the visual, audio, or motion queues. 9. The system of claim 8, where the remote system comprises a computer system coupled to a display, where the display presents load magnitude, position of load, and alignment measurement data. 10. An operating room system comprising: one or more sensors configured to measure load and position of load of a muscular-skeletal system;an inertial sensing device configured to measure alignment of the muscular-skeletal system; anda remote system having a display configured to wirelessly receive load, position of load, and alignment quantitative measurement data load where the remote system is configured to display load magnitude, position of load, and alignment data to support installation of at least one prosthetic component; anda detector unit coupled to the remote system for receiving visual, audio, or motion queues from within a surgica... | ORTHOSENSOR INC | G08B003/00 A61B005/00 G06F003/01 G06F003/03 G06F003/0485 H04R025/00 G06F003/0488 G06F003/0346 A61B005/103 A61B008/08 H04L029/08 A61B005/11
US20160143581 | A1 | 2016-05-26 | NIR IMAGE-GUIDED TARGETING | 2016-01-29 | A method includes determining skin characteristics in a region of a patient. An in-treatment optical scan is performed on a region of a patient, wherein the in-treatment optical scan comprises a near infrared (NIR) energy source. A plurality of detected signals is detected from the optical scan. The skin characteristics are filtered out from the plurality of detected signals. Skeletal anatomy positioning associated with the region is determined from the plurality of signals that is filtered. | 1. A method for obtaining anatomical information, comprising: performing an optical scan on a region of a patient, wherein said optical scan comprises a near infrared (NIR) energy source;detecting a plurality of detected signals from said optical scan;filtering out skin characteristics from said plurality of detected signals; anddetermining skeletal anatomy associated with said region from said plurality of signals that is filtered. 2. The method of claim 1, further comprising determining skin characteristics for skin, wherein said determining skin characteristics for skin comprises: performing a computed tomography (CT) scan to determine at least one of skin reflectance characteristics, skin transmittive characteristics, and skin fluorescence characteristics. 3. The method of claim 1, further comprising determining skin characteristics for skin, wherein said determining skin characteristics for skin comprises: performing a magnetic resonance imaging (MRI) scan to determine at least one of skin reflectance characteristics, skin transmittive characteristics, and a skin fluorescence characteristics. 4. The method of claim 1, further comprising determining skin characteristics for skin, wherein said determining skin characteristics for skin comprises: performing another optical scan on said region of said patient, wherein said another optical scan comprises another source emitting NIR energy;varying intensity of said NIR energy for a given frequency; anddetermining signaling characteristics for skin related to at least one of skin reflectance characteristics, skin transmittive characteristics, and a skin fluorescence characteristics, such that said signaling characteristics are filterable from said plurality of detected signals. 5. The method of claim 1, further comprising determining skin characteristics for skin, wherein said determining skin characteristics for skin comprises: performing another optical scan on said region of said patient, wherein said another optical scan comprises another source emitting NIR energy;varying frequency of said NIR energy;determining signaling characteristics for skin related to at least one of skin reflectance characteristics, skin transmittive characteristics, and a skin fluorescence characteristics, such that said signaling characteristics are filterable from said plurality of detected signals. 6. The method of claim 1, further comprising determining skin characteristics for skin, wherein said determining skin characteristics for skin comprises: performing another optical scan using another source emitting NIR energy on said region of said patient, wherein said optical scan is performed over varying frequencies and intensities; anddetermining signaling characteristics for skin related to at least one of skin reflectance characteristics, skin transmittive characteristics, and a skin fluorescence characteristics, such that said signaling characteristics are filterable from said plurality of detected signals. ... |  | A61B005/00 A61N005/10 A61B005/055 A61B006/03 A61B006/00
US20160128602 | A1 | 2016-05-12 | TISSUE IDENTIFICATION DEVICE AND METHOD | 2014-06-11 | A tissue identification device for determining the presence and borders of the cancerous tissues includes a measurement probe having a concentric, 2 port coaxial structure and enabling the 1 Port and 2 Port scattering parameters of the target tissue to be measured. An S parameter measurement unit is associated with the measurement probe via an RF/Microwave cable and allows the S parameters of the tissue to be measured. A computer identifies the tissue by means of an application software block and a tissue identification software block therein in line with the measurement results from the S parameter measurement unit. A hand tool carrying the measurement probe thereon during the measurement process operates simultaneously with the computer by being in communication therewith. | 1. A tissue identification device for identification of the presence/absence and borders of the cancerous tissues, characterized in comprising: a measurement probe having a concentric, 2 Port coaxial structure, which allows measuring 1 Port and 2 Port scattering parameters of the target tissue,an S parameter measurement unit, which is associated with said measurement probe by means of an RF/Microwave cable and allows measuring S parameters of the tissues, anda computer which identifies the tissue by way of an application software block and tissue identification software block that it comprises, in line with the measurement results from S parameter measurement unit. 2. The tissue identification device (10) according to Clam 1, characterized in comprising a hand tool, which carries said measurement probe thereon during the measurement process, operates simultaneously with the computer by being in communication therewith, and is suitable for being used with a single hand, either by right or left hand. 3. The tissue identification device according to claim 2, characterized in comprising a calibration and sterilization unit, which comprises the required components for providing calibration and sterilization of the measurement probe. 4. The tissue identification device according to claim 1, characterized in that the measurement probe comprises; a port 1 conductive live end provided in the center; a dielectric material Type 1 enclosing said port 1 conductive live end; and a Port 1 and Port 2 conductive ground connection enclosing said dielectric material Type 1, which form the 1st port thereof; andsaid Port 1 and Port 2 conductive ground connection; dielectric material Type 2 enclosing the ground connection; and port 2 conductive live end enclosing the dielectric material Type 2, which form the 2nd port thereof. 5. The tissue identification device according to claim 4, characterized in comprising a non-conducting cover enclosing said 1st and 2nd ports of the measurement probe. 6. The tissue identification device according to claim 1, characterized in comprising a guide which enables the probe to be directly directed to the target tissue by guiding it inside the tissue during the measurement of dielectric permittivity of deep tissues. 7. The tissue identification device according to claim 1, characterized in comprising a measurement button which, when pushed thereon, transfers the information determined by the hand tool embedded software block of the micro-control card to the computer. 8. The tissue identification device according to claim 2, characterized in comprising a screen on which measurement results are shown, which enables the hand tool to be directed and informed about the measurement process during the surgery, and which is located on said hand tool. 9. The tissue identification device according to claim 6, characterized in comprising an elastic stopper, which covers the end of the measurement probe by being located inside the guide an... |  | A61B005/05 A61B090/70 A61B005/00
US09327129 | B2 | 2016-05-03 | Blended posture state classification and therapy delivery | 2009-04-30 | Techniques for classification of a posture state of a patient using multiple posture state definitions, and for delivering therapy according to the patient's classified posture state. Detected parameter values describing a patient's posture state are compared to posture state definitions. This comparison is used to determine similarity values describing how similar the patient's posture state is to each of the posture states described by the posture state definitions. Weighting factors may be determined from the similarity values and used to weight therapy parameter values that are associated with each of the posture state definitions. The resulting weighted therapy parameter values may be used to derive a blended therapy parameter value for use in delivering therapy to the patient. The patient's posture state may be expressed in terms of a blending of the multiple posture state definitions. | 1. A method, comprising: obtaining a detected parameter value for use in classifying a posture state of a patient;for each of multiple posture state definitions, obtaining a respective defined parameter value and a respective therapy parameter value;determining, for each of a plurality of the posture state definitions, a similarity between the detected parameter value and the respective defined parameter value;determining a blended therapy parameter value based on the similarities and respective therapy parameter values for the plurality of the posture state definitions; anddelivering therapy to the patient via a medical device according to the blended therapy parameter value. 2. The method of claim 1, wherein the detected parameter value is indicative of a posture of the patient, each of the defined parameter values is indicative of a respective defined posture, and wherein determining, for each of a plurality of the posture state definitions, a similarity comprises determining, for each of the plurality of the posture state definitions, how close the posture of the patient is to the defined posture that is indicated by the respective defined parameter value. 3. The method of claim 2, wherein the detected parameter value comprises a detected posture vector, wherein the respective defined parameter value comprises a respective defined posture vector, and wherein determining, for each of a plurality of the posture state definitions, a similarity further comprises determining, for each of the plurality of the posture state definitions, a cosine of an angle between the detected posture vector and the respective defined posture vector. 4. The method of claim 1, wherein determining, for each of a plurality of the posture state definitions, a similarity between the detected parameter value and the respective defined parameter value comprises selecting the plurality of the posture state definitions. 5. The method of claim 1, further comprising determining a respective weighting factor W for each of the similarities and wherein determining a blended therapy parameter value comprises using the respective weighting factor to weight a respective therapy parameter value. 6. The method of claim 5, wherein each weighting factor is determined using a weighting function. 7. The method of claim 6, wherein determining, for each of a plurality of the posture state definitions, a similarity between the detected parameter value and the respective defined parameter value comprises determining a distance between the detected parameter value and the respective defined parameter value, and further comprising selecting the weighting function to map the distance to a value between 0 and a predetermined maximum positive number. 8. The method of claim 5, wherein each of the plurality of posture state definitions is included in a set I and the respective therapy parameter value is a therapy parameter value P, and wherein the blended therapy parameter value is determine... | Medtronic, Inc. | A61N001/00 A61N001/365 A61B005/11 A61B005/00
US09345385 | B2 | 2016-05-24 | Subject observation apparatus and subject observation method | 2009-07-20 | A subject observation apparatus of the invention includes: a light-emitting section for emitting a light to a subject; an optical modulation section for detecting a scattering angle of a return light from the subject and performing optical modulation on the return light in accordance with the scattering angle; and a signal output section for generating a signal to show a state of light scattering in the subject based on the light subjected to the optical modulation by the optical modulation section, and outputting the generated signal. | 1. A subject observation apparatus comprising: a light-emitting section for emitting a light to a subject;a spectral distribution switching section for switching the light emitted from the light-emitted section to either a first light having a first spectral distribution or a second light having a second spectral distribution different from the first spectral distribution;an optical modulation section for detecting a scattering angle of a return light from the subject and performing optical modulation on the return light in accordance with the scattering angle, when the first light or the second light switched by the spectral distribution switching section is emitted to the subject; anda signal output section for generating a signal to show a state of light scattering in the subject based on the light subjected to the optical modulation by the optical modulation section, and outputting the generated signal,wherein the optical modulation section is configured as a plate-shaped optical element which includes at a center portion thereof a light-transmitting portion for transmitting a light when the scattering angle is equal to or larger than zero degree and equal to or smaller than a predetermined angle in a case where a forward-scattering direction of the return light is zero degree, and includes a light selective-transmitting portion around the light transmitting portion, the light selective-transmitting portion transmitting only lights of wavelength bands coincident with the first spectral distribution and shielding lights of wavelength bands other than the wavelength bands coincident with the first spectral distribution, when the scattering angle is larger than the predetermined angle and smaller than 90 degrees. 2. The subject observation apparatus according to claim 1, wherein the signal output section is an image pickup device for generating an image pickup signal by photoelectrically converting the light subjected to the optical modulation by the optical modulation section and outputting the generated image pickup signal. 3. The subject observation apparatus according to claim 1, wherein the subject is a scattering medium including a plurality of regions each having a different light scattering characteristic. 4. The subject observation apparatus according to claim 1, wherein wavelength bands of a maximum light intensity of the first spectral distribution do not overlap with wavelength bands of a maximum light intensity of the second spectral distribution. 5. The subject observation apparatus according to claim 4, wherein the first spectral distribution and the second spectral distribution have comb shapes, and wavelength bands of the parts corresponding to the comb shapes do not overlap each other. | OLYMPUS CORPORATION | A61B001/00 A61B001/04 A61B001/045 A61B005/00 G01N021/47 G02B023/24
US20160135775 | A1 | 2016-05-19 | System And Method For Time-Resolved, Three-Dimensional Angiography With Physiological Information | 2014-11-17 | A system and method are provided for generating time resolved series of angiographic volume data having velocity or velocity-derived information integrated therewith. An image processing system is configured to receive angiographic volume data and flow sensitive imaging data and process the angiographic volume data and flow sensitive imaging data to generate a combined dataset. A display is configured to display the combined dataset as angiographic volumes having associated time-resolved, color-coded flow information. | 1. A system for generating time resolved series of angiographic volume data having velocity or velocity-derived information integrated therewith, the system comprising: an image processing system configured to receive angiographic volume data and flow sensitive imaging data and process the angiographic volume data and flow sensitive imaging data to generate a combined dataset; anda display configured to display the combined dataset as angiographic volumes having associated time-resolved, color-coded flow information. 2. The system of claim 1 wherein the flow sensitive imaging data includes velocity encoded MRI data. 3. The system of claim 1 wherein the flow sensitive imaging data includes 3D Doppler ultrasound data. 4. The system of claim 1 wherein the angiographic volume data includes x-ray projection data. 5. The system of claim 1 wherein the x-ray projection data includes four-dimensional (4D) digital subtraction angiography data. 6. The system of claim 1 wherein the image processing system is configured to co-register the angiographic volume data and flow sensitive imaging data in 3D space. 7. The system of claim 6 wherein the image processing system is further configured to perform co-registration using the angiographic volume data as a constraining volume. 8. The system of claim 1 wherein the image processing system is further configured to generate the combined dataset by multiplying the angiographic volume data and the flow sensitive imaging data. 9. The system of claim 1 wherein the flow sensitive imaging data includes a series of flow-sensitive 3D volumes. 10. The system of claim 9 wherein the series flow-sensitive 3D volumes includes a measure of velocity in one of three (x,y,z) directions or net velocity. 11. The system of claim 1 wherein the time-resolved, color-coded flow information includes at least one of velocity measurements, pressure gradient information, wall shear stress, or flow streamline information. 12. The system of claim 1 wherein the image processing system is configured to generate the combined dataset by multiplying components of the flow-sensitive imaging data by the angiographic volume data without changing spatial information of the angiographic volume data or flow information of the flow-sensitive imaging data. 13. The method of claim 1 wherein the image processing system is configured to generate the combined dataset by spatially convolving the flow-sensitive volume data to increase a signal to noise ratio (SNR). 14. A method for generating time resolved series of angiographic volume data having velocity or velocity-derived information integrated therewith, the method comprising: generating a series of 3D time-resolved vascular volumes from time resolved x-ray projection data;generating a series of flow-sensitive 3D volumes from one of magnetic resonance imaging data or ultrasound data; andintegrating the series of 3D time-resolved vascular volumes and the series flow-sensitive 3D volumes to genera... |  | A61B006/00 A61B008/06 A61B005/00 G06T007/00 G06T005/50 G06T015/08 G06T011/00 A61B005/055 A61B008/08
US20160143596 | A1 | 2016-05-26 | ASSESSING PATIENT RISK OF AN ACUTE HYPOTENSIVE EPISODE WITH VITAL MEASUREMENTS | 2016-02-02 | What is disclosed is a system and method for assessing patient risk for an occurrence of an acute hypotensive episode within the timeframe of a prospective prediction window using multiple vitals to improve predictive accuracy. In one embodiment, the present method involves the following. A training set is retrieved from a database. The training set comprises vital measurements for a plurality of intensive care patients. Each vital measurement xt has been obtained at time t, where t=1, . . . , N, and N is the number of measurements obtained for that patient, with t being reckoned from a start of the vital measurement at pre-defined time intervals. The training set is used to train the present classifier system. The classifier classifies a yet unclassified patient into a first at risk for an acute hypotensive episode or into a second class not at risk for the occurrence of an acute hypotensive episode. | 1. A method for assessing patient risk for an acute hypotensive episode, the method comprising: receiving records containing vital measurements from a plurality of intensive care patients, each vital measurement xt having been obtained at time t, where t=1, . . . , N, and N is the number of measurements obtained for that patient, with t being reckoned from a start of the vital measurement at pre-defined time intervals;using the vital measurements as a training set to train a classifier system which classifies an unclassified patient into one of: a first class where the patient is identified as being at risk for an acute hypotensive episode occurring within a timeframe of a prediction window of w minutes in the future, and a second class where the patient is identified as not being at risk for an acute hypotensive episode;obtaining at least one vital measurement of an unclassified patient; andusing the trained classifier system to classify the unclassified patient into one of the first and second classes based on the unclassified patient's obtained vital measurement. 2. The method of claim 1, wherein an acute hypotensive episode is defined as an interval [xi; xi+30] in which at least 27 measurements are not greater than 60. 3. The method of claim 1, wherein the vital measurements are retrieved from a database containing physiological signals, patient vitals, and clinical data of patients in intensive care. 4. The method of claim 1, further comprising discarding records that contain less than 6 hours of vital measurements and discarding records of patients that have had an acute hypotensive event within the first 5 hours of recorded data. 5. The method of claim 1, wherein using only vital measurements in a range of 10-60 minutes immediately preceding a start of the prediction window for both first and second classes. 6. The method of claim 1, wherein the vital measurements comprises mean arterial pressure and any of: body temperature, blood pressure, heart rate, and respiratory rate. 7. The method of claim 6, wherein a first vector consisting of vital measurements of patients in the first class is y1 and a second vector consisting of vital measurements of patients in the second class is y2, and where μ1=mean(y1), μ2=mean(y2), σ1=sd(y1), σ2=sd(y2), k is a value such that (μ1+kσ1&lt;μ2−kσ2), and where n1 is a number of y1 values above μ2−kσ2, n2 is a number of y1 values within (μ1+kσ1, μ2−kσ2), n3 is a number of y2 values below μ1+kσ1, and n4 is a number of y2 values within (μ1+kσ1, μ2−kσ2). 8. The method of claim 7, wherein, in response to a mean of the unclassified patient's vital measurements averaged over at least a 1 hour time interval immediately preceding a start of the prediction window is less than (μ1+k0σ1), where k0=min(n1+n2+n3+n4), classifying the unclassified patient into the first class. 9. The method of claim 7, wherein, in response to a mean of the unclassified patient's vital measurements averaged over at least a 1 hour time... |  | A61B005/00 G06N099/00 G06F019/00 A61B005/0205
US20160135732 | A1 | 2016-05-19 | SYSTEMS AND METHODS FOR USING ISOLATED VOWEL SOUNDS FOR ASSESSMENT OF MILD TRAUMATIC BRAIN INJURY | 2016-01-25 | A system and method of identifying an impaired brain functionality such as a mild traumatic brain injury using speech analysis. In one example, recordings are taken on a device from athletes participating in a boxing tournament following each match. In one instance, vowel sounds are isolated from the recordings and acoustic features are extracted and used to train several one-class machine learning algorithms in order to predict whether an athlete is concussed. | 1. A method of identifying a mild traumatic brain injury comprising: using a sound recording device to capture spoken sound recording data from at least one individual at a first point in time to establish a spoken sound baseline;storing the spoken sound baseline in a data repository;capturing a spoken sound from a patient at a second point in time subsequent to the first point in time;comparing the spoken sound to the spoken sound baseline retrieved from the data repository; andusing the comparison of the spoken sound to the spoken sound baseline retrieved from the data repository to determine if the patient has experienced a mild traumatic brain injury between the first point in time and second point in time. 2. A method as recited in claim 1, wherein the captured spoken sound recording data is from a single individual. 3. A method as recited in claim 2, wherein the patient is the single individual. 4. A method as recited in claim 1, wherein the spoken sound baseline is a normalization of captured spoken sound recordings from a plurality of individuals. 5. A method as recited in claim 1, further comprising removing unwanted noise from at least one of the recorded spoken sound baseline or the captured spoken sound. 6. A method as recited in claim 1, further comprising isolating a speech segment from at least one of the recorded spoken sound baseline or the captured spoken sound. 7. A method as recited in claim 6, wherein isolated speech segment is a vowel sound. 8. A method as recited in claim 6, wherein isolating the speech segment further comprises identifying the onset of the speech segment via an onset detection routine. 9. A method as recited in claim 1, further comprising identifying a speech feature in at least one of the recorded spoken sound baseline or the captured spoken sound. 10. A method as recited in claim 9, wherein the speech feature is at least one of pitch, formant frequencies F1-F4, jitter, shimmer, mel-frequency cepstral coefficients, or harmonics-to-noise ratio. 11. A method as recited in claim 1, wherein the comparison of the spoken sound to the spoken sound baseline comprises a learning model with an associated learning algorithm. 12. A method as recited in claim 11, wherein the learning model analyzes the comparison data and recognizes patterns for assessment and regression analysis. 13. A method as recited in claim 11, wherein comparison of the spoken sound to the spoken sound baseline is performed via a support vector machine. 14. A non-transient, computer-readable media having stored thereon instructions for assisting a healthcare provider in identifying a mild traumatic brain injury, the instructions comprising: receiving from a sound recording device, spoken sound recording data from at least one individual at a first point in time to establish a spoken sound baseline;storing the spoken sound baseline in a data repository;receiving spoken sound from a patient at a second point in time subsequent to the ... |  | A61B005/00 A61B007/00
US20160128625 | A1 | 2016-05-12 | METHOD AND APPARATUS FOR ADVISING PHYSICAL CONDITION AND RECORDING MEDIUM USING THE METHOD | 2015-11-11 | A method and an apparatus for advising physical condition and a recording medium using the method are provided. In the method, multiple sensing devices are configured for a user to measure data related to the physical condition of the user. The data measured by the sensing devices is collected and integrated to estimate a current physical condition of the user. An advice is concluded at least based on the current physical condition of the user, and prompted for the user. | 1. A method for advising physical condition, adapted for an electronic apparatus, comprising: configuring a plurality of sensing devices for a user to measure the data related to the physical condition of the user;collecting the data measured by the sensing devices;integrating the collected data to estimate a current physical condition of the user;concluding an advice at least based on the current physical condition of the user; andprompting the advice. 2. The method according to claim 1, wherein after the step of collecting data measured by each of the sensing devices, the method further comprises: uploading the collected data to a server comprising data collected from the other users to compare the collected data to the data from the other users according to a user information of the user and conclude the advice according to the collected data from the other users and the user information of the user; andreceiving the advice from the server for prompting. 3. The method according to claim 1, wherein the step of integrating the collected data to estimate the current physical condition of the user comprises: comparing the collected data with physiological data recorded in a history of the user; andestimating the current physical condition according to a comparison result of the collected data and the physiological data. 4. The method according to claim 1, wherein after the step of collecting the data measured by the sensing devices, the method further comprises: determining an exercise activity currently exercised by the user based on the collected data; andconcluding an exercise advice based on the estimated current physical condition and the determined exercise activity of the user. 5. The method according to claim 1, wherein the step of concluding the advice at least based on the current physical condition of the user further comprises: concluding the advice based on an environmental data around the user. 6. The method according to claim 1, wherein the sensing devices comprise a scale and the step of configuring the plurality of sensing devices for the user to measure the data of the user comprises: receiving a plurality of user information, wherein each user information including an electrocardiogram (ECG) of a corresponding user;detecting a first user steps on the scale;in response to detecting the first user steps on the scale, measuring the ECG of the first user by using electrodes disposed in the scale;comparing the measured ECG of the first user with the received user information and recognizing an identification of the first user at least based on a plurality of features analyzed from the ECG; andactivating the scale to measure data of the first user in response to the recognized identification. 7. The method according to claim 6, wherein the step of measuring data of the first user comprises: determining whether the data of first user measured by the scale comprises a body fat percentage data according to the received user informa... |  | A61B005/00 A61B005/0205 A61B005/117 A61B005/0402 A61B005/04
US20160128576 | A1 | 2016-05-12 | Needle Device and Needle System Having the Same | 2015-10-29 | A needle device includes a shell, a needle unit and an actuator. The needle unit includes an outer needle guide and a needle member. The outer needle guide has a piercing end for penetrating into a biological body part. The needle member is disposed in the shell and the outer needle guide, has a needle tip protruding from a front end of the shell and is movable relative to the shell and the outer needle guide. The actuator is responsive to a control signal to drive the needle member and move the needle tip between a protruding position, where the needle tip protrudes out of the piercing end of the outer needle guide, and a retracting position, where the needle tip is retracted back into the outer needle guide. | 1. A needle system capable of identifying tissue type of a tissue layer, the needle system comprising: a needle device for penetrating into a biological body part, the needle device including a needle member that extends and is movable along a needle axis and that has a needle tip, andan electrically operable actuator that is responsive to a control signal to drive the needle member to move along the needle axis against a resistance force from a tissue layer being penetrated by the needle device; anda control module including a displacement detecting unit that is operable to detect displacement of the needle member along the needle axis, anda control unit that is electrically coupled to the actuator and the displacement detecting unit, that is operable to output the control signal to the actuator, and that is operable to identify the tissue type of the tissue layer based on the control signal and the displacement of the needle member detected by the displacement detecting unit. 2. The needle system according to claim 1, wherein, the needle device further includes an outer needle guide that extends along the needle axis to surround the needle member and that has s a piercing end configured to penetrate into the biological body part, the needle member being movable relative to the outer needle guide between a protruding position, where the needle tip of the needle member protrudes out of the piercing end of the outer needle guide, and a retracting position, where the needle tip of the needle member is retracted back into the outer needle guide. 3. The needle system according to claim 2, wherein, when the needle member is driven by the actuator to move the needle tip toward the protruding position, the needle tip protrudes out of the piercing end of the outer needle guide by a protruding distance that corresponds to the displacement of the needle member detected by the displacement detecting unit, the control unit being operable to compare the detected displacement of the needle member with a predetermined distance, so as to identify the tissue type of the tissue layer based on the comparison and the control signal. 4. The needle system according to claim 3, further comprising a monitoring unit that is communicatively coupled to the control unit, the control unit being operable to calculate a resistance force value of the tissue layer based on the control signal and the comparison and to transmit the resistance force value to the monitoring unit for display. 5. The needle system according to claim 3, wherein the control unit is operable to adjust the control signal based on the displacement of the needle member, and to enable the actuator to drive the needle member to move the needle tip back to the protruding position. 6. The needle system according to claim 2, wherein the actuator drives the needle member to reciprocate between the protruding position and the retracting position. 7. The needle system according to claim 2, wherein the needle ... |  | A61B005/00 A61M019/00 A61B017/34
US09517018 | B2 | 2016-12-13 | Methods and systems for power optimization in a medical device | 2016-01-25 | A physiological monitoring system may use photonic signals to determine physiological parameters. The system may vary parameters of a light drive signal used to generate the photonic signal from a light source such that power consumption is reduced or optimized. Parameters may include light intensity, firing rate, duty cycle, other suitable parameters, or any combination thereof. In some embodiments, the system may use information from a first light source to generate a light drive signal for a second light source. In some embodiments, the system may vary parameters in a way substantially synchronous with physiological pulses, for example, cardiac pulses. In some embodiments, the system may vary parameters in response to an external trigger. | 1. A method of operating a pulse oximeter, the method comprising: receiving, using the pulse oximeter, a photoplethysmograph (PPG) signal from a pulse oximetry sensor, which comprises a light source and a detector;determining, using the pulse oximeter, pulse period information related to physiological pulses in the PPG signal;determining, using the pulse oximeter, at least one parameter of a light drive signal based on the pulse period information, wherein the at least one parameter is determined to vary substantially synchronously with the physiological pulses, and wherein the determined at least one parameter comprises at least one of light intensity, light source firing rate, and duty cycle; andgenerating, using the pulse oximeter, the light drive signal for activating the light source of the pulse oximetry sensor to emit a photonic signal according to the determined at least one parameter. 2. The method of claim 1, wherein the determined at least one parameter comprises the light intensity. 3. The method of claim 2, wherein the light intensity comprises a high value and a low value, and wherein the light intensity is varied so that the high value coincides with at least one of a systole period, a diastole period, a dicrotic notch, a peak, and a trough of the physiological pulses. 4. The method of claim 3, wherein the light intensity is varied according to a periodic waveform, wherein the periodic waveform comprises a plurality of peaks, and wherein the peaks are configured to coincide with at least one of a systole period, a diastole period, a dicrotic notch, a peak, and a trough of the physiological pulses. 5. The method of claim 1, wherein the determined at least one parameter comprises the light source firing rate. 6. The method of claim 5, wherein the light source firing rate comprises a high value and a low value, and wherein the light source firing rate is varied so that the high value coincides with at least one of a systole period, a diastole period, a dicrotic notch, a peak, and a trough of the physiological pulses. 7. The method of claim 6, wherein the light source firing rate is varied according to a periodic waveform, wherein the periodic waveform comprises a plurality peaks, and wherein the peaks are configured to coincide with at least one of a systole period, a diastole period, a dicrotic notch, a peak, and a trough of the physiological pulses. 8. The method of claim 1, further comprising: determining, using the pulse oximeter, a sampling rate of an analog-to-digital converter of the pulse oximeter based on the pulse period information, wherein the sampling rate is determined to vary substantially synchronously with the physiological pulses in the PPG signal; andsampling, using the analog-to-digital converter of the pulse oximeter, the PPG signal according to the determined sampling rate. 9. The method of claim 1, wherein the determined at least one parameter comprises the duty cycle. 10. The method of claim 9, wherein... | COVIDIEN LP | A61B005/00 A61B005/0205 A61B005/024 A61B005/0402 A61B005/1455 A61B005/0295 A61B005/026
US09770213 | B2 | 2017-09-26 | Device, system and method for extracting physiological information | 2015-08-20 | The present invention relates to a device, system and a method for extracting physiological information indicative of at least one vital sign of a subject from detected electromagnetic radiation reflected from the subject. The device comprises an input interface for receiving a data stream of detection data derived from detected electromagnetic radiation reflected from a skin region of a subject, wherein the detected electromagnetic radiation is detected by a polarized radiation detector, while the polarization angle of the polarized radiation detector is changed, a PPG extraction unit for extracting a photoplethysmographic, PPG, signal from said detection data, a signal quality determination unit for determining quality metrics from said PPG signal for different settings of the polarization angle of the polarized radiation detector, a selection unit for selecting the optimum quality metric value from the determined quality metrics and for generating polarization control information for use by said polarized radiation detector for setting the polarization angle to an angle value, at which said optimum quality metric value was obtained, for subsequent detection of radiation, and a processor for deriving physiological information indicative of at least one vital sign from the PPG signal. | 1. A device for extracting physiological information indicative of at least one vital sign of a subject from detected electromagnetic radiation reflected from the subject, said device comprising: an input interface for receiving a data stream of detection data from a polarized radiation detector for detecting electromagnetic radiation reflected from a skin region of a subject, wherein the detected electromagnetic radiation is detected while a polarization angle of the polarized radiation detector is changed,a PPG extraction unit for extracting a photoplethysmographic, PPG, signal from said detection data,a signal quality determination unit for determining quality metrics from said PPG signal for different settings of the polarization angle of the polarized radiation detector,a selection unit for selecting the optimum quality metric value from the determined quality metrics and for generating polarization control information for use by said polarized radiation detector for setting the polarization angle to an angle value, at which said optimum quality metric value was obtained, for subsequent detection of radiation, anda processor for deriving physiological information indicative of at least one vital sign from the PPG signal. 2. The device as claimed in claim 1, wherein said PPG extraction unit is configured to extract two or more PPG signals from detection data derived from detected electromagnetic radiation reflected from two or more different skin regions of the subject,wherein said signal quality determination unit is configured to determine quality metrics from each of said two or more PPG signals, andwherein said selection unit is configured to separately select the optimum quality metric value from the calculated quality metrics for each of said two or more PPG signals and to generate a polarization control information for use by said polarized radiation detector for setting the polarization angle to angle values, at which said optimum quality metric value was obtained for the respective PPG signal, for subsequent detection of radiation from the respective skin region. 3. The device as claimed in claim 1, wherein said PPG extraction unit, said signal quality determination unit and said selection unit are configured to regularly, continuously or from time to time perform their functions to regularly, continuously or from time to time adjust the polarization angle setting. 4. The device as claimed in claim 1, wherein said PPG extraction unit is configured to extract two or more PPG signals from different frequency components of said detection data,wherein said signal quality determination unit is configured to determine quality metrics from each of said two or more PPG signals, andwherein said selection unit is configured to separately select the optimum quality metric value from the calculated quality metrics for each of said two or more PPG signals and to generate a polarization control information for use by said polarized radiation de... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/024 A61B005/1455
US09332920 | B2 | 2016-05-10 | Rotor identification using sequential pattern matching | 2013-12-19 | An anatomical mapping system includes a plurality of mapping electrodes each having an electrode location and configured to detect activation signals of intrinsic physiological activity within an anatomical structure. A mapping processor is associated with the plurality of mapping electrodes and is configured to record the detected activation signals and associate one of the plurality of mapping electrodes with each recorded activation signal. The mapping processor is further configured to analyze the recorded activation signals to identify at least one recurring pattern based on a relationship between a timing of the detected activation signals and the electrode locations of the mapping electrode associated with each detected activation signal. | 1. An anatomical mapping system comprising: a plurality of mapping electrodes configured to detect activation signals of intrinsic physiological activity within an anatomical structure, each of the plurality of mapping electrodes having an electrode location;a mapping processor associated with the plurality of mapping electrodes, the mapping processor configured to record the detected activation signals and associate one of the plurality of mapping electrodes with each recorded activation signal, the mapping processor further configured to analyze the recorded activation signals to identify at least one recurring pattern of the detected activation signals sensed by mapping electrodes at different electrode locations based on a relationship between a timing of the detected activation signals and the electrode locations of the mapping electrode associated with each detected activation signal. 2. The anatomical mapping system according to claim 1, wherein the mapping processor is further configured to determine a location of the at least one recurring pattern with respect to the anatomical structure. 3. The anatomical mapping system according to claim 1, wherein the mapping processor is further configured to correlate the at least one recurring pattern to at least one pathology template to determine whether the at least one recurring pattern is indicative of a pathology. 4. The anatomical mapping system according to claim 3, wherein the mapping processor is configured to correlate the at least one recurring pattern to the at least one pathology template based on activation signals detected by a subset of the plurality of mapping electrodes. 5. The anatomical mapping system according to claim 1, wherein the recurring pattern is a recurring sequential pattern and the mapping processor is configured to run a sequential mining algorithm on the detected activation signals to identify at least one recurring pattern. 6. The anatomical mapping system according to claim 1, wherein the recurring pattern is a recurring association pattern and the mapping processor is configured to run a location association mining algorithm on the detected activation signals to identify at least one recurring pattern. 7. A catheter system comprising: a plurality of mapping electrodes configured to detect activation signals of intrinsic cardiac activity, each of the plurality of mapping electrodes having an electrode location;a mapping processor associated with the plurality of mapping electrodes, the mapping processor configured to record the detected activation signals and associate one of the plurality of mapping electrodes with each recorded activation signal, the mapping processor further configured to analyze the recorded activation signals to identify at least one recurring pattern of the detected activation signals sensed by mapping electrodes at different electrode locations based on a relationship between a timing of the detected activation signals and the elect... | Boston Scientific Scimed Inc. | A61B005/04 A61B005/042 A61B005/00 A61B005/0464 G06K009/00 A61B018/14 A61B018/00
US10265100 | B2 | 2019-04-23 | Transradial access devices, systems, and methods | 2016-01-20 | Transradial access devices, systems, and methods. An exemplary device of the present disclosure for providing access to a vessel or other luminal organ and to prevent overperforation of the same comprises a needle having a needle tip and at least one current element positioned thereon or embedded therein at or near the needle tip, a sheath positioned around at least part of the needle, and a protractor in communication with the needle, whereby engagement of the protractor causes the needle tip to protract from the sheath through the aperture, and whereby release of the protractor causes the needle tip to retract through the aperture into the sheath. Activation of the at least one current element can facilitation vasodilation of the vessel or other luminal organ. | 1. A device for providing access to a vessel or other luminal organ and to prevent overperforation of the same, the device comprising: a needle having a needle tip and at least one current element positioned thereon or embedded therein at or near the needle tip;a sheath positioned around at least part of the needle, the sheath defining an aperture at a distal sheath end, an interior space around the needle positioned therein, and having a blunt non-piercing tapered distal end configured for percutaneous insertion into a vessel or other luminal organ, wherein at least a portion of a wall of the sheath is comprised of translucent or transparent material; anda protractor in communication with the needle, whereby engagement of the protractor causes the needle tip to protract from the sheath through the aperture, and whereby release of the protractor causes the needle tip to retract through the aperture into the sheath,wherein the device is configured so that when the needle tip punctures the vessel or other luminal organ underneath skin after puncturing the skin and is advanced along with the sheath so that the distal sheath end is also positioned within the vessel or other luminal organ and so that fluid within the vessel or other luminal organ is in contact with the needle tip, activation of the at least one current element can facilitation vasodilation of the vessel or other luminal organ and the fluid can enter the interior space of the sheath through the aperture and be visually detected within the sheath through the portion of the sheath wall comprising the translucent or transparent material to indicate that the distal sheath end is positioned within the vessel or other luminal organ so to prevent overperforation of the vessel or other luminal organ. 2. The device of claim 1, wherein the at least one current element comprises an excitation electrode configured to generate an electric current. 3. The device of claim 1, wherein the protractor is selected from the group consisting of a button and a rotatable dial. 4. The device of claim 1, wherein the tapered distal end of the sheath is configured to allow blood or other bodily fluids to pass between an outside wall of the needle and the aperture when the needle is protracted from the sheath. 5. The device of claim 1, further comprising: a needle mechanism in communication with the needle and the protractor, the needle mechanism operable to facilitate protraction and retraction of the needle tip from the sheath. 6. The device of claim 5, wherein the needle mechanism is selected from the group consisting of a spring, a compliant polymer, a coil, a sponge, and a bladder. 7. The device of claim 1, further comprising: one or more wire apertures defined therethrough, the one or more wire apertures configured to receive a wire. 8. The device of claim 7, further comprising: a wire positioned through the one or more wire apertures, the wire configured to permit a second device to be delivered over... | DTherapeutics, LLC | A61B017/34 A61B005/053 A61B018/14 A61B005/1473 A61M025/09 A61B005/145 A61B017/00
US09357150 | B2 | 2016-05-31 | Image sensor with integrated power conservation control | 2013-12-03 | An integrated image sensor circuit with multiple power modes is disclosed. The integrated circuit comprises a pixel array, an analog block to process analog signal associated with the pixel array, where the analog block comprises an analog to digital convertor (ADC), and a first control circuit to enable/disable the analog block or to configure the analog block to a high/low-power mode depending on whether the pixel array is in a readout frame or in a reset frame with no active readout. The integrated image sensor circuit may further comprise a post-processing block and a second control circuit to enable/disable the post-processing block or to configure the post-processing block to the high-power mode or the low-power mode depending on whether the pixel array is in the readout frame or in the reset frame with no active readout. | 1. An integrated image sensor circuit, comprising: a pixel array to capture an image projected thereon;at least one analog block to process analog signal associated with the pixel array, wherein said at least one analog block comprises an analog to digital convertor (ADC); anda first control circuit to enable or disable said at least one analog block depending on whether the pixel array is in a readout frame, or in a reset frame with no active readout, or to configure said at least one analog block to a high-power mode or a low-power mode depending on whether the pixel array is in the readout frame, or in the reset frame with no active readout, wherein the readout frame is substantially the same as a first period from a first-row readout pulse to a last-row readout pulse, and the reset frame is substantially the same as a second period from a first-row reset pulse to a last-row reset pulse. 2. The integrated image sensor circuit of claim 1, wherein said at least one analog block is enabled or is configured to the high-power mode when the pixel array is in the readout frame, and wherein said at least one analog block is disabled or is configured to the low-power mode when the pixel array is in the reset frame. 3. The integrated image sensor circuit of claim 2, wherein said at least one analog block is powered up at a short period before the readout frame starts, and wherein the short period depends on a settling time associated with said at least one analog block. 4. The integrated image sensor circuit of claim 3, wherein the short period has a first range as small as one order of microsecond to a second range as large as one order of millisecond. 5. The integrated image sensor circuit of claim 1, wherein said at least one analog block further comprises one or more additional analog blocks selected from an analog group consisting of a bias circuit, a reference circuit and a gain amplifier. 6. The integrated image sensor circuit of claim 1, wherein the reset frame corresponds to a first period of time starting from a global reset to immediately before a first-row readout for the integrated image sensor circuit operated in a global shutter mode and wherein the reset frame corresponds to a second period of time starting from a first-row reset to immediately before the first-row readout for the integrated image sensor circuit operated in a rolling shutter mode. 7. The integrated image sensor circuit of claim 1, wherein said at least one analog block is powered down once the readout frame is completed. 8. The integrated image sensor circuit of claim 1, further comprising at least one post-processing block, and a second control circuit to enable or disable said at least one post-processing block or to configure said at least one post-processing block to the high-power mode or the low-power mode depending on whether the pixel array is in the readout frame or in the reset frame. 9. The integrated image sensor circuit of claim 8, wherein said at le... | Capso Vision Inc. | H04N005/378 H04N005/357 A61B005/00 A61B001/00 H04N005/232
US09339202 | B2 | 2016-05-17 | System for processing physiological data | 2013-06-28 | Physiological signals such as ECG signals are obtained from a patient. In accordance with one or more embodiments, an apparatus, system and/or method is directed to at least two ECG sensing electrodes that adhere to remote locations on a patient and sense ECG signals from the patient. An amplifier circuit amplifies the sensed ECG signals to provide amplified ECG signals, and a digitizing circuit digitizes the amplified ECG. A computing circuit processes the digitized ECG signals (e.g., by removing noise). A battery powers the aforesaid circuits, and a housing houses the battery and aforesaid circuit. A fastener mechanically fastens and electrically couples the housing to one of the electrodes, and the other electrodes are coupled to the amplifier via a flexible insulated lead wire. | 1. An apparatus comprising: at least two ECG sensing electrodes configured and arranged to adhere to remote locations on a patient and to sense ECG signals from the patient;a digitizing circuit configured and arranged to digitize the ECG signals to provide digitized ECG signals;a computing circuit configured and arranged to process the digitized ECG signals by reducing in-band noise as measured by the ANSI/AAMI EC57:1998 standard by at least 20 dB with a Quality of Signal Reconstruction of >95%;a housing that houses the digitizing circuit and the computing circuit;a fastener configured and arranged to mechanically fasten the housing to a first one of the ECG sensing electrodes and to electrically couple the first one of the ECG sensing electrodes to the digitizing circuit; andfor each additional one of the at least two ECG sensing electrodes that is in addition to the first one of the ECG sensing electrodes, a flexible insulated lead wire configured and arranged to couple the additional ECG sensing electrode to the digitizing circuit. 2. The apparatus of claim 1, wherein the fastener is configured and arranged to mechanically connect the housing to a first one of the ECG sensing electrodes by affixing the housing to the first one of the ECG sensing electrodes, wherein the computing circuit is configured and arranged to remove noise from the digitized ECG signals while maintaining morphology of the digitized ECG signal, therein providing denoised ECG signals, andprocess the denoised ECG signals by carrying out a function selected from the group consisting of: detecting an R-R interval, detecting a Q-T interval, detecting a QRS complex, and a combination thereof. 3. The apparatus of claim 1, further including a battery configured and arranged to power the digitizing and computing circuits, wherein the housing also houses the battery, and wherein the first one of the ECG sensing electrodes is configured and arranged to adhere to the patient and to support the entire weight of the housing via the fastener. 4. The apparatus of claim 1, further including an adhering component configured and arranged to adhere and fully support the first one of the ECG sensing electrodes and the housing to the skin of a patient. 5. The apparatus of claim 1, wherein each ECG signal is sampled at 300 Hz or less and the computing circuit is configured and arranged to reduce the in-band noise using an average power consumption that is less than 500 micro-Watts per ECG signal. 6. The apparatus of claim 1 wherein the housing has an average thickness of less than 12 mm and a volume of less than 18 cc. 7. The apparatus of claim 1, wherein the housing includes the first one of the electrodes. 8. The apparatus of claim 1, wherein the housing includes the fastener. 9. The apparatus of claim 1 wherein the flexible insulating lead wires are >10 cm in length. 10. The apparatus of claim 1 wherein the flexible insulating lead wires are >15 cm in length. 11. The apparatus of c... | VivaQuant LLC | A61B005/00 A61B005/0402 A61B005/0205 A61B005/0456
US20160143576 | A1 | 2016-05-26 | MRI IMAGE FUSION METHODS AND USES THEREOF | 2014-07-14 | A method for fusing a pre-operative MRI prostate image to an intra-operative TRUS or CT prostate image according to a least-cost affine transformation of the MRI contour onto the TRUS or CT contour, with smooth non-linear warping adjustment. MRI and CT processing may be performed as a pre-operational procedure for increased efficiency, while TRUS may be performed concurrent with a surgical procedure. | 1. A method for generating a Trans-Rectal Ultra-Sound (TRUS)-Magnetic Resonance Imaging (MRI) fusion image of a prostate gland of a subject, said method comprising the following steps: (a) inputting an MRI scan of the prostate gland of said subject;(b) segmenting the MRI scan to produce at least one segmented MRI contour surface of the organ, said contour comprising a plurality of three-dimensional (3D) landmark points;(c) inputting a TRUS scan of the prostate gland of said subject;(d) segmenting the TRUS scan to produce at least one segmented TRUS contour surface of the prostate gland, said contour comprising a plurality of 3D landmark points, wherein the at least one MRI contour surface and the at least one TRUS contour surface correspond to the same anatomical surface;(e) resampling the TRUS and MRI contours to a common geometric space;(f) computing a linear transformation that maps the MRI contour surface onto the TRUS contour surface, the linear transformation being an affine transformation estimated by minimization of the matching cost between the plurality of landmark points on the MRI contour and the plurality of landmark points on the TRUS contour;(g) applying said linear transformation to the MRI contour points to obtain linearly transformed (LT) MRI contour points;(h) computing a local shape descriptor for each LT landmark point of the MRI contour surface and each landmark point of the TRUS contour surface;(i) computing an optimal assignment between said LT landmark MRI and TRUS contour surface points that minimizes a matching cost criterion between the shape descriptors of the matched points, said optimal assignment defining a sparse vector field mapping MRI contour points onto TRUS contour points;(j) computing a dense deformation field by smooth interpolation of said sparse vector field to map any point of the whole MRI volume onto a point of the TRUS volume;and(k) applying the linear and non-linear mapping of steps (f) through (j) to map points of the MRI image into the TRUS image;wherein the performance of steps (a) through (k) generates a Trans-Rectal Ultra-Sound-Magnetic Resonance Imaging fusion image of the prostate of said subject. 2. The method of claim 1, wherein said at least one contour surface comprises an external surface of the prostrate or a portion thereof, a contour of an internal surface of the prostate or a portion thereof, a contour of a transitional zone of the prostate or a portion thereof, a contour of a central zone of the prostate or a portion thereof, a contour of a peripheral zone of the prostate or a portion thereof, a contour of an interface between a central zone and a peripheral zone of the prostate or a portion thereof, a contour of a surface bordering the prostate and the urethra or a portion thereof, a contour based on observable calcifications, or any combinations thereof, or any combination thereof. 3. (canceled) 4. The method of claim 1, wherein said minimizes a matching cost criterion of step (... |  | A61B005/00 A61B008/08 G01R033/48 A61B008/12
US09330523 | B2 | 2016-05-03 | Wagering game machine with biofeedback-aware game presentation | 2012-09-14 | A computerized wagering game system includes a gaming module comprising gaming code which is operable when executed on to conduct a wagering game on which monetary value can be wagered, and a biofeedback module operable to track at least one biometric characteristic of a game player. The wagering game is further operable to alter presentation of the wagering game based on changes in the at least one biometric characteristic. A mood enhancement module is operable to provide energy to a wagering game player at a frequency designed to provide a mood enhancing effect. | 1. A method of operating a gaming system, the gaming system including one or more processors and an electronic gaming machine for conducting a wagering game, the electronic gaming machine having a gaming cabinet with one or more electronic input devices and one or more electronic display devices coupled to the gaming cabinet, the method comprising: detecting, by at least one of the one or more electronic input devices of the electronic gaming machine, a physical item associated with a monetary value, the monetary value establishing a credit balance that changes based on play of the wagering game;receiving, by at least one of the one or more electronic input devices of the electronic gaming machine, a wager input from a player to play the wagering game, the wager input being covered by the credit balance;initiating, by at least one of the one or more processors, the wagering game in response to the wager input from the player;determining, by at least one of the one or more processors of the gaming system, an outcome of the wagering game based, at least in part, on one or more random elements generated by at least one of one or more electronic random element generators;collecting, via one or more biometric sensors operatively coupled to the electronic gaming machine, biofeedback data to determine a physiological state of the player during play of the wagering game;loading a baseline biometric history of the player from stored historical data;providing, by one or more electronic output devices of the electronic gaming machine, a selected frequency of energy to the player during play of the wagering game based on the loaded baseline biometric history;evaluating, by the one or more processors, the collected biofeedback data to determine a new physiological state of the player;adjusting the frequency of energy based on the evaluation; andreceiving, by at least one of the one or more electronic input devices, a cashout input from the player to initiate a payout from the credit balance. 2. The method of claim 1, wherein evaluating the collected biofeedback data to determine the physiological state of the player comprises evaluating player body temperature, player blood pressure, player pupil dilation, player blood flow, or player pulse rate, or any combination thereof, measured by the one or more biometric sensors. 3. The method of claim 1, wherein providing a selected frequency of energy to the player comprises: providing a selected frequency of light energy to the player via an electronic display device of the electronic gaming machine, wherein the selected frequency of light energy is provided by altering a refresh rate of the electronic display device. 4. The method of claim 1, wherein the selected frequency of energy comprises a selected frequency of sound energy provided by a speaker of the electronic gaming machine to supplement the vocal frequencies. 5. The method of claim 1, further comprising varying the frequency over time. 6. The method ... | Bally Gaming, Inc. | A63F009/24 A63F013/00 G06F017/00 G06F019/00 G09B019/00 G07F017/32 A61B005/16 A61B005/00
US09332932 | B2 | 2016-05-10 | Blood collecting puncture device and magazine used for the same | 2009-03-18 | Provided are a blood collecting puncture device having high safety which can hygienically maintain the inside of a magazine and prevent infection, and the magazine. A blood collecting puncture device has a blood collecting puncture device body removably incorporating a magazine for housing lancets for puncturing human skin. The blood collecting puncture device body has a first housing section for housing not-yet-used lancets and a second housing section for housing used lancets to be disposed of. When a rotor is rotated, individual lancets are sequentially conveyed on a one-by-one basis to a stock position, a puncture position, and a disposal position, and this causes a not-yet-used lancet to be housed in the first housing section and also causes a used lancet to be housed in the second housing section, so that the lancets are separately housed in spaces isolated from each other. | 1. A blood sampling puncturing device comprising: a magazine storing a plurality of lancets each having a needle to puncture skin; anda rotation rotor that is located outside of the magazine and configured to hold at least one lancet of the plurality of lancets, the magazine being removably mounted in the blood sampling puncturing device, whereinthe magazine has a first storing section that is configured to store lancets before use therein, a second storing section that is configured to store lancets after use therein, and a partition plate that closes the second storing section, the first storing section defining a space therein and the second storing section defining a space therein, the space of the first storing section and the space of the second storing section being physically separated from one another by the partition plate, and the partition plate being configured to open the second storing section upon storing the at least one lancet in the second storing section, andthe rotation rotor has at least one lancet holding hole provided at a circumferential surface of the rotation rotor, the lancet holding hole being configured to remove at least one lancet of the plurality of lancets stored in the space of the first storing section in the magazine, to an outside of the magazine, by rotation, to hold the removed lancet in the lancet holding hole and to further rotate the lancet held in the lancet holding hole into a puncturing position for a puncturing operation, and when the puncturing operation has been performed the rotation rotor is configured to further rotate the lancet held in the lancet holding hole to the second storing section in the magazine to release the lancet into the space of the second storing section. 2. The blood sampling puncturing device according to claim 1, further comprising one or more holding positions to hold the at least one lancet removed from the first storing section of the magazine. 3. The blood sampling puncturing device according to claim 2, wherein the holding positions include a stocking position to prepare and to hold the at least one lancet removed from the first storing section prior to the puncturing operation, the puncturing position for the puncturing operation, and a discarding position to hold the at least one lancet in the second storing section after the puncturing operation is performed. 4. The blood sampling puncturing device according to claim 3, wherein the rotation rotor is configured to convey the at least one lancet to the stocking position, the puncturing position and the discarding position in sequence. 5. The blood sampling puncturing device according to claim 4, further comprising: a needle feeding lever configured to separate the at least one lancet from the plurality of lancets stored in the first storing section and to hold the at least one lancet in the stocking position; anda stopper to lock the needle feeding lever. 6. The blood sampling puncturing device according to claim ... | PANASONIC HEALTHCARE HOLDINGS CO., LTD. | A61M005/32 A61B005/15 A61B005/151
US09351698 | B2 | 2016-05-31 | Vascular data processing and image registration systems, methods, and apparatuses | 2013-03-12 | In part, the invention relates to processing, tracking and registering angiography images and elements in such images relative to images from an intravascular imaging modality such as, for example, optical coherence tomography (OCT). Registration between such imaging modalities is facilitated by tracking of a marker of the intravascular imaging probe performed on the angiography images obtained during a pullback. Further, detecting and tracking vessel centerlines is used to perform a continuous registration between OCT and angiography images in one embodiment. | 1. A processor-based method of displaying an angiographic and an intravascular representation of a blood vessel comprising: generating a set of optical coherence tomography image data in response to distance measurements of the blood vessel obtained during a pullback of a probe through the blood vessel using an optical coherence tomography system, the set of optical coherence tomography image data comprising a plurality of cross-sectional image at a plurality of positions along the blood vessel;generating a set of angiography image data using an angiography system during the pullback of the probe through the blood vessel using an optical coherence tomography system, the set of angiography image data comprising a plurality of two-dimensional images obtained at different points in time during the pullback;displaying, on a graphical user interface, a first panel comprising a first longitudinal view of the blood vessel generated using the optical coherence tomography image data; anddisplaying, on a graphical user interface, a second panel comprising a frame of the angiography image data identifying the blood vessel using one or more points in the frame and a vessel centerline passing through the one or more points. 2. The method of claim 1, further comprising co-registering the optical coherence tomography image data and the angiography data using vessel centerlines to create a continuous registration of a tracked marker, wherein the tracked marker is disposed on an optical coherence tomography data collection probe. 3. The method of claim 1 further comprising co-registering the optical coherence tomography image data and the angiography data such that selecting a point along the vessel centerline through a user interface changes a frame identifier in the first longitudinal view. 4. The method of claim 2 further comprising using pullback speed or pullback length to perform an iterative search to reject candidates for the tracked marker based on the possible locations for such markers based upon the pullback length and/or pullback speed. 5. The method of claim 1 wherein the vessel centerline is generated using a shortest path technique and a plurality of processing steps from a Dijkstra algorithm. 6. The method of claim 1 further comprising the step of removing a guide catheter image from one or more frames of angiography data using superposition of an intensity profile. 7. The method of claim 1 wherein the vessel centerline is generated using path information generated from one or more angiography frames in the absence of contrast solution. 8. The method of claim 1 further comprising generating a confidence score for each detection and co-registration between angiography data and optical coherence tomography data. | LIGHTLAB IMAGING, INC. | G06K009/00 A61B006/00 A61B006/12 A61B008/12 A61B005/00 A61B006/03 A61B005/055 A61M025/10 A61B019/00 A61B008/00
US09327123 | B2 | 2016-05-03 | Endovascular nerve monitoring devices and associated systems and methods | 2012-11-06 | Endovascular nerve monitoring devices and associated systems and methods are disclosed herein. A nerve monitoring system configured in accordance with a particular embodiment of the present technology can include a shaft having a proximal portion and a distal portion and a nerve monitoring assembly at the distal portion. The shaft is configured to locate the distal portion intravascularly at a treatment site. The nerve monitoring assembly can include a bipolar stimulation electrode array and a bipolar recording electrode array disposed distal to the bipolar stimulation electrode assembly. | 1. A method of monitoring renal nerve activity, the method comprising: deploying a nerve monitoring assembly in a renal artery of a human patient, wherein the nerve monitoring assembly comprises a first electrode array at a distal portion of a shaft and a second electrode array spaced laterally apart from the first electrode array, wherein deploying the nerve monitoring assembly in the renal artery further comprises inflating a first balloon within the renal artery, wherein the inflated first balloon contacts an inner wall of the renal artery, and wherein the first and second electrode arrays are on an outer surface of the first balloon;delivering a neuromodulation assembly to a target site within the renal artery between the first and second electrode arrays, wherein delivering the neuromodulation assembly to the target site within the renal artery between the first and second electrode arrays further comprises inflating a second balloon within the first balloon, wherein the second balloon is configured to contact the first balloon to define a neuromodulation region;stimulating renal nerves with the first electrode array, wherein the stimulation energy is at a magnitude which induces renal nerve activity;recording, with the second electrode array, renal nerve activity resulting from the stimulating act; andapplying therapeutically-effective neuromodulation energy to the target site with the neuromodulation assembly, wherein the neuromodulation energy is at a level which inhibits, reduces, or blocks renal nerves, and wherein applying therapeutically-effective neuromodulation energy to the target site with the neuromodulation assembly further comprises applying cryogenic cooling to the renal artery proximate the neuromodulation region. 2. The method of claim 1 wherein: stimulating renal nerves with the first electrode array comprises delivering a plurality of stimulus pulses with the first electrode array, each stimulus pulse having an intensity of about 20-60 mA at a pulse length of about 100-400 μs;recording renal nerve activity with the second electrode array comprises recording electrograms of individual electrodes of the second electrode array, wherein the electrograms correspond to the nerve activity resulting from the stimulus pulses; andwherein the method further comprises detecting nerve locations proximate the second electrode array, wherein the second electrode array is proximal to the first electrode array. 3. The method of claim 1 wherein stimulating renal nerves with the first electrode array comprises delivering a current pulse having an intensity of at least 20 mA. 4. The method of claim 1 wherein deploying the nerve monitoring assembly in the renal artery comprises deploying the first electrode array proximal to the second electrode array. 5. The method of claim 1 wherein: stimulating renal nerves and recording the resultant nerve activity is performed after applying a first cycle of therapeutically-effective neuromodulation e... | Medtronic Ardian Luxembourg S.a.r.l. | A61B018/02 A61N001/36 A61B005/00 A61N001/05 A61B018/00 A61B018/14 A61B005/20
US10035850 | B2 | 2018-07-31 | Methods for treating neovascular age-related macular degeneration | 2015-11-06 | A method is provided for reducing the treatment burden for patients who have an intraocular neovascular disorder, the method comprising administering a therapeutically effective amount of VEGF antagonist on a dosing schedule that includes treatment intervals of 8 and/or 12 weeks. | 1. A method for treating neovascular age-related macular degeneration (nAMD) in a mammal, the method comprising: a) administering to the mammal three individual doses of a VEGF antagonist at 4-week intervals, wherein the VEGF antagonist is an anti-VEGF antibody that comprises the sequence of SEQ ID NO: 3;b) assessing the mammal for best corrected visual acuity (BCVA), visual acuity (VA), central subfield thickness (CSFT), and the presence of intraretinal cysts/fluid at week 12;c) assessing the mammal for BCVA, VA, CSFT, and the presence of intraretinal cysts/fluid at week 16;d) assessing the mammal for BCVA, VA, CSFT, and the presence of intraretinal cysts/fluid at week 20; ande) administering to the mammal an additional dose of the VEGF antagonist at week 20 and once every 8 weeks (q8 regimen) thereafter if the following criteria are met and every 12 weeks (q12 regimen) thereafter if the following criteria are not met: a) decrease in BCVA of ≥5 letters, due to nAMD disease activity, at Week 16 compared to Baseline,b) decrease in BCVA of ≥5 letters, due to nAMD disease activity, at Week 16 compared to Week 12,c) VA decline of ≥3 letters and CSFT increase ≥75 μm, at Week 16 compared to Week 12, andd) new or worsening intraretinal cysts (IRC)/intraretinal fluid (IRF) at Week 16 compared to Week 12. 2. The method of claim 1, further comprising assessing the mammal at Weeks 32 and 44 for BCVA, and administering to the mammals additional doses on a q8 regimen after the further assessment at Week 32 or 44 if BCVA is ≥5 letters due to nAMD disease activity compared to Week 12. 3. The method of claim 1, wherein the mammal is a human. 4. The method of claim 1, wherein the VEGF antagonist is administered by intravitreal injection. 5. A method for treating neovascular AMD (nAMD) comprising administering to a mammal three individual doses of a VEGF antagonist at 4-week intervals, followed by additional doses every 12 weeks (q12) and/or every 8 weeks (q8) depending on the outcome of disease activity assessments using pre-defined visual and anatomic criteria as assessed at Weeks 12, 16, 20, 32, and 44 after the first individual dose is administered, wherein the VEGF antagonist is an anti-VEGF antibody that comprises the sequence of SEQ ID NO: 3, wherein the mammal is treated every 8 weeks (q8 regimen) starting at Week 16 if the following criteria are met and every 12 weeks (q12 regimen) starting at Week 20 if the following criteria are not met: a) decrease in BCVA of ≥5 letters, due to neovascular AMD (nAMD) disease activity, at Week 16 compared to Baseline,b) decrease in BCVA of ≥5 letters, due to nAMD disease activity, at Week 16 compared to Week 12,c) VA decline of ≥3 letters and CSFT increase ≥75 μm, at Week 16 compared to Week 12, andd) new or worsening intraretinal cysts (IRC)/intraretinal fluid (IRF) at Week 16 compared to Week 12. 6. The method of claim 5, further comprising assessing a mammal selected for q12 regimen at weeks 20, 32, and 44 for B... | Novartis Pharma AG | C07K016/22 A61B003/12 A61K039/395 A61K009/00 A61B005/00 A61K039/00
US09342179 | B2 | 2016-05-17 | Strain sensing element, pressure sensor, microphone, blood pressure sensor, and touch panel | 2014-08-28 | According to one embodiment, a strain sensing element includes a film unit, and a sensing unit. The film unit has a film surface and is capable of being deformed. The sensing unit includes a first sensing element and a second sensing element. The first sensing element is provided between a part of the film unit and the second sensing element. The first sensing element includes a first magnetic layer having a changeable magnetization with a deformation of the film unit, a second magnetic layer provided apart from the first magnetic layer, and a first spacer layer provided between the first and second magnetic layers. The second sensing element includes a third magnetic layer having a changeable magnetization with the deformation of the film unit, a fourth magnetic layer provided apart from the third magnetic layer, and a second spacer layer provided between the third and fourth magnetic layers. | 1. A strain sensing element comprising: a film unit having a film surface and being deformable; anda sensing unit including a first sensing element and a second sensing element,the first sensing element being provided between a part of the film unit and the second sensing element,the first sensing element including a first magnetic layer, a second magnetic layer provided apart from the first magnetic layer in a first direction crossing the film surface, and a first spacer layer provided between the first magnetic layer and the second magnetic layer, a magnetization of the first magnetic layer being configured to change in accordance with a deformation of the film unit,the second sensing element including a third magnetic layer, a fourth magnetic layer provided apart from the third magnetic layer in the first direction, and a second spacer layer provided between the third magnetic layer and the fourth magnetic layer, a magnetization of the third magnetic layer being configured to change in accordance with the deformation of the film unit,whereina magnetization of the second magnetic layer and a magnetization of the fourth magnetic layer are fixed,the fourth magnetic layer is disposed between the third magnetic layer and the film unit, andthe second magnetic layer is disposed between the first magnetic layer and the film unit. 2. The element according to claim 1, further comprising: a first electrode; anda second electrode,the sensing unit being provided between the first electrode and the second electrode,a current flowing between the first electrode and the second electrode being configured to flow between the first magnetic layer and the second magnetic layer, andthe current being configured to flow between the third magnetic layer and the fourth magnetic layer. 3. A pressure sensor comprising: the strain sensing element according to claim 1; anda support supporting the film unit. 4. A strain sensing element comprising: a film unit having a film surface and being deformable; anda sensing unit including a first sensing element and a second sensing element,the first sensing element being provided between a part of the film unit and the second sensing element,the first sensing element including a first magnetic layer, a second magnetic layer provided apart from the first magnetic layer in a first direction crossing the film surface, and a first spacer layer provided between the first magnetic layer and the second magnetic layer, a magnetization of the first magnetic layer being configured to change in accordance with a deformation of the film unit,the second sensing element including a third magnetic layer, a fourth magnetic layer provided apart from the third magnetic layer in the first direction, and a second spacer layer provided between the third magnetic layer and the fourth magnetic layer, a magnetization of the third magnetic layer being configured to change in accordance with the deformation of the film unit,whereina magnetization of the... | KABUSHIKI KAISHA TOSHIBA | G01L001/12 G06F003/041 G01L009/16 A61B005/021 G01L009/00 H04R019/00 H04R019/04 H04R031/00 A61B005/00
US09339242 | B2 | 2016-05-17 | Systems, methods, components, and software for monitoring and notification of vital sign changes | 2014-04-04 | A system and method for monitoring and notification of injury wherein an individual wears a monitoring device, the monitoring device tracks biometric data such as blood oxygen levels, heart rate, and acceleration of the individual and sends the data to a base station, the base station analyzes the data to see if bodily changes such as decreased blood oxygen levels or significant changes in heart rate or acceleration has occurred, and the base station determines whether to sound an alarm to indicate death or severe bodily injury. | 1. An apparatus comprising a wearable device, wherein the wearable device includes: a circuit-board enclosure having a cradle containing a removable battery, wherein the cradle is in engagement with a transceiver module for contacting a wrist of a person;an adjustable strap for encircling the wrist and removably attaching the wearable device to the wrist, the adjustable strap having a first and a second portion sandwiched between a portion of the transceiver module and a portion of the cradle, with one of the adjustable strap portion being an end of the adjustable strap;a pulse oximeter for sensing change in a blood-related parameter;an accelerometer for sensing change in a body movement of the person; anda transmitter for transmitting the blood-related parameter and the body movement. 2. The apparatus of claim 1, wherein the pulse oximeter is a finger-mountable pulse oximeter. 3. The apparatus of claim 1, wherein the pulse oximeter is located on the inside of the adjustable strap. 4. The apparatus of claim 1, wherein the wearable device is also securably attachable to the person. 5. The apparatus of claim 1, further comprising a base station for bidirectional wireless communication with the wearable device. 6. The apparatus of claim 5, wherein biometric information for the person experiencing an emergency medical condition is inputted into the base station. 7. The apparatus of claim 6, wherein the biometric information inputted into the base station is baseline pulse data for use in determining whether the person is experiencing an emergency medical condition. 8. The apparatus of claim 1, wherein the adjustable strap encirlces a finger and includes a flexible circuit board supporting a processor coupled to a Bluetooth transmitter and the pulse-oximeter. 9. The apparatus of claim 8, further comprising a relay transceiver for receiving signals from the Bluetooth transmitter and relaying them to a base station. 10. The apparatus of claim 9, wherein the relay transceiver is mountable to the person or to a fixed location within a space containing the person. | Pacific Place Enterprises, LLC | A61B005/00 A61B005/1455 A61B005/11 G08B021/04
US10383604 | B2 | 2019-08-20 | Data analysis system and method | 2014-06-27 | A method and apparatus for determining at least one representative temperature value for a female human user for an extended period includes receiving at least a first, a second and a third plurality of temperature measurements obtained from a female human user during at least first, second and third respective extended periods, wherein each extended period includes at least one hour and wherein the start of each extended period is separated by at least 8 hours. At least one representative temperature value is calculated for the second extended period. The representative temperature value is calculated using at least one first temperature value obtained from a plurality of measurements taken during the first extended period, at least one second temperature value obtained from a plurality of measurements taken during the second extended period and at least one third temperature value obtained from a plurality of measurements taken during the third extended period. | 1. An apparatus for determining at least one representative temperature value for a female human user for an extended period, the apparatus comprising: a temperature sensor for obtaining a first, a second and a third plurality of temperature measurements from a human user during at least first, second and third respective extended periods;a communications interface for receiving at least the first, the second and the third plurality of temperature measurements obtained from a female human user;wherein each extended period comprises at least one hour and wherein the start of an extended period is separated by at least 8 hours from the beginning of a subsequent and a preceding extended period;a processor arranged to calculate, based on the temperature measurements, at least one representative temperature value for the second extended period, wherein the representative temperature value is calculated using: at least one first temperature value obtained from a plurality of measurements taken during the first extended period;at least one second temperature value obtained from a plurality of measurements taken during the second extended period; andat least one third temperature value obtained from a plurality of measurements taken during the third extended period; anda memory arranged to store the representative temperature value for further analysis. 2. A method of determining at least one representative temperature value for a female human user for an extended period, the method comprising: receiving at least a first, a second and a third plurality of temperature measurements obtained from a temperature sensor for obtaining a first, a second, and a third plurality of temperature measurements from a female human user during at least first, second and third respective extended periods, wherein each extended period comprises at least one hour and wherein the start of each extended period is separated by at least 8 hours from the beginning of a subsequent and a preceding extended period;receiving, via a communications interface, at least the first, the second, and the third plurality of temperature measurements obtained from a female human user;calculating, using a processor arranged to calculate based on the temperature measurements, at least one representative temperature value for the second extended period, wherein the representative temperature value is calculated using: at least one first temperature value obtained from a plurality of measurements taken during the first extended period;at least one second temperature value obtained from a plurality of measurements taken during the second extended period; andat least one third temperature value obtained from a plurality of measurements taken during the third extended period; andstoring the representative temperature value in a memory for further analysis. 3. The method according to claim 2 wherein each extended period comprises at least two time windows and wherein a representative temperature val... | FERTILITY FOCUS LIMITED | A61B005/00 A61B010/00 A61B005/01
US09330342 | B2 | 2016-05-03 | On-bed monitoring system for range of motion exercises with a pressure sensitive bed sheet | 2013-12-10 | A system includes a pressure sensitive material that provides an indication of applied pressure for multiple locations on the material, and an analysis device in communication with the pressure sensitive material. The analysis device receives the indication of applied pressure, determines, for each of multiple measurement periods, a pressure image from the indication of applied pressure such that a sequence of pressure images is determined, and constructs a manifold representing the sequence of pressure images. | 1. A system, comprising: a pressure sensitive material that provides an indication of applied pressure for a plurality of locations on the material; anda non-transitory computer-readable storage medium for a processing device in communication with the pressure sensitive material, including instructions to: receive the indication of applied pressure;determine, for each of a plurality of measurement periods, a pressure image from the indication of applied pressure such that a sequence of pressure images is determined; andconstruct a manifold representing the sequence of pressure images. 2. The system of claim 1, wherein the pressure sensitive material comprises an e-textile material layer. 3. The system of claim 1, wherein the pressure sensitive material is constructed in multiple layers, including at least: a first layer including conductive traces, a second layer of variable conductivity material positioned next to the first layer, and a third layer including conductive traces positioned next to the second layer. 4. The system of claim 3, wherein the plurality of locations on the material is a plurality of intersections of the conductive traces of the first layer and the third layer, and at each intersection, the conductive traces of the first layer and the third layer cross at an angle with respect to each other, and the conductive traces of the first layer and the third layer are separately selectable to allow for measuring a conductance at each intersection. 5. The system of claim 4, wherein the indication of applied pressure includes the conductance measured at multiple intersections. 6. The system of claim 1, wherein the instructions to construct the manifold include instructions to: search for the k nearest neighbors;perform a weighted reconstruction; andcreate a low dimensional embedding. 7. The system of claim 1, wherein the indication of applied pressure represents training data and the manifold is a first manifold, and the analysis device is further configured to: determine the similarity of a second manifold to the first manifold. 8. The system of claim 1, wherein the manifold represents a defined rehabilitation exercise. 9. The system of claim 8, wherein the defined rehabilitation exercise is one of: leg lift, heel slide, lateral roll, head lift, and sit-up. 10. A method, comprising: receiving data representing conductance measurements from a plurality of pressure sensors;determining a sequence of pressure maps from at least a portion of the data;for each pressure map in the sequence of pressure maps, searching for the k nearest neighbors;performing a weighted reconstruction from the k nearest neighbors; andcreating a low dimensional embedding of the weighted reconstruction. 11. The method of claim 10, wherein the low dimensional embedding is embedding in a manifold. 12. The method of claim 10, wherein the data represents pressure between a subject and a surface measured at a plurality of points on the surface, and the low ... | THE REGENTS OF THE UNIVERSITY OF CALIFORNIA | G06K009/62 G06K009/00 A61B005/00 A61B005/103 A61B005/11
US09332960 | B2 | 2016-05-10 | System and method for determining location and orientation of a device in-vivo | 2012-02-01 | A system and method for determining location and orientation of an in-vivo device, with respect to an external system in which the device is located include a frame with external magnets attached thereon. An in-vivo device is inserted into the patient's body, which is placed within the system, and the external magnets apply magnetic forces on the in-vivo device. A radio beacon transmitter is attached to the frame for transmitting a radio pulse. The in-vivo device includes an ultrasonic transmitter for transmitting an ultrasonic signal, which is triggered by the radio pulse. At least three transponders are placed on the patient's body, each transponder sending a first acoustic signal triggered by the radio pulse, and each sending a second acoustic signal triggered by the device's ultrasonic signal. At least three sonic detectors are located on the frame for detecting each of the transponders' first and second sonic signals, and a processor measures time of detection of the transponders' signals and thus calculates location of the device in frame coordinates, and in body coordinates. | 1. A system for determining a location of an in-vivo device, the system comprising: a frame comprising external magnets and magnetic actuators for surrounding a patient's body, said body comprising an in-vivo device, said in-vivo device comprising an internal magnet, wherein said external magnets are configured for applying magnetic forces on the internal magnet of said in-vivo device;a radio beacon transmitter attached to the frame for transmitting a radio pulse;wherein said in-vivo device comprises an ultrasonic transmitter for transmitting an ultrasonic signal that is triggered by said radio pulse;at least three transponders placed on the patient's body, each transponder sending a first acoustic signal triggered by the radio pulse, and each sending a second acoustic signal triggered by the in-vivo device's ultrasonic signal;at least three sonic detectors located on the frame for detecting each of the transponders' first and second acoustic signals; anda processor configured for: measuring a time of detection of said transponders' signals by said sonic detectors,measuring a time of detection of said in-vivo device's signal by said transponders,calculating a distance between said transponders and said sonic detectors, based on the time of detection of said transponders' signals,calculating a distance between the in-vivo device and said transponders, based on the time of detection of the in vivo device's signal, anddetermining the location of the in-vivo device in frame coordinates, and in body coordinates, based on both the distance between the transponders and the sonic detectors, and the distance between the in-vivo device and the transponders. 2. The system according to claim 1, wherein each transponder comprises an ultrasonic receiver and a sonic transmitter. 3. The system according to claim 1, wherein said sonic detectors are connected to the processor. 4. The system according to claim 1, wherein said in-vivo device comprises at least two ultrasonic transmitters, each transmitting an ultrasonic signal having a different frequency. 5. The system according to claim 1, wherein said in-vivo device is a swallowable capsule endoscope. 6. The system according to claim 1, wherein said device comprises tools for performing operations at a target location in-vivo. 7. A method for detecting a location of an in-vivo device, the method comprising: inserting an in-vivo device into a patient's body, said device comprising an ultrasonic transmitter and an internal magnet;placing the patient into a frame surrounding said patient, the frame comprising external magnets for applying magnetic forces on the internal magnet of the in-vivo device;sending a radio pulse through a radio beacon transmitter to at least three transponders placed on the patient, and to the in-vivo device, for triggering each of the transponders to transmit an acoustic signal and for triggering the device to transmit an ultrasonic signal at substantially the same time period;detecti... | GIVEN IMAGING LTD. | A61B005/05 A61B008/08 A61B001/00 A61B001/04 A61B005/06 A61B005/07 A61B008/00 A61B017/00
US09351652 | B2 | 2016-05-31 | Systems for safe and remote outpatient ECG monitoring | 2014-07-14 | A system and method providing outpatient ECG monitoring and safe home based cardiac tele-rehabilitation. The system includes a recordation module for recording ECG signals using at least one lead, a tele-rehabilitation module for home based exercise management for a patient's recovery from a heart disease, the tele-rehabilitation module including a processing module for recognizing erroneous data from the ECG signals and an analysis module for calculating beat-to-beat annotations and determining if an ECG event and/or if a QT interval duration change has occurred. The system can include an exercise module for guiding the patient during an exercise session, a visual display that informs the patient to start and/or to stop the tele-rehabilitation exercise, a visual display and/or audible signal that informs the patient of an incoming or a missed tele-rehabilitation exercise session, and/or a communication module for transmitting/receiving data between the a cardiac tele-rehabilitation module and a physician/monitoring center. | 1. A system for remote monitoring of a patient's state, the system comprising: a tele-rehabilitation or continuous arrhythmia diagnostics module configured to acquire continuous real-time beat-to-beat electrocardiogram (ECG) signals of the patient;an interactive user interface configured to acquire patient subjective symptoms, the patient subjective symptom association with an ECG fragment of the continuous real-time beat-to-beat ECG signals;an analysis module configured to automatically annotate the ECG signals with information describing each beat represented by the ECG signals and automatically detect a cardiac event, the analysis module further configured to correlate the ECG signals with the automatically annotated ECG signals, automatically detected cardiac event, and patient subjective symptoms;a communication module configured to transmit the continuous real-time beat-to-beat ECG signals, annotations, automatically detected events, and patient subjective symptoms to a control module;the control module configured to visual represent the continuous real-time beat-to-beat ECG signals, annotations, automatically detected events, and patient subjective symptoms; anda reporting module configured to report annotations, for every ECG beat and/or the automatically detected cardiac event, and/or noise level for each ECG beat and ADC network interference for each ECG beat. 2. A method for analyzing limited-lead electrocardiogram (ECG) system signals, comprising: recording ECG signals, by a recording module, using at least one lead;performing cardiac tele-rehabilitation, by a tele-rehabilitation module, wherein performing cardiac tele-rehabilitation includes:recognizing erroneous data from the ECG signals, wherein recognizing erroneous data from the ECG signals includes using a signal pre-processing algorithm (SPA) configured to i) divide the ECG signals into overlapping blocks, wherein block sizes are related to a calculated average hear rate (HR), ii) subtract, from each block, a mean value of each block, and iii) reconstruct the ECG signals by cross-fading the successive segments;calculating annotations for every ECG beat;reporting the calculated annotations, by a reporting module, for every ECG beat and/or the ECG events and/or noise level for each ECG beat and ADC network interference for each ECG beat. 3. The method of claim 2, wherein the ECG event is reported, by the reporting module, to medical personnel or a wearer of the recordation module. 4. The method of claim 2, wherein the calculated annotations representing each ECG beat are segmented and then reported, by the reporting module, to medical personnel. 5. The method of claim 2, further comprising receiving, via the reporting module, the annotations for each ECG beat and/or the ECG event at a remote location. 6. The method of claim 2, wherein the erroneous data from the ECG signals is detected using an analysis algorithm. 7. The method of claim 6, wherein the analysis algorithm inc... | Medicalgorithmics S.A. | A61B005/04 A61B005/0452 A61B005/0464 A61B005/0472 G06F019/00 A61B005/00 A61B005/0432
US09333326 | B2 | 2016-05-10 | Guidewire | 2014-06-25 | A guide wire has a coil body covered with a hydrophilic coating film that exhibits sufficient lubricity even when there is little moisture in its surroundings. In the coil-type guide wire, a coating portion is formed by a plurality of layers where the innermost layer includes a hydrophobic film and an outer layer includes a hydrophilic film. Portions of the innermost layer extend between adjacent coils of the coil body and spaces are formed between the outer layer thereof and the portions of the innermost layer between the coils. The spaces allow moisture to pool so that when moisture in the guide wire's surroundings decreases, it is possible to supply the moisture pooled in the spaces to the hydrophilic film. Thus, even when there is little moisture in the guide wire's surroundings, it is possible to have a sufficient amount of lubricity. | 1. A guide wire, comprising: a core shaft;a coil body having a plurality of coils and which covers the core shaft, each of the coils of the coil body is spaced apart from adjacent coils so as not to contact the adjacent coils; anda coating portion which covers the coil body and which extends between the adjacent coils of the coil body, the coating portion being formed by a plurality of layers that are stacked on each other, whereinthe plurality of layers comprises: an innermost layer comprising a hydrophobic film that extends continuously over an outer periphery of the coil body including over gaps between the adjacent coils; andan outer layer comprising a hydrophilic film that extends continuously over an outer periphery of the innermost layer; anda first space is formed between the outer layer and the innermost layer only in areas that extend between the adjacent coils of the coil body. 2. The guide wire according to claim 1, wherein: the outer layer comprises at least two layers, each of the at least two layers comprising a hydrophilic film; anda second space is formed between layers of the outer layer at axial positions of the guide wire corresponding to the portions of the innermost layer that extend between the adjacent coils of the coil body. 3. The guide wire according to claim 2, wherein: an outermost layer of the at least two layers of the outer layer has a higher degree of water retentivity than other layers of the outer layer. | ASAHI INTECC CO., LTD. | A61B005/00 A61M025/00 A61M025/09
US09351708 | B2 | 2016-05-31 | Automated doppler velocimetry using a low-cost transducer | 2011-12-16 | An automatic, stand-alone, hand-held ultrasound blood-vessel examination device includes a reduced number of transducer elements and presents a simplified user interface, without the need for displaying an image of any of the vessels. The probe, in one embodiment, acquires and examines a volume of interest, searches for a target vessel, tests the vessel for normality of blood flow, and reports the diagnosis, all automatically and without need for user intervention. In another embodiment, the probe finds a body vessel in a volume, and extracts a clinical Doppler parameter, all automatically and without need for user intervention. | 1. An ultrasound device having a processor configured to carry out a series of acts, said series comprising acts of: a) examining, using an ultrasound probe of the ultrasound device, a volume to find body vessels present within the volume; andb) if one or more vessels are found within the volume in said examining act, then b1) selecting, for fluid-flow analysis, a vessel from among said vessels found within the volume, b2) generating, for said fluid-flow analysis, information particular to the selected vessel, b3) determining, based on the generated information, whether said selected vessel matches a target vessel, and b4) if the selected vessel matches the target vessel, then b4) (i) providing, based on the generated information, an indication as to normality of fluid flow in the selected vessel, else b4) (ii) repeating the selecting, generating, and determining with a next vessel of the vessels found, until there is no further next vessel of the vessels found within the volume,wherein said acts of examining and selecting (i) do not rely on a display of ultrasound scan images, and (ii) are performed automatically and without need for user intervention to interpret an ultrasound scan, andwherein the processor is further configured to determine, based on the generated information, whether the selected vessel is a maternal vessel or a fetal vessel. 2. The ultrasound device of claim 1, wherein the generated information comprises a quantitative characteristic of a waveform, wherein said waveform is representative of magnitude, over time, of velocity of fluid in said selected vessel. 3. The ultrasound device of claim 1, wherein the processor is further configured to perform acts of performing said fluid-flow analysis of the generated information, and, subject to a result based on said fluid-flow analysis, providing, based on said generated information, the indication as to normality of fluid flow in said selected vessel. 4. The ultrasound device of claim 3, wherein the processor is further configured to perform an act of, at least until said providing is performed, repeating, with said selected vessel serving as a current vessel, act b) for each next vessel from among said vessels found. 5. The ultrasound device of claim 3, wherein the processor is further configured to perform an act of classifying said selected vessel based on said generated information, said providing act being subject to a result of said classifying. 6. The ultrasound device of claim 1, wherein the processor is further configured to perform an act of designating, prior to said examining act, the target vessel within a vessel category, said category being based on physiology. 7. The ultrasound device of claim 1, wherein the ultrasound emitted from the elements and reflected back to the ultrasound probe is used to acquire volumetric data to be examined. 8. The ultrasound device of claim 1, wherein the ultrasound device is implemented as a hand-held, stand-alone device. 9. Th... | Koninklijke Philips N.V. | A61B008/00 A61B008/08 A61B005/00 A61B008/06
US09675293 | B2 | 2017-06-13 | Image analyzing device, image analyzing method, and computer program product | 2015-11-17 | According to an embodiment, an image analyzing device includes a first acquirer, a constructor, a first calculator, a second calculator, and a third calculator. The first acquirer is configured to acquire image information on a joint of a subject and bones connected to the joint. The constructor is configured to construct a three-dimensional shape of the bones and the joint, and relation characteristics between a load and deformation in the bones and the joint from the image information. The first calculator is configured to calculate a positional relation between the bones connected to the joint. The second calculator is configured to calculate acting force of a muscle acting on the bones connected to the joint based on the positional relation. The third calculator is configured to calculate first stress acting on the joint based on the three-dimensional shape, the relation characteristics, and the acting force. | 1. An image analyzing device comprising: a first acquirer configured to acquire image information on a joint of a subject and bones connected to the joint;a constructor configured to construct a three-dimensional shape of the bones and the joint, and relation characteristics between a load and deformation in the bones and the joint from the image information;a first calculator configured to calculate a positional relation between the bones connected to the joint;a second calculator configured to calculate acting force of a muscle acting on the bones connected to the joint based on the positional relation; anda third calculator configured to calculate first stress acting on the joint based on the three-dimensional shape, the relation characteristics, and the acting force. 2. The device according to claim 1, wherein the third calculator calculates distribution of the first stress acting on the joint based on the three-dimensional shape, the relation characteristics, and the acting force. 3. The device according to claim 1, wherein the second calculator performs inverse dynamic calculation using the positional relation to calculate the acting force. 4. The device according to claim 1, wherein the first calculator extracts the muscle from the image information and further calculates a feature indicating a length of the muscle based on feature points of the muscle including a start portion and an end portion on the bones, andthe second calculator performs inverse dynamic calculation using the positional relation and the feature to calculate the acting force. 5. The device according to claim 1, wherein the constructor further constructs second relation characteristics between a load and deformation of a living tissue other than the bones or the joint, the living tissue being hardly deformed by a factor other than the load but deformation caused by the load is capable of being extracted, andthe third calculator calculates the first stress acting on the joint based on the three-dimensional shape, the relation characteristics, the second relation characteristics, and the acting force. 6. The device according to claim 1, wherein the acting force includes at least one of muscle tension of the muscle coupled to the bones, relation characteristics between a load and deformation of a soft tissue attached to the bones, and torque acting on the joint. 7. The device according to claim 1, further comprising: a second acquirer configured to acquire an artificial joint model indicating a three-dimensional shape and an arrangement position of an artificial joint; anda fourth calculator configured to calculate second stress acting on the artificial joint based on a second mechanical model obtained by adding the artificial joint model to a first mechanical model indicating the three-dimensional shape and the relation characteristics and based on the acting force. 8. The device according to claim 7, further comprising a receiver configured to receive a change ins... | Kabushiki Kaisha Toshiba | G06K009/00 A61B005/00 G06K009/46 G01N003/08 G06T019/00 G06T007/11 G06T007/136
US09355458 | B2 | 2016-05-31 | Image processing apparatus, method and program | 2015-03-13 | A path detection-use graph structure is generated based on a plurality of nodes representing the plurality of linear structures, and a path that is included in the generated path detection-use graph structure and connects a plurality of root nodes representing points of origin of the plurality of linear structures to each other is detected. Then, based on a predetermined condition representing a feature of an erroneous connection edge erroneously connecting two nodes that are to belong to different graph structures to each other, a connection cost is set for each of edges forming the path so that the erroneous connection edge is hard to connect, and based on the set connection costs, the plurality of graph structures corresponding respectively to the plurality of linear structures are generated. | 1. An image processing apparatus, comprising a graph structure generation unit configured to generate, based on medical image data including a plurality of linear structures each repeatedly branching from a point of origin and extending in directions away from the point of origin in such a manner as to become wider, graph structures by defining the plurality of linear structures through use of a plurality of nodes and a plurality of edges connecting the plurality of nodes to each other, wherein the graph structure generation unit comprises: a node extraction unit configured to extract, based on the medical image data, the plurality of nodes for defining the plurality of linear structures;a root node extraction unit configured to extract, from the extracted plurality of nodes, a plurality of root nodes corresponding respectively to the points of origin of the plurality of linear structures;a path detection unit configured to generate a path detection-use graph structure based on the plurality of nodes and detect a path that is included in the generated path detection-use graph structure and connects the plurality of root nodes to each other; anda cost setting unit configured to set, based on a predetermined condition representing a feature of an erroneous connection edge, which erroneously connects two nodes that are to belong to different graph structures to each other, a connection cost for each of edges forming the path so that the erroneous connection edge is hard to connect, andwherein the graph structure generation unit generates, based on the plurality of nodes and the set connection costs, a plurality of graph structures corresponding respectively to the plurality of linear structures. 2. An image processing apparatus according to claim 1, wherein the graph structure generation unit generates provisional graph structures corresponding to the plurality of linear structures based on the plurality of nodes, discriminate, for each path connecting the two root nodes to each other, the erroneous connection edge based on the connection costs, and disconnects the discriminated erroneous connection edge, to thereby generate the plurality of graph structures from the provisional graph structures. 3. An image processing apparatus according to claim 1, wherein the path comprises a shortest path connecting two root nodes that are selected from among the plurality of root nodes to each other. 4. An image processing apparatus according to claim 1, wherein the cost setting unit sets the connection cost for the each of the edges forming the path so that the each of the edges is hard to connect as a distance from the two root nodes becomes larger. 5. An image processing apparatus according to claim 1, wherein the cost setting unit acquires, based on the medical image data, thickness of each of the plurality of linear structures corresponding to the each of the edges forming the path, and sets the connection cost for the each of the edges forming the pat... | FUJIFILM Corporation | G06K009/00 G06T007/00 A61B005/055 A61B005/08 G06T017/05 G06T019/00
US09327141 | B2 | 2016-05-03 | Gantry image guided radiotherapy system and related treatment delivery methods | 2014-11-19 | Systems, methods, and related computer program products for image-guided radiation treatment (IGRT) are described. Provided according to one preferred embodiment is an IGRT apparatus including a barrel-style rotatable gantry structure that provides high mechanical stability, versatility in radiation delivery, and versatility in target tracking. Methods for treatment radiation delivery using the IGRT apparatus include conical non-coplanar rotational arc therapy and cono-helical non-coplanar rotational arc therapy. A radiation treatment head (MV source) and a treatment guidance imaging system including a kV imaging source are mounted to and rotatable with a common barrel-style rotatable gantry structure, or alternatively the MV and kV sources are mounted to separate barrel-style rotatable gantry structures independently rotatable around a common axis of rotation. Methods for intra-fraction target tracking in a gantry-style IGRT system based on comparisons between a pre-acquired planning image and intrafraction x-ray tomosynthesis images and/or intrafraction cone beam CT (CBCT) images are also described. | 1. A method of image guided radiation treatment (IGRT) of a body part, comprising: providing an IGRT apparatus including a rotatable gantry structure, a radiation treatment head, and a treatment guidance imaging system including a first x-ray cone beam imaging source mounted to and rotatable with said rotatable gantry structure, the treatment guidance imaging system further including a first imaging detector;receiving a pre-acquired image data set of the body part acquired in a reference frame generally independent of a reference frame of the IGRT apparatus;during a patient setup interval, positioning the body part into an initial treatment position relative to the IGRT apparatus, said initial treatment position being along an axis of rotation of said rotatable gantry structure;rotating the rotatable gantry structure through a range of gantry angles greater than 180 degrees plus a fan beam angle of the first x-ray cone beam imaging source;subsequent to said patient setup interval, operating said first x-ray cone beam imaging source and said first imaging detector to acquire a first population of x-ray cone beam projection images of the body part for a respective first population of gantry angles and acquisition times;processing the first population of x-ray cone beam projection images to compute therefrom a time sequence of sliding-window cone beam CT (CBCT) volumes characterized in that each subsequent member of said time sequence is computed using at least one same x-ray cone beam projection image as used in computing at least one previous member of said time sequence; andoperating the radiation treatment head to deliver treatment radiation to the body part based at least in part on a comparison between each of said time sequence of sliding-window CBCT volumes and said pre-acquired image data set. 2. The method of claim 1, further comprising: computing a latest member of said time sequence of sliding-window CBCT image volumes;comparing said latest member with said pre-acquired image data set; anddelivering treatment radiation to the body part based said comparison of said latest member with said pre-acquired image data set. 3. The method of claim 2, wherein said computing the latest member of said time sequence of CBCT image volumes comprises: selecting from said first population of x-ray cone beam projection images a first subset thereof having a corresponding range of gantry angles greater than 180 degrees plus said fan beam angle from a latest gantry angle and less than 360 degrees from the latest gantry angle; andcomputing said latest member from said first subset of said first population of x-ray cone beam projection images. 4. The method of claim 2, said first x-ray cone beam imaging source comprising an x-ray source array having a number of x-ray sources positioned thereacross, each of said first population of x-ray cone beam projection images having an angle of incidence upon the body part determined by the gantry angle corresponding... | Accuray Incorporated | A61N005/10 A61B006/00 A61B005/00 A61B005/055 A61B006/02 A61B006/03
US09867544 | B2 | 2018-01-16 | Spectrometer including vertical stack structure and non-invasive biometric sensor including the spectrometer | 2015-08-10 | Provided are a spectrometer that may be easily manufactured while having high resolution and sensitivity due to reduced light loss and a non-invasive biometric sensor including the spectrometer. The spectrometer includes: a stacked light absorbing structure including a plurality of absorbing layers stacked in a vertical direction and having different absorption wavelength bands, and a plurality of tunnel junction layers respectively interposed between the plurality of absorbing layers to electrically connect the plurality of absorbing layers; and an illuminating unit configured to provide the stacked light absorbing structure with an illumination light for saturation of the plurality of absorbing layers. | 1. A spectrometer comprising: a stacked light absorbing structure comprising a plurality of absorbing layers stacked in a vertical direction and having a plurality of absorption wavelength bands, and a plurality of tunnel junction layers respectively interposed between the plurality of absorbing layers to electrically connect the plurality of absorbing layers;a light source configured to generate an illumination light including the plurality of absorption wavelength bands of the plurality of absorbing layers;a notch filter array disposed between the light source and the stacked light absorbing structure and comprising a plurality of notch filters having a plurality of absorption characteristics to block only a wavelength band selected from among the plurality of absorption wavelength bands; anda controller configured to control the notch filter array to rotate according to a lapse of time to change, in a sequential order, the selected wavelength band to be blocked among the plurality of absorption wavelength bands. 2. The spectrometer of claim 1, wherein each of the plurality of notch filters has a sector shape, andthe notch filter array has a rotatable disk shape that is formed by connecting the plurality of notch filters on a rotation axis. 3. The spectrometer of claim 1, wherein each of the plurality of absorbing layers comprises a positive-negative (p-n) junction, andeach of the plurality of tunnel junction layers comprises a p-n junction that is doped at a higher doping concentration than the p-n junction of each of the plurality of absorbing layers. 4. The spectrometer of claim 1, wherein the plurality of absorbing layers are arranged in descending order of energy band gap from a light input side, andeach of the plurality of tunnel junction layers has a larger energy band gap than the plurality of absorbing layers. 5. The spectrometer of claim 1, wherein the plurality of absorbing layers and the plurality of tunnel junction layers comprise a compound semiconductor. 6. The spectrometer of claim 5, wherein the plurality of absorbing layers and the plurality of tunnel junction layers are formed of a material comprising any one selected from a combination comprising gallium arsenic phosphide (GaAsP), gallium arsenide (GaAs), and indium gallium arsenide (InGaAs), a combination comprising aluminum gallium arsenide (AlGaAs), GaAs, and InGaAs, a combination comprising InGaAs, GaAs, and gallium phosphide (GaP), and a combination comprising indium arsenide (InAs), GaAs, and gallium nitride (GaN). 7. The spectrometer of claim 1, wherein each of the plurality of absorbing layers has a thickness of about 10 nm to about 10 μm, andeach of the plurality of tunnel junction layers has a thickness of about 10 nm to about 100 nm. 8. The spectrometer of claim 1, further comprising: a first electrode disposed on a bottom surface of the stacked light absorbing structure; anda second electrode disposed on a top surface of the stacked light absorbing structu... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/00 H01L031/0304 G01J003/51 G01J003/02 G01J003/10 G01J003/32 G01J003/42 G01J003/44 H01L031/103 G01J003/12
US09918690 | B2 | 2018-03-20 | Synthetic data-driven hemodynamic determination in medical imaging | 2015-10-07 | In hemodynamic determination in medical imaging, the classifier is trained from synthetic data rather than relying on training data from other patients. A computer model (in silico) may be perturbed in many different ways to generate many different examples. The flow is calculated for each resulting example. A bench model (in vitro) may similarly be altered in many different ways. The flow is measured for each resulting example. The machine-learnt classifier uses features from medical scan data for a particular patient to estimate the blood flow based on mapping of features to flow learned from the synthetic data. Perturbations or alterations may account for therapy so that the machine-trained classifier may estimate the results of therapeutically altering a patient-specific input feature. Uncertainty may be handled by training the classifier to predict a distribution of possibilities given uncertain input distribution. Combinations of one or more of uncertainty, use of synthetic training data, and therapy prediction may be provided. | 1. A method for hemodynamic determination in medical imaging, the method comprising: acquiring medical scan data representing an anatomical structure of a patient;extracting a set of features from the medical scan data;inputting, by a processor, the features to a machine-trained classifier, the machine trained classifier trained only from synthetic data not specific to any patients; andoutputting, by the processor with application of the machine-trained classifier to the features, a hemodynamic metric. 2. The method of claim 1 wherein acquiring comprises acquiring angiography data. 3. The method of claim 1 wherein acquiring comprises acquiring with the medical scan data comprising a two or three-dimensional representation of the anatomical structure. 4. The method of claim 1 wherein extracting the set of the features comprises: extracting geometrical features of the anatomical structure; andextracting the features of one or more abnormalities of the anatomical structure. 5. The method of claim 1 wherein extracting the set of the features comprises extracting functional features representing operation of the anatomical structure, wherein the machine-trained classifier is trained from virtual representations of the operation of the anatomical structure. 6. The method of claim 1 wherein extracting the set of the features comprises extracting an ischemic weight and an ischemic contribution score, the ischemic contribution score being a function of the ischemic weight. 7. The method of claim 6 wherein extracting the ischemic weight comprises computing branch ischemic weights as a function of a global ischemic weight. 8. The method of claim 6 wherein extracting the ischemic contribution score comprises computing the ischemic contribution score as a function of the ischemic weight and a radius. 9. The method of claim 1 wherein extracting the set of the features comprises extracting branch interaction features. 10. The method of claim 1 wherein inputting comprises inputting to the machine-trained classifier trained only from the synthetic data, the synthetic data comprising an in vitro model with a ground truth of the hemodynamic metric measured form the in vitro model. 11. The method of claim 1 wherein inputting comprises inputting to the machine-trained classifier trained only from the synthetic data, the synthetic data comprising an in silico model with a ground truth of the hemodynamic metric computed with computation fluid dynamics. 12. The method of claim 1 wherein inputting comprises inputting a sub-set of the set of features, the sub-set for a sub-part of the anatomical structure, and wherein outputting comprises outputting the hemodynamic metric for the sub-part of the anatomical structure; and further comprising subsequently repeating the inputting and outputting for remaining features of the set for another part of the anatomical structure. 13. The method of claim 1 wherein outputting comprises outputting a value of the hemodynamic... | Siemens Healthcare GmbH | G06K009/00 A61B006/00 G06T007/00 G06K009/52 G06K009/62 A61B005/026 G06F019/00 A61B008/06 A61B008/08 G06T007/11 A61B006/03 A61B005/02 A61B005/00 A61B008/00
US09338409 | B2 | 2016-05-10 | System and method for home health care monitoring | 2013-01-17 | A home health care monitoring method and system are disclosed. In one embodiment, a method includes: capturing a plurality of video sequences from a plurality of respective cameras disposed in different locations within a patient's home, including capturing two-dimensional image data and depth data for each video sequence; defining a plurality of events to monitor associated with the patient, the events including, during a predetermined time period, at least (1) the patient's body being in a particular position for at least a predetermined amount of time; performing depth-enhanced video content analysis on the plurality of video sequences to determine whether the event (1) has occurred. The video content analysis may include: for each of at least two cameras: automatically detecting a human object based on the two-dimensional image data; using the depth data to confirm whether the human object is in the particular position; and based on the confirmation, tracking an amount of time that the human object is in the particular position; and then determining that the event (1) has occurred based on the collective tracked amount of time. | 1. A method of monitoring the health of a person in home, the method comprising: capturing a plurality of video sequences from a plurality of respective cameras disposed in different locations within the patient's home, including capturing two-dimensional image data and depth data for each video sequence;defining a plurality of events to monitor associated with the patient, the events including, during a predetermined time period, at least (1) the patient's body being in a particular position for at least a predetermined amount of time;performing depth-enhanced video content analysis on the plurality of video sequences to determine whether the event (1) has occurred, wherein the video content analysis includes:for each of at least two cameras:automatically detecting a human object based on the two-dimensional image data;using the depth data associated with the detected human object to confirm whether the human object is in the particular position, without using depth data associated with other parts of the two-dimensional image in the confirmation; andbased on the confirmation, tracking an amount of time that the human object is in the particular position; anddetermining that the event (1) has occurred based on the collective tracked amount of time. 2. The method of claim 1, further comprising, for each camera: determining a height of the human object based on the depth data; and using the height to confirm that the human object is in the particular position. 3. The method of claim 2, further comprising, for each camera: determining a volume of the human object based on the depth data, and using the volume to verify that the human object is the person. 4. The method of claim 3, further comprising, for each camera: verifying that the human object is the patient by comparing the volume of the human object to a stored known volume of the patient. 5. The method of claim 2, further comprising, for each camera: performing facial recognition on the human object; and verifying that the human object is the patient based on the facial recognition. 6. The method of claim 1, wherein the particular position is one or more of sitting, standing up, or lying down. 7. The method of claim 6, wherein the events include: (2) the person spending at least a predetermined amount of time in a particular room. 8. The method of claim 7, wherein the events include: (3) the person taking medication. 9. The method of claim 1, further comprising: setting an alarm to go off based on whether event (1) occurs. 10. A home health monitoring system, comprising: a plurality of cameras in different rooms in a home, the different cameras configured to capture two-dimensional image data and depth data from a plurality of respective video images;a central computer system configured to perform video content analysis based on the two-dimensional image data and the depth data; andone or more communication media configured to transmit data from the plurality of cameras to the centr... | AVIGILON FORTRESS CORPORATION | A61B006/00 H04N007/18 G06K009/00 A61B005/00 A61B005/107 A61B005/11 A61B005/117 G08B013/196 G08B021/04
US20160120420 | A1 | 2016-05-05 | VASCULAR MEASUREMENT SYSTEM | 2014-10-30 | A vascular measurement system is provided to perform various types of peripheral vascular measurements to evaluate at least one of arterial blood flow and venous blood flow. The system includes a plurality of ports configured to connect a plurality of inflatable cuffs, an inflation device, a deflation device, and a volume measuring device configured to withdraw a predetermined volume of air contained in at least one of the first and second inflatable cuffs and resupply the predetermined volume of air into the at least one of the first and second inflatable cuffs. | 1. A system for determining blood pressure of a subject, the system comprising: a first port configured to be connected to a first inflatable cuff, the first inflatable cuff coupled to a first pressure sensor;a second port configured to be connected to a second inflatable cuff, the second inflatable cuff coupled to a second pressure sensor;an inflation device configured to inflate at least one of the first and second inflatable cuffs;a deflation device configured to deflate at least one of the first and second inflatable cuffs; anda volume measuring device configured to withdraw a predetermined volume of air contained in at least one of the first and second inflatable cuffs and resupply the predetermined volume of air into the at least one of the first and second inflatable cuffs. 2. The system of claim 1, wherein the volume measuring device comprises: a barrel;a plunger defining a chamber within the barrel, the chamber being in fluid communication with at least one of the first and second inflatable cuffs through the first and second ports, and the plunger displaceable within the barrel; andan actuator configured to move the plunger within the barrel. 3. The system of claim 2, wherein the actuator operates to move the plunger within the barrel in a first longitudinal direction to withdraw the predetermined volume of air contained in at least one of the first and second inflatable cuffs into the chamber, and move the plunger within the barrel in a second longitudinal direction opposite to the first longitudinal direction to resupply the withdrawn air into the at least one of the first and second inflatable cuffs. 4. The system of claim 3, wherein the predetermined volume of air is calculated by a longitudinal displacement of the plunger within the barrel, which is multiplied by a projected area of a surface of the plunger exposed to the chamber. 5. The system of claim 4, wherein the longitudinal displacement of the plunger is determined by a sensor module configured to detect a position of the plunger. 6. The system of claim 4, wherein the longitudinal displacement of the plunger is calculated based upon a number of rotations of a driving end of the actuator. 7. The system of claim 4, wherein the volume measuring device is configured to establish a pressure-to-volume relationship from the predetermined volume of air and a pressure difference of the at least one of the first and second inflatable cuffs, the pressure difference detected by at least of the first and second pressure sensors associated with the at least one of the first and second inflatable cuffs. 8. The system of claim 1, further comprising: a monitoring device configured to monitor a pressure detected by at least one of the first and second pressure sensors. 9. An apparatus for evaluating vascular flow, the apparatus comprising: a housing;a plurality of ports arranged on the housing and configured to be connected to a plurality of inflatable cuffs, the plurality of inflatab... |  | A61B005/022 A61B005/00 A61B005/026
US09339175 | B2 | 2016-05-17 | System and method to improve mouth disease diagnosis | 2007-12-31 | Embodiments of the present invention address the deficiencies of the art in respect to oral cavity medical assessment and provide a method, system and apparatus for mouth disease diagnosis and treatment. In an embodiment of the invention, an apparatus for scanning an oral cavity can include a mouth guard spreader that has a partially transparent receiving member that defines a receiving cavity, a first spreader portion and a second spreader portion coupled to the partially transparent receiving member. The apparatus further can include an image capture device for capturing images of the oral cavity disposed within the partially transparent receiving member of the mouth guard spreader. The apparatus yet further can include a storage device coupled to the image capture device for storing each captured image of the oral cavity. | 1. An apparatus for scanning an oral cavity, the apparatus comprising: a mouth spreader comprising opposing first and second angular spreader members coupled to one another at a receiving member, the opposing first and second angular spreader members acting in concert to create a gap between a cheek and a gumline of an oral cavity of only a single left or right side of a face by a first one of the spreader members contacting the cheek and gumline of the single side of a face to separate the cheek from the gumline and by a second one of the spreader members contacting only the cheek and gumline of the single left or right side of the face to separate the cheek from the gumline;an image capture device configured to capture images of an oral cavity and disposed within the receiving member; and,a data store coupled to the image capture device storing captured image of the oral cavity. 2. The apparatus of claim 1, wherein the first spreader member and the second spreader member are a first prong and a second prong. 3. The apparatus of claim 2, wherein the first prong is configured to include a first angled prong segment and a first extending prong segment and the second prong is configured to include a second angled prong segment and a second extending prong segment. 4. The apparatus of claim 2, wherein the first prong and the second prong have a substantially flat geometry. 5. The apparatus of claim 1, wherein the data store is one of a compact disk (CD), a digital versatile disk (DVD), a flash card, a smart card and an electronic diskette. 6. The apparatus of claim 1, wherein the image capture device is a camera having a wide angle lens. 7. The apparatus of claim 1, further comprising a processor for analyzing the captured image. 8. The apparatus of claim 1, further comprising a monitor for displaying the captured image. 9. The apparatus of claim 1, wherein the receiving member is partially transparent. 10. The apparatus of claim 1, wherein the first and second members and the receiving member are one piece. 11. The apparatus of claim 1, wherein the first and second members each define at an outer edge opposite the receiving member a channel into which teeth are placed. 12. The apparatus of claim 1, wherein the image capture device is tubular in shape and the receiving member also is tubular in shape and adapted to receive the image capture device. | International Business Machines Corporation | A61B001/32 A61B001/00 A61B001/24 A61B005/00
US09351665 | B2 | 2016-05-31 | Method and system for detecting a respiratory signal | 2010-01-14 | A method of analyzing a plethysmograph signal is disclosed. The method comprises: extracting from the signal a plurality of features, thereby constructing a feature space. The method further comprises employing a path selection procedure to the feature space for determining at least one sequence of respiration frequencies and reconstructing a respiratory signal using the sequence. | 1. A method of analyzing a photoplethysmograph signal, comprising: segmenting the photoplethysmograph signal into at least a plurality of temporal segments;extracting from each temporal segment of said plurality of temporal segments a plurality of features including at least respiration parameters, to construct a multidimensional feature space;employing a path selection procedure to said multidimensional feature space to assign a respiration frequency or a range of respiration frequencies separately for each segment of said plurality of temporal segments, thereby determining at least one sequence of respiration frequencies, wherein each respiration frequency or range of respiration frequencies corresponds to a state of said multidimensional feature space and wherein said path selection procedure calculates probabilities of transitions between states;reconstructing a respiratory signal using said sequence; anddisplaying said reconstructed respiratory signal and adjusting a therapy based on said displayed reconstructed respiratory signal. 2. The method according to claim 1, further comprising using said sequence of respiration frequencies for selecting a corresponding sequence of respiration tidal volumes. 3. The method according to claim 1, wherein said plurality of features further comprises spectral parameters, and the method further comprises analyzing the spectral parameters for detecting temporal segments corresponding to a physiological event, wherein said path selection procedure is employed separately for different types of physiological events. 4. The method according to claim 3, wherein said physiological event is selected from the group consisting of atrial fibrillation, periodic limb movements and apnea. 5. The method according to claim 3, further comprising updating said feature space following identification of a segment corresponding to a physiological event. 6. The method according to claim 5, further comprising calculating a transfer function filter for each segment other than segments corresponding to physiological events, wherein said updating of said feature space comprises filtering a respective segment corresponding to a respective physiological event using a transfer function filter of a preceding segment. 7. The method according to claim 1, wherein said path selection procedure comprises an expectation maximization algorithm. 8. The method according to claim 1, wherein said path selection procedure features a hidden Markov model. 9. Apparatus for analyzing a photoplethysmograph signal, comprising: an input unit for receiving the photoplethysmograph signal; anda data processor configured for segmenting the photoplethysmograph signal into at least a plurality of temporal segments; extract from each temporal segment of said plurality of temporal segments a plurality of features including at least respiration parameter, thereby constructing a multidimensional feature space; employing a path selection procedure to said m... | Widemed Technologies Ltd. | A61B005/08 A61B005/024 A61B005/1455
US09332921 | B2 | 2016-05-10 | Anti-microbial electromyography needle | 2013-05-31 | A coated electromyography needle including an anti-microbial, electrically insulative coating applied to non-tip portion of an electrode of the electromyography needle. | 1. An electromyography needle comprising: an electrode including an elongated body having a sharpened tip portion and a non-tip portion; andan electrically insulative coating on a surface of the non-tip portion of said electrode, wherein the electrically insulative coating includes a base material having a plurality of anti-microbial particles interspersed in said base material, said base material including an epoxy and a fluoropolymer. 2. The electromyography needle of claim 1, wherein the anti-microbial particles include at least one selected from the group consisting of: silver particles, silver-oxide particles, glass-silver particles, ceramic particles, a combination of silver and ceramic particles, copper particles, glass-copper particles, a combination of copper and ceramic particles, a combination of copper and silver particles and a combination of copper, silver and ceramic particles. 3. The electromyography needle of claim 1, which includes a top coating including a plurality of particles of a low-friction material. 4. The electromyography needle of claim 3, wherein at least some of the plurality of anti-microbial particles at least partially protrude through the top coating. 5. The electromyography needle of claim 1, wherein the electrode includes a cylindrical electrode. 6. The electromyography needle of claim 1, wherein the electrically insulative coating includes an additive that improves the bonding capabilities of the coating. 7. An electromyography needle comprising: a handle;an electrode including an elongated body having a sharpened tip portion, a non-tip portion and a handle connecting portion connected to the handle; andan electrically insulative coating on a surface of the non-tip portion of said electrode, wherein the electrically insulative coating includes a base material having a plurality of anti-microbial particles interspersed in said base material, said base material including an epoxy and a fluoropolymer. 8. The electromyography needle of claim 7, wherein the anti-microbial particles include at least one selected from the group consisting of: silver particles, silver-oxide particles, glass-silver particles, ceramic particles, a combination of silver and ceramic particles, copper particles, glass-copper particles, a combination of copper and ceramic particles, a combination of copper and silver particles and a combination of copper, silver and ceramic particles. 9. The electromyography needle of claim 7, which includes a top coating including a plurality of particles of a low-friction material. 10. The electromyography needle of claim 9, wherein the plurality of anti-microbial particles at least partially protrude from the electrically insulative coating, through the top coating. 11. The electromyography needle of claim 7, wherein the electrically insulative coating is on a surface of the handle connection portion of the electrode. 12. The electromyography needle of claim 7, wherein the electrode includes a... | Innovatech, LLC | A61B005/0492 A61B005/00 A01N059/16 A01N059/20
US09332925 | B2 | 2016-05-10 | Magnetic resonance imaging apparatus having continuous readout gradient with gradient slew rate changed | 2011-06-07 | A magnetic resonance imaging apparatus according to an embodiment includes a sequence control unit that controls a gradient magnetic field power supply, thereby performing a pulse sequence including the application of a continuous readout gradient magnetic field pulse. The sequence control unit controls the gradient magnetic field power supply such that the slew rate of the gradient magnetic field pulse is reduced in stages as the output voltage of a gradient magnetic field amplifier is reduced in stages. | 1. A magnetic resonance imaging apparatus comprising: a gantry, static and gradient magnetic field generators, at least one radio frequency (RF) coil coupled to an imaging volume, RF transmitter and RF receiver circuitry coupled to the at least one RF coil, and a gradient magnetic field power supply including a gradient magnetic field amplifier coupled to the gradient magnetic field generators; andprocessing circuitry configured to:A) specify a waveform of a continuous readout gradient magnetic field pulse with a slew rate of the waveform being reduced according to a stepwise reduction in an output voltage of the gradient magnetic field amplifier; andB) perform an MRI pulse sequence including the application of the continuous readout gradient magnetic field pulse having the specified waveform of step A) with the processing circuitry also controlling the gradient magnetic field power supply. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the processing circuitry is further configured to: predict, when imaging conditions are received, a stepwise reduction in the output voltage on the basis of the received imaging conditions;specify the waveform of the gradient magnetic field pulse corresponding to the predicted stepwise reduction;generate an instruction signal on the basis of the specified waveform of the gradient magnetic field pulse; andperform the pulse sequence by controlling the gradient magnetic field power supply using the generated instruction signal. 3. The magnetic resonance imaging apparatus according to claim 1, wherein the processing circuitry is further configured to: measure the output voltage of the gradient magnetic field amplifier;specify, when the measured output voltage is received, the waveform of the gradient magnetic field pulse corresponding to the measured output voltage;change an instruction signal that controls the gradient magnetic field power supply on the basis of the specified waveform of the gradient magnetic field pulse; andperform the pulse sequence by controlling the gradient magnetic field power supply using the changed instruction signal. 4. The magnetic resonance imaging apparatus according to claim 1, wherein the processing circuitry is further configured to: reduce the slew rate of the waveform of the gradient magnetic field pulse at an interval of one period or an interval of a plurality of periods. 5. The magnetic resonance imaging apparatus according to claim 2, wherein the processing circuitry is further configured to: reduce the slew rate of the waveform of the gradient magnetic field pulse at an interval of one period or an interval of a plurality of periods. 6. The magnetic resonance imaging apparatus according to claim 3, wherein the processing circuitry is further configured to: reduce slew rate of the waveform of the gradient magnetic field pulse at an interval of one period or an interval of a plurality of periods. 7. The magnetic resonance imaging apparatus accordi... | TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/05 G01R033/54 A61B005/055
US09339603 | B2 | 2016-05-17 | Adhesive and peripheral systems and methods for medical devices | 2009-02-27 | A repeater system may control a pump by using a repeater and a user interface. An adhesive patch system may be used for affixing a pump or other object to a human body. Such an adhesive patch system may include two sets of adhesive members, each member including an adhesive material on at least one side so as to attach to the body. The members of the first set are spaced to allow the members of the second set to attach to the body in spaces provided between the members of the first set, and the members of the second set are spaced to allow members of the first set to detach from the body without detaching the members of the second set. Also, fill stations and base stations are provided for personal pump systems. | 1. A delivery device for delivering an infusion medium to a user, the device comprising: a disposable housing assembly adapted to be secured to a user's skin;a reusable housing assembly adapted to be secured to the disposable housing assembly, the reusable housing assembly comprising a volume sensing device, wherein the volume sensing device determines the volume of fluid delivered to the user;a reservoir having an interior for containing a fluidic medium, the reservoir having an outlet port and being operatively coupled to a pumping assembly to selectively drive fluid out the outlet port, from the reservoir interior; anda needle insertion device for inserting a cannula into the user's skin, the needle insertion device comprising an introduction needle for movement between a retracted position and an extended position; wherein the disposable assembly has a passage through which the introduction needle extends, andwherein the cannula has a central passage through which the introduction needle extends. 2. The delivery device of claim 1, the disposable assembly further comprising: a base portion comprising the reservoir the base having a reservoir cavity, and a fluid path, the reservoir cavity comprising the outlet port and a septum, the fluid path comprising a fluid channel from the outlet port to a cannula port;at least one flexible membrane portion attached onto the base portion and forming a top of the reservoir cavity and a top of the fluid channel; anda tubing connected to the base portion at the cannula port, the tubing in fluid communication with the reservoir cavity. 3. The delivery device of claim 2, wherein the tubing portion comprises a cannula port end and a cannula interface end, the cannula interface end configured to attach to a cannula. 4. The delivery device of claim 2, wherein the fluid path further comprises a series of regions exposed to the flexible membrane portion, at least one of the regions being a valve region. 5. The delivery device of claim 4, wherein the series of regions further comprises a pump chamber region. 6. The delivery device of claim 2, wherein the base portion and the flexible membrane portion constitute a fluidic assembly, and wherein the disposable housing assembly further comprises a disposable base into which the fluidic assembly fits. 7. The delivery device of claim 6, further comprising a pad coupled to a bottom of the disposable base, the pad for attaching the disposable housing assembly to a patient. 8. The delivery device of claim 2, wherein the reservoir port in fluid communication with the reservoir cavity and the septum for penetration by a needle so as to fill the reservoir cavity with therapeutic fluid. 9. The delivery device of claim 2, the disposable housing assembly further comprising a latching mechanism for permitting selective engagement and disengagement with the disposable assembly. 10. The delivery device of claim 2, the disposable housing assembly further comprising a mechanis... | DEKA Products Limited Partnership | A61M005/145 A61M005/158 A61M005/142 A61M005/168 G05D007/06 A61M005/172 G08C017/02 H04B007/26 G05B023/02 A61B005/00 A61M005/50
US09339194 | B2 | 2016-05-17 | System, method and article for normalization and enhancement of tissue images | 2011-03-08 | In medical imaging, a fiducial marker facilitates tissue image correlation that allows for image analysis, normalization and correction of the optical exposure and spectral and spatial distribution in order to compensate for the surface reflections, sub surface tissue interactions and spatial orientation of the excitation and imaging axes to the subject tissue. Using a cross comparison, clinicians can model tissue image data in different forms in order to reference and compare data from various spectral components and or from different images. This may enhance human interpretation between images including the variations between images even when the spectral, spatial and optical conditions or the image resolution or sensitivity are compromised. Such may be used to assess cosmetic, moisturizing, therapeutic materials and treatments. | 1. A system for use in tissue analysis, the system comprising: at least one processor; andat least one nontransitory storage medium that stores processor executable instructions which when executed cause the at least one processor to:compare an appearance of at least one shape of a virtual fiducial marker in a first digital image of a portion of a tissue to at least one defined actual shape of the virtual fiducial marker;compare an appearance of each of a plurality of sections of a physical fiducial marker in the first digital image to respective ones of defined sections of the physical fiducial marker, wherein the defined sections include a number of tissue phantoms each having a respective spectral characteristic that matches a respective spectral characteristic of tissue of a type represented in the first digital image; andat least one of correlate, normalize, or correct at least the first digital image, based at least in part on the comparisons;the system being characterized in that the physical fiducial marker includes a scatter layer that overlies at least some but not all of the tissue phantoms and which simulates an optical character of the type of tissue represented in the first digital image, and the instructions cause the at least one processor to compare the appearance of the sections which include the tissue phantoms which are overlaid by the scatter layer with a number of defined sections which include tissue phantoms that are not overlaid by the scatter layer. 2. The system of claim 1 wherein a number of sections of the physical fiducial marker include a respective color including at least one of black, white, a plurality of different shades of grey, and a plurality of additional colors that are not black, white or grey, and the instructions cause the at least one processor to compare the appearance of the sections which include the respective colors with respective ones of a defined set of respective colors. 3. The system of claim 1 wherein the instructions further cause the at least one processor to store the digital image as a multi-layer image file, including a first digital image layer that stores digital images and at least a second digital image layer that stores image metadata. 4. The system of claim 3 wherein the instructions further cause the at least one processor to store to a diagnostic layer of the digital image information indicative of at least one of an NADH fluorescence, a collagen fluorescence, a physical scattering of light from the tissue at a number of physical layers of the tissue due to tissue density, a spectral distribution due to a size of a cell nuclei, and a hemoglobin absorption due to increased blood flow or oxygenation; wherein normalizing a surface reflection in each digital image in a multiple image series, is achieved by normalizing color markers to allow accurate comparison of a Epidermal Layer to the Melanin Layer in the in proximity to lesion and where subsurface scatter is normalized to tho... | Cernoval, Inc. | A61B005/00 A61B005/0205 A61B005/11 G03B015/00 G03B015/14 A61B005/026 A61B005/145 A61B005/1455
US09351680 | B2 | 2016-05-31 | Method and apparatus for a variable user interface | 2004-10-14 | An analyte monitoring system is provided. The system may include a housing (200) and a visual display (206) on the housing (200), the visual display (206) having at lease one visual indicator position next to a corresponding marking (208, 210, 212, 214 and 216) on the housing (200). The system may include a processor (60) driving the visual display (206), wherein the processor (60) runs software that is modifiable to provide a variable user interface on the visual display (206). The system may include a wireless communication to allow applets or programs to be down-loaded by the processor (60). | 1. An analyte measurement device comprising: a housing;a plurality of penetrating members positionable in the housing;a penetrating member driver configured to be coupled to penetrating members;a processor configured to modify a user profile and modulate power from a power supply to the penetrating member driver through an amplifier;a visual display on said housing, said visual display having at least one visual indicator positioned next to a corresponding marking on the housing, the display configured to be customized based on a user's condition and include an indicator of a number of penetrating members available for use; andwherein the device is configured to receive telephone calls. 2. An analyte monitoring device comprising: a housing;a visual display on said housing , the display configured to be customized based on a user's condition and include an indicator of a number of penetrating members available for use,a plurality of penetrating members positionable in the housing;a penetrating member driver configured to be coupled to penetrating members;said visual display having at least one visual indicator positioned next to a corresponding marking on the housing;a voice recognition device coupled to the visual display;a processor configured to modify a user profile and modulate power from a power supply to the penetrating member driver through an amplifier; andwherein the device is configured to receive telephone calls. 3. A skin penetrating system, comprising: a housing;a plurality of penetrating members positioned in the housing,a penetrating member driver configured to be coupled to penetrating members;an analyte detecting member coupled to a sample chamber, the analyte detecting member being configured to determine a concentration of an analyte in a body fluid using a sample of less than 1 .mu.L of a body fluid disposed in the sample chamber;a visual display on said housing, said visual display having a screen saver which is activated after a preset period of nonuse by a user, the display configured to be customized based on a user's condition and include an indicator of a number of penetrating members available for use;a voice recognition device coupled to the visual display;a processor configured to modify a user profile and modulate power from a power supply to the penetrating member driver through an amplifier; andwherein the device is configured to receive telephone calls. 4. A tissue penetrating system, comprising: a housing;a penetrating member positioned in the housing,a visual display on said housing, said visual display having at least one visual indicator positioned next to a corresponding marking on the housing, the display configured to be customized based on a user's condition and include an indicator of a number of penetrating members available for use;a voice recognition device;a processor configured to modify a user profile and modulate power from a power supply to the penetrating member driver through an amplifier;a s... | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/00 A61B005/157 A61B005/15 A61B005/145
US09351667 | B2 | 2016-05-31 | Device, system, and method for determination of oral/lip stiffness | 2010-08-26 | A device for measuring orofacial stiffness in a subject can include: two lip saddle attachment components configured for attachment to the lip saddles of a patients mouth; two elongate members, each being coupled with one of the lip saddle attachment components; a pivot member that couples the two elongate members at a pivot point opposite from the two lip saddle attachment components; an electronic sensor configured to sense the stiffness of the lips by sensing movement of the elongate members with respect to the pivot point, where the electronic sensor is operably coupled to each of the elongate members; and a pressure component configured move with respect to the pivot point so as to provide pressure to and/or receive pressure from the lip saddle attachment components. | 1. A device for measuring orofacial stiffness in a subject having lips, the device comprising: two lip saddle attachment components configured for attachment to the lip saddles of the subject's mouth;two elongate members comprised of a first elongate member and a second elongate member, each being coupled with one of the lip saddle attachment components;a pivot member that couples the two elongate members at a pivot point at a distance from the two lip saddle attachment components;an electronic sensor mounted between the two elongate members and configured to sense stiffness of the subject's lips by sensing movement of the elongate members with respect to the pivot point, where the electronic sensor is coupled to each of the elongate members; anda pressure component mounted between the two elongate members and configured to move with respect to the pivot point so as to provide pressure-actuated displacement to increase a horizontal lip aperture of the subject's lips and/or receive pressure from the lip saddle attachment components,wherein one of the electronic sensor or pressure component is directly coupled to two movable coupling members comprised of a first movable coupling member and a second movable coupling member and the other of the electronic sensor or pressure component is indirectly coupled to the two movable coupling members, wherein the first movable coupling member is coupled to the first elongate member and the second movable coupling member is coupled to the second elongate member, and the first movable coupling member and second movable coupling member being coupled to opposite ends of the one of the electronic sensor or pressure component. 2. A device as in claim 1, further comprising: an electrode system for measuring changes in an orbicularis oris superior (OOS) and orbicularis oris inferior (OOI) of the subject, which electrode system being configured for determining whether the subject is actively moving, holding or tensing the lips. 3. A device as in claim 2, the electrode system further comprising an electrode pair for monitoring the OOS, an electrode pair for monitoring the OOI, and a reference electrode. 4. A device as in claim 3, further comprising a bite block. 5. A device as in claim 4, further comprising an anchor configured for anchoring the device to the subject. 6. A device as in claim 5, wherein the anchor is configured as a chin anchor or nose anchor. 7. A device as in claim 6, further comprising a computer system capable of being in communication with the sensor and the pressure component, and with the electrode system. 8. A device as in claim 7, characterized by at least one of the following: the two lip saddle attachment components each include a feature configured for attachment to the lip saddles of the subject's mouth;the two elongate members form an even arm cantilever in a “X” shape or a “V” shape;the pivot point forms an intersection of the two elongate members;the electronic sensor is a differen... | The University of Kansas | A61B005/117 A61B005/103 A61B005/11 A61B005/22 A61B005/0488 A61B005/00
US09339641 | B2 | 2016-05-17 | Method and apparatus for transdermal stimulation over the palmar and plantar surfaces | 2013-03-15 | Various methods and devices for providing transdermal electrical stimulation therapy to a patient are provided. In certain variations, a method may include positioning a stimulator electrode over a glabrous skin surface overlying a target nerve of a subject. Electrical stimulation may be delivered through or across the glabrous skin surface to the target nerve to stimulate the target nerve, while remaining safe and tolerable to the patient. Electrical stimulation may be delivered at frequencies that may be painful or intolerable when applied over non-glabrous surfaces of the body. Various applicators or devices for providing transdermal electrical stimulation therapy are also provided. | 1. A method for providing transdermal electrical stimulation therapy to a subject comprising: positioning a stimulator electrode over a glabrous skin surface overlying a palm of the subject;delivering, via a pulse generator, electrical stimulation transdermally through the glabrous skin surface and to a target nerve or tissue within the hand to stimulate the target nerve or tissue within the hand until pain felt by the subject is mitigated;providing feedback signals based on F-wave detection, andadjusting or optimizing the electrical stimulation delivered through the glabrous skin surface and to the target nerve or tissue within the hand based on the F-wave detection. 2. The method of claim 1, wherein the electrical stimulation is delivered at a frequency of about 5 Hz to about 60 Hz, while remaining safe and tolerable to the subject. 3. The method of claim 1, wherein the stimulator electrode is a surface electrode. 4. The method of claim 1. wherein the electrical stimulation is delivered intermittently or on a chronic basis. 5. The method of claim 1. further comprising: detecting nerve stimulation via at least one sensor positioned on the subject:receiving a signal from the sensor indicative of the detected electrical stimulation thereby providing a feedback about the efficacy of the applied electrical stimulation therapy such that the therapy is adjusted or optimized. 6. The method of claim 5, wherein the feedback is queried such that the electrical stimulation therapy is adjusted to ensure that a minimum amount of energy is being applied to stimulate the target nerve within the hand while reducing the risk of burns or intolerance. 7. The method of claim 5, wherein the feedback is queried such that the positioning of the stimulator electrode is adjusted to optimize the electrical stimulation therapy. 8. The method of claim 5, wherein the sensor is selected from the group consisting of a surface electrode, microneedle electrode, or motion sensor. 9. The method of claim 5, wherein afferent nerve stimulation is detected. 10. The method of claim 1, further comprising positioning a ground electrode on the subject to facilitate penetration of an electrical current through the glabrous surface to stimulate the target nerve within the hand. 11. The method of claim 1, wherein the stimulator electrode is attached to the glabrous skin surface with an adhesive. 12. The method of claim 1, wherein the stimulator electrode is positioned over the glabrous skin surface with an ergonomic applicator. 13. The method of claim 12, wherein the applicator is a glove or brace configured to he positioned against the palmar surface of a hand such that the electrode delivers electrical stimulation to a target nerve within the hand. 14. The method of claim 12, wherein the applicator and/or one or more electrodes are coated with a pain relief drug. 15. The method of claim 1, wherein the electrical stimulation is automatically paused for a preset amount of time... | EMKinetics, Inc. | A61N001/34 A61N001/04 A61N002/02 A61N001/36 A61N001/378 A61N002/00 A61N001/05 A61B005/00
US09332919 | B2 | 2016-05-10 | Heart monitoring apparatus | 2013-05-03 | An electrocardiographic device for sensing cardiac activity in a user, the electrocardiographic device comprising: plurality of electrodes provided at a surface of the device for electrical engagement with a pair of extremities of a user; one or more movement sensors arranged to, in use, detect movement of a user's extremities; and a control unit configured to receive electrical signals from the electrodes and to, in use, process said signals in dependence on the output of the one or more movement sensors so as to filter variations from the signals due to movement of the user's extremities. | 1. An electrocardiographic device for sensing cardiac activity in a user, the electrocardiographic device comprising: a plurality of electrodes provided at a surface of the device for electrical engagement with the hands of a user;a set of indicia indicating an intended position for a user's hands on the device, the set of electrodes being arranged relative to the set of indicia such that, in use when a user's hands are in the intended position, at least one electrode is located under the thenar eminence of each of the user's hands;one or more movement sensors arranged to, in use, detect movement of a user's hands; anda control unit configured to receive electrical signals from the electrodes and to, in use, process said signals in dependence on the output of the one or more movement sensors so as to filter variations from the signals due to movement of the user's hands. 2. An electrocardiographic device as claimed in claim 1, wherein the plurality of electrodes comprises a first set of one or more electrodes for engagement with a first hand of a user and a second set of one or more electrodes for engagement with a second hand of a user, the first set of electrodes having an associated first movement sensor arranged to detect movement of the first hand and the second set of electrodes having an associated second movement sensor arranged to detect movement of the second hand, the control unit being configured to perform filtering of signals from the first and second sets of electrodes in dependence on the output from the respective movement sensor. 3. An electrocardiographic device as claimed in claim 1, wherein said filtering comprises cancellation or suppression of features in electrical signals from the electrodes that are common to both the signals from the electrodes and the output of the one or more movement sensors. 4. An electrocardiographic device as claimed in claim 1, wherein the control unit is configured to provide said filtered electrical signals at an output of the device. 5. An electrocardiographic device as claimed in claim 1, wherein the output of the one or more movement sensors represents motion and/or vibration of the user's hands. 6. An electrocardiographic device as claimed in claim 1, wherein the one or more movement sensors include one or more of an accelerometer, a strain gauge, a microphone and a pressure sensor. 7. An electrocardiographic device as claimed in claim 1, wherein the electrodes of the set are provided at two electrically insulating user interaction regions of the electrocardiographic device, each user interaction region being adapted for engagement with a hand of a user. 8. An electrocardiographic device as claimed in claim 7, wherein each user interaction region comprises a layer of shape memory foam for engagement with a user's hand, the electrodes being provided at the shape memory foam. 9. An electrocardiographic device as claimed in claim 7, wherein each user interaction region is supported on ... | Cardiocity Limited | A61B005/0404 A61B005/0452 A61B005/00 A61B005/0408
US20160120455 | A1 | 2016-05-05 | Method for monitoring incontinence | 2014-11-04 | According to one aspect of the inventive concept there is provided a method of monitoring incontinence for a user, comprising: determining, using a urine sensitive circuit provided at an absorbent article and arranged to present a changed electrical characteristic when exposed to urine, whether the user has urinated on the absorbent article, performing, using a measurement circuit, a measurement on the urine bladder to determine a parameter which varies with a fill level of the urine bladder, and in response to determining that the user has urinated on the absorbent article, recording, by a processing circuit, data representing said parameter. | 1. A method of monitoring incontinence for a user, comprising: determining, using a urine sensitive circuit provided at an absorbent article and arranged to present a changed electrical characteristic when exposed to urine, whether the user has urinated on the absorbent article,performing, using a measurement circuit, a measurement on the urine bladder to determine a parameter which varies with a fill level of the urine bladder, andin response to determining that the user has urinated on the absorbent article, recording, by a processing circuit, data representing said parameter. 2. A method according to claim 1, further comprising, in response to determining that the user has urinated on the absorbent article, recording, by the processing circuit, data representing said parameter and an indication of that the user has urinated on the absorbent article. 3. A method according to claim 1, further comprising: performing, a plurality of times, using the measurement circuit, a measurement on the urine bladder to determine a parameter which varies with a fill level of the urine bladder, andrecording, by the processing circuit, an association between each determined parameter and an indication of whether the user has urinated on the absorbent article. 4. A method according to claim 1, further comprising determining, by the processing circuit, a parameter threshold based on the parameter represented by the recorded data. 5. A method according to claim 1, further comprising determining, by the processing circuit, a parameter threshold based on the parameter represented by the recorded data and a previously determined parameter threshold. 6. A method according to claim 4, further comprising: comparing, by the processing circuit, said parameter to the parameter threshold, andgenerating, by the processing circuit, an alert signal based on a result of the comparison. 7. A method according to claim 6, further comprising providing a visual, an audible or a tactile indication to the user in response to the alert signal. 8. A method according to claim 1, further comprising: estimating, by the processing circuit, an amount of urine released on to the urine sensitive circuit, and,in response to determining that the user has urinated on the absorbent article, recording by the processing circuit, data representing the estimated amount of urine released on to the urine sensitive circuit. 9. A method according to claim 8, wherein said amount of urine is estimated based on the changed electrical characteristics of the urine sensitive circuit. 10. A method according to any one of claims 1, further comprising: determining an orientation and/or a movement of the user using a sensor arranged on the user, andin response to determining that the urine sensitive circuit has been exposed to urine, recording, by the processing circuit, data representing an estimated movement and/or posture of the user based on the orientation and/or the movement determined by the sensor. ... |  | A61B005/20 A61B005/145 A61B005/01 A61B005/00 A61B005/11 A61B005/1477 A61F013/42
US20160135694 | A1 | 2016-05-19 | MEDICAL RADAR METHOD AND SYSTEM | 2014-06-06 | Radar is used to measure internal body motion. Radar reflections from the body are measured using a range of frequencies that includes a higher frequency band and a lower frequency band. In the higher frequency band, e.g. above 24 GHz, the radar signal hardly penetrates the skin, whereas it penetrates deeper into the chest in the lower frequency band e.g. below 10 GHz. Chest surface motion is estimated by means of the measurements using the higher frequency band. Effects of the estimated chest surface movement are subtracted from the measurements from the lower frequency band. The resulting response after removal is used to fit a model of a heart. Fitting may be performed in a series of fitting steps, including fitting parameters X of a geometric model to the measurements; determining a least square solution of fit errors between the measurements and a Taylor expansion from the grid model obtained with the fitted parameters X as a function of adaptions of the grid model; and subsequently determining a further adaptation of the grid model that best fits the measurements. | 1. A method of measuring internal body motion using radar, the method comprising measuring radar reflections using a range of frequencies that includes a higher frequency band and a lower frequency band;estimating chest surface motion using the measurements using the higher frequency band;removing effects of the estimated chest surface movement from the measurements from the lower frequency band. 2. A method according to claim 1, wherein the higher radar frequency band is a 24 Ghz band, 60 GHz band or 76 Ghz band and the lower radar frequency band is in a 2-10 Ghz band. 3. A method according to claim 1, wherein measuring the radar reflections comprises performing measurements in multiple discrete radar frequency bands. 4. A method according to claim 1, wherein the multiple radar frequency bands include a 2-10 Ghz band, a 24 Ghz band, a 60 GHz band and a 76 Ghz band. 5. A method according to claim 1, any one of the preceding claims, comprising fitting a model of a heart to the measurements. 6. A method according to claim 4, wherein said fitting comprises fitting parameters X of a geometric model to the measurements;determining a least square solution of fit errors between the measurements and a Taylor expansion from the grid model obtained with the fitted parameters X as a function of adaptions of the grid model;determining a further adaptation of the grid model that best fits the measurements. 7. A method according to claim 5, comprising displaying the model obtained by said fitting. 8. A method according to claim 1, selecting earliest received reflection peaks from the radar reflections of radar transmissions using the higher frequency band;estimating parameters of the selected reflections peaks;subtracting a modeled signal that corresponds to a peak with the estimated parameters from the measurements from the lower frequency band. 9. A method to claim 7, comprising tracking the received peaks in response to radar transmissions at successive time points and selecting the earliest peaks from peaks that lie along a closest detected track in a distance range of the radar. 10. A method to claim 7, comprising fitting a model of a heart to the result of said subtraction. 11. A method according to claim 10, wherein said fitting comprises using inverse modeling from the from the measurements from the lower one the frequency bands to which said subtracting has been applied to determine parameters of a model that defines locations of elements of a human heart as a function of the parameters. 12. A method according to claim 11, wherein said fitting comprises determining a least square solution of fit errors between the measurements from the lower one the frequency bands to which said subtracting has been applied and a linear expansion of the effect of displacements of locations of elements from the locations defined by said model using the parameters obtained by said inverse modeling;performing a search for displacements of locations of elements... |  | A61B005/02 A61B005/00
US10165988 | B2 | 2019-01-01 | Personalized detection system for detecting magnetic objects in the human organism | 2014-06-06 | The invention relates to a device with a sensor arrangement, which is able to detect magnetic or magnetized oral administration forms after oral take-up, moreover tracks the dissolution thereof via the reduction or disappearance of the magnetic field of the oral administration form, and with a log function, which records a subjective evaluation on the part of the human wearer of the sensor arrangement during or after the oral take-up. | 1. A device, comprising a detector system for registering magnetic bodies in the human organism, which detector system comprises:at least two sensor arrangements with an instrument for recording the magnetic flux density measured by each sensor arrangement, wherein each sensor arrangement has one, two or three anisotropic magnetoresistance sensors, the axes of easy magnetization of which point in pair-wise different directions, and each sensor arrangement has a distance of 0.5 to 50 cm from the remaining sensor arrangement or sensor arrangements, and at least two sensor arrangements are tilted at an angle of between 0 and 45° with respect to one another, anda log for registering a quantified evaluation before, during and/or after the registration of the magnetic body or bodies. 2. The device according to claim 1, wherein the sensor arrangements are integrated in at least one strap, the clothing and/or an item of jewelry, or worn or affixed directly on the body by means of a suction cup or fastening aid, andthe log is carried on the body. 3. The device according to claim 1, wherein the log is a mobile computer, a mobile telephone, PDA, small computer, data logger with transmitter, and/or an input unit, and the log is electronically connected to the sensor arrangement. 4. A method for registering the magnetic flux density generated by a magnetic body in the human organism by means of a device according to claim 1, comprising: (a) applying a set and reset pulse, at least once, to each anisotropic magnetoresi stance sensor,(b) amplifying the signals of each AMR sensor by means of suitable signal conditioning and by means of at least one low-pass filter,(c) determining and recording the difference in the magnitudes of the vectors of the magnetic flux densities of each sensor arrangement,and/or determining and recording the angle Φ between the vectors from the measurement signals of the AMR sensors,and, (d) simultaneously with one of the steps (a), (b) or (c) or after a time T has elapsed following step (c),registering the quantified evaluation, at least once, in the log, which is undertaken by the human wearer of the sensor arrangements. 5. The method according to claim 4, wherein the quantified evaluation is undertaken by virtue of the fact that a sense of well-being and/or physical fitness and/or at least one other subjective criterion, is assigned to alphanumerical signs or a grade scale, and entered into the log. 6. The method according to claim 5, wherein the values, obtained in step (c), for the difference and/or for the angle Φ are recorded as a function of time and the evaluation, registered in step (d), is recorded as a function of time. 7. A method for registering oral administration forms and determining the time or times of the disintegration of the magnetic component in the digestive tract comprising registering magnetic bodies in a human organism with the device according to claim 1. 8. The method according to claim 7, wherein the... | EVONIK ROEHM GmbH | A61B005/05 A61B005/00 G01R033/09 G01V003/08 A61J003/00
US10722301 | B2 | 2020-07-28 | Systems and methods for lesion assessment | 2015-11-03 | Ablation visualization and monitoring systems and methods are provided. In some embodiments, such methods comprise applying ablation energy to a tissue to form a lesion in the tissue, illuminating the tissue with a light to excite NADH in the tissue, wherein the tissue is illuminated in a radial direction, an axial direction, or both, monitoring a level of NADH fluorescence in the illuminated tissue to determine when the level of NADH fluorescence decreases from a base level in the beginning of the ablating to a predetermined lower level, and stopping ablation of the tissue when the level of NADH fluorescence reaches the predetermined lower level. | 1. A method for monitoring ablation of cardiac tissue of endocardium comprising: applying ablation energy to a cardiac tissue to form a lesion in the tissue;illuminating the cardiac tissue to excite nicotinamide adenine dinucleotide hydrogen (NADH) in the cardiac tissue and determining a base level of NADH fluorescence in the illuminated cardiac tissue;monitoring a level of NADH fluorescence in the illuminated cardiac tissue during the application of ablation energy to the illuminated cardiac tissue;reducing the NADH fluorescence in the illuminated cardiac tissue to between 60% and 80% of the base level and achieving a steady state NADH fluorescence; andupon achieving the steady state NADH fluorescence for between 5 and 10 seconds, stopping ablation of the illuminated cardiac tissue to limit an amount of ablation energy delivered to the cardiac tissue. 2. The method of claim 1 wherein the cardiac tissue is illuminated with light having a wavelength between about 300 nm and about 400 nm. 3. The method of claim 1 further comprising monitoring a level of the reflected light having a wavelength between about 450 nm and 470 nm. 4. The method of claim 1 wherein the ablation energy is selected from the group consisting of radiofrequency (RF) energy, microwave energy, electrical energy, electromagnetic energy, cryoenergy, laser energy, ultrasound energy, acoustic energy, chemical energy, thermal energy, electroporation energy, and combinations thereof. 5. The method of claim 1 further comprising advancing a catheter to the cardiac tissue, the catheter comprising: a catheter body;a distal tip positioned at a distal end of the catheter body for delivering ablation energy to the cardiac tissue, the distal tip defining an illumination cavity having one or more openings for exchange of light between the illumination cavity and the tissue; andone or more optical fibers extending through the catheter body into the illumination cavity of the distal tip, the one or more optical fibers being in communication with a light source and a light measuring instrument to illuminate the cardiac tissue and to relay light energy reflected from the cardiac tissue to the light measuring instrument. 6. The method of claim 1 further comprising illuminating the cardiac tissue in a radial direction and an axial direction with respect to a longitudinal axis of a catheter. 7. The method of claim 1 further comprising providing a real time visual feedback about the lesion formation by displaying the level of NADH fluorescence. 8. The method of claim 1, wherein the ablation energy is applied when a NADH fluorescence peak is detected. 9. The method of claim 1 further comprising performing an ultrasound evaluation of cardiac tissue in combination with monitoring the level of NADH fluorescence. 10. The method of claim 1 further comprising monitoring a level of the reflected light having a wavelength between about 375 nm and about 575 nm. 11. The method of claim 1, further compri... | The George Washington University 460Medical, Inc. | A61B018/14 A61B005/00 A61B018/12 A61B018/00 A61B090/00 A61B005/06
US10433816 | B2 | 2019-10-08 | Method and system for manipulating medical device operating parameters on different levels of granularity | 2014-11-20 | Various embodiments include a system and method for manipulating medical device operating parameters, stored in a hierarchical organization, at different levels of granularity. The method can include receiving, at a processor, first user input selecting an operating parameter of a medical device from a medical device user interface displayed at a display system. The method may include presenting, at the display system, a parameter presentation prompted by the operating parameter selection. In various embodiments, a plurality of setting values each corresponding to a different hierarchical level of the operating parameter is presented in the parameter presentation. The method may include receiving, at the processor, second user input manipulating a value of one of the setting values presented in the parameter presentation. The method can also include storing, at storage, the value of the manipulated setting value. | 1. A method, comprising: receiving, at a processor, first user input selecting an operating parameter of a medical device from a medical device user interface displayed at a display system, wherein the operating parameter is organized in a hierarchy of a plurality of different hierarchical levels including a higher hierarchical level and a lower hierarchical level that is dependent on the higher hierarchical level, and wherein the operating parameter is executable by the medical device to perform an examination using the medical device;presenting, at the display system, a parameter presentation prompted by the operating parameter selection, wherein a plurality of setting values each corresponding to a different one of the plurality of different hierarchical levels of the operating parameter is simultaneously presented in the parameter presentation, and wherein a lower hierarchical level setting value of the plurality of setting values presented in the parameter presentation is equal to or within a higher hierarchical level setting value of the plurality of setting values presented in the parameter presentation;receiving, at the processor, second user input manipulating a value of one of the plurality of setting values presented in the parameter presentation;storing, at storage, the value of the manipulated one of the plurality of setting values; andexecuting, by the medical device, the stored value of the manipulated one of the plurality of setting values during operation of the medical device. 2. The method according to claim 1, comprising, prior to receiving the first user input: receiving, at the processor, third user input selecting a medical device examination type from the medical device user interface displayed at the display system; andpresenting, at the display system, the operating parameters corresponding with the selected medical device examination type. 3. The method according to claim 2, comprising executing the selected medical device examination type. 4. The method according to claim 1, comprising automatically updating at least one other value of the plurality of setting values presented in the parameter presentation based on the second user input to maintain consistency between the plurality of setting values at the plurality of different hierarchical levels of the operating parameter. 5. The method according to claim 4, wherein the value of the manipulated one of the plurality of setting values is at the lower hierarchical level of the operating parameter, wherein the at least one other value of the plurality of setting values is at the higher hierarchical level of the operating parameter, and wherein the at least one other value of the plurality of setting values is broadened to include the value of the manipulated one of the plurality of setting values. 6. The method according to claim 4, wherein the value of the manipulated one of the plurality of setting values is at the higher hierarchical level of the operating param... | GENERAL ELECTRIC COMPANY | A61B006/00 A61B008/00 A61B008/13 A61B005/055 A61B006/03 G16H040/63 A61B005/00 G06F019/00
US09731126 | B2 | 2017-08-15 | Transcutaneous electrical nerve stimulator with automatic detection of leg orientation and leg motion for enhanced sleep analysis, including enhanced transcutaneous electrical nerve stimulation (TENS) using the same | 2015-12-28 | Apparatus for providing transcutaneous electrical nerve stimulation (TENS) therapy to a user, the apparatus comprising: a housing; an application unit for providing mechanical coupling between the housing and the user's body; a stimulation unit for electrically stimulating at least one nerve of the user; a sensing unit for sensing the user's body movement and body orientation; and a reporting unit for providing the user with feedback based on the user's sensed body movement and body orientation. | 1. Apparatus for providing transcutaneous electrical nerve stimulation (TENS) therapy to a user, said apparatus comprising: a housing;an application unit for providing mechanical coupling between said housing and the user's body;a stimulation unit for electrically stimulating at least one nerve of the user;a sensing unit for (i) sensing the user's body movement and body orientation to determine whether the user is in a sleep state or a wake state, and (ii) analyzing the sleep state of the user; anda feedback unit for at least one of (i) providing the user with feedback in response to the analysis of the sleep state of the user, and (ii) modifying the electrical stimulation provided to the user in response to the analysis of the sleep state of the user. 2. Apparatus according to claim 1 wherein said application unit is a flexible band. 3. Apparatus according to claim 1 wherein said application unit communicates whether said housing is mechanically coupled to the user's body. 4. Apparatus according to claim 3 wherein a mechanical element determines whether said housing is mechanically coupled to the user's body. 5. Apparatus according to claim 4 wherein said mechanical element is a tension gauge. 6. Apparatus according to claim 3 wherein an electrical element determines whether said housing is mechanically coupled to the user's body. 7. Apparatus according to claim 6 wherein said electrical element measures the voltage value of a voltage divider circuit formed by said housing and the user's body. 8. Apparatus according to claim 1 wherein said sensing unit uses data from an electromechanical sensor. 9. Apparatus according to claim 8 wherein whether said housing is mechanically coupled to the user's body determines the usability of the data from said electromechanical sensor. 10. Apparatus according to claim 8 wherein said electromechanical sensor is an accelerometer. 11. Apparatus according to claim 1 wherein said sensing unit determines the user's body orientation state with an analysis unit operating on earth gravitational acceleration measurements. 12. Apparatus according to claim 11 wherein the user's body orientation comprises one from the group consisting of upright and recumbent. 13. Apparatus according to claim 12 wherein said analysis unit is a programmable microprocessor that averages projections of the gravitational acceleration measurements on an axis of said electromechanical sensor and compares the average with a threshold. 14. Apparatus according to claim 13 wherein said axis is aligned with the longitudinal direction of the user's body. 15. Apparatus according to claim 13 wherein said threshold is one half of the standard earth gravitational acceleration value. 16. Apparatus according to claim 13 wherein said average is calculated over a period of one-minute. 17. Apparatus according to claim 13 wherein the user's body orientation is determined to be recumbent when said average is below said threshold. 18. Apparatus a... | Neurometrix, Inc. | A61N001/36 A61N001/04 A61B005/11 A61B005/00 A61N001/32
US09348976 | B2 | 2016-05-24 | Information processing method and electronic device | 2014-03-29 | An information processing method and an electronic device are provided. The method is applied to the electronic device. The electronic device includes a frame structure, a display module, a processor and a first type of sensor. The method includes: obtaining a first type of data which reflects a physical performance of the first user and is captured by the first type of sensor, in the case that the electronic device is fixed on an operation portion of the first user through the frame structure and the electronic device is functioning; obtaining configuration information at least based on the first type of data; and automatically setting a second type of sensor different from the first type of sensor, and/or a display screen, and/or at least one piece of application data which may be fed back by the electronic device, to suit the physical performance of the first user. | 1. An information processing method applied to an electronic device, wherein: the electronic device includes a frame structure, a display module, a processor, and a first type of sensor;the frame structure includes a fixing structure that fixes the electronic device on an operation portion of a first user;the display module includes a display screen and is fixed on the frame structurethe first type of sensor is fixed through the frame structure;the processor is connected with the first type of sensor and is fixed through the frame structure;the method comprises: obtaining a first type of data that reflects a physical performance of the first user and is captured by the first type of sensor, in a case that (i) the electronic device is fixed on the operation portion of the first user through the frame structure and (ii) the electronic device is functioning;obtaining configuration information at least based on the first type of data; andautomatically setting, based on the configuration information, a second type of sensor different from the first type of sensor, and/or the display screen, and/or at least one piece of application data that can be fed back by the electronic device, to suit the physical performance of the first user;the electronic device further includes a third type of sensor different from the first type of sensor;the method further includes, before obtaining the configuration information, obtaining a second type of data reflecting a location of the electronic device through the third type of sensor, andobtaining the configuration information includes: searching for first configuration information in a preset area around the location of the electronic device based on the second type of data; andselecting, from the first configuration information, the configuration information corresponding to the first type of data. 2. The method according to claim 1, wherein obtaining the configuration information comprises at least one of: obtaining, at least based on the first type of data, the configuration information through cloud matching; andobtaining, at least based on the first type of data, the configuration information through another electronic device bound to the electronic device. 3. The method according to claim 1, wherein the first type of data is a heart rate and/or a temperature of the first user. 4. An electronic device comprising: a frame structure, wherein the frame structure comprises a fixing structure that fixes the electronic device on an operation portion of a first user;a display module fixed on the frame structure, wherein the display module comprises a display screen;a first type of sensor fixed through the frame structure;a processor connected with the first type of sensor and fixed through the frame structure, wherein the processor comprises: a data obtaining unit configured to obtain a first type of data that reflects a physical performance of the first user and is captured by the first type of sensor, in a case th... | Beijing Lenovo Software Ltd. Lenovo (Beijing) Co., Ltd. | G08B001/08 G06F019/00 A61B005/145 A61B005/021 A61B005/024
US20160143535 | A1 | 2016-05-26 | SYSTEMS TO MONITOR BODY PORTIONS FOR INJURY AFTER IMPACT | 2014-11-21 | Systems are described for monitoring extremities for injury or damage following a physical impact. A device embodiment includes, but is not limited to, a deformable substrate integrated with a textile; a sensor assembly coupled to the deformable substrate, the sensor assembly configured to generate one or more sense signals based on detection of a physical impact to a body portion and based on detection of a physiological parameter; circuitry operably coupled to the sensor assembly and configured to receive the one or more sense signals based on detection of the physical impact and to determine whether the physical impact exceeds a threshold impact value, the circuitry configured to instruct the sensor assembly to detect one or more physiological parameters of the body portion when the physical impact exceeds the threshold impact value; and a reporting device operably coupled to the circuitry. | 1. A device, comprising: a deformable substrate integrated with a textile configured to interface with a body portion;a sensor assembly coupled to the deformable substrate, the sensor assembly configured to generate one or more sense signals based on detection of a physical impact to the body portion and based on detection of a physiological parameter of the body portion;circuitry operably coupled to the sensor assembly and configured to receive the one or more sense signals based on detection of a physical impact to the body portion and to determine whether the physical impact exceeds a threshold impact value, the circuitry configured to instruct the sensor assembly to detect one or more physiological parameters of the body portion when the physical impact exceeds the threshold impact value; anda reporting device operably coupled to the circuitry and configured to generate one or more communication signals based on instruction by the circuitry. 2. The device of claim 1, wherein the textile includes an inner surface and an outer surface, wherein at least a portion of the deformable substrate is positioned between the inner surface and the outer surface. 3. The device of claim 1, wherein the textile includes an inner surface and an outer surface, wherein the deformable substrate is attached to one of the inner surface or the outer surface. 4. The device of claim 1, wherein the deformable substrate is integral to a weave of the textile. 5. The device of claim 1, wherein one or more fiber of the textile secures the deformable substrate to at least a portion of the textile. 6. The device of claim 1, wherein the textile includes one or more electronic threads. 7.-10. (canceled) 11. The device of claim 1, wherein the circuitry includes a comparison module configured to compare the one or more sense signals based on detection of a physical impact to the body portion to reference data indicative of the threshold impact value. 12. The device of claim 11, wherein the reference data indicative of the threshold impact value correlates to a threshold for a traumatic injury. 13. The device of claim 1, wherein the sensor assembly includes a pressure sensor configured to detect the physical impact to the body portion and to generate one or more sense signals responsive to the detection of the physical impact. 14. The device of claim 13, wherein the circuitry is configured to instruct the sensor assembly to detect one or more physiological parameters of the body portion based upon the one or more sense signals generated by the pressure sensor. 15. The device of claim 1, wherein the sensor assembly includes a proximity sensor configured to detect the physical impact to the body portion and to generate one or more sense signals responsive to the detection of the physical impact. 16. The device of claim 15, wherein the circuitry is configured to instruct the sensor assembly to detect one or more physiological parameters of the body portion based upon the... |  | A61B005/00
US09940810 | B2 | 2018-04-10 | Person support apparatuses with patient mobility monitoring | 2015-10-30 | A person support apparatus, such as a bed, stretcher, cot, or the like, includes a fall detector. The fall detector is adapted to detect when a person associated with the person support apparatus has fallen and to issue a fall alarm. In some embodiments, the person support apparatus also includes an exit detection system that issues an exit alarm when the person exits from the person support apparatus. The fall alarm is given a higher priority than the exit alarm. The fall detector may include a camera, a thermal image sensor, a device worn by the person, or another sensor. The person support apparatus may also, or alternatively, include a timer for measuring how long an occupant remains out of the person support apparatus. | 1. A person support apparatus comprising: a support surface adapted to support a person thereon;an exit detector adapted to issue an alarm if the person exits the support surface;a fall detection sensor integrated into the person support apparatus, the fall detection sensor adapted to detect if the person falls; anda controller in communication with the exit detector and the fall detection sensor, the controller adapted to issue a first alarm if the person exits the support surface and to issue a second alarm if the person falls, the second alarm having a higher priority than the first alarm, and wherein both the first and second alarms are communicated by the controller to a remote location. 2. The person support apparatus of claim 1 wherein the exit detector comprises a plurality of load cells positioned to support the support surface. 3. The person support apparatus of claim 1 wherein the first alarm includes both a visual and an audio component. 4. The person support apparatus of claim 1 wherein the fall detection sensor comprises an image sensor adapted to detect images of the person after he or she exits from the person support apparatus. 5. The person support apparatus of claim 4 wherein the image sensor is a thermal image sensor. 6. The person support apparatus of claim 1 wherein the remote location includes a nurse call system, and wherein the person support apparatus further comprises a cable interface adapted to releasably couple to a cable of the nurse call system, the cable interface including a plurality of electrical connectors, and wherein the controller is adapted to open or close a first switch in communication with a first one of the electrical connectors when issuing the first alarm and to open or close a second switch in electrical communication with a second one of the electrical connectors when issuing the second alarm. 7. The person support apparatus of claim 1 wherein the remote location includes a local area network and the person support apparatus further comprises a wireless interface adapted to wirelessly communicate with the local area network, and the controller is further adapted to send a first message to the local area network via the wireless interface when issuing the first alarm, and to send a second message to the local area network via the wireless interface when issuing the second alarm. 8. The person support apparatus of claim 1 wherein the person support apparatus is one of a bed, stretcher, and chair. 9. A person support apparatus system comprising: a first person support apparatus comprising a first support surface, an exit detector, and a first controller;a second person support apparatus comprising a second support surface, a fall detection sensor integrated into the second person support apparatus, and a second controller; andwherein the first person support apparatus is in communication with the second person support apparatus and the first controller is adapted to issue a first alarm if a p... | Stryker Corporation | G08B023/00 A61B005/00 G08B021/04 A61B005/11 G08B021/22 G08B001/08 G06F011/00
US09339240 | B2 | 2016-05-17 | System and method of resolving outliers in NIRS cerebral oximetry | 2014-01-08 | A system and method for non-invasively estimating the tissue blood oxygen saturation level of a human subject, including so-called “outliers”, whose physiological make-up causes previously-known techniques to generate invalid tissue blood oxygen saturation estimations. The system includes a computing device and a sensor. The sensor includes a light source configured to emit light of at least four different wavelengths, one at a time. The sensor also includes two light detectors, each positioned a different distances from the light source. Optical density measurements are taken by the light detectors and provided to the computing device. A first tissue blood oxygen saturation value is computed using the optical density measurements associated with three of the four wavelengths, and a second tissue blood oxygen saturation value is computed using the optical density measurements associated with four of the wavelengths. The first and second tissue blood oxygen saturation values are compared, and the human subject is identified as an “outlier” based upon that comparison. | 1. A system for assessing tissue blood oxygen saturation levels, comprising: a computing device;a sensor in communication with said computing device, said sensor being configured to be attached to a human subject and having: a light source capable of selectively emitting a plurality of wavelengths of light, one wavelength at a time; anda first light detector positioned a first distance from said light source and a second light detector positioned a second distance from said light source, said second distance being greater than said first distance;said computing device having a memory for storing an algorithm and a processor for executing instructions associated with said algorithm, said algorithm comprising the following steps:causing said light source to emit light of a plurality of wavelengths, one wavelength at a time, into a human subject;receiving optical density measurements from said first and second light detectors for said plurality of wavelengths;calculating a space contrasted ratio of the wavelength contrast difference of the optical density based on the optical density measurements from said first and second light detectors for a first grouping of wavelengths emitted by said light source;estimating a first tissue blood oxygen saturation value based on said space contrasted ratio of the wavelength contrast difference of the optical density;calculating a second order contrasted optical density ratio based on the optical density measurements from said first and second light detectors for a second grouping of said wavelengths emitted by said light source;estimating a second tissue blood oxygen saturation value based on said second order contrasted optical density ratio; andcomparing said first blood oxygen saturation value and said second blood oxygen saturation value. 2. The system of claim 1, wherein said first tissue blood oxygen saturation value is determined to be a valid estimate of said actual tissue blood oxygen saturation level if said first tissue blood oxygen saturation value is about the same as said second tissue blood oxygen saturation value. 3. The system of claim 1, wherein said first tissue blood oxygen saturation value is determined to be an invalid estimate of said actual tissue blood oxygen saturation level if said first tissue blood oxygen saturation value is at least about 10% or more different than said second tissue blood oxygen saturation value. 4. The system of claim 1, wherein said space contrasted ratio of the wavelength contrast difference of the optical density is calculated by applying the following equation: OD′Deep-Shall(λ1)/OD′Deep-Shall(λ2)where OD′(λ)=OD(λ)−OD(λ+Δ) is the wavelength contrasted optical density that can be described as a wavelength contrast difference of the optical density OD(λ); and ODDeep-Shall(λ)=ODDeep(λ)−ODShallow(λ) is the spatial contrasted optical density that is the difference of the optical density measured by the far detector ODDeep(λ) and the near detector ODShallow(λ) at ... | Covidien LP | A61B005/1455 A61B005/00
US09351650 | B2 | 2016-05-31 | Image processing apparatus and image processing method | 2014-02-18 | An image processing apparatus obtains an image of an eye area, determines a measurement target among multiple vascular branches based on information regarding multiple vascular branches that include multiple vascular bifurcations in the obtained image, and measures the size of a blood cell aggregate in the determined measurement target. | 1. An image processing apparatus comprising: an image obtaining unit configured to obtain an image of an eye area;an information obtaining unit configured to obtain information regarding a plurality of vascular branches that include a plurality of vascular bifurcations in the obtained image;a determining unit configured to determine a measurement target from the plurality of vascular branches based on the obtained information; anda measuring unit configured to measure the size of a blood cell aggregate in the determined measurement target. 2. The apparatus according to claim 1, wherein the information obtaining unit obtains at least one value out of a length and a diameter of the plurality of vascular branches as the information regarding the plurality of vascular branches, and wherein the determining unit determines the measurement target from the plurality of vascular branches based on the obtained value. 3. The apparatus according to claim 2, further comprising: a vascular branch specifying unit configured to specify the plurality of vascular branches using a region that includes two neighboring vascular bifurcations out of the plurality of vascular bifurcations and a blood vessel between the two vascular bifurcations as one vascular branch,wherein the information obtaining unit obtains at least one out of the lengths and the vascular diameters of the specified plurality of vascular branches. 4. The apparatus according to claim 2, wherein the determining unit determines a vascular branch among the plurality of vascular branches whose vascular diameter is within a predetermined range as the measurement target. 5. The apparatus according to claim 2, wherein the image obtaining unit obtains a moving image of the eye area as the image, and wherein the measuring unit measures change in the size of the blood cell aggregate based on the obtained value. 6. The apparatus according to claim 5, further comprising: a vascular region specifying unit configured to specify a vascular region from the obtained moving image based on change in luminance values in the obtained moving image; anda vascular bifurcation specifying unit configured to specify the plurality of vascular bifurcations from the specified vascular region. 7. The apparatus according to claim 6, further comprising: a region-of-interest specifying unit configured to specify a region of interest from the obtained moving image based on the specified vascular region and a parafoveal region in the eye area,wherein the vascular bifurcation specifying unit specifies the plurality of vascular bifurcations from the specified region of interest,wherein the information obtaining unit furthermore obtains information regarding a luminance value in the specified region of interest, andwherein the determining unit determines the measurement target from the specified plurality of vascular branches based on the obtained value and the information regarding the luminance value. 8. The apparatus according ... | Canon Kabushiki Kaisha | A61B003/14 A61B003/00 A61B005/026 G06K009/00 A61B003/10 A61B003/12
US20160135699 | A1 | 2016-05-19 | REDUCING MOVEMENT AND ELECTROSTATIC INTERFERENCE IN A NON-RESISTIVE CONTACT SENSOR ASSEMBLY | 2015-10-10 | A non-resistive contact sensor assembly, includes a non-resistive contact sensor device, a first inner region of material in which the sensor device is surrounded or embedded, a boundary region, around the first inner region, that is at least partly comprised of a conducting material and that picks up electrical charge from the first inner region, a second inner region around the boundary region, an outer region around the second inner region, and at least one diode, disposed in the second inner region and connected between the boundary region through the second inner region to the outer region. | 1. A non-resistive contact sensor assembly, comprising: (a) a non-resistive contact sensor device;(b) a first inner region of material in which the sensor device is surrounded or embedded;(c) a boundary region, around the first inner region, that is at least partly comprised of a conducting material and that picks up electrical charge from the first inner region;(d) a second inner region around the boundary region;(e) an outer region around the second inner region; and(f) at least one diode, disposed in the second inner region and connected between the boundary region through the second inner region to the outer region. 2. The non-resistive contact sensor assembly of claim 1, wherein at least one of the first inner region, the second inner region, and the outer region includes a triboelectrically neutral or relatively neutral material. 3. The non-resistive contact sensor assembly of claim 1, wherein at least one of the first inner region, the second inner region, and the outer region includes an elastic material that acts to hold the sensor device in place, thereby minimizing the likelihood of sensor movement against the object surface and leading, in turn, to less chance of contact electrification and movement artifacts. 4. The non-resistive contact sensor assembly of claim 1, wherein at least one of the first inner region, the second inner region, and the outer region includes a biasing structure that acts to hold the sensor device in place, thereby minimizing the likelihood of sensor movement against the object surface and leading, in turn, to less chance of contact electrification and movement artifacts. 5. The non-resistive contact sensor assembly of claim 4, wherein the biasing structure includes a spring. 6. The non-resistive contact sensor assembly of claim 1, wherein at least one of the first inner region, the second inner region, and the outer region includes material that is relatively triboelectrically matched to the surface of the object being measured, tested, or the like, thereby minimizing the chances that charge separation will occur when this region and the surface of the object may move together. 7. The non-resistive contact sensor assembly of claim 1, wherein at least one of the first inner region, the second inner region, and the outer region includes material that is conductive, allowing for dissipation of any surface or external charge. 8. The non-resistive contact sensor assembly of claim 1, further comprising a housing that is treated with an ionized liquid. 9. A non-resistive contact sensor assembly, wherein the surface region of the entity where a sensor or sensors are housed may be separated from the rest of the surface of the entity using a conducting band, comprising: (a) a non-resistive contact sensor device, including a dry electrode component;(b) a first region of material; and(c) a second region of material, wherein at least part of the material is a conducting material;(d) wherein the first region is wal... |  | A61B005/04 A61B005/00
US09968688 | B2 | 2018-05-15 | Shielded targeting agents, methods, and in vivo diagnostic system | 2014-11-12 | A system is provided which includes nanoparticle conjugates configured to bind with a tumor cell, the nanoparticle conjugate comprising a nanoparticle, at least one targeting entity bound to the nanoparticle, and at least one shielding entity that shields at the at least one targeting entity, the nanoparticle, or both; a body-mountable device mounted on an external surface of a living body and configured to detect a tumor cell binding response signal transmitted through the external surface, wherein the tumor cell binding response signal is related to binding of the nanoparticle conjugates with one or more tumor cells; and a processor configured to non-invasively detect the one or more tumor cells based on the tumor cell response signal. Nanoparticle conjugates and methods for use for treating or imaging tumor cells are also provided. | 1. A shielded nanoparticle conjugate comprising: a nanoparticle;at least one targeting entity for binding to tumor cells or tissue in vivo, the at least one targeting entity bound to the nanoparticle; andat least one shielding entity for shielding the at least one targeting entity and for unshielding from the at least one targeting entity in a microenvironment of the tumor cells or tissue in vivo, the at least one shielding entity is selected from a protease sensitive bivalent peptide or a bivalent anti-idiotypic aptamer. 2. The shielded nanoparticle conjugate of claim 1, wherein the nanoparticle comprises a polymer material. 3. The shielded nanoparticle conjugate of claim 1, wherein the nanoparticle comprises a non-polymeric material. 4. The shielded nanoparticle conjugate of claim 1, wherein the nanoparticle comprises a magnetic or paramagnetic material. 5. The shielded nanoparticle conjugate of claim 1, wherein the at least one targeting entity comprises an antibody, peptide, protein, nucleic acid, small molecule, carbohydrate, or lipid. 6. The shielded nanoparticle conjugate of claim 1, wherein the bivalent anti-idiotypic aptamer is directed to a tumor metabolite. 7. The shielded nanoparticle conjugate of claim 1, wherein the at least one targeting entity is an antibody and the at least one shielding entity is a protease-sensitive bivalent peptide that binds to the antigen binding site of the antibody. 8. The shielded nanoparticle conjugate of claim 1, wherein the at least one targeting entity is an antibody and the at least one shielding entity is a bivalent anti-idiotypic aptamer with a tumor metabolite-binding domain, wherein the bivalent anti-idiotypic aptamer binds to the antigen binding site of the antibody. 9. The shielded nanoparticle conjugate of claim 1, further comprising at least one detection label. 10. The shielded nanoparticle conjugate of claim 1, further comprising at least one agent to be delivered to the tumor cells. 11. The shielded nanoparticle conjugate of claim 10, wherein the at least one agent comprises at least one anti-tumor agent. 12. The shielded nanoparticle conjugate of claim 10, wherein the at least one agent comprises a contrast imaging agent. 13. The shielded nanoparticle conjugate of claim 1, wherein said shielded nanoparticle conjugate comprises at least one moiety that exhibits fluorescence, luminescence, magnetic or paramagnetic properties. 14. A method for in vivo imaging in a mammal of tumor cells or tissue that express a selected marker; said method comprising the steps of: (a) administering to the mammal a composition as recited in claim 1, wherein the at least one targeting entity targets a selected marker of tumor cells or tissue;(b) allowing the at least one targeting entity to bind to the selected marker of the tumor cells or tissue in vivo to form a conjugate; and(c) imaging the tumor cells or tissue with a non-invasive imaging technique that has a resolution enhanced by a presence o... | Verily Life Sciences LLC | A61B005/05 A61K047/48 A61K049/00 A61K047/30 A61M005/00 A61B005/00 G01R033/56 A61B005/145 A61K049/18 A61K051/12 A61B008/08 A61B006/00
US09339257 | B2 | 2016-05-17 | Measuring apparatus and method thereof | 2013-11-22 | A measuring apparatus is provided. The measuring apparatus includes a measurement reference part setting unit configured to set-measurement reference parts to each of a plurality of sequential ultrasound images of a subject, a tracking unit configured to perform tracking of a plurality of regions in the ultrasound images, a first motion information calculating unit configured to calculate first motion information of the regions, based on the tracking performed by the tracking unit, and a second motion information calculating unit configured to calculate second motion information of the measurement reference parts, based on the first motion information, wherein the plurality of regions move with the measurement reference parts and include parts capable of capturing moving factors that are not able to be captured by tracking the measurement reference parts between a first ultrasound image and a second ultrasound image using the tracking unit. | 1. A measuring apparatus comprising: a measurement reference point setting unit configured to set a first reference point to each of a plurality of sequential ultrasound images of a subject;a region of interest setting unit configured to set a first region of interest to a first part including the first reference point, and to set a plurality of second regions of interest to a plurality of second parts in each of the plurality of sequential ultrasound images, where the plurality of second regions of interest are set to both sides of the first region of interest, and where each of the plurality of second regions of interest does not include the first reference point;a tracking unit configured to perform tracking of the first region of interest and the plurality of second regions of interest in the ultrasound images;a first motion information calculating unit configured to determine a first moving vector indicative of movement of the first part and a plurality of second moving vectors, where each of the plurality of second moving vectors is indicative of movement of one of the plurality of second regions of interest based on the tracking performed by the tracking unit;a second motion information calculating unit configured to calculate an average vector based on the first moving vector and the plurality of second moving vectors; anda display unit configured to display a graph indicative of movement of the first reference point based on the average vector. 2. The measuring apparatus according to claim 1, wherein the plurality of second regions of interest are set along a vascular wall. 3. The measuring apparatus according to claim 1, wherein the region of interest setting unit is further configured to identify a vascular wall in the ultrasound image and set the plurality of second regions of interest to the vascular wall. 4. A measuring method for tracking movement of parts, comprising: setting a first reference point to each of a plurality of sequential ultrasound images with a measurement reference point setting unit;setting a first region of interest to a first part, the first part including the first reference point;setting a plurality of second regions of interest to a plurality of second parts in each of the plurality of sequential ultrasound images, where the plurality of second regions of interest are set to both sides of the first region of interest, where each of the plurality of second regions of interest does not include the first reference point;tracking the first region of interest and the plurality of second regions of interest with a tracking unit;determining, with a first motion information calculating unit, a first moving vector indicative of movement of the first part based on the tracking;determining, with the first motion information calculating unit, a plurality of second moving vectors, where each of the plurality of second moving vectors is indicative of movement of one of the plurality of second regions of interest based ... | GENERAL ELECTRIC COMPANY | A61B008/08 G06T007/20 A61B005/026
US09345405 | B2 | 2016-05-24 | Electrophysiological mapping system using external electrodes | 2011-09-14 | A system and method for generating an electrophysiological map are provided. The system includes an electronic control unit (ECU) configured to receive a signal generated by an electrode disposed at a position on an external surface of the body and indicative of electric potential. The ECU is further configured to identify a surface boundary of an object of interest within the body using an image of the object. The ECU is further configured to identify intervening objects along a pathway between the position on the external surface and the surface boundary of the object of interest from one or more images of the pathway. The ECU is further configured to obtain an impedance value for each of the intervening objects and to determine an electric potential at the surface boundary of the object of interest responsive to the signal from the electrode and the impedance values of the intervening objects. | 1. A system for generating an electrophysiological map, comprising: a first electrode disposed at a first position on an external surface of a body;an imaging system configured for imaging an object of interest within said body;an electronic control unit electrically connected to said first electrode and the imaging system, the electronic control unit configured to: receive a signal generated by said first electrode, said signal indicative of an electric potential measured at said first position;identify a surface boundary of the object of interest using an image of said object of interest;obtain, from the imaging system, a tissue density for a first intervening object located within the body;calculate a first impedance value for the first intervening object from the obtained tissue density along a first pathway between said first position on said external surface of said body and said surface boundary of said object of interest; anddetermine an electric potential at said surface boundary of said object of interest responsive to said signal from said first electrode and said first impedance value of said first intervening object along said first pathway. 2. The system of claim 1 wherein said electronic control unit is further configured to identify said first intervening object along said first pathway from one or more images of said first pathway. 3. The system of claim 2 wherein said images of said first pathway comprise magnetic resonance images or computed tomography images. 4. The system of claim 1, wherein the electronic control unit further comprises a input/output interface configured to receive said first impedance value for said first intervening object along said first pathway. 5. The system of claim 1 wherein said electronic control unit is further configured to obtain a second impedance value for a second intervening object along said first pathway and to determine said electric potential at said surface boundary of said object of interest responsive to said second impedance value. 6. The system of claim 1 wherein said electronic control unit is further configured to obtain a second impedance value for a second intervening object along said first pathway, to obtain a total impedance value responsive to said first and second impedance values and to determine said electric potential at said surface boundary of said object of interest responsive to said total impedance value. 7. The system of claim 1, further comprising a display configured to generate an electrophysiological map of said object of interest responsive to signals generated by said electronic control unit. 8. The system of claim 1 wherein said electronic control unit is further configured to compare said electric potential measured at said first position to a previously acquired electric potential measured at said first position. 9. The system of claim 1 wherein said electronic control unit is further configured, in obtaining said first impedance value for said fir... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/00 A61B005/053 A61B005/0402 G06K009/00 A61B005/0408
US09326699 | B2 | 2016-05-03 | Methods for using electrical impedance tomography | 2013-04-16 | Methods for use of EIT. Disclosed are: (1) EIT used to obtain a final solution to an EIT inverse problem for localizing tissues undergoing changes in impedance, which is used as a constraint on solving an EEG source localization inverse problem; (2) EIT used with MREIT, where the MREIT is used to constrain the solutions to the EIT inverse problem for the distribution of static tissue impedance; (3) EIT used with MREIT, where the MREIT is used to constrain the solutions to the EIT inverse problem for localizing tissues undergoing changes in impedance; and (4) EIT according to any of (1)-(3) as feedback for modifying at least one of the location, magnitude, and timing of currents injected for the purpose of neurostimulation. | 1. A method for use of electrical impedance tomography, hereinafter referred to as EIT, comprising: performing an EIT procedure, including distributing a plurality of electrodes on a subject's head such that each electrode makes electrical contact with the head surface at a respective location that is spaced apart from the location associated with each other electrode, and performing an EIT current injecting step by injecting a first set of currents into the head through a first set of said electrodes, each of the currents of said first set of currents being of sufficiently small magnitude so as not to provoke a substantial brain response, either alone or in combination with other currents of said first set of currents, and, during a time when said first set of electrodes is injecting one or more of the currents of said first set of currents, utilizing a remaining second set of said electrodes to measure respective electrical potentials arising in response;performing a magnetic resonance electrical impedance tomography, hereinafter referred to as MREIT, current injecting step by injecting a second set of currents into the subject's head using the same first set of electrodes used for said EIT current injecting step; andobtaining a magnetic resonance image of said second currents, the method further comprising obtaining one or more EIT solutions to an EIT inverse problem, posed by said electrical potentials and said first set of currents, for either or both (a) static head tissue impedances, or (b) head tissues having changes in impedance, said obtaining including constraining one or more of said one or more EIT solutions by said magnetic resonance image. 2. The method of claim 1, limited to EIT solutions of type (a). 3. The method of claim 2, wherein said electrical potentials are measured simultaneously with the MREIT current injecting step. 4. The method of claim 3, wherein said step of constraining includes evaluating potential instances of said one or more EIT solutions for a Bayesian conditional probability therefor, the condition of said probability being based on said magnetic resonance image. 5. The method of claim 2, wherein said step of constraining includes evaluating potential instances of said one or more EIT solutions for a Bayesian conditional probability therefor, the condition of said probability being based on said magnetic resonance image. 6. The method of claim 1, limited to EIT solutions of type (b). 7. The method of claim 6, wherein said electrical potentials are measured simultaneously with the MREIT current injecting step. 8. The method of claim 7, wherein said step of constraining includes evaluating potential instances of said one or more EIT solutions for a Bayesian conditional probability therefor, the condition of said probability being based on said magnetic resonance image. 9. The method of claim 6, wherein said step of constraining includes evaluating potential instances of said one or more EIT solutions fo... | Electrical Geodesics, Inc. | A61B005/05 A61B005/053 A61B005/0476 G01R033/32 G01R033/48 A61B005/055
US10349880 | B2 | 2019-07-16 | Safety blood collection assembly with indicator | 2016-01-20 | A needle assembly is disclosed. The needle assembly includes a housing having a flash chamber, and having a distal end and a proximal end engageable with a specimen collection container. The assembly includes a cannula having a patient end, a non-patient end, and a sidewall extending therebetween defining a cannula interior. The patient end of the cannula projects at least partially from the distal end of the housing, and the cannula interior is in fluid communication with the flash chamber. The assembly further includes a shield restrainably engaged with a portion of the housing and axially transitionable over the patient cannula from a retracted position in which the patient end is exposed, to an extended position in which the patient end is shielded by at least a portion of the shield, wherein at least a portion of the flash chamber is visible in the retracted position. | 1. A needle assembly, comprising: a housing having a distal portion and a proximal portion engageable with a specimen collection container;a cannula having a patient end, a non-patient end, and a sidewall extending therebetween defining a cannula interior, the patient end of the cannula projecting at least partially from the distal portion of the housing; anda shield restrainably engaged with a portion of the housing and axially transitionable over the patient end of the cannula from a retracted position in which the patient end is exposed, to an extended position in which the patient end is shielded by at least a portion of the shield, wherein the shield comprises a first portion for slideably engaging an outer surface portion of the housing and a second portion at least partially surrounding a portion of the cannula in the extended position, wherein the second portion of the shield comprises a depending arm transitionable from a first position in which the depending arm is substantially perpendicular to a longitudinal axis of the cannula, to a second position in which the depending arm is oriented substantially along the longitudinal axis of the cannula. 2. The needle assembly of claim 1, wherein the depending arm comprises a first portion and a second portion, the first portion and the second portion pivotally or hingedly connected therebetween. 3. A needle assembly, comprising: a housing having a distal portion and a proximal portion engageable with a specimen collection container;a cannula having a patient end, a non-patient end, and a sidewall extending therebetween defining a cannula interior, the patient end of the cannula projecting at least partially from the distal portion of the housing; anda shield restrainably engaged with a portion of the housing and axially transitionable over the patient end of the cannula from a retracted position in which the patient end is exposed, to an extended position in which the patient end is shielded by at least a portion of the shield, wherein the shield comprises a first portion for slideably engaging a portion of the housing, and a second portion at least partially surrounding a portion of the cannula in the extended position, wherein the housing comprises a first portion and a second portion, the first portion of the housing distal to the second portion of the housing, and wherein the shield includes a distal end that is entirely disposed within an interior of the first portion of the housing in the retracted position, and a specimen collection container receiving port is defined within the second portion. 4. The needle assembly of claim 3, wherein the first portion and the second portion of the housing are separate components that are assembled together. 5. A needle assembly, comprising: a housing having a distal portion and a proximal portion engageable with a specimen collection container;a cannula having a patient end, a non-patient end, and a sidewall extending therebetween defining a cann... | Becton, Dickinson and Company | A61B005/00 A61B005/15 A61B005/154 A61B005/153
US10251587 | B2 | 2019-04-09 | Conditioning of chemo-optical sensors for transcutaneous application | 2014-06-06 | The present invention relates to a chemo-optical sensor unit for transcutaneous measurement of a concentration of a gas, comprising: at least one sensing layer adapted to be irradiated with a predetermined radiation; and at least one gas-permeable layer adjacent to one side of the at least one sensing layer, adapted to pass gas whose concentration is to be measured through the gas-permeable layer towards the sensing layer; wherein said chemo-optical sensor unit is adapted to operate with a contact medium between the gas-permeable layer and the skin, wherein said contact medium comprises a first compound other than water; wherein said chemo-optical sensor unit is characterized in that said at least one gas-permeable layer and said at least one sensing layer are permeable to said first compound; and wherein the chemo-optical sensor unit is adapted to measure an optical response of the at least one sensing layer, whose optical response depends on the concentration of the gas. The present invention also relates to a system comprising such a chemo-optical sensor, as well as to a method for conditioning a chemo-optical sensor unit for measuring a concentration of a gas and a thereby obtainable conditioned sensor. | 1. A chemo-optical sensor unit for transcutaneous measurement of a concentration of a gas, comprising: at least one sensing layer adapted to be irradiated with a predetermined radiation;at least one gas-permeable layer adjacent to one side of the at least one sensing layer and adapted to pass gas whose concentration is to be measured through the at least one gas-permeable layer towards the at least one sensing layer;a contact medium configured to contact skin and, when the contact medium is in contact with the skin, is positioned between the at least one gas-permeable layer and the skin, wherein the contact medium comprises a first compound other than water, wherein the first compound is an organic compound, and wherein the chemo-optical sensor unit is characterized in that the at least one gas-permeable layer and the at least one sensing layer are permeable to the first compound; andat least one photosensitive detector adapted to measure an optical response of the at least one sensing layer to the predetermined radiation, wherein the optical response depends on the concentration of the gas. 2. The chemo-optical sensor unit of claim 1, wherein the at least one gas-permeable layer and/or the at least one sensing layer comprises the first compound. 3. The chemo-optical sensor unit of claim 1, wherein the contact medium is at least one of: gas-permeable, biocompatible, thermally conductive, and capable of at least partially penetrating the chemo-optical sensor unit. 4. The chemo-optical sensor unit of claim 1, wherein the first compound is a carbohydrate, a reduced carbohydrate, a polysaccharide, a polyhydroxy alkane or any derivative thereof, or a C2-C10 alcohol or a derivative thereof. 5. The chemo-optical sensor unit of claim 4, wherein the first compound is propylene glycol, xylitol, sorbitol or glycerol, or a derivative thereof. 6. The chemo-optical sensor unit claim 1, wherein the contact medium additionally comprises (i) water and (ii) dissolved sodium chloride. 7. The chemo-optical sensor unit of claim 1, wherein the at least one gas-permeable layer and/or the at least one sensing layer comprises a silicon rubber. 8. The chemo-optical sensor unit of claim 1, wherein the first compound is present in the chemo-optical sensor unit in such a concentration that the optical response is stable when the chemo-optical sensor unit is in contact with the contact medium having a constant gas concentration. 9. The chemo-optical sensor unit of claim 1, wherein the at least one sensing layer comprises luminescent material, and wherein the at least one gas-permeable layer is adapted to prevent light from passing through the at least one gas-permeable layer. 10. The chemo-optical sensor unit of claim 1, wherein: the chemo-optical sensor unit is a transcutaneous sensor unit for measuring blood gas concentration; andthe blood gas concentration corresponds to a gas concentration of O2 and/or a gas concentration of CO2. 11. The chemo-optical sensor uni... | KONINKLIJKE PHILIPS N.V. | A61B005/1455 A61B005/145 G01N033/00
US10241161 | B2 | 2019-03-26 | Magnetic resonance imaging device and imaging parameter determination method | 2014-07-23 | Provided is a technique capable of achieving safety and image quality at the same time in an MRI device using a transmission coil including plural channels. An SAR distribution is calculated, and an imaging parameter is determined for optimizing a radio frequency magnetic field distribution in an imaging region while suppressing a maximum value of an SAR to be equal to or smaller than a predetermined threshold value. The determined imaging parameter includes a radio frequency magnetic field parameter for specifying radio frequency pulses to be transmitted through each of the plural channels. The SAR distribution is calculated using a database that retains an electric field distribution of each of the plural channels for each subject model retained in advance. | 1. A magnetic resonance imaging device comprising: a transmission coil that includes a plurality of channels that respectively transmit radio frequency pulses to a subject;a storage device which stores a database of data to calculate a specific absorption rate distribution for each of a plurality of predetermined subject models; anda processor programmed to determine a plurality of imaging parameters among imaging conditions used in imaging the subject, and execute imaging using the determined imaging parameters,wherein the processor is further programmed to:determine the imaging parameters to optimize a radio frequency magnetic field distribution in a region of interest while suppressing a maximum value of a local specific absorption rate to be equal to or smaller than a specific absorption rate threshold value using the data in the database, andwherein the imaging parameters includes a radio frequency magnetic field parameters which specify the radio frequency pulses to be transmitted through each of the plurality of channels and a specific absorption rate parameter which is different than the radio frequency magnetic field parameters and affects the specific absorption rate,wherein the processor is further programmed to:determine the radio frequency magnetic field parameters according to predetermined optimization conditions,calculate a specific absorption rate distribution using the determined radio frequency magnetic field parameters, a latest specific absorption rate parameter, and the data in the database when the radio frequency magnetic field parameters are determined,determine whether a maximum value of the calculated specific absorption rate distribution exceeds the specific absorption rate threshold value,when the maximum value exceeds the specific absorption rate threshold value, adjust the specific absorption rate parameter in a predetermined range,when the specific absorption rate parameter is changed, calculate the specific absorption rate distribution using the changed specific absorption rate parameter, determine whether a maximum value of the calculated specific absorption rate distribution using the changed specific absorption rate parameter exceeds the specific absorption rate threshold value, andrepeats, when the maximum value of the specific absorption rate distribution exceeds the specific absorption rate threshold value, a process of changing a threshold value for regulating a uniformity condition in the optimization conditions by a predetermined value, a process of determining the radio frequency magnetic field parameter again, and a process of calculating the specific absorption rate distribution using the re-determined radio frequency magnetic field parameter, until the maximum value of the specific absorption rate distribution becomes equal to or smaller than the specific absorption rate threshold value. 2. The magnetic resonance imaging device according to claim 1, wherein the optimization conditions include the uni... | Hitachi, Ltd. | G01R033/28 A61B005/055 G01R033/483 A61B005/00 G01R033/3875 G01R033/34 G01R033/561
US10238312 | B2 | 2019-03-26 | Wrist-type body component measuring apparatus and body component measuring method using the same | 2015-07-21 | Provided is a wrist-type body component measuring apparatus. The wrist-type body component measuring apparatus includes: a band configured to be worn on a wrist of a user; a first input electrode and a first output electrode disposed on an inside surface of the band and configured to be in contact with the wrist of the user; a second input electrode and a second output electrode disposed on an outside surface of the band; a measuring unit configured to apply a current to the first and second input electrodes and detect a voltage from the first and second output electrodes to measure a body impedance of the user; and an electrode converter configured to convert a disposition of the first and second input electrodes and the first and second output electrodes based on a determination of whether the band is worn on a left wrist or a right wrist of the user. | 1. A wrist-type body component measuring apparatus comprising: a band configured to be worn on a user;a first input electrode and a first output electrode which are disposed on an inside surface of the band and configured to contact a wrist of the user;a second input electrode and a second output electrode which are disposed on an outside surface of the band;a measuring unit configured to apply a current to the first and second input electrodes and detect a voltage from the first and second output electrodes to measure a body impedance of the user;an input interface configured to receive, from the user, information about whether the band is worn on a left wrist or a right wrist of the user; andan electrode converter configured to convert a disposition of the first and second input electrodes and the first and second output electrodes according to the information about whether the band is worn on the left wrist or the right wrist. 2. The wrist-type body component measuring apparatus of claim 1, wherein the measuring unit comprises: a current provider configured to apply the current to the first and second input electrodes;a voltage detector configured to detect the voltage from the first and second output electrodes; andan impedance calculator configured to calculate the body impedance from the current and the voltage. 3. The wrist-type body component measuring apparatus of claim 2, wherein the electrode converter comprises a plurality of switches configured to change a first connection between each of the first and second output electrodes and the voltage detector and a second connection between each of the first and second input electrodes and the current provider to a third connection between each of the first and second output electrodes and the current provider and a fourth connection between each of the first and second input electrodes and the voltage detector. 4. The wrist-type body component measuring apparatus of claim 1, wherein an arrangement direction of the first input electrode and the first output electrode and an arrangement direction of the second input electrode and the second output electrode are different from a lengthwise direction of the band. 5. The wrist-type body component measuring apparatus of claim 4, wherein the arrangement direction of the first input electrode and the first output electrode and the arrangement direction of the second input electrode and the second output electrode are perpendicular to the lengthwise direction of the band. 6. The wrist-type body component measuring apparatus of claim 5, wherein the electrode disposition is converted such that the first input electrode and the second input electrode are disposed on a hand side of the wrist and the first output electrode and the second output electrode are disposed on a body side of the wrist based on the information about whether the band is worn on the left wrist or on the right wrist. 7. The wrist-type body component measuring apparatus of cla... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/053 A61B005/00 A61B005/06 A61B005/024 A61B005/0402 A61B005/0205
US20160135750 | A1 | 2016-05-19 | METHODS AND SYSTEMS FOR DETECTION AND THERMAL TREATMENT OF LOWER URINARY TRACT CONDITIONS | 2016-01-21 | The invention relates generally to systems and methods for detecting and/or treating lower urinary tract conditions. One embodiment of the invention is directed to a medical device. The medical device includes an elongate member having a proximal end and a distal end and an end effector assembly that extends distally from the distal end of the elongate member. The end effector assembly includes a plurality of end effector units. Each end effector unit has a sensing element for detecting a location of abnormal organ function and a treatment element for treating the location of abnormal organ function. | 1. A medical device, comprising: an elongate member having a proximal end and a distal end; andan end effector assembly extending distally from the distal end of the elongate member, the end effector assembly including a plurality of end effector units each having a sensing element for detecting a location of abnormal organ function and a treatment element for treating the location of abnormal organ function. 2. The medical device of claim 1, wherein the end effector assembly includes a plurality of legs extending from a proximal end of the end effector assembly to a distal end of the end effector assembly, the plurality of legs forming a three-dimensional sphere in an expanded state. 3. The medical device of claim 2, wherein each leg includes a plurality of lumens, and wherein each lumen is in communication with a corresponding lumen of the elongate member. 4. The medical device of claim 3, wherein each lumen of each leg of the end effector assembly and corresponding lumen of the elongate member receives a bundle that terminates at one of the plurality of end effectors. 5. The medical device of claim 4, wherein the bundle includes a wire, a fluid conduit, a lead, a catheter, and an optical fiber. 6. The medical device of claim 1, wherein the plurality of end effector units are substantially uniformly disposed on the end effector assembly. 7. The medical device of claim 1, wherein each end effector unit further includes a second treatment element for treating the location of abnormal organ function. 8. The medical device of claim 1, wherein each end effector unit further includes a fluid conduit for delivering a fluid to the location of abnormal organ function. 9. A device for treating a lower urinary tract, comprising: an elongate member having a proximal end, a distal end, and one or more lumens;an end effector assembly extending distally from the distal end of the elongate member, the end effector assembly defining one or more exit apertures, wherein each exit aperture is in communication with a corresponding lumen of the elongate member; andone or more bundles, each bundle extending through one of the one or more lumens of the elongate member and terminating at an end effector unit disposed in one of the one or more exit apertures of the end effector assembly. 10. The device of claim 9, wherein the end effector unit includes one or more apertures and a fluid port. 11. The device of claim 10, wherein each bundle includes a wire that terminates at a sensing device provided in one of the one or more apertures of the end effector unit. 12. The device of claim 10, wherein each bundle includes an energy transmission medium that terminates at a treatment element provided in another of the one or more apertures of the end effector unit. 13. The device of claim 10, wherein each bundle includes a fluid conduit that terminates at the fluid port. 14. A device for treating a lower urinary tract, comprising: a support structure configured to c... | Boston Scientific Scimed, Inc. | A61B005/00 A61F002/00 A61B018/14
US20160143552 | A1 | 2016-05-26 | ELECTROCARDIOGRAPHY SIGNAL EXTRACTION METHOD | 2016-01-28 | An electrocardiography signal extraction method includes receiving an electrocardiography signal, detecting a peak of a waveform of the electrocardiography signal, separating the waveform into left and right waves, normalizing the left wave and a plurality of scales of Gaussian function, comparing the normalized left wave with a left part of the normalized scales of Gaussian function, acquiring a left part error function, indicating a left minimum comparative error, selecting a left scale of Gaussian function with the left minimum comparative error, obtaining a left duration of the waveform, normalizing the right wave, comparing the normalized right wave with a right part of the normalized scales of Gaussian function, acquiring a right part error function, indicating a right minimum comparative error, selecting a right scale of Gaussian function with the right minimum comparative error, obtaining a right duration of the waveform, and obtaining an extracted wave. | 1. An electrocardiography signal extraction method for reducing the effect of the baseline drift of an electrocardiography signal retrieved by a signal retriever, the electrocardiography signal extraction method being performed on a processor of a computer system along with a predetermined database, the electrocardiography signal extraction method comprising: receiving the electrocardiography signal by the processor of the computer system;detecting a peak of a waveform of the electrocardiography signal;separating the waveform into a left wave and a right wave, wherein the left wave is the portion of the waveform to the left of the detected peak and the right wave is the portion of the waveform to the right of the detected peak;normalizing the left wave and a plurality of scales of a Gaussian function;comparing the normalized left wave with a left part of the normalized scales of the Gaussian function;acquiring a left part error function according to the compared result of the normalized left wave and the left part of the normalized scales of the Gaussian function;indicating a left minimum comparative error;selecting a left scale of the Gaussian function with the left minimum comparative error;obtaining a left duration of the waveform according to the selected left scale of the Gaussian function and the peak;normalizing the right wave;comparing the normalized right wave with a right part of the normalized scales of the Gaussian function;acquiring a right part error function according to the compared result of the normalized right wave and the right part of the normalized scales of the Gaussian function;indicating a right minimum comparative error;selecting a right scale of the Gaussian function with the right minimum comparative error;obtaining a right duration of the waveform according to the selected right scale of the Gaussian function and the peak;obtaining an extracted wave from the detected peak, the selected left duration and the selected right duration; anddisplaying the extracted wave on a display of the computer system,wherein the Gaussian function is represented byg(μ,σ)(x)=12πσ-12(x-μσ)2,wherein μ is 0, wherein σ is 5 to 20, wherein x is represented byx=f1×(v1f2),wherein f1 is a sampling rate of the signal retriever, wherein f2 is a sampling rate of the predetermined database, wherein a ratio of v1 to f2 is 0.12-0.2. 2. The electrocardiography signal extraction method for reducing the effect of the baseline drift in the electrocardiography signal as claimed in claim 1, further comprising de-noising the waveform before separating the waveform. 3. The electrocardiography signal extraction method for reducing the effect of the baseline drift in the electrocardiography signal as claimed in claim 1, wherein the left wave and the right wave are normalized at the same time. 4. The electrocardiography signal extraction method for reducing the effect of the baseline drift in the electrocardiography signal as claimed in claim 1, ... |  | A61B005/0456 A61B005/00 A61B005/04
US20160136414 | A1 | 2016-05-19 | FLEXIBLE AND/OR TAPERED THERAPY ELECTRODE | 2016-01-20 | An electrode assembly includes a first surface to be placed adjacent a person's skin and a second surface including a plurality of reservoirs of conductive gel. The plurality of reservoirs of conductive gel are disposed on sections of the electrode assembly that are at least partially physically separated and may move at least partially independently of one another to conform to contours of a body of a patient. The electrode assembly is configured to dispense an amount of the electrically conductive gel onto the first surface in response to an activation signal and to provide for a defibrillating shock to be applied to the patient through the amount of the electrically conductive gel. | 1. A wearable medical device comprising: one or more electrode assemblies, at least one of the one or more electrode assemblies configured to facilitate proper placement of an electrode into the wearable medical device. 2. The wearable medical device of claim 1, wherein at least one of the one or more electrode assemblies comprises a therapy electrode configured to deliver a defibrillating shock to a patient. 3. The wearable medical device of claim 1, wherein the at least one of the one or more electrode assemblies comprises an ECG sensing electrode configured to monitor an ECG signal of a patient. 4. The wearable medical device of claim 1, wherein the wearable medical device comprises a pocket configured to receive the electrode. 5. The wearable medical device of claim 4, wherein the at least one of the one or more electrode assemblies is disposed on a garment wearable on a torso of a patient. 6. The wearable medical device of claim 1, wherein the at least one of the one or more electrode assemblies comprises at least one feature to discourage insertion of the electrode into the wearable medical device in an incorrect direction. 7. The wearable medical device of claim 1, wherein the wearable medical device comprises at least one feature to discourage incorrect insertion of the electrode into the wearable medical device. 8. The wearable medical device of claim 7, wherein the at least one feature is a shape of a pocket of the wearable medical device. 9. The wearable medical device of claim 8, wherein the shape of the pocket conforms to a shape of the electrode. 10. The wearable medical device of claim 8, wherein the pocket is tapered from a first end to a second end. 11. The wearable medical device of claim 7, wherein the at least one feature is a fastener having a first portion disposed in the wearable medical device and a complimentary portion disposed on the electrode. 12. The wearable medical device of claim 11, wherein the at least one feature is a snap. 13. The wearable medical device of claim 11, wherein the at least one feature is a magnet. 14. The wearable medical device of claim 11, wherein the at least one feature is a hook and loop fastener. 15. A therapy electrode configured to be inserted into a pocket of a garment of a wearable medical device, the therapy electrode comprising: at least one feature discouraging incorrect insertion of the therapy electrode into the pocket. 16. The therapy electrode of claim 15, wherein the at least one feature discourages insertion of the therapy electrode into a second pocket of the wearable medical device in an incorrect direction. 17. The therapy electrode of claim 16, wherein the at least one feature is a shape of the therapy electrode. 18. The therapy electrode of claim 17, wherein the shape of the therapy electrode conforms to a shape of the pocket of the garment. 19. The therapy electrode of claim 17, wherein the therapy electrode is tapered from a first end to a second end. ... |  | A61N001/04 A61B005/00 A61B005/0408
US20160120474 | A1 | 2016-05-05 | Wearable Device for the Ear with Electroencephalographic and Spectroscopic Sensors | 2016-01-11 | This invention is a wearable device for the ear comprising: an ear attachment with a forehead-projecting portion, an electromagnetic energy sensor which is configured to collect data concerning electromagnetic energy which is transmitted through and/or emitted by the person's body, a microphone, a spectroscopic sensor which is configured to collect data concerning light energy which is transmitted through and/or reflected from the person's body, a sound-emitting unit, a power source, and a data processor. This device configuration is advantageous for mobile electroencephalographic monitoring and spectroscopic analysis of body tissue. | 1. A wearable device for the ear comprising: an ear attachment, wherein this ear attachment further comprisesan inserted portion which is configured to be at least partially inserted into the ear canal,a frontal-connecting portion which is configured to span a portion of an outer ear or ear canal opening perimeter between the 11 o'clock and 4 o'clock positions, wherein clock hour positions around the perimeter of an outer ear or ear canal opening are defined relative to a 12 o'clock position that is vertically above the center of the outer opening of the ear canal,a forehead-projecting portion which is configured to span from the ear to the person's temple and/or forehead,a rear-spanning portion which is configured to span a portion of an outer ear or ear canal opening perimeter between the 5 o'clock and 1 o'clock positions, andan earlobe portion which is configured to be attached to the earlobe;an electromagnetic energy sensor which is configured to collect data concerning electromagnetic energy which is transmitted through and/or emitted by the person's body;a microphone;a spectroscopic sensor which is configured to collect data concerning light energy which is transmitted through and/or reflected from the person's body;a sound-emitting unit;a power source; anda data processor. 2. The device in claim 1 wherein the forehead-projecting portion is connected to or integrated with the frontal-connecting portion. 3. The device in claim 1 wherein the forehead-projecting portion is configured to curve toward the center of a person's forehead as it projects forward from the person's ear. 4. The device in claim 1 wherein the forehead-projecting portion is configured with a tensile mechanism or actuator such that its front portion presses gently against a person's temple and/or forehead. 5. The device in claim 1 wherein the electromagnetic energy sensor is an electroencephalographic sensor. 6. The device in claim 1 wherein the electromagnetic energy sensor is located on the forehead-projecting portion. 7. The device in claim 1 wherein the electromagnetic energy sensor is located on the inserted portion. 8. The device in claim 1 wherein the electromagnetic energy sensor is located on the earlobe portion. 9. The device in claim 1 wherein the spectroscopic sensor is located on the earlobe portion. 10. A wearable device for the ear comprising: an ear attachment, wherein this ear attachment further comprisesa frontal-connecting portion which is configured to span a portion of an outer ear or ear canal opening perimeter between the 11 o'clock and 4 o'clock positions, wherein clock hour positions around the perimeter of an outer ear or ear canal opening are defined relative to a 12 o'clock position that is vertically above the center of the outer opening of the ear canal,a forehead-projecting portion which is configured to span from the ear to the person's temple and/or forehead,a rear-spanning portion which is configured to span a portion of an outer e... | Medibotics LLC | A61B005/00 A61B005/0478 A61B007/04
US10123704 | B2 | 2018-11-13 | Method and system for dual-band active thermal imaging using multi-frequency currents | 2013-07-05 | A hybrid system is developed using thermal and electrical impedance imaging methods together. The innovation of the approach relies on the frequency dependence of the tissue's electrical impedance which facilitates the acquisition of multiple thermal images with currents at different frequencies injected to the region of the body under inspection. Proposed method without current application (in passive mode of operation) provides images which are obtained by standard thermal imagers. On the other hand, the application of the electrical current (in active mode of operation) increases the temperature contrast on the body surface depending on the electrical property of tissue. Therefore, the technique while increasing the thermal contrast provides frequency dependent conductivity distribution data which can be used as a basis for the detection of the breast carcinoma. The sensitivity of the technique is increased by an infrared camera with dual band (MWIR/LWIR) imaging capability. | 1. A medical electro-thermal imaging method, comprising the steps of: a. preparing a patient and the imaging room according to a thermal imaging standards;b. recording of a passive thermal image of the tissue in two different windows (MWIR and LWIR) using a dual band thermal camera;c. placing a plurality of non-invasive electrodes to the specified locations on the surface of the breast tissue;d. selecting frequency of an electric current that is inserted to the breast tissue using the control unit;e. applying the electric current to the targeted breast tissue for predetermined time period;f. recording an active thermal image of the stimulated area in, heating period (in both windows (MWIR and LWIR));g. recording the active thermal image of the stimulated area in cooling period (in both windows (MWIR and LWIR)) until the temperature of the breast tissue becomes stable;h. going to the item c and repeating the same procedure for different frequency values of the applied electric current;i. eliminating the position differences of the active thermal image and the passive thermal image and matching the active thermal image and the passive thermal image (image registration);j. obtaining the absolute temperature value of the breast and cancerous tissue for any screening time period using the dual band property of the thermal camera;k. comparing and determining the ratio/difference of the active thermal image and the passive thermal image (in single band and dual band imaging modes) obtained under different frequencies using the image processing algorithm which is embedded in the control unit to detect the cancerous tissue (tumor tissue);l. if the active thermal image is not enough to make a healthy diagnosis for the presence of the tumor after changing the frequency of the electric current or the positions of the electrodes; carrying out the previous d, e, f, g, h, i, j, k and l steps until a healthy diagnosis or detection is made;wherein the decision in this sense is made with the help of the algorithm embedded in the control unit; andwherein the patients disrobe to the waist and let the surface of the breasts cool to room temperature of 18-22° C. for 10-15 minutes. 2. A medical electro-thermal imaging system configured to perform the method of claim 1, which uses the electric current as an external heat source to improve the temperature contrast, comprising: an imaging room configured to satisfy the thermal imaging standards;a dual-band infrared camera configured to sense temperature emission and emissivity of the focused area to record a passive and an active thermal image of the tissue in two different windows (MWIR and LWIR);a plurality of non-invasive electrodes are configured to insert the specified locations on the surface of the breast tissue deliver the electric current to the breast tissue;a voltage controlled current source configured to provide the electric current that is injected to the breast tissue;a control unit configured to regulate ... | - | A61B005/00 A61B005/01 A61B005/05 G06T007/00 A61B005/053
US20160149306 | A1 | 2016-05-26 | MICROSTRIP ANTENNA STRUCTURE AND MICROWAVE IMAGING SYSTEM USING THE SAME | 2014-12-09 | A microstrip antenna structure is provided, which includes a substrate, a ring microstrip line and a signal transmission port. The substrate has opposite first and second surfaces. The ring microstrip line is disposed on the first surface of the substrate. The ring microstrip line has a short coupling gap ranged between 0.004λg and 0.06λg for forming a radiation bandwidth with high selectivity, where λg represents a guided wavelength of an electromagnetic wave in the ring microstrip line corresponding to a center frequency of the radiation bandwidth. The signal transmission port is disposed on the second surface of the substrate. The signal transmission port penetrates through the substrate and is electrically connected to the ring microstrip line. | 1. A microstrip antenna structure, comprising: a substrate having opposite first and second surfaces;a ring microstrip line disposed on the first surface of the substrate, the ring microstrip line having a short coupling gap ranged between 0.004λg and 0.06λg for forming a radiation bandwidth with high selectivity, where λg represents a guided wavelength of an electromagnetic wave in the ring microstrip line corresponding to a center frequency of the radiation bandwidth; anda signal transmission port disposed on the second surface of the substrate, the signal transmission port penetrating through the substrate and electrically connected to the ring microstrip line. 2. The microstrip antenna structure of claim 1, wherein the ring microstrip line has a width ranged between 0.01λg and 0.13λg. 3. The microstrip antenna structure of claim 1, wherein the ring microstrip line is a rectangular-shaped ring coaxial line, a rectangular-shaped ring coplanar waveguide line, a rectangular-shaped ring slotline or a rectangular-shaped ring stripline. 4. The microstrip antenna structure of claim 1, wherein the ring microstrip line comprises at least one material selected from the group consisting of titanium (Ti), cobaltum (Co), wolfram (W), hafnium (Hf), tantalum (Ta), molybdanium (Mo), chromium (Cr), agtentum (Ag), cuprum (Cu) and aluminium (Al). 5. The microstrip antenna structure of claim 1, wherein the substrate is a FR4 substrate, a RT/Duroid series substrate, an aluminum oxide substrate, a RO series substrate, a high temperature cofired ceramic (HTCC) substrate, a low temperature cofired ceramic (LTCC) substrate, a transparent conductive substrate or a semiconductor substrate; wherein the RO series substrate comprises at least one material selected from the group consisting of magnesium oxide, calcium oxide, strontium oxide and barium oxide. 6. The microstrip antenna structure of claim 1, wherein the signal transmission port comprises at least one material selected from the group consisting of titanium, cobaltum, wolfram, hafnium, tantalum, molybdanium, chromium, agtentum, cuprum and aluminium. 7. The microstrip antenna structure of claim 1, further comprising: a ground conductor disposed on the second surface of the substrate, the ground conductor electrically insulated from the signal transmission port. 8. The microstrip antenna structure of claim 7, wherein the ground conductor defines an inner space, and the signal transmission port is located in the inner space. 9. The microstrip antenna structure of claim 7, wherein the ground conductor comprises at least one material selected from the group consisting of titanium, cobaltum, wolfram, hafnium, tantalum, molybdanium, chromium, agtentum, cuprum and aluminium. 10. A microstrip antenna structure, comprising: a substrate having opposite first and second surfaces;a plurality of ring microstrip lines disposed on the first surface of the substrate, each of the ring microstrip lines having a short coup... |  | H01Q009/04 A61B005/05
US20160120418 | A1 | 2016-05-05 | BLOOD PRESSURE MEASUREMENT DEVICE | 2014-06-03 | A method for measuring blood pressure of a patient continuously and non-invasively comprises measuring a first pressure signal being proportional to arterial blood pressure of the patient at a first location by a first pressure sensor (P1). In addition ambient pressure is measured by a third ambient pressure sensor (P3). Furthermore the method comprises subtracting the signal derived from the third ambient pressure sensor from the signal derived from the first pressure sensor to compensate for alterations induced by alterations in measurement point altitude and atmospheric pressure changes thereby providing a signal representing relative systemic arterial blood pressure of the patient | 1. A device for measuring blood pressure of a patient continuously and non-invasively, wherein the device comprises: a first pressure sensor configured to measure a pressure signal being induced by arterial blood pressure of the patient at a first location,a second pressure sensor configured to measure a pressure signal being induced by arterial blood pressure of the patient at a second location, said second location being different than said first location,a third ambient pressure sensor configured to measure ambient pressure,wherein the device is configured to: subtract the signal derived from the third ambient pressure sensor from the signal derived from the first pressure sensor and the from the signal derived from the second pressure sensor to compensate for alterations induced by alterations in measurement points altitude and atmospheric pressure changes thereby providing a signal representing relative systemic arterial blood pressure of the patient. 2. The device of claim 1, wherein the first and second sensors are arranged in the device so that in use they are configured to be pressed against measurement location of the patient at a known fixed distance from each other, wherein the distance is between 0.5-5 cm. 3. The device of claim 2, wherein the device is configured to be placed on a measurement location on the course of distal radial artery. 4. The device of claim 1, wherein the pressure sensors are capacitive sensors. 5. The device of claim 1, wherein the device comprises at least one accelerometer for measuring movements of the device and thereby the movements of the user, comprising movements of a hand or changes in altitude, comprising falls and collapses. 6. The device of claim 14, wherein the device is configured to use acceleration data provided by at least two accelerometers for baseline calibration procedure yielding Δh and ΔPhydrostatic_calibration and subsequently ΔPWVcalibration. 7. The device of claim 1, wherein the pressure sensors are capacitive pressure sensors, and wherein the sampling resolution is at least 100 Hz. 8. The device of claim 1, wherein a maximum signal is derived of the pulse wave after said subtraction and wherein the first and second sensors are arranged to detect the signals so that the first proximal sensor detects the signal before the second distal one, whereupon the device is configured to provide this as a first quality control. 9. The device of claim 1, wherein the blood pressure is based on pulse wave velocity measurement, wherein the velocity of the pulse is determined based on the time difference between said first and second detectors detect the same pulse and the distance of said first and second sensors. 10. The device of claim 1, wherein the signal of the third ambient pressure sensor is used for calibration of the first or second sensors measurements so that the signals representing the absolute systemic arterial blood pressure of the patient is provided. 11. The device of clai... |  | A61B005/021 A61B005/00 A61B005/0205
US20160128582 | A1 | 2016-05-12 | A METHOD OF DETERMINING SYSTOLIC AND DIASTOLIC BLOOD PRESSURE AND THE UNIT FOR THIS METHOD | 2014-06-18 | The method of determining systolic and diastolic blood pressure where the set of received data of each curve of the plethysmographic curve gained from the unit for monitoring the life functions in a shape of the ring and/or the earring is divided into individual segments that respond to one heartbeat of the measured subject and after processing there are selected and recalculated segments meeting the defined criteria. The unit for monitoring and measuring the life functions comprises a supporting construction in a shape of a ring (29) and/or an earring (39). | 1. A method of determining systolic and diastolic blood pressure, wherein the received data set of separate curves of plethysmographic curve (1) is divided into individual segments (2a ) with two local minimums given by the beginning (M1) and the end (M2) and one global maximum (M3) that matches one heartbeat of the measured subject and represents one cycle of the heart activity, the course of segments (2a ) of the plethysmographic curve (1) is then divided into individual sections containing at least one segment (2a ) with the division up to 20 segments (2a) and for the exclusion of the saturated signal from further assessment are according to the constant Ka the saturated parts of segments (2a) excluded from further assessment and the recommended values are always determined according to a type of the scanning unit as the constant Ka respectively its limits, specifically KaMIN a KaMAX, and the course of the plethysmographic curve (1) is then filtered by making averages of consecutive values of the plethysmographic curve (1), each segment (2a) is checked for differences between the value Y of the coordination of the beginning (M1) and the end (M2) of the curve and only selected segments (2b) are left, these segments up to 25%, selected segments (2b) are being normed so that the value Y of the global maximum (M3) is matched with the value 1 and the value 0 is determined as a result of the arithmetic average of the beginning (M1) and the end (M2) for selected segments (2b) with the difference lower than 5% and/or the values of the plethysmographic curve (1) for the selected segments (2b) with the difference lower than 25% are recalculated in the manner that the beginning and the end of the selected segment (2b) have the same value, and are divided into 6 parts defined as the dividing vertical out of the maximum (M3) divided in the axis Y for values between 20 to 60% and 40 to 90% out of the maximum (M3), where the second dividing must be by percentage larger than the first one and the partial areas (S6, S7, S8, S9, S10, S11) are determined as well as the centers of gravity (Tx6, Ty6, Tx7, Ty7, Tx8, Ty8, Tx9, Ty9, Tx10, Ty10, Tx11, Ty11) of each part and the centers of gravity (Tx, Ty) of the whole selected segment (2b), at the same time the constant (Kind) is determined as the value of the statistic average of subject checking measurements example for systolic and diastolic pressure and the constant (Kn) as the value of the statistic average of the subject controlling measurements divided according to the age category and that together with a determination of the range constant (KR) of the measurement range, the quantity (D) is determined as the first derivation of the leading edge of each segment (2a) and now the systolic blood pressure is determined as a result of KRS* KnS*KindS*D*S8 and the diastolic blood pressure is determined as a result of KRD*KnD*KindD*((Ty*Tx)*S). 2. The method according to claim 1, characterized by determining the parti... | CZECH TRCHNICAL UNIVERSITY IN PRAGUE, FACULTY OF ELECTRICAL ENGINEERING DEPARTMENT OF TELECOMMUNICA | A61B005/021 A61B005/026 A61B005/00 A61B005/0295
US09332941 | B2 | 2016-05-10 | Body-worn sensor for characterizing patients with heart failure | 2013-12-31 | The invention provides a sensor for measuring both impedance and ECG waveforms that is configured to be worn around a patient's neck. The sensor features 1) an ECG system that includes an analog ECG circuit, in electrical contact with at least two ECG electrodes, that generates an analog ECG waveform; and 2) an impedance system that includes an analog impedance circuit, in electrical contact with at least two (and typically four) impedance electrodes, that generates an analog impedance waveform. Also included in the neck-worn system are a digital processing system featuring a microprocessor, and an analog-to-digital converter. During a measurement, the digital processing system receives and processes the analog ECG and impedance waveforms to measure physiological information from the patient. Finally, a cable that drapes around the patient's neck connects the ECG system, impedance system, and digital processing system. | 1. A method for monitoring an electrical impedance from a patient, the method comprising the following steps: providing a loop-shaped, flexible cable around a patient's neck, the flexible cable including: i) at least four electrodes, with a first set of electrodes configured to connect to the patient's chest proximal to opposite sides of their heart and inject electrical current into the patient near their chest, and a second set of electrodes configured to connect to the patient's chest proximal to opposite sides of their heart and measure electrical signals from the patient's chest; ii) an impedance-measuring system worn on the patient's body and in electrical contact with the second set of electrodes; and iii) a data-processing system worn on the patient's body and in electrical contact with the impedance-measuring system;injecting electrical current into the patient near their chest with at least one electrode in the first set of electrodes;measuring a voltage near the patient's chest with the second set of electrodes located on the patient's chest proximal to opposite sides of their heart, the voltage related to a product of the injected current and an impedance of the patient; andprocessing the voltage to determine an impedance value. 2. The method of claim 1, further comprising the step of measuring a voltage with the second set of electrodes using a differential amplifier. 3. The method of claim 2, wherein the differential amplifier is configured to measure a time-dependent voltage indicating the product of electrical impedance near the patient's chest and current injected by the second set of electrodes. 4. The method of claim 3, further comprising measuring a first time-dependent voltage with the differential amplifier, the first time-dependent voltage indicating how fluid levels and respiration affect electrical impedance in the patient's chest. 5. The method of claim 4, further comprising processing the first time-dependent voltage with a computer algorithm to estimate the fluid levels in the patient's chest. 6. The method of claim 4, further comprising processing the first time-dependent voltage with a computer algorithm to estimate the patient's respiration rate. 7. The method of claim 3, further comprising measuring a second time-dependent voltage with the differential amplifier, the second time-dependent voltage indicating how heartbeat-induced blood flow affects electrical impedance in the patient's chest. 8. The method of claim 7, further comprising processing the second time-dependent voltage with a computer algorithm to estimate the patient's stroke volume. 9. The method of claim 7, further comprising processing the second time-dependent voltage with a computer algorithm to estimate the patient's cardiac output. 10. The method of claim 7, further comprising processing the second time-dependent voltage with a computer algorithm to estimate the patient's heart rate. 11. The method of claim 1, further comprising measuri... | TOSENSE, INC. | A61B005/053 A61B005/00 A61B005/0452 A61B005/029 A61B005/0408 A61B005/11 A61B005/0205
US09330461 | B2 | 2016-05-03 | Image-based method for measuring elasticity of biological tissues and system thereof | 2012-08-29 | The present disclosure proposes an image-based method for measuring elasticity of biological tissues including following steps: obtaining N successive grayscale images of a testing biological tissue, where N is a positive integer; obtaining regions of interest (ROI) of the grayscale images; dividing the regions of interest into a grid of small sections; calculating relative displacement vector of each interrogation window using texture matching method; and calculating elastic modulus of each interrogation window according to the relative displacement vectors. The present disclosure also proposes an image-based system for measuring elasticity of biological tissues. The present disclosure applies to various resolutions of grayscale images obtaining in ultrasound imaging, optical imaging, photoacoustic imaging, CT imaging, magnetic resonance imaging ect. The system can be integrated into a traditional clinical imaging system as an image post-processing software module to characterize tissue elasticity. Because it's no need to update the traditional clinical imaging system hardware, updating cost is low. It's easy to be accepted by hospital and is convenient to popularize. | 1. An image-based method for measuring elasticity of biological tissues, comprising the following steps: step A, obtaining a cineloop consisting of N grayscale images of a testing biological tissue, where N is a positive integer; wherein step A further includes step A1, collecting successive N grayscale images by an imaging system, where N is the number of image frames collected by the imaging system within m cardiac cycles, where m is a positive integer;step B, obtaining regions of interest (ROI) of the grayscale images;step C, dividing the regions of interest into a grid of small sections known as interrogation windows;step D, calculating a relative displacement vector of each interrogation window using texture matching method; wherein step D further includes step D1, calculating the two-dimensional translational displacement of the texture of each interrogation window using the 2-D normalized cross-correlation technique which combines the sub-pixel method and the filter and interpolation method;step D2, estimating rotation and deformation according to the gradient of the two-dimensional translational displacement using a multiple iterative algorithm;step D3, applying the 2-D normalized cross-correlation technique to a reduced interrogation window to obtain higher spatial resolution based on the rotation and deformation displacements and using a spurious vector elimination algorithm to obtain more accurate displacement estimates; andstep E, calculating elastic modulus of each interrogation window according to the relative displacement vectors said displacement vectors including lateral and axial components stored in three-dimensional (3-D) arrays, where the three dimensions are lateral and axial positions within an image and frame number. 2. The method according to claim 1, wherein the step E further comprises the following steps: step E1, obtaining the displacement gradient of each layer with a constant thickness from the estimated axial displacement;step E2, obtaining the maximum displacement gradient of each layer during one cardiac cycle;step E3, obtaining the elastic modulus according to the maximum displacement gradient and the blood pressure applied normal to each layer. 3. The method according to claim 1, further comprising: slicing the biological tissue into sections;pathology staining the section;obtaining composition distributions of the biological tissue according to the pathology staining;obtaining a range of the elastic modulus of the different compositions according to the distributions of the composition and the elastic modulus. 4. The method according to claim 1, wherein the step E further comprises the following steps: step E1, obtaining the displacement gradient of each layer with a constant thickness from the estimated axial displacement;step E2, obtaining the maximum displacement gradient of each layer during one cardiac cycle;step E3, obtaining the elastic modulus according to the maximum displacement gradient and the b... | Shenzhen Institutes of Advanced Technology Chinese Academy of Sciences | G06K009/00 G06T007/00 G01N033/483 A61B008/08 A61B005/02 A61B006/00 G01S007/52 G06T007/20 A61B006/03 A61B005/00 A61B008/00
US09864780 | B2 | 2018-01-09 | Gait-based biometric data analysis system | 2015-11-19 | Systems and methods for diagnosing a user's condition based on his gait. A sensor module with multiple sensors is placed inside a user's shoe and biometric data is gathered from the sensors when the user takes a step or walks. The data is used to generate loops as the various sets of data is plotted against each other. The loops obtained from the data are then compared against stored loops previously obtained. Based on the results of the comparison, the user's condition is diagnosed using predetermined indicators of specific health issues. Using the biometric data and in conjunction with data from various databases, it can be determined whether the user has a specific condition, whether a specific condition is worsening, or whether a specific condition is improving. | 1. A method for determining changes in a user's condition, the method comprising: a) in a single insole worn by a user, the insole having a plurality of sensors for gathering gait-based pressure data and having at least one processor, using said at least one processor to select a first sensor and a second sensor from said plurality of sensors, said plurality of sensors being in said insole and each of said plurality of sensors being a discrete force sensor that detects an amount of force or pressure exerted on a section or region of said insole;b) using said at least one processor to gather pressure data from each of said first sensor and said second sensor selected in step a), said pressure data being for a single step of said user;c) using a server connected to at least one database and remote from said insole to correlate data gathered from said first and second sensors such that data points gathered from said first sensor and from said second sensor at similar instances are matched with each other to result in data point pairs, each data point pair having a first data point gathered from the first sensor and a second data point from the second sensor;d) using said server to determine at least one characteristic loop from said data point pairs, each characteristic loop being a loop formed when said data point pairs are plotted such that data points from said first sensor are on one axis and data points from said second sensor are on another axis, each one of said at least one characteristic loop having at least one geometric parameter extractable from said at least one characteristic loop;e) using said server to retrieve signature characteristic data, said signature characteristic data being previous pressure data resulting from biometric data previously gathered from said user, said signature characteristic data being pressure data previously gathered from said first sensor and from said second sensor for a previous single step of said user, said signature characteristic data being previous data point pairs previously gathered from said first sensor and from said second sensor, each previous data point pair having a previous data point from said first sensor and a previous data point from said second sensor;f) using said server to determine a signature characteristic loop from said signature characteristic data, said signature characteristic loop being a loop formed when said previous data point pairs are plotted such that previous data points from said first sensor are on one axis and previous data points from said second sensor are on another axis, said signature characteristic loop having geometric parameters extractable from said signature characteristic loop;g) using said server to calculate said geometric parameters from said at least one characteristic loop and from said signature characteristic loop, wherein said geometric parameters include at least one of: a length of said loops;a width of said loops;an angle of said loops with a gi... | AUTONOMOUS_ID CANADA INC. | A61B005/103 G06F017/30 A61B005/00 A43B003/00 A61B005/11
US09335167 | B2 | 2016-05-10 | Apparatus for measuring body size and method of doing the same | 2011-09-21 | Provided is a body measurement device capable of carrying out a body measurement, such as of abdominal girth or chest girth, with a simple configuration which eliminates moving components. In particular, the body measurement device includes computing device that converts two-dimensional coordinates of images of linear light beams included in images having been captured by a plurality of image capturing devices into two-dimensional coordinates on a certain plane in accordance with a relation stored in a database, and computes a peripheral length of a body of a target person within the certain plane. | 1. An apparatus for measuring a body size, comprising: light emitting devices emitting linear light beams so as to encircle a target person within a certain plane;a plurality of image capturing devices positioned around said target person such that a total circumference of said linear light beams having been emitted to said target person can be captured, and that an optical axis thereof extends obliquely relative to said certain plane;a database storing therein a relation between two-dimensional coordinates on images having been captured by said image capturing devices and two-dimensional coordinates on said certain plane; anda computing device converting two-dimensional coordinates of images of said linear light beams included in images having been captured by said image capturing devices into two-dimensional coordinates on said certain plane in accordance with said relation stored in said database, and computing a peripheral length of a body of said target person within said certain plane,wherein said database divides said certain plane into a plurality of ranges, assigns each of said ranges to each of said image capturing devices, and stores said relation therein for each of said ranges, andsaid computing device converts two-dimensional coordinates of images of said linear light beams included in images having been captured by said image capturing devices corresponding to said ranges into two-dimensional coordinates on said certain plane in accordance with said relation stored in said database, and computes a peripheral length of a body of said target person in each of said ranges within said certain plane. 2. The apparatus as set forth in claim 1, wherein said computing device a peripheral length of a body of said target person within said certain plane by totaling the peripheral lengths calculated for each of said ranges. 3. The apparatus as set forth in claim 1, wherein said database stores therein a second relation between two-dimensional coordinates on images having been captured by said image capturing devices and three-dimensional coordinates including two-dimensional coordinates on said certain plane, in place of or together with said relation. 4. The apparatus as set forth in claim 1, wherein said computing device forms an image displaying a periphery of said target person within said certain plane. 5. A method of measuring a body size, comprising: a first step of emitting linear light beams so as to encircle a target person within a certain plane;a second step of capturing an image of a total circumference of said linear light beams having been emitted to said target person, in a direction oblique to said certain plane;a third step of calculating a relation between two-dimensional coordinates on images having been captured in said second step and two-dimensional coordinates on said certain plane; anda fourth step of converting two-dimensional coordinates of images of said linear light beams included in images having been captured... | Marinex Corporation | H04N015/00 H04N013/02 G06T017/00 G01C011/10 A61B005/00 A61B005/107
US09326780 | B2 | 2016-05-03 | Patient selectable joint arthroplasty devices and surgical tools incorporating anatomical relief | 2014-01-06 | Disclosed herein are methods, compositions and tools for repairing articular surfaces repair materials and for repairing an articular surface. The articular surface repairs are customizable or highly selectable by patient and geared toward providing optimal fit and function. The surgical tools are designed to be customizable or highly selectable by patient to increase the speed, accuracy and simplicity of performing total or partial arthroplasty. | 1. A system for joint arthroplasty for repairing a joint of a patient, the system comprising: a first template, the first template including: a contact surface for engaging a first articular surface of the joint of the patient, the contact surface including shape information derived from electronic image data of at least a portion of the first articular surface; at least a relieved portion of the contact surface further including an anatomical relief configured such that when the contact surface engages the first articular surface, the relieved portion does not engage an anatomical structure of the first articular surface; andat least one guide for directing movement of a surgical instrument; andwherein the guide has a predetermined orientation relative to one of an anatomical and a biomechanical axis associated with the joint of the patient. 2. The system of claim 1, wherein the at least one guide is one of a guide aperture, a reaming aperture, a drill aperture and a cut plane. 3. A method of joint arthroplasty using the system of claim 1, comprising: engaging the contact surface of the first template with the first articular surface of the joint, andmoving the surgical instrument with the at least one guide to prepare the joint of the patient for receiving an implant. 4. The system of claim 1, wherein the joint of the patient is a knee joint, a shoulder joint, a hip joint, a vertebrae, an elbow joint, or an ankle joint. 5. The system of claim 1, wherein the anatomic relief accommodates one or more osteophytes of the joint. 6. The system of claim 1, wherein the anatomic relief accommodates an area of normal cartilage of the joint. 7. The system of claim 1, wherein the anatomic relief accommodates an area of diseased cartilage of the joint. 8. The system of claim 1, wherein the anatomic relief accommodates an area of subchondral bone of the joint. 9. The system of claim 1, wherein the anatomical structure includes a cut surface portion of the joint. 10. The system of claim 1, wherein the anatomical structure includes an uncut surface portion of the joint. 11. The system of claim 1, wherein the joint is a knee joint, and wherein the anatomical structure includes a trochlear portion of the knee joint. 12. The system of claim 1, further including one or more implant components for repairing the joint. | ConforMIS, Inc. | A61B017/17 A61F002/30 A61B019/00 A61B017/15 A61F002/38 A61B017/56 A61B005/00 A61B017/16
US09326751 | B2 | 2016-05-03 | Catheter guidance of external energy for renal denervation | 2011-11-14 | An in vivo apparatus includes a flexible shaft having a proximal end, a distal end, and a length sufficient to access a patient's renal artery relative to a percutaneous access location. An energy guide apparatus is provided at the distal end of the shaft and dimensioned for deployment within the renal artery. An ex vivo apparatus includes an arrangement configured to localize the energy guide apparatus within the renal artery, and an energy source configured to direct ablative energy to target tissue located a predetermined distance from the localized energy guide apparatus. The target tissue includes perivascular renal nerve tissue adjacent the renal artery. | 1. An apparatus, comprising: an in vivo apparatus comprising: a flexible shaft having a proximal end a distal end; andan energy guide apparatus provided at the distal end of the shaft and configured to generate a rotating energy beacon including a rotating magnetic field and a rotating beam of acoustic energy;an ex vivo apparatus comprising: a localizing arrangement configured to localize the energy guide apparatus within a renal artery, wherein the localizing arrangement is configured to: sense the rotating magnetic field;detect the rotating beam of acoustic energy; andlocalize the energy guide apparatus using at least the sensed rotating magnetic field and the detected rotating beam of acoustic energy; andan energy source configured to direct ablative energy to a position that is adjacent to and offset from the localized energy guide apparatus;wherein the localizing arrangement comprises an array of ultrasound transducers configured to receive the rotating beam of acoustic energy and operable in a low-intensity imaging mode to generate an image of the energy guide apparatus; andwherein the localizing arrangement is configured to localize the energy guide apparatus within the renal artery using at least the generated image. | Boston Scientific Scimed, Inc. | A61B005/05 A61B008/12 A61B008/08 A61B008/00 A61B017/22 A61B019/00 A61N007/02 A61B017/00 A61B018/00 A61N005/10 A61N007/00
US09782102 | B2 | 2017-10-10 | Electrocardiography to differentiate acute myocardial infarction from bundle branch block or left ventricular hypertrophy | 2015-03-24 | Acute myocardial infarction (AMI) is diagnosed if: (1) the ECG traces satisfy an ST Elevation Myocardial Infarction (STEMI) criterion and the ECG traces do not indicate the subject has a confounding cardiac condition, or (2) the ECG traces satisfy the STEMI criterion and the ECG traces also indicate the subject has the confounding cardiac condition and a three-dimensional vector cardiograph (3D-VCG) signal generated from the ECG traces includes an ST vector in the ST segment of the 3D-VCG signal and a terminal QRS vector of maximum magnitude in a terminal portion of the QRS complex of the 3D-VCG signal for which the angle between the ST vector and the terminal QRS vector is less than a threshold angle, e.g. in the range [130°, 170°] inclusive. The confounding cardiac condition may be bundle branch block (BBB), left ventricular hypertrophy (LVH), or interventricular conduction delay (IVCD). | 1. A cardiac monitoring system comprising: an electrocardiograph (ECG) monitor configured for operative connection with a plurality of ECG electrodes, the ECG monitor including a display device and an electronic data processing component configured to perform ECG analysis of ECG traces for a plurality of ECG leads acquired by the ECG monitor including the operations of:determining whether the ECG traces satisfy an ST Elevation Myocardial Infarction (STEMI) criterion;analyzing the ECG traces to detect a confounding cardiac condition that is not acute myocardial infarction (AMI) and that produces ST deviation;conditional on the STEMI criterion being satisfied and not detecting the confounding cardiac condition, displaying an AMI alarm on the display device of the ECG monitor;conditional on the STEMI criterion being satisfied and also detecting the confounding cardiac condition, performing the further operations of: converting the ECG traces to a three-dimensional vector cardiograph (3D-VCG) signal;computing an angle between (1) an ST vector characterizing the ST segment of the 3D-VCG signal and (2) a terminal QRS vector characterizing a terminal portion of the QRS complex of the 3D-VCG signal; andconditional on the computed angle being less than a threshold angle, displaying the AMI alarm on the display device of the ECG monitor. 2. The cardiac monitoring system of claim 1 wherein the converting operation comprises: converting the ECG traces to said 3D-VCG signal comprising a vector sum of lead vectors in a three-dimensional space wherein each lead vector has magnitude corresponding to a lead voltage and lies along an anatomical direction corresponding to a spatial orientation of the lead voltage. 3. The cardiac monitoring system of claim 2 wherein the ECG traces for the plurality of ECG leads acquired by the ECG monitor are 12-lead ECG traces and the converting operation further comprises: calculating the lead voltages comprising Frank lead voltages from the 12-lead ECG traces. 4. The cardiac monitoring system of claim 1 wherein the confounding cardiac condition includes one or more of: left bundle branch block (LBBB), right bundle branch block (RBBB), left ventricular hypertrophy (LVH), and Interventricular Conduction Delay (IVCD). 5. The cardiac monitoring system of claim 1 wherein the operation of analyzing the ECG traces to detect the confounding cardiac condition comprises: detecting the confounding cardiac condition if the QRS complex exceeds a threshold duration. 6. The cardiac monitoring system of claim 1 wherein: the operation of analyzing the ECG traces to detect the confounding cardiac condition includes analyzing the ECG traces to detect a bundle branch block (BBB) condition and analyzing the ECG traces to detect a left ventricular hypertrophy (LVH) condition; andthe AMI alarm is displayed on the display device of the ECG monitor conditional on the computed angle being less than (i) a threshold angle ThBBB if BBB is detected or (i... | KONINKLIJKE PHILIPS N.V. | A61B005/04 A61B005/0472 A61B005/00 A61B005/0432 A61B005/0408
US09326717 | B2 | 2016-05-03 | Adjustable connector and dead space reduction | 2014-04-21 | Methods and systems for determining the concentration of one or more analytes from a sample such as blood or plasma are described. The systems described herein can be configured to withdraw a certain volume of sample from a source of bodily fluid, direct a first portion of the withdrawn sample to an analyte monitoring system and return a second portion of the sample to the patient. The analyte monitoring system can be connected to the source of bodily fluid via a connector that is configured to maintain uniform velocity across the connector and reduce the dead space volume. | 1. An accommodating medical connector configured to reduce leaks and improve fluid flow between a medical device and a source of body fluid, the accommodating connector comprising: an outer housing having an elongate housing axis, a first side configured to accept standard medical tubing comprising an incoming fluid line, and a connecting side having an outer connecting structure configured to securely but removably interlock with an outer portion of a standard medical connector to form an interior volume;an elongate cylindrical tube having a stiffness greater than the standard medical tubing and having an elongate tube axis, the elongate cylindrical tube positioned so its elongate tube axis is generally aligned with and concentric to the elongate housing axis of the outer housing, an interior of the elongate cylindrical tube forming a fluid path and terminating at a port, the elongate cylindrical tube configured to extend from the accommodating connector through the outer connecting structure at the connecting side and protrude into the interior volume in a direction aligned with the elongate axis and fluid path such that the elongate cylindrical tube carries fluid through the interior volume and allows the fluid to flow out of the port and into a lumen of the standard medical connector, the interior of the elongate cylindrical tube aligned with the lumen and having a diameter that corresponds to a diameter of the lumen to improve fluid flow;a hub within the outer housing, the hub having a longitudinal hub being axially aligned with both the incoming fluid line and the elongate tube axis and positioned between them to form a secure connection such that any movement of the elongate cylindrical tube in a direction of the elongate tube axis and the elongate housing axis causes corresponding movement of the incoming fluid line, the fluid path passing through the hub; anda spring configured to exert force on the elongate cylindrical tube or hub in a direction aligned with the fluid path such that the elongate cylindrical tube protrudes into the interior volume and resiliently abuts the lumen of the standard medical connector such that when fluid passes through the interior volume and out of the port, it continues directly into the aligned lumen of the standard medical connector, minimizing leakage. 2. The accommodating connector of claim 1, wherein the standard medical connector comprises a standard Luer connector, and wherein the accommodating connector is configured to connect with a fluid line attached to a patient. 3. The accommodating connector of claim 2, wherein the fluid line comprises a central venous catheter. 4. The accommodating connector of claim 2, wherein the fluid line comprises a peripheral venous catheter. 5. The accommodating connector of claim 2, wherein the fluid line has a diameter of up to 16 gauge. 6. The accommodating connector of claim 1, wherein the standard medical connector is selected from a group consisting of conn... | OptiScan Biomedical Corporation | A61B005/05 A61B005/15 A61B005/00 A61B005/155 A61B005/145 A61B005/1455 A61B005/157 A61M039/10 A61B005/153
US09351649 | B2 | 2016-05-31 | System and method for determining video-based pulse transit time with time-series signals | 2013-09-13 | What is disclosed is a system and method for determining a subject of interest's arterial pulse transit time from time-varying source signals generated from video images. In one embodiment, a video imaging system is used to capture a time-varying source signal of a proximal and distal region of a subject of interest. The image frames are processed to isolate localized areas of a proximal and distal region of exposed skin of the subject. A time-series signal for each of the proximal and distal regions is extracted from the source video images. A phase angle is computed with respect to frequency for each of the time-series signals to produce respective phase v/s frequency curves for each region. Slopes within a selected cardiac frequency range are extracted from each of the phase curves and a difference is computed between the two slopes to obtain an arterial pulse transit time for the subject. | 1. A method for determining an arterial pulse transit time between a proximal and distal region of a subject of interest from source video images acquired using a video imaging system, the method comprising: receiving time-varying source images acquired over at least one channel of a video imaging system, said source images comprising video images captured of a proximal and distal region of an area of exposed skin of a subject of interest in which plethysmographic signals can be registered and at least one background region in which plethysmographic signals cannot be registered, and said proximal and distal regions and said background region being exposed to the same environmental factors;extracting, from said source images a first time-series signal for said proximal region and a second times series signal for said distal region and at least one background region time-series signal;subtracting said first time-series signal and said background region times-series signal to obtain environmentally compensated time-series signal for said proximal region;subtracting said second time-series signal and said background region times-series signal to obtain environmentally compensated time-series signal for said distal region;computing a first phase angle φ1(ω) with respect to frequency ω of said environmentally compensated time-series signal for said proximal region;computing a second phase angle φ2(ω) with respect to frequency ω of said environmentally compensated time-series signal for said distal region;plotting said computed phases with respect to frequency to obtain respective phase vs frequency curves for each of said proximal and distal regions;computing respective slopes T1=∂φ1/∂ω and T2=∂φ2/∂ω within a defined cardiac frequency range between 0.75 to 4.0 Hz from each of said respective phase vs frequency curves;computing a difference between said slopes T1 and T2, said difference comprising an arterial pulse transit time for said subject between said proximal and distal regions; andstoring, using a processor, said arterial pulse transit time to a storage device. 2. The method of claim 1, wherein said time-varying source signal comprises any combination of: NIR images, RGB images, RGB with NIR images, multispectral images, thermal images, and hyperspectral images. 3. The method of claim 1, further comprising determining, from said arterial pulse transit time, any of: a blood pressure in said subject's vascular network, a blood vessel dilation over time, a blood vessel blockage, a blood flow velocity, and the existence of a peripheral neuropathy. 4. The method of claim 1, further comprising: processing said video images to determine a region of skin; andisolating said proximal and distal regions from said skin regions. 5. The method of claim 1, wherein said video images are acquired by a single NIR channel, further comprising: partitioning frames from said channel for each region into a non-skin region and a skin region; andperforming source se... | Xerox Corporation | A61B005/024 A61B005/00 A61B005/021 G06T007/00
US09326886 | B2 | 2016-05-03 | Fluid filled implants for treating obstructive sleep apnea | 2009-10-29 | An implant for treating obstructive sleep apnea includes a flexible chamber, a fluid reservoir in fluid communication with the flexible chamber, and a fluid transfer assembly in communication with the fluid reservoir and the flexible chamber for transferring fluid therebetween for selectively modifying the rigidity of the flexible chamber. The flexible chamber is implantable within the soft tissue of an oropharyngeal airway of a patient, such as within the tongue, the soft palate, or the pharyngeal wall. The fluid reservoir and the fluid transfer assembly are implantable within the inframandibular region of the patient. The fluid transfer assembly is selectively engageable by the patient for transferring fluid between the two chambers for modifying the rigidity, flexibility, and/or shape of the flexible chamber with minimal or no change to the volume of the implant. | 1. An implant for treating obstructive sleep apnea comprising: a first chamber implantable within soft tissue of an oropharyngeal airway;a second chamber in fluid communication with said first chamber;a fluid transfer assembly in fluid communication with said first and second chambers for transferring fluid between said first and second chambers for modifying the rigidity of said first chamber, wherein said fluid transfer assembly adapted to be is squeezed for drawing fluid from said second chamber and introducing said drawn fluid into said first chamber, wherein said first chamber comprises a flexible chamber and said second chamber comprises an expandable fluid reservoir, and wherein said first chamber becomes more rigid and said expandable fluid reservoir shrinks as said fluid transfer assembly draws said fluid from said second chamber and introduces said fluid into said first chamber, and wherein said first chamber becomes less rigid and said second expandable fluid reservoir expands as said fluid transfer assembly removes said fluid from said first chamber and returns said fluid back to said second chamber. 2. The implant as claimed in claim 1, wherein said first chamber, said second chamber, and said fluid transfer assembly comprise a unitary structure, and wherein said implant does not have a fluid inlet. 3. The implant as claimed in claim 1, wherein said first chamber is implantable in a tongue. 4. The implant as claimed in claim 1, wherein said first chamber is implantable within the soft tissue at any selected angle relative to an anterior-posterior, lateral, vertical, or horizontal axis of a patient. 5. The implant as claimed in claim 1, wherein said first chamber is implantable within a soft palate. 6. The implant as claimed in claim 1, wherein said first chamber is implantable within a pharyngeal wall. 7. The implant as claimed in claim 1, wherein said second chamber and said fluid transfer assembly are implantable within an inframandibular region of a patient. 8. The implant as claimed in claim 1, wherein said fluid comprises a liquid or a gas. 9. An implant for treating obstructive sleep apnea comprising: a flexible chamber;a fluid reservoir in fluid communication with said flexible chamber;a fluid transfer assembly in fluid communication with said fluid reservoir and said flexible chamber for transferring fluid therebetween so as to selectively modify the rigidity of said flexible chamber, wherein said fluid transfer assembly adapted to be is squeezed for drawing fluid from said fluid reservoir and introducing said drawn fluid into said flexible chamber, wherein said flexible chamber becomes more rigid as said fluid transfer assembly draws said fluid from said fluid reservoir and introduces said fluid into said flexible chamber, wherein said flexible chamber becomes less rigid and the volume of said fluid in said fluid reservoir expands as said fluid transfer assembly removes said fluid from said flexible chamber and retur... | Ethicon, Inc. | A61F005/37 A61F013/00 A61M005/00 A61F002/06 A61B005/103 A61B005/117 A61F005/56 A61B005/00 A61F005/00
US09341627 | B2 | 2016-05-17 | Protein structural biomarkers to guide targeted chemotherapies | 2010-02-19 | A rapid, infrared spectroscopic method has been developed to assess the efficacy of targeted chemotherapeutics against the structure of the polypeptide target, based on the effect of natural polymorphic sequence variation on the conformation of the protein. This method has an advantage over the current genomics-based screening, as the new method provides a direct readout of the structural, and hence functional, outcome of polymorphisms to the protein region targeted by drugs. It allows rapid measurement of a protein's susceptibility to therapeutic targeted agents, prior to using the drug as treatment in the patient. This method can be used to identify biomarkers for a response for a protein to a drug which can be readily tested, interpreted, and used in a clinical setting. | 1. An in vitro analytical method for determining the response of one or more polymorphic variants x of a protein from a patient to a drug compound, wherein the drug compound inhibits at least one polymorphic variant y of the protein, said method comprising the steps of: (a) measuring the conformation of each polymorphic variant x in the absence of drug compound, of polymorphic variant y in the absence of drug compound, and of polymorphic variant y when bound to the drug compound by vibrational spectroscopy using a spectrometer;(b) determining whether the vibrational spectrum in an amide I region of each polymorphic variant x directly overlays onto the amide I band frequency when superimposed to: (i) the vibrational spectrum of polymorphic variant y, in the absence of the drug compound, or (ii) the vibrational spectrum of polymorphic variant y when bound to the drug compound;(c) resolving whether each of one or more polymorphic variants x: (i) will not be resistant to the therapeutic effect of the drug compound if the vibrational spectrum in the amide I region directly overlays onto the amide I band frequency when superimposed to the vibrational spectrum of polymorphic variant y, in the absence of the drug compound; or (ii) will be resistant to the therapeutic effect of the drug compound if the vibrational spectrum in the amide I region directly overlays onto the amide I band frequency when superimposed to the vibrational spectrum of polymorphic variant y, when bound to the drug compound; and(d) administering to the patient an effective amount of the drug compound identified in step (c)(i),wherein the vibrational spectroscopic spectrum in the amide I region is determined as absorption or transmittance at one or more wavelengths. 2. The method of claim 1, wherein step (a) is conducted by infrared or Raman spectroscopic measurement. 3. The method of claim 1, wherein step (a) is conducted by Fourier-transform infrared spectroscopy. 4. The method of claim 1, wherein the drug compound is a chemotherapeutic agent. 5. The method of claim 4, wherein the drug compound is a chemotherapeutic agent that acts as an allosteric inhibitor of the protein. 6. The method of claim 5, wherein the drug compound inhibits mitosis. 7. The method of claim 1, wherein the protein is an ATPase. 8. The method of claim 1, wherein the ATPase is selected from the group consisting of muscle myosin II, Na+,K+-ATPase, P glycoprotein, H+,K+-pump, vacuolar H+-ATPase, and DNA topoisomerase II. 9. The method of claim 1, wherein the protein is a kinesin selected from the group consisting of Eg5, Kif5A, Kif21A, Kif1B, and Kif6. 10. The method of claim 1, wherein the protein is kinesin Eg5, and wherein the drug compound inhibits kinesin Eg5. 11. The method of claim 10, wherein the drug is selected from the group consisting of monastrol, S-trityl-L-cysteine and its derivatives, quinazolines (e.g., ispinesib), adociasulfates, thetrahydroishquinolines, dihydropyrazoles, thiophenes, ... | Sunyoung Kim | G01N033/574 C12Q001/68 A61B005/00
US09347765 | B2 | 2016-05-24 | Real time SD-OCT with distributed acquisition and processing | 2007-10-05 | The present invention relates to distributed computing system to acquire, process, and display OCT images at real-time frame rates. | 1. A method for computing optical image data comprising the steps of: providing at least one graphics processing unit in communication, through a communication link, with at least one central processing unit for processing optical and ultrasound image data from a combined ultrasound (US)/optical coherence tomography (OCT) system, wherein the optical image data is generated by a swept-source OCT instrument;using the at least one graphics processing unit to perform a Fourier Transformation processing step on the optical image data;using the at least one central processing unit to perform a data acquisition step, a pre-processing step, and any other processing steps on the optical image data, wherein data acquisition step comprises synchronizing the swept-source optical coherence tomography instrument by clocking a light source, wherein one of the at least one central processing units performs at least one task on the optical image data and at least one task on the ultrasound image data; andtransmitting the processed optical image data and ultrasound image data to a display device to thereby display images at real-time frame rates. 2. The method of claim 1, wherein the any other processing steps comprise the steps of post-processing, image display, and post-display interaction. 3. The method of claim 1, wherein data acquisition comprises triggering instrumentation and digitization of the optical imaging signal. 4. The method of claim 1, wherein pre-processing comprises noise removal, windowing and apodization, and interpolation to linear wavenumber space. 5. The method of claim 2, wherein post-processing further comprises at least one of filtering, edge detection, feature extraction, and mapping polar representations. 6. The method of claim 2, wherein post-display interaction comprises at least one of requesting additional post-processing, interactive processing, measurement of image features, and playback control. 7. The method of claim 1, further comprising using the at least one graphics processing unit and the at least one central processing unit for performing ultrasound imaging. 8. A processing system for computing image data comprising: a combined ultrasound (US) and swept-source optical coherence tomography (OCT) system configured to generate ultrasound image data and optical image data;a graphics processing unit configured to perform a Fourier Transformation processing step on the optical image data; anda central processing unit configured to perform data acquisition, pre-processing, and all other processing steps on the optical image data, wherein data acquisition comprises synchronizing the swept-source optical coherence tomography instrument by clocking a light source, wherein the central processing unit performs at least one task for both the ultrasound image data and the optical image data;a display for displaying the processed optical image data; anda communication link providing data communication between the graphics processi... | Volcano Corporation | G01B009/02 G01N021/47 A61B005/00 G06T011/00 G06T001/20 A61B019/00
US09339277 | B2 | 2016-05-17 | Computer-assisted hip replacement surgery | 2003-10-03 | A CAS system and method for guiding an operator in inserting a femoral implant in a femur as a function of a limb length and orientation of the femoral implant with respect to the femur, comprising a reference tool for the femur, a registration tool, a bone altering tool and a sensing apparatus. A controller is connected to the sensing apparatus to: i) register a frame of reference of the femur by calculating surface information provided by the registration tool as a function of the position and orientation of the registration tool provided by the sensing apparatus, and/or retrieving in a database a model of the femur; ii) calculate a desired implant position with respect to the frame of reference as a function of the limb length; and iii) calculate a current implant position and orientation in relation to the desired implant position with respect to alterations being performed in the femur with the bone altering tool, as a function of the position and orientation of the bone altering tool provided by the sensing apparatus and of a digital model of a femoral implant provided by the database. The database is connected to the controller for the controller to store and retrieve information relating to an operation of the controller. The computer-assisted system may be used to guide an operator in inserting a pelvic implant in an acetabulum as a function of an orientation of the pelvic implant with respect to the pelvis. | 1. A method of doing surgical treatment with a tracking system in computer-assisted surgery for guiding an operator in inserting a femoral implant of a hip joint implant in a femur as a function of a limb length and orientation of the femoral implant, comprising: obtaining a frame of reference of the femur, the frame of reference being trackable in space;obtaining a digital model of a geometry of a femoral implant, and a digital model of a geometry of a tooling end of a rasp;calculating a desired implant insertion for the femoral implant with respect to the frame of reference of the femur, as a function of the limb length;calculating alterations being made to a femur with a real-time tracking of the rasp altering the femur for receiving the femoral implant and with the digital model of the geometry of the tooling end of the rasp; anddetermining a current implant insertion with respect to the desired implant insertion, the current implant insertion being calculated as a function of the digital model of the femoral implant and the calculated alterations. 2. The method according to claim 1, wherein determining the current implant insertion includes determining at least one of the varus/valgus angle of the femoral implant, the anteversion of the femoral implant, and an offset of the femoral implant. 3. The method according to claim 2, wherein the varus/valgus angle of the femoral implant is calculated as a function of an axis of the intramedullary canal in relation to an axis of the femoral implant with respect to the frame of reference, and obtaining the axis of the intramedullary canal is performed by: performing an opening in the femur;providing a tool trackable in space, said tool having a leading end thereof being insertable in a determined way with respect to a surface of the intramedullary canal; andobtaining the axis of the intramedullary canal with respect to the frame of reference by calculating and relating reference points in the intramedullary canal by inserting the leading end of the tool at given depths in the intramedullary canal and calculating a reference point of the intramedullary canal for each said given depths as a function of a tracking of said tool having the leading end inserted in said determined way. 4. The method according to claim 1, wherein determining the current implant insertion is repeated following a change in the desired implant insertion caused by a change in an acetabular center of rotation. 5. The method according to claim 1, further comprising guiding the operator in inserting the femoral implant in the intramedullary canal by tracking the femoral implant during the insertion. 6. The method according to claim 1, wherein obtaining the frame of reference of the femur is intraoperative. 7. The method according to claim 6, wherein the frame of reference of the femur includes a center of the femoral head, and obtaining the center of the femoral head is performed by: positioning a reference system trackable in... | ORTHOSOFT HOLDINGS INC. | A61B019/00 A61B017/16 A61B005/06 A61B005/103 A61B017/15 A61B017/17 G09B023/28 G09B023/30 A61B005/00 A61F002/34 A61F002/46 A61F002/30 A61F002/36
US09345418 | B2 | 2016-05-24 | EEG net with transmission capabilities | 2008-03-27 | An electroencephalography system is provided. The electroencephalography system includes a support structure that covers at least predetermined areas on a patient's head, a plurality of electrodes mounted on the support structure so that, the electrodes are distributed around the patient's head and measure usable electrical signals that are representative of electrical activity or activity in a patient's head, a plurality of transmitting wires incorporated into the support structure, a multi-pin connector having an output and an input, each of the plurality of transmitting wires directly connected between a respective one of each of the plurality of electrodes and a corresponding pin of the multi-pin connector so that, electrical signals are sent to the output of the connector and a first transmitter that is capable of being operatively coupled to the output of the connector. | 1. An electroencephalography system comprising: a support structure that is adapted to cover at least predetermined areas on a patient's head when it is mounted on a patient's head;a plurality of electrodes that are mounted on the support structure so that, when the support structure is mounted on a patient's head, the electrodes are distributed around the patient's head and are capable of measuring usable electrical signals that are representative of electrical activity or activity in a patient's head;a plurality of transmitting wires incorporated into the support structure;at least one multi-pin connector having an output and an input, with the plurality of transmitting wires hard-wired into the multi-pin connector and where each of the plurality of transmitting wires is directly connected between a respective one of each of the plurality of electrodes and a corresponding pin of the input of the multi-pin connector so that, when the headpiece is mounted on a patient's head and when the electrodes are detachably coupled to the predetermined portions of a patient's head, electrical signals that are representative of electrical brain activity of the patient are sent to the output of the connector; anda first transmitter that is capable of being operatively coupled to the output of the connector. 2. The system of claim 1, wherein each one of the plurality of electrodes are detachably coupled to the support structure. 3. The system of claim 1, where in the support structure includes at least one strap. 4. The system of claim 1, wherein the first transmitter is adapted to transmit the electrical signals through cellular transmission. 5. The system of claim 1, wherein the first transmitter is mounted directly on the support structure. 6. The system of claim 5, further comprising a second transmitter, the first transmitter being designed to send electrical signals representative of brain activity inside the patient's head to the second transmitter which is designed to retransmit the electrical signals to a remote location where the signals can be shown on a display and viewed by medical personnel. 7. The system of claim 6, wherein the second transmitter is adapted to retransmit the electrical signals to the remote location through cellular transmission. 8. The system of claim 6, wherein the second transmitter is adapted to retransmit the electrical signals to the remote location through the internet. 9. The system of claim 1, wherein the connector is adapted to convey electrical signals representative of an EEG measurement, audio signals, and/or an EKG measurement to its output. 10. The system of claim 1, wherein each one of the plurality of electrodes comprise needle electrodes, disc electrodes or pronged electrodes. 11. The system of claim 1, wherein each one of the electrodes is plated with stainless steel, platinum, ruthenium, silver, silver chloride, tin or gold. 12. The system of claim 1, wherein each electrode comprises a plurality of... | Dignity Health | A61B005/0476 A61B005/0478 A61B005/00
US09339222 | B2 | 2016-05-17 | Particle-containing membrane and particulate electrode for analyte sensors | 2013-05-31 | Systems and methods of use involving sensors having a particle-containing domain are provided for continuous analyte measurement in a host. In some embodiments, a continuous analyte measurement system is configured to be wholly, transcutaneously, intravascularly or extracorporeally implanted. | 1. A system for continuous in vivo measurement of an analyte in a host, the system comprising: a continuous analyte sensor configured for implantation in a host and configured to generate a signal indicative of an analyte concentration; anda layer disposed over at least a portion of the sensor, wherein the layer comprises a particle-containing domain comprising a metal, wherein the particle-containing domain is configured to react with at least one interfering species. 2. The system of claim 1, wherein the particle-containing domain is more distal to the sensor than other domains of the layer. 3. The system of claim 1, wherein the particle-containing domain comprises at least one material selected from the group consisting of platinum, platinum-iridium, iridium, palladium, graphite, gold, silver, silver chloride, carbon, and conductive polymers. 4. The system of claim 1, wherein the metal comprises sputtered metal. 5. The system of claim 4, wherein the sputtered metal is platinum. 6. The system of claim 1, wherein the particle-containing domain comprises a conductive particles concentration that is sufficient for the particle-containing domain to function as a conductive film. 7. The system of claim 1, wherein the particle-containing domain is configured to reduce flux therethrough of hydrogen peroxide. 8. The system of claim 1, wherein the particle-containing domain is non-powered. 9. The system of claim 1, wherein the particle-containing domain is powered. 10. The system of claim 9, wherein the sensor electronics are configured to apply a potential to the particle-containing domain. 11. The system of claim 1, wherein the layer comprises an analyte-permeable polymer. 12. The system of claim 11, wherein the analyte-permeable polymer comprises a hydrophilic polymer. 13. The system of claim 11, wherein the analyte-permeable polymer comprises at least one of polyurethane or silicone. 14. The system of claim 1, wherein the interfering species interferes with measurement of the analyte concentration to produce a signal that does not accurately represent the analyte concentration. 15. The system of claim 1, wherein the interfering species is a reactive metabolic species. 16. The system of claim 1, wherein the interfering species is a reactive oxygen species. 17. The system of claim 1, wherein the interfering species is a reactive nitrogen species. 18. The system of claim 1, wherein the interfering species is H2O2. | DexCom, Inc. | A61B005/00 A61B005/1473 A61B005/145 A61B005/1486
US09326710 | B1 | 2016-05-03 | Contact lenses having sensors with adjustable sensitivity | 2012-09-20 | Apparatus and methods employing contact lenses having electrochemical sensors that have adjustable analyte-sensing sensitivity are provided. In some aspects, level of the analyte or biological substance is detected and a number of sub-electrodes of the sensor are turned on or off based on the level of the analyte or biological substance. The sensitivity of the sensor is adjusted based on the number of sub-electrodes that are turned on or off. In some aspects, the current level from the sensor is employed in determining the number of sub-electrodes to turn on or off. In one aspect, a contact lens includes: a substrate; a power component; a circuit and an adjustment circuit. The circuit can include a potentiostat; and a sensor electrically coupled to the potentiostat, and configured to sense a level of an analyte. The sensitivity of the sensor can be adjustable by the adjustment circuit. | 1. A contact lens, comprising: a substrate;a power component, disposed on or within the substrate, configured to output power;a potentiostat, disposed on or within the substrate, and configured to be powered by the power component;a sensor, disposed on or within the substrate, electrically coupled to the potentiostat, and configured to sense a level of an analyte in an eye over which the contact lens is configured to be worn, wherein the sensor comprises a working electrode having a plurality of sub-electrodes, and wherein a sensitivity of the sensor is adjustable; andan adjustment circuit, disposed on or within the substrate, electrically coupled to the sensor, and configured to control the sensitivity of the sensor by turning one or more of the plurality of sub-electrodes on or off to the potentiostat. 2. The contact lens of claim 1, wherein the adjustment circuit is further configured to control the sensitivity of the sensor by initiation of calibration of the sensor prior to operation of the contact lens. 3. The contact lens of claim 1, wherein the adjustment circuit is further configured to dynamically control the sensitivity of the sensor based, at least, on feedback associated with a level of current output from the sensor. 4. The contact lens of claim 1, wherein the adjustment circuit is further configured to dynamically control the sensitivity of the sensor based, at least, on feedback associated with the level of the analyte being sensed by the sensor. 5. The contact lens of claim 1, wherein the contact lens further comprises a measurement circuit configured to measure a current signal from one or more of the plurality of sub-electrodes switched on to the potentiostat. 6. A contact lens, comprising: an adjustment circuit configured to adjust a level of sensitivity of a sensor of the contact lens; anda measurement circuit configured to: determine a current output from the sensor, wherein the sensor comprises a plurality of sub-electrodes of a working electrode;determine whether the current meets a predefined condition for adjustment of a number of the plurality of sub-electrodes, wherein adjustment comprises activation or de-activation of at least one of the plurality of sub-electrodes; andoutput a feedback signal to the adjustment circuit in response to the predefined condition being met, wherein the feedback signal includes at least one of information indicative of a level associated with the current output from the sensor or information indicative of an adjustment to be performed by the adjustment circuit. 7. The contact lens of claim 6, wherein the information indicative of the adjustment to be performed by the adjustment circuit comprises a number of the plurality of sub-electrodes to activate or de-activate. 8. The contact lens of claim 6, wherein the predefined condition comprises the current being within a predefined range of values less than a maximum output current for which the measurement circuit is configured to perfo... | Verily Life Sciences LLC | A61B005/02 A61B005/145 A61B005/00 A61B003/10 H05K003/46 A61B005/1473 H05K001/18 H05K003/04
US09330237 | B2 | 2016-05-03 | Pattern recognition and filtering in a therapy management system | 2008-12-24 | A method of diabetes analysis includes receiving a plurality of glucose level readings for a user. A common event occurrence in at least two of the glucose level readings is determined. The at least two glucose level readings from the common event occurrence onwards in time for a time period is analyzed. A glucose level pattern formed by the at least two glucose level readings having a similar shape is determined. At least one anomalous glucose level reading having the similar shape and not conforming to the glucose level pattern is analyzed. The at least one anomalous glucose level reading is adapted to the pattern to form an adapted glucose level pattern. An insulin dosage for the time period beginning at the common event occurrence is calculated based on the adapted glucose level pattern. | 1. A method of diabetes analysis, comprising: receiving, using a processor and a glucose sensor, a plurality of glucose level readings over a plurality of days for a user, each of the plurality of glucose level readings having multiple glucose level measurements for one of the plurality of days;determining, using the processor, a common event occurrence in at least two of the glucose level readings;analyzing, using the processor, the at least two glucose level readings from the common event occurrence onwards in time for a time period;identifying, using the processor, a glucose level pattern formed by the at least two glucose level readings having a similar shape;analyzing, using the processor, an anomalous glucose level reading not conforming to the glucose level pattern to determine whether the anomalous glucose level reading has the similar shape;adjusting, using the processor, the anomalous glucose level reading to adapt to the glucose level pattern to form an adapted glucose level pattern if the anomalous glucose level reading has the similar shape;filtering out, using the processor, the anomalous glucose level reading if the anomalous glucose level reading does not have the similar shape;calculating, using the processor, an insulin dosage for the time period beginning at the common event occurrence based on the adapted glucose level pattern; andinitiating an action, wherein the action includes at least one of recommending the calculated dosage and automatically delivering the calculated insulin dosage,wherein the method is implemented on a portable medical system. 2. The method of claim 1, wherein the common event occurrence is selected from the group consisting of: breakfast, lunch, dinner, a meal bolus, a correction bolus, and a bedtime. 3. The method of claim 1, wherein the insulin dosage is for a basal insulin dosage. 4. The method of claim 1, wherein the at least one anomalous glucose level reading having the similar shape has at least one of: a greater or lesser magnitude, and a higher or lower basal glucose level than the at least two glucose level readings forming the glucose level pattern. 5. The method of claim 1, wherein adjusting, using the processor, the anomalous glucose level reading to adapt to the glucose level pattern if the anomalous glucose level reading has the similar shape includes at least one of: shifting the anomalous glucose level reading in time compared to the at least two glucose level readings forming the glucose level pattern, normalizing a magnitude of the anomalous glucose level reading toward a target range of glucose level measurements, compressing or stretching the anomalous glucose level reading in time compared to the at least two glucose level readings forming the glucose level pattern, and shifting the anomalous glucose level reading toward the target range of glucose level measurements. 6. The method of claim 1, wherein the plurality of glucose level readings for the user excludes readings from... | Medtronic MiniMed, Inc. | G06F019/00 A61B005/00 A61M005/172
US09351657 | B2 | 2016-05-31 | Cardiac activity visualization with frequency discrimination | 2013-07-19 | A method includes measuring electrical activity at multiple points on a surface of a heart of a patient. User input indicative of a spectral slice selected from a frequency band is received. Respective levels of the electrical activity within the selected spectral slice are calculated. The calculated levels are displayed on a map of the heart. | 1. A method, comprising: measuring electrical activity at multiple points on a surface of a heart of a patient using one or more sensors attached to a body of the patient positioned near the heart for receiving electro-cardiac signals from the heart;percutaneously inserting a catheter into a chamber of the heart and measuring local electrical activity at multiple points on the surface of the heart with the catheter;tracking a position of the catheter in the heart at each of the multiple points on the surface of the heart and recording the local electrical activity at each of the multiple points on the surface of the heart using a mapping and navigation system operatively connected to the catheter;receiving heart image data as well as the measured local electrical activity and the position of the catheter recorded at each of the multiple points on the surface of the heart and user input indicative of a spectral slice selected from a frequency band of the electro-cardiac signals from the heart by the one or more sensors attached to the body of the patient;calculating respective levels of the electrical activity within the selected spectral slice; anddisplaying the calculated levels of the electrical activity within the selected spectral slice on a 3-D map of the heart on a display. 2. The method according to claim 1, wherein calculating the levels of the electrical activity within the selected spectral slice comprises computing a frequency spectrum of the electro-cardiac signals at the respective multiple points. 3. The method according to claim 1, wherein displaying the levels comprises measuring respective positions of the catheter while the catheter touches the points, and displaying the levels at the respective positions on the map of the heart. 4. The method according to claim 1, wherein receiving the user input comprises receiving the selected spectral slice from a slide bar input device. 5. The method according to claim 1, wherein displaying the calculated levels comprises assigning the levels respective colors, and coloring the map of the heart in accordance with the colors. 6. An apparatus, comprising: a display;one or more sensors attached to a body of a patient positioned near a heart for receiving electro-cardiac signals from the heart;a catheter for percutaneous insertion into a chamber of a heart, the catheter being used to measure local electrical activity at multiple points on a surface of the heart;a mapping and navigation system operatively connected to the catheter and configured to track a position of the catheter in the heart at each of the multiple points on the surface of the heart and to record the local electrical activity at each of the multiple points on the surface of the heart;an image of the heart, the image being digitized as heart image dataan electro-cardio signal frequency discrimination system comprising: (i) an interface, which is configured to receive the heart image data as well as the measured local elect... | Biosense Webster (Israel) Ltd. | A61B005/00 A61B005/044 A61B005/042 A61B005/04 A61B005/0452 A61B017/00
US09449496 | B2 | 2016-09-20 | System for convergence of alarms from medical equipment | 2014-11-20 | A device for processing alarms includes a processor and a computer-readable medium storing instructions which, when executed by the processor, cause the processor to perform various operations. The operations include receiving a first alarm from a first diagnostic device and determining a location associated with a first user. When the location associated with the first user is determined to be proximate to a device, the operations further include presenting a first verbal alert indicating that the first diagnostic device is presenting the first alarm. When the location associated with the first user is determined to be not proximate to the device, the operations further include sending a first notification message to a first mobile device of the first user indicating that the first diagnostic device is presenting the first alarm. | 1. A device, comprising: a processor; anda computer-readable medium storing instructions which, when executed by the processor, cause the processor to perform operations, the operations comprising: receiving a first alarm from a first diagnostic device;determining a location associated with a first user;when the location associated with the first user is determined to be proximate to the device, presenting a first verbal alert indicating that the first diagnostic device is presenting the first alarm; andwhen the location associated with the first user is determined to be not proximate to the device, sending a first notification message to a first mobile device of the first user indicating that the first diagnostic device is presenting the first alarm. 2. The device of claim 1, further comprising: an audio output unit, for outputting the first verbal alert. 3. The device of claim 2, wherein the first verbal alert comprises a synthetic speech that is generated by the device based upon the first alarm. 4. The device of claim 1, further comprising: a wireless transceiver, wherein the determining the location associated with the first user comprises detecting the first mobile device of the first user via the wireless transceiver. 5. The device of claim 4, wherein the determining the location associated with the first user comprises: determining that the location associated with the first user is proximate to the first diagnostic device when a strength of a wireless signal from the first mobile device exceeds a predefined threshold signal strength. 6. The device of claim 1, further comprising: a radio frequency identification reader, wherein the determining the location associated with the first user comprises detecting a radio frequency identification tag of the first user via the radio frequency identification reader. 7. The device of claim 1, wherein the operations further comprise: determining a location associated with a second user;presenting the first verbal alert indicating that the first diagnostic device is presenting the first alarm, when the location associated with the second user is determined to be proximate to the device and when the first verbal alert has not previously been presented; andwhen the location associated with the second user is determined to be not proximate to the device, sending a second notification message to a second mobile device of the second user indicating that the first diagnostic device is presenting the first alarm. 8. The device of claim 1, wherein the operations further comprise: receiving a second alarm;determining that the first alarm has a priority over the second alarm;clearing the first alarm; andwhen the location associated with the first user is determined to be proximate to the device and after the first alarm is cleared, presenting an additional verbal alert that is based upon the second alarm; andwhen the location of the first mobile device is determined to be not proximate to the device and ... | AT&T Mobility II LLC | G08B005/22 G08B027/00 A61G012/00 A61B005/00 G08B021/18 G06Q050/24 G08B021/22
US09763614 | B2 | 2017-09-19 | Wearable device and system for monitoring physical behavior of a vehicle operator | 2015-11-04 | A system uses a wearable electronic device to monitoring behavior of an operator of a vehicle. The wearable device is worn on the head of the operator of the vehicle and includes one or more motion sensors. A processor of the wearable device or of a proximate portable electronic device receives data from the one or more motion sensors and detects a pattern of motion based on the data. The system then may determine a physiological state or physical behavior of the driver based on the detected patterns or motion. Detected physiological states may include head bobs associated with sleepiness. Detected physical behaviors may include dangerous behaviors such as phone usage, prolonged gazes, sudden lane changes, or hard braking or steering. | 1. A system for monitoring behavior of an operator of a vehicle, comprising: a wearable electronic device configured to be worn on the head of an operator of the vehicle, wherein the wearable electronic device comprises a motion sensor; a processor; anda non-transitory computer-readable medium containing programming instructions that, when executed by the computer, cause the processor to:receive a first set of data from the motion sensor,analyze the first set of data to detect a pattern of motion that comprises a head rotational motion of the operator of the vehicle by:determining a forward-facing direction position as a position corresponding to the most common orientation of the wearable device,determining, based on the first set of data from the motion sensor, angular displacement of the operator's head, anddetermining whether the angular displacement corresponds to a phone usage event, wherein the phone usage event remained constant for at least a threshold period of time,access a data set corresponding to a plurality of patterns of motion,use the data set corresponding to the plurality of patterns of motion to analyze the detected pattern of motion and determine a physical behavior of the operator of the vehicle based on the detected pattern of motion by determining that the operator exhibited dangerous operator behavior if the system detected a phone usage event for at least a threshold period of time, andgenerate an alert if the determined physical behavior corresponds to a dangerous operator behavior. 2. The system of claim 1, further comprising a reference device that is attached to the vehicle and that includes: a transmitter, anda reference device motion sensor that is configured to collect a second set of data corresponding to motion of the vehicle. 3. The system of claim 2, wherein the instructions are further configured to cause the processor to: receive the second set of data from the transmitter; anddetermine that the physical behavior also corresponds to a hard braking event if: the first set of data indicates that the motion sensor of the wearable electronic device is experiencing a negative acceleration above a threshold value, andat the same time, the second set of data also indicates that the motion sensor of the reference device is experiencing a negative acceleration above a threshold value. 4. The system of claim 2, wherein the instructions are further configured to cause the processor to: receive the second set of data from the transmitter; anddetermine that the physical behavior also corresponds to a sudden lane change event if: the first set of data indicates that the motion sensor of the wearable electronic device is experiencing a rate of change of direction in a z-plane relative to the earth exceeding a threshold value within a threshold time, andat the same time, the second set of data also indicates that the motion sensor of the reference device is experiencing is experiencing a rate of change of direction in th... | Maven Machines, Inc. | A61B005/18 A61B005/00 A61B005/11 B60R001/04 G08B021/24 G08B021/06
US09901251 | B2 | 2018-02-27 | System and method of workflow management | 2015-11-19 | Wireless systems and methods include a plurality of peripheral electronic devices each having a wireless communication system. A processor is configured to establish an association confidence level indicative of a likelihood that a peripheral electronic device is associated to a monitored subject for each peripheral electronic device based on association criteria. Indicators are configured to communicate the association the association confidence level. | 1. A wireless patient monitoring system, comprising: a plurality of peripheral electronic devices each having a wireless communication system and a sensor, each of the plurality of peripheral electronic devices adapted to be secured to a monitored patient, the sensors of each of the plurality of electronic devices acquire physiological data from the monitored patient, the physiological data acquired from the monitored patient being transmitted by the wireless communication device of a respective peripheral electronic device;a processor that registers each of the plurality of peripheral electronic devices to the monitored patient, the processor receives the physiological data from the respective peripheral electronic devices, the processor establishes an association confidence level for each peripheral electronic device based upon at least the received physiological data from each of the peripheral electronic devices, the association confidence level indicative of a likelihood that the physiological data acquired by the peripheral electronic device is from the monitored patient, wherein the processor compares each of the association confidence levels established for each peripheral device to a predetermined threshold and if an association confidence level is below the predetermined threshold, the processor disassociates that peripheral electronic device from the monitored patient; andat least one automated therapeutic delivery device communicatively connected to the processor;wherein the processor operates the at least one automated therapeutic delivery device based upon the physiological data from peripheral devices with association confidence levels above the predetermined threshold. 2. The wireless patient monitoring system of claim 1, wherein the plurality of peripheral electronic devices is a first plurality of peripheral electronic devices and the monitored patient is a first monitored patient, the system further comprising: a second plurality of peripheral electronic devices registered by the processor to a second monitored patient, wherein the processor further receives physiological data from the second plurality of peripheral electronic devices, establishes an association confidence level for each of the peripheral electronic devices of the second plurality of peripheral electronic devices, and uses the established association confidence levels to separate physiological data acquired from the first monitored patient from physiological data acquired from the second monitored patient. 3. The wireless patient monitoring system of claim 1, wherein the at least one automated therapeutic delivery device comprises at least one of a defibrillator, a drug pump, an intravenous solution pump, and a ventilator. 4. The wireless patient monitoring system of claim 1, wherein if the association confidence level for one or more of the plurality of peripheral devices is below the predetermined threshold, processor disables operation of the at least one... | General Electric Company | A61B005/00 A61M005/172 A61N001/39 A61B005/0402 A61B005/0476 A61B005/11 A61M016/00 A61B005/021 A61B005/024 G06F019/00 H04W004/00 A61B005/01 A61B005/0488 A61B005/145 A61B005/0205 A61B005/08
US10271799 | B2 | 2019-04-30 | Patient alarm data application | 2015-11-20 | A patient alarm application provides a way for allowing clinical personnel to monitor and control patient-specific alarm settings. A user interface allows clinical personnel to compare graphical representations of unit population vital signs data with patient-specific vital signs data as well timelines of alarms for the patient. Analysis of the alarm data allows the user interface to show clinical personnel the estimated effects of changes to alarm threshold settings. | 1. A method of managing alarms in a clinical environment, comprising: receiving alarm data for a patient in a clinical unit during a predetermined time interval;receiving, during the predetermined time interval, physiological data for the patient and physiological data for a population of the clinical unit corresponding to the patient;displaying in a patient-specific user interface a timeline of alarms for the patient corresponding to a physiological data type during the predetermined time interval;calculating a plurality of numerical time in alarm duration estimates for the patient corresponding to a plurality of alarm threshold values for the physiological data type;displaying the plurality of numerical time in alarm duration estimates for the patient in the patient-specific user interface; anddisplaying in the patient-specific user interface a graphical representation of the physiological data for the patient corresponding to the physiological data type; anddisplaying in the patient-specific user interface a graphical representation aggregating physiological data corresponding to the physiological data type for the population of the clinical unit as a whole. 2. The method of claim 1, further comprising: displaying in the patient-specific user interface a total number of alarms for the patient and total time in alarm duration for the patient corresponding to the physiological data type. 3. The method of claim 1, further comprising: displaying in the patient-specific user interface a plurality of timelines of alarms for the patient corresponding to a plurality of physiological data types during the predetermined time interval. 4. The method of claim 1, further comprising: receiving a user selection of the timeline of alarms; anddisplaying the numerical time in alarm duration estimates responsive to the user selection of the timeline of alarms. 5. The method of claim 1, further comprising: identifying a current alarm threshold for the physiological data type;displaying the current alarm threshold in the patient-specific user interface; anddisplaying alarm threshold changes in the patient-specific user interface. 6. The method of claim 1, further comprising: instantiating the patient-specific user interface automatically responsive to a clinical environment policy. 7. A non-transitory machine readable medium, on which are stored instructions, comprising instructions that when executed cause a clinical patient monitoring computer to: receive alarm data for a patient during a predetermined time interval;receive physiological data for the patient and physiological data for a population of a clinical unit corresponding to the patient during the predetermined time interval;display in a patient-specific user interface a timeline of alarms for the patient corresponding to a physiological data type during the predetermined time interval;calculate a plurality of numerical time in alarm duration estimates for the patient corresponding to a plurality o... | Medical Informatics Corporation | G08B023/00 A61B005/00
US10426362 | B2 | 2019-10-01 | Deep-brain probe and method for recording and stimulating brain activity | 2015-11-10 | A probe suitable for deep-brain recording and stimulation is provided. The probe comprises a wire bundle that includes a plurality of wires, an integrated circuit having a plurality of electrodes, and an interposer that joins the wire bundle and the integrated circuit such that each of the plurality of electrodes is electrically connected with a different wire of the plurality of wires. | 1. A deep-brain probe comprising: (A) a wire bundle comprising a first plurality of wires that are arranged in a first two-dimensional arrangement, wherein each of the first plurality of wires has a first end and a second end, and wherein each of the first plurality of wires includes: a core that is electrically conductive;an electrical insulator that surrounds the core along the length of the wire; andan electrode that is electrical-insulator-free;wherein the wire bundle has a first portion in which the first plurality of wires is mechanically joined together to collectively define a stalk, the first end of each wire of the plurality thereof terminating at a first surface of the stalk; andwherein the wire bundle has a second portion in which each wire of the first plurality thereof is unrestrained, the second end and the electrode of each wire of the plurality thereof being in the second portion;(B) a first integrated circuit comprising a plurality of bond pads that is arranged in a second two-dimensional arrangement; and(C) an interposer that includes: (i) a body having a first surface and a second surface, the body comprising a first material that is an electrical insulator, and the body being compressible, wherein the body includes;(a) a first layer comprising the first material, the first material being characterized by a first modulus of elasticity;(b) a second layer that includes the first surface; and(c) a third layer that includes the second surface, the second and third layers comprising a third material having a second modulus of elasticity that is higher than the first modulus of elasticity;wherein the first layer is between the second and third layers; and(ii) a plurality of conductors that extend through the body from the first surface to the second surface, wherein each of the plurality of conductors comprises a second material that is electrically conductive;wherein the wire bundle, first integrated circuit, and interposer are arranged such that each of the plurality of bond pads is electrically connected with each of one or more wires of the first plurality thereof via at least one of the plurality of conductors. 2. The probe of claim 1 wherein a first bond pad of the plurality of bond pads is electrically connected to a second plurality of wires via a second plurality of conductors, and wherein a second bond pad of the plurality of bond pads is electrically connected to a third plurality of wires via a third plurality of conductors, the second plurality of wires and third plurality of wires being mutually exclusive, the first plurality of wires including the second plurality of wires and the third plurality of wires, and the first plurality of conductors including the second plurality of conductors and the third plurality of conductors. 3. The probe of claim 1 wherein the second two-dimensional arrangement is characterized by an inter-bond-pad spacing in a first dimension, and wherein the first two-dimensional arrangement is c... | The Board of Trustees of the Leland Stanford Junior University THE FRANCIS CRICK INSTITUTE | A61B005/04 A61B005/00 A61N001/05 A61B005/0478 A61N001/36
US10080502 | B2 | 2018-09-25 | Solenoid valve and electronic blood pressure monitor equipped with the same | 2016-01-08 | In a solenoid valve, a bobbin and a solenoid coil are accommodated in a casing. A rod-shaped plunger is slidably inserted into the bobbin. A core provided with a flow port through which a fluid flows is disposed on one side of the casing. A valve body is attached to one end opposing the flow port of the plunger. In a non-operating period during which a solenoid coil is in a non-energized state, the valve body moves away from the flow port due to the biasing force applied by the biasing portion and another end of the plunger protrudes outward from the casing through the through hole, abuts against lock portions disposed outside of the casing, and is thereby locked. The lock portions can be elastically deformed by an external force such that separation of the plunger and the valve body from the casing is allowed. | 1. A solenoid valve capable of variably controlling a flow amount of a fluid, comprising: a casing comprising a through hole formed therethrough;a bobbin accommodated in the casing and including a lock portion extending from an outer surface of the casing;a solenoid coil wound around the bobbin;a rod-shaped plunger slidably inserted in the bobbin;a core that is provided with a flow port through which the fluid flows and is disposed on a side of the casing opposing one end of the rod-shaped plunger;a valve body that is provided on the one end of the rod-shaped plunger and is disposed so as to oppose the flow port; anda biasing portion that biases the rod-shaped plunger in a direction of moving away from the core,wherein, in a non-operating period, during which the solenoid coil is in a non-energized state, the valve body provided on the one end of the rod-shaped plunger moves away from the flow port and another end opposite to the one end of the rod-shaped plunger protrudes outward from the casing through the through hole, abuts against the lock portion disposed outside of the casing, and is thereby locked,wherein in a working period during which the solenoid coil is in an energized state, the flow amount of the fluid flowing through the flow port is adjusted due to the rod-shaped plunger and the valve body being moved inside the bobbin against a biasing force applied by the biasing portion, due to a magnetic force generated by the solenoid coil,wherein the lock portion can deform elastically due to an external force such that separation of the rod-shaped plunger and the valve body from the casing is allowed,wherein the casing has a U-shaped yoke and a plate-shaped yoke lid that closes an open end of the U-shaped yoke, and the through hole is formed in the plate-shaped yoke lid,wherein the bobbin has a main body portion accommodated in the casing, a first extended portion and a second extended portion formed integrally with the main body portion, andwherein the first extended portion and the second extended portion come into contact with a pair of opposing sides of the plate-shaped yoke lid and extend outward of the casing through the through hole parallel to a movement path of the rod-shaped plunger, and constitute the lock portion. 2. The solenoid valve according to claim 1, wherein each of the first and second extended portions includes: an arm portion that extends outward of the casing through the through hole along the movement path of the rod-shaped plunger, a leading end side of the arm portion being able to bend in a direction of moving away from the movement path of the rod-shaped plunger due to the external force; anda hook portion that extends substantially orthogonally toward the movement path of the rod-shaped plunger from the leading end of the arm portion, and wherein the hook portion locks the other end of the rod-shaped plunger during the non-operating period. 3. The solenoid valve according to claim 2, wherein the through hole ... | OMRON HEALTHCARE Co., Ltd. | A61B005/02 A61B005/0235 F16K031/06
US09333345 | B2 | 2016-05-10 | Electrical stimulation for a functional electrical stimulation system | 2014-10-03 | An electrical stimulation system and method for generating a stimulation signal for at least two electrodes coupled to a body part in a functional electrical stimulation system. One example embodiment includes a controller unit operable for receiving stimulation parameters and a trigger signal, and in response to receiving the trigger signal, outputting control signals based on the stimulation parameters. A voltage conversion module coupled to the unit receives the control signals and converts a supply voltage based on the received control signals. A switch receives the converted supply voltage at a first terminal and outputs a simulation signal at a second terminal. Outputting of the converted supply voltage at the second terminal by the switch is controlled by a driver module based on the received control signals. | 1. An electrical stimulation system for generating a stimulation signal for at least two electrodes for stimulating a body part in a functional electrical stimulation system, the electrical stimulation system comprising: a controller unit operable for receiving at least one set of stimulation parameters, receiving a trigger signal; and in response to receiving the trigger signal, outputting at least a first control signal and a second control signal based on the at least one set of stimulation parameters;a voltage conversion module coupled to the controller unit, the voltage conversion module for receiving at least the first control signal and converting a supply voltage based on the received first control signal; andat least one switch receiving the converted supply voltage at a first terminal and selectively outputting based on the second control signal a stimulation signal at a stimulation output terminal. 2. The electrical stimulation system of claim 1, further comprising a driver module controlling the at least one switch based on at least the second control signal, wherein the at least one switch is configured to output the converted voltage as the stimulation signal at the output terminal when the driver module configures the at least one switch to a closed position. 3. The system of claim 1, wherein the at least one set of stimulation parameters comprise a desired amplitude of the stimulation signal and wherein the first control signal is adjusted based on the desired amplitude to correspondingly control certain parameters of the voltage conversion module. 4. The system of claim 3, wherein the at least one set of stimulation parameters further comprise a desired rise time, a desired hold time, a desired drop time, and a desired idle time of a cycle of the stimulation signal and wherein the first control signal is adjusted based on the desired rise time, desired hold time, desired drop time and desired idle time. 5. The system of claim 1, wherein the voltage conversion module comprises: a DC/DC voltage converter having a feedback terminal and an output terminal to output the converted voltage;a feedback resistor coupling the voltage output terminal with the feedback terminal; anda variable resistor coupling the feedback terminal to a reference; and wherein the converted voltage is based on the resistance value of the feedback resistor and the resistance value of the variable resistor. 6. The system of claim 5, wherein the first control signal comprises a value of the variable resistor for outputting the desired amplitude of the stimulation signal and wherein the value of the variable resistor is varied in time according to the received first control signal. 7. The system of claim 3, wherein the at least one set of stimulation parameters further comprise a desired period and a desired pulse width of the stimulation signal and wherein the second control signal is adjusted based on the desired period and pulse width. 8. The system of c... | Ensilver Canada | A61N001/36 A61N001/378 A61B005/0488 A61B005/00 A61B005/11
US09337901 | B2 | 2016-05-10 | Medical image diagnosis apparatus | 2013-02-04 | The medical image diagnosis apparatus has a body, a power transmitting part, a power receiving part, and a device. The power transmitting part is housed in the body, has a transmission coil, and generates an oscillating field due to resonance from the transmission coil upon receiving electric power. The power receiving part has a reception coil, resonates with a frequency substantially equal to the resonant frequency of the transmission coil, and generates electric power upon receiving the oscillating field. The device houses the power receiving part, deriving organization information through operation via non-contact with the body by electric power from the power receiving part. The apparatus comprises a status checker and a controller that determines whether or not electric power should be generated depending on at least one of the statuses to control the generation of electric power by the transmission coil. | 1. A medical image diagnosis apparatus that receives signals from a test body and derives organization information thereof so as to check the test body, comprising: a body;a power transmitting circuit having a transmission coil housed in the body, receiving electric power from the outside and generating an oscillating field due to resonance from the transmission coil;a power receiving circuit having a reception coil that resonates at a frequency substantially equal to the resonant frequency of the transmission coil and generates electric power upon receiving the oscillating field generated by the transmission coil;a device that houses the power receiving circuit and operates by electric power from the power receiving circuit via non-contact with the body so as to derive the organization information;a status checker configured to derive at least one of the operation statuses of the body and the device; anda controller configured to determine whether or not the electric power should be generated depending on at least one of the statuses of the body and the device, and based on the determination results, to control the generation of the electric power by the transmission coil. 2. The medical image diagnosis apparatus according to claim 1, wherein the device is a wireless X-ray detector of an X-ray diagnosis apparatus, a high frequency coil unit of an MRI apparatus, or an ultrasound probe of an ultrasound diagnosis apparatus. 3. The medical image diagnosis apparatus according to claim 1, wherein the power transmitting circuit comprises the transmission coil and first capacitor so as to generate resonance, while the power receiving circuit comprises the reception coil and second capacitor so as to generate resonance. 4. The medical image diagnosis apparatus according to claim 1, wherein the device has a battery for storing the electric power generated by the reception coil and operates using the electric power stored in the battery. 5. The medical image diagnosis apparatus according to claim 4, further comprising an indicator configured to indicate the remaining power level of the battery. 6. The medical image diagnosis apparatus according to claim 1, comprising the two or more devices, wherein the two or more power receiving circuits included in the two or more devices receive an oscillating field generated by the one power transmitting circuit to generate electric power. 7. The medical image diagnosis apparatus according to claim 4, wherein the status checker has a first detector configured to detect the start trigger of the check, and the controller is configured to stop the supply of electric power by the power transmitting circuit according to detection of the start trigger. 8. The medical image diagnosis apparatus according to claim 7, wherein the start trigger includes at least one of the input of patient information into the body, execution of the operation for instructing start of the check, and arrangement of the device at a predeterm... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | H02J007/00 A61B005/05 H04B005/00 A61B006/00 A61B008/00 H02J007/02 H02J005/00
US20160131724 | A1 | 2016-05-12 | MRI SCANNER BORE COVERINGS | 2014-06-19 | A bore covering for sealing an MRI scanner bore to create a controlled environment within the MRI scanner bore is disclosed. The bore coverings used with climate control devices may create an incubator-like environment to allow for improved preparation, control, and imaging of infants. A system and kit including one or more bore coverings, and a method of using the bore coverings to create a controlled environment within the MRI scanner bore are also described. | 1. A bore covering for creating a controlled environment within an MRI scanner bore, wherein the bore covering comprises a covering sheet with opposed outer and inner surfaces and a fastening means attached to the inner surface, wherein the inner surface is positioned over a bore opening of the MRI scanner bore and the fastening means is reversibly attached to an exposed edge of the bore opening to form a seal over the bore opening. 2. The bore covering of claim 1, wherein the attachment means is situated near an outer edge of the bore covering. 3. The bore covering of claim 2, wherein the attachment means extend completely around the outer edge, forming a continuous closed shape. 4. The bore covering of claim 2, wherein the attachment means include one or more gaps around the outer edge. 5. The bore covering of claim 4, wherein the one or more gaps are distributed around the outer edge. 6. The bore covering of claim 4, wherein the attachment means include one gap for accessing the MRI scanner bore from outside the bore covering. 7. The bore covering of claim 1, wherein the bore covering further comprises at least one opening formed within the covering sheet and extending from the outer surface to the inner surface. 8. The bore covering of claim 7, wherein the at least one opening comprises a first opening operatively connected to a climate control device. 9. The bore covering of claim 8, wherein the at least one opening further comprises a second opening comprising a vent to exhaust air out of the MRI scanner bore. 10. The bore covering of claim 9, wherein the vent comprises a valve to exhaust an amount of air out of the MRI scanner bore when a bore air pressure exceeds a predetermined level. 11. The bore covering of claim 1, wherein at least a portion of the covering sheet is transparent. 12. The bore covering of claim 1, wherein the covering sheet further comprises a plastic material. 13. The bore covering of claim 1, wherein the fastening means comprises a pressure-sensitive adhesive. 14. The bore covering of claim 13, wherein the pressure-sensitive adhesive is chosen from a weak-tack adhesive and a low-tack adhesive. 15. A system for creating a controlled environment within an MRI scanner bore comprising: at least one bore covering, each bore covering comprising a covering sheet with opposed outer and inner surfaces, a fastening means attached to the inner surface and at least one opening formed within the covering sheet and extending between the outer and inner surfaces; andat least one climate control device for regulating at least one environmental parameter within the MRI scanner bore, wherein the at least one climate control device is operatively connected to the at least one opening of the at least one bore covering;wherein each of the at least one bore coverings is sealed over a bore opening to seal the MRI scanner bore from the outside environment. 16. The system of claim 15, further comprising a controller operatively... |  | G01R033/31 A61B005/055
US20160144155 | A1 | 2016-05-26 | OCCLUSION SHEATH FOR IMAGING CATHETER | 2014-07-01 | An occlusion sheath having an expandable member is described herein. The occlusion sheath is configured to occlude or limit blood flow to a vessel while providing a central lumen to allow a catheter, such as an imaging catheter and/or an atherectomy catheter, to travel therethrough. | 1. An occlusion sheath, comprising: an outer elongate member;an inner elongate member extending within the outer elongate member;a central lumen within the inner elongate member, the central lumen having a diameter of greater than 0.055 inches and configured to accommodate a catheter; andan occlusion balloon attached to the outer elongate member and configured to expand to occlude or reduce the flow through a vessel. 2. The occlusion sheath of claim 1, wherein a proximal end of the balloon is sealed against the outer elongate member, and wherein a distal end of the balloon is sealed against the outer elongate member or the inner elongate member. 3. The occlusion sheath of claim 1, wherein an annular inflation lumen is formed between the outer elongate member and the inner elongate member, the inflation lumen in communication with the balloon. 4. The occlusion sheath of claim 3, further comprising a hole through the outer elongate member between the proximal end of the balloon and the distal end of the balloon to allow communication between the inflation lumen and the balloon. 5. The occlusion sheath of claim 4, wherein the hole is positioned closer to a proximal end of the balloon than a distal end of the balloon. 6. The occlusion sheath of claim 1, wherein the sheath further comprises a tapered tip. 7. The occlusion sheath of claim 6, wherein the tapered tip is attached to the outer elongate member with a lap joint. 8. The occlusion sheath of claim 6, wherein the tapered tip is radiopaque. 9. The occlusion sheath of claim 1, wherein an outer diameter of the balloon when fully inflated and unconstrained is 4-20 mm. 10. The occlusion sheath of claim 1, wherein an outer diameter of the outer elongate tube is less than 0.165 inches in diameter. 11. The occlusion sheath of claim 1, further comprising a radiopaque marker configured to indicate a position of the balloon. 12. The occlusion sheath of claim 1, further comprising a hub attached to the inner and outer elongate members, the hub including a first port for passage of the catheter and a second port for passage of air or fluid to inflate the balloon. 13. The occlusion sheath of claim 1, wherein the catheter is an atherectomy catheter. 14. The occlusion sheath of claim 1, wherein the catheter is an imaging catheter. 15. A method of performing atherectomy within a vessel, comprising: inserting an occlusion sheath into a vessel;inserting an atherectomy catheter having an imaging sensor and a cutter through a central lumen of the occlusion sheath;inflating a balloon of the occlusion sheath to reduce the flow through the vessel;imaging the vessel with the imaging sensor while the vessel flow is reduced; andcutting tissue inside the vessel with the cutter. 16. The method of claim 15, wherein inflating the balloon comprises injecting fluid or air into an annular space between an inner and outer elongate member of the occlusion sheath such that the fluid or air flows down the annular spac... |  | A61M025/10 A61B017/3207 A61B005/00
US09345423 | B2 | 2016-05-24 | Method for measuring breath alcohol concentration and apparatus therefor | 2013-06-18 | A method and apparatus for measuring user breath alcohol concentration. A flow of an expired breath sample is passed through a fuel cell sensor giving an output signal proportional to the amount of alcohol present in the sample. By measuring pressure, the volume of the sample may be calculated by integrating pressure over expiration time of the sample, whereas breath alcohol concentration is calculated based on the fuel cell output signal. Both sample volume and breath alcohol concentration values are continually updated by integrating measured instantaneous pressure and fuel cell output signal over time, irrespective of breath sample volume. When the user stops blowing, volume compensation is performed to obtain a compensated fuel cell output signal using a stored calibration volume. Hence, an improved method for accurately measuring breath alcohol concentration of a user is achieved, capable of handling varied expired volumes of breath, obviating the need for sampling mechanism. | 1. A method for measuring breath alcohol concentration (BrAC) of a user, comprising the steps of: receiving a flow of an expired breath sample from the user;measuring the instantaneous pressure of the flow of the expired breath sample;recording the expiration time of the breath sample;leading the breath sample into a fuel cell sensor; andcalculating the breath alcohol concentration (BrAC) based on an output signal (FCout) of the fuel cell sensor;calculating the volume (Vb) of the breath sample based on the measured pressure;updating continually the breath sample volume (Vb) and the breath alcohol concentration (BrAC) by integrating the measured instantaneous pressure and the fuel cell output signal (FCout) over time, irrespective of the breath sample volume (Vb);calculating the flow rate (Q) of the breath sample based on the breath sample volume (Vb) and the recorded expiration time; andwhen the user stops blowing, performing the following steps before calculating the final breath alcohol concentration (BrAC): compensating the fuel cell output signal (FCout) using a stored calibration volume (Vcal) to obtain a volume compensated fuel cell output signal (FCVcomp),compensating the volume compensated fuel cell output signal (FCVcomp) to obtain a flow rate compensated fuel cell output signal (FCQcomp) using a stored flow rate adjustment factor (Qf) corresponding to the calculated flow rate (Q)measuring a temperature (T); andcompensating the compensated fuel cell output signal using a stored temperature adjustment factor (Tf) corresponding to the measured temperature. 2. The method according to claim 1, further comprising the steps of, if no measurements have been made for a predetermined period of time: performing a calibration by making a measurement of a sample of predetermined volume and concentration;repeating the calibration step at least once; andstoring the average value of the fuel cell output signal (FCout) as the calibration volume (Vcal). 3. The method according to claim 1, further comprising the step of: determining the blood alcohol concentration (BAC) based on the breath alcohol concentration (BrAC). 4. The method according to claim 3, further comprising the step of: displaying the resulting blood alcohol concentration (BAC). 5. The method according to claim 1, wherein said volume compensating is performed using the formula:FCcomp=FCout·Vca⁢⁢lVb. 6. The method according to claim 1, further comprising the step of: preventing start-up of a vehicle if the calculated breath alcohol concentration (BrAC) exceeds a predetermined threshold value. 7. The method according to claim 1, wherein the pressure is measured by means of a pressure-based pressure sensor. 8. Apparatus for measuring breath alcohol concentration (BrAC), comprising: means for receiving an expired breath sample of a user;means for measuring the instantaneous pressure of the flow of the expired breath sample;means for recording the expiration time of the breath sample;a f... | Alco Systems Sweden AB | G01N033/497 A61B005/097 A61B005/085 G01N033/98 G01N033/48 A61B005/08 A61B005/087 A61B005/00 B60W040/08
US09339619 | B2 | 2016-05-17 | Pressure adjustment method for CPAP machine | 2015-03-30 | A positive airway pressure apparatus is automatically adjusting. Pressure increases in response to apnea events when the apparatus is in one or more responsive states. Pressure does not increase in response to apnea events when the apparatus is in a non-responsive state. The apparatus switches between responsive and nonresponsive states depending upon any of a number of different criteria that help differentiate between open airway apnea events and closed airway apnea events. | 1. A method for controlling an automatically adjusting PAP apparatus, comprising: operating a controller in a first responsive state in which the controller is configured to respond to a first series of non-isolated apneas by increasing an output pressure of a gases flow source;entering a pause state when a total number of pressure increases in response to non-isolated apneas exceeds a threshold, the pause state having a predefined duration, wherein during the pause state, the controller increases the output pressure of the gases flow source in response to flow-limited breaths and non-flow limited hypopneas but does not increase the output pressure in response to apneas; andafter the pause state, operating the controller in a second responsive state in which the controller is configured to respond to a second series of non-isolated apneas by increasing an output pressure of a gases flow source. 2. The method of claim 1, wherein increasing the output pressure of the gases flow source comprises increasing the output pressure by 1 cm H2O in response to each non-isolated apnea. 3. The method of claim 1, further comprising entering a suppressed state when a total number of pressure increases in response to non-isolated apneas since exiting the pause state exceeds a second threshold, wherein the controller is configured to remain in the suppressed state until the output pressure falls below a threshold, wherein during the suppressed state the controller is configured to not increase the output pressure in response to apneas but is configured to increase the output pressure in response to flow-limited breaths and non-flow limited hypopneas. 4. The method of claim 1, wherein the controller is configured to enter the pause state when a total number of non-isolated apneas or non-flow limited hypopneas less nominal periods of normal breathing exceeds the threshold. 5. The method of claim 1, further comprising sensing pressure or flow or both supplied by the gases flow source or supplied to a patient. 6. The method of claim 5, further comprising determining apneas, non-flow limited hypopneas, and flow limited breaths based on the sensed pressure or flow. 7. The method of claim 1, wherein the controller is configured to increase the output pressure in response to flow-limited breaths subject to a maximum therapy pressure. 8. A method for controlling an automatically adjusting PAP apparatus, comprising: operating a controller in a first responsive state in which the controller is configured to respond to a first series of non-isolated apneas by increasing an output pressure of a gases flow source, wherein in the first responsive state the controller is also configured to respond to non-flow limited hypopneas by increasing the output pressure; andentering a pause state when a total number of pressure increases in response to non-flow limited hypopneas and non-isolated apneas less nominal periods of normal breathing exceeds a threshold, the pause state ha... | Fisher & Paykel Healthcare Limited | A61M016/00 A61B005/04 A61B005/00 A61M016/16
US09355474 | B2 | 2016-05-31 | Method of processing optical coherence tomography images | 2013-06-28 | An automatic or semiautomatic method of processing intravascular Optical Coherence Tomography (OCT) images, comprising: pre-processing the primary image by segmenting the guidewire in the lumen to exclude it from the image, determining whether the lumen is open or closed; if the lumen is determined to be closed, performing a first automatic lumen segmentation procedure on the pre-processed image to define a perimeter line of the lumen, corresponding to the inner edge of the vessel wall, and determine a centroid CL of the lumen area; if the lumen is determined not to be closed, sequentially performing an automatic lumen closing procedure on the pre-processed image, and a second automatic lumen segmentation procedure, to define a perimeter line of the lumen, corresponding to the inner edge of the vessel wall, and determine a centroid CL of the lumen area; automatically finding an outer edge of the vessel wall and defining a first region of interest ROI1 between the inner edge and the outer edge of the lumen; automatically segmenting a fibrous plaque in the first region of interest ROI1 thereby defining a second region of interest ROI2 that contains the fibrous plaque, and segmenting a hyporeflective plaque, contained in a radial region that falls within ROI1 and is external to ROI2. | 1. A semiautomatic method of processing intravascular Optical Coherence Tomography (OCT) images, comprising: receiving in a computer a primary cross-sectional OCT image visualising a section of a blood vessel having a lumen and a vessel wall;pre-processing the primary image by the computer, wherein the pre-processing step comprises segmenting a guide wire in the lumen and excluding the guide wire from the primary image to obtain a pre-processed image;determining whether the lumen is closed;if the lumen is determined to be closed, performing a first automatic lumen segmentation procedure on the pre-processed image to define a perimeter line of the lumen, corresponding to the inner edge of the vessel wall, and determine a centroid CL of the lumen area;if the lumen is determined not to be closed, sequentially performing an automatic lumen closing procedure on the pre-processed image, and a second automatic lumen segmentation procedure, to define a perimeter line of the lumen, corresponding to the inner edge of the vessel wall, and determine a centroid CL of the lumen area;automatically finding an outer edge of the vessel wall and defining a first region of interest ROI1 between the inner edge and the outer edge of the lumen;automatically segmenting a fibrous plaque in the first region of interest ROI1, thereby defining a second region of interest ROI2 that contains the fibrous plaque, andsegmenting a hyporeflective plaque, contained in a radial region that falls within ROI1 and is external to ROI2. 2. The method of claim 1, wherein determining if the lumen is closed is carried out by an automatic process. 3. The method according to claim 1, wherein segmenting a hyporeflective plaque is carried out by a semi-automatic process in which a first detection of the image region containing the hyporeflective plaque is manually carried out by an operator. 4. The method according to claim 1 wherein, in the step of pre-processing the primary image, the step of segmenting the guide wire comprises: creating a sub-image in a central portion of the primary image;transforming said sub-image into a binary image using a digital binarization filter that defines a light intensity threshold value or a range of light intensity threshold values to generate a sample image highlighting one or more regions having a high light intensity relative to a background;applying a morphological filter to the sample image, which filter sequentially uses a morphological operator of recognition and extraction of objects having a round shape and a morphological erosion operator to eliminate small non-round objects so as to identify a round object associated with the guide wire; andcalculating the area of the round object so identified and calculating a local centroid CW of said area. 5. The method as claimed in claim 4 wherein, in the step of pre-processing the primary image, the step of excluding the guide wire from the primary image comprises: restoring the primary image, andsubtrac... | Consiglio Nazionale delle Ricerche Fondazione Toscana Gabriele Monasterio per IA Ricerca Medica e di Sanita Pubblica | G06K009/00 G06T011/00 A61B005/00 G06T007/00 G06T007/60 A61B005/02
US09326666 | B2 | 2016-05-03 | Multi-channel medical imaging system | 2013-11-18 | A medical imaging system provides simultaneous rendering of visible light and fluorescent images. The system may employ dyes in a small-molecule form that remain in a subject's blood stream for several minutes, allowing real-time imaging of the subject's circulatory system superimposed upon a conventional, visible light image of the subject. The system may provide an excitation light source to excite the fluorescent substance and a visible light source for general illumination within the same optical guide used to capture images. The system may be configured for use in open surgical procedures by providing an operating area that is closed to ambient light. The systems described herein provide two or more diagnostic imaging channels for capture of multiple, concurrent diagnostic images and may be used where a visible light image may be usefully supplemented by two or more images that are independently marked for functional interest. | 1. A system comprising: a light source providing visible light, a first excitation wavelength and a second excitation wavelength, the light source adapted to illuminate a subject; a lens that captures light from the subject;a first dichroic mirror that reflects visible light from the light toward a visible light camera as a visible light image and transmits wavelengths above a visible light range toward a second dichroic mirror;the second dichroic mirror reflecting a diagnostic image having a first wavelength toward a camera response to the first wavelength, and the second dichroic mirror transmitting a second diagnostic image having a second wavelength toward a camera responsive to the second wavelength; anda display that renders two or more of the visible light image, the diagnostic image, and the second diagnostic image superimposed on one another. 2. The system of claim 1 wherein the diagnostic image is a blood flow image. 3. The system of claim 1 wherein the diagnostic image includes a tumor. 4. The system of claim 1 wherein the first wavelength is a near-infrared wavelength. 5. The system of claim 1 wherein the first wavelength is substantially 700 nm and the second wavelength is substantially 800 nm. 6. The system of claim 1 wherein the light source includes a plurality of high-power light emitting diodes. | Beth Israel Deaconess Medical Center | A61B001/04 A61B001/00 A61B001/06 A61B005/00 B82Y005/00 B82Y010/00 A61B001/313 A61B005/026 A61B019/00 A61K049/00
US10467382 | B2 | 2019-11-05 | Conceivable basal body temperatures and menstrual cycle | 2014-11-14 | The present invention discloses software programs, systems and methods for increasing the chances of users conception by maximizing the user's fertility potential by providing a software program that combines identifying the underlying causes of infertility and using best practices as described by current medical literature for fertility changes along with herbal recommendations to provide personalized wellness recommendations that give the individual user the power to improve their natural fertility. | 1. A method of increasing chances of conception, comprising: providing a portable electronic device having a processor and a memory that executes a mobile application that is configured to:receive, by the portable electronic device, data about a fertility state of a user;assess personalized fertility state of the user based on the received data of the fertility state of the user;wherein assessing the user's personalized fertility state includes collecting data on menstrual parameters, lifestyle factors, and Basal body temperatures (“BBT's”);assigning the user, one or more hurdles of one or more symptoms that represents underlying fertility issues of the user;determine a deviation between the user's personalized fertility state and an ideal fertility state;algorithmically identify the user's underlying fertility issues based on the determined deviation from the ideal fertility state;store, in the mobile application, a plurality of herbal formulas targeted at moving the user's personalized fertility state toward the ideal fertility state;compare, by the mobile application, the underlying fertility issues with the stored herbal formulas;select, by the mobile application, based on the underlying fertility issues and a phase of a menstrual cycle of the user, a personalized herbal formula recommendations for the user to take to move the user's personalized fertility state toward the ideal fertility state;wherein after the one of more hurdles are assigned, the personalized herbal formula recommendations for the user are recommended based on the underlying fertility issues and phases of the menstrual cycle of the user; anddisplay the personalized herbal formula recommendations for the user. 2. The method of claim 1, wherein said collecting data comprises collecting data on high, low, or erratic BBT's across the menstrual cycle of the user, menstrual clotting, menstrual pain, insufficient menstrual flow, hemorrhagic menstrual flow, insufficient cervical discharge, irregular ovulation, symptoms associated with premenstrual syndrome (PMS) including one or more of bloating, irritability, headaches, breast tenderness, and cramping, and spotting. 3. The method of claim 1, wherein said collecting data further comprises identifying subclinical factors and their causes that affect the user's personalized fertility state and the user's ability to conceive. 4. The method of claim 3, wherein the one or more hurdles and their main associated symptoms are selected from the group consisting of: a) hot—high follicular phase basal body temperatures, scanty cervical discharge, insomnia (harder to stay asleep);b) cold—low luteal phase basal body temperatures, cold hands and feet, frequent urination;c) stuck—premenstrual syndrome (PMS), erratic temperatures across the menstrual cycle of the user, irritability;d) pale—scanty menstrual blood, anxiety, pale skin; ande) tired—fatigue, gas and bloating, loose stools, weight gain or retention. 5. The method of claim 4, wherei... | Brazen Incorporated | G06F019/00 A61B005/00 G09B019/00 A61B005/01 A61B010/00
US10278646 | B2 | 2019-05-07 | Measuring myocardial physiological parameters | 2015-10-29 | A method for measuring a myocardial physiologic parameter according to an embodiment includes placing an at least partially convex portion of a spectral sensor against an intercostal space of a human over a heart of the human and measuring the physiologic parameter of a myocardium of the heart with the spectral sensor over time during an emergency medical event. The spectral sensor may be configured to determine and visually display a suggested position adjustment for directing the spectral radiation more directly toward the tissue of interest (e.g. the myocardium), and/or for placing the operative elements of the spectral sensor closer to the tissue of interest (e.g. the myocardium). | 1. A system for determining an effectiveness of non-invasively measuring a physiologic status of a myocardium of a patient, the system comprising: a probe having a housing, wherein the housing is shaped to conform to a general shape of an indentation of an intercostal space of the patient;an optical spectroscope, at least partially disposed within the housing, the optical spectroscope comprising at least one light source capable of emitting light at a range of wavelengths, a wavelength-sensitive sensor capable of detecting light intensity, at two or more distinct wavelengths, of light scattered and/or reflected by tissue of the myocardium; anda processor communicably coupled to a memory, the memory including instructions that, when executed by the processor, cause the processor to (1) receive light intensity measurements from the optical spectroscope, (2) determine a first magnitude of effectiveness of determining the physiologic status of the myocardium of the patient, (3) calculate positional information associated with the first magnitude of effectiveness, (4) determine a second magnitude of effectiveness of determining the physiologic status of the myocardium of the patient, (5) calculate positional information associated with the second magnitude of effectiveness, (6) compare the first magnitude of effectiveness with the second magnitude of effectiveness, and (7) direct the user, based on the comparison, to place the probe in an optimal position using directional commands to the user. 2. The system of claim 1, wherein the housing has a generally convex outer profile to facilitate maintaining substantially continuous surface contact between the housing and the intercostal space when the housing is rotated. 3. The system of claim 1, wherein the first magnitude of effectiveness is based on an effectiveness of the at least one light source in illuminating the tissue of the myocardium. 4. The system of claim 1, wherein the memory further includes instructions that, when executed by the processor, cause the processor to cause a first indication to be communicated to the user, wherein the first indication indicates the first magnitude of effectiveness. 5. The system of claim 4, wherein the first indication is communicated using at least one of a visual display, an audio tone, a verbal communication, and a haptic vibration of the probe. 6. The system of claim 1, wherein the memory further includes instructions that, when executed by the processor, cause the processor to cause a second indication to be provided to the user, wherein the second indication indicates which of the first and second magnitudes of effectiveness is larger. 7. The system of claim 6, wherein the second indication is communicated using at least one of a visual display, an audio tone, a verbal communication, and a haptic vibration of the probe. 8. The system of claim 1, wherein the probe contains at least one inertial sensor. 9. The system of claim 8, wherein the inertial... | ZOLL MEDICAL CORPORATION | A61B005/00 A61N001/04 A61B005/0402 A61B005/02 A61B005/145 A61N001/39
US09345404 | B2 | 2016-05-24 | Mobile device that monitors an individuals activities, behaviors, habits or health parameters | 2013-08-06 | A mobile device is provided that obtains monitored information about an individual. A display is coupled to a housing. A memory is coupled to a memory and one or more processing units (CPU's), a peripherals interface, a Network Systems circuitry and an input/output (I/O) subsystems. A system is included with one or more sensors that detects or measure an individual's activities. | 1. A monitoring system, comprising: a monitoring device that includes ID circuitry with an ID storage that contains a unique individual ID, a communication system which reads and transmits the unique individual ID from the ID storage, a power source, a pathway system, a microphone, an RF transmitter and sensors to determine air quality, sound level, sound quality, light quality and ambient temperature near an individual, the RF transmitter serving as the communication system;an accelerometer or a movement detection device configured to detect movement of the individual, the accelerator and the monitoring device configured to calculate or derive sleep onset and wake time, sleep interruptions, and quality and depth of sleep from the movement of the individual, the microphone configured to record an individual's movement sounds detected by the accelerometer or the movement detection device, the accelerometer or movement detection device configured to cause the microphone to stop recording the individual's movement sounds when the movement sounds are not directed to a sleep related parameter;a telemetry system, the telemetry system in operation receiving sleep information and sleep related parameter information data of one or more individual's, the telemetry system or the monitoring system at least partially analyzing the one or more individual's sleep information and sleep related parameter information data to calculate or derive the sleep onset and the wake time, the sleep interruptions, and the quality and depth of the sleep, the telemetry system having a database that includes base individual information related to the sleep information and the sleep related parameter information data and information regarding a degree to which the one or more individual's has followed a routine; andan alert, wherein the alert is activated only when the one or more individual's is awake. 2. The device of claim 1, further comprising: one or more analysis tools at the telemetry system. 3. The device of claim 2, further comprising: a feedback system or subsystem at the telemetry system that provides feedback to the individual. 4. The device of either claim 3, wherein the feedback is in graphical form. 5. The device of claim 3, wherein the feedback system or subsystem in operation generates one or more web pages containing the feedback. 6. The device of claim 2, wherein the one or more analysis tools in operation compares the one or more individual's sleep information and sleep related parameter information data to the base individual information in the database. 7. The device of claim 6, wherein the base individual information includes threshold values. 8. The device of claim 7, the one or more analysis tools in operation compares the one or more individual's sleep information and sleep related parameter information data from the telemetry system with one or more sensors to the threshold values. 9. The device of claim 8, wherein an alert at the monitoring de... | Hello Inc. | G08B023/00 A61B005/00 H02J007/02 H02J017/00 A61B005/11 H04W004/00 H02J007/00 H04L029/08 H04W004/02
US09330678 | B2 | 2016-05-03 | Voice control device, voice control method, and portable terminal device | 2013-06-21 | A voice control device includes a calculation section configured to calculate a response time representing a time difference between a voice in a received signal and a voice in a sending signal; a hearing estimate section configured to estimate hearing of a user based on the calculated response time; and a voice control section configured to control the received signal by a compensation quantity responsive to the estimated hearing. | 1. A voice control device inclusive of a semiconductor integrated circuit, comprising: a memory;a processor coupled to the memory and configured to: calculate a time length of a voice in a received signal;calculate a difference between a high range power and a low range power of the voice in the received signal as an articulation rate;calculate a response time representing a time difference between the voice in the received signal and a voice in a sending signal when the calculated time length of the voice in the received signal is longer than a first threshold value while the calculated articulation rate is greater than a third threshold;estimate hearing of a user based on the calculated response time;control the received signal by a compensation quantity depending on the estimated hearing; andproduce sound responsive to the controlled received signal. 2. The voice control device as claimed in claim 1, wherein the processor is configured to: calculate a time length of the voice in the sending signal, andcalculate the response time when the calculated time length of the voice in the sending signal is longer than a second threshold value. 3. The voice control device as claimed in claim 1, wherein the processor adjusts the compensation quantity depending on a time change in the response time to control the received signal using the adjusted compensation quantity. 4. The voice control device as claimed in claim 3, wherein the processor executes control to emphasize the received signal when the response time after the control is shorter than the response time before the control, or executes control to decrease the received signal when the response time after the control is longer than the response time before the control. 5. The voice control device as claimed in claim 1, wherein the processor does not calculate the response time when a starting time of the voice in the sending signal precedes an ending time of the voice in the received signal. 6. The voice control device as claimed in claim 1, wherein the processor estimates the hearing when the response time is less than a fourth threshold value. 7. The voice control device as claimed in claim 1, wherein the processor calculates an average response time by averaging a plurality of the calculated response times, estimates the hearing based on the average response time. 8. A method of controlling voice in a voice control device inclusive of a semiconductor integrated circuit, comprising: calculate a time length of a voice in a received signal;calculate a difference between a high range power and a low range power of the voice in the received signal as an articulation rate;calculating a response time representing a time difference between the voice in the received signal and a voice in a sending signal when the calculated time length of the voice in the received signal is longer than a first threshold value while the calculated articulation rate is greater than a third threshold;estimating a he... | FUJITSU LIMITED | G10L021/04 H04M001/60 A61B005/00 A61B005/12
US20160148535 | A1 | 2016-05-26 | Tracking Nutritional Information about Consumed Food | 2015-11-18 | A system for tracking nutritional information about consumed food includes a processor and memory. The memory includes programmed instructions executable by the processor to receive a first input from a first sensor indicating a swallowing and/or chewing activity, receive a second input from a second sensor indicating a physiological response of the user, and generate a nutritional value consumed at least in part based on the first input and the second input. | 1. A system for tracking nutritional information about consumed food, comprising: a processor and memory, the memory comprising programmed instructions executable by the processor to: receive a first input from a first sensor indicating a swallowing and/or chewing activity;receive a second input from a second sensor indicating a physiological response of a user; andgenerate a nutritional value consumed at least in part based on the first input and the second input. 2. The system of claim 1, wherein the first sensor comprises a microphone capable of recording sounds representative of swallowing and/or chewing activity. 3. The system of claim 1, wherein the first sensor is incorporated into eye glasses. 4. The system of claim 1, wherein the first sensor comprises an accelerometer. 5. The system of claim 1, wherein the first sensor is attachable to a neck of the user. 6. The system of claim 1, wherein the programmed instructions are further executable to determine an amount of food consumed based on the first input. 7. The system of claim 1, wherein the programmed instructions are further executable to determine a type of food based on the second input. 8. The system of claim 1, wherein the second sensor is incorporated into an implant. 9. The system of claim 1, wherein the second sensor comprises a non-invasive mechanism to measure a physiological characteristic indicative of the physiological response. 10. The system of claim 1, wherein the swallowing and/or chewing activity includes a number of swallows. 11. The system of claim 1, wherein the swallowing and/or chewing activity includes a chewing duration. 12. The system of claim 1, wherein the programmed instructions are further executable to send a message when a nutritional goal is exceeded. 13. The system of claim 1, wherein the programmed instructions are further executable to send the nutritional information to a database. 14. The system of claim 1, wherein the first sensor and/or the second sensor are incorporated into a wearable computing device. 15. The system of claim 1, wherein the physiological response is a glycemic response. 16. A system for tracking nutritional information about consumed food, comprising: a processor and memory, the memory comprising programmed instructions executable by the processor to: receive a first input from a first sensor capable of recording swallowing and/or chewing activity;receive a second input from a second sensor indicating a glycemic response;determine an amount of food consumed based on the first input;determine a type of food based on the second input; andgenerate a number of calories consumed at least in part based on an amount and the type of food. 17. The system of claim 16, wherein the programmed instructions are further executable to send the number of calories to a database. 18. The system of claim 16, wherein the first sensor and/or the second sensor are incorporated into a wearable computing device. 19. The system of claim... |  | G09B019/00 A61B005/145 A61B007/02 A61B007/00 G09B005/02 A61B005/00
US20160128624 | A1 | 2016-05-12 | THREE DIMENSIONAL IMAGING OF THE MOTION OF TEETH AND JAWS | 2015-11-05 | Pre-calibrated position sensors, and a computer or cloud based service are used to show motion in real time of the teeth and jaws depicting six degrees of freedom. The computer receives anatomically correct 3D images of any portion of the craniomandibular system with a position sensor(s) attached. The position sensor data is then transferred with wireless attachments to the patient to generate true to life motion of the digital 3D models of the teeth and jaws. The 3D models are then employed to provide various techniques for treating bite disorders. | 1. A method for visualizing jaw function comprising: affixing a plurality of wireless position sensors in proximity to two or more pairs of opposing teeth of a patient's teeth/jaw;affixing an sensor board in proximity to the plurality of wireless position sensors;generating a three dimensional (3D) model the of kinesiologic activity of the craniomandibular complex based upon a sensing of the plurality of wireless position sensors by the sensor board while an upper and lower jaw of the teeth/jaw are in motion with respect to each other; 2. The method of claim 1, further comprising calibrating the plurality of sensors with respect to the sensor board by selecting corresponding data points while the upper and lower jaw are in a known fixed position. 3. The method of claim 1, wherein the wireless position sensors are affixed in proximity to the two or more pairs of opposing teeth while affixed to a delivery jig, further comprising calibrating the plurality of sensors with respect to the sensor board based upon a known position of the wireless position sensors with respect to the delivery jig. 4. The method of claim 1, wherein the sensor board comprises: a magnetometer; andan accelerometer. 5. The method of claim 1, wherein the plurality of wireless position sensors are magnets. 6. The method of claim 1, further comprising generating a occlusal appliance based upon the 3D model. 7. The method of claim 1, further comprising: collecting physiologic data corresponding to the patient; andgenerating a treatment plan for the patient based upon the 3D model and the physiologic data. 8. The method of claim 7, wherein the physiologic data comprises a selection of measurements of a list, the list comprising: a pulse oximetry;a heart rate;an EKG;an EEG;a blood pressure; anda skeletal muscle EMG. 9. The method of claim 1, further comprising using the 3D model for determining a craniomandibular relationship for a treatment bite position, wherein the treatment bite position can be modified or maintained throughout treatment with inter-occlusal appliances as determined by and correlated to one or more physiologic measurements. 10. The method of claim 1, further comprising: calculating a space between the jaws and teeth corresponding to a desired positioning of the jaws and teeth and based upon the 3D model; andgenerating a mouth piece in conformity with the space to hold the jaws and teeth in the desired position. |  | A61B005/00 A61C007/00 A61B005/0476 A61B005/0488 A61B005/0205 A61B005/0402
US20160135702 | A1 | 2016-05-19 | Techniques for Predicting Cardiac Arrhythmias Based on Signals from Leads of Electrocardiography | 2014-06-20 | Techniques for predicting cardiac arrhythmia includes obtaining first data that indicates an electrocardiography recording from a patient; and, automatically deriving, on a processor, P-wave characteristics on a plurality of leads of the electrocardiography recording. A value for a first parameter, Pindex3, is determined based on a standard deviation of P-wave duraations automatically derived from only three leads of the plurality of leads. A risk of incidence of cardiac arrhythmia for the patient is determined based, at least in part, on the first parameter, Pindex3. | 1. A method comprising: obtaining first data that indicates an electrocardiography recording from a patient;automatically deriving, on a processor, P-wave characteristics on a plurality of leads of the electrocardiography recording;determining a value for a first parameter, called Pindex3, based on a standard deviation of P-wave durations automatically derived from only three leads of the plurality of leads;determining a risk of incidence of cardiac arrhythmia for the patient based, at least in part, on the first parameter, Pindex3; andpresenting on a display device output data that indicates the risk of incidence. 2. A method as recited in claim 1, wherein the only three leads are orthogonal to each other. 3. A method as recited in claim 1, wherein the only three leads consist of lead I, lead aVF and lead V1 of the electrocardiography recording from the patient. 4. A method as recited in claim 1, further comprising treating the patient based on the determined risk of incidence of cardiac arrhythmia. 5. A method as recited in claim 4, wherein: determining the risk of incidence of cardiac arrhythmia for the patient further comprises determining that the patient is at elevated risk for atrial fibrillation if the value for the first parameter exceeds 45; andtreating the patient further comprises treating the patient to reduce risk of atrial fibrillation if the patient is determined to be at elevated risk for atrial fibrillation. 6. A method as recited in claim 1, wherein: the method further comprises determining a value for a different second parameter based on P-wave characteristics automatically derived from at least one of the plurality of leads; anddetermining the risk of incidence of cardiac arrhythmia for the patient further comprises determining the risk of incidence of cardiac arrhythmia for the patient based at least in part on the second parameter. 7. A method as recited in claim 6, wherein the different second parameter is selected from a group comprising: abnlPaxis that indicates a first value for an abnormal P-wave axis and a different second value for a normal P-wave axis;anyPAC that indicates a third value in the presence of a premature atrial contraction (PAC) within a randomly chosen 10-second period and a different fourth value in the absence of the premature atrial contraction (PAC);PRdurlong that indicate a PR interval greater than 200 miliseconds;Plongthree that indicates a maximum P-wave duration of greater than 120 ms;PVCs that indicates a fifth value for a premature ventricular contraction within a randomly chosen 10-second period and a different sixth value for the absence of such a ventricular premature complex; andLBBB that indicates a seventh value in the presence of a left bundle branch block and a different eighth value in the absence of a left bundle branch block. 8. A method as recited in claim 7, wherein: the first value is about 1, the second value is about 0, the third value is about 1, the fourth value is a... | The Board of Trustees of the Leland Stanford Junior University | A61B005/046 A61B005/00 A61B005/044 A61B005/0408 A61B005/04
US09326697 | B2 | 2016-05-03 | Long-term monitoring for discrimination of different heart rhythms | 2014-06-25 | A method, system, and device for detection of an arrhythmia, and discrimination between different types of arrhythmia, for example to determine whether to administer an electric shock to the heart, the device comprising a wearable monitor with electrodes that detect the electrical activity of a beating heart, attached to an embedded monitoring system having an amplifier, a microprocessor, a data storage device, and a power supply, all disposed on a substrate having large distal end portions that attach to the electrodes and a narrow intermediate portion that attaches to the monitoring system. | 1. A method for calculating an arrival time for a QRS complex, comprising: obtaining with a wearable or an implantable arrhythmia monitoring and detection device a time sequence comprising equally spaced amplitude signals representing the electrical activity of a heart;selecting a time window defined to include a fixed number of the equally spaced amplitude signals such that the time window has a window width;obtaining from the time sequence a first segment comprising the fixed number of equally spaced amplitude signals, and calculating for the first segment the sum of the absolute value of the differences between each adjacent pair of amplitude signals in the first segment;obtaining from the time sequence subsequent overlapping segments of amplitude signals, and calculating for the subsequent overlapping segments the sum of the absolute value of the differences between each adjacent pair of amplitude signals in the segment, wherein the step of obtaining subsequent overlapping segments continues until the sum exceeds a predetermined threshold value;when the sum exceeds the predetermined threshold value, setting a flag and recording the corresponding calculated sum in a maximum sum parameter;obtaining from the time sequence subsequent overlapping segments, and calculating for the subsequent overlapping segments the sum of the absolute value of the differences between each pair of amplitude signals in the time window, and if the calculated sum is greater than the maximum sum parameter, resetting the maximum sum parameter to the calculated sum and recording an identifier parameter that identities the segment, wherein the step of obtaining segments continues until the number of subsequent overlapping segments that have been obtained without resetting the maximum sum parameter exceeds a predetermined number; andcalculating the arrival time of the QRS complex as the center of the time window corresponding to the segment identified by the identifier parameter. 2. The method of claim 1, wherein the identifier parameter comprises the end time of the segment corresponding to the identifier parameter, and the center of the time window is calculated by subtracting half of the window size from the identifier parameter. 3. The method of claim 2, wherein the arrival time of the QRS complex is used to calculate a heartbeat RR interval. 4. The method of claim 3, wherein the arrival time of the QRS complex is used to calculate an instantaneous heart rate. 5. The method of claim 2, further comprising resetting the flag and the maximum sum, and repeating the method to calculate a next arrival time of a QRS complex from the time sequence. 6. The method of claim 2, wherein the time sequence comprising equally spaced amplitude signals comprises a digitized continuously recorded electrocardiogram. 7. The method of claim 6, wherein the electrocardiogram is a single channel electrocardiogram recording at least 100 samples per second. 8. The method of claim 1, where... | University of Washington | A61B005/0452 A61B005/046 A61B005/024 A61B005/0432 A61B005/0408 A61B005/0468 A61B005/0472 A61B005/00 A61N001/362 A61N001/37
US09326727 | B2 | 2016-05-03 | On-body medical device securement | 2014-05-15 | Devices and methods for maintaining a medical device on-body are provided. Embodiments include medical device securement systems having first and second on-body securement elements. Also provided are systems and kits for use maintaining a medical device on-body. | 1. A device, comprising: a non-adhesive backing having a first surface and a second surface; andan adhesive disposed on at least a portion of the first surface of the non-adhesive backing forming an adhesive portion and a non-adhesive portion on the first surface of the non-adhesive backing, the non-adhesive portion positionable over a medical device and sized and configured so that the medical device is not contacted with the adhesive portion when the non-adhesive portion is positioned over the medical device, the non-adhesive portion including an opening, and a transparent shield covering the opening. 2. The device of claim 1, further including a removable release liner adjacent the adhesive. 3. The device of claim 1, further including a removable casting layer adjacent the second surface of the non-adhesive backing. 4. The device of claim 1, wherein the adhesive portion is opaque. 5. The device of claim 1, wherein the non-adhesive portion comprises a pouch. 6. The device of claim 1, wherein the adhesive portion is configured for adherence to an adhesive patch on which the medical device is positioned. 7. The device of claim 1, wherein the medical device includes one of sensor electronics, an RF receiver unit, a data monitoring device, an infusion device, or a remote control unit. 8. A system, comprising: sensor electronics having a housing and a sensor extending from the housing, the housing including a surface covered with an adhesive to adhere to a skin surface; anda device configured to be positionable over at least a portion of the housing, wherein the device comprises: a non-adhesive backing having a first surface and a second surface; andan adhesive disposed on at least a portion of the first surface of the non-adhesive backing forming an adhesive portion and a non-adhesive portion on the first surface of the non-adhesive backing, the non-adhesive portion positionable over a medical device and sized and configured so that the medical device is not contacted with the adhesive portion when the non-adhesive portion is positioned over the medical device, the non-adhesive portion including an opening, and a transparent shield covering the opening. 9. The system of claim 8, wherein the adhesive and the housing are water resistant or waterproof. 10. The system of claim 8, further including a removable release liner adjacent the adhesive. 11. The system of claim 10, wherein the removable release liner comprises a plurality of removable release liners adjacent the adhesive covered surface. 12. The system of claim 8, further including a removable casting layer adjacent the second surface of the non-adhesive backing. 13. The system of claim 8, wherein the transparent shield is configured to be positioned over the sensor electronics. 14. The system of claim 8, wherein the transparent shield is compressible. | ABBOTT DIABETES CARE INC. | A61B005/00 A61B005/145 A61B005/1486
US10383570 | B2 | 2019-08-20 | Methods for detection of respiratory effort and sleep apnea monitoring devices | 2015-07-31 | A sleep apnea diagnostic system includes a housing that is configured to be attached to near the nose of a patient's face to sense physiological information of a patient. The housing includes sensors to sense the physiological information. The physiological information may be, for example, air flow through the nose or the mouth or both. The physiological information further may be, for example, blood volume. The sleep apnea diagnostic system includes at least one processor in the housing or external to the housing or both to analyze the physiological information to determine whether the patient has experienced irregular or abnormal respiratory activity and to detect respiratory effort. The analysis may be real time or delayed. | 1. A sleep apnea monitoring device comprising: a housing shaped to fit the contours of a patient's face;one or more photosensors positioned on a nasal area of the patient and coupled to the housing, the one or more photosensors configured to generate output signals related to peripheral photoplethysmography data of the patient;one or more air-flow sensors positioned near or inside one or more nostrils of the patient and coupled to the housing, the one or more air-flow sensors configured to generate output signals conveying airflow information related to an airflow of the patient;a memory disposed in the housing and configured to store physiological information related to the peripheral photoplethysmography data and store the airflow information; andone or more processors disposed in the housing and configured to: receive, from the one or more photosensors, the peripheral photoplethysmography data;receive, from the one or more air-flow sensors, the airflow information;identify an apneic event based on the airflow information being indicative of an absence of airflow;obtain, via filtering of the peripheral photoplethysmography data, a DC component of the peripheral photoplethysmography data;determine, based on the peripheral photoplethysmography data, a change in vascular tone from a time prior to the apneic event to a time during the apneic event;determine whether the DC component indicates respiratory effort from the patient; andresponsive to the DC component indicating no respiratory effort from the patient, determine a diagnosis of the apneic event based on an adjusted component, the adjusted component being based on the DC component and the change in vascular tone. 2. The sleep apnea device of claim 1, wherein the one or more processors are configured to: determine an effect of vasoconstriction on the DC component based on the change in vascular tone;subtract the effect of vasoconstriction from the DC component to obtain an adjusted component;determine a respiratory effort of the patient based on the adjusted component; anduse the respiratory effort determined based on the adjusted component to determine the diagnosis of the apneic event. 3. The sleep apnea device of claim 1, wherein the one or more processors are configured to: responsive to the DC component indicating no respiratory effort from the patient, perform the following operations: subtracting an effect of vasoconstriction from the DC component to obtain an adjusted component, the effect of vasoconstriction being (i) an effect on the DC component and (ii) based on the change in vascular tone;determining a respiratory effort of the patient based on the adjusted component; andusing the respiratory effort determined based on the adjusted component to determine the diagnosis of the apneic event. 4. The sleep apnea device of claim 1, wherein the housing is shaped to fit the skin surface of the patient's nose, upper lip, and the area of maxillary sinuses. 5. The sleep apnea device of ... | KONINKLIJKE PHILIPS N.V. | A61B005/00 A61B005/0205 A61B005/1455 A61B005/02 A61B005/0295
US09872638 | B2 | 2018-01-23 | Method and apparatus for determining information indicative of cardiac malfunctions | 2014-06-13 | An apparatus for determining information indicative of cardiac malfunctions includes a processing device (602) configured to detect whether the time-trend of a signal indicative of cardiovascular motion manifests an indicative phenomenon where the amplitude of a first heartbeat is greater than the amplitude of a second heartbeat and the temporary heartbeat rate is greater prior to the first heartbeat than prior to the second heartbeat. The processing device is further configured to produce an indicator of cardiac malfunction, e.g. an indicator of atrial fibrillation, in response to a situation in which the indicative phenomenon is detected to be present. | 1. An apparatus comprising: a signal interface configured to receive a first signal indicative of cardiovascular motion; anda processing device coupled to the signal interface, the processing device being configured to: determine, based on the first signal, a correlation between amplitudes of heartbeats and quantities indicative of heartbeat rates prior to the heartbeats to detect whether a time-trend of the first signal manifests an indicative phenomenon where an amplitude of a first heartbeat is greater than an amplitude of a second heartbeat and a heartbeat rate is greater prior to the first heartbeat than prior to the second heartbeat,produce an indicator of cardiac malfunction in response to detection of the indicative phenomenon, andoutput the indicator to a user to treat or diagnose a patient. 2. The apparatus according to claim 1, wherein the processing device is configured to use a temporal length of a heartbeat period preceding the first heartbeat as a quantity indicative of the heartbeat rate prior to the first heartbeat, and a temporal length of a heartbeat period preceding the second heartbeat as a quantity indicative of the heartbeat rate prior to the second heartbeat. 3. The apparatus according to claim 1, wherein the processing device is configured to determine the correlation between the amplitudes of the heartbeats and temporal lengths of the heartbeat periods preceding the heartbeats, a negative correlation indicating presence of the indicative phenomenon. 4. The apparatus according to claim 1, wherein the processing device is configured to determine the correlation between the amplitudes of the heartbeats and heartbeat rates on heartbeat periods preceding the heartbeats, a positive correlation indicating presence of the indicative phenomenon. 5. The apparatus according to claim 1, wherein the processing device is configured to compute a correlation coefficient C indicative of the degree of the correlation between the amplitudes of the heartbeats and the quantities indicative of the heartbeat rates prior to the heartbeats: C=E{(A−μA)×(Q−μQ)},where E is the expected value operator, A represents the amplitudes of the heartbeats, mA is the mean of the amplitudes, Q represents the quantities indicative of the heartbeat rates, and mQ is the mean of the quantities indicative of the heartbeat rates. 6. The apparatus according to claim 1, wherein the processing device is configured to extract or produce, from or on the basis of the first signal, one of the following which represent the amplitudes of the heartbeats: heights of AO-peaks caused by opening of the aortic valve, maximum peak-to-peak values within time windows each involving one of the AO-peaks, root-mean-square values computed over the time windows each involving one of the AO-peaks. 7. The apparatus according to claim 6, wherein the processing device is configured to extract, from the first signal, time periods between successive AO-peaks, temporal lengths of the time... | PRECORDIOR OY | A61B005/04 A61B005/11 A61B005/046 A61B005/0452 A61B005/0456 A61B005/00
US09344688 | B2 | 2016-05-17 | Recall device | 2014-10-15 | A small wearable recall device is provided to capture images triggered by a combination of a detection of a capture condition (e.g., changes in motion, temperature or light level) followed by a relatively stable period, as detected by an accelerometer. By triggering on the combination of a detected capture condition followed by a detected stability condition, a clearer image of the environment of an interesting event is expected to be captured. The small size of the recall device makes it possible to integrate it into common portable consumer products, such as MP3 players, purses, clothing, hats, backpacks, necklaces, collars, and other human-wearable products. | 1. A portable recall device configured to be carried by a wearer comprising: a camera configured to capture an image;at least one accelerometer configured to detect acceleration of the camera;a memory configured to store images; anda controller operably connected to the camera, wherein the controller is configured to: detect a stable condition based on the at least one accelerometer;trigger the camera to capture the image based on the detected stable condition, wherein the image is stored in the memory;receive a sensor signal and, based on the sensor signal, detect occurrence of a capture condition and whether the capture condition occurred in temporal proximity of the capture of the image by the camera; andmaintain storage of the image in the memory if the capture condition occurred in temporal proximity of the capture of the image by the camera, and to delete the image from the memory if no capture condition is detect in temporal proximity of the capture. 2. The portable recall device of claim 1 wherein the image is one image of a video sequence. 3. The portable recall device of claim 1 further comprising: an audio recording circuit recording ambient sounds, wherein the memory further maintains storage of the recorded ambient sounds if the capture condition is detected in temporal proximity of the capture; the memory deleting the storage of the recorded ambient sounds if no capture condition is detected in the temporal proximity of the capture. 4. The portable recall device of claim 1, wherein the capture condition is detected based a sensor signal. 5. The portable recall device of claim 4, and further comprising one or more sensors configured to generate the sensor signal, wherein the controller is configured to receive a sensor signal from the one or more sensors and to detect the capture condition based on the sensor signal. 6. The portable recall device of claim 1, wherein the controller is configured to detect the stable condition based on a determination that an acceleration of the camera is below a defined threshold. 7. The portable recall device of claim 1, wherein the defined temporal proximity comprises a time period after the capture of the image. 8. The portable recall device of claim 1, wherein image capture by the camera is triggered such that the camera captures images continuously during the stable condition. 9. The portable recall device of claim 1, wherein image capture by the camera is triggered such that the camera captures images periodically during the stable condition. 10. The portable recall device of claim 1, wherein, if a capture condition is detected, the image is stored in memory in association with the capture condition. 11. The portable recall device of claim 1, and further comprising: an environmental sensor that monitors ambient conditions to detect the capture condition. 12. The portable cal device of a 11, wherein the capture condition comprises at least one of: a change in ambient light;a change in ... | Microsoft Technology Licensing, LLC | H04N005/225 H04N007/18 H04N005/232 A61B005/11 G01S019/34 G08B013/196 G08B015/00 H04L029/08 H04L029/06 H04L012/54 H04L029/12 G06F011/20 G06F021/62 G07F017/16 G06Q020/10 G06Q020/12 G06Q020/38 G06Q020/40 G06Q020/42 G06Q030/02 H01M016/00 H04B001/04 H04B001/40 H04B007/06 H04B007/08 H04J013/12 H04L001/00 H04L001/06 H04L001/08 H04L001/18 H04L005/00 H04L009/32 H04L012/12 H04L012/28 H04L012/40 H04L012/44 H04L012/64 H04L027/26 H04M001/02 H04M001/253 H04M003/22 H04M003/42 H04M007/00 H04N001/00 H04N007/01 H04N007/08 H04N007/14 H04N021/61 H04N021/647 H04W024/00 H04W028/14 H04W052/14 H04W052/24 H04W052/46 H04W052/48 H04W074/02 H04W076/02 G06Q030/06 H04L012/24 H04L012/801 H04L012/815 G06Q050/18 H04L012/863 H04L012/851 H04L012/841 H04N005/77 G06F011/16 H04B007/155 H04N001/32 H04W008/00 H04W008/04 H04W028/04 H04W048/16 H04W052/22 H04W072/08 H04W080/10 H04W084/12 H04W088/16
US09939501 | B2 | 2018-04-10 | Iterative minimization procedure with uncompressed local SAR estimate | 2014-11-12 | A method for providing an image of a subject in a magnetic resonance imaging MRI system with parallel transmission (pTx) is provided. A localizer scan of the subject is provided. Body anatomy is determined from the localizer scan. The body anatomy is matched with at least one body model of a plurality of body models. B1+ and Bo maps are calculated from the body anatomy. The electric fields from the at least one body model and constraints are used to simultaneously determine a plurality of excitation pulses for a pTx. MRI is performed on the subject using the determined excitation pulses. Global SAR is monitored in real time to ensure all time-averaged constraints are satisfied simultaneously. | 1. A method for providing an image of a subject in a magnetic resonance imaging MRI system with parallel transmission (pTx), comprising: a) providing a localizer scan of the subject;b) determining body anatomy from the localizer scan;c) matching the body anatomy with at least one body model of a plurality of body models;d) calculating B1+ and B0 maps from the body anatomy;e) using electric fields from the at least one body model and constraints to simultaneously determine a plurality of excitation pulses for a pTx performed as an optimization problem with a single objective term consisting of all time-averaged quantities including local SAR, global SAR, and average power and constraints on instantaneous quantities including flip angle inhomogeneity resulting from each pulse and instantaneous power applied to each channel, wherein local SAR matrices at all voxels are included in the objective term without compression;f) performing MRI on the subject using the determined excitation pulses;g) monitoring global SAR in real time to ensure all time-averaged constraints are satisfied simultaneously; andh) providing an image from the performing the MRI on the subject and monitoring the global SAR in real time. 2. The method, as recited in claim 1, wherein more than one body model is used for SAR estimation, performed by concatenating SAR matrices from the voxels across all body models. 3. The method, as recited in claim 1, wherein only the magnitude of excitation is determined and the phase of excitation is not determined. 4. The method, as recited in claim 1, wherein flip angle inhomogeneity is nonconvex with freedom in choice of spokes locations and target phase. 5. The method, as recited in claim 1, wherein the optimization problem is decomposed into two subproblems to determine the pulses iteratively. 6. The method, as recited in claim 5, wherein a first subproblem of the two subproblems consists of an unconstrained minimization of a piecewise positive semidefinite function. 7. The method, as recited in claim 6, wherein a second subproblem of the two subproblems consists of a projection operation onto a nonconvex set of feasible pulses. 8. The method, as recited in claim 6, wherein the first subproblem is solved with Kelley's cutting plane algorithm. 9. The method, as recited in claim 8, wherein a subgradient and value of the time-averaged cost function in each iteration of Kelley's cutting plane algorithm is computed with a vectorized oracle. 10. The method, as recited in claim 9, wherein the vectorized oracle exists on a GPU for additional acceleration. 11. The method, as recited in claim 7, wherein the projection operation is performed in parallel for each RF pulse. 12. The method, as recited in claim 7, wherein the nonconvex projection operation is accomplished through sequential convex projections onto ellipsoids with fixed k-space locations and target phases. 13. The method, as recited in claim 1, wherein the step of calculating B1+... | The Board of Trustees of the Leland Stanford Junior University | A61B005/05 G01R033/28 A61B005/055 G01R033/561 G01R033/24
US09345422 | B2 | 2016-05-24 | Method of locating the tip of a central venous catheter | 2014-10-03 | A method of locating a tip of a central venous catheter (“CVC”) having a distal and proximal pair of electrodes disposed within the superior vena cava, right atrium, and/or right ventricle. The method includes obtaining a distal and proximal electrical signal from the distal and proximal pair of electrodes and using those signals to generate a distal and proximal P wave, respectively. A deflection value is determined for each of the P waves. A ratio of the deflection values is then used to determine a location of the tip of the CVC. Optionally, the CVC may include a reference pair of electrodes disposed within the superior vena cava from which a reference deflection value may be obtained. A ratio of one of the other deflection values to the reference deflection value may be used to determine the location of the tip of the CVC. | 1. A method of locating a tip of a central venous catheter based on a reference P wave detected by a reference pair of electrodes, a first P wave detected by a first pair of electrodes disposed near the tip of the central venous catheter, and a second P wave detected by a second pair of electrodes disposed on the central venous catheter and spaced proximally from the first pair, the method comprising: determining a first deflection value as a function of the first P wave;determining a second deflection value as a function of the second P wave;calculating a reference deflection value as a function of the reference P wave;comparing the first deflection value to the second deflection value;comparing the reference P wave with at least one of the first and second P waves; anddetermining a location of the tip of the central venous catheter based on the comparison of the first and second deflection values and the comparison of the reference P wave with at least one of the first and second P waves. 2. The method according to claim 1, wherein comparing the first deflection value to the second deflection value comprises calculating a first ratio of the first deflection value to the second deflection value, and wherein determining the location of the tip of the central venous catheter based on the comparison of the first and second deflection values and the comparison of the reference P wave with at least one of the first and second P waves comprise comparing the first ratio to a predetermined maximum threshold. 3. The method according to claim 1, wherein comparing the first deflection value to the second deflection value comprises calculating a first ratio of the first deflection value to the second deflection value, and wherein determining the location of the tip of the central venous catheter based on the comparison of the first and second deflection values and the comparison of the reference P wave with at least one of the first and second P waves comprise comparing the first ratio to a predetermined minimum threshold. 4. The method according to claim 1, wherein comparing the reference P wave with at least one of the first and second P waves comprises calculating a second ratio of one of the first and second deflection values to the reference deflection value, and wherein determining the location of the tip of the central venous catheter based on the comparison of the first and second deflection values and the comparison of the reference P wave with at least one of the first and second P waves comprise comparing the second ratio to a predetermined maximum threshold. 5. The method according to claim 1, wherein the first, second, and reference P waves each comprise a series of discrete numerical values, and wherein calculating the first, second and reference deflection values as a function of the first, second, and reference P waves comprise: determining a maximum value within the series of discrete numerical values of the P wave;determining a minimum... | Bard Acess Systems, Inc. | A61B005/06 A61B005/0452 A61B005/042
US09351684 | B1 | 2016-05-31 | Ketone measurement system with user interface for efficient categorization of measurements | 2015-07-23 | Systems and related methods are provided for measuring a ketone in the fluid of a patient or other user. The systems include a portable measurement device for analyzing a plurality of fluid samples and for generating a corresponding plurality of ketone measurement results. The systems also include an electronic or communications device, for example, such as a cell phone or smartphone, configured to communicate with the measurement device. The system further includes a software application. The systems and related methods according to one aspect provide reminders, preferably interactive reminders that facilitate user compliance with a weight management or diabetic ketoacidosis (DKA) monitoring or prevention program. The systems and related methods according to another aspect incorporate ketone tags and/or trigger points to improve accuracy of results and to advance user compliance. | 1. A system for measuring ketone levels in breath of a user, the system comprising: a portable measurement device configured to analyze breath samples of the user and to generate ketone measurements based on the breath samples, the portable measurement device comprising a breath input port fluidly coupled to a ketone sensor, and comprising a wireless transceiver capable of communicating the ketone measurements to a mobile communications device associated with the user; anda mobile application stored in computer storage, the mobile application comprising executable instructions that direct the mobile communications device to at least: receive, via a wireless communications link between the portable measurement device and the mobile communications device, the ketone measurements generated by the portable measurement device;display a user interface that includes a personalized set of tags associated with the user, each tag representing a respective condition that can affect ketone measurements, each tag individually selectable by the user via the user interface, on a measurement-by-measurement basis, to tag individual ketone measurements with information that enables masking effects of the conditions to be considered during subsequent analysis of the ketone measurements;receive a ketone measurement from the portable measurement device via the wireless communications link;receive, via the user interface, and in association with the ketone measurement, a selection by the user of a tag; andrecord the selection of the tag in association with the ketone measurement. 2. The system of claim 1, wherein the breath input port is configured to fluidly connect to a breath sample container. 3. The system of claim 2, wherein the breath sample container is a breath bag. 4. The system of claim 2, wherein the portable measurement device includes a pump that pumps a breath sample from the breath sample container to the ketone sensor. 5. The system of claim 1, wherein the mobile application further includes a second user interface that provides functionality for the user to define the personalized set of tags. 6. The system of claim 5, wherein the second user interface prompts the user to select tags from a master set of tags. 7. The system of claim 6, wherein the master set of tags includes a tag representing a sleep disruption condition. 8. The system of claim 6, wherein the master set of tags includes a tag representing intake of food within a defined time period of taking a ketone measurement. 9. The system of claim 6, wherein the master set of tags includes a tag representing intake of a medication within a defined time period of taking a ketone measurement. 10. The system of claim 1, wherein the executable instructions further direct the mobile communications device to display the ketone measurement in association with the selected tag. 11. The system of claim 1, wherein the executable instructions further direct the mobile communications device to gen... | Invoy Technologies, Inc. | G01N033/48 A61B005/00 A61B005/08
US10512420 | B2 | 2019-12-24 | Systems to monitor body portions for injury after impact | 2014-11-21 | Systems are described for monitoring extremities for injury or damage following a physical impact. A device embodiment includes, but is not limited to, a deformable substrate; a sensor assembly coupled to the deformable substrate, the sensor assembly configured to generate one or more sense signals based on detection of a physical impact to a body portion and based on detection of a physiological parameter; circuitry operably coupled to the sensor assembly and configured to receive the one or more sense signals based on detection of the physical impact and to determine whether the physical impact exceeds a threshold impact value, the circuitry configured to instruct the sensor assembly to detect one or more physiological parameters of the body portion when the physical impact exceeds the threshold impact value; and a reporting device operably coupled to the circuitry. | 1. A device, comprising: a deformable substrate comprising a conformable structure configured to conform to a skin surface of a body portion;a sensor assembly coupled to the deformable substrate, the sensor assembly including a motion sensor configured to detect a movement of the body portion and to generate one or more sense signals responsive to the detection of the movement,a strain sensor configured to detect a deformation of the skin surface of the body portion and to generate one or more sense signals based on detection of a physical impact to the body portion via detection of the deformation of the skin surface by the strain sensor, andan optical sensor configured to detect one or more physiological parameters of the body portion including an optical property of the body portion and to generate one or more sense signals based on detection of the optical property of the body portion, the optical property of the body portion including at least one of a rubor property or a discoloration property;circuitry operably coupled to the sensor assembly, the circuitry including circuitry configured to receive the one or more sense signals responsive to the detection of the movement,circuitry configured to instruct the sensor assembly to begin monitoring for the physical impact to the body portion based upon the received one or more sense signals responsive to the detection of the movement,circuitry configured to receive the one or more sense signals based on detection of a physical impact to the body portion,circuitry configured to determine whether the physical impact exceeds a threshold impact value,circuitry configured to instruct the sensor assembly to detect the physiological parameter of the body portion including the optical property of the body portion when the physical impact exceeds the threshold impact value, andcircuitry configured to compare the optical property of the body portion to reference optical property data indicative of a traumatic injury; anda reporting device operably coupled to the circuitry and configured to generate one or more communication signals based on instruction by the circuitry. 2. The device of claim 1, wherein the circuitry includes a comparison module configured to compare the one or more sense signals based on detection of a physical impact to the body portion to reference data indicative of the threshold impact value. 3. The device of claim 2, wherein the reference data indicative of the threshold impact value correlates to a threshold for a traumatic injury. 4. The device of claim 1, wherein the sensor assembly includes a pressure sensor configured to detect the physical impact to the body portion and to generate one or more sense signals responsive to the detection of the physical impact. 5. The device of claim 4, wherein the circuitry is configured to instruct the sensor assembly to detect one or more physiological parameters of the body portion based upon the one or more sense signals generated by the ... | ELWHA LLC | A61B005/11 A41D013/12 A42B003/04 A61B005/00
US09877658 | B2 | 2018-01-30 | Passive method and system for contact and/or non-contact with or without intervening materials for detection and identification of the incidence, traverse and physiological condition of a living human at any instant | 2015-01-06 | A low input current amplifier has a voltage spectral density curve to operate at 50 Hz or less and is connected to dielectric materials to receive a signal irrespective of ground reference. The amplifier outputs a first output. A multi-stage amplifier includes a stage connected in series with the low input current amplifier to amplify the first signal to distinguish the incidence, traverse and physiological condition of a living human. The resulting signal is then processed by an algorithm and displayed as human specific motion, heart rate and respiratory rate. | 1. A method of detecting a presence or traverse of a living human, comprising: (a) monitoring spatial dynamics of a polarized entity including tissue, organs and a living body of the living human;(b) filtering extraneous ambient electromagnetic background noise;(c) detecting and quantifying amplitude changes in volts per meter per second or current; and(d) outputting voltage or current amplitude signals,wherein steps (a) and (b) are practiced by connecting a current amplifier having a low frequency voltage spectral density curve to dielectric materials which are proximate to or in contact with the living human or in contact with intervening dielectric materials proximate to or in direct contact with the living human. 2. A method according to claim 1, wherein step (c) is practiced by identifying a response to frequency changes in a volts per meter pattern on the intervening dielectric materials or the proximate/contact dielectric materials. 3. A method according to claim 2, wherein step (c) is further practiced by characterizing the frequency changes in the volts per meter pattern and resetting and responding to a succeeding voltage pattern change. 4. A method of detecting a presence or traverse of a living human, comprising: (a) monitoring spatial dynamics of a polarized entity including tissue, organs and a living body of the living human;(b) filtering extraneous ambient electromagnetic background noise;(c) detecting and quantifying amplitude changes in volts per meter per second or current;(d) outputting voltage or current amplitude signals;(e) detecting the presence of the living human based on the voltage amplitude signals output in step (d); and(f) distinguishing the presence of the living human from a presence of a non-human living entity based on the voltage amplitude signals output in step (d). 5. A method according to claim 4, further comprising detecting a traverse direction based on a pattern of the voltage amplitude signals output in step (d). 6. A method according to claim 4, further comprising detecting a speed of travel. 7. A method of detecting a presence or traverse of a living human, comprising: (a) monitoring spatial dynamics of a polarized entity including tissue, organs and a living body of the living human;(b) filtering extraneous ambient electromagnetic background noise;(c) detecting and quantifying amplitude changes in volts per meter per second or current;(d) outputting voltage or current amplitude signals; and(e) categorizing the voltage amplitude signals according to the living human and a non-human living entity. 8. A method according to claim 7, wherein the categorizing step comprises: generating a plurality of discriminators using empirical data that spans the types of data to be detected; andapplying the discriminators as templates or filters to the voltage amplitude signals output in step (d), thereby yielding output corresponding to a degree to which the voltage amplitude signals correlate with each of the ... | DKL INTERNATIONAL, INC. | G08B023/00 A61B005/0205 A61B005/0428 A61B005/00
US09345885 | B2 | 2016-05-24 | Pacemaker for unilateral vocal cord autoparalysis | 2015-05-04 | A method for treating a human subject having unilateral vocal cord paralysis includes providing a stimulating electrode configured to directly contact a recurrent laryngeal nerve or vagus nerve of the subjects, sensing electrical activity of a synkinetically reinnervated dysfunctional muscle of the subject, generating at least one stimulation parameter, using a processor, based on the sensed electrical activity, and stimulating the recurrent laryngeal nerve or vagus nerve, using the stimulating electrode, based on the at least one stimulation parameter. A system is also provided. | 1. A method of treating a human subject having unilateral vocal cord autoparalysis, the method comprising: providing a stimulating electrode configured to directly contact a recurrent laryngeal nerve or vagus nerve of the subject;sensing electrical activity of a synkinetically reinnervated dysfunctional muscle of the subject;generating at least one stimulation parameter, using a processor, based on the sensed electrical activity; andstimulating the recurrent laryngeal nerve or vagus nerve, using the stimulating electrode, based on the at least one stimulation parameter. 2. The method of claim 1, wherein the stimulating electrode is configured to sense the electrical activity of the synkinetically reinnervated dysfunctional muscle. 3. The method of claim 1, wherein the stimulating electrode is configured to be central to an insertion point of the recurrent laryngeal nerve or vagus nerve into innervated muscle. 4. The method of claim 3, wherein central to the insertion point is where the recurrent laryngeal nerve passes an aortic arch. 5. The method of claim 1, wherein the synkinetically reinnervated dysfunctional muscle is a vocalizing muscle that opens or closes a vocal fold. 6. The method of claim 5, wherein the vocalizing muscle is selected from the group consisting of thyroarytenoid muscles, vocalis muscles, posterior cricoarytenoid muscle, lateral cricoarytenoid muscles, and combinations thereof. 7. The method of claim 1, wherein stimulating the recurrent laryngeal nerve or vagus nerve stimulates the abductor muscle. 8. The method of claim 1, wherein stimulating the recurrent laryngeal nerve or vagus nerve stimulates the adductor muscles. 9. The method of claim 1, wherein stimulating the recurrent laryngeal nerve or vagus nerve stimulates the abductor muscle and the adductor muscles. 10. The method of claim 1, wherein sensing includes detecting electromyographic (EMG) activity of a vocalizing muscle, detecting movement related to voice production, or both. 11. The method of claim 1, further comprising: determining when the sensed electrical activity has reached a predetermined level; andgenerating the at least one stimulation parameter when the predetermined level is reached. 12. The method of claim 1, further comprising activating a glottic closure reflex, using the stimulating electrode, based on the at least one stimulation parameter. 13. The method of claim 1, further comprising activating a glottic closure reflex based on the at least one stimulation parameter, wherein activating includes electrical stimulation, mechanical stimulation, or a combination thereof. 14. The method of claim 1, wherein the at least one stimulation parameter includes a stimulation voltage that is above a threshold for activation of vocal cord adductor muscles of the subject and below a threshold for activation of vocal cord abductor muscles of the subject. 15. The method of claim 1, wherein the at least one stimulation parameter includes a stimulat... | MED-EL Elektromedizinische Geraete GmbH | A61N001/36 A61H023/02 A61B005/0492 A61N001/05
US09339252 | B2 | 2016-05-17 | X-ray CT imaging device and X-ray CT imaging method | 2014-03-07 | According to an X-ray imaging device for performing CT imaging, a main body control section controls at least one of an X-ray revolution plane formed by an X-ray cone beam along with the revolution of a revolving arm and a revolution range of the revolving arm in accordance with a CT imaging area accepted by an imaging area setting screen. Such control is performed in order to decrease X-ray radiation to a high sensitivity site in the area that revolves during the CT imaging, the high sensitivity site being positionally set in a biological body as a site that is sensitive to X-rays. | 1. An X-ray CT imaging device, comprising: a revolution section that causes an X-ray generation source and an electric X-ray detector, facing each other while having a subject therebetween, to revolve about a revolution shaft as a center of revolution;a revolution driving section that drives driving the revolution section to revolve with respect to the subject;an X-ray radiation area setting section that accepts a setting of a local area of the subject as an X-ray radiation area;a control section that controls at least the revolution driving section; anda high sensitivity site specification section to specify a high sensitivity site in a biological body that is highly sensitive to an X-ray;wherein the revolution section is revolved with respect to the X-ray radiation area to perform CT imaging; andwherein, when radiation of an X-ray flux in a revolving movement is performed during the CT imaging in accordance with the X-ray radiation area accepted by the X-ray radiation area setting section, the control section controls at least one of an X-ray revolution plane formed by the X-ray flux along with the revolution of the revolution section, and a revolution range of the revolution section, such that an amount of X-ray radiation toward the high sensitivity site is decreased. 2. An X-ray CT imaging device according to claim 1, wherein the high sensitivity site specification section specifies the high sensitivity site in a schematic view of a head of the subject. 3. An X-ray CT imaging device according to claim 1, wherein the control section performs revolution range control of changing at least one of an X-ray radiation start position and an X-ray radiation end position in the revolution range in accordance with the X-ray radiation area set by the X-ray radiation area setting section and the high sensitivity site. 4. An X-ray CT imaging device according to claim 1, wherein the control section performs control of changing at least one of an X-ray radiation angle with respect to a predetermined reference plane and a radiation position on the subject to control the x-ray revolution plane in accordance with the X-ray radiation area set by the X-ray radiation area setting section and the high sensitivity site. 5. An X-ray CT imaging device according to claim 4, further comprising: a support section that supports the revolution section via the revolution shaft; anda support section moving section that moves the support section;wherein the support section moving section includes at least one of a shaft moving mechanism that moves the revolution shaft with respect to the support section and a relative moving mechanism that moves the support section with respect to the subject; andwherein the control section controls the movement of the support section moving section thus to control the X-ray revolution plane. 6. An X-ray CT imaging device according to claim 4, further comprising an X-ray radiation range restriction section that restricts an X-ray radiati... | J. Morita Manufacturing Corporation | A61B005/06 A61B005/08 G01N023/04 A61B006/00 A61B006/03 A61B006/14
US10130127 | B2 | 2018-11-20 | Surgical suit | 2014-06-13 | A body surgical suit to be worn by surgery patients during surgical procedures, said body surgical suit being configured to encase the limbs and torso of a patient is described. The body surgical suit has one or more openings allowing to access one or more portions of the body of a patient wearing the body surgical suit, one or more sensors configured to monitor the patient's vital signs, a control monitor configured to control operation of the sensors, and selectively operatable contractive elements arranged in the legs and/or arms portions and configured to allow peristaltic movement of blood in the arms and legs of a patient wearing the body surgical suit. | 1. A surgical suit to be worn by surgery patients during surgical procedures, said surgical suit being configured to encase the limbs and torso of a patient and comprising: one or more openings allowing to access one or more portions of the body of a patient wearing the surgical suit;one or more sensors configured to monitor the patient's vital signs;a control monitor configured to control operation of said sensors; andselectively operatable contractive elements arranged in the legs and/or arms portions and configured to allow peristaltic movement of blood in the arms and legs of a patient wearing the surgical suit,wherein:said sensors are built-in components of the surgical suit, andelectrical wires and cables of the sensors and said control monitor are embedded in the suit fabric structure and arranged so as to have a single electric output allowing connection of all the built-in components to the respective devices arranged in an operation roomand wherein:said surgical suit further includes a vacuum system comprising a plurality of sucking ducts ending with suctions cups, said suction cups being arranged on portions of the surgical suit intended to face a patient supporting surface. 2. The surgical suit of claim 1, wherein the sucking ducts are connected to a single common sucking port arranged on a portion of the surgical suit and configured to be attached to a vacuum system provided in an operation room. 3. The surgical suit of claim 1, wherein the ducts of the sucking system are formed in a polymeric material layer sandwiched between two fabric layers and wherein suction cups are embedded or simply connected to the end of the ducts and so configured to protrude outwards relative to the surgical suit. 4. The surgical suit of claim 1, wherein the suction cups are arranged on a back side of the surgical suit at wrist and shoulder portions, at back and bottom portions, as well as at leg portions. 5. The surgical suit of claim 1, wherein each suction cup comprises means configured to allow to adjust the flow rate of the air sucked by a vacuum system provided in an operation room to which the surgical suit is attached. 6. The surgical suit of claim 1, wherein said sensors comprise electrocardiology (ECG) sensors, body temperature sensors, patient position sensors, as well as blood pressure and heart beat sensors. 7. The surgical suit of claim 1, further comprising defibrillator pads, wherein said defibrillator pads are arranged at the portions configured to contact the right side of the patient's sternum just below the clavicle and the left anterior axillary line between the fifth and the sixth ribs and wherein the defibrillator pads are operably connected to the control monitor. 8. The surgical suit of claim 1, further comprising an electronic unit with a microprocessor operatively connected to the sensors and devices and configured for wireless transmission of data to and from said sensors and devices. 9. The surgical suit of claim 1, f... | THE BOARD OF TRUSTEES OF THE UNIVERSITY OF ILLINOIS | A61B005/00 A41D013/12 A61B090/98 A61B005/0205 A61H009/00 A61M035/00 A61N001/04 A41D001/00 A61B005/021 A61B005/024 A61B005/0402 A61B005/11
US09351656 | B2 | 2016-05-31 | System for determining the condition of a heart valve | 2015-07-15 | An implantable medical device has an impedance processor for determining atrial impedance data reflective of the cardiogenic impedance of an atrium of a heart during diastole and/or systole of heart cycle. Ventricular impedance data reflective of the cardiogenic impedance of a ventricle during diastole and/or systole are also determined. The determined impedance data are processed by a representation processor for estimating a diastolic and/or a systolic atrial impedance representation and a diastolic and/or a systolic ventricular impedance representation. A condition processor determines the presence of any heart valve malfunction, such as valve regurgitation and/or stenosis, of at least one heart valve based on the estimated atrial and ventricular impedance representations. | 1. A system comprising: an implantable medical device;a non-implantable communication and processing device comprising: a receiver for wirelessly receiving data transmitted by said implantable medical device;said implantable medical device comprising:a lead connecting arrangement connectable to multiple cardiac leads each having at least one electrode;a signal generator that generates a first electric signal applicable over at least a portion of an atrium of a first side of a heart by a pair of electrodes of said multiple cardiac leads and generating a second electric signal applicable over at least a portion of a ventricle of said first side of said heart by a pair of electrodes of said multiple cardiac leads;a transmitter that wirelessly transmits data to said non-implantable communication and processing device;an impedance processor that determines, based on said first electric signal and a first resulting electric signal measured over at least a portion of said atrium by a pair of electrodes of said multiple cardiac leads, atrial impedance data reflective of a cardiogenic impedance of said atrium during at least a sub-phase of at least one heart cycle and determining, based on said second electric signal and a second resulting electric signal measured over at least a portion of said ventricle by a pair of electrodes of said multiple cardiac leads, ventricular impedance data reflective of a cardiogenic impedance of said ventricle during said at least a sub-phase of said at least one heart cycle;a representation processor that estimates an atrial impedance representation based on said atrial impedance data and a ventricular impedance representation based on said ventricular impedance data;a condition processor that determines a condition of a valve of said heart based on said atrial impedance representation and said ventricular impedance representation;wherein said impedance processor determines, based on said first electric signal and said first resulting electric signal, atrial impedance data representing a cardiogenic impedance of said atrium during a diastolic phase and a systolic phase of said at least one heart cycle and determining, and based on said second electric signal and said second resulting electric signal, ventricular impedance data representing a cardiogenic impedance of said ventricle during said diastolic phase and said systolic phase of said at least one heart cycle;wherein said representation processor estimates i) a diastolic, atrial impedance representation based on said atrial aid impedance data reflective of said cardiogenic impedance of said atrium during said diastolic phase, ii) a systolic, atrial impedance representation based on said atrial impedance data reflective of said cardiogenic impedance of said atrium during said systolic phase, iii) a diastolic, ventricular impedance representation based on said ventricular impedance data reflective of said cardiogenic impedance of said ventricle during said diastolic pha... | ST. JUDE MEDICAL AB | A61B005/042 A61B005/02 A61B005/04 A61B005/0432 A61B005/00 A61B005/07 A61N001/362 A61N001/365 A61N001/368 A61N001/37
US10004409 | B2 | 2018-06-26 | System for calibrating a blood pressure measurement based on vascular transit of a pulse wave | 2015-08-22 | The invention provides a system and method for measuring vital signs (e.g. SYS, DIA, SpO2, heart rate, and respiratory rate) and motion (e.g. activity level, posture, degree of motion, and arm height) from a patient. The system features: first and second sensors configured to independently generate time-dependent waveforms indicative of one or more contractile properties of the patient's heart; and a cuff-based oscillometric module. A processing component, typically worn on the patient's body and featuring a microprocessor, receives the time-dependent waveforms generated by the different sensors and processes them to determine patient-specific calibration values for use in a continuous blood pressure measurement based on pulse wave velocity (PWV). | 1. A method of determining patient-specific calibration values for use in a continuous blood pressure measurement based on pulse wave velocity (PWV), the method comprising: positioning on the patient (i) a sensor device on a patient, the sensor device configured to measure sensor data indicative of pulse arrival time (PAT) at a first point on an extremity of the patient, the PAT defined as a time difference between ventricular depolarization occurring in a cardiac cycle, and arrival of a pressure wave at the first point resulting from blood flow caused by the ventricular depolarization through a vascular path of length (Lt), wherein the sensor device is operably connected to a processing component, and(ii) an inflatable cuff device on the patient's extremity at a second point on the extremity of the patient at a position on the vascular path between the heart and the first point, the inflatable cuff device comprising an inflatable cuff which is configured to cover a portion of the vascular path of length (Lc) and an inflation pump, the inflatable cuff device configured to measure a waveform indicative of cuff pressure on the patient's extremity during inflation, wherein the inflatable cuff device is operably connected to the processing component;receiving at the processing component the sensor data and the waveform indicative of cuff pressure, and causing the processing component to calculate therefrom a series of pulse arrival times (“PATs”) as a function of cuff pressure during inflation of the inflatable cuff to above the patient's systolic blood pressure, wherein the inflation causes a compliance change in the vascular path of length L2, thereby causing the PAT measured by the sensor device to lengthen as cuff pressure is increased; anddetermining using the processing component one or more coefficients which relate PWV to mean arterial pressure (MAP) by modeling the series of PATs measured as a function of cuff pressure as a nonlinear relationship, wherein the nonlinear relationship is expressed asPAT⁡(t)=(Lt-LcaMAPcal+pwv0)+(Lca⁡(MAPcal-Pcuff⁡(t))+pwv0)wherein pWV0 is the PWV through the portion of the vascular path covered by the inflatable cuff, and a is a constant. 2. A method according to claim 1, wherein the sensor device comprises (a) an electrical sensor comprising: at least two electrodes configured to detect electrical signals from the patient's body, and an electrical circuit operably connected to the electrodes and configured to process the detected electrical signals, the electrical sensor generating a time dependent electrical waveform indicative of one or more contractile properties of the patient's heart, and (b) an optical sensor comprising a source of electromagnetic radiation configured to irradiate tissue of the patient with radiation emitted therefrom, and a detector configured to detect one or more properties of the electromagnetic radiation after it irradiates soft tissue, the optical sensor generating a time dependent... | SOTERA WIRELESS, INC. | A61B005/021 A61B005/0205 A61B005/11 A61B005/02 A61B005/022 A61B005/026 A61B005/00 A61B005/0402 A61B005/024 A61B005/0404
US10497473 | B2 | 2019-12-03 | Systems, devices, and methods for revising or supplementing ROM-based RF commands | 2015-11-18 | Systems, devices, and methods are provided that enable the revision of RF command handling software stored in ROM, and that enable to supplementation of RF command handling software stored in ROM. Examples of the systems, devices, and methods make use of a lookup data structure stored within writable non-volatile memory. | 1. A method of producing a sensor control device that comprises a processor, radio frequency (RF) communication circuitry, a non-transitory and non-volatile writable memory, and a non-transitory and non-volatile read-only memory (ROM) having a first command handler stored thereon, wherein the first command handler is capable of executing a first RF command, the method comprising: inputting a lookup data structure and a second command handler into the writable memory of the sensor control device; andprogramming the sensor control device to determine whether a received RF command is to be executed by the second command handler stored in the writable memory or by the first command handler stored in the ROM based at least in part by reference to the received RF command in the lookup data structure,wherein the lookup data structure includes an identifier for the second command handler and an identifier for an RF command that is executable by the second command handler, andwherein the sensor control device is programmed to execute the received RF command with the second command handler if the identifier for the received RF command is located in the lookup data structure and is associated with the identifier for the second command handler. 2. The method of claim 1, wherein the identifier for the second command handler is a memory pointer. 3. The method of claim 1, wherein the second command handler is a replacement command handler for the first command handler, and wherein the sensor control device is programmed to execute the received RF command with the replacement command handler. 4. The method of claim 1, wherein the second command handler is capable of executing a second RF command not executable by the first command handler. 5. The method of claim 1, wherein the sensor control device comprises a serial interface and wherein the second command handler and the lookup data structure are input into the writable memory through the serial interface. 6. The method of claim 5, wherein the sensor control device comprises a boundary scan architecture, and the serial interface provides access to the boundary scan architecture. 7. The method of claim 5, wherein the writable memory has a memory input/output (I/O) bus that provides read and write access to the memory, wherein the serial interface is directly externally accessible and the memory I/O bus is not directly externally accessible. 8. The method of claim 1, wherein the second command handler and the lookup data structure are input into the writable memory using the RF communication circuitry. 9. A method of executing a received radio frequency (RF) command in a sensor control device, wherein the sensor control device comprises RF communication circuitry, a processor, a non-transitory and non-volatile read-only memory (ROM), and a non-transitory writable memory, wherein a first command handler is stored in the ROM, the first command handler being capable of executing a first RF command, and wher... | ABBOTT DIABETES CARE INC. | G16H040/63 G06F013/42 A61B005/145 A61B005/00 G06F013/16 G06F019/00
US09326690 | B2 | 2016-05-03 | Patient monitor with visual reliability indicator | 2008-10-14 | A patient monitor configured to display patient information in a display area. The patient information is indicative of at least one physiological parameter of a patient and includes at least one unitary display element in the display area that indicates simultaneously both a quantitative value and a reliability indication of the physiological parameter. An appearance of the display element transitions so as to represent a plurality of different reliability levels, including at least a first reliability level, a second reliability level, and at least one intermediary reliability level between the first and second reliability levels. | 1. A patient monitor comprising: a display device; anda processor configured to acquire a physiological parameter of a patient and to direct the display device to display a value of the physiological parameter,wherein the displayed value comprises text having a circumscribing outline and a fill populating the outline, andwherein the processor is further configured to direct the display device to smoothly shrink a level of the fill populating the outline over time to indicate a continuum of decreasing intermediate levels of reliability for the physiological parameter, such that the circumscribing outline of the displayed text value has a first level of fill at a first time and the circumscribing outline of the displayed text value has a second, lower level of fill at a second time,wherein the levels of reliability indicate at least in part an elapsed time since the physiological parameter was acquired. 2. A method for patient monitoring comprising: using a processor to acquire a physiological parameter of a patient;displaying on a display device a value of the physiological parameter, wherein the displayed value comprises text having a circumscribing outline and a fill populating the outline; andsmoothly shrinking a level of the fill populating the outline over time to indicate a continuum of decreasing intermediate levels of reliability for the physiological parameter, such that the circumscribing outline of the displayed text value has a first level of fill at a first time and the circumscribing outline of the displayed text value has a second, lower level of fill at a second time,wherein the levels of reliability indicate at least in part an elapsed time since the physiological parameter was acquired. 3. A patient monitor comprising: a display device; anda processor configured to acquire a physiological parameter of a patient and to direct the display device to display a value of the physiological parameter,wherein the displayed value comprises text having a fill of a first fill color, andwherein the processor is further configured to direct the display device to smoothly shrink a level of the fill over time to indicate a continuum of decreasing intermediate levels of reliability for the physiological parameter by replacing the fill with an upper fill portion having a second fill color and with a lower fill portion having the first fill color, such that a circumscribing outline of the displayed text value has a first level of fill with the first fill color at a first time and the circumscribing outline of the displayed text value has a second, lower level of fill with the first fill color at a second time, and wherein the levels of reliability indicate, at least in part, an elapsed time since the physiological parameter was acquired and wherein the reliability level is represented by the relative heights of the upper and lower portions. 4. A method for patient monitoring comprising the steps of: using a processor to acquire a physiological p... | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO. LTD. | G06F019/00 A61B005/00 A61B005/0205 G06F003/0481
US09326736 | B2 | 2016-05-03 | Method and system for visualizing mechanical ventilation information | 2011-05-20 | Methods and systems for visualizing mechanical ventilation information are provided. One system includes a user interface for a medical system having a ventilator is provided. The user interface includes a memory storing one or more sensor measurements for a patient and a processor. The processor is programmed to display a visualization of a tidal volume determined from one of the sensor measurements of the patient and display a visualization of lung compliance for the patient, wherein the tidal volume visualization and the lung compliance visualization are displayed concurrently. | 1. A user interface for a medical system having a ventilator, the user interface comprising: a memory storing one or more sensor measurements for a patient; anda processor programmed to:display a virtual gauge on a tidal volume and lung compliance (VT+C) viewer, the virtual gauge includes a VT (tidal volume) indicator and a VT and compliance indicator, the VT indicator corresponds to a tidal volume determined from one of the sensor measurements, the VT and compliance indicator is positioned along both sides of the virtual gauge indicating a degree of VT and compliance relationship with respect to a position on the virtual gauge, wherein the processor is further programmed to display a patient breath indicator corresponding to a measured tidal volume bordered on two opposing sides by the VT and compliance indicator, the patient breath indicator moving between the VT and compliance indicator of the virtual gauge based on one of the sensor measurements. 2. The user interface of claim 1, wherein the virtual gauge includes a scale portion that defines a range of lung compliance. 3. The user interface of claim 1, wherein the VT and compliance indicator includes a color pattern, the color pattern is configured to indicate different ranges of acceptable compliance levels with respect to positions on the virtual gauge, the color pattern is updated based on one of the sensor measurements. 4. The user interface of claim 1, wherein the processor is further programmed to display threshold indicators corresponding to lung inflation for the patient. 5. The user interface of claim 4, wherein the threshold indicators are a lung over inflation indicator and a lung under inflation indicator. 6. The user interface of claim 1, wherein the patient breath indicator moves upward and downward within the virtual tube structure corresponding to breaths of the patient. 7. The user interface of claim 1, wherein the processor is further programmed to display a maximum breath indicator corresponding to a maximum tidal volume for a previous patient breath. 8. The user interface of claim 1, wherein the processor is further programmed to update the display of the virtual gauge based on changes in the one or more sensor measurements. 9. The user interface of claim 8, wherein the changes in the one or more sensor measurements are based on user changes to at least one ventilator setting. 10. The user interface of claim 1, wherein the processor is programmed to display the VT and compliance indicator on both sides of the virtual gauge. 11. The user interface of claim 1, wherein the VT indicator is positioned adjacent to the VT and compliance indicator. 12. A method for presenting ventilator data, the method comprising: displaying a virtual gauge on a tidal volume and lung compliance (VT+C) viewer, the virtual gauge includes a VT (tidal volume) indicator and a VT and compliance indicator, the VT indicator corresponds to a tidal volume determined from one of the sensor measu... | General Electric Company | A61M016/00 A61B005/091 A61M016/01 A61B005/00 A61M016/10
US09352171 | B2 | 2016-05-31 | Nerve treatment system | 2013-01-27 | A method of treatment includes delivering an agent into a patient, and providing energy inside the patient so that the energy can interact with the agent to treat a nerve. A system for treatment includes a source of agent for delivery into a patient, and an energy source for providing energy inside the patient so that the energy can interact with the agent to treat a nerve. | 1. A therapeutic ultrasonic energy system, comprising: a transducer; anda processor coupled to the transducer;wherein the processor is configured to control the transducer so that the transducer delivers energy to treatment areas that are arranged approximately in a spherical or ellipsoidal configuration around a reference position associated with a blood vessel. 2. The system of claim 1, wherein the processor is configured to control one or more operation parameter for the transducer so that the transducer delivers the energy to nerve tissue on at least two sides of the blood vessel. 3. The system of claim 1, further comprising an intravascular catheter, wherein the transducer is located on the catheter. 4. The system of claim 3, further comprising a positioning system for determining a position of the intravascular catheter. 5. The system of claim 3, wherein the transducer on the catheter is configured to deliver therapeutic energy. 6. The system of claim 1, wherein the transducer is configured to deliver the energy from outside a patient to inside the patient. 7. The system of claim 1, wherein the transducer is configured to deliver the energy from inside the blood vessel to outside the blood vessel. 8. The system of claim 1, further comprising a motion tracking module to track motion of the blood vessel. 9. The system of claim 8, wherein the motion tracking module is configured to track speckles. 10. The system of claim 8, wherein the motion tracking module is configured to determine a three dimensional coordinate of the blood vessel. 11. The system of claim 8, wherein the motion tracking module is configured to track a doppler ultrasound signal from the blood vessel. 12. The system of claim 1, further comprising a patient interface configured to position the transducer so that regions of bone are avoided by the energy from the transducer. 13. The system of claim 1, wherein the transducer is configured to deliver the energy at an energy level sufficient to create a reduction in a norepinephrine output of an autonomic nerve region. 14. The system of claim 1, wherein the processor is configured to determine a position of the blood vessel, and to control the transducer to deliver the energy to one or more of the treatment areas that are about ten millimeters away from a position of the blood vessel. 15. A therapeutic energy system, comprising: an energy source; anda processor coupled to the energy source;wherein the processor is configured to control the energy source so that the energy source delivers energy to treatment areas that are arranged approximately in a spherical or ellipsoidal configuration around a reference position associated with a blood vessel. 16. The system of claim 15, wherein the processor is configured to control one or more operation parameter for the energy source so that the energy source delivers the energy to nerve tissue on at least two sides of the blood vessel. 17. The system of claim 15, further com... | Kona Medical, Inc. | A61B005/05 A61N005/06 A61B005/00 A61B008/06 A61N007/02 A61B008/08 A61B019/00 A61N007/00 A61B018/00 A61B006/00 A61B008/00
US09351689 | B2 | 2016-05-31 | Sphygmomanometer having function of calculating risk degree of circulatory system disease | 2012-02-24 | An electronic sphygmomanometer having a function to calculate an index value of risk of circulatory system disease from measured blood pressure information includes a CPU, a memory, and a display. The CPU calculates the blood pressure of a person being examined based on pressure changes of an air bladder as measured by a pressure sensor. The memory stores data of the measured blood pressure together with time information of the measurement. The CPU calculates an index value of risk of circulatory system disease of the person being examined based on a difference between a blood pressure value obtained from a measurement and a blood pressure value obtained from a previous measurement and an elapsed time between the measurement and the previous measurement, A display part displays the calculated index value of risk of circulatory system disease of the person being examined together with the measured blood pressure. | 1. An electronic sphygmomanometer comprising: an air bladder;a pressure sensor that detects an internal pressure change of the air bladder;a processor that a measures a current measurement and at least one previous measurement of blood pressure of a person being examined based on the detected internal pressure change of the air bladder, the current measurement and the previous measurement are effected at a same portion of the person being examined, the processor comprising:a judging part that judges whether to calculate an index value of risk of circulatory system disease of the person being examined based on the comparison of the current measurement and the previous measurement;an index value calculating part that calculates the index value of risk of circulatory system disease of the person being examined based on: the judging result of the judging part;a difference between a blood pressure value of the person being examined obtained from the current measurement and a blood pressure value of the person being examined obtained from the previous measurement; andan elapsed time between the current measurement and the previous measurement;a memory that stores the at least one previous measurement of blood pressure of the person being examined together with a time information of the at least one previous measurement; anda display that displays the calculated index value of risk of circulatory system disease of the person being examined and the current measured blood pressure. 2. The electronic sphygmomanometer according to claim 1, wherein the index value of risk of circulatory system disease is calculated by dividing the difference between the blood pressure value of the person being examined obtained from the current measurement and the blood pressure value of the person being examined obtained from the previous measurement by the elapsed time between the current measurement and the previous measurement. 3. The electronic sphygmomanometer according to claim 1, wherein the blood pressure value of the person being examined is systolic blood pressure obtained from each measurement. 4. The electronic sphygmomanometer according to claim 1, wherein the processor measures the current measurement and the at least one previous measurement of blood pressure of the person being examined based on the detected internal pressure change of the air bladder during an inflating process of a cuff; andwherein the processor measures the current measurement and the at least one previous measurement of blood pressure of the person being examined based on the detected internal pressure change of the air bladder during a deflating process of the cuff that follows after the inflating process,wherein the index value of risk of circulatory system disease of the person being examined is calculated by dividing a difference of the blood pressure value obtained from the measurement at the inflating process and the blood pressure value obtained from the measurement at the defl... | OMRON HEALTHCARE Co., Ltd. | A61B005/02 A61B005/00 A61B005/022
US09339195 | B2 | 2016-05-17 | Apparatus, system, and method for seizure symptom detection | 2011-01-27 | An apparatus, system, and method are disclosed for detecting seizure symptoms in an individual 102. A sensor module 202 receives physiological data for an individual 102 from one or more sensors 110, 112, such as a heart activity sensor. A feature detection module 204 detects a predefined feature 500, 520, 530, 540 in the physiological data. The predefined feature 500, 520, 530, 540 is associated with a seizure or another medical condition. An alert module 206 broadcasts an alert in response to the feature detection module 204 detecting the predefined feature 500, 520, 530, 540. | 1. An apparatus to detect seizure symptoms, the apparatus comprising: a sensor module that receives heart activity data for an individual from a heart activity sensor and that receives additional data from one or more additional sensors, the one or more additional sensors comprising one or more of: a respiration sensor that receives respiration activity from the individual;an electromyography (“EMG”) sensor measures electric activity of one or more muscles of the individual;a skin conductivity sensor that measures skin conductivity of the individual;a skin temperature sensor that measures skin temperature of the individual;a peripheral blood oxygen saturation sensor that measures saturation of peripheral blood oxygen of the individual; andan audible noise sensor that measures audible vocalizations of the individual;a feature detection module that detects a predefined feature in the heart activity data, the predefined feature associated with a seizure;a confirmation feature module that, in response to the feature detection module detecting a predefined feature in the heart rate activity, detects one or more additional predefined features associated with a seizure in additional data for the individual from the one or more additional sensors; andan alert module that broadcasts an alert in response to the feature detection module detecting the predefined feature and the confirmation feature module detecting at least one additional predefined feature. 2. The apparatus of claim 1, wherein the one or more additional predefined features comprise one or more of a predefined increase in an integrated respiration waveform, a predefined increase in respiration amplitude, a predefined change in respiration rate, a breath of at least a predefined duration, a breath of at least a predefined amplitude, a breath of at least a predefined volume, and a predefined change in a period of a respiration interval. 3. The apparatus of claim 1, wherein the one or more additional predefined features comprise one or more of a predefined change in magnitude of an integrated EMG waveform, a predefined increase in rhythmic muscle activity, a predefined increase in mean firing rate of a muscle, a predefined change in an accumulated EMG waveform within a predefined time window, and a predefined decrease in signal energy of an EMG waveform preceding the predefined feature. 4. The apparatus of claim 3, wherein the one or more additional sensors further include an opposing EMG sensor, the EMG sensor measuring muscle activity in a first muscle group and the opposing EMG sensor measuring muscle activity in an opposing muscle group, wherein the one or more additional predefined features comprise a predefined amount of correlation between muscle activity in the first muscle group and in the opposing muscle group. 5. The apparatus of claim 1, wherein the one or more additional features comprise one or more of: a predefined increase in a skin conductivity of the individual;a predefine... | RESEARCH TRIANGLE INSTITUTE | A61B005/0205 A61B005/01 A61B005/0488 A61B005/145 A61B005/04 A61B005/0402 A61B005/0432 A61B005/044 A61B007/00 A61B005/00 A61B005/024 A61B005/0245 A61B005/053 A61B005/08 A61B005/11 A61B005/113
US09327131 | B2 | 2016-05-03 | Detecting implanted medical electrical lead dislodgement using cardiac signals | 2010-12-06 | Evaluation of an implanted electrical lead condition includes comparing electrogram template features with test electrogram features. The evaluating also includes determining the implanted electrical lead condition based solely on the electrogram comparison. The compared test electrogram features and template electrogram features may be atrial amplitudes and ventricular amplitudes. The sensing may be with a quad polar lead. The compared test electrogram features and electrogram template features may account for different patient postures and/or may account for respiration modulation. | 1. A method for evaluating an implanted electrical lead condition, the method comprising: sensing a test electrogram having test atrial cardiac potentials and test ventricular cardiac potentials using an electrode coupled to the implanted electrical lead;comparing one or more features of an atrial template with one or more features of the test atrial potentials;comparing one or more features of a ventricular template with one or more features of the test ventricular potentials;comparing the one or more features of the test atrial potentials to the one or more features of the test ventricular potentials; anddetermining the implanted electrical lead condition based on the comparison of the one or more features of the atrial template with one or more features of the test atrial potentials, the comparison of the one or more features of the ventricular template with one or more features of the test ventricular potentials, and the comparison of the test atrial potentials and test ventricular potentials. 2. The method of claim 1, in which the compared test electrogram features and template electrogram features comprise atrial amplitudes and ventricular amplitudes. 3. The method of claim 2, further comprising sensing the compared test electrogram features and electrogram template features with a quad polar lead. 4. The method of claim 2, in which the compared test electrogram features and electrogram template features comprise features for a plurality of different patient postures. 5. The method of claim 2, in which the compared test electrogram features and electrogram template features comprise features modulating with respiration. 6. The method of claim 1, in which the test electrogram features and the template electrogram features comprise timing between atrial and ventricular events. 7. The method of claim 1, in which the electrogram template features are collected when the implanted electrical lead is known to be in a correct position. 8. An implantable medical device (IMD) comprising: an electrical lead having an electrode adapted to sense atrial cardiac potentials and to measure ventricular cardiac potentials;a programmable microcontroller coupled to the electrical lead, the programmable microcontroller controlling operation of the IMD;a memory coupled to the programmable microcontroller, the memory storing an atrial electrogram template and a ventricular electrogram template; anda lead condition analysis module stored in the memory, wherein, when executed by the programmable microcontroller, the lead condition analysis module configures the IMD: to measure a test electrogram having test atrial cardiac potentials and test ventricular cardiac potentials using the electrode;to compare one or more features of the atrial electrogram template with one or more features of the test atrial potentials;to compare one or more features of the ventricular electrogram template with one or more features of the test ventricular potentials;to compare the o... | PACESETTER, INC. | A61N001/37 A61B005/0452 A61B005/08 A61B005/11 A61B005/00 A61N001/39 A61N001/08
US09332922 | B2 | 2016-05-10 | Solution for internal monitoring of body | 2007-06-29 | A device, method, and system for monitoring the status of an internal part of a body using an electromagnetic transceiver operating in the microwave regime; a processing unit compares measured data with reference data to determine the status of the internal part. The solution may be arranged to continuously monitor the internal part of the body and the system may be arranged so as to be wearable or portable. | 1. A device for obtaining a representation of the status of internal parts of a body part of a patient, comprising: a processing unit (35, 800);an interface (805) to a microwave transceiver (34);a display unit (36);wherein the processing unit is arranged to:repeatedly send a control signal to the microwave transceiver for generating a pulse of microwave radiation to be radiated from an antenna system comprised of at least one antenna (31) into the body part, the body part moving during said radiation of microwave radiation;receive signals, in frequency domain, indicative of received microwave radiation from the microwave transceiver;analyze the received signals with respect to time and amplitude, a first portion of the received signals being electromagnetic waves from radiation scattered directly off the skin of the patient and a second portion of the received signals being other electromagnetic waves from the other radiation penetrating into the patient and scattered off the internal tissues of the patient, the first portion of the received signals being a first-received reflection and the second portion of the received signals being a second-received reflection, the analysis including performing an inverse Fourier transform of received signals into time domain to time-separate and extract the first-received reflection from the second-received reflection;compensate for patient movements and extract data of interest indicating reflections from the internal tissues by assigning the extracted first-received reflection as an indication of patient body movement, and assigning the extracted second-received reflection as an indication of the internal tissues;compare the analyzed signals with theoretically or measured reference data; anddetermine the status of the internal part of the body from the comparison. 2. The device according to claim 1, wherein the device is arranged to analyze the received signals with reference to a brain. 3. The device according to claim 1, wherein the device is arranged to analyze the signals with reference to a hip joint. 4. The device according to claim 1, wherein the device is arranged to analyze the signals with reference to an initial measurement of the internal part of the body. 5. The device according to claim 1, wherein the processing unit is arranged to use information relating to polarization, amplitude, and phase in analysis. 6. A method for obtaining a representation of the status of internal parts of a body, comprising the steps of repeatedly sending a control signal from a processing unit to a microwave transceiver, the body moving during said radiation of microwave radiation;receiving signals, in frequency domain, indicative of microwave signals from the microwave transceiver;analyzing the received signals with respect to time and amplitude, a first portion of the received signals being electromagnetic waves from radiation scattered directly off the skin of the patient and a second portion of the receive... | MEDFIELDS DIAGNOSTICS AB | A61B006/00 A61B005/05 A61F002/46 A61B005/00 A61F002/32 A61F002/30 A61F002/48
US09329253 | B2 | 2016-05-03 | Method to generate magnetic resonance measurement data with image contrast selected and produced by preparation pulses | 2012-09-28 | In a method and magnetic resonance apparatus to generate magnetic resonance measurement data, k-space corresponding to the examination region is subdivided into a first region and a second region. The first region includes the k-space center, and a ratio of preparation pulses radiated by the radio-frequency antenna of the magnetic resonance apparatus to RF excitation pulses radiated by the radio-frequency antenna is greater in the first region than in a second region of k-space corresponding to the examination region that is not covered by the first region of k-space. Preparation pulses in the entirety of k-space corresponding to the examination region are therefore not radiated uniformly often, but instead a reduced number of preparation pulses is applied in the second region in comparison to the first region. The total measurement time to generate magnetic resonance data, as well as a SAR exposure, are reduced. | 1. A method to generate magnetic resonance (MR) data of an examination region of an examination subject, comprising: operating an MR data acquisition unit having a measurement volume in which the examination region of the examination subject is located, to generate echo signals originating from said examination region by radiating a plurality of radio-frequency (RF) excitation pulses successively with an RF antenna of said MR data acquisition unit, and generating respectively different magnetic field gradients with a gradient coil system of said MR data acquisition unit to spatially encode said echo signals until data entries corresponding to said examination region have been entered into k-space in a memory, and radiating preparation pulses into the examination region of the examination subject with said RF antenna with a ratio of said preparation pulses to said RF excitation pulses being larger in a first region of k-space corresponding to said examination region and that comprises at least a center of k-space, than a second region of k-space corresponding to said examination region, said second region not being covered by said first region and wherein said first region is a region surrounding said k-space center with a radius depending on a time t1 that follows after a radiated RF excitation pulse up to acquisition of a following echo signal;acquiring said echo signals with an RF reception antenna after each radiated RF excitation pulse; andmaking said entries in k-space available from said memory as a measurement data file. 2. The method as claimed in claim 1 comprising selecting pulses for radiation as said preparation pulses from the group consisting of T1 preparation pulses, T2 preparation pulses, and preparation pulses configured to suppress unwanted signals. 3. The method as claimed in claim 1 comprising radiating no preparation pulses in said second region. 4. The method as claimed in claim 1 comprising selecting said ratio to be at least twice as large in said first region as in said second region. 5. The method as claimed in claim 1 comprising radiating said preparation pulses with a uniform distribution in said first region and in said second region. 6. The method as claimed in claim 1 wherein said first region comprises lines of k-space with a predetermined maximum distance from said center of k-space. 7. A magnetic resonance (MR) apparatus to generate MR data of an examination region of an examination subject, comprising: an MR data acquisition unit having a measurement volume in which the examination region of the examination subject is located;a control unit configured to operate said MR data acquisition unit to generate echo signals originating from said examination region by radiating a plurality of radio-frequency (RF) excitation pulses successively with an RF antenna of said MR data acquisition unit, and generating respectively different magnetic field gradients with a gradient coil system of said MR data acquisition un... | Siemens Aktiengesellschaft | G01R033/56 A61B005/055 G01R033/28 G01R033/48 G01R033/561
US09326708 | B2 | 2016-05-03 | Ambient temperature sensor systems and methods | 2010-03-26 | A first sensor may be configured to measure a sensed amount of a physiological parameter and to generate a first signal based on the sensed amount of the physiological parameter measured by the first sensor. A second sensor may be configured to measure a temperature and to generate a second signal based on the temperate measured by the second sensor. A housing may have heat-generating electronics including a processor that may be configured to determine an overall amount of the physiological parameter based on the first signal generated from the first sensor and the second signal generated from the second sensor. The second sensor may be thermally insulated from the heat-generating electronics. | 1. A monitoring system for measuring a physiological parameter, the system comprising: a housing having electronics, the electronics including heat-generating electronics, a first sensor and a second sensor;the first sensor configured to measure a sensed amount of a physiological parameter and to generate a first signal based on the sensed amount of the physiological parameter measured by the first sensor; andthe second sensor electrically connected to the electronics in the housing and configured to measure a temperature and to generate a second signal based on the temperature measured by the second sensor;the heat-generating electronics comprising a processor configured to determine an overall amount of the physiological parameter based on the first signal generated from the first sensor and the second signal generated from the second sensor;the second sensor held by a support structure, the second sensor thermally insulated from the heat-generating electronics and the first sensor and the second sensor are non-abutting; anda thermal insulation material arranged to thermally insulate the second sensor from the heat-generating electronics and the first sensor, while the second sensor remains electrically connected to the heat-generating electronics in the housing, wherein the thermal insulation material is configured to thermally insulate the second sensor so that the temperature measured by the second sensor is thermally insulated from heat emitted by the first sensor, and the thermal insulation material is arranged between the first sensor and the second sensor through all direct line-of-sight convection paths from the first sensor to the second sensor such that the temperature measured by the second sensor is isolated from the heat emitted by the first sensor through all direct line-of-sight convection paths. 2. The system of claim 1, wherein the physiological parameter comprises a concentration of blood glucose. 3. The system of claim 1, wherein the thermal insulation material surrounds at least a portion of the second sensor, the thermal insulation material is adapted to thermally insulate the second sensor from heat generated by the heat-generating electronics. 4. The system of claim 1, the housing having a chamber, the second sensor arranged within the chamber;the chamber being coupled to the thermal insulation material; andthe chamber sufficiently located away from the heat-generating electronics to minimize an affect of the heat-generating electronics on the temperature measured by the second sensor. 5. The system of claim 4, the chamber located at a location corresponding to approximately a furthest distance within the housing from the heat-generating electronics. 6. The system of claim 4, the chamber located at a location corresponding to a coolest portion of the housing during operation of the monitoring system. 7. The system of claim 1, wherein the thermal insulation material surrounds at least a portion of the second sensor, ... | MEDTRONIC MINIMED, INC. | A61B005/00 A61B005/145
US09339662 | B2 | 2016-05-17 | Methods and devices for accurately classifying cardiac activity | 2013-10-18 | Methods, systems, and devices for signal analysis in an implanted cardiac monitoring and treatment device such as an implantable cardioverter defibrillator. In illustrative examples, captured data including detected events is analyzed to identify likely overdetection of cardiac events. In some illustrative examples, when overdetection is identified, data may be modified to correct for overdetection, to reduce the impact of overdetection, or to ignore overdetected data. New methods for organizing the use of morphology and rate analysis in an overall architecture for rhythm classification and cardiac signal analysis are also discussed. | 1. A method of cardiac signal analysis in an implantable medical device (IMD), the IMD including a number of implantable electrodes and operational circuitry coupled to the implantable electrodes for performing the method of cardiac signal analysis, the method comprising: a) the operational circuitry detecting events in a signal;b) the operational circuitry determining which detected events to certify for use in rhythm identification by: i) applying interval analysis on a set of detected events in order to identify overdetected events such that, if a likely overdetection is identified, the likely overdetection is not certified;ii) applying analysis to identify overdetected events using interval and shape characteristics of pairs of consecutive events such that, if a likely overdetection is identified, the likely overdetection is not certified; andiii) certifying at least some detected events of the set of detected events; andc) the operational circuitry determining whether a malignant cardiac rhythm is occurring using only the certified detected events. 2. The method of claim 1 further comprising the operational circuitry performing waveform appraisal on detected events and, for each event, either passing the detected event or marking the detected event as suspect, wherein step b) is performed only on detected events that pass waveform appraisal; wherein waveform appraisal operates to determine whether a detected event is noisy and, if so, to prevent any noisy detected events from passing further into the analysis. 3. The method of claim 1 wherein step c) comprises the following: determining that an X/Y condition has been met, wherein X represents how many malignant events have been identified in a set of Y certified detected events; anddetermining that the X/Y condition meets a persistence condition. 4. The method of claim 1 wherein the signal in which events are detected is created by: x) capturing a subject signal from electrodes configured to be implanted in a patient subject; andy) conditioning the signal for use in cardiac signal analysis;wherein steps x) and y) are performed on a continuous basis, while step a) is performed iteratively using the conditioned signal. 5. The method of claim 1 wherein step b)iii) is performed by certifying all events that pass steps i) and ii) without being identified as overdetections. 6. The method of claim 1 further comprising d) the operational circuitry delivering therapy in response to an identified malignant rhythm. 7. An implantable cardiac stimulus device (ICSD) comprising electrodes for capturing data from a recipient of the ICSD and operational circuitry coupled to the electrodes to analyze signal captured from the patient, wherein the operational circuitry is configured to perform a method comprising the following steps: a) detecting events in a signal;b) determining which detected events to certify for use in rhythm identification by: i) applying interval analysis on a set of detected events... | CAMERON HEALTH, INC. | A61B005/04 A61N001/39 A61B005/0452 A61N001/37 A61B005/00
US10010290 | B2 | 2018-07-03 | Signal processing apparatus, signal recovery apparatus, signal processing method, and signal recovery method | 2015-04-27 | A method and apparatus for processing a signal, wherein the apparatus identifies a sample interval satisfying a predetermined reference using a change in a plurality of samples based on a downsampling, acquires, based a result of the identifying, a feature point of a signal before the downsampling is applied to the signal, and acquires time information corresponding to a position of the feature point in a feature sample interval including the feature point is provided. | 1. A signal processing apparatus comprising: a processor and a memory, the memory containing computer readable code that, when executed by the processor, configures the processor to, extract second samples from first samples based on a downsampling,identify a sample interval satisfying a reference using a change the second samples,acquire, a feature point based on a sample value of at least one sample included in the sample interval, the at least one sample being a member of the first samples,acquire time information corresponding to a time difference between a time associated with the feature point on the sample interval and a time associated with an adjacent one of the first samples adjacent to the feature point on the sample interval,add multiple of the feature point to the time information, andgenerate a data stream including the value of the feature point and the time information having added thereto the multiple of the feature point. 2. The signal processing apparatus of claim 1, wherein the processor is configured to transmit the data stream to a remote apparatus through a communication interface. 3. The signal processing apparatus of claim 1, wherein the at least one sample represents a motion of a user wearing a device including the signal processing apparatus. 4. The signal processing apparatus of claim 1, wherein the signal processing apparatus is embedded in a walking assistance device. 5. The signal processing apparatus of claim 1, wherein the feature point is an inflection point of the at least one sample. 6. A signal recovery apparatus comprising: a processor and a memory, the memory containing computer readable code that, when executed by the processor, configures the processor to, receive a data stream from an external apparatus,detect a feature point of a signal sensed by the external apparatus based on data acquired from the data stream,identify time information from data having a greatest value among the acquired data,determine an index of the feature point based on the time information, andinterpolate the data based on a result of the processor determining the index to generate restored data, wherein the time information is a result of adding multiple of the feature point to time information corresponding to a time difference between a time associated with the feature point and a time associated with an adjacent one of first samples adjacent to the feature point. 7. The signal recovery apparatus of claim 6, wherein the processor is configured to upsample the acquired data. 8. The signal recovery apparatus of claim 6, wherein the external apparatus is a walking assistance device. 9. The signal recovery apparatus of claim 6, wherein the processor is configured to output the restored data. 10. The signal recovery apparatus of claim 6, wherein the processor is configured to analyze walking features of a user wearing the external apparatus based on the restored data. 11. A signal processing method comprising: extracting ... | Samsung Electronics Co., Ltd. | H04L007/00 A61B005/00 G01R025/00 H04Q009/00 A61B005/11
US09636080 | B2 | 2017-05-02 | Systems for monitoring the cardiovascular system using a heart lung sound | 2015-07-31 | The present invention relates to A system for monitoring the cardiovascular system using a heart sound, which comprises a heart sound receiving means(100) for receiving a heart sound(HS); DSP (Digital Signal Processor) module(200) for converting the heart sound(HS) received from said heart-sound receiving means(100) to a digital signal; and an arithmetic section(300) for calculating phonocardiogram(PCG) from the digital signal converted in the DSP module (200) and calculating information related to a preload or information related to a cardiac contractile force or information related to a respiratory change. | 1. A system for providing treatment information using phonocardiogram, the system comprising: a heart sound receiving means for receiving heart sounds;DSP (Digital Signal Processor) module for converting the heart sounds received from said heart-sound receiving means to digital signals;an arithmetic section for calculating phonocardiogram from the digital signals converted in the DSP module and calculating treatment information for fluid administration; andan output section for outputting the calculated phonocardiogram and the calculated treatment information,wherein the arithmetic section: i) calculates a time and an amplitude of a first heart sound and a time and an amplitude of a second heart sound from the calculated phonocardiogram,ii) calculates a systolic time interval using the difference between the time of the first heart sound and the time of the second heart sound and calculates an amplitude conversion ratio using the amplitude of the first heart sound and the amplitude of the second heart sound,iii) determines whether the calculated systolic time interval exceeds a predetermined first reference value and whether the calculated amplitude conversion ratio exceeds a predetermined second reference value, andiv) calculates the treatment information for fluid administration corresponding to the calculated systolic time interval and the calculated amplitude conversion ratio, based on the determination in the step of iii), the treatment information for fluid administration being one of predetermined four treatment informations. 2. The system for providing treatment information using phonocardiogram according to claim 1, wherein the arithmetic section calculates information related to cardiac contractile force using the calculated phonocardiogram, and wherein the arithmetic section calculates as a trend at least one of the treatment information for fluid administration and the information related to the cardiac contractile force. 3. The system for providing treatment information using phonocardiogram according to claim 1, wherein the arithmetic section calculates the systolic time variation in accordance with a predetermined method utilizing the calculated systolic time interval. 4. The system for providing treatment information using phonocardiogram according to claim 1, wherein the predetermined four treatment informations are to a first, second, third and fourth treatment information, and the above arithmetic section: calculates the first treatment information when the systolic time interval exceeds the predetermined first reference value and the amplitude conversion ratio does not exceed the predetermined second reference value;calculates the second treatment information when the systolic time interval exceeds the predetermined first reference value and the amplitude conversion ratio exceeds the predetermined second reference value;calculates the third treatment information when the systolic time interval does not exceed the predetermi... | THE ASAN FOUNDATION | A61B005/04 A61B007/02 A61B005/0402
US09730650 | B2 | 2017-08-15 | Alarm characterization for analyte monitoring devices and systems | 2016-01-15 | Methods and apparatus including determining a rate of occurrence of a glycemic excursion event, determining a frequency of an alarm activation associated with the glycemic excursion event, determining an analyte level associated with the alarm activation, and setting an alarm parameter based on one or more of the determined rate of occurrence of the glycemic excursion event, the frequency of the alarm activation associated with the glycemic excursion event or the determined analyte level are provided. | 1. A method performed by one or more processors, the method comprising: determining a rate of occurrence of a glycemic excursion event;determining a frequency of an alarm activation associated with the glycemic excursion event;determining an analyte level associated with the alarm activation;setting an alarm parameter based on at least two of (i) the determined rate of occurrence of the glycemic excursion event, (ii) the frequency of the alarm activation associated with the glycemic excursion event or (iii) the determined analyte level; andcontrolling an administration of a therapy based on the alarm parameter. 2. The method of claim 1, wherein the glycemic excursion event includes one of a hypoglycemic event or a hyperglycemic event. 3. The method of claim 1, further including determining an analyte level associated with the alarm activation based at least in part on analyte sensor data from an in vivo analyte sensor. 4. The method of claim 3, including determining a rate of change of the analyte level associated with the glycemic excursion event or the alarm activation, or a combination thereof. 5. The method of claim 3, wherein the in vivo analyte sensor comprises a plurality of electrodes including a working electrode, wherein the working electrode comprises an analyte-responsive enzyme and a mediator, wherein at least one of the analyte-responsive enzyme and the mediator is chemically bonded to a polymer disposed on the working electrode, and wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer. 6. The method of claim 1, wherein determining the rate of occurrence of the glycemic excursion event includes: monitoring a glucose level for a predetermined time period;detecting the monitored glucose level transcending a predetermined threshold glucose level; anddetermining the rate of occurrence based on the detected monitored glucose level transcending the predetermined threshold glucose level within the predetermined time period. 7. The method of claim 1, including programming an alarm function based on the parameter for the alarm activation. 8. The method of claim 1, wherein the frequency of the alarm activation is a number of alarm occurrences over a predetermined time period, and wherein setting the parameter for the alarm activation includes modifying a threshold level of the alarm activation based on determining the number of alarm activations. 9. The method of claim 1, wherein modifying the parameter for the alarm activation is programmable to reduce the rate of occurrence of the glycemic excursion event. 10. The method of claim 1, including generating a notification associated with an analyte monitoring system based on the alarm parameter. 11. An apparatus, comprising: an interface component;one or more processors operatively coupled to the interface component; anda memory storing instructions which, when executed by the one or more processors, causes the one or more processors t... | Abbott Diabetes Care Inc. | A61B005/05 A61B005/00 A61B005/145
US09730598 | B2 | 2017-08-15 | Systems and methods for performing electrocardiograms | 2015-11-14 | A system for performing an electrocardiogram (ECG) can include a handheld electrocardiograph device having a right arm electrode, a left arm electrode, and a left leg electrode, and can be configured to receive signals from the electrodes and to send data based on the electrode signals to a mobile electronic device. The mobile electronic device can be configured to process and analyze the receive information to provide ECG data, such as 6-lead ECG data. The mobile electronic device can analyze the ECG data to provide diagnostic information. The mobile electronic device can transfer the ECG data to a remote computing system, which can analyze the ECG data to provide diagnostic information. | 1. An apparatus for performing an electrocardiogram comprising: a housing comprising three electrodes, wherein the three electrodes are not coupled to the apparatus with wires exterior to the housing, wherein the three electrodes are disposed on a single surface of the housing;an attachment mechanism for removably attaching the housing to a mobile device;a communication interface;one or more computer readable storage media;program instructions stored on the one or more computer readable storage media that, when executed by a controller, direct the controller to: receive signals from the three electrodes;analyze the signals to determine signal-related data; andtransmit the signal-related data, via the communication interface, to the mobile device. 2. The apparatus of claim 1, further comprising: one or more amplifiers configured to amplify analog signals received from the three electrodes; andprogram instructions stored on the one or more computer readable storage media that, when executed by the controller, direct the controller to receive and analyze amplified analog signals from the one or more amplifiers. 3. The apparatus of claim 1, further comprising: an analog signal processor configured to perform analog signal processing on analog signals received from the three electrodes; andprogram instructions stored on the one or more computer readable storage media that, when executed by the controller, direct the controller to receive and analyze processed analog signals from the analog signal processor. 4. The apparatus of claim 1, further comprising: an analog-to-digital converter configured to convert analog signals from the three electrodes to digital signals; andprogram instructions stored on the one or more computer readable storage media that, when executed by the controller, direct the controller to receive and analyze the digital signals from the analog-to-digital converter. 5. The apparatus of claim 1, further comprising: second program instructions stored on second computer readable storage media that, when executed by a second controller in the mobile device, direct the second controller to: receive the signal-related data transmitted by the communication interface; andprovide 6-lead electrocardiogram data based at least in part on the signal-related data. 6. The apparatus of claim 5, further comprising program instructions stored on the second computer readable storage media that, when executed by the second controller, direct the second controller to: display user interface elements configured to receive input from a user to initiate an electrocardiogram procedure; andin response to receiving the input from the user, sending an instruction via the communication interface to initiate the electrocardiogram procedure. 7. The apparatus of claim 6, wherein the user interface elements allow the user to set a configurable delay time for initiating the electrocardiogram procedure, and wherein the mobile device delays the sending of the ... | - | A61B005/04 A61B005/0404 A61B005/00 A61B005/0408 A61B005/0428 A61B005/0432 A61B005/044
US09351651 | B2 | 2016-05-31 | Magnetic field measurement apparatus | 2014-04-02 | A magnetic field measurement apparatus includes a probe beam irradiating unit that radiates a probe beam. A gas cell is arranged on an optical axis of the probe beam and shows a linear dichroism with respect to the probe beam between a first axis and a second axis. Another gas cell is arranged on the optical axis of the probe beam on the opposite side of the probe beam irradiating unit with respect to the gas cell, and shows a linear dichroism with respect to the probe beam between a third axis and a fourth axis, which are different from the first and the second axes. A measuring unit measures a difference between magnetic fields in the gas cells and on the basis of amounts of change of a plane of polarization of the probe beam that has passed through the gas cell and the gas cell. | 1. A magnetic field measurement apparatus comprising: a probe beam irradiating unit comprising optics that radiate a probe beam;a first medium arranged on an optical axis of the probe beam and showing a linear dichroism with respect to the probe beam between a first axis and a second axis;a second medium arranged on the optical axis of the probe beam on the opposite side of the irradiating unit with respect to the first medium, and showing a linear dichroism with respect to the probe beam between a third axis and a fourth axis, which are different from the first axis and the second axis; anda measuring unit comprising at least one sensor that measures a difference between a magnetic field in the first medium and a magnetic field in the second medium on the basis of amounts of change of a plane of polarization of the probe beam that has passed through the first medium and the second medium. 2. The magnetic field measurement apparatus according to claim 1, comprising: a first irradiating unit comprising optics that irradiate the first medium with a first pump beam; anda second irradiating unit comprising optics that irradiate the second medium with a second pump beam, whereinthe first medium includes an atom group composed of a plurality of atoms excited by the first pump beam in the interior thereof, andthe second medium includes an atom group composed of an plurality of atoms excited by the second pump beam in the interior thereof. 3. The magnetic field measurement apparatus according to claim 2, wherein the first irradiating unit radiates the first pump beam in a direction which is not in parallel to the direction in which the second irradiating unit radiates the second pump beam. 4. The magnetic field measurement apparatus according to claim 1, comprising: a beam splitter arranged at a position interposed between the first medium and the second medium, and configured to separate a radiated light beam and generate a first pump beam that travels toward the first medium and a second pump beam that travels toward the second medium; anda third irradiating unit comprising optics that irradiate the beam splitter with the light beam, whereinthe first medium includes an atom group composed of a plurality of atoms excited by the first pump beam in the interior thereof, andthe second medium includes an atom group composed of an plurality of atoms excited by the second pump beam in the interior thereof. 5. The magnetic field measurement apparatus according to claim 1, comprising: a plurality of beam splitters arranged at positions interposed between the first medium and the second medium, and configured to separate a radiated light beam and generate a first pump beam that travels toward the first medium and a second pump beam that travels toward the second medium;a plurality of third irradiating units provided in one-to-one correspondence with the plurality of beam splitters, and configured to irradiate the corresponding beam splitters with the light b... | SEIKO EPSON CORPORATION | A61B005/05 A61B005/04 G01R033/032
US10117577 | B2 | 2018-11-06 | Visual field calculation apparatus and method for calculating visual field | 2016-01-14 | A visual field calculation apparatus capable of accurately calculating a visual field range of a user without using a complex configuration is provided. The visual field calculation apparatus includes a saccade detector that detects a saccade on the basis of a first gaze direction detected at a first timing and a second gaze direction detected at a second timing, a saccade speed calculator that calculates speed of the saccade on the basis of a time difference between the first timing and the second timing, the first gaze direction, and the second gaze direction, and a visual field range calculator that calculates a displacement vector of a saccade whose speed exceeds a first threshold, and calculates an area including a final point of the displacement vector as the visual field range of the user. | 1. A visual field calculation apparatus comprising: a saccade detector that obtains gaze direction information including gaze directions of a user in a saccade detection period, and detects a first plurality of saccades on a basis of a moving amount of each of the gaze directions in unit time;a saccade speed calculator that calculates a speed of each of the plurality of first saccades in the unit time; anda visual field range calculator that extracts a second plurality of saccades whose speeds exceed a first threshold from among the first plurality of saccades, calculates displacement vectors of the second plurality of saccades, and calculates an area including an initial point and a final point of each of the displacement vectors as the visual field range of the user,wherein the first threshold is a value which separates exogenous saccades and endogenous saccades, andthe visual field range of the user is calculated in a range of a surface, and information is displayed on the surface relative to the visual field range. 2. The visual field calculation apparatus according to claim 1, further comprising: a model holder that holds at least one of visual field models, each visual field model having a predetermined shape,wherein the visual field range calculator calculates the visual field range of the user by enlarging or reducing one of the at least one visual field models on a basis of the initial point and the final point of each of the displacement vectors. 3. The visual field calculation apparatus according to claim 2, wherein the predetermined shape s one of a circle, an ellipse for which a ratio of a major axis and a minor axis is set to a certain value, and an arbitrary shape for which a reference point is defined. 4. The visual field calculation apparatus according to claim 1, further comprising: a visual field calculation time setter that sets a saccade detection period, in which the first plurality of saccades is detected,wherein the visual field calculation time setter estimates a second point in time when the visual field range of the user has changed before a first point in time when the calculation of the visual field range of the user has started, andwherein the visual field calculation time setter sets a period from the second point to the first point as the saccade detection period. 5. The visual field calculation apparatus according to claim 1, further comprising: a face direction detector that detects directions of a face of the user in the saccade detection period,wherein the visual field range calculator calculates, for each saccade of the first plurality of saccades, a difference in the direction of the face of the user between a point in time when the saccade starts and a point in time when the saccade ends as a face direction displacement,wherein the visual field range calculator extracts, from among the second plurality of saccades, only a third plurality of saccades in which the face direction displacement is equal to or... | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | A61B003/113 A61B005/00 G02B027/01 A61B003/024 G06K009/00 G08G001/16 A61B003/00 A61B005/0496 G06F003/01
US09345433 | B1 | 2016-05-24 | Affixation of objects to garments | 2015-05-01 | Described herein are methods and apparatuses to mate the object with a cap through a fabric portion of the garment. The methods and apparatuses described herein may address the needs for a safe, secure, and detachable method and apparatus for affixing an object (such as an encased wireless sensor) to a garment worn by a subject. | 1. A method of safely, securely, and detachably affixing an object onto a portion of a garment, wherein the object is configured to mate with a cap through a fabric region of the garment, the method comprising: placing the fabric region between the object and the cap; andmating the object and the cap together with the fabric region therebetween,further wherein the object is shaped in a disc with one or a plurality of spiral-shaped threads configured on its side surface, and the cap is shaped in a ring with one or a plurality of spiral-shaped tracks configured on its inner surface, so that the threads and the tracks are engaged when screwing the object into the cap through the fabric, further comprising an interlocking region with teeth configured on the threads of the object and tabs configured on the cap, wherein the teeth are interlocked with the tabs when the object and the cap are screwed together through the fabric, wherein the teeth are arranged with a decreasing gap between adjacent teeth so that different thicknesses of fabric may be held between the object and the cap. 2. The method of claim 1, further comprising removing the object from the garment by disengaging the object from the cap. 3. The method of claim 1, further comprising removing the object from the garment by squeezing the cap to release the interlocking teeth and tabs and then unscrewing the object from the cap. 4. The method of claim 1, wherein placing comprises placing the fabric region between the object and the cap wherein a body flexion angle sensor is coupled to the object. 5. The method of claim 4, further wherein the body flexion angle sensor includes a microcontroller and is configured to produce a signal corresponding to a flexion angle of a body of a subject wearing the garment and to communicate the signal to the microcontroller to detect a change in the subject's body flexion angle, and to generate, transmit or generate and transmit an alert when the change in the subject's body flexion angle exceeds a threshold indicating that the subject has gotten up form a bed or out of a chair. 6. An apparatus configured to safely, securely, and detachably affix an object to a portion of a garment, the apparatus comprising: an object shaped in a disc having a thickness, comprising one or a plurality of spiral-shaped threads on a side surface,an annular cap comprising one or a plurality of spiral-shaped tracks on an inner surface configured to mate with the spiral-shaped threads of the object through a fabric of the garment, andan interlocking region with teeth on the spiral-shaped threads of the object and tabs on the cap, configured so that the teeth interlock with the tabs when the object and the cap are screwed together through the fabric of the garment, wherein the teeth are arranged with a decreasing gap between the teeth so that different thicknesses of fabric may be held between the object and the cap. 7. The method of claim 6, further comprising a body flexio... | SensaRx, LLC | A61B005/11 A61B005/00
US09339372 | B2 | 2016-05-17 | Breast implant with regionalized analyte sensors responsive to external power source | 2014-06-12 | Breast implants including sensor modules and related methods are described herein. Breast implants include those with: a shell configured to be substantially filled with a viscous material; a plurality of projections extending from an external surface of the shell, the projections forming a plurality of compartments adjacent to the external surface of the shell; at least one fluid-permeable cover attached to the projections, the cover completely enveloping the shell and the plurality of projections; and a plurality of sensor modules attached to the shell and positioned at a distance from each other, each of the sensor modules oriented to detect one or more analytes in a fluid within one of the plurality of compartments, wherein each of the plurality of sensor modules includes a unique identifier and is configured to utilize energy transmitted from an external source. | 1. A method of monitoring information from a breast implant, comprising: receiving first information from a first sensor module attached to a shell of a breast implant within an individual, wherein the first information includes a first unique sensor module identifier and sensor data from the first sensor module;receiving second information from a second sensor module attached to the shell of the breast implant within the individual, wherein the second information includes a second unique sensor module identifier and the sensor data from the second sensor module;forming an initial record from the first information and the second information, the initial record including positional information regarding the first sensor module and the second sensor module;calculating deviation limits regarding the initial record;setting deviation parameters based on the deviation limits and a predetermined set of standards;saving the initial record and the deviation parameters in a memory in a computing device;receiving third information from the first sensor module attached to the shell of the breast implant within the individual, wherein the third information includes the first unique sensor module identifier and the sensor data from the first sensor module;receiving fourth information from the second sensor module attached to the shell of the breast implant within the individual, wherein the fourth information includes the second unique sensor module identifier and the sensor data from the second sensor module;updating the initial record with the third information and the fourth information;saving the updated record in memory in the computing device;comparing the updated record to the initial record and to the deviation parameters; andindicating if the updated record is within the deviation parameters of the initial record and the positional information for each of the first sensor module and the second sensor module. 2. The method of claim 1, comprising: receiving fifth information from the first sensor module attached to the shell of the breast implant within the individual, wherein the fifth information includes the first unique sensor module identifier and the sensor data from the first sensor module;receiving sixth information from the second sensor module attached to the shell of the breast implant within the individual, wherein the sixth information includes the second unique sensor module identifier and the sensor data from the second sensor module;updating the initial record with the fifth information and the sixth information;saving the updated record in the memory in the computing device;comparing the updated record to the initial record and to the deviation parameters;indicating if the updated record is within the deviation parameters of the initial record. 3. A method of monitoring information from a breast implant, comprising: directing, from an ex-vivo device, a non-electrical power source to a breast implant in-vivo;sending, from a remote devi... | Elwha LLC | A61F002/12 A61B005/00 A61B005/145 A61B005/1473
US10537275 | B2 | 2020-01-21 | System and method for non-invasive, intracranial brain motion monitoring | 2014-06-25 | A system and a method is provided for assessing motion of a biological tissue of a subject including one or more superficial biological layers and a targeted biological layer. An optical perturbation is introduced within the one or more superficial biological layers but not within the targeted biological layer. A set of optical signal data is acquired preceding, during, or following the optical perturbation and, using the set of optical signal data, a set of optical characteristics is determined that is representative of light transiting the biological layers. Using the set of optical characteristics and a model of the biological layers, a target optical signal consistent with a target biological layer is separated and a movement of the desired biological tissue is determined using the target optical signal. | 1. A method for assessing a motion of a biological tissue of a subject, the biological tissue comprising one or more superficial biological layers and a targeted biological layer, the method comprising: a) arranging one or more wearable elements on the subject, the one or more wearable elements configured to transmit or receive optical signals at one or more near-infrared wavelengths;b) acquiring a set of optical signal data using at least one of the one or more wearable elements preceding, during, or following an optical perturbation within the one or more superficial biological layers at a first acquisition time relative to the optical perturbation;c) determining, using the set of optical signal data, a set of optical characteristics representative of light transiting the one or more superficial biological layers and the targeted biological layer;d) separating, using the set of optical characteristics and a model of the one or more superficial biological layers and the targeted biological layer, determined to be from the targeted biological layer, wherein the model relates photon propagation through the one or more superficial biological layers and the targeted biological layer and optical properties of each superficial biological layer and the targeted biological layer; ande) generating a report indicative of the motion of the targeted biological layer within the subject, wherein the motion of the targeted biological layer is calculated using the target optical signal. 2. The method of claim 1, wherein step b) further comprises acquiring a set of unperturbed optical signal data using at least one of the one or more wearable elements in the absence of the optical perturbation or preceding, during, or following the optical perturbation at a second acquisition time relative to the optical perturbation that is distinct from the first acquisition time, and wherein step c) comprises determining, using the set of optical signal data and the set of unperturbed optical signal data, the set of optical characteristics representative of light transiting the one or more superficial biological layers and the targeted biological layer. 3. The method of claim 1, wherein the one or more near-infrared wavelengths are between about 650 nanometers and about 950 nanometers. 4. The method of claim 1, wherein the one or more superficial biological layers comprise a scalp tissue layer and a skull tissue layer. 5. The method of claim 1, wherein the targeted biological layer comprises a brain tissue layer. 6. The method of claim 1, wherein the one or more wearable elements comprise a first source-detector distance and a second source-detector distance that is different than the first source-detector distance. 7. The method of claim 6, wherein the one or more wearable elements are configured to transmit and receive electromagnetic signals at a first near-infrared wavelength over the first source-detector distance and electromagnetic signals at the first near-infra... | The General Hospital Corporation | A61B006/00 A61B005/00 A61B005/1455
US20160120500 | A1 | 2016-05-05 | FETAL HEART RATE MONITORING ASSEMBLY | 2014-06-26 | A fetal heart rate measuring assembly (1) comprising an ultrasound transducer (71) adapted to perform fetal heart rate measurements when attached to the abdomen of a pregnant woman, a logic unit (31) connected to the ultrasound transducer (71), a battery (39), and a measurement presentation unit (35), wherein the entire assembly (1) is adapted to be attached to said pregnant woman. The assembly further comprises an accelerometer (73) adapted to sense movement of the pregnant woman. | 1. A fetal heart rate measuring assembly comprising an ultrasound transducer adapted to perform fetal heart rate measurements when attached to the abdomen of a pregnant woman, a logic unit connected to the ultrasound transducer, a battery, and a measurement presentation unit, wherein the entire assembly is adapted to be attached to said pregnant woman, wherein it further comprises an accelerometer adapted to sense movement of the pregnant woman, wherein the accelerometer is adapted to sense orientation of the pregnant woman with respect to gravity. 2. The fetal heart rate measuring assembly according to claim 1, wherein the fetal heart rate measuring assembly comprises an electrocardiography heart rate unit adapted to measure the heart rate of the pregnant woman, connected to electrocardiography electrodes, wherein the electrocardiography electrodes are dry electrodes and are arranged on the transducer housing. 3. The fetal heart rate measuring assembly according to claim 1, wherein the logic unit is adapted to perform at least one of the following: i) discard fetal heart rate measurements performed during excessive movements; andii) perform said fetal heart rate measurements only during periods having permitted movements. 4. The fetal heart rate measuring assembly according to claim 3, wherein the fetal heart rate measuring assembly comprises a timekeeping unit and a memory unit, and that it is adapted to record fetal heart rate measurement results in the memory unit, and allocate times (T) at which the measurements took place and orientation of pregnant woman to the results, and present the fetal heart rate measurement results with the allocated time (T) and orientation with the measurement presentation unit. 5. The fetal heart rate measuring assembly according to claim 1, wherein the fetal heart rate measuring assembly comprises a proximity sensor connected to the logic unit, wherein the logic unit is adapted to initiate a fetal heart rate measurement upon receiving an initiation signal from the proximity sensor. 6. The fetal heart rate measuring assembly according to claim 5, wherein the initiation signal comprises a plurality of succeeding separate proximity measurements according to a predetermined initiation signal pattern. 7. The fetal heart rate measuring assembly according to claim 5, wherein the fetal heart rate measuring assembly further comprises a light sensor connected to the logic unit, wherein the logic unit is adapted to discard initiation signals from the proximity sensor if the light sensor senses light values below a predetermined value. 8. The fetal heart rate measuring assembly according to claim 1, wherein the ultrasound transducer is arranged in a transducer housing, and that the battery and the logic unit are arranged in a main housing, wherein the ultrasound transducer is connected to the logic unit through a wire extending between the transducer housing and the main housing. 9. The fetal heart rate measuring as... | Laerdal Global Health AS | A61B008/08 A61B005/00 A61B005/11 A61B005/024 A61B008/02 A61B008/00
US20160135687 | A1 | 2016-05-19 | VEIN IMAGING SYSTEMS AND METHODS | 2015-06-04 | Some embodiments of this disclosure relates to systems and methods for imaging a patient's vasculature. For example, near infrared (NIR) light can be used to illuminate a target area and light that is reflected or scattered from the target area can be used for generate an image of the target area. In some embodiments, the system can be configured such that the image shows the presence, absence, or extent of infiltration or extravasation in the target area. The system can be configured to document that presence, absence, or extend of infiltration or extravasation at an infusion site. In some embodiments, an imaging system can be mounted onto a patient so that the imaging system can monitor an infusion site, and the imaging system can be configured to automatically detect the presence of infiltration or extravasation. | 1.-194. (canceled) 195. A method of detecting infiltration or extravasation in tissue surrounding an infusion site on a patient, the method comprising: illuminating the tissue surrounding the infusion site with near infrared (NIR) light of a first wavelength during a first time;receiving light of the first wavelength reflected or scattered from the tissue onto one or more light sensors;illuminating the infusion site with near infrared (NIR) light of a second wavelength during a second time;receiving light of the second wavelength reflected or scattered from the tissue onto the one or more light sensors;generating one or more images of the tissue using the light of the first wavelength received by the one or more light sensor and using the light of the second wavelength received by the one or more light sensors;displaying the one or more images of the tissue to a medical practitioner; anddetecting the presence of infiltration or extravasation in the tissue based at least in part on the one or more images of the tissue. 196. The method of claim 195, further comprising infusing an imaging enhancement agent through the infusion site, wherein detecting the presence of infiltration or extravasation in the tissue comprises identifying imaging enhancement agent that has leaked out of a vein. 197. The method of claim 195, wherein the NIR light of the first wavelength and the NIR light of the second wavelength is absorbed by hemoglobin in blood of the patient such that the one or more images are configured to distinguish between blood and the tissue, and wherein detecting the presence of infiltration or extravasation in the tissue comprises identifying blood that has leaked out of a vein. 198. The method of claim 195, wherein detecting the presence of infiltration or extravasation comprises performing image processing on the one or more images using a computer processor to detect the presence of infiltration or extravasation. 199. The method of claim 195, further comprising: associating the one or more images with a patient identifier and with time information;storing the one or more images, the associated patient identifier, and the associated time information in a patient treatment archive in a computer-readable memory device;receiving a notification of a claim of medical error for the patient; andretrieving, using one or more computer processors in communication with the computer-readable memory device, the one or more images of the tissue surrounding the infusion site on the patient from the patient treatment archive. 200. The method of claim 195, wherein receiving the light of the first wavelength onto the one or more light sensors comprises receiving the light of the first wavelength onto a first light sensor; and wherein receiving the light of the second wavelength onto the one or more light sensors comprises receiving the light of the second wavelength onto a second light sensor. 201. The method of claim 195, wherein generating one or more i... |  | A61B005/00
US10149624 | B2 | 2018-12-11 | Method and device for measuring intracranial pressure, ICP, in a subject | 2015-11-02 | According to an aspect, there is provided a method for measuring the intracranial pressure, ICP, in a subject, the method comprising detecting whether spontaneous retinal venous pulsations, SRVPs, are occurring in an eye of the subject as the orientation of the head of the subject changes; identifying the orientation of the head of the subject at which SRVPs start to occur or stop occurring; and using the identified orientation of the head of the subject at which SRVPs start to occur or stop occurring to determine the ICP in the subject. | 1. A method for measuring the intracranial pressure (ICP) in a subject, the method comprising: with a first sensor, detecting whether spontaneous retinal venous pulsations (SRVPs) are occurring in an eye of the subject as the orientation of the head of the subject changes;with a second sensor attached to a table, bed or chair on which the subject is sitting or lying, measuring the orientation of the head of the subject at which SRVPs start to occur and stop occurring, the measured orientation of the head of the subject being an angle of rotation of the head about an axis that is perpendicular to a sagittal plane through the head or body of the subject; andwith a control unit, using the measured orientation of the head of the subject at which SRVPs start to occur or stop occurring to determine the ICP in the subject by determining a pressure difference between an intraocular pressure (IOP) in the eye of the subject and the ICP at a predetermined orientation using the measured orientation. 2. The method as claimed in claim 1, wherein using the measured orientation to determine the ICP in the subject further comprises: determining the ICP from a measurement of the TOP in the eye of the subject at the predetermined orientation and the determined pressure difference at the predetermined orientation. 3. The method as claimed in claim 2, wherein, in the event that no orientation is identified at which SRVPs start to occur or stop occurring, the method further comprises: providing an output indicating that the ICP is beyond a threshold value. 4. The method as claimed in claim 1, wherein the method further comprises: changing the orientation of the table, bed or chair to change the orientation of the head of the subject while detecting, with the first sensor, whether SRVPs are occurring in the eye of the subject. 5. The method as claimed in claim 1, wherein the second sensor is one of an accelerometer, a gyroscope or a camera, and the first sensor is a pulse oximeter. 6. A non-transitory computer readable medium having computer readable code embodied therein, the computer readable code being configured such that, on execution by a suitable computer, processor or control unit, the computer, processor or control unit is caused to perform the method of claim 1. 7. A device for measuring the intracranial pressure (ICP) in a subject, the device comprising: a first sensor configured to detect whether spontaneous retinal venous pulsations (SRVPs) are occurring in the eye of the subject as the orientation of the head of the subject changes;a second sensor configured to measure an angle of rotation of the head about an axis that is perpendicular to a sagittal plane through the head or body of the subject; anda control unit configured to:identify the orientation of the head of the subject at which SRVPs start to occur and stop occurring from outputs from the first sensor and second sensor; anduse the measured orientation of the head of the subject at which SR... | KONINKLIJKE PHILIPS N.V. | A61B005/103 A61B005/117 A61B005/03 A61B003/16 A61B003/12 A61B005/00 A61B005/0205 A61B005/11 A61B005/024 A61B005/107
US09774589 | B2 | 2017-09-26 | Management method and arrangement for outpatient electrocardiography on a patient | 2015-10-29 | A management method for outpatient electrocardiography on a patient using an ECG recorder, an administration network, a first mobile device and at least one second mobile device is provided. In order to initialize outpatient electrocardiography on a patient, the first mobile device is connected to the ECG recorder and to the administration network and in the process an identification code assigned to the ECG recorder is transmitted from the ECG recorder to the administration network. After the verification of the authorizations of the mobile device, at least one certificate assigned to the ECG recorder is provided by the administration network and is transmitted to the ECG recorder for storage via the mobile device. The second mobile device is connected to the ECG recorder and to the administration network and in the process the identification code assigned to the ECG recorder is transmitted from the ECG recorder to the administration network. | 1. A management method for outpatient electrocardiography on a patient which method, with the use of an electrocardiography (ECG) recorder, an administration network, a first mobile device and at least one second mobile device, wherein software for initializing an controlling the ECG recorder and/or for displaying the ECG profile is installed on the mobile devices, comprises the following steps: storing the mobile devices in the administration network as mobile device each with stipulated authorizations;connecting the first mobile device to the ECG recorder and to the administration network for initializing outpatient electrocardiography on a patient;in the process transmission of an identification code assigned to the ECG recorder from the ECG recorder to the administration network via the first mobile device, and,after a verification of the authorizations of the first mobile device by the administration network, providing at least one certificate assigned to the ECG recorder by the administration network which is transmitted to the ECG recorder for storage via the first mobile device;connecting a second mobile device to the ECG recorder and to the administration network, so that the profile of the ECG can be checked and/or displayed by the second mobile device;in the process transmission of the identification code assigned to the ECG recorder from the ECG recorder to the administration network via the second mobile device; andtransmission of one of the previously generated certificates assigned to the ECG recorder from the administration network to the second mobile device according to the authorizations of the second mobile device; andtransmission of this certificate from the second mobile device to the ECG recorder and, after a comparison with the at least one certificate previously transmitted by the first mobile device, enabling the access by the second mobile device for the purpose of controlling the ECG recorder and/or displaying the ECC profile according to the authorizations of the second mobile device. 2. The method according to claim 1, wherein the authorizations of the mobile devices, which are stipulated in the administration network, determine which mobile devices are allowed to gain access to which functions of the ECG recorder. 3. The method according to claim 1, wherein the mobile devices are identified inside the administration network using identification codes which are assigned to the mobile devices and are transmitted from the mobile devices to the administration network when making the connection. 4. The method according to claim 1, wherein one of the second mobile devices is assigned to the patient as a patient mobile device, the authorizations allocated to the patient mobile device in the administration network being restricted to the display of the ECG profile in a framework stipulated for the patient. 5. The method according to claim 1, wherein the mobile devices use the same application software which can be downl... | GETEMED MEDIZIN—UND INFORMATIONSTECHNIK AG | G06F003/00 H04L029/06 A61B005/0432 A61B005/00 H04W012/06 H04W012/08
US09345891 | B2 | 2016-05-24 | Systems and methods for positioning implanted devices in a patient | 2015-01-15 | Systems and methods for positioning implanted devices in a patient are disclosed. A method in accordance with a particular embodiment includes, for each of a plurality of patients, receiving a target location from which to deliver a modulation signal to the patient's spinal cord. The method further includes implanting a signal delivery device within a vertebral foramen of each patient, and positioning an electrical contact carried by the signal delivery device to be within ±5 mm. of the target location, without the use of fluoroscopy. The method can still further include, for each of the plurality of patients, activating the electrical contact to modulate neural activity at the spinal cord. In further particular embodiments, RF signals, ultrasound, magnetic fields, and/or other techniques are used to locate the signal delivery device. | 1. A patient treatment system, comprising: a signal delivery system that includes a patient-implantable signal delivery device coupleable to a signal generator to deliver a neural modulation signal;a signal transmission system carried by the signal delivery system to transmit a locator signal; andan array of external signal detector elements and display elements that is movable relative to the signal delivery device, the external signal detector elements being positionable to detect the locator signal wirelessly through the patient's skin, the display elements being operatively coupled to the signal detector elements to display data based on information provided by the detector elements, wherein the array defines a plane, and wherein individual detector elements and corresponding individual display elements are co-located to have the same coordinates in the plane. 2. The treatment system of claim 1 wherein the signal delivery system includes an implanting tool releasably carried by the signal delivery device, and wherein the implanting tool includes a locator signal transmitter. 3. The treatment system of claim 2 wherein the signal delivery device has an aperture, and wherein the implanting tool includes a stylet slideably and releasably received in the aperture of the signal delivery device. 4. The treatment system of claim 1, wherein the display elements are operatively coupled to the external signal detector elements to present an indication of a signal strength detected at the external signal detector elements. 5. The treatment system of claim 1, further comprising an insertion tracker operably coupleable to the signal delivery system to track a length of the signal delivery system inserted into the patient. 6. The treatment system of claim 5 wherein the insertion tracker includes an electromechanical device that is releasably engageable with the signal delivery system. 7. The treatment system of claim 5, further comprising a computer-readable medium containing instructions that, when executed: compare a first inserted length of the signal delivery device based on information received from the external signal detector with a second inserted length of the signal delivery device based on information received from the insertion tracker; andprovide an indication when the first and second inserted lengths differ by at least a threshold amount. 8. The treatment system of claim 1 wherein the array has at least one releasable fixation element positioned to releasably secure the array to the patient's skin. 9. The treatment system of claim 1 wherein the array is carried by a flexible, conformable support member. 10. The treatment system of claim 1 wherein the signal transmission system includes an electromagnetic signal emitter. 11. The treatment system of claim 1 wherein the transmission system includes an RF signal emitter. 12. The treatment system of claim 1 wherein the transmission system includes a magnet. 13. The treatment system of ... | Nevro Corporation | A61B005/06 A61N001/372 A61B019/00 A61N001/36 A61B005/00
US09332914 | B2 | 2016-05-10 | Coaxial dual lumen pigtail catheter | 2013-11-11 | A coaxial dual lumen pigtail catheter utilizes coaxial construction incorporating a thin wall guiding catheter technology for the outer lumen and using a strong braided diagnostic technology for the central lumen to accommodate high-pressure injections. The catheter includes a manifold body to provide for connection to each of the dual lumens. The distal end of the coaxial dual lumen pigtail catheter tapers to a more flexible portion that is perforated by spiral side holes to provide for more undistorted pressure readings in the left ventricle. The coaxial dual lumen pigtail catheter also utilizes proximal straight sideholes at the end of the dual lumen portion and a taper between the dual lumen portion and the single lumen portion. | 1. A method, comprising: inserting a catheter having a coaxial dual lumen portion, including an inner lumen and a surrounding outer lumen that share a common axis, and a single lumen portion, including an extension of the inner lumen, into a vascular lumen;advancing the catheter through the vascular lumen, including positioning a segment of the single lumen portion, having one or more distal side holes, on a first side of a lesion and positioning a segment of the dual lumen portion, having one or more proximal side holes, on an opposing second side of the lesion;fluidly coupling a pressure measuring device to the one or more distal side holes of the single lumen portion and the one or more proximal side holes of the dual lumen portion, thereby operatively coupling the pressure measuring device to the first side of the lesion and the second side of the lesion, respectively;obtaining simultaneous pressure measurements from the first side of the lesion and the second side of the lesion; anddetermining a pressure differential across the lesion,wherein inserting the dual lumen portion, including the inner lumen and the outer lumen, into the vascular lumen includes inserting exactly two lumens that extend along an entire portion length from a manifold to the single lumen portion. 2. The method of claim 1, wherein inserting the dual lumen portion, including the inner lumen and the outer lumen, into the vascular lumen includes inserting a stronger braided inner tube material surrounded by a thinner-walled outer tube material into the vascular lumen. 3. The method of claim 1, wherein inserting the catheter into the vascular lumen includes inserting a catheter having a collective length of the dual lumen portion and the single lumen portion of at least 110 centimeters. 4. The method of claim 1, wherein inserting the dual lumen portion of the catheter into the vascular lumen includes inserting an outer diameter of six, seven, or eight French into the vascular lumen. 5. The method of claim 1, wherein inserting the single lumen portion of the catheter into the vascular lumen includes inserting an outer diameter of five French into the vascular lumen. 6. The method of claim 1, wherein positioning the segment of the single lumen portion on the first side of the lesion includes positioning a pigtail catheter configuration on the first side of the lesion. 7. The method of claim 6, wherein positioning the pigtail catheter configuration on the first side of the lesion includes anchoring the single lumen portion on the first side of the lesion. 8. The method of claim 6, wherein positioning the pigtail catheter configuration on the first side of the lesion includes presenting a blunt rounded structure to the internal vascular lumen wall. 9. The method of claim 1, wherein positioning the segment of the single lumen portion on the first side of the lesion includes positioning a bend catheter configuration on the first side of the lesion or across the lesion. 1... |  | A61B005/02 A61B005/0215 A61B005/00
US09345414 | B1 | 2016-05-24 | Method for providing dynamic gain over electrocardiographic data with the aid of a digital computer | 2016-01-15 | In one embodiment, ECG data collected during the long-term monitoring are compressed through a two-step compression algorithm executed by an electrocardiography monitor. Minimum amplitude signals may become indistinguishable from noise if overly inclusive encoding is employed in which voltage ranges are set too wide. The resulting ECG signal will appear “choppy” and uneven with an abrupt slope. The encoding used in the first stage of compression can be dynamically rescaled on-the-fly when the granularity of the encoding is too coarse. In a further embodiment, offloaded ECG signals are automatically gained as appropriate on a recording-by-recording basis to preserve the amplitude relationship between the signals. Raw decompressed ECG signals are filtered for noise content and any gaps in the signals are bridged. The signal is then gained based on a statistical evaluation of peak-to-peak voltage (or other indicator) to land as many ECG waveforms within a desired range of display. | 1. A method for providing dynamic gain over electrocardiographic data with the aid of a digital computer, comprising the steps of: continuously sensing and recording cutaneous action potentials with an ambulatory monitor recorder and representing each of the action potentials as a compression code in a sequence stored into the monitor recorder's memory;electronically interfacing the monitor recorder with an external download station and offloading the sequence of compression codes from the memory;decompressing the sequence of compression codes on the download station into a time series that comprises a digital representation of a raw electrocardiography signal;identifying cardiac waveforms within the time series on the download station and measuring peak-to-peak voltages in the cardiac waveforms that fall within a preferred voltage range inside of a temporal window in the time series;statistically representing a correlation between the peak-to-peak voltages on the download station and determining a signal gain that places the correlation between the peak-to-peak voltages within the preferred range; andgaining the cardiac waveforms on the download station by the signal gain. 2. A method in accordance with claim 1, further comprising the step of: defining the correlation between the peak-to-peak voltages within the preferred range as one of average voltage, maximum observed voltage, minimum observed voltage, and mean voltage. 3. A method in accordance with claim 1, further comprising the step of: selecting the signal gain such that the correlation between the peak-to-peak voltages is substantially centered within the preferred range. 4. A method in accordance with claim 1, further comprising the steps of: identifying and marking noise in the digital representation of the raw electrocardiography signal; andignoring those portions of the raw electrocardiography signal that have been marked as noise when measuring the peak-to-peak voltages in the cardiac waveforms. 5. A method in accordance with claim 4, further comprising the step of: applying a filter in the identification of the noise selected from the group comprising a high pass filter, a low pass filter, an amplitude change filter, and a zero-crossings filter. 6. A method in accordance with claim 5, further comprising the step of: triggering the high pass filter when at least eight seconds of substantially sub-0.2 mV peak-to-peak voltages are detected. 7. A method in accordance with claim 5, further comprising the steps of: triggering the low pass filter when supra-20 mV peak-to-peak voltages are detected; andmarking a span of substantially 60 seconds as noise, 30 seconds preceding and 30 seconds following the triggering of the low pass filter. 8. A method in accordance with claim 5, further comprising the steps of: triggering the amplitude change filter when peak-to-peak voltage strengths of adjoining segments of substantially three seconds change by at least 300% in strength; andmarking... | BARDY DIAGNOSTICS, INC. | A61B005/0428 A61B005/04 A61B005/0432 A61B005/00 A61B005/0452 A61B005/0408
US10362813 | B2 | 2019-07-30 | Athletic band with removable module | 2015-11-19 | A band includes a tubular body defining a central passage, where the tubular body is configured to be worn on a user's body such that a portion of the user's body is received in the central passage, and a housing formed of a polymer material and connected to the tubular body. The housing defines a cavity and an access opening providing access to the cavity, and the housing is configured to removably receive an electronic module in the cavity through the access opening. The housing further includes a slot in communication with the cavity, with the slot configured to permit passage of moisture away from the housing. The sides of the tubular body may also define a slope that is from 0-0.75. | 1. A band comprising: a tubular body having an inner surface and an outer surface, with the inner surface defining a central passage, wherein the tubular body is configured to be worn on a user's body such that a portion of the user's body is received in the central passage, the tubular body comprising a first layer and a second layer comprising an elastic fabric material; anda housing formed of a polymer material and connected to the tubular body such that a portion of the housing is received between the first and second layers, wherein the housing has an outer wall and an inner wall, a cavity defined between the outer wall and the inner wall of the housing, and an access opening providing access to the cavity, wherein the housing is configured to removably receive an electronic module in the cavity through the access opening, wherein the polymer material of the housing has greater rigidity than the elastic fabric material, and wherein the housing is positioned such that the access opening is open to the inner surface of the tubular body and accessible for insertion and removal of the electronic module from the inner surface of the tubular body,wherein the housing further comprises a slot in communication with the cavity, the slot configured to permit passage of moisture away from the housing. 2. The band of claim 1 wherein the tubular body has a top end and a bottom end, with the central passage extending between the top and bottom ends, such that the top end is configured to be located above the bottom end when the tubular body is worn on the user's body, and wherein the slot is located at an end of the housing most proximate to the bottom end. 3. The band of claim 2, wherein the housing has two opposed side walls extending from the end of the housing, and wherein the housing further comprises a second slot located in one of the side walls. 4. The band of claim 2, wherein the housing has two opposed side walls extending from the end of the housing, and wherein the housing further comprises additional slots located in each of the side walls. 5. The band of claim 1, wherein the housing further defines a sensor opening in communication with the cavity, and wherein the sensor opening is configured to permit a sensor of the electronic module to be placed in close proximity to skin of the user when the electronic module is received in the cavity. 6. The band of claim 1, wherein the tubular body has a top end and a bottom end, with the central passage extending between the top and bottom ends, such that the top end is configured to be located above the bottom end when the tubular body is worn on the user's body, wherein the housing includes two side walls extending in a direction between the top and bottom ends, and wherein the slot is located in one of the side walls. 7. The band of claim 1, wherein the housing is elongated between a first end and a second end, and the slot is located at the first end of the housing. 8. The band of claim 1, w... | NIKE, Inc. | A61B005/00 A41D001/00 A63B024/00 A41D020/00 G08B006/00 G06F003/02 G06F003/01 G06F001/16 A61B005/0205 A61B005/024 A61B005/11 A41D027/20 B29C065/02 B29C065/52 B29C065/78 B29C065/00 A41H043/04 G06F019/00 A45F005/00 G01C021/36 G16H020/40 B29K623/00 B29L029/00 B29L031/48
US09351674 | B2 | 2016-05-31 | Method for enhancing pulse oximetry calculations in the presence of correlated artifacts | 2014-08-04 | Methods and systems for determining a physiological parameter in the presence of correlated artifact are provided. One method includes receiving two waveforms corresponding to two different wavelengths of light from a patient. Each of the two waveforms includes a correlated artifact. The method also includes combining the two waveforms to form a plurality of weighted difference waveforms, wherein the plurality of weighted difference waveforms vary from one another by a value of a multiplier. The method further includes identifying one of the weighted difference waveforms from the plurality of weighted difference waveforms using a characteristic of one or more of the plurality of weighted difference waveforms and determining a characteristic of the correlated artifact based at least in part on the identified weighted difference waveform. | 1. A medical monitor, comprising: a light drive configured to drive one or more light emitters;circuitry configured to receive signals from one or more photodetectors, wherein the signals comprise a correlated artifact;a memory storing instructions configured to:combine the plurality of received signals to form a plurality of weighted difference waveforms;identify a weighted difference waveform from the plurality of weighted difference waveforms that most or least closely resembles a plethysmographic waveform; anddetermine a physiological parameter based at least in part on the identified weighted difference waveform; anda processor configured to execute the instructions. 2. The medical monitor of claim 1, wherein the memory stores instructions configured to identify a weighted difference waveforms from the plurality of weighted difference waveforms by identifying the waveform that least closely resembles a plethysmographic waveform. 3. The medical monitor of claim 2, wherein the memory stores instructions configured to determine a least representative physiological parameter based at least in part on the waveform that least closely resembles a plethysmographic waveform. 4. The medical monitor of claim 3, wherein the memory stores instructions configured to identify physiological parameter measurements that are associated with the correlated artifact. 5. The medical monitor of claim 4, wherein the memory stores instructions configured to disregard physiological parameter measurements that are associated with the correlated artifact when determining the physiological parameter. 6. The medical monitor of claim 1, wherein the correlated artifact comprises an artifact caused by patient motion, a respiratory artifact, an electronic interference artifact, or a combination thereof. 7. A pulse oximetry system comprising: a sensor configured to detect electromagnetic radiation signals corresponding to different wavelengths of light; anda processor configured to: receive the signals corresponding to the different wavelengths of light from the sensor;combine the signals to form a plurality of weighted difference waveforms that vary by a multiplier;determine a physiological parameter measurement based at least in part on the plurality of weighted difference waveforms; anddetermine that the physiological parameter is valid based on its association with a waveform of the plurality of weighted difference waveforms that most closely resembles a plethysmographic waveform. 8. The system of claim 7, wherein the processor is configured to determine the physiological parameter is invalid based on its association with a waveform of the plurality of weighted difference waveforms that least closely resembles a plethysmographic waveform. 9. The system of claim 8, wherein the waveform of the plurality of weighted difference waveforms that least closely resembles a plethysmographic waveform has a characteristic skew associated with an artifact. 10. The system of c... | COVIDIEN LP | A61B005/1455 A61B005/00 A61B005/0205 A61B005/024
US09354287 | B2 | 2016-05-31 | System and apparatus for receiving magnetic resonance (MR) signals from an imaging subject | 2007-09-07 | An apparatus for receiving magnetic resonance (MR) signals emitted by an imaging subject includes a receiver coil configured to detect the MR signals and a frequency translating preamplifier coupled to the receiver coil. The frequency translating preamplifier is configured to amplify the MR signals and to convert a frequency of the MR signals to an intermediate frequency. The frequency translating preamplifier may include an amplifier having a predefined gain, a frequency filter configured to filter at least one predetermined frequency and a mixer configured to convert the frequency of the MR signals to the intermediate frequency. | 1. An apparatus for receiving magnetic resonance (MR) signals emitted by an imaging subject, the apparatus comprising: a receiver coil configured to detect the MR signals;a frequency translating preamplifier coupled to the receiver coil and configured to amplify the MR signals and to convert a frequency of the MR signals to an intermediate frequency; anda receiver coupled to an output of the frequency translating preamplifier and configured to receive the amplified and frequency converted MR signals. 2. An apparatus according to claim 1, wherein the intermediate frequency is different than the frequency of the MR signals. 3. An apparatus according to claim 2, wherein the intermediate frequency is lower than the frequency of the MR signals. 4. An apparatus according to claim 2, wherein the intermediate frequency is higher than the frequency of the MR signals. 5. An apparatus according to claim 1, wherein the frequency translating preamplifier comprises: an amplifier having a predefined gain;a frequency filter coupled to the amplifier and configured to filter at least one predetermined frequency; anda mixer coupled to the frequency filter and configured to convert the frequency of the MR signals to the intermediate frequency. 6. An apparatus according to claim 5, wherein the mixer is configured to convert the frequency of the MR signals to the intermediate frequency based on a local oscillator signal frequency. 7. An apparatus according to claim 5, wherein the amplifier is a low noise amplifier. 8. A system for receiving magnetic resonance (MR) signals emitted by an imaging subject, the system comprising: at least one receiver coil configured to detect the MR signals;at least one frequency translating preamplifier coupled to the at least one receiver coil and configured to amplify the MR signals and to convert a frequency of the MR signals to an intermediate frequency; anda receiver coupled to the at least one frequency translating preamplifier and configured to receive and process the amplified and frequency converted MR signals. 9. A system according to claim 8, wherein the at least one receiver coil is included in an array of receiver coils. 10. A system according to claim 8, wherein the receiver is a transceiver. 11. A system according to claim 8, further comprising a connector that couples the receiver to the at least one frequency translating preamplifier. 12. A system according to claim 11, wherein the connector is a low frequency compatible connector. 13. A system according to claim 12, wherein the connector is a twisted pair cable. 14. A system according to claim 8, wherein the at least one frequency translating preamplifier comprises: an amplifier having a predefined gain;a frequency filter coupled to the amplifier and configured to filter at least one predetermined frequency; anda mixer coupled to the frequency filter and configured to convert the frequency of the MR signals to the intermediate frequency. 15. A system accor... | General Electric Company | A61B005/05 G01R033/36 A61B005/055
US09326709 | B2 | 2016-05-03 | Systems, devices and methods for managing glucose levels | 2011-03-09 | Systems, devices and methods for the management of glucose levels in the body of patient featuring user interface input mechanisms configured to provide haptic feedback to the user are provided. | 1. A glucose management device, comprising: a processor; anda user interface controllable by the processor, the user interface comprising a user input mechanism configured for tactile contact and movement by a user for entering values of one or more parameters for controlling the glucose management device;wherein the user input mechanism is configured to provide a haptic feedback response to the user, wherein the user input mechanism is controllable by the processor to adjust the haptic feedback response as the entered value of a selected parameter approaches a predefined limit. 2. The glucose management device of claim 1, wherein the haptic feedback response comprises a damping of movement of the user input mechanism wherein the extent of damping increases the closer the entered value approaches the predefined limit. 3. The glucose management device of claim 2, wherein the haptic feedback response further comprises locking the user input mechanism against further movement upon the entered value of the selected parameter exceeding the predefined limit. 4. The glucose management device of claim 2, wherein the user input mechanism is movable in at least two directions wherein movement in a first direction increases the value of the selected parameter and movement in a second direction decreases the value of the selected parameter. 5. The glucose management device of claim 4, wherein the user input mechanism is rotatable about an axis and the movement of the user input mechanism is rotational. 6. The glucose management device of claim 5, wherein the user input mechanism is further moveable in a direction along a line orthogonal to the axis of rotation. 7. The glucose management device of claim 1, wherein the user input mechanism is a toggle button. 8. The glucose management device of claim 1, wherein the user interface comprises a display for displaying the values of the one or more parameters for controlling the glucose management device, wherein the display is configured to display the entered value of the selected parameter. 9. The glucose management device of claim 8, wherein the display is configured to provide visual feedback to the user when the entered value of the selected parameter exceeds the predefined limit. 10. The glucose management device of claim 1, wherein the haptic feedback response further comprises providing a detented sensation to the user wherein one detent corresponds to a predefined unit of the entered value of the selected parameter. 11. A glucose management device, comprising: a processor; anda user interface controllable by the processor, the user interface comprising a user input mechanism configured for tactile contact and movement by a user for entering values of one or more parameters for controlling the glucose management device;wherein the user input mechanism is configured to provide a haptic feedback response to the user, wherein the user input mechanism is controllable by the processor to adjust the ha... | ABBOTT DIABETES CARE INC. | A61B005/00 A61B005/145 G06F003/01 H01H003/00
US09532743 | B2 | 2017-01-03 | Apparatus for detecting tissue hardness of living body | 2014-11-06 | An apparatus for detecting tissue hardness of a living body includes a housing, a cap, a bottom cover, a driving module, a buffering module, a spring pipeline, a push-rod and a detecting module, and the driving module includes a firing spring and a restoring spring, and the two springs have different rigidities so as to form a reciprocating structure of the firing mechanism. After the detecting module of a firing mechanism is triggered, the resilience of the restoring spring compresses the firing spring back to its firing length, so as to restore its position by a single driving point and allow users to complete the firing operation by one hand and simplify the operation of doctors to improve convenience and efficiency. | 1. An apparatus for detecting tissue hardness of a living body, comprising: a housing, being a first hollow tube, and having an outer embedding wall formed around an outer periphery of the housing, an inner embedding wall formed inside the housing, and a fixed base formed in a tube body of the housing;a cap, embedded into the outer embedding wall of the housing, and having an opening;a bottom cover, coupled to the housing, the bottom cover including a pipeline formed therein, a connecting portion disposed on a side of the bottom cover, and a through hole formed at the connecting portion; anda driving module, comprising: a pressing device, being a second hollow tube, and having a triggering rod extended separately from both sides of the pressing device;a pull-rod, installed in the second hollow tube of the pressing device;a restoring spring, installed between the pressing device and the fixed base;a trigger, fixed to the pull-rod, and having an elastic brace formed separately on both sides of the trigger, such that the elastic brace supports and abuts the inner embedding wall of the housing; anda firing spring, installed between the fixed base of the housing and the trigger, and having a coefficient of rigidity greater than a coefficient of rigidity of the restoring spring; anda buffering module, comprising: a seat body, disposed at a first end of the housing, and having a penetrating hole; anda buffer spring, sheathed on the seat body; anda spring pipeline, installed in the pipeline of the bottom cover;a push-rod, installed in the spring pipeline, the push-rod having a front end fixed to the trigger; anda detecting module, combined with a back end of the push-rod, and including: a front carrier, having a perforation and a plurality of front carrier holes;a spring duct, having a core wire fixed therein;a rear carrier, having a plurality of rear carrier holes, and a conductive silver paste filled into the rear carrier holes;a sensor, combined with the rear carrier;a plurality of conductive wires, wound around the core wire, such that after a proximal end of the conductive wires passes through the rear carrier holes of the rear carrier, the proximal ends of the conductive wires are coupled to the sensor, and after a distal end of the conductive wires passes through the front carrier holes of the front carrier, said distal end is passed out of the perforation and connected to a back-end system for reading a signal; anda waterproof layer, coated to encapsulate a periphery of the rear carrier and the spring duct. 2. The apparatus for detecting tissue hardness of a living body as claimed in claim 1, further comprising a handle cover sheathed on the outer periphery of the housing. 3. The apparatus for detecting tissue hardness of a living body as claimed in claim 1, further comprising an adhesive layer disposed at a combining position of the sensor and the push-rod. 4. The apparatus for detecting tissue hardness of a living body as claimed in claim 1,... | Southern Taiwan University of Science and Technology Kuang Tai Metal Industrial Co., Ltd. | A61B005/151 G01N033/66 A61B005/00
US09326701 | B2 | 2016-05-03 | Method and magnetic resonance system to automatically determine imaging planes | 2013-03-15 | In a method and imaging apparatus to automatically determine imaging planes for an imaging procedure to visualize a percutaneous intervention of a patient along a trajectory to a target point, data are obtained that represent a reference coordinate system with three orthogonal reference planes, and data also are acquired that represent the trajectory. One of the reference planes is determined as a selected reference plane based on the arrangement thereof relative to the trajectory. A first imaging plane is defined such that the trajectory lies therein, and such that it has a defined arrangement in relation to the selected reference plane. A second imaging plane is defined such that the trajectory lies therein, and such that it is orthogonal to the first imaging plane. A third imaging plane is defined such that it is orthogonal to the first imaging plane and second imaging plane, and includes the target point. | 1. A computerized method to automatically determine imaging planes in an imaging procedure to visualize a percutaneous intervention of a patient whom an instrument proceeds along a trajectory within the patient from a starting point to a target point within the patient while the patient is situated in an imaging apparatus, said method comprising: providing reference coordinate system data to a processor that represent a reference coordinate system of the patient, said reference coordinate system data defining three orthogonal reference planes of said reference coordinate system;providing said processor with trajectory data representing the spatial arrangement of said trajectory and said target point, said trajectory having a defined spatial arrangement between said starting point and said target point in relation to said reference planes;in said processor, automatically selecting one of said reference planes as a selected reference plane based on the arrangement of the reference planes relative to the trajectory;in said processor, defining a first imaging plane such that said trajectory lies in said first imaging plane, and such that said first imaging plane has a defined spatial arrangement in relation to the selected reference plane;in said processor, defining a second imaging plane such that the trajectory also lies in said second imaging plane and such that said second imaging plane is orthogonal to said first imaging plane;in said processor, defining a third imaging plane such that said third imaging plane is orthogonal to said first imaging plane and is orthogonal to said second imaging plane, and such that said third imaging plane comprises said target point; andfrom said processor, generating and emitting an electrical signal in a format to operate said imaging apparatus in order to acquire image data from the patient in said first, second and third imaging planes, said image data comprising data that visually represent said trajectory between said starting point and said target point. 2. A method as claimed in claim 1 comprising providing said reference coordinate system data to said processor with said three orthogonal planes respectively being a transverse plane, a sagittal plane and a frontal plane of said patient examined person. 3. A method as claimed in claim 1 comprising determining said first imaging plane such that a plane normal of said first imaging plane is parallel to a component of a plane normal of the selected reference plane that is perpendicular to said trajectory. 4. A method as claimed in claim 1 comprising determining said first imaging plane such that said first imaging plane has a minimized clearance with said selected reference plane. 5. A method as claimed in claim 4 wherein said clearance is an angle between said first imaging plane and said selected reference plane. 6. A method as claimed in claim 1 comprising determining a clearance in said processor for at least two of the reference planes relative to sa... | Siemens Aktiengesellschaft | A61B005/05 A61B005/06 A61B005/055 A61B017/34 A61B006/12 A61B006/00 G01R033/28 G01R033/54 A61B010/02 A61B019/00 A61B006/03
US09338379 | B2 | 2016-05-10 | System, device, and method for obtaining an image of a person's skin | 2013-03-18 | There is provided an imaging device including a lighting unit whose lighting directions to a subject are able to be switched, and a control unit that performs focus adjustment on the subject for every lighting directions to calculate evaluation values in accordance with focus states, and determines a direction in which a focus state becomes best as a lighting direction based on the evaluation values to capture the subject. | 1. An imaging device comprising: a lighting unit configured to provide light to a subject from a direction obtained from a plurality of switchable lighting directions which are switchable therebetween; anda control unit configured to perform focus adjustment on the subject for every one of more than one of the lighting directions to calculate evaluation values in accordance with focus states, and to determine a lighting direction to capture the subject which has an optimal focus state associated therewith as compared to the focus states based on the evaluation values. 2. The imaging device according to claim 1, wherein the lighting unit switches the lighting directions by emitting illumination light from different positions in a circumferential direction with respect to an optical axis of an imaging optical system. 3. The imaging device according to claim 2, wherein the lighting unit switches irradiation angles of the illumination light. 4. The imaging device according to claim 2, wherein the control unit thins out the lighting directions in the circumferential direction, performs focus adjustment on the subject for every lighting directions after thinning out to calculate evaluation values in accordance with focus states, and calculates evaluation values of the thinned-out lighting directions from the calculated evaluation values. 5. The imaging device according to claim 2, wherein the control unit groups the lighting directions in the circumferential direction, and performs focus adjustment on the subject for every lighting directions in one group to calculate evaluation values in accordance with focus states. 6. The imaging device according to claim 5, wherein the control unit includes lighting directions crossing at right angles in the one group. 7. The imaging device according to claim 1, wherein the lighting unit provides lighting using a light source having an emission wavelength suitable for an analysis process performed using a captured image of the subject. 8. The imaging device according to claim 1, wherein the lighting unit switches the lighting directions by switching light sources emitting illumination light, or controlling passage of illumination light emitted from a light source using a shutter. 9. The imaging device according to claim 1, wherein the control unit moves a focus position by a predetermined distance with respect to a focus adjustment position at which the focus state becomes best based on the evaluation values. 10. An imaging method comprising: providing lighting to a subject from a direction obtained from a plurality of switchable lighting directions which are switchable therebetween; andperforming focus adjustment on the subject for every one of more than one of the lighting directions to calculate evaluation values in accordance with focus states, and to determine a lighting direction to capture the subject which has an optimal focus state associated therewith as compared to the focus states based on the ... | Sony Corporation | H04N005/369 H04N005/232 H04N005/235 A61B005/00
US09355309 | B2 | 2016-05-31 | Generation of medical image series including a patient photograph | 2013-01-08 | Systems, methods, and computer-readable storage media relate to generate an image series that includes a patient image and a medical image. The patient image and the medical image may be associated based on identification information. | 1. A computer-implemented method for generating an image series, comprising: receiving at least one medical image of a patient acquired by a medical imaging device and recorded onto a recording device;receiving identification information associated with the at least one medical image when acquisition of the at least one medical image is detected by an image-activated circuitry,the identification information including an acquisition time of the at least one medical image and an identifier for the recording device; receiving at least one patient image of the patient acquired by a camera and identification information associated with the at least one patient image when the acquisition of the at least one medical image is detected, the identification information including an acquisition time of the patient image and the identifier for the recording device;associating the at least one patient image and the at least one medical image based on the identification information; andgenerating an image series that includes the at least one patient image and the at least one medical image. 2. The method according to claim 1, wherein the identification information includes at least one of an identifier for the medical imaging device that acquires the at least one medical image, image information, image practitioner information, or patient information. 3. The method according to claim 1, wherein: the identification information is included in a DICOM object; andthe recording device is a radiological cassette. 4. The method according to claim 1, further comprising: receiving a plurality of different medical images based on the identification information for each medical image,wherein the identification information for each medical image includes an acquisition time within a specific time period and an identifier for the recording device on which the each medical image is recorded. 5. The method according to claim 4, further comprising: comparing the plurality of different medical images to the at least one patient image;receiving the identification information for each of the plurality of medical images; anddetermining one or more medical images of the plurality of different medical images that corresponds to the at least one patient image based on the identifier of the recording device and the acquisition time. 6. The method according to claim 1, further comprising confirming that the image series is properly associated with at least one of a medical record, patient information, or work order. 7. The method according to claim 6, further comprising: comparing the patient of the image series to a patient associated with at least one of a medical record, patient information, or work order to determine whether the patient of the image series corresponds to the patient associated with the medical record, patient information, or work order. 8. The method according to claim 7, wherein: the comparing includes at least one of comparing the generated image series t... | Emory University Georgia Tech Research Corporation | G06K009/00 A61B005/117 G06F019/00
US09332917 | B2 | 2016-05-10 | System for non-invasive cardiac output determination | 2013-01-17 | A method determines cardiac output or stroke volume by receiving signal data representing multiple parameters of a patient concurrently acquired over a particular time period and comprising at least one of, (a) a parameter derived from an ECG waveform of the patient, (b) a parameter derived from a blood pressure signal of the patient, (c) a parameter derived from signal data representing oxygen content of blood of the patient and (d) a parameter derived from a patient cardiac impedance value. A selected parameter of the multiple concurrently acquired parameters is used in calculating a heart stroke volume of the patient comprising volume of blood transferred through the blood vessel in a heart cycle, in response to, a combination of a weighted summation of values of the selected parameter over the particular time period. Data representing the calculated heart stroke volume is provided to a destination device. | 1. A method for determining cardiac output or stroke volume, comprising the activities of: receiving, by a processing device, signal data representing a plurality of parameters of a patient, said plurality of parameters being concurrently acquired from a patient over a particular time period and comprising at least two of (a) a parameter derived from an ECG waveform of said patient,(b) a parameter derived from a blood pressure signal of said patient,(c) a parameter derived from signal data representing oxygen content of blood of said patient, and(d) a parameter derived from a patient cardiac impedance value;combining, by the processing device via a configured artificial neural network, said plurality of parameters and other information of the patient to generate coefficients as output of the artificial neural network, wherein the configured artificial neural network includes weights that are tuned using a training data set including said plurality of concurrently acquired parameters and the coefficients are generated by applying the weights between multiple layers of the artificial neural network;calculating, by the processing device, a heart stroke volume of said patient comprising volume of blood transferred through the blood vessel in a heart cycle by determining a summation of the coefficients multiplied by selected parameters of the plurality of concurrently acquired parameters and statistical values of said selected parameters over said particular time period;characterizing, by the processing device, cardiac arrhythmia using the calculated heart stroke volume; andproviding, by the processing device, data representing the characterized cardiac arrhythmia to a destination device. 2. The method according to claim 1, wherein the signal data representing oxygen content of blood of said patient comprises a blood oxygen saturation (SPO2) signal. 3. The method according to claim 1, wherein said plurality of parameters further comprises at least one of, (a) temperature of said patient, (b) invasive blood pressure of said patient, (c) non-invasive blood pressure of said patient and (d) a rate of change of blood pressure of said patient. 4. The method according to claim 1, wherein the statistical values comprise a standard deviation or a higher order statistical value. 5. The method according to claim 1, wherein calculating the heart stroke volume of said patient further comprises using a weighted combination of one or more parameters and a function of a form,∑Vital⁢⁢signs⁢⁢signals_parameters⁢(•)=∑i∈ROI_paramters⁢_number⁢αi⁡(t)·fi⁡(t)where, Vital signs signals_parameters is a function for calculating indices based on patient parameters, ROI_paramters_number represents a set of integers ranging from 1 to a total number of the patient parameters, αi(t) represents the coefficient which determines relative contribution from a particular patient parameter, fi(t) represents the patient parameters, and (t) represents time. 6. The method according to cla... | Siemens Medical Solutions USA, Inc. | A61B005/029 A61B005/00 A61B005/021 A61B005/0205 A61B005/0456 A61B005/053 A61B005/1455
US09326720 | B2 | 2016-05-03 | Wireless, implantable electro-encephalography system | 2012-02-09 | An electro-encephalography system includes an internal device for implantation below a scalp and above a skull of a patient, and an external device to be worn or carried outside the patient's body. The internal device includes a first electrode for receiving neurological signals originating from the patient's brain and a second electrode for receiving artifacts originating from sources other than the patient's brain. The external device includes a processing unit for receiving data from the internal device, mitigating the effects of the artifacts, and determining a neurological state of the patient. A therapy device is included to provide a therapy to the patient based on the patient's neurological state. | 1. An electro-encephalography system, comprising: (a) an internal device for implantation below a scalp and above a skull of a patient, the internal device comprising: (i) a first electrode disposed on a first side of the internal device for receiving a first signal comprising neurological signals originating from the patient's brain and artifacts originating from sources other than the patient's brain, the first signal comprising a first ratio of the neurological signals to the artifacts;(ii) a second electrode disposed on a second side of the internal device for receiving a second signal comprising the artifacts; and(iii) a therapy device for providing therapy to the patient; and(b) an external device in wireless communication with the internal device, the external device to be worn or carried outside the patient's body, the external device comprising a processing unit for: (i) providing power to the internal device;(ii) receiving data from the first and second electrodes, the data comprising the first and second signals;(iii) processing the received data by subtracting at least a portion of the second signal received from the second electrode disposed on the second side of the internal device from the first signal received from the first electrode disposed on the first side of the internal device;(iv) determining a neurological state of the patient based on the processed signal; and(v) transmitting instructions to the therapy device for providing therapy to the patient according to the neurological state. 2. The system of claim 1, wherein the processing unit records at least one of the processed signal or the data received from the first and second electrodes. 3. The system of claim 1, wherein the first side of the internal device is opposite the second side of the internal device. 4. The system of claim 1, wherein the therapy device comprises at least one of the first electrode or the second electrode. 5. The system of claim 1, wherein the therapy provided to the patient is selected from the group consisting of electrical stimulation, acoustic stimulation, RF stimulation, ultrasound stimulation, magnetic stimulation, electro-magnetic stimulation, photon stimulation, and a pharmaceutical compound. 6. The system of claim 1, wherein the second electrode is configured for implantation in a muscle artifact location. 7. The system of claim 1, wherein the processing unit determines the neurological state of the patient based on data received from the first and second electrodes in response to stimulation of the patient's brain. 8. The system of claim 1, wherein the neurological state is indicative of a propensity for a seizure. 9. The system of claim 1, wherein the processing unit provides trans-cutaneous wireless power to the internal device. 10. The system of claim 1, wherein the artifacts originate from a source selected from the group consisting of the patient's scalp, the patient's eyes, electromyographic signals, and motion. 11. The ... | THE CHARLES STARK DRAPER LABORATORY, INC. | A61B005/04 A61B005/00 A61B005/0478 A61N001/05 A61N001/36
US09333051 | B2 | 2016-05-10 | Method for manufacturing appliances for repositioning teeth | 2012-03-02 | A method for generating digital data sets is provided. The method includes: obtaining a first digital data set representing a first tooth arrangement; providing at least one operator for modifying a digital data set representing a tooth arrangement; providing target data which represents orthodontic treatment requirements; generating a second digital data set representing a second tooth arrangement based on the first digital data set using a state space search method and the at least one operator; determining whether the second tooth arrangement meets the treatment requirements; generating a Nth digital data set representing a Nth tooth arrangement based on a N−1th data set representing a N−1th tooth arrangement using the state space search method and the at least one operator, where N is an integer greater than 2; and determining whether the Nth tooth arrangement meets the treatment requirements. | 1. A method for manufacturing an appliance for repositioning tooth, the method comprising: obtaining a first digital data set representing a first tooth arrangement;providing at least one operator for modifying the first digital data set;providing a heuristic search method based on a cost function, wherein the cost function contains an actual cost and an estimated cost;generating a second digital data set based on the heuristic search method using the first digital data set and the at least one operator, wherein the second digital data set represents a second tooth arrangement;providing orthodontic treatment requirements;determining whether the second tooth arrangement meets the orthodontic treatment requirements;producing a positive model of the second tooth arrangement when the second tooth arrangement meets the orthodontic treatment requirement;forming a negative model based on the positive model as the appliance for repositioning tooth;generating a Nth digital data set representing a Nth tooth arrangement based on the heuristic search method using a N−1th digital data set representing a N−1th tooth arrangement and the at least one operator if the N−1th tooth arrangement does not meet the orthodontic treatment requirement, wherein N is an integer greater than 2;determining whether the Nth tooth arrangement meets the orthodontic treatment requirements;forming a first digital data set sequence having a plurality of successive digital data sets from the second digital data set to the Nth digital data set if the Nth tooth arrangement meets the orthodontic treatment requirement;producing a series of successive positive models of a corresponding series of successive tooth arrangements using the first digital data set sequence;forming a series of successive negative models based on the series of successive positive models as a series of successive appliances for repositioning tooth; andmanufacturing a plurality of appliances using the Nth digital data set. 2. The method of claim 1 further comprising: selecting in the first tooth arrangement at least one tooth needs to be repositioned. 3. The method of claim 2 further comprising: selecting at least one operator for said at least one selected tooth. 4. The method of claim 3 further comprising: setting a limitation for said at least one selected operator for said at least one selected tooth. 5. The method of claim 1 further comprising: setting a step for said at least one operator. 6. The method of claim 1, wherein the actual cost function or the estimated cost function comprises a parameter selected from the group consisting of the amount of crowding, overjet, overbite, dental arch convexity, depth of curve of Spee, Bolton ratio, dental arch width, arch symmetry, angulation of crown, tongue, midline and a combination thereof. 7. The method of claim 1, wherein said at least one operator is selected from the group consisting of movement along positive x-axis, movement along negative x-axis, movemen... | WUXI EA MEDICAL INSTRUMENTS TECHNOLOGIES LIMITED | G06F019/00 A61C007/00 G06F017/10 G06F017/50 A61B005/00
US09354290 | B2 | 2016-05-31 | Method and magnetic resonance system to generate an MR image with a tracking factor | 2012-03-02 | For the generation of a magnetic resonance (MR) image of a predetermined volume segment of a living subject, a tracking factor is determined for each of different regions of the volume segment, from which the position of the respective region can be determined depending on a position of a moving area of the subject. MR image data of the volume segment are acquired for different positions of the moving area. The position of the moving area is calculated depending on the position and the tracking factor of the respective region, and the MR image data of the respective region are reconstructed using the MR image data of the volume segment corresponding to the calculated position of the moving area. The MR image of the predetermined volume segment is generated as a combination of the constructed MR image data of the regions. | 1. A method to generate a magnetic resonance (MR) image of a predetermined volume segment of a living subject, comprising: operating an MR scanner to acquire MR data from a predetermined volume segment of a living subject, for different positions of a moving area of the living subject, said volume segment comprising a plurality of different regions within said volume segment;providing said MR data to a computer, and using said computer to determine the different position of the moving area and a respective tracking factor for each of said plurality of different regions of the volume segment that correlates movement of each different region to said moving area;for each of said different regions, using said computer to calculate a position of the respective region dependent on the position of the respective region and the tracking factor for the respective region, by determining the position of the respective region from the tracking factor dependent on a position of the moving area; andusing the MR data and the calculated positions of the respective regions to reconstruct a motion-corrected image of each respective region wherein the calculated positions of the respective regions correct for motion of the respective regions due to said moving area, and making the reconstructed motion-corrected images available in electronic form from the computer as a data file. 2. A method as claimed in claim 1 wherein the living subject is a mammal and wherein the moving area is a diaphragm of the mammal. 3. A method as claimed in claim 1 wherein the living subject is a mammal, and wherein said predetermined volume segment comprises at least a portion of a region of the heart of the mammal, including an immediate environment of the heart of the mammal. 4. A method as claimed in claim 3 wherein said predetermined volume segment comprises a coronary region of the mammal. 5. A method as claimed in claim 1 wherein the living subject is a mammal and wherein the moving area is a diaphragm of the mammal and wherein the predetermined volume segment is at least a portion of a region of the heart of the mammal, including an immediate environment of the heart of the mammal. 6. A method as claimed in claim 5 wherein said regions of the volume segment are located at least partially atop one another in a direction of a basic magnetic field of the MR scanner, and wherein the position of the respective regions corresponds to a position along the direction of the basic magnetic field. 7. A method as claimed in claim 1 comprising acquiring said MR data of said volume segment for different positions of the moving area by a step-by-step automatic entry by said computer of said data unto an electronic memory organized as k-space that corresponds to a volume that comprises the predetermined volume segment, independent of the position of the moving area, and for each step in entering the MR data into k-space, acquiring only MR data of a partial segment of k-space for the corresp... | Siemens Aktiengesellschaft | A61B005/00 G01R033/567 A61B005/055 G01R033/563
US09333047 | B2 | 2016-05-10 | Method for generating position data of an instrument | 2010-06-25 | A method for generating position data of an instrument (1), in particular a medical instrument, wherein the method includes the following steps: capturing position measurement data by way of at least one position sensor (7, 9) arranged on the instrument, determining a movement state of the instrument on the basis of the captured position measurement data, filtering the captured position measurement data on the basis of the determined movement state. A system for presenting an instrument, a computer program product including a program code, a computer program and a data storage medium including a program code. | 1. A method for generating position data of a medical instrument, the method comprising the following steps: capturing position measurement data of at least two electromagnetic position sensors which are arranged on the medical instrument, the position measurement data indicating a measured position of the medical instrument;implementing by a data processor: determining a movement state of the medical instrument on the basis of the captured position measurement data, the movement state being defined by how the medical instrument is being moved and a speed the medical instrument is being moved, the movement state being one of a positioning state in which the medical instrument is kept still in a position or moved or rotated with a speed lower than a predetermined value, a rotation state in which the medical instrument is swung around its own axis, a moving state in which the medical instrument is moved with a speed higher than a predetermined value, and an interference state in which the measured position of the medical instrument is influenced by another object, the determination comprising:calculating variables describing a correlation between the position signals of the position sensors on the basis of the captured position measurement data, the calculated variables describing a correlation between the position signals of the position sensors are correlation coefficients for each degree of freedom of the medical instrument,calculating variables describing a dimension of the divergence of captured and previously measured position measurement data, the calculated variables describing a dimension of the divergence of captured and previously measured position measurement data are variances for each degrees of freedom of the medical instrument,determining the respective movement state of the medical instrument by means of an algorithm for classification comprising comparisons of the calculated variables with predetermined values where the algorithm for classification based on a decision tree comprising decision criteria to be checked, each decision criterium comparing the calculated variables to predefined values, depending on the result of a first decision criterium, at least one further decision criterium is checked and at the end of the decision tree a different threshold value is allocated to each of the determined movement states, andfiltering the captured position measurement data on the basis of at least one comparison of the allocated threshold value with a predefined plausibility value to determine if the predefined plausibility value exceeds the allocated threshold value; anddisplaying an image data set based on the filtered captured position measurement data on a display, the image data set including a graphical depiction of the medical instrument, the graphical depiction being displayed on the display only when the predefined plausibility value does not exceed the allocated threshold value. 2. The method as recited in claim 1 further co... | Flagon GmbH | A61B005/05 A61B019/00
US09345849 | B2 | 2016-05-24 | Device and method for capturing a surface of an object | 2012-11-09 | A device (10) for capturing a surface of an object (30) is provided, wherein the device (10) comprises a flexible membrane (14), which is configured to be deformed from a first shape to a second shape, when at least a part of the flexible membrane (14) is brought into contact with the surface of the object (30), wherein the device (10) comprises a plurality of strain-gauges (18) coupled to the flexible membrane (14), wherein the strain-gauges (18) are configured to measure a plurality of strain signals that depend on the shape of the flexible membrane (14), and wherein the device (10) comprises an output unit (24) for outputting the strain signals for reconstruction of the surface of the object (30) as a function of the strain signals. | 1. A device for capturing a surface of an object comprising: a flexible membrane, which is configured to be deformed from a first shape to a second shape, when at least a part of the flexible membrane is brought into contact with the surface of the object, wherein the flexible membrane is held by a rigid frame;a plurality of strain-gauges coupled to the flexible membrane, wherein the strain-gauges are configured to measure a plurality of strain signals that depend on the shape of the flexible membrane; andan output unit for outputting the strain signals for reconstruction of the surface of the object as a function of the strain signals. 2. The device according to claim 1, wherein the first shape conforms to a plane. 3. The device according to claim 1, wherein the flexible membrane is mechanically coupled with the rigid frame by means of the strain-gauges, which are arranged all around a peripheral part of the flexible membrane. 4. The device according to claim 1, wherein the strain-gauges are mounted on a surface of the flexible membrane or wherein the strain-gauges are arranged within the flexible membrane. 5. The device according to claim 1, wherein the flexible membrane is made of silicone or an elastomer. 6. An arrangement for reconstructing a surface of an object comprising: a device according to claim 1 for providing a plurality of strain signals; anda processing unit coupled to the device for reconstructing the surface of the object as a function of the strain signals. 7. A method of providing a mask for a patient, including using the device according to claim 1 to generate 3D information of the pantient's face, using the 3D information of the patient's face for (i) manufacturing the mask for the patient, (ii) determining the shape of the mask for the patient and/or (iii) selecting the mask suitable for the patient from a predetermined set of masks. 8. A method for capturing a surface of an object by means of a flexible membrane that is coupled with a plurality of strain-gauges, the method comprising: bringing at least a part of the flexible membrane into contact with the surface of the object, wherein the flexible membrane is deformed from a first shape to a second shape, wherein the flexible membrane is held by a rigid frame;measuring a plurality of strain signals by means of the strain-gauges, wherein the strain signals depend on the shape of the flexible membrane; andreconstructing the surface of the object as a function of the strain signals. 9. The method according to claim 8, wherein the first shape conforms to a plane and wherein the second shape conforms to the surface of the object (30). 10. A method of providing a mask for a patient, including receiving 3D information of the patient's face by means of the method according to claim 8, and customizing a design of the mask for the patient based on the 3D information or identifying the mask in a mask identification program using the 3D information. | KONINKLIJKE PHILIPS N.V. | A61B005/117 A61B005/103 A61M016/06 G01B007/16 G01B007/28 A61B005/107 A61B005/00
US10226608 | B2 | 2019-03-12 | Optical capsule and spectroscopic method for treating and diagnosing the intestinal tract | 2015-09-24 | A device and method for mapping, diagnosing and treating disorders or other diseases, disorders or conditions (e.g., bleeding, ischemic or necrotic tissue, and presence of certain chemicals or substances) of the intestinal tract is provided using a capsule passing through the intestinal tract and sensing optical characteristics as the capsule passes through. Further, a capsule tracking system is provided for tracking a capsule's location along the length of an intestinal tract as various treatment and/or sensing modalities are employed. In one variation, an acoustic signal is used to determine the location of the capsule. A map of optical characteristics may be derived from the pass of a capsule to diagnose the disorder or disease. The capsule or subsequently passed capsules may treat, further diagnose or mark the intestinal tract at a determined location along its length. | 1. A system for diagnosing a gastrointestinal condition in an intestinal tract of a patient, the system comprising: an autonomous capsule sized to pass through the intestinal tract of a patient, the autonomous capsule comprising: a light source configured to emit light from the autonomous capsule so that the emitted light is reflected from intestinal tract tissue when the autonomous capsule is moving through the intestinal tract; and a sensor configured to sense reflected light for at least one wavelength at a first location within the intestinal tract and to output a signal representative of light sensed by the sensor at the at least one wavelength, each at least one wavelength associated with a corresponding condition of intestinal tract tissue;a processor operatively coupled to the sensor and configured to receive a signal representative of light sensed by the sensor, the processor further configured to associate each of the at least one wavelength with the corresponding condition of intestinal tract tissue, to select a gastrointestinal condition of intestinal tract tissue for detection, to control the emission of light from the light source for detecting the selected condition, to control wavelengths of light sensed by the sensor, and to determine a presence of the condition of intestinal tract tissue based at least in part on the received signal, the condition being at least one of a presence of a substance, an absence of a substance, or a condition of tissue of the intestinal tract; andan autonomous capsule tracking system, the autonomous capsule tracking system including multiple ultrasound receivers and an acoustic reference receiver, wherein the multiple ultrasound receivers and the acoustic reference receiver are configured to establish multiple differential times between the multiple ultrasound receivers and the acoustic reference receiver; anda treatment capsule configured to treat the condition determined to be present by the autonomous capsule, wherein the treatment capsule includes a first electrode located on a capsule body of the treatment capsule, and a second electrode located on an elongate member extending from the capsule body. 2. The system of claim 1, wherein the light source emits light at the at least one wavelength. 3. The system of claim 1, wherein the autonomous capsule further comprises a filter operatively coupled to the sensor, wherein the filter is configured to filter light of the at least one wavelength into the sensor. 4. The system of claim 1, wherein the autonomous capsule tracking system is configured to track a location of the autonomous capsule within a three-dimensional coordinate system. 5. The system of claim 4, wherein the processor is operatively coupled to the autonomous capsule tracking system to receive information on the location of the autonomous capsule within the intestinal tract, and wherein the processor is configured to identify a location of the selected gastrointestinal condition of in... | Entrack, Inc. | A61B005/1459 A61B005/06 A61B005/00 A61M031/00 A61B001/04 A61B005/07 A61N001/36 A61B005/03 A61B005/11 A61B005/145 A61M005/172 A61B005/053 A61B005/01 A61B018/14 A61B001/00 A61N001/372 A61B018/00
US09560992 | B2 | 2017-02-07 | Methods and systems for observation of tympanic function | 2015-10-09 | Methods and systems for endoscopic observation of tympanic function. | 1. A system for observation of two dimensional deformation and two dimensional shape of an inner ear membrane, the system comprising: a transducer that emits a predetermined acoustic input to excite an inner ear;a microphone to detect a response to inner excitation by the predetermined acoustic input;a speculum configured to examine the inner ear, the speculum comprising: a laser diode operatively connected to an optical fiber to provide an incident light into the inner ear while the inner ear is responding to the predetermined acoustic input;a gradient-index (GRIN) lens to receive a portion of the incident light, said portion being reflected or scattered due to the light inner ear excitation; andan imaging lens disposed to receive the portion of the incident light from said GRIN lens:a reference optical fiber light source that emits a reference light beam and a micromirror disposed to receive the reference light beam and configured to modulate the reference light beam;a beam splitter wedge disposed to receive the modulated light from said micromirror and the portion of the incident light from said imaging lens, said beam splitter wedge further providing light to a digital camera; wherein the digital camera obtains an interferometric recording from the received portion of the incident light, the reference lightbeam and the light received from the beam splitter wedge;a computer having computer code instructions stored thereon, the computer code instructions being configured to cause the computer to: analyze the interferometric recording to determine spatial and temporal information related to inner ear function; the spatial and temporal information comprising the two dimensional deformation and two dimensional shape of the inner ear membrane;wherein obtaining the interferometric recording further comprises providing control signals to said micromirror to: a) vary, by controllable deflection, a length of a path of the reference lightbeam by a predetermined fraction of a wavelength;b) obtain an intensity distribution corresponding to the interferometric recording between the reflected or scattered portion of the incident light and the reference lightbeam after traversing said length of said path; andc) repeat steps a) and b) for a number of predetermined fractions of the wavelength; thereby obtaining a plurality of intensity distributions in order to obtain a predetermined function characteristic of the inner ear which correlates with spatial and temporal information. 2. The system of claim 1 wherein said computer code instructions, in causing said computer to analyze the interferometric recording, further causes said computer to: apply a phase unwrapping algorithm to said predetermined function. 3. The system of claim 1 wherein said computer code instructions, in causing said computer to analyze the interferometric recording, causes said computer to: a) provide control signals to said micromirror, said control signals resulting in varying a lengt... | Worcester Polytechnic Institute | A61B006/00 A61B005/12 A61B005/00 A61B001/227 A61B001/32 A61B001/04 A61B001/00 A61B007/00 A61B001/06
US10993637 | B2 | 2021-05-04 | Operating apparatus for medical apparatus | 2015-01-14 | There is provided an operating apparatus for a medical apparatus, the operating apparatus including a ring-shaped magnet magnetically polarized in a circumferential direction or a radial direction and configured to rotate along with rotating operation by a user, and a sensor unit configured to detect a magnetic field and to output a signal depending on the detected magnetic field. The sensor unit outputs two-phase signals having different phases. | 1. An operational apparatus for a medical device, the operational apparatus comprising: a ring-shaped magnet polarized in a predetermined direction, the ring-shaped magnet to rotate in response to manual manipulation by a user;at least one sensor to detect a magnetic field and to output a signal based on the detected magnetic field, the at least one sensor being arranged in an airtight and waterproof shielded area, wherein the at least one sensor is entirely encircled by the ring-shaped magnet;a partition that forms the airtight and waterproof shielded area and separates the ring-shaped magnet from the at least one sensor, but does not shield a magnetic field, wherein the partition between the ring-shaped magnet and the at least one sensor is a non-ferromagnetic material; andan operating ring that covers the ring-shaped magnet. 2. The operational apparatus for the medical device according to claim 1, wherein the medical device is an endoscope. 3. The operational apparatus for the medical device according to claim 1, wherein the medical device is a surgical device. 4. The operational apparatus for the medical device according to claim 1, wherein the at least one sensor outputs two-phase signals having different phases. 5. The operational apparatus for the medical device according to claim 4, wherein the at least one sensor outputs signals with each respective signal level based on the detected magnetic field, as the two-phase signals. 6. The operational apparatus for the medical device according to claim 1, wherein the ring-shaped magnet includes an N-pole magnet and an S-pole magnet alternately arranged in a circumferential direction. 7. The operational apparatus for the medical device according to claim 1, wherein the ring-shaped magnet includes a magnetized portion and a void alternately arranged in a circumferential direction, the magnetized portion being magnetically polarized so that an N pole and an S pole are arranged in a radial direction. 8. The operational apparatus for the medical device according to claim 1, wherein the at least one sensor is placed so as not to contact the ring-shaped magnet, and is immovable by magnetic force of the ring-shaped magnet. 9. The operational apparatus for the medical device according to claim 4, wherein the at least one sensor outputs the two-phase signals. 10. The operational apparatus for the medical device according to claim 4, wherein the at least one sensor includes two sensors a first one of which detects a first one signal of the two-phase signals, and a second one of which detects a second one signal of the two-phase signals. 11. The operational apparatus for the medical device according to claim 10, wherein the ring-shaped magnet includes an N-pole magnet and an S-pole magnet alternately arranged in a circumferential direction, and wherein the two sensors included in the at least one sensor detect either a magnetic field of one polarity or magnetic fields of different polarities. 12.... | SONY OLYMPUS MEDICAL SOLUTIONS INC. | A61B005/06 G01D005/245 A61B001/00 G06F003/0362 A61B001/045 A61B005/05 A61B005/00 A61B017/00
US09347880 | B2 | 2016-05-24 | Method and apparatus for imaging of semi-transparent matter | 2009-08-31 | The invention relates to a method and apparatus for measuring the spatial and velocity distribution of scattering structures of a sample in semi-transparent matter. The acquired spectral data I(k, t) are transformed in two steps to image data I(z, v). A Doppler shift is imposed on the light to a separate real and mirror images in the v space to suppress the complex ambiguity artifact. | 1. Method for measuring the spatial and velocity distribution of scattering structures of a sample of semitransparent matter comprising the steps of acquiring sequentially a plurality of spectral optical coherence tomography spectra I(k) comprising data belonging to a real image and data belonging to a mirror image,for which time intervals between two recordings are known so that a data set I(k, t) representing the time dependence of the recorded spectra is generated,transforming said data set I(k, t) to a data set I(z, v) by performing a k-z transformation and a t-v transformation, where z represents the positions of back-scattering structures of the sample along the axis of the incident light and v represents the velocity component along said the axis of back-scattering structures of the sample,the method comprising:selecting a velocity vD such that, after said t-v transformation, the range for v for all data belonging to the real image does not overlap with the range for v for all data belonging to the mirror image, wherein vMin≦vD&lt;vDetMax wherein vMin is the minimum velocity producing a detectable phase difference between the data belonging to the real image and the data belonging to the mirror image and wherein for a time interval Δt between recordings of the spectra I(k), the maximum velocity vDetMax is π/(2k0Δt), where k0 is the central wavenumber, or maximum velocity vDetMax is π(2kMaxΔt), where kmax is the maximum wave number of the spectrum used for the measurement,imposing a Doppler shift on a light, whereby a sample light and a reference light used in the method appear at a point of recombination as originating from sources moving relatively to each other with the selected velocity vD, characterized in that the Doppler shift is imposed by reflecting the light with a mirror moving with a velocity component vD along the axis of the reflected light. 2. Method according to claim 1, characterized in that said Doppler shift is imposed on the reference light. 3. Method according to claim 1, characterized in that said time intervals are shorter than π/(2ko vD). 4. Method according to claim 3, characterized in that said time intervals are shorter than π/(2kmax vD). 5. Method according to claim 1, characterized in that the t-v transformation is performed first and that the k-z transformation is performed only for a portion of the data set I(k, v) for which v is beyond a certain value. 6. Apparatus for measuring the spatial and velocity distribution of scattering structures of a sample of semi-transparent matter comprising an arrangement for spectral optical coherence tomography comprising a sample path and a reference path, wherein said arrangement is configured to record spectra sequentially after known time intervals so that a data set I(k, t) representing the time dependence of the recorded spectra may be generated, said recorded spectra comprising data belonging to a real image and data belonging to a mirror image, characterized in ... | CANON OPHTHALMIC TECHNOLOGIES SP. Z O.O | G01B009/02 G01N021/47 A61B005/00 A61B003/10 A61B003/12
US09345510 | B2 | 2016-05-24 | Atherectomy catheters with longitudinally displaceable drive shafts | 2011-07-01 | Described herein are atherectomy catheters, systems and methods that include longitudinally displaceable drive shafts that drive actuation of one or more cutters at the distal end of the catheter. The catheters described herein may include one or more imaging sensors for imaging before, during or after cutting tissue. In some variations the imaging sensor may be rotated around the perimeter of the catheter independently of the rotation of the cutter. Also describe herein are imaging catheters that may be used without cutters. | 1. An atherectomy catheter device configured to visualize and to cut tissue, the device comprising: an elongate catheter body;a distal tip attached to the elongate catheter body;a cutter having a distal cutting edge, the cutter configured to rotate relative to the elongate catheter body;a cutter drive shaft within the elongate catheter body and connected to the cutter, the cutter drive shaft configured to rotate the cutter, wherein the cutter drive shaft is further configured to be longitudinally displaced proximally or distally to deflect the distal tip to expose the cutting edge of the cutter; andan optical fiber within the cutter drive shaft extending a length of the elongate catheter body, a distal end of the optical fiber forming an imaging sensor;wherein the distal end of the optical fiber is rotationally fixed to the cutter and configured to rotate therewith during imaging, a proximal end of the optical fiber is rotationally fixed to a portion of the catheter device, and an area between the proximal end of the optical fiber and the distal end of the optical fiber is free to rotate within the cutter drive shaft. 2. The device of claim 1, further comprising a ramped slide surface between the distal tip and a region within the elongate body proximal to the cutter, wherein the ramped slide surface is configured to guide deflection of the distal tip as the cutter drive shaft is moved longitudinally. 3. The device of claim 1, wherein the imaging sensor comprises an OCT imaging sensor. 4. The device of claim 1, wherein the cutter comprises a ring cutter. 5. The device of claim 1, wherein the distal tip is configured to collect tissue cut by the cutter. 6. The device of claim 1, further comprising a proximal handle having a first driver for driving rotation of the cutter and a second driver for driving rotation of the imaging sensor. 7. The device of claim 1, further comprising a proximal handle having a first driver for driving rotation of the cutter between 100 and 4000 rpm, and a second driver for driving rotation of the imaging sensor at less than 100 rpm. 8. The device of claim 1, further comprising a proximal handle having a first driver for driving rotation of the cutter in a first direction and a second driver for alternately driving rotation of the imaging sensor in a first rotational direction and a second rotational direction. 9. The device of claim 1, wherein the portion of the catheter device is the handle. 10. The device of claim 1, wherein the cutter drive shaft is configured to be moved longitudinally to pack tissue into the distal tip. 11. The device of claim 1, wherein the area of the optical fiber between the proximal end and the distal end is not fixed to any portion of the device. | Avinger, Inc. | A61B017/22 A61B017/3207 A61B001/313 A61B001/00 A61B005/00
US09326724 | B2 | 2016-05-03 | Devices and methods for monitoring physiological information relating to sleep with an implantable device | 2014-10-28 | Described here are implantable devices and methods for monitoring physiological information relating to sleep. The implantable devices are generally designed to include at least one sensor for sensing physiological information, a processor for processing the physiological information using low computational power to detect a sleep stage, and a battery. The detected sleep stage information may then be used to indicate sleep quality, identify or monitor a medical condition, or guide treatment thereof. | 1. A method of detecting a sleep state or sleep stage by analyzing electrical signals from a patient's brain with an implantable device, the method comprising: sensing electrical signals on a plurality of sensing channels;applying a first sensing channel of the plurality of sensing channels to a first detection channel, the first detection channel configured to determine whether at least one analysis of the electrical signals sensed on the first sensing channel results in an affirmative outcome;applying a second sensing channel of the plurality of sensing channels to a second detection channel, the second detection channel configured to determine whether at least one analysis of the electrical signals sensed on the second sensing channel results in an affirmative outcome; andapplying the result of the at least one analysis of the first detection channel and the result of the at least one analysis of the second detection channel to a logic operation, wherein an output of the logic operation indicates whether a sleep state or sleep stage has been detected by the implantable device. 2. The method of claim 1, wherein the at least one analysis of the electrical signals sensed on the first sensing channel determines the sleep stage or sleep state, and the at least one analysis of the electrical signals sensed on the second sensing channel determines the sleep stage or sleep state. 3. The method of claim 2, wherein the logic operation is such that the output of the logic operation indicates the sleep state or sleep stage has been detected when the result of the at least one analysis of the first detection channel and the result of the at least one analysis of the second detection channel are both affirmative outcomes. 4. The method of claim 2, wherein the logic operation is such that the output of the logic operation indicates the sleep state or sleep stage has been detected when at least one of the results of the at least one analysis of the first detection channel and the at least one analysis of the second detection channel, is an affirmative outcome. 5. The method of claim 1, wherein the at least one analysis of the electrical signals sensed on the first sensing channel determines a noise level below a threshold, and the at least one analysis of the electrical signals sensed on the second sensing channel determines the sleep stage or sleep state. 6. The method of claim 5, wherein the logic operation is such that the output of the logic operation indicates the sleep state or sleep stage as has been detected when the result of the at least one analysis of the first detection channel and the result of the at least one analysis of the second detection channel are both affirmative outcomes. 7. The method of claim 1, wherein the implantable device is further configured so that the result of the at least one analysis of the first detection channel persists for a period of time so that the first detection channel result and the second detection channe... | NeuroPace, Inc. | A61B005/048 A61B005/00 A61B005/0478 A61N001/36 A61B005/01 A61B005/0205 A61B005/026 A61B005/0452 A61B005/1455 A61B005/04 A61B005/0476 A61B005/024 A61B005/11
US09724003 | B2 | 2017-08-08 | Ultra-low power continuous heart rate sensing in wearable devices | 2014-11-14 | Systems and methods may provide for a piezoelectric film that generates an excitation signal in response to pressure variations on a surface of the piezoelectric film and an analog front end coupled to the piezoelectric film, wherein the analog front end generates a first measurement signal based on the excitation signal. Additionally, a heart rate monitor may be coupled to the analog front end, wherein the heart rate monitor generates a heart rate measurement based on the first measurement signal. In one example, the analog front end includes a single stage amplifier. | 1. An apparatus comprising: a piezoelectric film to generate an excitation signal in response to pressure variations on a surface of the piezoelectric film;an analog front end coupled to the piezoelectric film, the analog front end to generate a first measurement signal based on the excitation signal;a heart rate monitor coupled to the analog front end, the heart rate monitor to generate a continuous heart rate measurement based on the first measurement signal;an optical module, implemented at least partly in one or more of configurable logic or fixed functionality logic hardware, coupled to the heart rate monitor; andan activity level monitor coupled to the optical module, the activity level monitor to activate the optical module in response to a physical activity condition being met and deactivate the optical module in response to the physical activity condition not being met, wherein the optical module is to generate a second measurement signal when the optical module is activated, and wherein the heart rate monitor is to generate the continuous heart rate measurement based on the second measurement signal when the second measurement signal is received from the optical module. 2. The apparatus of claim 1, wherein the analog front end includes a single stage amplifier. 3. The apparatus of claim 1, wherein the activity level monitor is to determine whether the physical activity condition is met based on the first measurement signal. 4. The apparatus of claim 1, wherein the activity level monitor is to determine whether the physical activity condition is met based on an accelerometer signal. 5. A wearable system comprising: a wristband including a piezoelectric film with a surface exposed to an interior region of the wristband, the piezoelectric film to generate an excitation signal in response to pressure variations on the surface of the piezoelectric film; andan electrical housing coupled to the wristband, the electrical housing including: an analog front end coupled to the piezoelectric film, the analog front end to generate a first measurement signal based on the excitation signal,a heart rate monitor coupled to the analog front end, the heart rate monitor to generate a heart rate measurement based on the first measurement signal, anda user interface coupled to the heart rate monitor, the user interface to output the heart rate measurement. 6. The wearable system of claim 5, wherein the analog front end includes a single stage amplifier. 7. The wearable system of claim 5, further including: an optical module, implemented at least partly in one or more of configurable logic or fixed functionality logic hardware, coupled to the heart rate monitor; andan activity level monitor coupled to the optical module, the activity level monitor to activate the optical module in response to a physical activity condition being met, wherein the optical module is to generate a second measurement signal when the optical module is activated, and wherein th... | Intel Corporation | A61B005/02 A61B005/024 A61B005/00 A61B005/11
US09936880 | B2 | 2018-04-10 | Camera based photoplethysmogram estimation | 2015-11-25 | A system for estimating a photoplethysmogram waveform of a target includes an image processor configured to obtain images of the target and a waveform analyzer. The waveform analyzer is configured to determine a weight of a portion of the target. The weight is based on a time variation of a light reflectivity of the portion of the target. The time variation of the light reflectivity of the target is based on the images. The waveform analyzer is further configured to estimate a PPG waveform of the target based on the weight of the portion and the time variation of the light reflectivity of the portion. | 1. A method of estimating a photoplethysmogram (PPG) waveform of a target, comprising: obtaining a plurality of images, wherein an image of the plurality of images comprises the target;subdividing the target into a plurality of portions;determining, based on the plurality of images, a time variation of a light reflectivity of a portion of the plurality of portions;determining, based on the time variation of the light reflectivity, a weight associated with the portion of the plurality of portions; andestimating the PPG waveform of the target based on the weight and the image of the plurality of images,wherein estimating the PPG waveform of the target comprises multiplying the weight of the portion of the plurality of portions by the light reflectivity of the portion of the plurality of portions to form a weighted light reflectivity of the portion. 2. The method of claim 1, wherein obtaining the plurality of images comprises: activating an illumination source; andcapturing a first image of the target while the illumination source is activated. 3. The method of claim 2, wherein obtaining the plurality of images further comprises: capturing a second image of the target while the illumination source is activated,wherein the second image is captured at a time that is different than a time of capture of the first image. 4. The method of claim 1, wherein determining the time variation of the light reflectivity of the portion of the plurality of portions comprises: identifying the target in each image of the plurality of images;generating an alignment transformation for each image of the plurality of images based on the location of the target in the image of the plurality of images;applying the alignment transform to each image of the plurality of images to associate at least one pixel of each image of the plurality of images with the target; anddetermining the time variation of the light reflectivity of the portion of the plurality of portions based on the at least one pixel of each image of the plurality of each image and a capture time of each image of the plurality of images. 5. The method of claim 1, wherein determining the weight associated with the portion of the plurality of portions comprises: generating an estimate of a frequency of the PPG waveform based on the light reflectivity of the portion of plurality of portions;determining a first spectral power of the time variation of the light reflectivity of the portion of the plurality of portions in a spectral region including the estimate of the frequency of the PPG waveform;determining a second spectral power of the time variation of the light reflectivity of the portion of the plurality of portions in a spectral region that does not include the estimate of the frequency of the PPG waveform; anddetermining the weight based on the first spectral power and the second spectral power. 6. The method of claim 1, further comprising: generating an estimate of a frequency of the PPG waveform based o... | William Marsh Rice University | A61B005/00 A61B005/024 A61B005/11 G06K009/00 G06K009/20 G06K009/46
US09872626 | B2 | 2018-01-23 | Printed circuit board assembly with ferrite for medical device | 2015-11-06 | A printed circuit board assembly includes a printed circuit board and a ferrite. The printed circuit board has a substantially planar top surface and a substantially planar bottom surface. The ferrite has an annular cross section that defines a channel. The ferrite is positioned on the printed circuit board such that a portion of the plurality of traces passes through the channel. | 1. A printed circuit board assembly, comprising: a printed circuit board, comprising: a plurality of traces;two slots defining a tongue portion of the printed circuit board;a substantially planar top surface; anda substantially planar bottom surface; anda ferrite having an annular cross section defining a channel, wherein the ferrite extends through the two slots and is positioned on and surrounds at least part of the tongue portion of the printed circuit board, such that a portion of the plurality of traces passes through the channel of the ferrite. 2. The printed circuit board assembly of claim 1, wherein the annular cross section is formed by two separate pieces that are joined together. 3. The printed circuit board assembly of claim 1, wherein the annular cross section is substantially rectangular. 4. A medical device, comprising: a housing; anda printed circuit board assembly supported by the housing, the printed circuit board assembly comprising: a printed circuit board including a plurality of traces and two slots defining a tongue portion, wherein the printed circuit board has a substantially planar top surface and a substantially planar bottom surface; andwherein at least some of the plurality of traces carry sensitive signals;a ferrite extending through the two slots and surrounding at least part of the tongue portion, the ferrite having an annular cross section, the annular cross section defining a channel, wherein the ferrite is positioned on the tongue portion such that a portion of the plurality of traces passes through the channel. 5. The printed circuit board assembly of claim 1, further comprising a connector assembly positioned on a distal end of the tongue portion,wherein the ferrite is positioned on the distal end of the tongue portion. 6. The printed circuit board assembly of claim 5, wherein at least some of the plurality of traces carry a sensitive signal, and wherein, when the plurality of traces pass through the channel, the plurality of traces are positioned between the substantially planar top surface and the substantially planar bottom surface of the printed circuit board. 7. The printed circuit board assembly of claim 6, wherein the sensitive signal includes vital signs data. 8. The printed circuit board assembly of claim 7, wherein the connector assembly is configured to receive vital signs data from an external device. 9. The printed circuit board assembly of claim 1, further comprising a plurality of mounting elements configured to mate with a plurality of complementary mounting elements positioned on a medical device main assembly. 10. A medical device, comprising: a printed circuit board assembly in communication, the printed circuit board assembly comprising: a printed circuit board including a plurality of traces, a tongue portion defined by two slots in the printed circuit board, a substantially planar top surface and a substantially planar bottom surface; anda ferrite extending through the two slots ... | WELCH ALLYN, INC. | A61B005/022 A61B005/021 H05K001/02 A61B005/0205 A61B005/00 H02J001/00 H02J007/00 G02F001/1333 H05K001/18 A61B005/0235 A61B050/13
US10881340 | B2 | 2021-01-05 | Medical device inserters and processes of inserting and using medical devices | 2016-01-15 | An apparatus for insertion of a medical device in the skin of a subject is provided, as well as methods of inserting medical devices. | 1. A method of inserting an analyte sensor into a subject, the method comprising: positioning a distal surface of an insertion assembly on skin of the subject such that an adhesive layer adheres to the skin, the insertion assembly comprising a spring in a loaded position and an interior; andpressing a button of the insertion assembly in a first lateral direction toward the interior of the insertion assembly, wherein pressing the button releases the spring from the loaded position such that the spring moves a sliding member lengthwise along a length of a track in the insertion assembly, wherein the length of the track is oriented in a second lateral direction different from the first lateral direction,wherein movement of the sliding member along the length of the track causes an actuator to advance in a distal longitudinal direction, wherein advancement of the actuator causes advancement of a sharp and the analyte sensor in the distal longitudinal direction,wherein the sliding member applies a force on the actuator in a proximal longitudinal direction after the sharp and the analyte sensor have been advanced in the distal longitudinal direction, andwherein the distal longitudinal direction and the proximal longitudinal direction are defined along an insertion axis. 2. The method of claim 1, wherein the spring is a first spring, the method further comprising retracting the sharp with a second spring of the insertion assembly. 3. The method of claim 2, wherein the length of the track is a first length of the track, and wherein further movement of the sliding member along a second length of the track causes retraction of the sharp with the second spring, wherein the second length of the track is oriented in the second lateral direction. 4. The method of claim 3, wherein the sliding member is connected to a rotatable member such that the sliding member remains in contact with the rotatable member after the movement of the sliding member along the first and the second lengths of the track. 5. The method of claim 2, wherein the second spring is a compression spring. 6. The method of claim 1, wherein the second lateral direction is perpendicular to the insertion axis. 7. The method of claim 1, wherein the sliding member is restrained to motion along the track. 8. The method of claim 1, wherein the analyte sensor is a glucose sensor. 9. The method of claim 1, wherein at least a portion of the analyte sensor is housed within the sharp. 10. The method of claim 1, wherein the insertion assembly comprises an on-body device. 11. The method of claim 1, further comprising installing on-body electronics, without use of the insertion assembly, after the advancement of the analyte sensor into the subject. 12. The method of claim 11, further comprising wirelessly communicating, by the on-body electronics, analyte related data. 13. The method of claim 1, wherein the first lateral direction is perpendicular to the second lateral direction. 14. An insertio... | ABBOTT DIABETES CARE INC. | A61B017/32 A61B005/151 A61B005/15 A61B005/145 A61B005/00 A61M005/158 A61B005/157 A61B017/34 A61B005/01
US20160143597 | A1 | 2016-05-26 | PROCESSING DEVICE | 2014-11-24 | A processing device including an interface unit, a first slot, and a transmission unit is provided. The interface unit is configured to receive a plurality of physiological monitoring signals. The first slot is configured to hold a first analysis module. When the first analysis module is inserted into the first slot, the first analysis module analyzes a first signal group to generate a first analysis result. The first signal group includes at least two of the physiological monitoring signals. The transmission unit outputs the first analysis result to at least one external device. | 1. A processing device, comprising: an interface unit configured to receive a plurality of physiological monitoring signals;a first slot configured to hold a first analysis module, wherein when the first analysis module is inserted into the first slot, the first analysis module analyzes a first signal group to generate a first analysis result, and the first signal group comprises at least two of the physiological monitoring signals; anda transmission unit outputting the first analysis result to at least one external device. 2. The processing device as claimed in claim 1, wherein when a second analysis module is inserted into the first slot, the second analysis module analyzes a second signal group to generate a second analysis result different from the first analysis result, and the second signal group comprises at least two of the physiological monitoring signals. 3. The processing device as claimed in claim 2, wherein at least one physiological monitoring signal among the first signal group is the same as at least one physiological monitoring signal among the second signal group. 4. The processing device as claimed in claim 1, further comprising: a second slot configured to hold a second analysis module, wherein when the second analysis module is inserted into the second slot, the second analysis module analyzes a second signal group to generate a second analysis result, and the second signal group comprises at least two of the physiological monitoring signals. 5. The processing device as claimed in claim 4, wherein at least one physiological monitoring signal among the first signal group is the same as at least one physiological monitoring signal among the second signal group. 6. The processing device as claimed in claim 4, wherein the transmission unit outputs the first analysis result to a first external device and outputs the second analysis result to a second external device. 7. The processing device as claimed in claim 4, wherein the transmission unit outputs the first and second analysis results to the external device. 8. The processing device as claimed in claim 1, wherein the external device displays the first analysis result as text or an image. 9. The processing device as claimed in claim 1, wherein the first analysis module comprises a analysis unit to analyzes the first signal group to generate an output signal. 10. The processing device as claimed in claim 9, wherein the transmission unit serves the output signal as the first analysis result and outputs the output signal to the external device, and the external device further comprises a report unit collects the output signal to generate a report result. 11. The processing device as claimed in claim 9, wherein the external device is a storage device storing the report result. 12. The processing device as claimed in claim 9, wherein the external device is a display device displaying the report result as text or an image. 13. The processing device as claimed in claim 9, ... |  | A61B005/00 A61B005/0488 A61B007/04 A61B005/0402 A61B005/0205 A61B005/0476
US20160135748 | A1 | 2016-05-19 | Inflatable Electroencephalogram Measuring Device | 2015-10-28 | An inflatable electroencephalogram measuring device comprises an inflatable helmet applicable for an user's head and at least one measuring electrode. A plurality of supporting strips are disposed within the inflatable helmet and the at least one measuring electrode is well distributed upon the supporting strips. When the inflatable helmet is inflated, the measuring electrode accordingly moves to make contact with the use's head so as to measure an electroencephalogram signal. By designing the inflatable electroencephalogram measuring device with an inflatable helmet, the proposed invention is widely utilized for measuring various users with a variety of head sizes, and thus is characterized with a potential for becoming an important tool of medical measurement. | 1. An inflatable electroencephalogram measuring device for measuring a testee's EEG signals, comprising: an inflatable helmet for said testee to wear over his or her head, comprising an inflatable body and a plurality of supporting strips which are disposed on an inner side of said inflatable body; andat least one measuring electrode disposed on said plurality of supporting strips, such that when said inflatable helmet is inflated, said measuring electrode is induced to make contact with said head to measure said testee's EEG signals. 2. The inflatable electroencephalogram measuring device of claim 1, wherein said inflatable helmet further comprises an inlet, which is disposed on one side of said inflatable body, and air can be blown into said inflatable body through said inlet for inflating said inflatable body. 3. The inflatable electroencephalogram measuring device of claim 1, further comprising a plurality of wires and a printed circuit board, wherein said plurality of wires are connected between said measuring electrode and said printed circuit board, and said testee's EEG signals are transmitted to said printed circuit board through said plurality of wires for further analyses. 4. The inflatable electroencephalogram measuring device of claim 1, wherein said measuring electrode is brush-electrode, dry electrode or wet electrode. 5. The inflatable electroencephalogram measuring device of claim 1, wherein each of said plurality of supporting strips is bendable so as to absorb stress generated by said moving measuring electrode. 6. The inflatable electroencephalogram measuring device of claim 1, wherein each of said plurality of supporting strips is made of flexible material. 7. The inflatable electroencephalogram measuring device of claim 1, further comprising a housing, and said inflatable helmet being disposed in an inner space of said housing for said testee to wear so as to measure said EEG signals. 8. The inflatable electroencephalogram measuring device of claim 7, wherein said housing further comprises an upper cover and a lower hoop, said upper cover connects said lower hoop and extends upwardly from said lower hoop, said lower hoop surrounds said testee's head, and said inflatable helmet is disposed on an inner side of said upper cover. 9. The inflatable electroencephalogram measuring device of claim 8, wherein said upper cover further comprises at least two fixing strips which make across to each other, and said fixing strips are commonly connected to said lower hoop. 10. The inflatable electroencephalogram measuring device of claim 8, wherein said lower hoop further comprises an adjustment element for controlling said housing tighter or looser. 11. The inflatable electroencephalogram measuring device of claim 1, wherein flow inside said inflatable body is designed to be asymmetric. 12. The inflatable electroencephalogram measuring device of claim 1, further comprising a pressure measuring device, which is designed for cont... |  | A61B005/00 A61B005/0478
US20160120407 | A1 | 2016-05-05 | MICROWAVE NEARFIELD RADAR IMAGING (NRI) USING DIGITAL BREAST TOMOSYNTHESIS (DBT) FOR NON-INVASIVE BREAST CANCER DETECTION | 2014-06-18 | In some aspects, the disclosure is directed methods and systems for granular imaging of a distribution of tissues. A tomographic device may acquire a first image of a distribution of tissues, the first image including a plurality of pixels. A tomographic image processor may translate at least some of the plurality of pixels of the first image into a plurality of values representative of a distribution of dielectric constants corresponding to the distribution of tissues. A nearfield radar imaging (NRI) system may generate a second image of the distribution of tissues based on the plurality of values. | 1. A method for granular imaging of a distribution of tissues, the method comprising: (a) acquiring, by a tomographic device, a first image of a distribution of tissues, the first image including a plurality of pixels;(b) translating at least some of the plurality of pixels of the first image into a plurality of values representative of a distribution of dielectric constants corresponding to the distribution of tissues; and(c) generating, by a nearfield radar imaging (NRI) system, a second image of the distribution of tissues based on the plurality of values. 2. The method of claim 1, wherein (a) comprises acquiring the first image of the distribution of tissues, the distribution of tissues comprising breast tissues or some other portion of a human body. 3. The method of claim 1, further comprising identifying cancerous tissue or other feature from the second image. 4. The method of claim 1, further comprising electromagnetically scanning the distribution of tissues, by the NRI system, within a predetermined period of time from the acquisition of the first image. 5. The method of claim 1, further comprising maintaining at least one of: a shape or a size of the distribution of tissues, across time periods during which the first image is acquired and during which the distribution of tissues is electromagnetically scanned by the NRI system. 6. The method of claim 1, further comprising locating the distribution of tissues within a container containing a bolus matching liquid. 7. The method of claim 1, wherein (b) comprises determining fat content corresponding to a first pixel of the plurality of pixels. 8. The method of claim 7, further comprising generating a first value of the plurality of values based on the determined fat content corresponding to the first pixel. 9. The method of claim 1, wherein (c) comprises generating a non-uniform function associated with at least one antenna of the NRI system. 10. The method of claim 9, further comprising generating the non-uniform function based on a finite difference frequency domain (FDFD) method or other full-wave electromagnetic method. 11. A system for granular imaging of a distribution of tissues, the system comprising: a tomographic device, the tomographic device configured to acquire a first image of a distribution of tissues, the first image including a plurality of pixels, wherein at least some of the plurality of pixels of the first image are translated into a plurality of values representative of a distribution of dielectric constants corresponding to the distribution of tissues; anda nearfield radar imaging (NRI) system, the NRI system configured to generate a second image of the distribution of tissues based on the plurality of values. 12. The system of claim 11, wherein the tomographic device is configured to acquire the first image of the distribution of tissues, the distribution of tissues comprising breast tissues or some other portion of a human body. 13. The system of claim ... |  | A61B005/00 A61B006/02 A61B006/00 A61B005/05
US20160128613 | A1 | 2016-05-12 | BODY INFORMATION SENSING DEVICE | 2015-11-04 | A body information sensing device for obtaining body information of a user. The body information sensing device includes a housing, a sensing module and a processing module. The housing is detachably disposed at the outer surface of a daily necessity. The sensing module is disposed at the outer surface of the housing, wherein the sensing module is for sensing the body information of the user. The processing module is disposed at the housing, and the sensing module is electrically connected with the processing module. | 1. A body information sensing device for obtaining body information of a user, the body information sensing device comprising: a housing detachably disposed at the outer surface of a daily necessity;a sensing module disposed at the outer surface of the housing, wherein the sensing module is for getting close to or contacting with the body of the user to sense the body information of the user;a processing module disposed at the housing, the sensing module being electrically connected with the processing module; anda message transceiver module disposed at the housing, the message transceiver module being electrically connected with the processing module,wherein the processing module controls the message transceiver module to send the body information sensed by the sensing module to a mobile electronic device. 2. The body information sensing device according to claim 1, wherein the housing is made of elastically-retractable material. 3. The body information sensing device according to claim 2, wherein the elastically-retractable material comprises a steel tape structure, a spring steel or an elastic metal bar. 4. The body information sensing device according to claim 2, wherein the sensing module comprises a body temperature sensing unit, a blood oxygen sensing unit or a heartbeat sensing unit. 5. The body information sensing device according to claim 4, wherein the heartbeat sensing unit is for being contacted by a thumb of the user, the blood oxygen sensing unit if for being contacted by an index finger of the user. 6. The body information sensing device according to claim 5, wherein the body temperature sensing unit is for being contacted by a palm of the user. 7. The body information sensing device according to claim 6, further comprising a light-emitting module, the light-emitting module being disposed at the housing, the light-emitting module being electrically connected with the processing module, the light-emitting module including a shield-sensing unit. 8. The body information sensing device according to claim 7, further comprising a vibration unit, the vibration unit being disposed at the housing, the vibration unit being electrically connected with the processing module. 9. The body information sensing device according to claim 8, further comprising an acceleration sensing unit, the acceleration sensing unit being electrically connected with the processing module. 10. The body information sensing device according to claim 9, further comprising a projection module, the projection module being electrically connected with the processing module, the processing module controlling the projection module to project information sensed by the sensing module. 11. The body information sensing device according to claim 1, further comprising a light-emitting display unit, the light-emitting display unit being disposed at the outer surface of the housing, the light-emitting display unit being for prompting a message. |  | A61B005/1455 A61B005/024 A61B005/00 A61B005/01
US09339412 | B2 | 2016-05-17 | Maintenance of core mammalian body temperature | 2011-11-29 | The present application relates to methods for maintaining core body temperature in a subject. Heat is applied to peripheral thermoregulatory control tissue of the subject to increase or maintain perfusion of blood in glabrous tissue of the subject, and a warming stimulus is applied to the glabrous tissue. The application of heat to the peripheral thermoregulatory control tissue occurs discretely and simultaneously with application of the warming stimulus to the glabrous tissue, without application of a heating device or a warming device to tissue other than the tissue subjected to the applied heat or to the applied warming stimulus, thereby maintaining the core body temperature of the subject. | 1. A method for maintaining the core body temperature of a subject, comprising: applying heat from a heating device, wherein the heating device is positioned over the cervical spinal region of the subject, the lumbar spinal region of the subject, or both, such that the heating device applies heat localized to the cervical spinal region of the subject, the lumbar spinal region of the subject, or both, to cause heating of peripheral thermoregulatory control tissue located in the cervical spinal region of the subject, the lumbar spinal region of the subject, or both, wherein the applied localized heat increases or maintains perfusion of blood in glabrous tissue of the subject; andapplying a warming stimulus from a warming device to the glabrous tissue;wherein the heat is applied to the peripheral thermoregulatory control tissue located in the cervical spinal region of the subject, the lumbar spinal region of the subject, or both, simultaneously with application of the warming stimulus to the glabrous tissue, thereby maintaining the core body temperature of the subject. 2. The method according to claim 1, wherein the glabrous tissue is located in a palmar region of the subject. 3. The method according to claim 1, wherein the glabrous tissue is located in a plantar region of the subject. 4. The method according to claim 1, wherein the subject is subjected to a hypothermia inducing event or activity. 5. The method according to claim 1, further comprising assessing the temperature of the glabrous tissue, the peripheral thermoregulatory control tissue, or the core body temperature of the subject. 6. The method according to claim 1, wherein applying the warming stimulus to a portion of skin of the subject concurrently applies heat to the glabrous tissue. 7. The method according to claim 1, wherein applying heat to the peripheral thermoregulatory control tissue comprises applying heat localized to tissue overlying or proximate to the peripheral thermoregulatory control tissue within the cervical spinal region of the subject, the lumbar spinal region, or both of the subject. 8. The method according to claim 7, wherein the heating device is positioned on the skin to apply the heat localized to tissue overlying or proximate to the peripheral thermoregulatory control tissue within the cervical spinal region of the subject, the lumbar spinal region, or both of the subject. 9. The method according to claim 1, wherein applying warming to the glabrous tissue comprises applying warming to tissue overlying or proximate to the glabrous tissue, such that by applying warming to said overlying tissue or said proximate tissue, warmth is applied to the glabrous tissue. | Board of Regents, The University of Texas System | A61F007/00 A61B005/01 A61F007/02 A61B005/00
US09406211 | B2 | 2016-08-02 | Wearable posture regulation system and method to regulate posture | 2015-06-29 | A wearable posture monitoring and feedback system and related method configured to monitor and provide feedback regarding a user's posture while operating and viewing a portable electronic device includes: a wearable frame configured to be fit on or about a head of the user; a sensor mounted on the wearable frame configured to monitor at least a head position of the user; and a transmitter configured to transmit data related to at least the head position wirelessly from the sensor to a receiver on or in the portable electronic device. The system is configured to (1) provide feedback to the user based on the data from the sensor; and (2) send a signal to alter an output provided to the user by the portable electronic device based on the data from the sensor. | 1. A wearable posture monitoring system configured to monitor and provide feedback regarding a user's posture while operating and viewing a portable electronic device, the system comprising: a wearable component configured to be fit on or about a head of the user;a sensor mounted on the wearable component configured to monitor posture of the user based on at least a head position or head orientation of the user; anda transmitter configured to transmit data related to posture wirelessly from the sensor to a receiver on or in the portable electronic device;wherein:the data transmitted by the transmitter to the portable electronic device is configured for use as input by a computer executable application program to:provide progressive negative feedback regarding posture during use of the portable electronic device to the user based on data from the sensor, by generating a first signal that is configured to warn the user of abnormal or poor posture, and thereafter if the posture of the user fails to be corrected, by generating a second signal configured to alter the output provided to the user by the portable electronic device,provide positive reinforcement to the user via the portable electronic device if the user's posture is acceptable, based on the data transmitted from the transmitter, andgenerate for display a screen listing available programs, and identify available programs which require use of the wearable component to be executed. 2. The wearable posture monitoring system of claim 1, wherein the system is configured to continuously monitor the user's reaction to the feedback supplied to the user during use of the portable electronic device. 3. The wearable posture monitoring system of claim 1, wherein the feedback is provided in the form of auditory, visual, or tactile cues or the feedback is provided by causing the portable electronic device to display a warning message. 4. The wearable posture monitoring system of claim 1, wherein the sensor comprises a tilt sensor configured to monitor the gaze angle of the user during use of the portable electronic device. 5. The wearable posture monitoring system of claim 1, wherein the system is configured to allow the portable electronic device to operate for time in excess of a finite, predetermined amount of time in response to observed good posture. 6. The wearable posture monitoring system of claim 5, wherein the system is configured to prevent the portable electronic device from operating for the entirety of the finite, predetermined amount of time in response to observed poor posture. 7. The wearable posture monitoring system of claim 1, wherein the portable electronic device is a hand-held portable electronic device, and wherein the system is configured to monitor posture based on the head position or the head orientation of the user during use of the hand-held portable electronic device. | Medical Wearable Solutions Ltd. | G08B023/00 G08B021/04 A61B005/103
US09339189 | B2 | 2016-05-17 | Method and device for transmitting sensor data of an implantable sensor to an external data processing unit | 2012-07-20 | During the transmission of sensor data of an implantable sensor with a measurement data recorder, an energy storage unit and a transmitter unit to an external data processing unit with a transmitter/receiver unit, a high-frequency energy carrier signal is firstly emitted from an energy supply unit. At least a part of the energy contained in the high-frequency energy carrier signal is stored in the energy storage unit of the implantable sensor. After the ending of the emission of the high-frequency energy carrier signal, a sensor measurement is carried out using the measurement data recorder of the implantable sensor. The measured sensor data are transmitted from the transmitter unit of the implantable sensor to the transmitter/receiver unit of the external data processing unit. The result is a transmission method, in which interference from the HF feed is avoided, with a compact design. | 1. A method for transmitting sensor data of an implantable sensor (2) to an external data processing unit (3) having the following steps: emitting a high-frequency energy carrier signal from an energy supply unit (18),storing at least a part of the energy contained in the high-frequency energy carrier signal in an energy storage unit (8) of the implantable sensor (2),ending the emission of the high-frequency energy carrier signal, wherein said ending the emission depends on a transmission spacing between the implantable sensor and the external data processing unit,after the ending of the emission of the high-frequency energy carrier signal, carrying out at least one sensor measurement using a measurement data recorder (4) of the implantable sensor (2),transmitting the measured sensor data from a transmitter unit (13) of the implantable sensor (2) to a transmitter/receiver unit (17) of the external data processing unit (3). 2. The method according to claim 1, comprising a monitoring of an energy supply of the measurement data recorder (4) by the energy storage unit (8) using a monitoring unit (16) of the implanted sensor (2). 3. The method according to claim 1, comprising a conversion of recorded analog sensor data to digital sensor data to be transmitted before the transmission. 4. The method according to claim 1, wherein after the sensor measurements have been carried out and before the measured sensor data have been transmitted, the emission of the high-frequency energy carrier signal is resumed. 5. A device (1) for transmitting sensor data, comprising: an implantable sensor (2) with a measurement data recorder (4), an energy storage unit (8) and a transmitter unit (13),an external data processor unit (3) comprising a transmitter/receiver unit (17), a control unit (19) to control a transmission method and to control starting and ending of the transmission of a high frequency signal between the transmitter/receiver unit (17) and the transmitter unit (13), the ending of the transmission being dependent on a transmission spacing between the implantable sensor and the external data processing unit (3), and an energy supply unit (18). 6. The device according to claim 5, wherein the energy storage unit (8) is configured as a capacitor. 7. The device according to claim 5, wherein the implanted sensor (2) has an A/D converter (11) to convert recorded analog sensor data into digital sensor data to be transmitted. 8. The device according to claim 5, wherein the implanted sensor (2) has a monitoring unit (16) to monitor an energy supply of the measurement data recorder (4) by the energy storage unit (8). 9. The device according to claim 5, comprising a start unit (36, 37) for starting an HF field-free measurement. 10. The device according to claim 9, wherein, the start unit is configured as a voltage sensor (36). 11. The device according to claim 9, wherein the start unit is configured as a time sensor (37). 12. A method for transmitting sensor... | RAUMEDIC AG | G08C019/04 G08B021/00 G08B023/00 G08C019/06 G01P015/00 A61B005/00 A61B001/00 A61B005/02 A61B005/103 A61B005/07 A61B005/01 A61B005/03 A61B005/1455
US09332886 | B2 | 2016-05-10 | Fiducial placement system and splayed stylet | 2013-01-28 | A fiducial deployment system may include one or more of a needle including a dimpled retention structure configured to releasably retain a plurality of fiducials, each of which includes a central aperture (which may be embodied as a groove or similar opening); a retention/deployment stylet that traverses through each fiducial's central aperture and releasably engages a distalmost fiducial of the plurality of fiducials with a splayed distal stylet end; a pusher cannula member; and echogenic dimpling and/or other echogenic-enhancing features on at least a distal portion of the needle. | 1. A fiducial deployment system comprising: a needle includinga generally tubular cannula body defining a needle lumen disposed through at least a lengthwise portion of the cannula body anda distal needle end region, the distal end region comprising a distal needle end opening at a distal end of the needle lumen;at least one fiducial comprising a fiducial body slidably disposed in the needle lumen, which fiducial body includes a fiducial aperture extending longitudinally therethrough;a pusher slidably disposed longitudinally through the needle lumen proximal of and contacting the at least one fiducial, where the pusher is configured to advance the at least one fiducial out of the distal needle end opening; anda stylet including at least one splayed distal end portion, the stylet extending longitudinally through a portion of the needle lumen and through the aperture of the at least one fiducial,wherein the at least one splayed distal end portion is biased apart and engaging a distal end surface of the at least one fiducial to releasably retain the at least one fiducial, and the at least one splayed distal end portion is sufficiently collapsible to pass proximally through a cross-sectional area of the fiducial aperture to release the at least one retained fiducial. 2. The fiducial deployment system of claim 1, where the pusher is embodied as a pusher cannula including a pusher lumen longitudinally disposed therethrough, through which a portion of the stylet slidably extends. 3. The fiducial deployment system of claim 2, where the pusher cannula includes a tubular length encompassing the pusher lumen. 4. The fiducial deployment system of claim 1, where the at least one fiducial comprises a generally cylindrical portion. 5. The fiducial deployment system of claim 1, where the needle comprises at least one echogenically-enhanced region. 6. The fiducial deployment system of claim 1, further comprising at least one protrusion configured as a retaining detent extending radially into the needle lumen to form a restricted inner diameter that is less than an inner diameter of a major length of the needle lumen, where the at least one needle protrusion is provided by a dimple in the tubular cannula body. 7. The fiducial deployment system of claim 6, where the stylet at least one splayed distal end portion is configured to engage a proximal surface of the at least one fiducial and pushingly advance the at least one fiducial distally through the needle lumen past the at least one protrusion. 8. The fiducial deployment system of claim 1, where the at least one fiducial comprises a plurality of fiducials, each including a proximal fiducial end and a distal fiducial end, where the proximal end of at least one of the fiducials is immediately adjacent the distal end of another of the fiducials. 9. The fiducial deployment system of claim 1, where the at least one splayed portion comprises a generally transverse split in a distal end of the stylet. 10. The f... | Cook Medical Technologies LLC | A61B017/00 A47L011/14 A61B001/018 A61B001/24 A61B001/273 A61B001/303 A61B001/31 A61B001/313 A61B005/00 A61B005/06 A61B006/12 A61B006/00 A61B008/08 A61N005/10 A47L011/38 B08B001/04 A46B013/00
US09345447 | B2 | 2016-05-24 | Right atrium indicator | 2013-03-14 | A method for determining the location of a medical device within a body is provided. The method includes transmitting from the medical device an acoustic signal; receiving with the medical device a reflected acoustic signal; advancing the medical device based on a first algorithm, the first algorithm including a first weighting factor and a first feature extracted from the reflected acoustic signal; determining a first location of the medical device based on the first algorithm; and moving the medical device to a second location based on a second algorithm, the second algorithm based on the determined first location and including at least one of a second weighting factor and a second feature extracted from the reflected acoustic signal. Also disclosed are systems and devices for performing the methods described herein. | 1. A method for determining that a distal region of a medical device has entered the right atrium of a patient, the method comprising: transmitting, from a non-imaging ultrasound transducer of the medical device, an acoustic signal;receiving, using the non-imaging ultrasound transducer of the medical device, a reflected acoustic signal;detecting, using an electrocardiogram sensor of the medical device, an electrocardiogram signal;extracting one or more acoustic features from the reflected acoustic signal;extracting one or more electrocardiogram features from the electrocardiogram signal;determining that the distal region of the medical device has entered the right atrium of the patient based on the one or more acoustic features and the one or more electrocardiogram features; andvisually indicating, in response to the determination that the distal region of the medical device has entered the right atrium of the patient, that the distal region of the medical device has entered the right atrium of the patient. 2. The method of claim 1, wherein the one or more acoustic features comprise a frequency content of the reflected acoustic signal over time. 3. The method of claim 2, wherein determining that the distal region of the medical device has entered the right atrium of the patient based on the one or more acoustic features and the one or more electrocardiogram features comprises determining that the distal region of the medical device has entered the right atrium of the patient when the frequency content of the acoustic signal over time changes between antegrade dominant and retrograde dominant while the medical device is not being advanced or retracted. 4. The method of claim 3, wherein the frequency content of the acoustic signal over time changes between antegrade dominant and retrograde dominant at a frequency of 7 kHz to 15 kHz. 5. The method of claim 1, wherein the one or more acoustic features comprise an antegrade power signal, a retrograde power signal, a ratio of the antegrade power signal to the retrograde power signal, or an overall power of the reflected acoustic signal. 6. The method of claim 5, wherein determining that the distal region of the medical device has entered the right atrium of the patient based on the one or more acoustic features and the one or more electrocardiogram features comprises determining that the distal region of the medical device has entered the right atrium of the patient when an average value of the antegrade power signal is equal to the retrograde power signal over a cardiac cycle. 7. The method of claim 1, further comprising: generating a Mel-frequency cepstrum of the reflected acoustic signal; anddetermining one or more coefficients of the Mel-frequency cepstrum. 8. The method of claim 7, further comprising comparing the determined one or more coefficients with a database of coefficients based on turbulent and non-turbulent blood flow. 9. The method of claim 1, further comprising monitoring respi... | VasoNova, Inc. | A61B008/08 A61B005/06 A61B008/12 A61B005/04 A61B005/0456 A61B008/06
US09330459 | B2 | 2016-05-03 | Thermometer having a digital infrared sensor on a circuit board that is separate from a microprocessor | 2014-08-11 | A non-touch thermometer that includes a first circuit board that has a microprocessor, a battery operably coupled to the microprocessor and a single button operably coupled to the microprocessor, the non-touch thermometer also includes a camera operably coupled to the microprocessor that provides at least two images to the microprocessor, the non-touch thermometer also includes a second circuit board that has a digital infrared sensor operably coupled to the microprocessor with no analog-to-digital converter operably coupled between the digital infrared sensor and the microprocessor, the digital infrared sensor having only digital readout ports, the digital infrared sensor having no analog sensor readout ports, the microprocessor is operable to receive from the digital readout ports a digital signal that is representative of an infrared signal detected by the digital infrared sensor and the microprocessor is operable to determine a temperature from the digital signal that is representative of the infrared signal. | 1. A non-touch thermometer to measure temperature, the non-touch thermometer comprising: a first circuit board having a microprocessor;a battery operably coupled to the microprocessor;a single button operably coupled to the microprocessor;a camera operably coupled to the microprocessor and providing at least two images to the microprocessor;a second circuit board having a digital infrared sensor operably coupled to the microprocessor with no analog-to-digital converter operably coupled between the digital infrared sensor and the microprocessor, the digital infrared sensor having only digital readout ports, the digital infrared sensor having no analog sensor readout ports; anda display device operably coupled to the microprocessor,wherein the microprocessor is operable to receive from the digital readout ports a digital signal that is representative of an infrared signal detected by the digital infrared sensor and the microprocessor is operable to determine the temperature from the digital signal that is representative of the infrared signal. 2. The non-touch thermometer of claim 1, wherein the display device further comprises: a green traffic light operable to indicate that the temperature is good;an amber traffic light operable to indicate that the temperature is low; anda red traffic light operable to indicate that the temperature is high. 3. The non-touch thermometer of claim 1 further comprising: the microprocessor including a pixel-examination-module configured to examine pixel values of the at least two images, a temporal-variation module to determine temporal variation of the pixel values between the at least two image, a signal processing module configured to amplify the temporal variation resulting in amplified temporal variation, and a visualizer to visualize a pattern of flow of blood in the amplified temporal variation in the at least two images. 4. The non-touch thermometer of claim 3, wherein the signal processing module is further configured to amplify variations of the pixel values between the at least two images. 5. The non-touch thermometer of claim 3, wherein the signal processing module is further configured to perform temporal processing. 6. The non-touch thermometer of claim 5, wherein the temporal processing is a bandpass filter. 7. The non-touch thermometer of claim 6, wherein the bandpass filter is configured to analyze frequencies over time. 8. The non-touch thermometer of claim 3, wherein applying signal processing includes spatial processing. 9. The non-touch thermometer of claim 8, wherein the spatial processing removes noise. 10. The non-touch thermometer of claim 1, further comprising: the first circuit board including: a first digital interface operably coupled to the microprocessor;the second circuit board including: a second digital interface, the second digital interface being operably coupled to the first digital interface; andthe digital infrared sensor operably coupled to the second digital interface... | Arc Devices Limited | G06K009/00 G01J005/00 G06T007/00 A61B005/026 G01J005/32 G01J005/02 G01J005/08 A61B005/00 A61B005/01 G06T005/20 A61B005/0205 A61B005/0402 A61B005/1455 G06K009/46 G06K009/62 G06T007/20 G06T003/40 G06T005/10 G06F019/00 G01J005/30 G06T007/40 H04N005/33 G01K001/02 G01K013/00 G06F003/042 H04N005/378 G06T005/00 A61B005/021 A61B005/024 A61B005/08
US10660699 | B2 | 2020-05-26 | Ablation therapy to disrupt coordinated bladder contractions | 2015-10-30 | The disclosure describes devices, systems, and techniques for identifying and treating bladder dysfunction. In one example, a method includes identifying one or more focal points at respective locations of bladder tissue of a bladder of a patient, the one or more focal points initiating coordinated contractions of a detrusor muscle. The method may also, or alternatively, include ablating, for each of the one or more focal points, a respective portion of the bladder tissue at the respective location of the focal point. Ablation of these targeted portions of the bladder tissue may reduce the coordinated contractions of the detrusor muscle and alleviate overactive bladder symptoms. | 1. A method for treating bladder dysfunction in a patient, the method comprising: identifying one or more focal points at respective locations of bladder tissue of a bladder of the patient, the one or more focal points initiating coordinated contractions of a detrusor muscle in the bladder;injecting a substance into at least one respective portion of the bladder tissue at the respective locations of the one or more focal points, the substance configured to temporarily disable initiation of contractions;confirming that the substance reduces coordinated contractions of the detrusor muscle; andbased on the confirmation, ablating, for each of the one or more focal points, the respective portion of the bladder tissue at the respective location of the focal point. 2. The method of claim 1, wherein identifying the one or more focal points comprises: imaging propagation of contractions in the detrusor muscle; andmapping, based on the imaged propagation of contractions, the one or more focal points at the respective locations of the bladder tissue. 3. The method of claim 2, wherein imaging the propagation of contractions comprises at least one of optically determining the propagation of contractions, obtaining electrogram information for the propagation of contractions, optically determining calcium movement for the propagation of contractions, magnetic resonance imaging the propagation of contractions, or ultrasound imaging of the propagation of contractions. 4. The method of claim 1, further comprising inserting an ablation instrument into the bladder via a urethra, wherein ablating the respective portions of the bladder tissue comprises directing the ablation to an inner surface of the bladder via the ablation instrument located within the bladder. 5. The method of claim 1, wherein ablating a respective portion of the bladder tissue comprises ablating a portion of at least one of a urothelium or the detrusor muscle. 6. The method of claim 1, wherein ablating a respective portion of the bladder tissue comprises ablating a portion of a nerve residing external to the detrusor muscle, wherein the nerve controls at least a portion of the detrusor muscle associated with at least one of the one or more focal points. 7. The method of claim 1, wherein ablating the respective portions of the bladder tissue comprises ablating the respective portions of the detrusor muscle associated with the focal points and not ablating portions associated with the non-focal points of the detrusor muscle. 8. The method of claim 1, wherein ablating the respective portions of the bladder tissue comprises ablating the respective portions of the bladder tissue via at least one of chemical ablation, radio frequency ablation, ultrasound ablation, mechanical ablation, heat ablation, or cyroablation. 9. The method of claim 1, wherein the one or more focal points are a subset of a plurality of focal points that initiate respective coordinated contractions of the detrusor muscle. ... | Medtronic, Inc. | A61B018/14 A61B005/20 A61B005/0488 A61B008/08 A61B005/11 A61B018/00 A61B018/08 A61B018/06 A61B018/02 A61B005/00
US09320939 | B1 | 2016-04-26 | Modular mirror box therapy system for the lower extremity | 2014-10-29 | The disclosed systems and methods provide for a mirror box therapy system as a therapeutic tool for treating patients having one healthy and one unhealthy limb. This can include stroke patients or those whom otherwise have full or partial paralysis. Also suitable patients include amputees with phantom limb pain or those with intact legs whom have unilateral chronic lower extremity pain. The system is designed to be used with patients in the short sit position or long sit position and can be alternated between each position including switching between left side and right side limb treatment. The methods of treatment with the disclosed system are designed to provide for positive visual feedback and induce brain plasticity with the improvement of neural connections leading to pain and/or motor rehabilitation of the impaired limb. | 1. A mirror box, comprising: a) three interconnected panels defining a C-shaped internal space configured to hold a patient's unhealthy limb, comprising: i. a rectangular mirror panel having first and second long and short sides; wherein the mirror faces away from the internal space;ii. an opaque long rectangular panel, having first and second long and short sides, wherein the first long side is perpendicularly hinged to the first long side of a mirror frame such that the long rectangular panel can fold toward the mirror panel to be substantially parallel with it;iii. a square or rectangular sliding panel, comprising a sliding side that is 25-75% less than a length of the first and second long sides of the long rectangular panel and is perpendicularly hinged to and configured to slide along the second long side of the mirror frame; such that the sliding panel can fold toward the mirror panel to be substantially parallel with it; andb) an opaque flap that can be releasably attached to the mirror box to block the patient from viewing their unhealthy limb when positioned within the internal space of the mirror box. 2. The mirror box of claim 1, further comprising a support arm substantially the same height of the first and second short sides of a mirror frame; positioned between the long rectangular panel and the sliding panel to keep them from folding toward the mirror panel. 3. The mirror box of claim 2, wherein the support arm is releasably attachable to the mirror box and comprises a middle frame rail that is about the same height of the first and second short sides of a mirror panel and is connected to two short frame rails that are shorter in length than the first and second short sides of a long rectangular panel frame, positioned between the long rectangular panel and the sliding panel to keep them from folding toward the mirror panel. 4. The mirror box of claim 3, wherein the sliding panel and the long rectangular panel have releasable fasteners to secure the support arm. 5. The mirror box of claim 4, wherein both a first and second non-sliding side of the sliding panel and each of the first and second short sides of the long rectangular panel comprise a releasable fastener for coupling the two short frame rails of the support arm. 6. The mirror box of claim 5, wherein the releasable fasteners are hook and loop straps. 7. The mirror box of claim 1, wherein the opaque flap is longer than the first and second short sides of the mirror panel. 8. The mirror box of claim 7, wherein the opaque flap is configured to be releasably attachable to an outer surface of the long rectangular panel. 9. The mirror box of claim 8, wherein the opaque flap is configured to be releasably attachable near the first short side of the long rectangular panel and the second short side of the long rectangular panel. 10. A method of treatment with the use of the mirror box of claim 1, comprising: a) identifying a patient with one unhealthy limb and a healthy l... | Amanda Hyslop | G09B019/00 A63B023/04 A61B005/00 A61F002/50
US09345435 | B1 | 2016-05-24 | Esophageal introducer | 2014-01-26 | An esophageal introducer is provided which facilitates the passage of medical probes into the esophagus or gastric cavity. It consists of an introduced tube 12 made of collapsible elastomeric material 21 with a flexible laminated steel at the bottom 20. The proximal opening is attached to a bite block 10. The distal end continues with an extension of the laminated steel 20a that tapers to a tip with a 45 degree bend 16. This extension is covered with elastomeric material 21a that overlaps the distal tip creating a flat soft end 19. Prior to insertion the introducer body is bended 20-30 degrees 17, the tip is advanced till it makes contact with the posterior pharyngeal wall FIG. 3a. ;Further pressure will advance the tip effortlessly, sliding on the posterior esophageal wall that is an anatomic continuation of the posterior pharyngeal wall. FIG. 3b, 3c. | 1. An article of medical use being an esophageal introducer to facilitate the passage of medical instruments form the mouth into the esophagus that comprises: (a) an elongated, fully collapsible within the anatomical boundaries of use introducer tube made with means for structural consistency along a bottom of said introducer tube, said introducer tube having a proximal end, a distal end, and a lumen extending between; said introducer tube having a predetermined internal diameter large enough to allow medical instruments to pass unobstructed through a length thereof, said length enough to traverse the distance from the mouth incisors line to the esophageal opening;(b) a bite block attached to the proximal end of said fully collapsible introducer tube; and(c) a distal extension of the introducer tube having a flat, tapered shape of a predetermined length that is a continuation of said bottom of the introducer tube with means for structural consistency. 2. The introducer assembly of claim 1 wherein said collapsible introducer tube is made of elastomeric material able to completely collapse during introduction into the pharynx and stretch out during passage of medical probes into the esophagus. 3. The introducer assembly of claim 2 wherein said assembly is made of a high tensile elastomer like nitrile rubber. 4. The introducer assembly of claim 2 wherein said collapsible introducer tube tapers from said proximal end to said distal end. 5. The introducer assembly of claim 2 wherein said distal end of the introducer tube includes edges that are beveled or thinned out. 6. The introducer assembly of claim 2 wherein said elastomeric material has a durometer of about 60 Shore A or less. 7. The introducer assembly of claim 2 wherein the inner surface of said collapsible introducer tube is covered by dry lubricant and viscous lubricant. 8. The introducer assembly of claim 1 wherein the means for providing structural consistency further comprises a steel laminate of flexible material with enough rigidity to maintain said structural consistency. 9. The introducer assembly of claim 8 wherein said steel laminate is attached at the bottom of the introducer tube and has a thickness of approximate 0.018 in. 10. The introducer assembly of claim 1 wherein said distal extension is comprised of a continuation of said steel laminate at the bottom of said elastomeric collapsible introducer tube as a means for structural consistency. 11. The introducer assembly of claim 8 wherein said steel laminate has a tapered shape towards a distal end thereof. 12. The introducer assembly of claim 7 wherein said steel laminate is covered by elastomeric material overlapping the edges thereof, thereby providing non traumatic contact surfaces with tissue. 13. The introducer assembly of claim 1 wherein a distal 3 mm of said steel laminate angulates 45 degrees. 14. The introducer assembly as in claim 1 wherein said elastomeric introducer tube is bent about 20-30 degrees in th... | Cesar Del Aguila | A61B005/00 A61B005/02 A61B017/02
US09332910 | B2 | 2016-05-10 | Absentminded state determination apparatus | 2012-10-03 | An absentminded state determination apparatus includes: a data obtaining element for obtaining a time series data of a physiological characteristic value of a participant, wherein a fluctuation component is overlapped on the time series data, and depends on a state of the participant including normal and absentminded states; a detection element for detecting reflecting portions of the time series data, which reflect on the fluctuation component; a counting element for counting the number of the reflecting portions of the time series data in a determination time period between a determination time and a certain past time; and a determination element for determining according to the number of the reflecting portions at the determination time whether the participant is in the absentminded state. The determination time period has a predetermined time width from the certain past time to the determination time. | 1. An absentminded state determination apparatus comprising: a data obtaining element for obtaining time series data, wherein at least portions of data points in the time series data are time intervals between adjacent pulse beats or heart beats of a participant;a detection element for detecting a plurality of peaks and troughs of the time series data;a counting element for counting the number of the peaks and troughs detected by the detection element in a determination time period; anda determination element for determining, without using a Fourier transform, according to the number of the peaks and troughs counted by the counting element whether the participant is in an absentminded state,wherein each peak is a data point, which is larger than each adjacent data point,wherein each trough is a data point, which is smaller than each adjacent data point,wherein each peak represents the largest time interval in a region of the time series data,wherein each trough represents the smallest time interval in a region of the time series data. 2. The absentminded state determination apparatus according to claim 1, further comprising: a threshold setting element for setting a threshold for the number of the peaks and troughs,wherein the determination element determines that the participant is in the absentminded state when the number of peaks and troughs is smaller than the threshold. 3. The absentminded state determination apparatus according to claim 2, further comprising: a standard time period determination element for determining a standard time period, which has a same time length as the determination time period, wherein the standard time period is defined in such a manner that the participant is deemed to be in a normal state in the standard time period; anda standard number counting element for counting the number of peaks and troughs in the standard time period, which is defined as a standard number,wherein the threshold setting element sets the threshold according to the standard number. 4. The absentminded state determination apparatus according to claim 3, wherein the participant is a driver of a vehicle,the absentminded state determination apparatus further comprising:an information obtaining element for obtaining at least one of vehicle acceleration information and vehicle speed information,wherein the standard time period determination element includes a standard time period judgment element for judging a time period, in which the vehicle runs stably, according to the at least one of vehicle acceleration information and vehicle speed information, andwherein the standard time period determination element selects one of time periods, which are judged as stable running periods by the standard time period judgment element, as the standard time period. 5. The absentminded state determination apparatus of claim 1, the peaks and troughs of the time series data, in which the counting element counts the number of peaks and troughs, include noise... | DENSO CORPORATION Koji Oguri | G08B023/00 A61B005/02 A61B005/18 A61B005/00 A61B005/0402 A61B005/0205 B60W040/08
US09339229 | B2 | 2016-05-17 | Analyte monitoring devices and methods | 2011-12-13 | On-body units are provided and include a transcutaneous analyte sensor, electronic circuitry electrically coupled to the sensor, and an on-body housing having electrical contacts disposed thereon. The on-body housing comprises a substantially flat surface and a curved and convex surface opposite the substantially flat surface. Systems and methods including the on-body units are also provided. Methods for performing continuity measurements using on-body units are also provided. The methods include positioning an on-body unit (OBU) on a skin of a subject, contacting a continuity test instrument to the OBU; and performing a continuity measurement with the continuity test instrument. An AC signal may be provided across two electrical contacts, or a DC signal may be provided across two electrical contacts in opposite directions for the same amount of time. The measurement may be performed with one electrical contact, or with isolated electrical contacts. | 1. An analyte monitoring system, comprising: an on-body unit (OBU) for analyte monitoring, the OBU comprising: a transcutaneous analyte sensor;electronic circuitry coupled to the analyte sensor; andan on-body housing comprising a substantially flat surface and a curved and convex surface opposite the substantially flat surface, wherein the on-body housing has electrical contacts disposed on the curved and convex surface, the electrical contacts electrically coupled to the electronic circuitry; anda hand-held reader comprising a plurality of electrical contacts that electrically couple to the electrical contacts on the on-body housing when the reader is positioned in physical contact with the on-body housing. 2. The analyte monitoring system of claim 1, wherein the on-body housing is dome-shaped. 3. The analyte monitoring system of claim 2, wherein the electronic circuitry comprises first and second RC pairs in series, and wherein the electrical contacts comprise: a first concentric electrical contact electrically connected between the first and second RC pairs;a second concentric electrical contact electrically connected to a first end of the first and second RC pairs in series; anda third concentric electrical contact electrically connected to a second end of the first and second RC pairs in series;wherein the analyte sensor comprises a working electrode and counter electrode, the working electrode coupled to the second concentric electrical contact, and the counter electrode coupled to the third concentric electrical contact. 4. The analyte monitoring system of claim 1, wherein the on-body housing is cone-shaped. 5. The analyte monitoring system of claim 4, wherein the electronic circuitry comprises first and second RC pairs in series, and wherein the electrical contacts comprise: a first concentric electrical contact electrically connected between the first and second RC pairs;a second concentric electrical contact electrically connected to a first end of the first and second RC pairs in series; anda third concentric electrical contact electrically connected to a second end of the first and second RC pairs in series;wherein the analyte sensor comprises a working electrode and counter electrode, the working electrode coupled to the second concentric electrical contact, and the counter electrode coupled to the third concentric electrical contact. 6. The analyte monitoring system of claim 1, wherein the electrical contacts comprise first, second, and third concentric electrical contacts disposed on an external surface of the on the on-body housing. 7. The analyte monitoring system of claim 1, wherein the on-body housing comprises a base perimeter with flat edges. 8. The analyte monitoring system of claim 1, wherein the on-body housing comprises a recessed base perimeter. 9. The analyte monitoring system of claim 1, wherein the electronic circuitry comprises first and second RC pairs in series, and wherein the electrical contacts comprise:... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/145 A61B005/1473
US09332943 | B2 | 2016-05-10 | Flexible surface parameter measurement system for the muscular-skeletal system | 2011-09-23 | At least one embodiment is directed to an insert for measuring a parameter of the muscular-skeletal system. The insert can be temporary or permanent. In one embodiment, the insert is prosthetic component for a single compartment of the knee. The insert comprises a support structure and a support structure respectively having an articular surface and a load bearing surface. The height of the insert can be less than 10 millimeters. At least one internal cavity is formed when support structures are coupled together for housing electronic circuitry, sensors, and the power source. The insert includes a flexible articular surface. Flexible articular surface transfers loading to sensors internal to the insert. | 1. A measurement system for measuring loading applied by a muscular-skeletal system comprising: an insert configured to measure loading value and loading location, where the insert includes:a first support structure having an articular surface and a major interior surface, wherein the major interior surface is planar, wherein the first support structure has a first region with a first thickness and a second region with a second thickness, and wherein the first and second thicknesses are different;a load plate wherein the load plate has a major planar surface, wherein the major planar surface of the load plate couples to the major interior surface of the first support structure;at least three sensors, where the articular surface is configured to transfer loading to the load plate, wherein the load plate is configured to transfer loading to the at least three sensors, wherein data from the at least three sensors is configured to be used by a processor to determine loading location and loading value on the articular surface, wherein the first support structure includes a circumferential groove formed in a periphery of the first support structure, the circumferential groove allowing the articular surface to flex under loading. 2. A measurement system for measuring loading applied by a muscular-skeletal system comprising: an insert configured to measure loading value and loading location, where the insert includes:a first support structure having an articular surface and a major surface wherein the first support structure includes a circumferential groove formed in a periphery of the first support structure, the circumferential groove allowing the articular surface to flex under loading;at least three sensors underlying the articular surface of the first support structure, wherein measurement data from the at least three sensors is configured to be used by a processor to determine loading location and loading value on the articular surface;a load plate coupled between the major surface of the first support structure and the at least three sensors; anda second support structure coupled to the first support structure where the first and second support structures are configured to form a housing that isolates the at least three sensors from an external environment. 3. The measurement system of claim 2 further including: electronic circuitry coupled to the at least three sensors configured to receive measurement data from the at least three sensors and transmit the measurement data from the insert; anda power source configured to provide power to the electronic circuitry and the at least three sensors where the electronic circuitry and power source are in the housing, where the first support structure has a first region with a first thickness and a second region with a second thickness, and where the first and second thicknesses are different. 4. A measurement system for measuring loading applied by a muscular-skeletal system comprising: a housing conf... | ORTHOSENSOR INC | A61B005/00 A61B005/103 A61F002/38 A61F002/46 A61F002/30
US09332934 | B2 | 2016-05-10 | Analyte sensor with lag compensation | 2013-02-08 | In particular embodiments, methods, devices and systems including calibrating analyte data associated with a monitored analyte level received from an analyte sensor based on a reference measurement, determining a lag time constant associated with the calibrated analyte data, and performing lag correction of the calibrated analyte data based on the determined time lag constant are disclosed. | 1. A computer implemented method for execution using one or more processors, the method comprising: receiving analyte data obtained from an analyte sensor in fluid contact with interstitial fluid under a skin layer;determining an analyte level in the interstitial fluid based at least in part on a nominal sensitivity associated with the analyte sensor and the received analyte data;determining a lag time constant associated with the determined analyte level of a user;performing lag correction of the determined analyte level based at least in part on the determined lag time constant;updating a rate of change of the analyte level;determining that the received analyte data has reached a predetermined count; andmodifying the lag corrected analyte level based on the updated rate of change of the analyte level and the determination that the received analyte data has reached the predetermined count. 2. The method of claim 1, further including outputting an indication associated with the lag corrected analyte level. 3. The method of claim 1, wherein the analyte level is a glucose level. 4. The method of claim 1, wherein the nominal sensitivity associated with the analyte sensor is determined prior the analyte sensor positioned in fluid contact with the interstitial fluid. 5. The method of claim 1, wherein performing lag correction of the analyte level is further based at least in part on the determined rate of change of the analyte level. 6. The method of claim 1, wherein determining the analyte level is further based at least in part on an average of at least two received analyte data. 7. The method of claim 1, wherein the subsequently received analyte data is updated based on the modified lag corrected analyte level. 8. A method performed using one or more processors, the method comprising: receiving analyte data obtained from an analyte sensor in fluid contact with interstitial fluid under a skin layer;determining an analyte level of the interstitial fluid based at least in part on a nominal sensitivity associated with the analyte sensor and the received analyte data;estimating a lag time based on a time constant component or a time shift component associated with the determined analyte level of the interstitial fluid;updating a rate of change of the analyte level;determining a count has reached a predetermined count; andmodifying the lag corrected analyte level based on the updated rate of change of the analyte level and the determination that the count has reached the predetermined count. 9. The method of claim 8, further including performing lag correction of the determined analyte level based on one or more of the estimated lag time or the rate of change of the analyte level. 10. The method of claim 8, wherein determining the analyte level is based at least in part on an average of at least two received analyte data. 11. The method of claim 8, wherein subsequently received analyte data is updated based on the modified lag corrected analyte... | Abbott Diabetes Care Inc. | A61B005/1486 A61B005/1495 A61B005/145 G01N033/487 G06F019/00 A61B005/00 A61B005/01
US09349100 | B2 | 2016-05-24 | Method for providing a prompt for real-time cognitive assistance | 2012-08-30 | Techniques for providing a prompt for real-time cognitive assistance. A method includes analyzing input from at least one environmental sensor to identify context information pertaining to a user situation, identifying a likely subsequent cognitive task of the user in the user situation based on the context information and use of a learned model, determining an action with respect to information to be suggested to the user via a corresponding prompt, wherein the determining is based on the likely subsequent cognitive task, the context information and information learned from at least one previous user situation, computing a confidence value to represent a level of certainty in the action, and providing the prompt to the user if the action has a confidence value greater than a threshold value corresponding to the action. | 1. A method for providing a prompt for real-time cognitive assistance, the method comprising: analyzing input from at least one environmental sensor to identify context information pertaining to a user situation;forecasting a subsequent cognitive task of the user in the user situation based on the context information and use of a learned model;determining a combination of two or more words to be communicated by the user in anticipation of the user communicating said combination of two or more words, wherein the determining is based on the forecasted subsequent cognitive task, the context information and information learned from at least one previous user situation;computing a confidence value to represent a level of certainty in the combination of two or more words to be communicated by the user in connection with the forecasted subsequent cognitive task;updating the confidence value at a specified interval until the confidence value reaches a threshold value; andproviding a prompt to the user, wherein said prompt contains the combination of two or more words to be communicated by the user, upon the confidence value reaching the threshold value;wherein at least one of the steps is carried out by a computer device. 2. The method of claim 1, wherein the threshold value corresponding to the combination of two or more words to be communicated by the user is modified based on feedback pertaining to the combination of two or more words to be communicated by the user. 3. The method of claim 1, further comprising: analyzing user input to identify a cognitive state of the user. 4. The method of claim 3, wherein the user input comprises brain wave activity or an electroencephalogram (EEG). 5. The method of claim 3, wherein the user input comprises information from a stored user profile that specifies at least one aspect of the user. 6. The method of claim 5, wherein the at least one aspect of the user includes a state of a user disease, nationality of the user, location of the user, language of the user, or employment of the user. 7. The method of claim 1, wherein the input includes audio, video or text-based input. 8. The method of claim 1, further comprising: analyzing input provided by an individual to supplement identification of the user situation. 9. The method of claim 8, wherein the individual comprises the user or an additional individual. 10. The method of claim 1, further comprising: using a learned model to identify a state of an environment involved in the user situation. 11. The method of claim 1, wherein said forecasting comprises performing an analysis of at least one word combination or word association of the user. 12. The method of claim 1, wherein said forecasting comprises analyzing at least one word combination used by at least one additional individual. 13. The method of claim 1, wherein said forecasting comprises determining at least one set of words commonly used in a particular context. 14. The method of claim 1, wher... | International Business Machines Corporation | G06F015/18 G06N007/00 G06F019/00 A61B005/00 A61B005/0476
US09345389 | B2 | 2016-05-24 | Additional systems and methods for providing real-time anatomical guidance in a diagnostic or therapeutic procedure | 2011-11-14 | A system and method for intra-operatively providing anatomical guidance in a diagnostic or therapeutic procedure is disclosed. In embodiments, the system includes multiple light sources configured to emit different frequencies, multiple electronic imaging devices to detect various frequencies of reflected, emitted, or scattered light. The system and method incorporate an optical probe is integral to an endoscopic device or a therapeutic laser system, optically coupled to a light source; a display for displaying at least one visual representation of data; and a controller programmed to generate at least one real-time integrated visual representation of an area of interest and to display the real-time visual representation on the display for guidance during the diagnostic or therapeutic procedure. | 1. A system for intra-operatively providing anatomical guidance in a diagnostic or therapeutic procedure, comprising: (a) a lamp configured to emit a beam of visible light to an area of interest of a living subject;(b) a laser configured to emit a beam of near-infrared light to the area of interest;(c) an optical probe optically coupled to the laser but not optically coupled to the lamp, comprising an optical fiber configured to deliver the emitted beam of near-infrared light to illuminate the area of interest and configured to collect light that is scattered or emitted from a contrast agent introduced into target tissues in the area of interest, in response to illumination by the laser;(d) a spectrometer optically coupled to the optical probe and configured to detect the collected light and to generate a corresponding signal that comprises collected light data, and wherein the optical probe is further configured to transmit the collected light to the spectrometer through the optical fiber;(e) a first CCD or CMOS camera configured to detect visible light that is emitted from the area of interest in response to illumination by the lamp and to generate a corresponding signal comprising visible light data;(f) a second CCD or CMOS camera configured to detect near-infrared light having a first predetermined wavelength that is emitted from the area of interest in response to illumination by the laser and to generate a corresponding signal comprising a first set of near-infrared light data;(g) a third CCD or CMOS camera configured to detect near-infrared light having a second predetermined wavelength that is different from the first predetermined wavelength and that is emitted from the area of interest in response to illumination by the laser, and to generate a corresponding signal comprising a second set of near-infrared light data;(h) a display for displaying at least one visual representation of data; and(i) a central processing unit, a processor, or a microprocessor in communication with each of the lamp, laser, spectrometer, first CCD or CMOS camera, second CCD or CMOS camera, third CCD or CMOS camera, and display, and programmed to generate at least one real-time integrated visual representation of the area of interest from each of the collected light data, visible light data, first set of near-infrared light data, and second set of near-infrared light data and to display the at least one real-time visual representation on the display, for guidance during the diagnostic or therapeutic procedure. 2. A system according to claim 1, further comprising a speaker, wherein the programmer is further programmed to generate at least one real-time aural representation of the area of interest from each of the collected light data, first set of near-infrared light data, and second set of near-infrared light data and to emit a sound for the at least one real-time aural representation through the speaker, for guidance during the diagnostic or therapeutic proced... | EMORY UNIVERSITY | A61B001/04 A61B001/00 A61B005/00 A61B005/06 G01J003/02 G01J003/44 A61B018/20 A61B019/00 A61B017/00
US09339199 | B2 | 2016-05-17 | Physiological signal detection device | 2013-02-26 | A physiological signal detection device has a battery disposed in the physiological signal detection device, a first input terminal, a second input terminal, and a charging detection terminal A physiological signal detection circuit generates a physiological signal according to a detection result of the first input terminal and the second input terminal A charging control circuit is electrically coupled to the first input terminal, the second input terminal and the charging detection terminal, wherein, when the first input terminal and the second input terminal are coupled to a power supply supplied by a charging device, the charging detection terminal receives a charging indication signal of the charging device and according to the charging indication signal the charging device is enabled so as to charge the battery with the power supply. | 1. A physiological signal detection device, comprising: a battery, disposed in the physiological signal detection device, comprising a positive terminal and a negative terminal;a first input terminal;a second input terminal;a charging detection terminal, for receiving a charging indication signal from a charging device;a physiological signal detection circuit, electrically coupled to the first input terminal and the second input terminal, wherein, when the first input terminal and the second input terminal are electrically connected to a measure target, the physiological signal detection circuit generates a physiological signal according to a detection result of the first input terminal and the second input terminal; anda charging control circuit, electrically coupled to the first input terminal, the second input terminal and the charging detection terminal, wherein, when the first input terminal and the second input terminal are coupled to a power supply supplied by the charging device, the charging detection terminal receives the charging indication signal of the charging device and according to the charging indication signal the charging control circuit is enabled so as to charge the battery with the power supply,wherein:the charging control circuit further comprises a first resistor coupling the charging detection terminal to a power terminal of the charging control circuit;when the first input terminal, the second input terminal and the charging detection terminal are coupled to the charging device but the charging device is not supplying the power supply, the first resistor and internal resistors within the charging device and coupled to the charging detection terminal form a voltage divider that divides a voltage at the power terminal of the charging control circuit to generate a divided voltage at the charging detection terminal for recognizing that the physiological signal detection device is operated in a storage mode;the charging control circuit further comprises a second resistor coupled between the charging detection terminal and a ground terminal;the charging device comprises a third resistor and a fourth resistor to be coupled in series between the charging detection terminal and a ground terminal; andthe charging device further comprises a diode having an anode coupled to a connection node between the third resistor and the fourth resistor and having a cathode to be coupled to the first input terminal. 2. The physiological signal detection device as claimed in claim 1, further comprising: a microprocessor unit, electrically coupled to the charging detection terminal and the physiological signal detection circuit, wherein the microprocessor unit receives the physiological signal generated from the physiological signal detection circuit and calculates a plurality of physiological information according to the physiological signal. 3. The physiological signal detection device as claimed in claim 2, further comprising: a power switch... | QUANTA COMPUTER INC. | A61B005/0452 A61B005/024 A61B005/00
US09345432 | B2 | 2016-05-24 | Systems and methods for sleep apnea detection from breathing sounds | 2012-06-19 | A method and an apparatus are disclosed for extending the service life of a battery source by fully utilizing the energy in batteries and/or providing higher energy reserve using multiple batteries. An universal battery bank consists of an ensemble of possibly heterogenous batteries all connected in series to provide a usable operational energy source. For a universal battery bank consists on N batteries, the minimum operational voltage for the individual batteries (Vmin) can be lowered to Vmin/N on average, thereby extending the time interval between battery replacements and draining the individual batteries more efficiently. To exploit this power supply that may vary over a wide voltage range, electronics are designed to monitor the individual batteries, guard against excessive voltage and current, and provide constant current sources and/or constant voltage sources for various appliances. | 1. A method of detecting obstructive sleep apnea in a subject, said method comprising the steps of: placing a head of an electronic stethoscope over the subject's suprasternal notch when said subject is in an awake state;providing an electrical signal representative of breathing sounds detected by the head of the electronic stethoscope within a frequency range, the breathing sounds corresponding to a patient breathing through their mouth;determining, by a processor, an obstructive sleep apnea index based on a ratio of an amount of energy of the electrical signal that is associated with a frequency below a cut-off frequency with respect to a total amount of energy associated with the entire frequency range, wherein the amount of energy associated with the frequency below the cut off frequency and the total amount of energy associated with the entire frequency range are each determined from a frequency spectrum of the electrical signal, wherein information regarding the frequency associated with the electrical signal are assessed in the frequency domain where frequency information from the frequency domain is used for calculating a spectral energy distribution of frequencies of the electrical signal to detect obstructive sleep apnea based on the spectral energy distribution of frequencies of the electrical signal; andidentifying, by the processor, the subject as having obstructive sleep apnea if the obstructive sleep apnea index is above a window maximum or below a window minimum. 2. The method as claimed in claim 1, wherein said cut-off frequency is about 100 Hz. 3. The method as claimed in claim 1, wherein said frequency range is between about 20 Hz and about 800 Hz. 4. The method as claimed in claim 1, wherein said window maximum is about 80%. 5. The method as claimed in claim 1, wherein said window minimum is about 25%. 6. The method as claimed in claim 1, wherein said method further includes the step of identifying, by the processor, the subject as not having obstructive sleep apnea if the obstructive sleep apnea index is within a window defined between the window maximum and the window minimum. 7. A method of detecting obstructive sleep apnea in a subject, said method comprising the steps of: placing a head of an electronic stethoscope over the subject's suprasternal notch when said subject is in an awake state;providing an analog electrical signal representative of breathing sounds detected by the head of the electronic stethoscope within a frequency range, the breathing sounds corresponding to a patient breathing through their mouth;receiving, by a recorder, the analog electrical signal and providing a digitized electrical signal to a processor that is representative of the analog electrical signal with the frequency range;applying, by the processor, a fast Fourier transform to the digitized electrical signal;determining, by the processor, an obstructive sleep apnea index based on a ratio of an amount of energy of the electrical signa... | Rhode Island Board of Education, State of Rhode Island and Providence Plantations | A61B007/00 A61B005/00 A61F005/56
US09326707 | B2 | 2016-05-03 | Alarm characterization for analyte monitoring devices and systems | 2009-11-10 | Methods and apparatus including determining a rate of occurrence of a glycemic excursion event, determining a frequency of an alarm activation associated with the glycemic excursion event, determining an analyte level associated with the alarm activation, and setting an alarm parameter based on one or more of the determined rate of occurrence of the glycemic excursion event, the frequency of the alarm activation associated with the glycemic excursion event or the determined analyte level are provided. | 1. A method, comprising: positioning an in vivo analyte sensor in fluid contact with interstitial fluid under a skin surface;monitoring an analyte level in the interstitial fluid with the in vivo analyte sensor;determining, using one or more processors, a rate of occurrence of a glycemic excursion event based on the monitored analyte level;determining, using the one or more processors, a number of alarm activations associated with the glycemic excursion event based on the monitored analyte level;setting, using the one or more processors, a parameter for the alarm activation based on the determined rate of occurrence of the glycemic excursion event and the number of alarm activations associated with the glycemic excursion event, wherein setting the parameter for the alarm activation is programmable; andcontrolling administration of therapy based on the parameter for the alarm activation. 2. The method of claim 1, wherein the glycemic excursion event includes one of a hypoglycemic event or a hyperglycemic event. 3. The method of claim 1, further including determining an analyte level associated with the alarm activation based at least in part on analyte sensor data from the in vivo analyte sensor. 4. The method of claim 3, including determining a rate of change of the analyte level associated with the glycemic excursion event or the alarm activation, or a combination thereof. 5. The method of claim 3, wherein the transcutaneous analyte sensor comprises a plurality of electrodes including a working electrode, wherein the working electrode comprises an analyte-responsive enzyme and a mediator, wherein at least one of the analyte-responsive enzyme and the mediator is chemically bonded to a polymer disposed on the working electrode, and wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer. 6. The method of claim 1, wherein determining the rate of occurrence of the glycemic excursion event includes: monitoring a glucose level for a predetermined time period;detecting the monitored glucose level transcending a predetermined threshold glucose level; anddetermining the rate of occurrence based on the detected monitored glucose level transcending the predetermined threshold glucose level within the predetermined time period. 7. The method of claim 1, including programming an alarm function based on the parameter for the alarm activation. 8. The method of claim 1, wherein the number of alarm activations is a number of alarm occurrences over a predetermined time period, and wherein setting the parameter for the alarm activation includes modifying a threshold level of the alarm activation based on determining the number of alarm activations. 9. The method of claim 1, wherein modifying the parameter for the alarm activation is programmable to reduce the rate of occurrence of the glycemic excursion event. 10. The method of claim 1, including generating a notification function associated with an analyte m... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/145
US09352152 | B2 | 2016-05-31 | Equine airway disorders | 2007-12-21 | Airway disorders in a horse are relieved by electrical stimulation of the airway tissue. Particular disorders and techniques include electrical stimulation to relieve laryngeal hemiplegia. A pacemaker processor generates an electrical treatment signal to be applied to upper airway tissue of the horse for treating the upper airway disorder. One or more stimulation electrodes interfaces with the upper airway tissue for delivering the treatment signal to the upper airway tissue. | 1. A device for treating laryngeal hemiplegia in horses, the device comprising: a pacemaker processor configured to generate an electrical signal to be applied to upper airway tissue of the horse and configured to cause complete and continuous abduction of the vocal fold during a period of minutes up to hours in order to treat the laryngeal hemiplegia; andone or more stimulation electrodes adapted to interface with the upper airway tissue for delivering the electrical signal to the upper airway tissue. 2. The device of claim 1, wherein at least a portion of the device is implanted in the horse. 3. The device of claim 2, wherein the implanted portion of the device communicates transcutaneously or percutaneously with a portion of the device located externally to the horse. 4. The device of claim 3, wherein communication transcutaneously is based on at least one of electromagnetic induction, acoustic energy, optical energy, and capacitor coupling. 5. The device of claim 2, wherein a portion of the device is configured to be placed temporarily on the surface of the horse when the device is operating to provide external signals to the implanted portion of the device. 6. The device of claim 2, wherein the implanted portion includes a power source which is charged percutaneously or transcutaneously. 7. The device of claim 1, wherein the electrical signal is derived from at least one of an electromyogram, an electronystagmograph, an electroglottograph, an electroencephalograph, a biopotential sensor, an ultrasound sensor, a hall sensor, a microphone, a pressure sensor, a strain transducer, a mechanical deformation sensor, and a motion sensor. 8. The device of claim 1, wherein the electrical signal is applied to the upper airway tissue of the horse using a biphasic waveform. 9. The device of claim 1, wherein the one or more stimulation electrodes are based on at least one of a cuff electrode, a multipolar cuff electrode, a tripolar cuff electrode, a flat nerve electrode, an epineural electrode, a shaft electrode, a longitudinal intrafascicular electrode, a thin wire electrode, a micro-machined electrode, a sieve electrode, and a staple electrode. 10. The device of claim 1, wherein the upper airway tissue includes one or more nerves of an airway structure. 11. The device of claim 10, wherein the one or more nerves includes the recurrent laryngeal nerve of the horse. 12. The device claim 11, wherein the upper airway tissue includes one or more axons of the abductor branch of the recurrent laryngeal nerve. 13. The device of claim 1, further comprising one or more treatment sensors configured to sense at least one therapy parameter related to the horse, wherein the at least one therapy parameter relates to at least one of air flow characteristics of the airway tract of a horse, contractile characteristics of the airway tissue of a horse, electrical characteristics of a portion of the body of the horse, temperature of a portion of the body of the h... | MED-EL Elektromedizinische Geraete GmbH Cornell University | A61N001/36 A61B005/00 A61N001/372 A61B005/0476 A61B005/0488 A61B005/085 A61B007/00 A61N001/05 A61B005/11
US10638969 | B2 | 2020-05-05 | Wireless pressure ulcer alert methods and systems therefor | 2015-11-09 | Pressure monitoring methods and systems for warning a patient or caregiver that soft tissue pressure has exceeded some predetermined level that over time would necessitate moving the patient to prevent or at least reduce a risk of soft tissue damage. The methods and systems entail the use of a pressure sensing unit adapted to be applied to or near a surface of the patient's body, and a sensor associated with the sensing unit to generate electrical outputs corresponding to soft tissue pressure sensed at the surface. The electrical outputs are wirelessly monitored over a preselected time period to generate a cumulative output signal based on the electrical outputs and corresponding to whether or not the soft tissue pressure has exceeded a predetermined pressure level during the preselected time period. An alarm is generated if the cumulative output signal exceeds a predetermined cumulative threshold until the soft tissue pressure drops below the predetermined pressure level. | 1. A pressure monitoring system for providing a warning to a patient or caregiver that the patient should be moved to at least reduce a risk of soft tissue damage to the patient, the system comprising: a plurality of pressure sensing units, each pressure sensing unit adapted to be applied on or near a surface of the patient's body that is susceptible to damage from soft tissue pressure, each pressure sensing unit comprising a vibration device and at least one sensor configured to generate electrical outputs corresponding to soft tissue pressure sensed by the sensor at the surface of the patient's body;means for wirelessly monitoring a plurality of the electrical outputs generated by the sensors of the pressure sensing units;counters associated with the monitoring means and the sensors, each of the counters generating a separate counter value for a corresponding one of the sensors that increases from an initial value while the soft tissue pressure sensed by the corresponding sensor exceeds a predetermined pressure level and decreases toward the initial value while the soft tissue pressure sensed by the corresponding sensor does not exceed the predetermined pressure level;a display associated with the monitoring means that visually displays a human outline, visual indications of locations and identities of the sensors on the human outline, the counter values of the sensors, and a visual warning if the counter value of any of the sensors exceeds a predetermined counter value;means for generating an alarm while any of the counter values exceed the predetermined counter value; andwherein the vibration device of any one of the pressure sensing units is activated to vibrate to notify the patient of the alarm if both the counter value for a corresponding one of the sensors exceeds the predetermined counter value and the soft tissue pressure sensed by the corresponding one of the sensors exceeds the predetermined pressure level, and is deactivated once the soft tissue pressure is no longer sensed by the corresponding one of the sensors or no longer exceeds the predetermined pressure level. 2. The pressure monitoring system according to claim 1, wherein the counter values increase at a first ratio relative to actual elapsed time, the counter values decrease at a second ratio relative to actual elapsed time, and the second ratio is less than the first ratio. 3. The pressure monitoring system according to claim 1, wherein the visual warning comprises elapsed time progress bars that each increase and decrease with a corresponding one of the counter values. 4. The pressure monitoring system according to claim 1, wherein the sensors are adapted to generate the electrical outputs only if the soft tissue pressure exceeds the predetermined pressure level. 5. The pressure monitoring system according to claim 4, wherein the sensors generate the electrical outputs at an identical output level regardless of how much the soft tissue pressure exceeds the predetermin... | Walgreens Health Solutions, LLC | A61B005/00
US10206613 | B2 | 2019-02-19 | Syringe-based fluid diversion mechanism for bodily fluid sampling | 2015-10-12 | A syringe-based device includes a housing, a pre-sample reservoir, and an actuator. The housing defines an inner volume between a substantially open proximal end portion and a distal end portion that includes a port couplable to a lumen-defining device. The pre-sample reservoir is fluidically couplable to the port to receive and isolate a first volume of bodily fluid. The actuator is at least partially disposed in the inner volume and has a proximal end portion that includes an engagement portion and a distal end portion that includes a sealing member. The engagement portion is configured to allow a user to selectively move the actuator between a first configuration such that bodily fluid can flow from the port to the pre-sample reservoir, and a second configuration such that bodily fluid can flow from the port to a sample reservoir defined at least in part by the sealing member and the housing. | 1. A method of sequestering bodily fluid withdrawn from a patient with a bodily fluid sequestration device, the method comprising: establishing fluid communication between the patient and a channel of the bodily fluid sequestration device via a port, the channel defining a first fluid flow path and including a valve configured to form a substantially fluid tight seal with walls defining the channel;the valve being operative to move from a closed configuration to an open configuration in response to a difference in pressure between a valve inlet and a valve outlet,the valve being further operative to return to the closed configuration in response to equalization of pressure between the valve inlet and the valve outlet;receiving an initial volume of bodily fluid from the patient into the bodily fluid sequestration device via the port and the first fluid flow path, the first fluid flow path extending from a proximal end of the port when the initial volume of bodily fluid is received;sequestering the initial volume of bodily fluid;after the initial volume of bodily fluid is sequestered, receiving a subsequent volume of bodily fluid from the patient into a second fluid flow path of the bodily fluid sequestration device, at least a portion of the second fluid flow path formed separately from the first fluid flow path; andproviding via the second fluid flow path a volume of bodily fluid with reduced contamination due to the sequestration of the initial volume of bodily fluid. 2. The method of claim 1, wherein the valve is a one-way check valve. 3. The method of claim 1, wherein sequestering the initial volume of bodily fluid includes sequestering the initial volume of bodily fluid in a pre-sample reservoir, the pre-sample reservoir at least partially defined by the channel. 4. The method of claim 3, wherein the initial volume of bodily fluid within the pre-sample reservoir is fluidically isolated from the second fluid flow path. 5. The method of claim 3, wherein the initial volume of bodily fluid is fluidically isolated in the pre-sample reservoir. 6. The method of claim 1, wherein the valve is disposed at an inlet of the channel. 7. The method of claim 1, wherein the valve disposed in the channel is configured to sequester the initial volume of bodily fluid from the second fluid flow path. 8. The method of claim 7, wherein the valve disposed in the channel is configured to fluidically isolate the initial volume of bodily fluid from the second fluid flow path. 9. The method of claim 1, wherein the valve disposed in the channel is configured to prevent the sequestered initial volume of bodily fluid from flowing out of the channel toward a sample reservoir. 10. The method of claim 9, wherein the valve disposed in the channel is configured to fluidically isolate the initial volume of bodily fluid from the second fluid flow path. 11. The method of claim 1, wherein the valve is configured to transition from a first operating mode in which the initi... | Magnolia Medical Technologies, Inc. | A61B005/00 A61B005/15 A61B005/153 A61B010/00 A61B005/155 A61B005/154
US09636026 | B2 | 2017-05-02 | Multi-layered structure | 2016-01-08 | One aspect relates to a layered structure with a substrate, a first layer over the substrate, and a second layer over the first layer. The substrate and the second layer are an electrically conductive material and the first layer is an insulating material or the substrate and the second layer are insulating material and the first layer is electrically conductive material. At least one of the first and second layers comprises an electrically conductive polymer. | 1. A layered structure comprising: a substrate;a first layer over the substrate; anda second layer over the first layer; characterized in that the substrate and the second layer are an electrically conductive material and the first layer is an insulating material or the substrate and the second layer are insulating material and the first layer is electrically conductive material, wherein at least one of the first and second layers comprises an electrically conductive polymer, and wherein each of the first and second layers are one tenth of the thickness of the substrate such that the layered structure substantially maintains the size and configuration of the substrate;wherein at least one of the first and second layers comprises electrically conductive Poly(3,4-ethylenedioxythiophene) poly(styrenesulfonate); andwherein the substrate and the first and second layers are biocompatible such that the layered structure is configured for implantation into the human body. 2. The layered structure of claim 1, wherein the layered structure has a length defining first and second ends and wherein at least one of the first and second layers extends substantially the length of the layered structure and includes contact areas proximate to the first and second ends that are configured for attaching an electrically conducting contact. 3. The layered structure of claim 1, wherein the thickness of at least one of the first and second layers is in a range from 0.05 μm to 10 μm. 4. The layered structure of claim 1, wherein the substrate comprises one of a metal wire and metal insertion needle. 5. The layered structure of claim 1, wherein the substrate has a sheet resistance of less than 10 kΩ. 6. A layered structure comprising: a conductive substrate;a first layer over the substrate, wherein the first layer comprises an insulating material and includes at least one access area; anda contact area at least partially within the access area and in conductive contact with the substrate;characterized in that the first layer has a thickness that is no more than one tenth of the thickness of the substrate such that the layered structure substantially maintains the size and configuration of the substrate and wherein the substrate, the first layer and the contact area are biocompatible such that layered structure is configured for implantation into the human body; andwherein the contact area is an electrically conductive Poly(3,4-ethylenedioxythiophene) poly(styrenesulfonate) material; andwherein the access area is a hole in the first layer having a diameter of less than 0.4 mm. 7. The layered structure of claim 6, wherein each of the first layer and the contact area are no more than one tenth the thickness of the substrate. 8. The layered structure of claim 6 characterized in that the thickness of the first layer is no more than one fiftieth of the thickness of the substrate. 9. A layered implanted medical device comprising: a substrate includinga plurality of conduct... | Heraeus Deutschland GmbH & Co. KG | H01L051/00 A61B005/00 A61B005/04 A61N001/05 H01B007/04 H01B001/12 H01B003/30 B21F019/00 A61B005/042 A61M025/09
US09629550 | B2 | 2017-04-25 | Spring type firing mechanism applied to endoscope accessory | 2014-11-11 | A spring type firing mechanism applied to an endoscope accessory includes a housing, a cap, a bottom cover, a driving module, a buffering module and a spring pipeline, and the spring pipeline includes a push-rod coupled to the sensor, and the driving module includes a firing spring and a restoring spring, and the two springs have different rigidities so as to form a reciprocating structure of the firing mechanism. After the sensor of the firing mechanism is triggered, the resilience of the restoring spring compresses the firing spring back to its firing length, so as to restore its position by a single driving point and allow users to complete the firing operation by one hand and simplify the operation of doctors to improve convenience and efficiency. | 1. A spring type firing mechanism applied to an endoscope accessory, comprising: a housing having a hollow tube body contour, and having an outer embedding wall formed around the periphery of the housing, an inner embedding wall formed inside the housing, and a fixed base formed in a tube body of the housing;a cap, embedded into the outer embedding wall of the housing, and having an opening;a bottom cover, coupled to the housing, and a pipeline extending from a bottom portion of the bottom cover;a driving module, including: a pressing device having a hollow-contoured tube, and having a triggering rod extended separately from both sides of the pressing device;a pull-rod, installed in the hollow-contoured tube of the pressing device;a restoring spring, installed between the pressing device and the fixed base;a trigger, fixed to the pull-rod, and having an elastic brace formed on both sides of the trigger, such that the elastic brace supports and abuts the inner embedding wall of the housing; anda firing spring, installed between the fixed base of the housing and the trigger, and having a coefficient of rigidity greater than a coefficient of rigidity of the restoring spring;a buffering module, including: a seat body, disposed at an end of the housing, and having a penetrating hole; anda buffer spring, sheathed on the seat body; anda spring pipeline, including a push-rod with a front end of the push-rod fixed to the trigger and a rear end coupled to a metal conductive wire, and a sensor combined with the metal conductive wire; anda plug installed to a top of the pressing device. 2. The spring type firing mechanism applied to the endoscope accessory as claimed in claim 1, further comprising a handle cover sheathed on an outer periphery of the housing. 3. The spring type firing mechanism applied to the endoscope accessory as claimed in claim 1, further comprising a nut secured to an end of the trigger. 4. A spring type firing mechanism applied to an endoscope accessory, comprising: a housing having a hollow tube body contour, and having an outer embedding wall formed around the periphery of the housing, an inner embedding wall formed inside the housing, and a fixed base formed in a tube body of the housing;a cap, embedded into the outer embedding wall of the housing, and having an opening;a bottom cover, coupled to the housing, and a pipeline extending from a bottom portion of the bottom cover;a driving module, including: a pressing device having a hollow-contoured tube, and having a triggering rod extended separately from both sides of the pressing device;a pull-rod, installed in the hollow-contoured tube of the pressing device;a restoring spring, installed between the pressing device and the fixed base;a trigger, fixed to the pull-rod, and having an elastic brace formed on both sides of the trigger, such that the elastic brace supports and abuts the inner embedding wall of the housing; anda firing spring, installed between the fixed base of the ho... | Southern Taiwan University of Science and Technology Kuang Tai Metal Industrial Co., Ltd. | A61B005/00 A61B001/00 F03G001/10
US10506973 | B2 | 2019-12-17 | Apparatus for injecting liquid into primo-node and primo-vascular system tracing system comprising the same | 2015-11-18 | Provided are an apparatus for injecting a liquid into a primo-node and a primo-vascular system tracing system including the apparatus. The apparatus includes a liquid injection tube to inject a liquid into a primo-node of an object, and a suction tube to hold the primo-node to inject the liquid injection tube into the primo-node. The liquid injection tube and the suction tube are formed as one body. The liquid injection tube is inserted into an inner side of the suction tube and an edge unit of the liquid injection tube protrudes to the outside from an inner side of a suction opening of the suction tube. | 1. An apparatus for injecting a liquid into a primo-node, the apparatus comprising: a suction tube comprising a suction opening formed therein and configured to hold the primo-node;a liquid injection tube disposed within the suction tube, wherein an edge of the liquid injection tube protrudes outward from the suction opening; anda bridge that fixes the liquid injection tube to the suction tube. 2. The apparatus of claim 1, wherein the liquid injection tube has an outer diameter smaller than an inner diameter of the suction tube. 3. The apparatus of claim 1, further comprising a suction pipeline connected to a side of the bridge. 4. The apparatus of claim 3, further comprising a negative pressure pump connected to the suction pipeline. 5. A primo-vascular system (PVS) tracing system comprising: the injector of claim 1 that injects a liquid into a primo node of an object; andan image taker that takes an image of the object. 6. The primo-vascular system tracing system of claim 5, wherein the injector comprises: a height controller; anda micro manipulator mounted on an edge of the height controller;and wherein the liquid injector is mounted on the micro manipulator. 7. The primo-vascular system tracing system of claim 6, wherein a weight center of the micro manipulator is coupled to the height controller. 8. The primo-vascular system tracing system of claim 6, wherein the micro manipulator further comprises an extension unit, wherein the liquid injection apparatus is mounted on the extension unit. 9. The primo-vascular system tracing system of claim 6, further comprising: a display to display an image taken by the image taker; anda controller. 10. A method of injecting liquid into a primo-node, the method comprising contacting a primo node with an apparatus of claim 1, wherein the liquid injection tube penetrates the primo node;applying a suction to the primo node through the suction tube, whereby the suction holds the primo node to the suction tube; andinjecting liquid into the primo node through the liquid injection tube while applying suction to the primo node through the suction tube. 11. The method of claim 10, wherein the liquid injection tube of the apparatus has an outer diameter smaller than an inner diameter of the suction tube. 12. The method of claim 10, wherein the apparatus further comprises a suction pipeline connected to a side of the bridge. 13. The method of claim 12, wherein the apparatus further comprises a negative pressure pump connected to the suction pipeline. | SAMSUNG ELECTRONICS CO., LTD. AICT (ADVANCED INSTITUTES OF CONVERGENCE TECHNOLOGY) | A61M001/00 A61M005/32 A61B005/00 A61M005/00
US10123718 | B2 | 2018-11-13 | Methods, systems, and assemblies for measuring bioelectrical signals of intra-abdominal organs | 2015-10-26 | The present disclosure discuses systems and methods for detecting and recording bioelectric signals, and specifically bioelectrical signals generated by abdominal organs, such as the uterus. The disclosure discusses area electrodes and arrays of area electrodes. The area electrodes are defined in a metal layer and include an inner and outer diameter. The area electrodes are configured to detect electrical signals generated substantially perpendicular to a surface of the area electrode while rejecting electrical signals generated substantially parallel to the surface of the area electrode. | 1. A method for measuring at least one bioelectrical signal, the method comprising: providing a first area electrode comprising: a planar metal layer defining an electrode body having an inner diameter and an outer diameter, the inner diameter defining an opening in the planar metal layer, wherein the electrode body has a surface area between 5 cm2 and 900 cm2 and a width between 0.5 cm and 5 cm; anda connector in electrical communication with the metal layer;coupling the first electrode to an abdomen of a patient; anddetecting a first bioelectrical signal with the first area electrode generated from a muscle-containing intra-abdominal organ in a patient. 2. The method of claim 1, wherein the muscle-containing intra-abdominal organ is a uterus. 3. The method of claim 1, further comprising detecting with the first area electrode an electrical signal generated by at least one of a fetal heart and a maternal heart. 4. The method of claim 1, further comprising: detecting, with a second area electrode, a second bioelectrical signal generated from the muscle-containing intra-abdominal organ in the patient, wherein the first bioelectrical signal is generated by a first contractile unit of the intra-abdominal organ and the second bioelectrical signal is generated by a second contractile unit of the intra-abdominal organ;detecting a synchronization pattern of the first and the second bioelectrical signals; anddetermining a labor status responsive to the synchronization pattern. 5. The method of claim 1, wherein the electrode body has one of a circular shape, a square with rounded corners shape, a hexagonal shape, or a rectangle with rounded corners shape. 6. The method of claim 1, wherein the first area electrode is configured to detect electrical signals originating substantially perpendicular to a surface of the electrode body, and to substantially reject a noise signal. 7. The method of claim 6, wherein the noise signal comprises at least one of electrical signals originating substantially lateral to the surface of the electrode body and environmental electrical noise. 8. The method of claim 1, further comprising determining a direction of an origination of the first bioelectrical signal. 9. The method of claim 1, further comprising referencing the first bioelectrical signal with a second bioelectrical signal from an electrode placed remote to the opening in the planar metal layer. 10. An electrode system comprising: a first area electrode comprising:a planar metal layer defining an electrode body having an inner diameter and an outer diameter, the inner diameter defining an opening in the planar metal layer, wherein the electrode body has a surface area between 5 cm2 and 900 cm2 and a width between 0.5 cm and 5 cm; anda connector in electrical communication with the metal layer. 11. The electrode system of claim 10, wherein the electrode body has one of a circular shape, a square with rounded corners shape, a hexagonal shape, or a rectangle ... | University of Tenessee Research Foundation | A61B005/0492 A61B005/00 A61B005/024 A61B005/0488
US09351672 | B2 | 2016-05-31 | Multiplexed pathlength resolved noninvasive analyzer apparatus with stacked filters and method of use thereof | 2014-09-22 | A noninvasive analyzer apparatus and method of use thereof is described using a plurality of sample illumination zones optically coupled to at least two optically stacked two-dimensional optical filter arrays. Sectioned pixels and/or zones of a detector array are optionally filtered for different light throughput and/or are passed through various pathlengths using the stacked two-dimensional optical filter arrays. Resulting pathlength resolved/wavelength controlled groups of spectra are subsequently analyzed to determine an analyte property. | 1. An apparatus for noninvasive analyte property estimation of a constituent of a human subject, comprising: a near-infrared analyzer, comprising: a near-infrared source;a detector array;a photon transport system configured to transport photons along an optical train from said near-infrared source to said detector array via an analyzer-subject interface;a first two-dimensional filter array comprising a first set of optical filters, said first set of optical filters comprising at least two optical filters; anda second two-dimensional filter array comprising a second set of optical filters,said optical train sequentially passing through said first two-dimensional filter array and said second two-dimensional filter array;a first filter of said first two-dimensional filter array, said first filter comprising an edge; anda second filter of said second two-dimensional filter array, said second filter comprising a non-edge, surface filter area,wherein said edge of said first filter optically overlaps said non-edge, surface filter area of said second filter. 2. The apparatus of claim 1, said first set of optical filters further comprising: a first optical filter optically linked to a first set of at least three detector elements of said detector array; anda second optical filter optically linked to a second set of at least three detector elements of said detector array. 3. The apparatus of claim 2, said first set of optical filters and said second set of optical filters further comprising: a first optical connection to a common detector element of said detector array. 4. An apparatus for noninvasive analyte property estimation of a constituent of a human subject, comprising: a near-infrared analyzer, comprising: a near-infrared source;a detector array;a photon transport system configured to transport photons along an optical train from said near-infrared source to said detector array via an analyzer-subject interface;a first two-dimensional filter array comprising a first set of optical filters, said first set of optical filters comprising at least two optical filters, said first set of optical filters further comprising: a first optical filter optically linked to a first set of at least three detector elements of said detector array; anda second optical filter optically linked to a second set of at least three detector elements of said detector array;a second two-dimensional filter array comprising a second set of optical filters; anda first detector element of said set of at least three detector elements further comprising a first surface area, said second detector element of said set of at least three detector elements further comprising a second surface area, said second surface area at least thirty percent larger than said first surface area,said optical train sequentially passing through said first two-dimensional filter array and said second two-dimensional filter array. 5. The apparatus of claim 4, said second detector element, comprising sai... |  | G01J005/02 A61B005/00 A61B005/1455 A61B005/145
US09352159 | B2 | 2016-05-31 | Cardiac resynchronization therapy utilizing P-wave sensing and dynamic anticipative left ventricular pacing | 2014-04-09 | A method for operating a pacemaker comprises the procedures of building a database of a cardiac cycle of a patient suffering from bundle branch block and artificially pacing a ventricle of the patient using the pacemaker according to anticipative atrioventricular (AV) delays in the database which are based on measured P-P intervals in the database. | 1. Method for operating a subcutaneous pacemaker, comprising the procedures of: building a database of a cardiac cycle of a patient suffering from bundle branch block, said database comprising time stamped entries indicating an onset of a P wave and a time duration of said P wave of said cardiac cycle and an onset of a QRS complex and a time duration of said QRS complex of said cardiac cycle immediately following said P wave, said database being built using at least two subcutaneous sensors for sensing said onset and said time duration of said P wave and said onset and said time duration of said QRS complex;determining atrioventricular (AV) delays of said cardiac cycle corresponding to measured P-P intervals of said cardiac cycle based on said time duration of said P wave and said onset of said QRS complex, said AV delays being stored in said database; andartificially pacing a first ventricle of said patient suffering from bundle branch block using said pacemaker according to said AV delays in said database, thereby enabling native activation of a second ventricle of said patient not suffering from bundle branch block and thereby also anticipating a correct time during said cardiac cycle to pace said first ventricle to maximize a likelihood of simultaneous activation of said first ventricle and said second ventricle. 2. The method according to claim 1, wherein at least one of said at least two sensors detects at least one P wave of said cardiac cycle and measures a time duration of said at least one P wave and wherein at least another one of said at least two sensors detects at least one QRS complex corresponding to said detected at least one P wave and measures a time duration of said at least one QRS complex. 3. The method according to claim 1, further comprising the procedure of determining said measured P-P intervals for each said P wave and a next detected P wave. 4. The method according to claim 1, wherein said procedure of building said database is executed over the course of a plurality of hours. 5. The method according to claim 1, wherein said procedure of building said database is executed using an electrocardiogram and computer software. 6. The method according to claim 1, further comprising the procedure of subtracting an offset measurement from each respective corresponding QRS complex detected in said cardiac cycle. 7. The method according to claim 6, wherein said offset measurement accounts for a delay in detecting each said respective corresponding QRS complex in said first ventricle in said patient. 8. The method according to claim 1, further comprising the procedure of adding a blanking period to said at least one sensor for detecting P waves. 9. The method according to claim 8, wherein said blanking period is programmable. 10. The method according to claim 8, wherein said blanking period substantially lasts an average duration of a T wave in said cardiac cycle. 11. The method according to claim 1, wherein said procedu... | NewPace Ltd. | A61N001/362 A61N001/368 A61N001/05 A61N001/365 A61N001/375 A61B005/0452
US10136842 | B2 | 2018-11-27 | Footwear apparatus with technique feedback | 2015-11-18 | A footwear apparatus includes a sensor in the footwear apparatus positioned to monitor an activity technique of the user when the footwear apparatus is joined with a foot of the user and a feedback generator is arranged to produce at least one of a haptic communication or an optical communication to the user in response to detecting a pattern in the activity technique. | 1. A footwear apparatus comprising: a sensor in the footwear apparatus positioned to monitor an activity technique of a user when the footwear apparatus is joined with a foot of the user, anda feedback generator arranged to produce a haptic communication to the user in response to detecting a pattern in the activity technique, the haptic communication arranged to cause at least some discomfort to the user, wherein the pattern includes a pronation element. 2. The footwear apparatus of claim 1, wherein the pattern is undesirable for the activity technique. 3. The footwear apparatus of claim 1, wherein the footwear apparatus comprises a shoe. 4. The footwear apparatus of claim 1, wherein the haptic communication is an electric shock. 5. The footwear apparatus of claim 1, wherein the haptic communication includes a mechanical mechanism arranged to come into contact with the foot. 6. The footwear apparatus of claim 1, wherein the feedback generator comprises an actuator capable of moving an impingement head into contact with the foot. 7. The footwear apparatus of claim 6, wherein the impingement head is one of multiple impingement heads. 8. The footwear apparatus of claim 7, wherein the multiple impingement heads are distributed throughout a support member of the footwear apparatus in a heel region of the support member, a toe region of the support member, a mid-foot region of the support member or combinations thereof. 9. The footwear apparatus of claim 6, wherein the impingement head is connected to an impingement body at least partially submerged in a thickness of a support member of the footwear apparatus, the impingement head having a longitudinal axis, the impingement body having a longitudinal axis, the impingement head having a larger lateral dimension than a lateral dimension of the impingement body. 10. The footwear apparatus of claim 9, wherein the impingement head is positioned such that the foot lands on the impingement head when the impingement head is caused to protrude beyond a foot engagement surface of the support member such that the impingement head one or more of pinches, stabs, and pricks the foot of the user. 11. The footwear apparatus of claim 1, wherein the pattern includes a toe striking element. 12. The footwear apparatus of claim 1, wherein the pattern includes a heel striking element. 13. The footwear apparatus of claim 1, wherein the pattern includes a mid-foot striking element. 14. A footwear apparatus, comprising: a sensor in the footwear apparatus positioned in one or more of an inner arch portion and an outer arch portion of the footwear apparatus to monitor pronation of a foot of a user during an activity technique of the user when the footwear apparatus is joined with the foot of the user; anda feedback generator arranged to produce a haptic communication to the user in response to detecting whether the user's foot has an inward or outward rolling motion;wherein the haptic communication includes a mecha... | ICON Health & Fitness, Inc. | A61B005/00 A61B005/11 A43B003/00 A61B005/103
US09327130 | B2 | 2016-05-03 | Implantable pacemakers control and optimization via fractional calculus approaches | 2014-04-14 | Method and system for non-linear modeling of physiological behavior, such as R-R intervals, in implantable devices, such as a rate responsive pacemakers, comprising a comprehensive modeling and optimization methodology based on fractional calculus and constrained finite horizon optimal control theory that allows for allows for fine-grain optimization of pacemaker response to heart rate variations; and the theoretical basis on which a hardware implementation of the fractional optimal controller that can respond to changes in the heart rate dynamics. Present invention describes a fractal approach to pacemaker control based on the constrained finite horizon optimal control problem. This is achieved by modeling the heart rate dynamics via fractional differential equations. Also, by using calculus of variations, the invention describes how the constrained finite horizon optimal control problem can be reduced to solving a linear system of equations. Finally, the invention describes the theoretical basis on which a hardware implementation become possible. | 1. A method for non-linear (fractional dynamics) modeling of physiological behavior measured by an implantable device, comprising the steps of: measuring the physiological behavior;determining magnitude of deviation between the measured physiological behavior and a reference value yref;modeling dynamics of the measured physiological behavior with continuous time fractional differential equations to identify parameters of a non-linear fractal model;selecting a constrained fractal optimal control problem;deriving optimality conditions for the constrained fractal optimal control problem when the magnitude of deviation exceeds a predetermined threshold;discretizing the optimality conditions for the constrained fractal optimal control problem; andsolving the constrained fractal optimal control problem corresponding to the optimality conditions to identify an optimal physiological behavior to operate the implantable device. 2. The method according to claim 1, wherein in the physiological behavior is an R-R interval. 3. The method according to claim 1, wherein the implantable device is a pacemaker. 4. The method according to claim 1, wherein the step of deriving optimality conditions for the constrained fractal optimal control problem comprises using wavelets and linear regression. 5. The method according to claim 1, wherein the step of selecting a constrained fractal optimal control problem comprises the step of selecting a finite time fractal optimal control with integral of squared tracking error criterion. 6. The method according to claim 1, wherein the step of selecting a constrained fractal optimal control problem comprises the step of selecting a finite time fractal optimal control with integral of squared time multiplied by squared tracking error criterion. 7. The method according to claim 1, wherein the step of determining magnitude of deviation between the measured physiological behavior and a reference value comprises analyzing the magnitude of a state variable y(t) minus the reference valve yref. 8. The method according to claim 1, wherein the state variable y(t) is an R-R interval. 9. The method according to claim 1, wherein the step of deriving optimality conditions for the constrained fractal optimal control problem when the magnitude of deviation exceeds a predetermined threshold, comprises the step of solving the following equations:minf⁡(t)⁢∫titf⁢{w2⁡[y⁡(t)-yref⁡(t)]2+z2⁢f2⁡(t)}⁢⁢ⅆtⅆα⁢y⁡(t)ⅆtα=a⁡(t)⁢y⁡(t)+b⁡(t)⁢f⁡(t),⁢y⁡(ti)=y0,y⁡(tf)=y0,0≤ymin≤y⁡(t)⁢≤max≤1fmin≤f⁡(t)⁢fmaxwhere y(t) represents the heart rate activity seen as a state variable, yref(t) denotes the reference value, f(t) is the pacing frequency, w and z are the weighting coefficients for the quadratic error and magnitude of the control signal, respectively, in the cost function, α is the exponent of the fractional order derivative characterizing the dynamics of the heart rate activity y(t), a(t) and b(t) are weighting coefficients for the heart activity and pacing f... | Carnegie Mellon University, a Pennsylvania Non-Profit Corporation | A61N001/00 A61N001/365 A61B005/024 G06F019/00 A61B005/00
US09681838 | B2 | 2017-06-20 | Monitoring a condition of a subject | 2016-01-13 | Apparatus for monitoring a subject is described. A sensor monitors the subject during a sleeping session of the subject, and generates a sensor signal in response thereto. A control unit analyzes the sensor signal. In response to analyzing the sensor signal, the control unit performs an action selected from the group consisting of: identifying that the subject is currently undergoing an apnea episode, and predicting that the subject is going to undergo an apnea episode. In response thereto, the control unit changes a position of at least a portion of a body of the subject by activating a device. Other applications are also described. | 1. Apparatus comprising: a sensor configured to monitor a subject during a sleeping session of the subject, and to generate a sensor signal in response thereto;a control unit configured to: analyze the sensor signal;in response to analyzing the sensor signal, perform an action selected from the group consisting of: identifying that the subject is currently undergoing an apnea episode, and predicting that the subject is going to undergo an apnea episode; andin response thereto, change a position of at least a portion of a body of the subject by activating a device. 2. The apparatus according to claim 1, wherein the control unit is configured to determine a posture of the subject by analyzing the sensor signal. 3. The apparatus according to claim 1, wherein the device includes an inflatable pillow, and wherein the control unit is configured to change the position of at least the portion of the subject's body by changing a level of inflation of the pillow. 4. The apparatus according to claim 1, wherein the device includes a bed, and wherein the control unit is configured to change the position of at least the portion of the subject's body by changing an angle of at least a portion of the bed. 5. The apparatus according to claim 1, wherein the device includes a mattress, and wherein the control unit is configured to change the position of at least the portion of the subject's body by changing an angle of at least a portion of the mattress. 6. The apparatus according to claim 1, wherein, in response to analyzing the sensor signal, the control unit is configured to identify that the subject is currently undergoing an apnea episode, and, in response thereto, to change the position of at least the portion of the subject's body by activating a device. 7. The apparatus according to claim 1, wherein, in response to analyzing the sensor signal, the control unit is configured to predict that the subject is going to undergo an apnea episode, and, in response thereto, to change the position of at least the portion of the subject's body by activating a device. | EARLYSENSE LTD. | A61B005/00 A61B005/04 A61B005/11 A61B005/0205 A61B005/024 A61B005/08 A61B005/113 A61B005/01 A61B005/02 A61F005/56
US10092236 | B2 | 2018-10-09 | Emergency medical services smart watch | 2016-01-05 | A system comprising: a wrist-worn device configured to be worn on the wrist of a rescuer performing cardiopulmonary resuscitation (CPR), the wrist-worn device including: one or more sensors coupled with the wrist-worn device, the one or more sensors being configured to sense one or more parameters indicative of a fatigue level of the rescuer; and a sensor interface configured to provide the sensed parameters to one or more external computing devices via an interface; and a wearable computing device configured to be worn by a rescuer, the wearable computing device including: a device interface for receiving information related to CPR; and a display for displaying an indication of the received information. | 1. A medical system for assisting a rescuer in resuscitation of a patient comprising: a head mounted device comprising a visual display;a chest compression sensor including one or more accelerometers configured to sense motion associated with cardio pulmonary resuscitation (CPR) chest compressions and to generate one or more signals indicative of chest motion;a wrist-worn device configured to be disposed on a wrist of the rescuer administering cardiopulmonary resuscitation (CPR) chest compressions, the wrist-worn device comprising one or more sensors configured to monitor at least one physical parameter of the rescuer;a patient monitor having at least one processor communicatively coupled to the chest compression sensor, the wrist-worn device, and the head mounted device, the at least one processor configured to: receive at least one of (i) the one or more signals indicative of the chest motion from the chest compression sensor and (ii) the at least one physical parameter of the rescuer from the wrist-worn device,analyze at least one of (i) the one or more signals indicative of the chest motion from the chest compression sensor and (ii) the at least one physical parameter of the rescuer from the wrist-worn device, anddetermine at least one of whether a quality of the chest compressions has deteriorated, and whether a physical status of the rescuer has declined, anddisplay on the head mounted device, visual feedback comprising an indication of at least one of the deterioration in the quality of the chest compressions, and the decline of the physical status of the rescuer. 2. The system of claim 1, wherein the patient monitor provides an indication to switch rescuers based on at least one of the deterioration in the quality of chest compressions, and the decline in the at least one physical parameter of the rescuer, wherein the deterioration in the quality of chest compressions is based on the determination of whether the at least one chest compression parameter is outside of an acceptable range. 3. The system of claim 1, wherein the head mounted device includes wearable glasses. 4. The system of claim 3, wherein the display is formed on at least one lens of the wearable glasses. 5. The system of claim 3, wherein the wearable glasses include a camera to capture the specific rescuer motions of the rescuer. 6. The system of claim 1, wherein the at least one physical parameter of the rescuer includes at least one of a heart rate, blood oxygen saturation, and a blood pressure of the rescuer. 7. The system of claim 1, wherein the patient monitor is part of an external defibrillator. 8. The system of claim 1, wherein the wrist-worn device comprises a memory configured to store a unique identifier associated with the wrist-worn device. 9. The system of claim 1, wherein the wrist-worn device is configured to power on when the wrist-worn device wraps around the wrist. 10. The system of claim 1, wherein the display of the head mounted device is conf... | ZOLL Medical Corporation | A61B005/00 A61H031/00 A61B005/0205 A61B005/021 A61B005/024 A61N001/39 G06F019/00 G16H040/63
US10548517 | B2 | 2020-02-04 | Spectroscopic apparatus for biological material and spectroscopic method using the same | 2015-11-06 | A spectroscopic apparatus and method for analyzing a biological material are provided. The spectroscopic apparatus may analyze a biological material which has an internal non-uniform tissue depending on a position thereof. The apparatus may include at least one detector configured to obtain respective detection spectrums corresponding to a plurality of measurement regions that are at mutually different positions of the biological material, and an information processor to determine whether the measurement regions are normal by mutually comparing the detection spectrums, or converting contribution degrees of data for a specific component of the biological material by differentiating the detection spectrums. | 1. A spectroscopic apparatus for analyzing a biological material having an internal non-uniform tissue depending on a position thereof, the spectroscopic apparatus comprising: a detector having a plurality of apertures, each of the plurality of apertures configured to obtain a respective detection spectrum of respective detection spectrums corresponding to a plurality of measurement regions that are disposed at mutually different positions of the biological material and correspond to each of the plurality of apertures, respectively; andan information processor configured to differentiate the respective detection spectrums and analyze the biological material based on the differentiating. 2. The spectroscopic apparatus of claim 1, wherein each of the plurality of measurement regions is formed of a multi-layered structure including two or more layers which are stacked, and each of the two or more layers is formed of a non-uniform thickness depending on the mutually different positions of the plurality of measurement regions. 3. The spectroscopic apparatus of claim 2, wherein the two or more layers forming each of the plurality of measurement regions differ from one another in contribution degrees of data for the respective detection spectrum of each of the plurality of measurement regions. 4. The spectroscopic apparatus of claim 1, wherein the information processor is further configured to determine whether at least one of the plurality of measurement regions is abnormal by mutually comparing the respective detection spectrums and determining whether at least one of the respective detection spectrums corresponding to the at least one of the plurality of measurement regions includes a noise signal or an abnormal signal. 5. The spectroscopic apparatus of claim 1, wherein the information processor is further configured to convert differential spectrums obtained by differentiating the respective detection spectrums into contribution degrees of data for a target analyte of the biological material. 6. The spectroscopic apparatus of claim 5, wherein the information processor is further configured to calculate a target measurement value for the target analyte of the biological material at an optional time using the contribution degrees of data for the target analyte of the biological material. 7. The spectroscopic apparatus of claim 1, wherein the detector is configured to obtain the respective detection spectrums using a Raman spectroscopy or a near infrared spectroscopy. 8. The spectroscopic apparatus of claim 1, wherein each of the plurality of apertures is further configured to be disposed adjacent to a surface of the biological material and receive a light signal from a measurement region among the plurality of measurement regions. 9. The spectroscopic apparatus of claim 1, wherein the biological material comprises a skin, and the information processor is further configured to measure a blood glucose in real time using the respective detection ... | SAMSUNG ELECTRONICS CO., LTD. IMEC VZW | A61B005/1455 A61B005/145
US10456061 | B2 | 2019-10-29 | Holding arrangement for a surgical access system | 2015-10-16 | A holding arrangement for a surgical access assembly is disclosed. One holding arrangement includes a holder member and a support member configured to engage with the holder member. The support member supports an outer sheath of a surgical assembly. The support member is compatible with a magnetic resonance imaging device. Another holding arrangement includes a holder member configured to hold a medical or diagnostic device such as a catheter. | 1. A holder member of a holding arrangement for a surgical access assembly comprising: one or more legs, each leg having a base connected thereto; anda pair of jaws,wherein the one or more legs, and/or each base of the one or more legs are selectively moveable, wherein the pair of jaws are selectively moveable with respect to each other;wherein each jaw within the pair of jaws comprises an inner edge and an outer edge,wherein each inner edge comprises a cut out extending inwardly from the inner edge,wherein the cut out of each inner edge oppose one another to collectively define a notch; andwherein the jaws are generally parallel to each other and define a channel, wherein the channel serves as a stop to retain a support member or a medical or diagnostic device within the notch. 2. The holder member of claim 1, wherein the one or more legs, and each base connected to the one or more legs, or the pair of jaws, are malleable. 3. The holder member of claim 2, wherein each base comprises a base opening for a fastener. 4. The holder member of claim 1, wherein the notch serves as a natural stop to retain a support member or a medical or diagnostic device in place. 5. The holder member of claim 4, wherein the support member or the medical or diagnostic device is compatible with a magnetic resonance imaging device. 6. The holder member of claim 1, further comprising a neck portion that is malleable, wherein the neck portion is positioned between the legs and a top portion that carries the jaws. 7. A holding arrangement for a surgical access assembly comprising: a holder member,a support member selectively engageable with the holder member,wherein at least the support member is compatible with a magnetic resonance imaging device;wherein the holder member further comprises:one or more legs, each of the one or more legs having a base, a neck portion, and a pair of jaws that are selectively moveable with respect each other, wherein the one or more legs, each base of the one or more legs, the neck portion, and/or the pair of jaws are malleable, and wherein each base is malleable to conform to a non-linear surface and comprises a base opening so that the holder member is attachable to the non-linear surface; andwherein the jaws are parallel to each other and define a channel, wherein the channel serves as a stop to retain the support member within a notch that is in communication with the channel. 8. The holding arrangement for the surgical access assembly of claim 7, wherein the support member is configured to support an outer sheath. 9. The holding arrangement of claim 8, wherein the holder member, the support member, the outer sheath, or a combination thereof, are formed from a non-ferrous material. 10. The holding arrangement of claim 9, wherein the holder member, the support member, or both, are formed from a polycarbonate, acrylonitrile butadiene styrene, a blend of polycarbonate and acrylonitrile butadiene styrene, polyolefins, polystyrene, bio... | Nico Corporation | A61B017/34 A61B005/055 A61B017/02 A61B050/20 A61B090/50 A61B017/00 A61B090/10 A61B090/11
US09865180 | B2 | 2018-01-09 | Heart phantom assembly | 2014-06-20 | A heart phantom assembly comprises a mechanical actuator system comprising a support operatively supporting at least two actuators and a fixing device. A phantom interface is operatively connected to the at least two actuators to be displaceable in at least one translational degree of freedom and at least one rotational degree of freedom as a function of actuation from the at least two actuators. A deformable heart phantom is connected at an end to the phantom interface, and at another end to the fixing device, to be subjected to compression/expansion in response to movements of the phantom interface in the translational degree of freedom, and to torsion in response to movements of the phantom interface in the rotational degree of freedom. | 1. A heart phantom assembly comprising: a mechanical actuator system comprising a support operatively supporting at least two actuators and a fixing device, and a phantom interface operatively connected to the at least two actuators to be displaceable in at least one translational degree of freedom and at least one rotational degree of freedom as a function of actuation from the at least two actuators; anda deformable heart phantom connected at an end to the phantom interface, and at another end to the fixing device, the deformable heart phantom made in a systolic condition to be subjected to compression/expansion in response to movements of the phantom interface in the translational degree of freedom, and to torsion in response to movements of the phantom interface in the rotational degree of freedom. 2. The heart phantom assembly according to claim 1, further comprising liquid injection means connected to the phantom interface and in fluid communication with the deformable heart phantom to inject liquid therein to inflate the deformable heart phantom. 3. The heart phantom assembly according to claim 2, further comprising a conduit in the phantom interface in fluid communication with an inner cavity of the deformable heart phantom and with the liquid injection means, for injection of fluid therethrough. 4. The heart phantom assembly according to claim 2, wherein the mechanical actuator system comprises three of the actuators, with one of the actuators configured to provide the one translational degree of freedom, another of the actuators configured to provide the one rotational degree of freedom, and another of the actuators configured to actuate the liquid injection means. 5. The heart phantom assembly according to claim 1, wherein the one of the actuators configured to provide the one translational degree of freedom is a linear actuator, with a direction thereof parallel to a longitudinal axis of the heart phantom. 6. The heart phantom assembly according to claim 1, wherein the one of the actuators configured to provide the one rotational degree of freedom is a linear actuator, with a direction thereof transverse to a longitudinal axis of the heart phantom, with a rack and pinion assembly operatively connecting the actuator to the phantom interface. 7. The heart phantom assembly according to claim 1, wherein the heart phantom has a flange at an end thereof configured for connection with the fixing device. 8. The heart phantom assembly according to claim 7, wherein the flange is a monolithic part of the heart phantom. 9. The heart phantom assembly according to claim 1, wherein the heart phantom has a ring-shaped part at an end thereof configured for connection with the fixing device. 10. The heart phantom assembly according to claim 9, wherein the ring-shaped part is made of a material having a greater hardness than that of a deformable body of the heart phantom. 11. The heart phantom assembly according to claim 9, wherein the ring-sh... | VAL-CHUM, LIMITED PARTNERSHIP | G09B023/28 G09B023/32 A61B006/00 A61B008/00 G09B023/34 A61B005/055 G01D018/00
US09726872 | B2 | 2017-08-08 | Three-dimensional optical coherence tomography apparatus and its application | 2015-08-03 | Provided herein are devices and systems that apply full-field optical coherence tomography (OCT) technology to three-dimensional skin tissue imaging. A special designed Mirau type objective and an optical microscope module allowing both OCT mode and orthogonal polarization spectral imaging (OPSI) mode are disclosed. | 1. A device comprising a light source module configured to provide a source light to the device, an optical microscope module, which handles the source light from the light source module and processes light signal; a Mirau type objective module, which handles light from the optical microscope module and processes light signal generated from a tissue optionally held by a tissue translation module holding said tissue sample; and a data processing unit for analyzing light signals from the tissue sample, wherein said Mirau type objective module comprises an interference objective immersed in a media having optical characteristics similar to the tissue sample to be analyzed, and wherein said optical microscope module comprises an optical switch configured to toggle between optical coherence tomography (OCT) mode and orthogonal polarization spectral imaging (OPSI) mode. 2. The device of claim 1, wherein the light source module comprises an amplified spontaneous emission light source, a superluminescent diode, a light emitting diode (LED), a broadband supercontinuum light source, a mode-locked laser, a tunable laser, a Fourier-domain Mode-locking light source, an optical parametric oscillator (OPO), a halogen lamp, a Ce3+:YAG crystal fiber, a Ti3+:Al2O3 crystal fiber, or a Cr4+:YAG crystal fiber. 3. The device of claim 2, wherein the light source module comprises a Ce3+:YAG crystal fiber or Ti3+:Al2O3 crystal fiber. 4. The device of claim 1, wherein said Mirau type objective module comprises an interference objective lens immersed in a media, a first glass plate, a second glass plate in a sealed container filled with one or more media. 5. The device of claim 4, wherein the one or more media comprises a first media and a second media. 6. The device of claim 1, wherein the optical characteristics is refractive index. 7. The device of claim 6, wherein the media has a refractive index in a range of about 1.2 to about 1.8. 8. The device of claim 7, wherein said media comprises water, silicone oil, or glycerol. 9. The device of claim 8, wherein said media comprises silicone oil. 10. The device of claim 1, wherein said first media comprises water and the second media comprises silicone oil. 11. The device of claim 1, wherein the optical switch is a quarter-wave plate where the fast axis of the quarter-wave plate is set to an angle of 45° allowing a portion of a signal light reflected or scattered from the sample passing through with polarization state unchanged. 12. The device of claim 1, wherein the optical switch is a quarter-wave plate where the fast axis of the quarter-wave plate is set to an angle of 0° or 90° allowing a portion of a signal light reflected or scattered from the sample passing through with polarization state rotated by 90°. 13. The device of claim 1, wherein the optical microscope module further comprises an objective lens, an optical long-wave-pass filter, and a polarization beam splitter. 14. The device of claim 1, wherein t... | NATIONAL TAIWAN UNIVERSITY | G02B021/00 A61B005/00 G02B005/30 G02B021/26 G02B021/02 G02B021/33
US20160121074 | A1 | 2016-05-05 | System with a Heart Rate Adjusting Mechanism | 2015-11-04 | A system for adjusting a heart rate includes an interface for communicating with a heart rate monitor and a speaker. The system also includes a processor and memory. The memory includes programmed instructions to cause the processor to determine a natural heart rate of a user, determine a target rate for the user, and cause a sound to be emitted from the speaker to adjust the natural heart rate to the target heart rate. | 1. A system for adjusting a heart rate of a user, comprising: a heart rate monitor;a transducer;an interface for communicating with the heart rate monitor and the transducer; and a processor and memory, the memory comprising programmed instructions to cause the processor to: determine a natural heart rate of the user;determine a target heart rate for the user; andcause a vibration to be emitted from the transducer to adjust the natural heart rate to the target heart rate. 2. The system of claim 1, wherein the transducer comprises a speaker in communication with the interface. 3. The system of claim 2, wherein the vibration comprises a sound. 4. The system of claim 1, wherein the target heart rate is configured to assist the user sleep. 5. The system of claim 1, wherein the vibration is adjusted over a time period. 6. The system of claim 5, wherein the vibration at a beginning of the time period is closer to the natural heart rate and the vibration at an end of the time period is closer to the target heart rate. 7. The system of claim 5, wherein the vibration is changed incrementally during the time period. 8. The system of claim 1, wherein the heart rate monitor is incorporated into a bed. 9. The system of claim 1, wherein the transducer is incorporated into a bed. 10. The system of claim 9, wherein the vibration is propagated through a medium of the bed. 11. The system of claim 1, wherein the heart rate monitor is in wireless communication with the processor. 12. The system of claim 1, wherein the target heart rate is determined based on a percentage reduction of a resting heart rate of the user. 13. A system for adjusting a heart rate, comprising: a communication interface;a heart monitor in communication with the communication interface;a speaker in communication with the communication interface; anda processor and memory, the memory comprising programmed instructions to cause the processor to: determine a natural heart rate of a user;determine a target heart rate configured to assist the user with sleeping; andcause a sound to be emitted from the speaker to adjust the natural heart rate to the target heart rate where the sound is adjusted over a time period such that the sound at a beginning of the time period is closer to the natural heart rate and the sound at an end of the time period is closer to the target heart rate. 14. The system of claim 13, wherein the sound is changed incrementally during the time period. 15. The system of claim 13, wherein the heart monitor is incorporated into a bed. 16. The system of claim 13, wherein the speaker is incorporated into a bed. 17. The system of claim 16, wherein the sound is propagated through a medium of the bed. 18. The system of claim 13, wherein the heart monitor is in wireless communication with the processor. 19. The system of claim 13, wherein the target heart rate is determined based on a percentage reduction of a resting heart rate of the user. 20. A system for adjusting... |  | A61M021/02 A61B005/00 A61B005/024
US20160120477 | A1 | 2016-05-05 | BIOLOGICAL-INFORMATION PROCESSING APPARATUS AND BIOLOGICAL-INFORMATION PROCESSING METHOD | 2016-01-08 | A biological-information processing apparatus includes a plurality of filters that detects a biological signal including a pulse wave component and a body motion noise component, detects a body motion signal correlated to the body motion noise component, and separates the pulse wave component and the body motion noise component included in the biological signal on the basis of the body motion signal. The filters have different filter characteristics. The biological-information processing apparatus calculates, for each of the filters, a correlation coefficient between the body motion signal and at least one of the pulse wave component and the body motion noise component and determines, on the basis of the correlation coefficient, as a pulse wave signal, the pulse wave component separated using any one filter. | 1. A biological-information processing apparatus comprising: a biological-signal detecting unit that detects a biological signal including a pulse wave component and a body motion noise component;a body-motion-signal detecting unit that detects a body motion signal; anda body-motion-noise removing unit that separates the pulse wave component and the body motion noise component from the biological signal on the basis of the body motion signal, whereinthe body-motion-noise removing unit includes: a plurality of filter units having different learning characteristics;a correlation-information calculating unit that calculates correlation information indicating a correlation degree between the body motion signal and output signals from the plurality of filter units; anda selecting unit that selects the output signal from the plurality of filter units on the basis of the correlation information. 2. The biological-information processing apparatus according to claim 1, wherein the learning characteristics include a step size for controlling a following characteristic to fluctuation in the body motion signal. 3. The biological-information processing apparatus according to claim 1, wherein the correlation-information calculating unit calculates, for each of the output signals from the filter units, the correlation information on the basis of the body motion signal, andthe selecting unit selects the output signal from the filter unit in which an absolute value of the correlation information is smallest. 4. The biological-information processing apparatus according to claim 3, wherein the output signals from the filter units are estimated pulse wave signals for estimating the pulse wave component. 5. The biological-information processing apparatus according to claim 1, wherein the correlation-information calculating unit calculates the correlation information on the basis of each of the output signals from the filter units and the body motion signal, andthe selecting unit selects the output signal from the filter unit in which an absolute value of the correlation information is largest. 6. The biological-information processing apparatus according to claim 5, wherein the output signals from the filter units are estimated body motion noise signals for estimating the body motion noise component. 7. The biological-information processing apparatus according to claims 1, wherein, when a difference between the correlation information of the filter unit that outputs the selected output signal and the correlation information of the other filter units exceeds a predetermined threshold, the selecting unit sets, as the learning characteristics of the other filters, the learning characteristics of the filter units that outputs the selected output signal. 8. The biological-information processing apparatus according to claims 1, wherein the body motion signal includes acceleration signals in one axial direction or at least two axial directions crossing each other, and... |  | A61B005/00 A61B005/024
US20160135697 | A1 | 2016-05-19 | PORTABLE ELECTRONIC HEMODYNAMIC SENSOR SYSTEMS | 2016-01-26 | Systems and methods are provided for extracting hemodynamic information, optionally employing portable electronic devices with optional User Interface (UI) features for system implementation. The systems and methods may be employed for acquiring hemodynamic signals and associated electrophysiological data and/or analyzing the former or both in combination to yield useful physiological indicia or results. Such hardware and software is advantageously used for non-invasively monitoring cardiac health. | 1-25. (canceled) 26. A system comprising: a vibration sensor adapted to capture a vibrational signal, and a computer processor,wherein the computer processor is adapted to resolve each of a pulse pressure waveform and an Embedded Frequency corresponding to sound of the heart from the vibrational signal, andwherein the processor is further adapted to, using the Embedded Frequency and the pulse pressure waveform, calculate at least one physiological parameter. 27. The system of claim 26, wherein the physiological parameter is selected from at least one of Dicrotic Notch (DN) position of the pulse pressure wave form, Ejection Fraction (EF) and systolic time intervals. 28. The system of claim 27, wherein the computer processer is further adapted to calculate Intrinsic Frequency (IF) parameters ω1 and ω2 on each side of the DN. 29. The system of claim 26, wherein the vibration sensor comprises a light source and a light sensor. 30. The system of claim 29, wherein the light sensor is an LED in a smartphone camera. 31. The system of claim 29, wherein the vibration sensor further comprising a membrane, the membrane made of a material selected to be at least partially reflective to the light source on an inner surface of the membrane. 32. The system of claim 31, wherein the membrane comprises metal or is metalized on the inner surface. 33. The system of claim 31, wherein the membrane material is selected to reduce light passing from an outer surface of the membrane to the sensor. 34. The system of claim 33, wherein the material substantially eliminates light passing from the outer surface. 35. The system of claim 33, wherein the membrane comprises metal or is metalized on at least one surface. 36. The system of claim 26, further comprising an electrocardiogram (ECG) sensor, wherein the processor is further adapted for producing an ECG signal. 37. The system of claim 36, wherein the computer processor is further adapted to calculate Ejection Fraction (EF) using the Embedded Frequency, the pressure waveform and the ECG signal. 38. A method comprising: positioning a vibration sensor on a subject's skin at a location peripheral to a the subject's heart,capturing a vibrational signal of skin motion,resolving the vibrational signal into each of a pressure waveform signal and an Embedded Frequency signal corresponding to sound of the heart, andcalculating with a computer processor, using the Embedded Frequency signal and the pulse pressure waveform signal, at least one physiological parameter 39. The method of claim 38, further comprising determining a Dichroitic Notch (DN) position within the pressure waveform signal using the Embedded Frequency signal with the computer processor. 40. The method of claim 38, further comprising calculating Intrinsic Frequency (IF) parameters ω1 and ω2 with the computer processor. 41. The method of claim 40, further comprising calculating Ejection Fraction (EF) with the computer processor using ω1 and ω2. 42. The... |  | A61B005/024 A61B007/02 A61B005/02 A61B005/00 A61B005/0408
US10062184 | B2 | 2018-08-28 | Medical image display control device, method of operation for same, and medical image display control program | 2016-01-27 | There are provided: a region extraction unit that extracts a plurality of regions from a three-dimensional image of a subject; an internal tissue information acquisition unit that sets, as a crossing region, a region that a light beam used when projecting the three-dimensional image onto the two-dimensional projection plane crosses first among the plurality of regions, sets, as a first intersection, a point on a region crossing the light beam first when the crossing region is excluded, sets, as a second intersection, a point crossing the crossing region when there is an extension in an opposite direction to a traveling direction of the light beam from the first intersection, and acquires information of an internal tissue included between the first and second intersections; and a display control unit that displays the internal tissue information so as to be superimposed on the three-dimensional image. | 1. A medical image display control device, comprising: a memory; anda processor that is connected to the memory, wherein the processor is configured to: extract a plurality of regions, which overlap each other in a viewing direction when projecting a three-dimensional image of a subject onto a two-dimensional projection plane, from the three-dimensional image of the subject;set, as a crossing region, a region that a light beam used when projecting the three-dimensional image of the subject onto the two-dimensional projection plane crosses first among the plurality of regions, set, as a first intersection, a point on a region crossing the light beam first when the crossing region is excluded, set, as a second intersection, a point crossing the crossing region when there is an extension in an opposite direction to a traveling direction of the light beam from the first intersection, and acquire internal tissue information of an internal tissue included between the first and second intersections;display the internal tissue information on a display unit so as to be superimposed on a part of the three-dimensional image corresponding to the crossing region projected onto the two-dimensional projection plane;extract at least three regions of the plurality of regions;acquire at least two pieces of the internal tissue information between at least the three regions: anddisplay the at least two pieces of the internal tissue information that are superimposed on different three dimensional images expressing different areas in different display modes. 2. The medical image display control device according to claim 1, wherein the processor is further configured to detect a specific tissue region from the internal tissue information, and wherein the processor is further configured to display the specific tissue region so as to be superimposed on the part of the three-dimensional image corresponding to the crossing region. 3. The medical image display control device according to claim 2, wherein the processor is further configured to detect the specific tissue region by detecting a value within a range set in advance from the internal tissue information. 4. The medical image display control device according to claim 1, wherein the processor is further configured to acquire a maximum value, a minimum value, a sum value, an average value, or a volume rendering value of pixels on the light beam between the first and second intersections as the internal tissue information. 5. The medical image display control device according to claim 1, wherein the processor is further configured to perform switching between display and non-display of a three-dimensional image of the plurality of regions. 6. The medical image display control device according to claim 1, wherein the processor is further configured to display the internal tissue information in a different display mode according to a position between the first and second intersections. 7. The medical image display ... | FUJIFILM CORPORATION | G06T011/00 A61B005/00 A61B006/00 G06T007/00 G06T015/10 G06T015/08 G06T019/00
US10206595 | B2 | 2019-02-19 | Intrauterine balloon apparatus, system, and method for augmenting uterine birthing forces during parturition | 2015-11-16 | An intrauterine balloon apparatus to augment uterine birthing forces and assist fetal descent during parturition is provided. The body of the balloon apparatus begins packaged in a compressed state minimizing volume, enabling delivery through the birth canal. The apparatus is advanced to a proximal uterine location and deployed for operation by introducing a pressurized agent through its conduit, causing the body to expand and apply directional forces towards the infant and birth canal. The balloon body is shaped to apply pressures dispersed towards the infant and the bidirectional conduit enables fluid conduction as the infant descends. After the infant has successfully descended and been delivered through the birth canal, pressure within the balloon body is relieved through the conduit thereby collapsing the volume of the balloon body and enabling refraction through the birth canal. | 1. A medical balloon apparatus configured to transition between a collapsed state and an expanded state, the medical balloon apparatus comprising: a conduit body having a proximal region, a distal region, and an internal lumen; anda balloon body coupled to the distal region of the conduit body and having an internal chamber in fluid communication with the internal lumen of the conduit body, the balloon body having a first wall region and a second wall region opposite the first wall region, and a third wall region between the first wall region and the second wall region, the third wall region being more elastic than the first wall region and the second wall region, the balloon body configured to expand to a shape wherein: the balloon body comprises an axis passing through the first wall region and the second wall region,during expansion of the balloon body, the first wall region is displaced relative to the axis in a first direction, and the second wall region is displaced relative to the axis in a second direction opposite the first direction; anda degree of displacement of the first wall region is greater than a degree of displacement of the second wall region. 2. The medical balloon apparatus of claim 1, wherein the third wall region has a thickness that is less than the thickness of the first wall region and the thickness of the second wall region. 3. The medical balloon apparatus of claim 1, wherein the third wall region is formed of a material having a modulus of elasticity less than the modulus of elasticity of the first wall region and the modulus of elasticity of the second wall region. 4. The medical balloon apparatus of claim 1, wherein the distal region of the conduit body has a longitudinal axis, and the first wall region extends away from the distal region of the conduit body, to one side of the longitudinal axis. | 3VO Medical, Inc. | A61B017/42 A61B005/04 A61B005/00 A61B005/06 A61B005/03 A61B017/00
US10820859 | B2 | 2020-11-03 | Systems and methods for detecting pulse wave velocity | 2015-10-15 | A non-invasive sensor unit adapted to be coupled to a patient includes a pair of light emitters spaced apart a known distance, and a pair light detectors. The light detectors detect light emitted from the emitters and scattered by a patient. The unit determines one or more cardiovascular characteristics of the patient from the scattered light, such as the patient's pulse wave velocity; a saturation of peripheral oxygen (SpO2) level; a temperature; a respiration rate; a heart rate; and a blood pressure. The light emitters emit light that may have wavelengths between 600 and 1000 nanometers. The unit, in some embodiments, is integrated into a patch adapted to be secured to the skin of the patient. Readings from the unit may be transmitted to a separate device spaced from the unit, such as via Bluetooth, WiFi, or by other means. | 1. A sensor unit adapted to be non-invasively coupled to a patient, the sensor unit comprising: a first light emitter;a plurality of first light detectors, the plurality of first light detectors adapted to detect a portion of first light that is emitted by the first light emitter and scattered by direct contact with the patient, the plurality of first light detectors including a first light detector spaced a first distance from the first light emitter and a second light detector spaced a second distance from the first light emitter, the second distance being different from the first distance;a second light emitter spaced from the first light emitter a known distance;a plurality of second light detectors, the plurality of second light detectors adapted to detect a portion of second light emitted by the second light emitter and scattered by direct contact with the patient, the plurality of second light detectors including a third light detector spaced the first distance from the second light emitter and a fourth light detector spaced the second distance from the second light emitter; anda controller in communication with the plurality of first light detectors and with the plurality of second light detectors, the controller adapted to automatically select the first and third light detectors or the second and the fourth light detectors and to use outputs from the selected light detectors, but not the non-selected light detectors, in combination with the known distance to determine a pulse wave velocity of the patient. 2. The sensor unit of claim 1 wherein the controller selects between the first and third light detectors or the second and fourth light detectors a based upon an orientation of the sensor unit relative to the patient's adjacent arterial structure. 3. The sensor unit of claim 1 further including a transmitter, the transmitter adapted to transmit the pulse wave velocity to at least one of the following: a cell phone, a computer, a patient support apparatus, and a local area network. 4. The sensor unit of claim 1 wherein the controller is adapted to repetitively determine the pulse wave velocity of the patient. 5. The sensor unit of claim 4 wherein the controller is further adapted to detect changes in the pulse wave velocity of the patient over time. 6. The sensor unit of claim 1 wherein the first light and second light have wavelengths between 600 and 1000 nanometers. 7. The sensor unit of claim 1 wherein the controller is further adapted to measure at least one of the following patient conditions: a saturation of peripheral oxygen (SpO2) level; a temperature; a respiration rate; a heart rate; and a blood pressure. 8. The sensor unit of claim 1 wherein the first and second light emitters and the pluralities of first and second light detectors are all integrated into a surface of one of a mattress and a cushion. 9. The sensor unit of claim 1 wherein the first light emitter, the plurality of first light detectors, the second light ... | Stryker Corporation | A61B005/00 A61B005/0205 A61B005/021 A61B005/1455 A61B005/024 A61B005/08
US09332937 | B2 | 2016-05-10 | Systems and methods for monitoring health and delivering drugs transdermally | 2013-09-25 | The present invention pertains to a system and method for transdermal sampling, comprising: at least one sampler for retrieving and transferring at least one analyte obtained transdermally from the skin of a subject; at least one detector system for identifying and quantifying said at least one analyte; and at least one logic module for (i) receiving and storing input data from said at least one detector, (ii) relating the input data to other data obtained from the subject, (iii) displaying output information, (iv) transmitting the output information to another system, and (v) controlling the operation of said at least one sampler and at least one detector. | 1. A transdermal monitoring and fluid delivery microsystem, comprising: a substrate including a power source, logic module and at least one connection receptacle;at least one removable transdermal monitoring subsystem mounted to the at least one connection receptacle, the at least one transdermal monitoring subsystem including at least one analyte detector and multiple single-use sampling units, each of the multiple single-use units including: a microablation heater for ablating a portion of the stratum corneum of a subject's skin in order to access interstitial fluid from the underlying epidermis; anda microchannel for transmitting the accessed interstitial fluid to the at least one analyte detector, wherein the at least one analyte detector communicates detection results to the logic module; andat least one removable microelectromechanical (MEMS) fluid delivery subsystem mounted to the at least one connection receptacle for delivering a fluid to the subject in response to the detection of the analyte by the at least one analyte detector, the MEMS fluid delivery subsystem including means for transdermal fluid delivery to the subject;wherein the logic module provides fluid delivery instructions to the at least one MEMS fluid delivery subsystem in accordance with the detection results from the at least one analyte detector. 2. The microsystem of claim 1, wherein the at least one analyte detector is configured to detect an analyte of interest selected from the group consisting of glucose, bilirubin, D-amino acids, an insecticide, atropine, pralidoxime, cytokine, dextromethorphan, caffeine, antihistamines, an organophosphate, microbial toxin, inflammatory sequeli to microbial toxin, spore metabolite, prealbumin, C-reactive protein, troponin I, estrogen, and testosterone. 3. The microsystem of claim 2, wherein the at least one analyte detector is configured to detect an amount of analyte of interest in the interstitial fluid of the subject and, further wherein, the logic module determines if the detected amount is at a predetermined level and, when the detected amount is below the predetermined level, the logic module instructs the MEMS fluid delivery subsystem to administer an analyte amount to the subject. 4. The microsystem of claim 1, wherein the fluid is selected from the group consisting of a drug, chemical and target probe. 5. The microsystem of claim 4, wherein the fluid further comprises a physiologically compatible solution. 6. The microsystem of claim 1, wherein the logic module controls operation of each of the multiple single-use sampling units and the MEMS fluid delivery subsystem on-demand. 7. The microsystem of claim 1, further comprising at least one environmental detector in communication with the logic module for sampling one or more target analytes in one of the air and water to which the subject is being exposed and communicating outward analyte detection results to the logic module, wherein the logic module provides fluid ... | Georgetown University Leidos, Inc. | A61M005/172 A61B005/145 A61B005/15 A61B005/00 A61B005/1455 A61B005/1486
US20160143564 | A1 | 2016-05-26 | MEASURING DEVICE AND MEASURING METHOD FOR NON-INVASIVE DETERMINATION OF THE D-GLUCOSE CONCENTRATION | 2014-06-23 | The invention relates to a measuring method and a measuring instrument for measuring raw data for determination of a blood parameter, in particular for non-invasive determination of the D-glucose concentration. The measuring device (1) comprises an excitation source (2) for generating electromagnetic radiation, a coupling arrangement (5-8) which is configured to couple in the radiation emitted by the excitation source (2) into a body surface of an object to be measured, and a sensor arrangement (13) which is configured to detect infrared (IR) radiation which is excited by the coupled-in radiation of the excitation source (2) in the body surface. The coupling arrangement (5-8) is configured to couple in the radiation emitted by the excitation source (2) extensively at a plurality of measuring points into the body surface, and the sensor arrangement (13) is configured to detect the IR radiation generated in the body surface at a plurality of measuring points. | 1-15. (canceled) 16. A measuring device adapted for measuring raw data for determining a blood parameter, comprising a) at least one excitation source for generating electromagnetic radiation,b) a coupling-in apparatus which is configured to couple the electromagnetic radiation emitted by the at least one excitation source into a body surface of a measurement object, andc) a sensor apparatus which is configured to detect an infrared (IR) radiation which is stimulated in the body surface by the electromagnetic radiation coupled-in to the body surface, whereind) the coupling-in apparatus is configured to couple into the body surface the electromagnetic radiation emitted by the at least one excitation source areally at a plurality of measurement positions, ande) the sensor apparatus is configured to detect the IR radiation generated in the body surface at a plurality of measurement positions. 17. The measuring device according to claim 16, wherein the at least one excitation source is a tuneable excitation source for generating electromagnetic radiation in at least one of the visible or IR range and the measuring device is configured to tune the at least one excitation source across a pre-determined spectral range during a measuring procedure. 18. The measuring device according to claim 16, wherein the coupling-in apparatus comprises a scanning unit which is configured to irradiate the plurality of measurement positions of the body surface time-sequentially. 19. The measuring device according to claim 17, wherein the sensor apparatus comprises an IR area sensor for detecting the IR radiation emitted at the plurality of measurement positions. 20. The measuring device according to claim 19, wherein the IR radiation generated at each measurement position of the plurality of measurement positions is imaged on a different region of the IR area sensor. 21. The measuring device according to claim 16, further comprising an evaluation unit for determining a blood parameter value depending on the IR radiation detected and on stored reference spectra. 22. The measuring device according to claim 19, wherein during tuning of the at least one excitation source across the pre-determined spectral range, the IR radiation detected is imaged for each of the plurality of measurement positions on a different region of the IR area sensor in order to measure an intensity sequence in a tuned spectral sequence. 23. The measuring device according to claim 21, wherein the evaluation unit is configured to identify, by at least one of comparison with reference spectra or mean value formation, those measurement positions which are suitable for determining a D-glucose concentration. 24. The measuring device according to claim 21, wherein the sensor apparatus comprises at least one of a spectrometer or an optical grating which is configured to image different wavelength ranges of the IR radiation generated in the body surface onto different columns of an IR area sensor, w... |  | A61B005/1455 A61B005/145
US20160120461 | A1 | 2016-05-05 | INFORMATION PROCESSING SYSTEM, PROGRAM, AND INFORMATION PROCESSING METHOD | 2014-04-17 | An information processing system 100 includes an information acquiring unit 110 that performs acquisition processing for information, a processing unit 120 that performs processing on the basis of the acquired information, and an output processing unit 130 that performs output processing for information created by the processing unit 120. The information acquiring unit 110 performs the acquisition processing for at least two kinds of information among staying time information, body information, and target information concerning a weight loss. The processing unit 120 creates advice information concerning the weight loss on the basis of a relational expression representing a relation between fluctuation in body weight or a body fat amount and the staying time information and the at least two kinds of information. The output processing unit 130 performs the output processing for the advice information. | 1. An information processing system comprising: an information acquiring unit that performs acquisition processing for information;a processing unit that performs processing on the basis of the information acquired by the information acquiring unit; andan output processing unit that performs output processing for the information created by the processing unit, whereinthe information acquiring unit performs the acquisition processing for at least two kinds of information among staying time information representing time in which a value represented by pulse wave information of a user is a value within a given fat burning zone, body information of the user, and target information concerning a weight loss of the user,the processing unit performs processing for creating advice information on the basis of a relational expression representing a relation between fluctuation in body weight or a body fat amount of the user and the staying time information and the at least two kinds of information acquired by the information acquiring unit, andthe output processing unit performs the output processing for the advice information created by the processing unit. 2. The information processing system according to claim 1, wherein, when a value corresponding to an integrated value of the staying time information in a given period is represented as T_zone, a fluctuation amount of the body weight or the body fat amount of the user in the given period is represented as W, and a fat burning coefficient is represented as Kfat, the relational expression is T_zone×Kfat=W. 3. The information processing system according to claim 1, wherein the information acquiring unit performs the acquisition processing for the staying time information and the body information, andthe processing unit performs processing for creating the advice information concerning an intake calorie of the user on the basis of the staying time information, the body information, and the relational expression. 4. The information processing system according to claim 3, wherein the information acquiring unit performs the acquisition processing for the staying time information up to given advice timing within an advice period and performs the acquisition processing for acquiring, as the body information, a measured value of the body weight or the body fact amount of the user up to the advice timing, andthe processing unit performs processing for calculating an estimated value of the body weight or the body fat amount at the advice timing corresponding to the staying time information on the basis of the staying time information and the relational expression and creating the advice information concerning the intake calorie on the basis of comparison processing of the calculated estimated value and the measured value. 5. The information processing system according to claim 1, wherein the information acquiring unit performs the acquisition processing for the staying time information and the target informatio... |  | A61B005/00 A61B005/0205 A61B005/107
US20160120447 | A1 | 2016-05-05 | SENSOR INSERTION METHOD, ANALYSIS METHOD, AND SENSOR INSERTION DEVICE | 2015-10-22 | A sensor insertion method for inserting a sensor into a user's body is provided. The method includes a unit fixing step of regulating a relative position between a sensor control unit and an inserting unit, where the sensor control unit is in a sensor insertion allowing state, and the inserting unit contains the sensor. In this step, the sensor control unit and the inserting unit are fixed to an insertion target portion of the user. Then, in a sensor inserting step, the inserting unit is caused to insert the sensor into the insertion target portion. In a sensor connecting step, the sensor control unit is shifted from the sensor insertion allowing state to a sensor connecting state, whereby electrical conduction is established between the sensor in the insertion target portion and the sensor control unit. | 1. A sensor insertion method for inserting a sensor into a user's body, the method comprising: a unit fixing step of regulating a relative position between a sensor control unit and an inserting unit, the sensor control unit being in a sensor insertion allowing state, the inserting unit containing the sensor therein, and of fixing the sensor control unit and the inserting unit to an insertion target portion of the user;a sensor inserting step of causing the inserting unit to insert the sensor into the insertion target portion; anda sensor connecting step of shifting the sensor control unit from the sensor insertion allowing state to a sensor connecting state, thereby establishing electrical conduction between the sensor in the insertion target portion and the sensor control unit. 2. The sensor insertion method according to claim 1, wherein, in the unit fixing step, after the inserting unit is fixed to the insertion target portion, the relative position of the sensor control unit with respect to the inserting unit is regulated and the sensor control unit is fixed to the insertion target portion. 3. The sensor insertion method according to claim 1, wherein, in the unit fixing step, after the sensor control unit is fixed to the insertion target portion, the relative position of the inserting unit with respect to the sensor control unit is regulated and the inserting unit is fixed to the insertion target portion. 4. The sensor insertion method according to claim 1, wherein, in the unit fixing step, after the relative position between the sensor control unit and the inserting unit is regulated, the sensor control unit and the inserting unit are fixed to the insertion target portion. 5. The sensor insertion method according to claim 4, wherein, in the unit fixing step, after the relative position between the sensor control unit and the inserting unit is regulated, the inserting unit is fixed to a unit base secured to the insertion target portion. 6. The sensor insertion method according to claim 4, wherein, in the unit fixing step, after the relative position between the sensor control unit and the inserting unit is regulated and the inserting unit is fixed to a unit base, the sensor control unit and the inserting unit are fixed to the insertion target portion by securing the unit base to the insertion target portion. 7. The sensor insertion method according to claim 1, wherein, in the sensor connecting step, the sensor and the sensor control unit are electrically connected to each other by moving a part or the whole of the sensor control unit. 8. The sensor insertion method according to claim 1, further comprising a unit releasing step of releasing the inserting unit from the insertion target portion after the sensor connecting step. 9. An analysis method comprising an analysis step of performing an analysis using the sensor after the sensor insertion method according to claim 1 is performed. 10. A sensor insertion device for inserting a sens... | ARKRAY, INC. | A61B005/145
US20160143539 | A1 | 2016-05-26 | Measuring probe, an apparatus and a method for label free attenuated reflection infrared spectroscopy | 2015-11-09 | Disclosed herein is a measuring probe, an apparatus, and a method for infrared spectroscopy. In some embodiments the measuring probe may have an elongated form with a first end for coupling and decoupling infrared light into and out of the measuring probe and a second end. In other embodiments, the measuring probe may comprise an attenuated total reflection (ATR) prism arranged at the second end of the measuring probe. The ATR prism may include at least a first surface having at least one measuring portion configured to be brought in optical contact with a measured object. The ATR prism may include at least a second surface having at least one reflective portion. In some embodiments, the ATR prism may include a cutting portion for cutting through the measured object. | 1. A measuring probe for infrared spectroscopy, the measuring probe having a first end and a second end, wherein the measuring probe comprises: an attenuated total reflection (ATR) prism arranged at the second end of the measuring probe, the ATR prism having at least a first surface having at least one measuring portion configured to be brought in optical contact with a measured object, at least a second surface having at least one reflective portion, and a cutting portion for cutting through the measured object, the cutting portion having a cutting tip with a cutting angle of equal to or less than a predetermined value,the ATR prism being configured such that: at least a portion of infrared light entering the measuring probe undergoes an attenuated total reflection at the at least one measuring portion of the first surface,at least a portion of the totally reflected light is reflected back towards the first surface by the at least one reflective portion of the second surface, and at least a portion of the light reflected back by the at least one reflective portion of the second surface undergoes an attenuated total reflection at the at least one measuring portion of the first surface and is decoupled from the ATR prism. 2. The measuring probe according to claim 1, wherein an angle between the first surface and a longitudinal axis of the measuring probe is equal to or less than 60°. 3. The measuring probe according to claim 1, wherein the cutting portion is formed by an intersection of the first surface and the second surface of the ATR prism. 4. The measuring probe according to claim 1, wherein a difference between a magnitude of an angle between the first surface and the second surface of the ATR prism and a magnitude of a critical total reflection angle is equal to or less than 12°. 5. The measuring probe according to claim 1, wherein the ATR prism is made of a material comprising at least one of: a diamond, a crystalline silicon, and a crystalline germanium. 6. The measuring probe according to claim 1, wherein at least one of the first surface and the second surface comprises at least one reflective portion configured to reflect incident light by means of non-attenuated total reflection towards the other surface of the ATR prism. 7. The measuring probe according to claim 1, wherein the first surface comprises a plurality of measuring portions. 8. The measuring probe according to claim 1, wherein the second surface comprises a transparent portion. 9. The measuring probe according to claim 1, wherein the ATR prism further comprises: a plurality of first surfaces, each first surface having at least one measuring portion configured to be brought in optical contact with a measured object. 10. The measuring probe according to claim 1, wherein the ATR prism further comprises: a first optical coupling surface for coupling infrared spectroscopic light into the ATR prism; anda second optical coupling surface for coupling light for a further op... |  | A61B005/00 G01J003/42 G01J003/02
US09326887 | B2 | 2016-05-03 | Methods and devices for treating conditions associated with autonomic dysfunction | 2014-11-10 | Methods and devices treating an autonomic nervous system associated disease condition in a subject are provided. Aspects of the invention include inducing one or more physiological response selected from the group consisting of sweating, gastric emptying, enhanced heart rate variability and enhanced quantitative sensory test responsiveness in a manner sufficient to modify the autonomic nervous system so as to treat the subject for the disease condition. The methods and devices find use in a variety of applications, e.g. in the treatment of subjects suffering from conditions arising from disorders of the autonomic nervous system. | 1. A device comprising: a physiological response stimulator configured to enhance one or more physiological responses of a subject having a condition arising from autonomic dysfunction, the physiological responses selected from the group consisting of: sweating,gastric emptying,heart rate variability,ability to rapidly transition from prone to suspended orientation, andquantitative sensory test responsiveness; anda controller configured to operate the device in a manner sufficient to modify the autonomic nervous system according to a predetermined protocol or dosing schedule so as to treat the subject for the condition. 2. The device according to claim 1, wherein the device is an implantable device. 3. The implantable device according to claim 2, wherein the implantable device is an implantable pulse generator. 4. The implantable device according to claim 2, wherein the implantable device is an implantable active agent delivery device. 5. The device according to claim 1, wherein the device is configured to be stably associated with a topical surface of the subject. 6. The device according to claim 5, wherein the device is configured to locally induce sweating in a subject. 7. The device according to claim 6, wherein the physiological response stimulator is a pharmacological delivery device. 8. The device according to claim 7, wherein the pharmacological delivery device is an iontophoretic delivery device. 9. The device according to claim 7, wherein the pharmacological delivery device administers a cholinergic agent. 10. The device according to claim 9, wherein the cholinergic agent is selected from the group consisting of: acetylcholine, pilocarpine and methacholine. 11. The device according to claim 1, wherein the physiological response stimulator is an electrical stimulator. 12. The device according to claim 1, wherein the device monitors the enhancement of the one or more physiological responses. | Palo Alto Investors | A61F007/00 A61M005/00 A61B005/145 A61B005/00 A61H023/00 A61K031/137 A61K031/166 A61K031/21 A61K031/4468 A61K031/454 A61K031/7048 A61N001/30 A61N001/36 A61N001/362 A61B005/024 A61B005/11
US09326685 | B2 | 2016-05-03 | Device for evaluating condition of skin or hair | 2012-09-14 | The present invention relates to a device for evaluating condition on skin or hair. The device provides one extension supporting a hydration meter suitable for measuring hydration values in areas difficult to measure (e.g., skin in scalp area) and; a separate extension supporting a camera for capturing images of an area of skin or hair and transferring said image to computer. Data from these sources are evaluated and scored. The multiple data points are used to provide a highly targeted product best suited for each individual consumer. | 1. A device, comprising: a housing, containing a hydration meter, a camera and an electrical wire or cord for transmitting data and images, wherein the housing has a first projection, a second projection and a third projection, and a central gripping portion,wherein the first projection, the second projection, and the third projection extend radially from the central gripping portion,wherein the first projection is configured to contain the hydration meter at the terminus of the projection, wherein the hydration meter comprises a surface evaluation area having an aspect ratio from 1:1 to 20:1,wherein the hydration meter is configured to measure skin and/or hair hydration and capture data corresponding to the measured hydration data,wherein the second projection is configured to contain the camera at the terminus of the projection, wherein the camera is configured to capture images,wherein the third projection is configured to comprise the electrical wire or cord that is configured to transmit the data from the hydration meter and the images captured by the camera to a computer. 2. The device according to claim 1, wherein the aspect ratio is greater than or equal to 2:1. 3. The device according to claim 2, wherein the aspect ratio is greater than or equal to 3:1. 4. The device according to claim 1, wherein the surface evaluation area is in the shape of an ellipse and the width of the ellipse defines the major axis and the height of the ellipse defines the minor axis. 5. The device according to claim 4, wherein the aspect ratio is greater than or equal to 2:1. 6. The device according to claim 1, wherein the surface evaluation area is oblong in shape. 7. The device according to claim 6, wherein the oblong shape is a rectangle with rounded corners. 8. The device according to claim 1, wherein the surface evaluation area is obround in shape. 9. The device according to claim 1, where the images and data are stored in the computer. 10. The device according to claim 1, wherein the camera is in communication with a printed circuit board. 11. The device according to claim 1, further comprising light emitting diodes. 12. The device according to claim 11, wherein the diodes are positioned on a ring supporting the diodes. 13. The device according to claim 12, wherein there is also a diffuser and the diodes face the diffuser towards an opening of the housing on which the camera is supported. 14. The device according to claim 1, wherein the device communicates with the computer via a USB port. 15. The device according to claim 1, wherein the hydration meter comprises a transducer and an interface circuit which together measure capacitance. 16. The device according to claim 15, wherein a change in capacitance is used to measure hydration. 17. A method for recommending a particular consumer product for skin or hair to a consumer, comprising: obtaining data corresponding to the hydration of skin and capturing an image of at least one of the scalp or... | Conopco, Inc. | A61B005/00 A61B005/053
US10028694 | B2 | 2018-07-24 | Non-invasive systems and methods to detect cortical spreading depression for the detection and assessment of brain injury and concussion | 2015-11-23 | The present invention provides systems and methods for detection and diagnosis of concussion and/or acute neurologic injury comprising a portable headwear-based electrode array and computerized control system to automatically and accurately detect cortical spreading depression and acute neurological injury-based peri-infarct depolarization (CSD/PID). The portable headwear-based electrode system is applied to a patient or athlete, and is capable of performing an assessment automatically and with minimal user input. The user display indicates the presence of CSD/PID, gauges its severity and location, and stores the information for future use by medical professionals. The systems and methods of the invention use an instrumented DC-coupled electrode/amplifier array which performs real-time data analysis using unique algorithms to produce a voltage intensity-map revealing the temporally propagating wave depressed voltage across the scalp that originates from a CSD/PID on the brain surface. | 1. A method of detecting concussion-based cortical spreading depression (CSD) non-invasively in a subject, the method comprising: using a portable, instrumented head-mountable detection unit comprising a plurality of scalp surface DC-electroencephalogram (DC-EEG) electrodes configured in a high density electrode array spaced between about 0.5 cm and about 1.5 cm apart from one another to detect a set of real-time scalp DC electrical potential signals non-invasively from a brain area of the subject and to transmit the real-time scalp DC electrical potential signals to an amplifier unit configured to amplify the detected electrical potential signals, said portable, instrumented head-mountable detection unit placed on a surface of a scalp area of the subject;using a computing device to detect the presence of the concussion-based CSD, said computing device comprised of a processing unit which includes one or more processors and a computer-readable memory containing programming instructions that, when executed, cause the processing unit to receive the amplified electrical potential signals to implement the programming instructions to determine that values of a first set of the electrical potential signals corresponding to a first portion of the brain area, representing a concussion-based CSD, are depressed as compared to values of a second set of the electrical potential signals corresponding to a second portion of the brain area, representing normal brain; andgenerating an output in response to the determining step, said output indicating that the subject is experiencing a concussion-based CSD. 2. The method of claim 1, wherein the plurality of scalp surface DC-EEG electrodes are spaced about 1.0 cm apart from one another. 3. The method of claim 1, further comprising: receiving a ground voltage from a ground electrode which is in contact with a non-brain area of the subject;using the received ground voltage to determine the primary DC voltage, Vi, for each of the plurality of scalp electrodes;using the primary DC voltage, Vi, to calculate a derived virtual ground voltage (Vg) for the plurality of scalp electrodes;determining a derived electrode voltage (Vig) based on the mean of the derived virtual ground voltages (Vg) for each of the scalp electrodes;determining whether the derived electrode voltages indicates a depression in a voltage level among a set of adjacent electrodes; andoutputting a report indicating a level of the depression and the area to which the depression corresponds. 4. The method of claim 1, wherein the computing device determines severity, morphology, velocity of propagation and trajectory of the concussion-based CSD, said severity proportional to the magnitude of DC voltage depression and area over which the DC voltage depression occurs, said morphology exhibited as an expanding ring, and said trajectory comprising an initiation point, an extinguish point, and a branching pattern. 5. The method of claim 1, wherein a velocity... | SCIENCEPLUSPLEASE, LLC | A61B005/04 A61B005/00 A61B005/0476
US09339246 | B2 | 2016-05-17 | Association of wireless detector with an imaging apparatus | 2010-10-07 | A method for directing image data to a dental treatment site wherein a sensor controller is positioned near a programmed radio-frequency identification device at the dental treatment site. A command entry configures the sensor controller to encode and transmit image data content for delivery to a receiving address, according to information obtained from the radio-frequency identification device. A digital sensor that is associated with the sensor controller is positioned in proximity to a subject. Image data is acquired from the digital sensor and transmitted from the sensor controller to a wireless access point, and from the wireless access point to a host computer at the receiving address. The acquired transmitted image data is stored in a computer-accessible electronic memory. | 1. A method for directing image data to a dental treatment site, comprising: positioning a sensor controller near a programmed radio-frequency identification device located at the dental treatment site;writing, into the radio-frequency identification device, information concerning an assignment of a wireless access point to the dental treatment site, the assignment being modifiable;providing a command that configures the sensor controller to encode and transmit image data content for delivery to a receiving address according to information obtained from the radio-frequency identification device, the obtained information associated with the dental treatment site and with the assigned wireless access point;positioning a digital sensor that is associated with the sensor controller in proximity to a subject;acquiring image data from the digital sensor and wirelessly transmitting the acquired image data from the sensor controller to the wireless access point;transmitting the data from the wireless access point to a host computer at the receiving address; andstoring the acquired transmitted image data in a computer-accessible electronic memory. 2. The method of claim 1, wherein the digital sensor is an x-ray sensor. 3. The method of claim 1, wherein the digital sensor is taken from the group consisting of an intra-oral camera, a digital camera, a video camera, and a sensor for an ultrasound imaging apparatus. 4. The method of claim 1, wherein providing the command to configure comprises pressing a control button on the controller. 5. The method of claim 1, wherein providing the command to configure the sensor controller comprises entering an instruction on a control monitor. 6. The method of claim 1, wherein providing the command to configure the sensor controller comprises moving the sensor controller to within a predetermined proximity to the programmed radio-frequency identification device. 7. The method of claim 1, further comprising: displaying the acquired image data. 8. A system for obtaining image data from a treatment site location within a network that has two or more wireless access points, comprising: a wireless identifier element modifiably encoded with information concerning the treatment site and with an assignment to a wireless access point;a wireless communication device energizable to interrogate the wireless identifier element and obtain the information concerning the treatment site and the assignment to the wireless access point;a controller associated with the wireless communication device and in communication with a detector element, the controller configured to obtain digital image data from the detector element, and to wirelessly transmit the obtained digital image data to the assigned wireless access point; anda host computer, comprising a processor, associated with the treatment site and in communication with the assigned wireless access point, that receives the transmitted digital image and stores the received digital ... | TROPHY | A61B006/14 A61B006/00 H04W004/00 A61B005/00 H04W004/04 H04W084/12
US09351641 | B2 | 2016-05-31 | Mobile processing device system for patient monitoring data acquisition | 2013-01-17 | A mobile processing device system for patient monitoring data acquisition includes a repository of information. The information associates a particular patient monitoring device type for displaying a particular patient parameter with a particular text label identifying the particular patient parameter. A portable processing device includes an imaging device for acquiring image data representing an image presenting patient parameter data from the particular patient monitoring device type. An image recognition processor, uses the information, for analyzing the image data to identify the particular text label identifying the particular patient parameter and a value of the particular patient parameter. An output processor communicates data representing the particular patient parameter and the value to a destination. | 1. A mobile processing device system for patient monitoring data acquisition, comprising: a repository of information associating a particular patient monitoring device type for displaying a particular patient parameter with a particular image type displayed by said device type and with a particular text label identifying said particular patient parameter in said particular image type;a portable processing device including an imaging device for acquiring image data representing an image presenting patient parameter data from the particular patient monitoring device type;an image recognition processor, using said information, for analyzing said image data to identify said particular text label identifying said particular patient parameter and a value of said particular patient parameter, and wherein said image recognition processor recognizes said particular image type of a plurality of different images types presented by said device type; andan output processor for communicating data representing said particular patient parameter and said value to a destination. 2. The system according to claim 1, wherein said repository, said image recognition processor and said output processor are within said portable processing device. 3. The system according to claim 1, wherein at least one of, said repository, said image recognition processor andsaid output processor are located remotely from said portable processing device. 4. The system according to claim 1, wherein said image recognition processor analyzes said image data to identifysaid particular text label using image data translation, rotation and scaling operations. 5. The system according to claim 1, wherein said repository of information further associates said particular patient monitoring device type with a mask identifying a particular area within said particular image containing said particular patient parameter andsaid image recognition processor analyzes said image data to identify said particular patient parameter by using said mask to identify said particular area within said particular image containing said particular patient parameter. 6. The system according to claim 5, wherein said image recognition processor analyzes said image data to recognize said patient parameter value in said particular area using at least one of, (a) an optical character recognition method and (b) one or more of image data translation, rotation and scaling operations. 7. The system according to claim 1, wherein said repository of information further associates said particular patient monitoring device type with a mask identifying a particular area within said particular image containing said particular text label andsaid image recognition processor analyzes said image data to identify said particular text label by using said mask to identify said particular area within said particular image containing said particular text label. 8. The system according to claim 7, wherein said image recognition processor... | Cerner Innovation, Inc. | H04N009/47 H04N007/18 A61B005/00 G06K009/32
US09355445 | B2 | 2016-05-31 | Breast cancer pathological image diagnosis support system, breast cancer pathological image diagnosis support method, and recording medium recording breast cancer pathological image diagnosis support program | 2008-02-20 | A breast cancer pathological image diagnosis support system for supporting a diagnosis of breast cancer based on a pathological image is provided. The breast cancer pathological image diagnosis support system includes an image obtaining unit which obtains an HE-stained image and an IHC image as pathological images to be diagnosed; an information obtaining unit which obtains information of a tumor area in the HE-stained image; a matching unit which calculates a matching position of the HE-stained image and the IHC image obtained by the image obtaining unit; a specifying unit which specifies a tumor area in the IHC image based on the information of the tumor area in the HE-stained image obtained by the information obtaining unit and information of the matching position calculated by the matching unit; and a calculating unit which calculates a staining positive cell content rate in the tumor area based on information of the tumor area in the IHC image specified by the specifying unit. | 1. A breast cancer pathological image diagnosis support device for supporting a diagnosis based on pathological images by using an HE-stained image and an IHC image as said pathological images of breast cancer, the device comprising: a computer processor; anda memory which stores a program implementing:a matching unit executed by the computer processor, which calculates a matching position of said HE-stained image and said IHC image;a specifying unit executed by the computer processor, which specifies, based on information of a tumor area specified in said HE-stained image and information of said matching position calculated by said matching unit, an area in the IHC image corresponding to the tumor area specified in the HE-stained image and specifies the specified area as a tumor area in the IHC image; anda calculating unit executed by the computer processor, which calculates a staining positive cell content rate in said tumor area based on information of the tumor area in said IHC image specified by said specifying unit,wherein said calculating unit converts image data of said tumor area of said IHC image into an HSV color space, calculates a threshold value by discrimination analysis of a value of H (hue) in the HSV color space of the image data, discriminates a brown area and a blue area in the tumor area of the IHC image on the basis of the calculated threshold value, extracts, in the tumor area of the IHC image, nucleus in the brown area to count the number of nuclei to obtain the number of staining positive cell nuclei, and extracts nucleus in the blue area to count the number of nuclei to obtain the number of staining negative cell nuclei, and calculates the staining positive cell content rate in said tumor area of the IHC image. 2. The breast cancer pathological image diagnosis support device as set forth in claim 1, wherein said matching unit binarizes a color scale of each pathological image of said HE-stained image and said IHC image, and calculates a position of the nucleus of said HE-stained image and said IHC image as the matching position, andsaid specifying unit matches the positions of the nucleus of said HE-stained image and said IHC image based on said matching position, and specifies the tumor area in said IHC image. 3. The breast cancer pathological image diagnosis support device as set forth in claim 1, wherein said matching unit calculates said matching position by phase-only correlation method. 4. The breast cancer pathological image diagnosis support device as set forth in claim 1, further comprising: an information obtaining unit which obtains information of said tumor area in said HE-stained image. 5. The breast cancer pathological image diagnosis support device as set forth in claim 1, further comprising: an image obtaining unit which obtains said HE-stained image and said IHC image as the pathological images to be diagnosed. 6. The breast cancer pathological image diagnosis support device as set forth in claim 1,... | NEC CORPORATION | G06K009/00 G06T007/00 A61B005/055 G01N001/30 G01N033/574
US09326694 | B2 | 2016-05-03 | Mobile terminal for measuring biological signal | 2009-01-08 | A mobile terminal for measuring a biological signal is disclosed. The mobile terminal for measuring a biological signal includes an electrocardiogram electrode unit having a first electrode provided in a front surface of a body of the mobile terminal and a second electrode provided in a first side surface of the mobile terminal body for measuring an electrocardiogram signal through the first electrode and the second electrode, a sensor unit provided in a second side surface of the mobile terminal body for measuring a pulse wave signal, a controller provided within the mobile terminal for analyzing the electrocardiogram signal received from the electrocardiogram electrode unit and the pulse wave signal received from the sensor unit and a radio frequency unit for transmitting an analyzed result by the control of the controller. | 1. A mobile terminal for measuring biological signals, comprising: an electrocardiogram electrode unit having a first electrode provided on a front surface of a body of the mobile terminal to enable contact with a first finger of a hand, and a second electrode and a third electrode each provided on a first side surface of the mobile terminal body to enable contact with an arm opposite to the hand and for measuring an electrocardiogram signal as a first biological signal through the first electrode, the second electrode, and the third electrode, wherein the second and third electrodes are positioned at the first side surface of the mobile terminal body;a reflective optical sensor unit having a light source for emitting light at tissue, and having a light receiver for receiving the emitted light reflected by the tissue, with the reflective optical sensor unit positioned at a second side surface on an opposite side of the mobile terminal body from the first side surface at which the second and third electrodes are positioned to enable contact with a second finger of the hand simultaneously with contact of the first electrode with the first finger and for optically measuring a pulse wave signal of the tissue from the received reflected light, and for generating a second biological signal from the pulse wave signal;a controller provided within the mobile terminal for analyzing the electrocardiogram signal received from the electrocardiogram electrode unit and the second biological signal received from the optical sensor unit to determine a pulse wave transmission time period; anda radio frequency (RF) unit for transmitting an analyzed result under the control of the controller,wherein at least one of the first electrode, second electrode, and third electrode is configured as a switch to trigger measuring the electrocardiogram signal and the pulse wave signal if one of the first electrode second electrode, and third electrode is touched by user. 2. The mobile terminal of claim 1, wherein the electrocardiogram electrode unit has the third electrode provided on the first side surface of the mobile terminal body adjacent to the second electrode, wherein the first, second, and third electrodes represent a positive (+) electrode, negative (−) electrode, and ground electrode, respectively. 3. The mobile terminal of claim 1, further comprising a display unit provided on the front surface of the mobile terminal body adjacent to the first electrode, wherein the controller controls the display unit to display the analyzed result. 4. The mobile terminal of claim 1, wherein a portion of the second side surface is substantially planer; and wherein the optical sensor unit comprises the light source and the light receiver presenting exterior planar surfaces which are substantially flush with attic substantially planar portion of the second side surface. 5. A mobile terminal for measuring biological signals, comprising: an electrocardiogram electrode unit having:a... | Samsung Electronics Co., Ltd. | A61B005/00 G06F019/00 A61B005/0404 A61B005/026 A61B005/024
US09339226 | B2 | 2016-05-17 | Systems and methods for joint replacement | 2011-01-21 | Systems and methods for joint replacement are provided. The systems and methods include a surgical orientation device, a reference sensor device, and at least one orthopedic fixture. The surgical orientation device, reference sensor device, and orthopedic fixtures can be used to locate the orientation of an axis in the body, to adjust an orientation of a cutting plane or planes along a bony surface, or otherwise to assist in an orthopedic procedure(s). | 1. A method of determining an anatomical feature during a knee procedure, comprising: coupling an orientation system with a distal aspect of a femur, the orientation system comprising a housing, a first orientation sensor and a second orientation sensor disposed within the housing, and a user interface operably coupled with the first orientation sensor and the second orientation sensor;wherein the first orientation sensor is selected from the group of sensor types consisting of: gyroscopic sensor, accelerometer sensor, tilt sensor, and magnetometer, and the second orientation sensor is selected from the group of sensor types consisting of: gyroscopic sensor, accelerometer sensor, tilt sensor, and magnetometer;wherein the first orientation sensor is a different sensor type than the second orientation sensor;estimating a first error associated with the first orientation sensor based in part upon the output of the second orientation sensor;estimating a second error associated with the second orientation sensor based in part upon the output of the first orientation sensor;moving the orientation system to collect data indicative of the anatomical feature relevant to the knee procedure; andreducing error from the data. 2. The method of claim 1, wherein estimating the second error includes estimating static bias in the second orientation sensor. 3. The method of claim 2, wherein estimating static bias further comprises maintaining the position of the orientation system for a minimum period. 4. The method of claim 1, further comprising, at the end of the moving step, returning the orientation system toward the position of the orientation system prior to the beginning of the moving step. 5. The method of claim 4, wherein the orientation system is configured to determine whether the user has returned the orientation system to a position that is approximately the same as the position of the orientation system prior to the beginning of the moving step. 6. The method of claim 5, wherein the orientation system comprises a filter with a minimum threshold to prevent a false stop indication. 7. The method of claim 1, wherein the moving step comprises moving in a loop configuration to collect data indicative of the anatomical feature relevant to the knee procedure. 8. The method of claim 7, wherein the moving step produces abduction of at least about 20 degrees while moving in the loop configuration. 9. The method of claim 1, wherein the first orientation sensor includes an accelerometer and estimating the first error includes estimating a gravity component of the accelerometer output. 10. The method of claim 9, further comprising determining, based upon a portion of the accelerometer output excluding the gravity component, the orientation of one or more of the housing and the femur. 11. The method of claim 1, further comprising: coupling an orthopedic orientation assembly with the femur, the orthopedic orientation assembly comprising a reference device ... | OrthAlign, Inc. | A61B005/103 A61B005/00 A61B017/15 A61B017/17 A61B019/00 A61B005/107
US10653335 | B2 | 2020-05-19 | Magnetic resonance imaging device and RF coil assembly with matching switch circuit and tuning switch circuit | 2014-06-23 | In a magnetic resonance imaging device provided with a deformable RF coil permanently mounted on a patient table, the RF coil can constantly maintain a matching state and a tuned state in a state of being flat on a top plate and in a state of being wound around a test object. Coil-side connectors 306-2 and 306-3 provided in a deformable RF coil 300 form a detector in cooperation with any one of fixture-side connectors 506-1 to 506-6, in which the detector detects a fitted state therebetween. A tuning circuit and a matching circuit of coil elements of the RF coil 300 are switched by an output of the detector to change tuning and matching parameters, and to maintain a matching state and a tuned state with respect to a state change indicating that the RF coil 300 is in a wound state. | 1. A magnetic resonance imaging device, comprising: a gantry that is provided with a static magnetic field generator that generates a static magnetic field and a gradient magnetic field generation coil that generates a gradient magnetic field;a table that is disposed in the gantry and on which a test object is placed;an RF coil that is configured so that a part thereof is fixed to the gantry or the table and at least another part thereof is formed of a flexible material and is deformable to be in contact with the test object, and receives a nuclear magnetic resonance signal generated from the test object;a fixture that fixes the RF coil to the test object; anda detector that detects the shape of the RF coil,wherein the RF coil includes a matching switch circuit and a tuning switch circuit configured to change a matching constant and a tuning constant of the RF coil according to an output of the detector and configured to be in an optimal matching state and an optimal tuned state when the test object is placed on the RF coil in a state where the RF coil is flatly mounted on the table,wherein the detector includes an RF coil-side connector disposed ancillary to the RF coil, a fixture-side connector disposed ancillary to the fixture, and a mechanism switch that turns on or off a connection between terminals when the RF coil-side connector and the fixture-side connector are connected to each other, andwherein the matching switch circuit and the tuning switch circuit have a circuit configuration that varies according to whether the mechanism switch turns the connection on or off, and respectively switch the matching switch circuit and the tuning switch circuit into the optimal matching state and the optimal tuned state in a shape of the RF coil corresponding to each of on the on state and the off state of the mechanism switch. 2. The magnetic resonance imaging device according to claim 1, wherein the fixture-side connector is provided in a plurality, andthe mechanism switch includes a switching button pressed when the RF coil-side connector is fitted to any one of the plurality of fixture-side connectors, and a switch that turns on or off the connection between the terminals by the switching button. 3. The magnetic resonance imaging device according to claim 1, wherein the RF coil is a RF coil used both as a transmission coil and a reception coil that emits a radio frequency magnetic field to the test object and receives the nuclear magnetic resonance signal. 4. The magnetic resonance imaging device according to claim 1, wherein the part of the RF coil fixed to the table is fixed in the table. 5. The magnetic resonance imaging device according to claim 1, wherein the fixture has an end thereof which is fixed to the table. 6. A magnetic resonance imaging device, comprising: a gantry that is provided with a static magnetic field generator that generates a static magnetic field and a gradient magnetic field generation coil that generates a gradient ... | Hitachi, Ltd. | A61B005/055 G01R033/34 G01R033/36
US09337900 | B2 | 2016-05-10 | Electronic device | 2012-11-12 | To improve the usability of an electronic device, the electronic device includes: a first communication unit located near a user; a receive unit configured to receive data through communication via a body of the user or near field communication between a second communication unit located in a member used by the user and the first communication unit; and a recording unit configured to record data pertaining to the member when communication is established between the first communication unit and the second communication unit. | 1. An electronic device comprising: an acceleration sensor that detects acceleration information of equipment used to perform a swing; an angular velocity sensor that detects angular velocity information of the equipment; and a communication unit configured to transmit the acceleration information and the angular velocity information for the swing at a predetermined interval when the communication unit receives a request from an external device every few seconds. 2. The electronic device according to claim 1, wherein when the communication unit receives the request from the external device after communication with the external device is established, the communication unit transmits the acceleration information and the angular velocity information for the swing after the communication is established. 3. The electronic device according to claim 1, wherein the acceleration sensor is a three-axis acceleration sensor, andthe angular velocity sensor is a three-axis angular velocity sensor. 4. The electronic device according to claim 1, further comprising an LED that emits light when the equipment is used. | NIKON CORPORATION | H04B005/00 H04B013/00 H04N005/232 G01S019/49 G01P001/12 A61B005/00 A61B005/11 H04N005/91 A61B005/0205 A61B005/024 H04N021/442
US09948881 | B2 | 2018-04-17 | Simultaneous display of two or more different sequentially processed images | 2016-01-19 | A medical imaging system having a processor with software executing thereon is provided for processing and display of multiple bandwidths of video in multiple display areas. The system receives a video signal with a plurality of portions and generates at least two signals there from. Each of the two signals has a bandwidth for display in a different display area. The two signals are updated so that each component displays a different portion of the input video signal, and the two signals may be combined for display on a single display device having two display areas. | 1. A medical imaging system comprising a signal processor operable or configured to: receive an incoming signal, the incoming signal including a plurality of images arranged in a sequence, each of the plurality of images having a plurality of color channels;alternately process, between a first signal processing mode and a second signal processing mode, sequential images of the plurality of images, thereby alternately generating first images into a video signal according to the first signal processing mode and second images into a video signal according to the second signal processing mode, wherein the first images are generated according to a first spectral bandwidth and the second images are generated according to a second spectral bandwidth, the first spectral bandwidth differing from the second spectral bandwidth; andalternately update an output signal as a video signal with the first images and the second images. 2. The medical imaging system of claim 1 wherein the first images comprise white light image signals and the second images comprise computed image signals. 3. The medical imaging system of claim 1, wherein the incoming signal is a video signal and the signal processor operable or configured to alternately update the output signal with the first images and the second images is operable or configured to update an output video signal with the first images at a first rate and update the output video signal with the second images at the first rate, wherein the first rate is half of the frame rate of the video signal. 4. The medical imaging system of claim 1, further comprising at least one monitor for displaying the output signal, the system configured such that one of the first images is displayed in a first display area at the same time as one of the second images is displayed in a second display area. 5. The medical imaging system of claim 4 wherein the first and second display areas are on a single monitor. 6. The medical imaging system of claim 1 further comprising: an interface;wherein selection of at least one of the first and second signal processing modes is received by a user through the interface. 7. The medical imaging system of claim 1 further comprising: a signal processor operable or configured to: receive an another incoming video signal including an another plurality of images arranged in a sequence, each of the another plurality of images having a plurality of color channels;alternately process sequential images of the another plurality of images, thereby alternately generating third images according to a third signal processing mode and fourth images according to a fourth signal processing mode; andalternately update an another output signal with the third images and the fourth images. 8. The medical imaging system of claim 7 wherein: one of the first images is displayed in a first display area;one of the second images is displayed in a second display area;one of the third images is displayed in a third display a... | Karl Storz Imaging, Inc. | H04N005/66 H04N005/445 H04N005/45 H04N005/262 A61B001/00 A61B001/06 A61B005/00 H04N009/12
US09668660 | B2 | 2017-06-06 | Active medical device for monitoring the status of a patient suffering from a risk of heart failure | 2016-01-15 | An active medical device is configured to receive inputs and to calculate a hemodynamic parameter representative of myocardium contractility determined from an endocardial acceleration signal. The microcontroller acquires heart rate and hemodynamic parameter pairs of values during a plurality of cardiac cycles. The microcontroller is configured to distribute the pairs of values into discrete bins to develop a profile for analysis. The microcontroller is configured to conduct an analysis comprising calculating an index representative of the patient's clinical status. The hemodynamic parameter representative of the myocardial contractility is a time interval separating the first and the second peak of endocardial acceleration. | 1. A computerized method for evaluating heart status of a patient, comprising: using a microcontroller coupled to heart rate and endocardial acceleration sensors to calculate a hemodynamic parameter;storing a plurality of heart rate and hemodynamic parameter pairs in bins, the bins grouped by heart rate, for both a long-term patient profile and a short-term patient profile;calculating a difference between the long-term profile and the short-term profile for each bin and multiplying each resultant difference by a weight for the bin;using the resulting products to calculate an index representative of the weighted difference between the long-term profile and the short-term profile; andgenerating an output from a device using a determination based on the index. 2. The computerized method of claim 1, wherein the hemodynamic parameter is the duration between two peaks of endocardial acceleration. 3. The computerized method of claim 1, wherein the hemodynamic parameter is the endocardial acceleration amplitude. 4. The computerized method of claim 1, wherein the determination based on the index comprises comparing the index to a threshold. 5. The computerized method of claim 4, wherein generating an output from the device comprises generating a graphical representation of a bar graph, wherein each heart rate bin forms a bar and each bar illustrates the long term profile data and the short term profile data for the bin. 6. The computerized method of claim 4, wherein generating an output comprises transmitting an alert from the device to a remote system. 7. The computerized method of claim 1, further comprising modelling the short-term profile using a linear regression model or a quadratic regression model. 8. The computerized method of claim 7, wherein parameters generated from the linear regression model or the quadratic regression model are used to calculate the index. 9. The computerized method of claim 1, wherein the bins grouped by heart rate are indicative of a discrimination between different activity phases of the patient. 10. The computerized method of claim 9, wherein the different activity phases of the patient include phases of effort, rest and recovery. 11. The computerized method of claim 9, further comprising defining triplets of values during different cardiac cycles, the triplets of values including a first value indicating an activity, as second value indicating the heart rate, and a third value indicating a time interval separating a first peak and a second peak of endocardial acceleration. 12. The computerized method of claim 11, further comprising developing a three dimensional graphical profile of the triplets, and placing the triplets in predetermined frequency classes of the three dimensional graphical profile. 13. The computerized method of claim 1, further comprising modifying a configuration of the device based on the index representative of the weighted difference between the long-term profile and the short-term pro... | SORIN CRM SAS | A61B005/0255 A61B005/00 A61B005/02 A61B005/024
US10463304 | B2 | 2019-11-05 | Integrated multimodal sensor device for intracranial neuromonitoring | 2014-06-12 | There is provided a monolithically integrated multimodal sensor device for intracranial neuromonitoring, the sensor device including: a single substrate; a temperature sensor formed on a first portion of the single substrate for detecting temperature; a pressure sensor formed on a second portion of the single substrate for detecting intracranial pressure; and an oxygen sensor formed on a third portion of the single substrate for detecting oxygen concentration. In particular, sensing portions of the temperature sensor, the oxygen sensor and the pressure sensor, respectively, are formed at different layers of the sensor device. There is also provided an integrated multimodal sensor system incorporating the sensor device and the associated methods of fabrication. | 1. A monolithically integrated multimodal sensor device for intracranial neuromonitoring, the sensor device comprising: a single substrate;a temperature sensor formed on a first portion of the single substrate for detecting temperature;a pressure sensor formed on a second portion of the single substrate for detecting intracranial pressure; andan oxygen sensor formed on a third portion of the single substrate for detecting oxygen concentration,wherein the temperature sensor, the pressure sensor, and the oxygen sensor are formed on the single substrate such that sensing portions of the temperature sensor, the oxygen sensor, and the pressure sensor, respectively, are formed at different layers of the sensor device as seen from a cross-sectional side view of the sensor device, andwherein the sensing portion of the temperature sensor is formed at a top of the single substrate as seen from the cross-sectional side view, the sensing portion of the pressure sensor being formed at intra-metal layers of the sensor device that exist between the top of the single substrate and a top of the sensor device as seen from the cross-sectional side view, and the sensing portion of the oxygen sensor being formed at a top metal layer of the sensor device that is above both the sensing portions of the temperature sensor and the pressure sensor as seen from the cross-sectional side view. 2. The sensor device according to claim 1, wherein the temperature sensor comprises one or more transistors, and the sensing portion of the temperature sensor includes an active region of the respective one or more transistors. 3. The sensor device according to claim 2, wherein the transistor is a parasitic vertical PNP transistor. 4. The sensor device according to claim 1, wherein the pressure sensor comprises one or more capacitive MEMS pressure sensors, and the sensing portion of the pressure sensor includes a movable electrode of the respective one or more capacitive MEMS pressure sensors. 5. The sensor device according to claim 1, wherein the oxygen sensor comprises a switchable array of oxygen sensor elements, and the sensing portion of the oxygen sensor comprises a working electrode of the respective oxygen sensor elements. 6. The sensor device according to claim 5, wherein the oxygen sensor is configured to switch to one of the oxygen sensor elements to provide the oxygen level reading which satisfies a predetermined sensitivity level. 7. The sensor device according to claim 1, further comprises one or more application-specific integrated circuits formed on the single substrate for processing readings from the temperature sensor, the oxygen sensor, and the oxygen sensor. 8. A method of fabricating a monolithically integrated multimodal sensor device for intracranial neuromonitoring, the method comprising: providing a single substrate;fabricating, on a first portion of the single substrate, a temperature sensor that detects temperature;fabricating, on a second portion of t... | AGENCY FOR SCIENCE, TECHNOLOGY AND RESEARCH SINGAPORE HEALTH SERVICES PTE. LTD. | A61B005/00 A61B005/01 A61B005/03 A61B005/145 A61B005/1486
US11235175 | B2 | 2022-02-01 | Automatic correlation modeling of an internal target | 2016-01-08 | A method and apparatus to automatically control the timing of an image acquisition by an imaging system in developing a correlation model of movement of a target within a patient. | 1. An apparatus, comprising: a data storage device; anda processing device, operatively coupled to the data storage device, the processing device to: automatically trigger image acquisition of a plurality of pretreatment images of a target, wherein the plurality of pretreatment images are internal to a patient, wherein acquisition of each of the plurality of pretreatment images is triggered at a different time; andgenerate a correlation model distinguishing between inspiration and expiration movements that maps the movements of an external marker to a target location of the target using the plurality of pretreatment images, wherein said automatically triggering comprises automatically triggering image acquisition at substantially evenly-distributed points of a periodic cycle, and wherein said automatically triggering each of the plurality of pretreatment images comprises: dividing historical data of movement of the target into a plurality of regions based on magnitude of movement of the target;establishing a plurality of model metric regions using the plurality of regions and inspiration and expiration movement in the historical data, wherein the plurality of model metric regions correspond to a plurality of phases of the periodic cycle;determining whether a current sample of the target movement is in a desired phase of the periodic cycle by categorizing the current sample into one of the plurality of model metric regions; andsending an imaging command to an imaging system to acquire the pretreatment image when the current sample is in the desired phase of the periodic cycle. 2. The apparatus of claim 1, wherein the processing device is to acquire a plurality of data points representative of a corresponding plurality of positions over time of the external marker associated with the patient, wherein the plurality of positions of the external marker define an external path of movement of the external marker, and the external path of movement defines the periodic cycle of the patient. 3. The apparatus of claim 2, wherein in generating the correlation model, the processing device is to: identify a path of movement of the target based on the plurality of data points and the plurality of pretreatment images at the substantially evenly-distributed points of the periodic cycle; anddevelop the correlation model using the path of movement of the target. 4. The apparatus of claim 1, wherein the processing device is further to: derive a target position of the target based on the correlation model;send a position signal associated with the target position to a beam generator controller; andcontrol the beam generator controller to direct a beam at the target. 5. The apparatus of claim 1, wherein the processing device is further to: periodically generate positional data about the target by automatically acquiring additional images of the target during treatment; andcontinuously generate positional data about external motion of the external marker during tre... | Accuray Incorporated | A61N005/10 A61B006/00 A61B005/11 A61B005/113 A61B005/00 A61B090/00 A61B034/10
US20160135782 | A1 | 2016-05-19 | FINGER JOINT ULTRASOUND IMAGING | 2014-11-14 | A method of acquiring ultrasound images for a rheumatoid arthritis examination comprises receiving a hand in a scanning assembly including a transducer array and a fluid providing an acoustic coupling between the transducer array and the hand and identifying locations of a plurality of finger joints of the hand while the hand is held stationary in the scanning assembly. Ultrasound images of the plurality of finger joints are acquired with the transducer array while the hand is held stationary in the scanning assembly, wherein the ultrasound images are of an area less than an entire area of the hand based on the identified locations of the finger joints of the hand. | 1. A method of acquiring ultrasound images for a hand examination, the method comprising: receiving a hand in a scanning assembly including a transducer array and a fluid providing an acoustic coupling between the transducer array and the hand;identifying locations of a plurality of finger joints of the hand while the hand is held stationary in the scanning assembly;acquiring ultrasound images of the plurality of finger joints with the transducer array while the hand is held stationary in the scanning assembly, the ultrasound images being of an area less than an entire area of the hand based on the identified locations of the finger joints of the hand. 2. The method of claim 1, wherein the transducer array is mounted to a movable support. 3. The method of claim 1, further comprising displaying at least one of the images of the finger joints. 4. The method of claim 1, wherein the ultrasound images comprise Power Doppler Imaging (PDI) images or High Res PDI images. 5. The method of claim 1, wherein identifying the locations of the plurality of finger joints comprises manually identifying the locations of the plurality of finger joints. 6. The method of claim 1, further comprising sensing a position of the hand, and wherein identifying locations of the plurality of finger joints comprises automatically identifying the locations of the plurality of finger joints based on the sensed position of the hand. 7. The method of claim 1, wherein identifying the locations of the plurality of finger joints comprises acquiring a b-mode ultrasound image of the hand and automatically identifying the positions of the plurality of finger joints based on the b-mode ultrasound image. 8. The method of claim 1 further comprising receiving signals representing an optical image of the hand, wherein identifying the locations of the plurality of finger joints comprises digitally analyzing the optical image. 9. The method of claim 1 further comprising: reconstructing a three-dimensional image of the hand; andautomatically adjusting positioning of the transducer array with respect to the hand based upon the three-dimensional image of the hand. 10. The method of claim 1 further comprising: presenting an image of at least a portion of the hand; andsuperimposing a representation of inflammation on the image based upon the acquired ultrasound images. 11. The method of claim 1, wherein the transducer array is large enough to scan the whole hand without moving the transducer array. 12. The method of claim 1, wherein the scanning assembly comprises a moveable support connected to the transducer array and an actuator connected to the transducer array, and further comprising controlling the actuator to automatically position the transducer array based on the identified locations of the finger joints. 13. The method of claim 12, wherein acquiring ultrasound images of the plurality of finger joints comprises capturing ultrasound images over a range of less than or equal to 3... |  | A61B008/08 A61B005/00 A61B008/00
US10105078 | B2 | 2018-10-23 | Method for managing health condition, health condition management apparatus, and health condition management system | 2015-10-07 | A method for causing a computer to execute a process for managing health condition, the process includes: detecting a grounding timing of a leg of a quadruped and a timing of weight-shift to a rear side of the quadruped; and determining whether or not the leg of the quadruped has a disease or injury, based on the detected grounding timing of the leg of the quadruped and the detected timing of a weight-shift to the rear side of the quadruped. | 1. A method for managing health condition, comprising: detecting, by a computer, a grounding timing of a leg of a quadruped by using a first sensor;detecting, by the computer, a timing of weight-shift to a rear side of the quadruped by using a second sensor which is mounted at a bottom position of a breast collar hung around a neck of the quadruped; anddetermining whether or not the leg of the quadruped has a disease or injury, based on the detected grounding timing of the leg of the quadruped and the detected timing of a weight-shift to the rear side of the quadruped. 2. The method for managing health condition according to claim 1, wherein the process includes detecting a peak acceleration in a vertically upward direction in a vertical direction component of an acceleration data, as the timing of the weight-shift to the rear side of the quadruped based on the acceleration data measured by an acceleration sensor corresponding to the second sensor. 3. The method for managing health condition according to claim 1, wherein the process includes detecting a peak acceleration in a vertically downward direction in a vertical direction component of an acceleration data as the timing of the grounding timing of the leg of the quadruped based on the acceleration data measured by an acceleration sensor mounted on the abdomen or chest of the quadruped and corresponding to the first sensor. 4. The method for managing health condition according to claim 3, wherein the quadruped is a horse, and further comprising: the process includes determining which leg of the right and left front legs of the quadruped is the leg being grounded based on information relating to a positional relationship in a front-back direction of a horse body between a right side leg and a left side leg of the horse, and an acceleration change in the vertical direction between peaks of acceleration in the vertically downward direction of the quadruped. 5. The method for managing health condition according to claim 1, wherein specifying a leg in which the grounding timing is detected immediately before the timing of the weight-shift to the rear side as a leg having a disease or injury. 6. The method for managing health condition according to claim 1, wherein the first sensor includes a right sensor mounted a right side of an abdomen of the quadruped and a left sensor mounted a left side of the abdomen of the quadruped. 7. A health condition management apparatus comprising: a first sensor mounted on a quadruped and configured to detect a grounding timing of a leg of a quadruped;a second sensor mounted at a bottom position of a breast collar hung around a neck of the quadruped and configured to detect a timing of weight-shift to a rear side of the quadruped;a processor configured to determine whether or not the leg of the quadruped has a disease or injury based on a detected grounding timing of the leg of the quadruped and a detected timing of weight-shift to the rear side of the quadrup... | FUJITSU LIMITED | A61B005/103 A61B005/00 A61B005/11
US20160128598 | A1 | 2016-05-12 | ELECTRONIC APPARATUS HAVING MEASUREMENT ELECTRODES | 2015-03-02 | A measurement device comprising includes a housing, an electronic unit included in the housing and configured to measure a biological state of a subject, and first and second electrodes disposed on a surface of the housing and are aligned along a first direction that is not orthogonal to one of first and second adjacent edges of the housing. | 1. A measurement device comprising: a housing;an electronic unit included in the housing and configured to measure a biological state of a subject; andfirst and second electrodes disposed on a surface of the housing and are aligned along a first direction that is not orthogonal to one of first and second adjacent edges of the housing. 2. The measurement device according to claim 1, wherein the surface of the housing has a first corner region, a second corner region, a third corner region, and a fourth corner region that is diagonal to the first corner region, andthe first electrode is located closest to the first corner region and the second electrode is located closest to the fourth corner region. 3. The measurement device according to claim 1, wherein each of the first and second electrodes has a first portion and a second portion that has a smaller outer perimeter than the first portion. 4. The measurement device according to claim 1, wherein the housing includes a flexible portion between a first less flexible portion of the housing in which the first electrode is disposed and a second less flexible portion of the housing in which the second electrode is disposed. 5. The measurement device according to claim 1, wherein the housing is formed of a flexible material. 6. The measurement device according to claim 1, wherein the housing includes a protruding portion between a first portion of the housing on which the first electrode disposed and a second portion of the housing on which the second electrode is disposed. 7. The measurement device according to claim 6, wherein the housing includes an output terminal on the protruding portion by which a measurement signal is transmitted to an external device. 8. The measurement device according to claim 7, further comprising: a battery, whereinthe housing further includes a battery terminal on the protruding portion by which power to charge the battery is received. 9. The measurement device according to claim 1, wherein the electronic unit is configured to generate an electrocardiogram of the subject based on a current flowing through the first and second electrodes. 10. The measurement device according to claim 9, further comprising: third and fourth electrodes for detecting a heart rate of the subject, that is disposed on a portion of the housing between the first and second electrodes. 11. The measurement device according to claim 1, further comprising: a substrate having a first surface on which the electronic unit is disposed and a second surface that is opposite to the first surface and with which the first and second electrodes are in electrical contact. 12. The measurement device according to claim 11, wherein the substrate is formed of a flexible material. 13. A measurement device comprising: a housing;an electronic unit included in the housing and configured to measure a biological state of a subject; andfirst and second electrodes disposed on a surface of the housing, whereina di... |  | A61B005/0408 A61B005/0245 A61B005/0404 A61B005/024
US09326737 | B2 | 2016-05-03 | Method and system for providing injury-based participation management for team activities | 2013-12-23 | An approach for providing diagnostic information regarding one or more impacts sustained by one or more users while the one or more users are participating in a team activity is described. A diagnostic platform detects one or more impacts sustained by a user, wherein the user is part of a team with one or more other users engaged in a team activity. The diagnostic platform generates a recommendation for a level of participation of the user, the one or more other users, or a combination thereof in the team activity based on the one or more impacts, a severity level of the one or more impacts, or a combination thereof. The diagnostic platform further presents the recommendation, a representation of the severity level, a representation of the level of participation, or a combination thereof in a user interface. | 1. A method comprising: detecting one or more impacts sustained by a user, wherein the user is part of a team with one or more other users engaged in a team activity;determining a substitution of the user by the one or more other users participating in the team activity based on the one or more impacts;generating a recommendation for a level of participation of the user in the team activity based on the one or more impacts, a severity level of the one or more impacts, or a combination thereof, wherein the recommendation includes the substitution; andpresenting the recommendation, a representation of the severity level, a representation of the level of participation, or a combination thereof in a user interface. 2. A method of claim 1, further comprising: monitoring the one or more impacts sustained by the user over a period of time, wherein the recommendation specifies the level of participation of at least a portion of the period of time. 3. A method of claim 2, wherein the period of time represents a duration of the team activity, a duration of a season of the team activity, a duration of a career of the team activity, or a combination thereof. 4. A method of claim 1, further comprising: monitoring biometric data, contextual data, or a combination thereof associated with the user, the one or more other users, the team, the team activity, or a combination thereof, wherein the recommendation is further based on the biometric data, the contextual data, or a combination thereof. 5. A method of claim 1, wherein the substitution includes disengaging the user from the team activity. 6. A method of claim 1, further comprising: calculating the severity level of a single one, a plurality, or an entirety of the one or more impacts,wherein the recommendation is further based on comparing the severity level against impact threshold criteria. 7. A method of claim 6, wherein the severity level is based on sensor information collected from equipment worn by the user, the one or more other users, or a combination thereof. 8. A method of claim 7, wherein the sensor information includes g-force data, vibration data, acceleration data, rotational data, or a combination thereof. 9. An apparatus comprising: at least one processor; andat least one memory including computer program code for one or more programs,the at least one memory and the computer program code configured to, with the at least one processor, cause the apparatus to perform at least the following, detect one or more impacts sustained by a user, wherein the user is part of a team with one or more other users engaged in a team activity,determine a substitution of the user by the one or more other users participating in the team activity based on the one or more impacts,generate a recommendation of a level of participation of the user in the team activity based on the one or more impacts, a severity level of the one or more impacts, or a combination thereof, wherein the recommendation includes t... | VERIZON PATENT AND LICENSING INC. | A61B005/00 G06F019/00 A61B005/11 A42B003/04
US20160128589 | A1 | 2016-05-12 | NERVOUS SYSTEM INTERFACE DEVICE | 2015-11-11 | A nervous system interface device can include a flexible active layer, a power coil, a communication module, and an antenna. More specifically, the flexible active layer can included a plurality of electrodes that are positioned to take neural measurements of the nervous system. The power coil can be electrically coupled to the flexible active layer and can be configured to receive wireless power from a power transfer device. The communication module can be electrically coupled to the flexible active layer and can be configured to receive power from the power coil. The antenna can also be adapted to wirelessly communicate neural measurement information to another device (e.g. an external computing device, receiver or the like). The nervous system interface device has two functional configurations which include a rolled and an unrolled configuration. In the rolled configuration the nervous system electrode device is rolled along a longitudinal axis for insertion into a hypodermic needle. In the unrolled configuration the nervous system electrode device is substantially flat. These nervous system electrode devices can dramatically improve access to brain function information, increase brain-machine interface functionality, and decrease foreign body response. | 1. A nervous system interface device, comprising: a flexible active layer including a plurality of electrodes that are positioned to take neural measurements of a nervous system;a power coil electrically coupled to the flexible active layer and configured to receive wireless power from a power transfer device;a communication module electrically coupled to the flexible active layer and configured to receive power from the power coil; andan antenna adapted to wirelessly communicate neural measurement information to another device;wherein the nervous system interface device has two functional configurations: a rolled configuration, wherein the electrode device is rolled along a longitudinal axis for insertion into a hypodermic needle; andan unrolled configuration, wherein the electrode device is substantially flat. 2. The device of claim 1, further comprising: a flexible support member encompassing the device and includes a dissolvable polymeric matrix for providing a selected structural integrity to the flexible support member for a selected period of time and which dissolves after insertion into the nervous system. 3. The device of claim 2, wherein the dissolvable polymeric matrix comprises at least one of a parilyne substrate and a polydimethylsiloxane (PDMS) substrate. 4. The device of claim 2, wherein the dissolvable polymeric matrix further includes at least one of growth hormones, dexamethasone, and minocycline. 5. The device of claim 1, further comprising a flexible polymeric battery that is: coupled to the power coil;configured to receive energy from the power coil; andconfigured to store the energy in one or more cells of the flexible polymeric battery. 6. The device of claim 5, wherein the energy is stored in two or more cells. 7. The device of claim 1, wherein at least one of a top surface of the flexible active layer, a bottom surface of the flexible active layer, and an outer surface of the communication module have a porous texture. 8. The device of claim 6, wherein the porous texture includes a plurality of holes from about 5 μm to about 15 μm in diameter. 9. The device of claim 1, wherein the flexible active layer has a coiled shape. 10. The device of claim 1, wherein the plurality of electrodes are arranged in an array along the coiled shape. 11. The device of claim 9, wherein the array of plurality of electrodes is about 30 μm in width and 30 μm in length. 12. The device of claim 1, wherein the plurality of electrodes are independently addressable. 13. The device of claim 1, wherein the plurality of electrodes are coupled to a top surface of the flexible active layer. 14. The device of claim 1, wherein the flexible active layer further includes distributed electronics. 15. The device of claim 14, wherein the distributed electronics interconnect the plurality of electrodes. 16. The device of claim 1, wherein the communications module and the flexible active layer each have a coiled shape which are coincident. 17. Th... |  | A61B005/04 A61B005/0478 A61B005/00
US09352160 | B2 | 2016-05-31 | Method and apparatus for phrenic stimulation detection | 2015-03-23 | Approaches for characterizing a phrenic stimulation threshold, a cardiac capture threshold, a maximum device parameter, and a minimum device parameter are described. A plurality of cardiac pacing pulses can be delivered by using a cardiac pacing device, a pacing parameter of the plurality of cardiac pacing pulses being changed between delivery of at least some of the pulses. One or more sensor signals can be evaluated to detect stimulation of the phrenic nerve by one or more of the plurality of cardiac pacing pluses. The evaluation of the one or more sensor signals and the pacing parameter can be compared to determine if a phrenic stimulation threshold is at least one of higher than a maximum device parameter and lower than a minimum device parameter. | 1. A method, comprising; delivering a pulse with a predetermined pulse energy;determining if the delivered pulse produced phrenic stimulation;if it is determined that the delivered pulse did not produce phrenic stimulation, performing a capture threshold test without performing a phrenic stimulation threshold test; andif it is determined that the delivered pulse did produce phrenic stimulation, performing a capture threshold test and a phrenic stimulation threshold test. 2. The method of claim 1, wherein the predetermined pulse energy corresponds to a cardiac pacing device capability limit. 3. The method of claim 1, wherein the predetermined pulse energy corresponds to a programmed cardiac pacing device limit. 4. The method of claim 1, wherein a cardiac pacing device includes a plurality of electrode combinations where each electrode combination defines a pacing vector, the method comprising performing the method for each of two or more pacing vectors. 5. The method of claim 4, further comprising: for each of the two or more pacing vectors, displaying on a display of a patient external device an indication of whether the pulse produced phrenic stimulation; andfor each of the two or more pacing vectors, displaying on the display of the patient external device an indication of a capture threshold as determined by the capture threshold test. 6. The method of claim 1, wherein the capture threshold test comprises: delivering a plurality of cardiac pacing pulses; andchanging a pacing parameter for one or more of the plurality of cardiac pacing pulses between delivery of at least some of the plurality of cardiac pacing pulses until a capture threshold is detected. 7. The method of claim 6, wherein the pacing parameter comprises an amplitude and/or width of a cardiac pacing pulse, and wherein the pacing parameter is reduced between delivery of at least some of the plurality of cardiac pacing pulses until the capture threshold is detected. 8. The method of claim 6, wherein the phrenic stimulation threshold test comprises: delivering a plurality of cardiac pacing pulses;changing a pacing parameter for one or more of the plurality of cardiac pacing pulses between delivery of at least some of the plurality of cardiac pacing pulses until the phrenic stimulation threshold is detected. 9. The method of claim 8, wherein at least some of the plurality of cardiac pacing pulses are part of both the capture threshold test and the phrenic stimulation threshold test. 10. The method of claim 1, wherein the phrenic stimulation threshold test comprises: delivering a plurality of cardiac pacing pulses; andchanging a pacing parameter for one or more of the plurality of cardiac pacing pulses between delivery of at least some of the plurality of cardiac pacing pulses until a phrenic stimulation threshold is detected. 11. The method of claim 10, wherein the pacing parameter comprises an amplitude and/or width of a cardiac pacing pulse, and wherein the pacing paramet... | Cardiac Pacemakers, Inc. | A61N001/18 A61N001/37 A61B005/00 A61N001/36 A61N001/365 A61B005/04 A61B005/042 A61B005/0488 A61B005/11 A61N001/368 A61N001/372
US09339066 | B2 | 2016-05-17 | Personal cooling and heating system | 2015-09-04 | Embodiments of the invention provide a cooling garment including a garment material configured to be worn on a user's upper body. The garment material includes an outer surface with gel-filled chambers, an inner surface, and a hole therethrough. The garment material is configured to compress against the user's body when worn so that protrusions extending from the inner surface contact the user's body and define air channels between the protrusions, the inner surface, and the user's body. The cooling garment also includes a fixing ring coupled to the garment material around the hole, and a suction device configured to be coupled to the fixing ring, wherein the suction device is configured to circulate air through the air channels. | 1. A cooling upper torso garment comprising: a garment material configured to be worn on a user's upper body, the garment material including an outer surface, an inner surface, and a plurality of holes therethrough;the outer surface comprising a plurality of chambers;gel material within the plurality of chambers;protrusions extending from the inner surface, wherein the garment material is configured to compress against the user's body when worn so that the protrusions contact the user's body and define air channels between the protrusions, the inner surface, and the user's body;a plurality of fixing rings coupled to the garment material around each of the plurality of holes; anda portable suction device configured to be detachably coupled to one of the plurality of fixing rings; the portable suction device configured to circulate air through the air channels when the cooling upper torso garment is donned. 2. The cooling upper torso garment of claim 1 wherein the suction device is detachable from the fixing ring. 3. The cooling upper torso garment of claim 2 and further comprising a cover configured to be coupled to the fixing ring when the suction device is removed. 4. The cooling upper torso garment of claim 1 wherein the suction device comprises a fan mechanism including a fan, a power source, and a controller. 5. The cooling upper torso garment of claim 4 wherein the fan mechanism includes a piezoelectric power generator configured to generate power from the user's movements while wearing the cooling garment. 6. The cooling upper torso garment of claim 1 wherein the garment material is a plastic material. 7. The cooling upper torso garment of claim 1 and further comprising reflectors coupled to the outer surface and configured to reflect at least one of light and heat away from the outer surface. 8. The cooling upper torso garment of claim 1 and further comprising tightening straps coupled to the outer surface and configured to tighten the cooling garment against the user's body. 9. The cooling upper torso garment of claim 1 and further comprising an air inlet extending through the garment material from the outer surface to the inner surface and a hose coupling the air inlet to the portable suction device. 10. The cooling upper torso garment of claim 1 and further comprising a sensor configured to sense one of a temperature and a pulse of the user. 11. The cooling upper torso garment of claim 10 wherein the portable suction device comprises a fan mechanism including a fan, a power source, and a controller, and the sensor is in communication with the controller. 12. The cooling upper torso garment of claim 10 and further comprising a Bluetooth transmitter in communication with the sensor and configured to wirelessly transmit sensor information to a smartphone. 13. A cooling upper torso garment comprising: a garment material configured to be worn on an user's upper body, the garment material including an outer surface, an inner surfa... |  | A41D013/002 A61F007/10 A61F007/00 A61B005/024 A61B005/00 A61B005/0205
US10588522 | B2 | 2020-03-17 | Biological sensor | 2016-02-02 | A biological sensor capable of improving the signal to noise ratio of a detection signal obtained by a light-receiving element and amplified by an amplifier is provided.;The biological sensor includes a microcontroller that generates a driving signal, a light-emitting element that emits light in accordance with the driving signal, a light-receiving element that outputs a current detection signal based on an intensity of received light, and an amplifying circuit that converts the current detection signal into a voltage detection signal, amplifies an alternating current component of the voltage detection signal, and outputs an amplified detection signal. Furthermore, the microcontroller generates an offset signal that is applied to an offset circuit to offset the direct current component of the voltage detection signal and to obtain biological information by processing the amplified detection signal. | 1. A biological sensor comprising: a light-emitting element that emits light based on a driving signal;a light-receiving element that outputs a current detection signal based on an intensity of received light including light emitted by the light-emitting element;an amplifying circuit configured to convert the current detection signal into a voltage detection signal, amplify an alternating current component of the voltage detection signal, and output an amplified detection signal; anda microcontroller configured to: generate the driving signal as a pulse-form driving signal,generate an offset signal as a pulse-form offset voltage that is synchronized with the pulse-form driving signal,apply the pulse-form offset voltage to an offset circuit coupled to the amplifying circuit to offset a direct current component of the voltage detection signal, andobtain biological information based on the amplified detection signal. 2. The biological sensor according to claim 1, wherein the amplifying circuit comprises: a first operational amplifier that converts the current detection signal into the voltage detection signal;a second operational amplifier that amplifies the voltage detection signal to output the amplified detection signal; anda capacitor connected between an output terminal of the first operational amplifier and an input terminal of the second operational amplifier. 3. The biological sensor according to claim 2, wherein the offset circuit is coupled to a path between the output terminal of the first operational amplifier and the input terminal of the second operational amplifier. 4. The biological sensor according to claim 3, wherein the offset circuit is coupled to a path between the output terminal of the first operational amplifier and the capacitor. 5. The biological sensor according to claim 1, wherein the offset circuit comprises a voltage dividing circuit that includes a plurality of resistors and that generates the offset voltage by voltage-dividing the offset signal generated by the microcontroller. 6. The biological sensor according to claim 1, further comprising a light-emitting element configured to output light of a different wave length from the light-emitting element, wherein the microcontroller is further configured to generate pulse-form driving signals having mutually different timings for each of the plurality of light-emitting elements, and to generate pulse-form offset voltages independent from each other in synchronization with the respective pulse-form driving signals outputted at mutually different timings. 7. The biological sensor according to claim 1, wherein the microcontroller is further configured to: extract a direct current component of the amplified detection signal, anddetermine the offset voltage outputted by the offset circuit so that a deviation of the extracted direct current component from a target value decreases. 8. The biological sensor according to claim 7, wherein the microcontroller stops generatin... | MURATA MANUFACTURING CO., LTD. | A61B005/024 A61B005/00 A61B005/1455
US10786178 | B2 | 2020-09-29 | Breathing tube for use in ultrasonic flow measurement systems | 2015-07-24 | This invention relates to a breathing tube for use in ultrasonic flow measurement systems for determining the volume flow and/or the molar mass of the respiration of humans and animals. According to the invention, the breathing tube at least partly has a polygonal cross-section. Furthermore at least one indicator is formed on the breathing tube, which can be read out via an external optical device. | 1. A breathing tube for use in an ultrasonic flow measurement system for determining a volume flow and/or a molar mass of respiration of humans and animals, wherein the breathing tube at least partly has a polygonal transversal cross-section extending along a longitudinal axis of the breathing tube, and the breathing tube has at least a pair of beveled outer sides opposing one another and extending along the longitudinal axis that permit insertion of the breathing tube into the measurement system when the breathing tube is in a predetermined orientation in relation to the measurement system and that prevent insertion of the breathing tube into the measurement system if the breathing tube is rotated 180 degrees about the longitudinal axis relative to the predetermined orientation, and wherein on the breathing tube at least one indicator is formed, which is readable out via an external optical device, wherein an outer surface of the breathing tube comprises a comb-like shape to form the at least one indicator in a region of at least one edge of the breathing tube, wherein openings in the breathing tube, which serve for passing through ultrasonic pulses, are arranged on the pair of beveled outer sides. 2. The breathing tube according to claim 1, wherein the at least one indicator is formed by areas on the outer surface of the breathing tube that differ from a remainder of the outer surface by a color and/or a degree of reflection. 3. The breathing tube according to claim 1, wherein the polygonal transversal cross-section comprises a substantially trapezoidal interior cross-section. 4. The breathing tube according to claim 1, wherein the polygonal transversal cross-section comprises a substantially trapezoidal cross-section. 5. The breathing tube according to claim 4, wherein the breathing tube has at least one circumferential sealing lip on the outer surface of the breathing tube. 6. The breathing tube according to claim 5, wherein the outer surface of the breathing tube has a curved shape such that in cooperation with the measurement system, into which the breathing tube is insertable, the at least one circumferential sealing lip effects secure sealing due to a resulting uniform force acting onto an outside of the at least one sealing lip. 7. The breathing tube according to claim 6, wherein the openings in the breathing tube are closed via a fabric-like net. 8. The breathing tube according to claim 7, wherein the fabric-like net is made from a same injectable plastic material as the breathing tube as to be injectable from the same injectable plastic material during an injection-molding operation for producing the breathing tube. 9. The breathing tube according to claim 7, wherein the fabric-like net is made of a material different from the breathing tube. 10. The breathing tube according to claim 9, wherein the breathing tube includes a mouthpiece to be plugged on separately. 11. The breathing tube according to claim 10, wherein the breat... | NDD MEDIZINTECHNIK AG | A61B005/087 G01F001/66 A61B005/08
US09339188 | B2 | 2016-05-17 | Methods from monitoring health, wellness and fitness with feedback | 2013-08-06 | A method is provided for obtaining monitored information about an individual to create life activity data. Individual information is detected or measured. The individual information is selected from of at least one of, an individual's activities, behaviors and habit information, and an individual's health. A monitoring device is used to detect or measure the individual information. The monitoring device includes ID circuitry with an ID storage that contains a unique individual ID, a communication system which reads and transmits the unique individual ID from the ID storage, a power source and a pathway system. The individual information is received from the monitoring device at a telemetry system that includes a database. The individual information is analyzed using life activity data with one or more analysis tools at the telemetry system. Life activity data is created for the individual. | 1. A method for obtaining monitored information about an individual and create life activity data, comprising: using a monitoring device that includes a microphone, an RF transmitter and sensors to determine air quality, sound level, sound quality, light quality, ambient temperature near the individual, a power source, a pathway system, ID circuitry with an ID storage that contains a unique individual ID, the RF transmitter serving as a communication system which reads and transmits the unique individual ID from the ID storage;using at least an accelerometer or a movement detection device configured to detect movement of the individual, the accelerator and the monitoring device configured to calculate or derive sleep onset and wake time, sleep interruptions, and a quality and depth of sleep of the individual, the microphone configured to record an individual's movement sounds detected by the accelerometer or the movement detection device, the accelerometer or the movement detection device configured to cause the microphone to stop recording the individual's movement sounds when the movement sounds are not directed to a sleep related parameter;receiving the individual monitored information from the monitoring device at a telemetry system that includes a database, the telemetry system in operation receiving the individual monitored information that includes the movement of the individual, the movement sounds of the individual, sleep information and sleep related parameter information data of the individual, the telemetry system or the monitoring device at least partially analyzing the individual's monitored information to calculate or derive the sleep onset and the wake time, the sleep interruptions, and the quality and the depth of the sleep of the individual;analyzing the individual sleep information and sleep related parameter information using the life activity data of one or more analysis tools at the telemetry system; andcreating life sleep information activity data and life sleep related parameter information activity data for the individual. 2. The method of claim 1, further comprising: creating caloric burn data for the individual, wherein the life activity data is information relating to caloric intake for the individual. 3. The method of claim 1, further comprising: receiving an individual survey from the individual at the telemetry system. 4. The method of claim 1, further comprising: providing feedback to the individual from the telemetry system. 5. The method of claim 1, further comprising: providing feedback to the individual. 6. The method of either claim 1, further comprising: providing feedback to the individual in graphical form. 7. The method of claim 1, further comprising: generating one or more web pages containing feedback for the individual. 8. The method of claim 7, further comprising: including an analysis element in the one or more analysis tools. 9. The method of claim 8, further comprising: including in the data... | James Proud Hello Inc. | G08B023/00 A61B005/00 H02J007/02 H02J017/00 A61B005/11 H04W004/00 H02J007/00 A61B005/0205 G01C022/00 H04L029/08
US09351687 | B2 | 2016-05-31 | Intravascular devices, systems, and methods | 2013-06-28 | Intravascular devices, systems, and methods are disclosed. In some embodiments, the intravascular devices include at least one electronic, optical, or electro-optical component positioned within a distal portion of the device and one or more connectors positioned at a distal portion of the device. In some instances, the connectors are flexible coils, such as a ribbon coil, formed of a conductive material. In some particular instances, the conductive coil is embedded within a polymer tubing. Further, in some embodiments the electronic, optical, or electro-optical component is positioned within a flexible element at the distal portion of the device. In some instances the flexible element is a coil. Methods of making and/or assembling such intravascular devices/systems are also provided. | 1. An intravascular device, comprising: a first flexible element;a second flexible element coupled to the first flexible element in a position proximal to the first flexible element;a third flexible element coupled to the second flexible element in a position proximal to the second flexible element;a distal core extending within the first flexible element;a mounting structure positioned within the second flexible element and fixedly secured to the distal core;at least one sensing component mounted to the mounting structure;a proximal core fixedly attached to the mounting structure and extending proximally from the mounting structure through the second and third flexible elements; andat least one conductor having a proximal section and a distal section, wherein the distal section of the at least one conductor is coupled to the at least one sensing component and the proximal section of the at least one conductor is coupled to at least one connector. 2. The intravascular device of claim 1, wherein the at least one connector comprises a coil. 3. The intravascular device of claim 2, wherein a wire forming the coil has a rectangular cross-sectional profile. 4. The intravascular device of claim 2, wherein a wire forming the coil has a circular cross-sectional profile. 5. The intravascular device of claim 2, wherein a wire forming the coil has a semi-circular cross-sectional profile with a rounded portion of the semi-circular cross-sectional profile extending outwardly. 6. The intravascular device of claim 2, wherein a wire forming the coil has a semi-circular cross-sectional profile with a rounded portion of the semi-circular cross-sectional profile extending inwardly. 7. The intravascular device of claim 2, wherein the proximal section of the at least one conductor is soldered to an inner portion of the coil. 8. The intravascular device of claim 2, wherein the coil is at least partially embedded within a polymer tubing. 9. The intravascular device of claim 1, wherein the second flexible element comprises a ribbon coil. 10. The intravascular device of claim 9, wherein the ribbon coil is embedded in a polymer tubing, the polymer tubing having a thickness between about 0.0005″ and about 0.003″. 11. The intravascular device of claim 9, wherein a section of the proximal core positioned within the ribbon coil is formed of nitinol. 12. The intravascular device of claim 11, wherein an insulating layer is positioned between the ribbon coil and the section of the proximal core positioned within the ribbon coil. 13. The intravascular device of claim 12, wherein the insulating layer is a polymer in which the ribbon coil is embedded. 14. The intravascular device of claim 1, wherein the proximal core includes a first section that is fixedly attached to the mounting structure and a second section extending proximally from the first section, wherein the first section is formed of a first material and the second section is formed of a second material diffe... | Volcano Corporation | A61B005/02 A61B005/00 A61B008/12 A61B005/0215 A61M025/00 A61M025/09
US09345411 | B2 | 2016-05-24 | Devices and methods for monitoring cerebral hemodynamic conditions | 2011-10-03 | Devices, and methods for measuring bioimpedance signals are disclosed. One aspect may include a headset apparatus including a retainer and electrodes. The retainer may be configured to position the electrodes on the head of a subject so as to obtain bioimpedance signals indicative of hemodynamic conditions associated with an MCA territory. A processor may be included to measure and analyze the obtained bioimpedance signals, and to output information for predicting hemodynamic conditions associated with an MCA territory. | 1. A headset apparatus, comprising: a retainer;a first electrode configured to connect to the retainer at a first position;a second electrode configured to connect to the retainer at a second position;the retainer comprising a substantially inflexible portion shaped to fit a temporal region of a head of a subject, the temporal region being adjacent to an ear of the subject, the portion being shaped to maintain, when the retainer is oriented on the head of the subject: a position of the first electrode anterior to a middle cerebral artery (MCA) of the subject and opposing an anterior horn of a lateral cerebral ventricle, anda position of the second electrode posterior to the MCA, above the ear of the subject, and opposing a posterior horn of the lateral cerebral ventricle; andat least one electrical conductor enabling the first electrode and the second electrode to be connected in a circuit, such that when the first electrode is in the first position and the second electrode is in the second position and the first and second electrodes are connected to a circuit through which current flows, the retainer, the first electrode, and the second electrode cooperate to enable current to flow from one of the first and second electrodes on a first side of the MCA through the lateral cerebral ventricle and to another of the first and second electrodes on a second side of the MCA, opposite the first side, wherein the retainer is adapted to maintain relative positions of the first electrode and the second electrode with respect to one another. 2. The headset apparatus of claim 1, wherein the portion maintains a spacing of between 70 mm and 100 mm between the first electrode and the second electrode. 3. The headset apparatus of claim 1, wherein the portion maintains a spacing of between 75 and 95 mm between the first electrode and the second electrode. 4. The headset apparatus of claim 1, wherein the portion positions the first and second electrodes according to a sensitivity map. 5. The headset apparatus of claim 1, wherein the portion positions the first electrode to be glued to exposed skin on a forehead of the subject and positions the second electrode to be glued to exposed skin above the ear of the subject. 6. The headset apparatus of claim 1, wherein the portion positions the first and second electrodes such that the first and second electrodes are not directly above external blood vessels or muscled areas. 7. The headset apparatus of claim 1, wherein the electrodes are oriented to enable current to pass from the first electrode, through the lateral cerebral ventricle and to the second electrode. 8. The headset apparatus of claim 1, further comprising a third electrode connected to the retainer at a third position and a fourth electrode connected to the retainer at a fourth position, such that when the retainer is oriented on the head of the subject, the third electrode is located on a side of the head opposite to the first electrode, and the fou... | Orsan Medical Technologies, Ltd. | A61B005/04 A61B005/00 A61B005/0265 A61B005/0295 A61B005/053 A61B005/02 A61B005/05
US09332918 | B1 | 2016-05-10 | Patient monitoring devices and systems | 2013-05-06 | In one embodiment, the invention relates to systems, methods, and apparatus relating to the detection of a neuropathy such as a periopertive neuropathy. In one embodiment, a wristband comprising a plurality of anodes and cathodes is used. The wristband can be a component in a electrode array that includes a plurality of reference or recording electrodes. The electrode array can be configured to stimulate and collect responsive signals from an ulnar, a median, radial and posterior tibial nerve. The simulation and signal collection can be performed on a continuous basis for time periods of interest such as a given perioperative time period using a monitoring device. | 1. A processor-based method of detecting a neuropathy in a patient using a nerve monitoring system comprising: providing a wristband comprising: a first pair of electrodes,a second pair of electrodes,a third pair of electrodes, andan elongate flexible substrate, said first, second, and third electrode pairs disposed in or on the flexible substrate; anda first electrical lead having an electrode contacting end and a monitoring device contacting end, the electrode contacting end in electrical communication with at least one electrode in the first pair of electrodes, a second electrical lead having an electrode contacting end and a monitoring device contacting end, the electrode contacting end in electrical communication with at least one electrode in the second pair of electrodes, a third electrical lead having an electrode contacting end and a monitoring device contacting end, the electrode contacting end in electrical communication with at least one electrode in the third pair of electrodes;placing said wristband over a wrist of said patient such that said first electrode pair is disposed over a first nerve in the wrist, said second electrode pair is disposed over a second nerve in the wrist, and said third electrode pair is disposed over a third nerve in the wrist;electrically stimulating the first, second, and third nerves using the first, second and third electrode pairs, and noninvasively monitoring an action potential in each of the first nerve, second nerve, and third nerve passing through the brachial plexus using a reference electrode;detecting a deviation relative to a baseline signal with respect to a responsive signal generated by one or more of the first, second, and third nerves following the electric stimulation using a processor and comparator connected to said first, second and third electrode pairs;comparing the deviation to a predetermined threshold using a processor and comparator; andgenerating an alert indicative of the neuropathy when the deviation exceeds a predetermined threshold. 2. The method of claim 1, wherein each electrode in the first pair is positioned above a median nerve. 3. The method of claim 2, wherein each electrode in the second pair is positioned above a radial nerve. 4. The method of claim 3, wherein each electrode in the third pair is positioned above an ulnar nerve. 5. The method of claim 1, wherein the flexible substrate has one or more demarcations configured to identify a boundary between one or more nerves or bones disposed relative to nerves. 6. The method of claim 1, further comprising a monitoring device in electrical communication with the monitoring device contacting end of the first, second and third electrical leads, wherein the monitoring device stimulates one or more electrodes in the first, second, and third electrode pairs and monitors responsive signals from one or more of a radial, ulnar or median nerve. 7. The method of claim 6, wherein the monitoring device comprises: a housing,o... |  | A61B005/05 A61B005/04 A61B005/00 A61N001/05 A61N001/04
US09339197 | B2 | 2016-05-17 | Intravascular implantable medical device introduction | 2012-05-29 | In one example, this disclosure is directed to a kit for intravascular implantation of an implantable medical device within a patient, the kit comprising an elongated outer sheath forming an inner lumen with a distal opening, the outer sheath sized to traverse a vasculature of the patient, and an elongated inner sheath with an inflatable member at its distal portion. The inflatable member is inflatable from a proximal end of the inner sheath to close-off the distal opening of the outer sheath when inflated. The inner sheath further includes a stopper proximally located relative to the inflatable member. The inflatable member is remotely controllable from a proximal end of the inner sheath to retract in a proximal direction towards the stopper. The inflatable member can be retracted in a proximal direction towards the stopper and past an implantable medical device positioned within a distal portion of the outer sheath. | 1. A kit for intravascular implantation of an implantable medical device within a patient, the kit comprising: an elongated outer sheath forming an inner lumen with a distal opening, the outer sheath sized to traverse a vasculature of the patient; andan elongated inner sheath with an inflatable member at a distal portion of the inner sheath, wherein the inflatable member is selectively inflatable from a proximal end of the inner sheath, wherein the inflatable member is configured to substantially close-off the distal opening of the outer sheath when inflated, wherein the inner sheath is slidable within the inner lumen of the outer sheath,wherein the inner sheath further includes a stopper proximally located relative to the inflatable member, wherein the inflatable member is remotely controllable from a proximal end of the inner sheath to retract in a proximal direction towards the stopper,wherein the kit is configured such that the inflatable member can be retracted in a proximal direction towards the stopper and past an implantable medical device positioned within a distal portion of the outer sheath. 2. The kit of claim 1, wherein the stopper is remotely controllable from a proximal end of the inner sheath to push the implantable medical device from the distal portion of the outer sheath when the inflatable member has been retracted towards the stopper to a position proximal to the implantable medical device within the inner lumen of the outer sheath. 3. The kit of claim 1, wherein the inner sheath further includes a tether that is remotely controllable from a proximal end of the inner sheath to release the implantable medical device from the distal portion of the outer sheath. 4. The kit of claim 3, wherein the tether includes a helical element, wherein remotely controlling the tether to release the implantable medical device from the distal portion of the outer sheath includes rotating the helical element such that the helical element releases a looped element of the implantable medical device. 5. The kit of claim 4, wherein remotely controlling the tether to release the implantable medical device from the distal portion of the outer sheath further includes pushing the implantable medical device out of the distal portion of the outer sheath with the helical element. 6. The kit of claim 1, further comprising a guidewire, wherein the inner sheath includes a guidewire lumen extending throughout the length of the inner sheath,wherein the outer sheath and the inner sheath are configured to follow a guidewire as an assembly with the guidewire lumen of the inner sheath sliding along the guidewire,wherein, in the assembly, the inner sheath is within the lumen of the outer sheath and the inflatable member of the inner sheath extends beyond the distal opening of the outer sheath and is inflated to substantially fill the inner lumen and close-off the distal opening of the outer sheath. 7. The kit of claim 6, wherein the inner sheath further includ... | Medtronic, Inc. | A61B005/02 A61B005/0215 A61B019/00 A61N001/365 A61N001/375 A61N001/05 A61N001/372 A61B005/042 A61B005/053
US09336611 | B2 | 2016-05-10 | Multi-contrast image reconstruction with joint bayesian compressed sensing | 2011-09-13 | A method for reconstructing multiple images of a subject depicting multiple different contrast characteristics from medical image data acquired with a medical imaging system is provided. Multiple image data sets are acquired with one or more medical imaging systems and the image data sets used to estimate hyperparameters drawn from a prior distribution, such as a prior distribution of image gradient coefficients. These hyperparameters and the acquired image data sets are utilized to produce a posterior distribution, such as a posterior distribution of image gradients. From this posterior distribution, multiple images with the different contrast characteristics are reconstructed. The medical imaging system may be a magnetic resonance imaging system, an x-ray computed tomography imaging system, an ultrasound system, and so on. | 1. A method for reconstructing images of a subject with a magnetic resonance imaging (MRI) system, the steps of the method comprising: a) acquiring with the MRI system, multiple image data sets with different contrast characteristics;b) estimating from the multiple image data sets, hyperparameters related to a prior distribution;c) producing a posterior distribution of image gradient coefficients using the estimated hyperparameters; andd) producing multiple images using the posterior distribution and the multiple image data sets, each of the multiple images having an image contrast associated with a contrast characteristic of the respective one of the multiple image data sets. 2. The method as recited in claim 1 in which step d) includes deriving image gradients from the posterior distribution. 3. The method as recited in claim 2 in which the image gradients are derived in step d) as the mean of the posterior distribution. 4. The method as recited in claim 2 in which step d) includes producing transformed image gradients by performing a Fourier transform on the image gradients. 5. The method as recited in claim 4 in which step d) includes producing multiple refined image data sets using the transformed image gradients. 6. The method as recited in claim 5 in which the multiple images are reconstructed by performing a Fourier transform on the multiple refined image data sets. 7. The method as recited in claim 1 in which step c) includes using Bayes' theorem to produce the posterior distribution using the hyperparameters and the acquired multiple image data sets. 8. The method as recited in claim 1 in which the hyperparameters are estimated in step b) by maximizing a log likelihood. 9. The method as recited in claim 1 in which the prior distribution related to the hyperparameters is a Gamma distribution. 10. The method as recited in claim 1 in which the hyperparameters are estimated in step b) using an expectation-maximization algorithm. 11. The method as recited in claim 10 in which step a) includes acquiring at least one image data set that samples k-space in satisfaction of a Nyquist criterion, and reconstructing a prior image from the at least one image data set. 12. The method as recited in claim 11 in which the prior image is used to initialize the expectation-maximization algorithm. 13. A method for reconstructing images of a subject from medical image data acquired with a medical imaging system, the steps of the method comprising: a) acquiring at least two medical image data sets with at least one medical imaging system, each of the at least two medical image data sets being acquired with a selected contrast characteristic;b) estimating from the at least two medical image data sets, hyperparameters related to a prior distribution of image gradient coefficients;c) producing a posterior distribution of image gradients using the estimated hyperparameters and the at least two medical image data sets; andd) reconstructing images of th... | Massachusetts Institute of Technology | G06T007/00 G06T011/00 A61B005/055 A61B005/00 G01R033/561 G01R033/48 G01R033/56 G01R033/563
US09345403 | B2 | 2016-05-24 | Wireless monitoring system with activity manager for monitoring user activity | 2013-08-14 | A system is disclosed that uses telemetry data based on a user habit information or user monitoring. A user monitoring device has a unique user ID. The monitoring device acquires user information selected from of at least one of, a user's activities, behaviors and habit information, and user monitoring. ID circuitry is at the user monitoring device. The ID circuitry includes ID storage, a communication system that reads and transmits the unique ID from ID storage, a power source and a pathway system to route signals through the circuitry. A telemetry system is in communication with the user monitoring device. The telemetry system includes a database of user ID's. The telemetry system is in communication with the user monitoring device. An activity manager is provided is used with the other elements to manage lifestyle activities of the user. | 1. A system for using telemetry data based on a user habit information or user monitoring, comprising: a monitoring device with a unique user ID, the monitoring device including a microphone, an RF transmitter and sensors to determine air quality, sound level, sound quality, light quality and ambient temperature near a user, the RF transmitter serving as a communication system with the unique user ID;ID circuitry at the monitoring device, the ID circuitry including ID storage, the communication system that reads and transmits the unique user ID from the ID storage, a power source and a pathway system, wherein the pathway system to route signals through the ID circuitry;an accelerometer configured to detect a user's movement information, the accelerometer and the monitoring device configured to assist to determine a user's sleep information and sleep behavior information, the microphone configured to record user movement sounds detected by the accelerometer, the accelerometer configured to cause the microphone to stop recording the user movement sounds when the movement sounds are not directed to a sleep related parameter;a telemetry system in communication with the monitoring device, the telemetry system including a database of user ID's, the telemetry system in communication with the monitoring device, the telemetry system analyzing telemetry data received from the monitoring device and creating personalized information about the user; andan activity manager is included in the monitoring device, the telemetry system or is separate from each and used for managing lifestyle activities of the user, and the activity manager includes an activity profile database. 2. The system of claim 1, wherein the activity manager in operation associates one or more contexts of a user activity. 3. The system of claim 2, wherein the one or more contexts is selected from at least one of, time, location, type of user activity, a duration of the user activity and a status of the user activity. 4. The system of claim 1, wherein at least one of the sensors is a lifestyle sensor. 5. The system of claim 1, wherein at least one of the sensors is a physiological sensor. 6. The system of claim 1, wherein at least one of the sensors monitors at least one of, (i) cardiac activity (ii) Hemodynamics, (iii) respiratory function, (iv) neurological monitoring, (v) a body fluid component, (vi) body temperature, (vii) stress and (viii) toxicity. 7. The system of claim 1, further comprising: a user interface at the monitoring device. 8. The system of claim 7, wherein the user interface includes an entry field to enter a name of an activity. 9. The system of claim 1, wherein the activity manager in operation manages and monitors a user activity. 10. The system of claim 1, wherein the activity manager in operation provides a user reminder when to start an activity. 11. The system of claim 1, wherein the activity manager in operation provides a user reminder when to stop an act... | Hello Inc. | G08B023/00 A61B005/00 G08B021/24 H02J007/02 H02J017/00 A61B005/11 H02J007/00 A61B005/0205
US10548991 | B2 | 2020-02-04 | IR dyes for fluorescence imaging | 2015-11-23 | A method for organ imaging, comprising: administering to a subject a diagnostic effective amount of 2-((E)-2-((E)-3-(2-((E)-3,3-dimethyl-5-sulfonato-1-(4-sulfonatobutyl)indolin-2-ylidene)ethylidene)-2-phenoxycyclohex-1-en-1-yl)vinyl)-3,3-dimethyl-1-(4-sulfonatobutyl)-3H-indol-1-ium-5-sulfonate or 2-((E)-2-((E)-3-(2-((E)-3,3-dimethyl-5-sulfonato-1-(4-sulfonatobutyl)indolin-2-ylidene)ethylidene)-2-(4-sulfonatophenoxy)cyclohex-1-en-1-yl)vinyl)-3,3-dimethyl-1-(4-sulfonatobutyl)-3H-indol-1-ium-5-sulfonate. In one embodiment, the organ includes one or more of kidney, bladder, liver, gall bladder, spleen, intestine, heart, lungs and muscle. | 1. A solid form, which is a Form A polymorph of a compound of Formula 1:having an X-ray powder diffraction pattern comprising peaks at 2-theta at 4.3°±0.2°, 9.6°±0.2° and 12.9°±0.2°. 2. The solid form of claim 1, further comprising peaks at 2-theta at about 18.3°±0.2° and about 20.8°±0.2°. 3. The solid form of claim 1, wherein the X-ray powder diffraction pattern for the Form A polymorph is as shown in FIG. 1E. 4. The solid form of claim 1, wherein the Form A polymorph is substantially pure. 5. A solid form, which is a Form B polymorph of a compound of Formula 1:having an X-ray powder diffraction pattern comprising peaks at 2-theta at 5.3°±0.2°, 14.2°±0.2° and 21.2°±0.2°. 6. The solid form of claim 5, further comprising peaks at 2-theta at about 4.3°±0.2° and about 20.7°±0.2°. 7. The solid form of claim 5, wherein the X-ray powder diffraction pattern for the Form B polymorph is as shown in FIG. 1H. 8. The solid form of claim 5, wherein the Form B polymorph is substantially pure. 9. A method for making a polymorphic compound of Formula 1, the method comprising: dissolving a compound of Formula 1 in a mixture of water and isopropanol with heating to form a dissolution, wherein the compound of Formula 1 has the formula:cooling the dissolution to ambient temperature to form a solid crystalline material; anddrying the solid crystalline material to form a Form A polymorph, which has an X-ray powder diffraction pattern comprising peaks at 2-theta at 4.3°±0.2°, 9.6°±0.2° and 12.9°±0.2°. 10. The method of claim 9, further comprising subjecting the Form A polymorph to a relative humidity of about 75% while heating to change the Form A polymorph to a Form B polymorph, wherein the Form B polymorph has an X-ray powder diffraction pattern comprising peaks at 2-theta at 5.3°±0.2°, 14.2°±0.2° and 21.2°±0.2°. | LI-COR, INC. | A61K049/00 C09B023/10 A61B005/00 A61B005/20
US09345439 | B2 | 2016-05-24 | Monte carlo and iterative methods for determination of tissue oxygen saturation | 2013-05-03 | A method for determining oxygen saturation includes emitting light from sources into tissue; detecting the light by detectors subsequent to reflection; and generating reflectance data based on detecting the light. The method includes determining a first subset of simulated reflectance curves from a set of simulated reflectance curves stored in a tissue oximetry device for a coarse grid; and fitting the reflectance data points to the first subset of simulated reflectance curves to determine a closest fitting one of the simulated reflectance curves. The method includes determining a second subset of simulated reflectance curves for a fine grid based on the closest fitting one of the simulated reflectance curves; determining a peak of absorption and reflection coefficients from the fine grid; and determining an absorption and a reflectance coefficient for the reflectance data points by performing a weighted average of the absorption coefficients and reflection coefficients from the peak. | 1. A method for determining oxygen saturation of tissue via a tissue oximetry device comprising: emitting light having at least two wavelengths into tissue from a first source structure and a second source structure that are optically coupled respectively to the first and second light sources;detecting the light by a plurality of detector structures subsequent to reflection of the light from the tissue, wherein the plurality of detector structures comprises a first detector structure, a second detector structure, a third detector structure, and a fourth detector structure,the first and second source structures and the first, second, third, and fourth detector structures are arranged in a circular arrangement on a probe face of the tissue oximetry device,a first distance is from the first detector structure to the first source structure,a second distance is from the first detector structure to the second source structure,the first distance is greater than the second distance,a third distance is from the second detector structure to the first source structure,a fourth distance is from the second detector structure to the second source structure,the fourth distance is greater than the third distance,a fifth distance is from the third detector structure to the first source structure,a sixth distance is from the third detector structure to the second source structure,the fifth distance is different from the first distance and the second distance,the sixth distance is different from the first distance and the second distance,a seventh distance is from the fourth detector structure to the first source structure,an eighth distance is from the fourth detector structure to the second source structure,the seventh distance is different from the first, second, fifth, and sixth distances,the eighth distance is different from the first, second, fifth, and sixth distances,the first distance is greater than the second, third, fifth, sixth, seventh, and eighth distances, andthe second distance is less than the fifth, sixth, seventh, and eight distances;generating digital reflectance data points for the tissue from detected light;retrieving a plurality of simulated data points from a memory of the tissue oximetry device for each simulated reflectance curve in a set of simulated reflectance curves, wherein the simulated data points are simulated reflectance intensities of the simulated reflectance curves for the first, second, third, fourth, fifth, sixth, seventh, and eighth distances;from the simulated data points, for the simulated reflectance intensities for the first, second, third, fourth, fifth, sixth, seventh, and eighth distances, for the set of simulated reflectance curves, selecting, by a processor of the tissue oximetry device, a first of the simulated reflectance curves as a first selected simulated reflectance curve;from first simulated data points of the plurality of simulated data points for simulated reflectance intensities for the first, second, thi... | ViOptix, Inc. | A61B005/1455 A61B005/00 A61B005/1495 A61B019/00 A61M035/00
US20160139039 | A1 | 2016-05-19 | IMAGING SYSTEM AND IMAGING METHOD | 2015-11-25 | An imaging system includes an infrared camera 10 that is sensitive to light beams having wavelengths in an infrared region, a lighting unit 20 that emits light beams having multiple wavelengths in an infrared region in a region including the wavelengths to which the infrared camera is sensitive, and a control unit 30 that controls capture of an image by the infrared camera 10 and emission of a light beam by the lighting unit 20. | 1. An imaging system comprising: an infrared camera that is sensitive to light beams having wavelengths in an infrared region;a lighting unit that emits light beams having a plurality of wavelengths in an infrared region in a region including the wavelengths to which the infrared camera is sensitive; and a control unit that controls capture of an image by the infrared camera and emission of a light beam by the lighting unit. 2. The imaging system of claim 1, wherein the lighting unit emits light beams having a plurality of wavelengths of 800 nm or more and 2500 nm or less. 3. The imaging system of claim 2, wherein the lighting unit emits light beams having a plurality of wavelengths of 1000 nm or more and 1600 nm or less. 4. The imaging system of claim 1, further comprising: a visible camera that is sensitive to a light beam having a wavelength in a visible region; anda visible lighting unit that emits a light beam having the wavelength to which the visible camera is sensitive. 5. The imaging system of claim 1, wherein the control unit comprises a lighting drive unit that causes the lighting unit to emit light beams having different wavelengths sequentially or simultaneously. 6. The imaging system of claim 5, wherein the control unit synchronizes switching between emission wavelengths by the lighting drive unit and capture of an image by the infrared camera. 7. The imaging system of claim 1, wherein the control unit assigns a frame number to each wavelength in a manner corresponding to the degrees of sensitivity of the infrared camera to the wavelengths. 8. The imaging system of claim 1, wherein the infrared camera comprises a plurality of infrared cameras disposed in such a manner that the infrared cameras look into an identical portion of a subject in different fields of view; and the lighting unit comprises lighting units disposed so as to correspond to the infrared cameras, whereinthe infrared cameras capture images of light beams emitted from the lighting units and then reflected from the subject or light beams emitted from the lighting units and then transmitted through the subject. 9. The imaging system of claim 1, further comprising a drive unit that moves the infrared camera and the lighting unit so that the infrared camera looks into an identical portion of a subject in different fields of view. 10. The imaging system of claim 1, wherein the control unit comprises an image processing unit that combines a plurality of images captured by the infrared camera. 11. The imaging system of claim 10, wherein the image processing unit generates a stereoscopic image of the subject on the basis of the images captured by the infrared camera. 12. The imaging system of claim 1, wherein the lighting unit emits the light beams having the wavelengths comprising at least one infrared wavelength of 1000 nm or more and is mounted on at least one of a mammotome device, mammography, a dermoscope, and a tooth row transmission device. 13. An imaging... | NATIONAL INSTITUTE OF ADVANCED INDUSTRIAL SCIENCE AND TECHNOLOGY NIKON CORPORATION | G01N021/359 A61B005/00 H04N005/33 H04N005/232 G01N021/3563 H04N013/02
US09351663 | B2 | 2016-05-31 | Method for locating a catheter tip using audio detection | 2014-05-08 | A method for locating a catheter tip within a human body is disclosed. An audio sensor is positioned at a site on the human body. An audio signal is detected by the audio sensor, and transmitted to an audio signal processing unit. The audio signal processing unit determines if the audio signal corresponds to a target location of the catheter tip, and transmits a notification signal to a user notification unit. An infusion of fluid or an audio emitting element can be used to generate the audio signal at the catheter tip. | 1. A method for locating a catheter tip within a human body comprising: positioning a first audio sensor at a first site on or near the human body,infusing fluid through an opening in the catheter tip,the first audio signal caused by the infusing fluid,detecting a first audio signal through the first audio sensor,transmitting the first audio signal to an audio signal processing unit,determining if the first audio signal corresponds to a target location of the catheter tip, andtransmitting a notification signal to a user notification unit. 2. The method of claim 1, wherein the first site is located substantially over a lower ⅓ of a superior vena cava of the human body. 3. The method of claim 1, wherein the first audio sensor comprises a piezo film element. 4. The method of claim 1, further comprising: detecting a first baseline audio signal through the first audio sensor,transmitting the first baseline audio signal to the audio signal processing unit, andcomparing the first baseline audio signal to the first audio signal. 5. The method of claim 1, further comprising: detecting a first electrocardiogram signal through a first electrocardiogram electrode,transmitting the first electrocardiogram signal to an electrocardiogram signal processing unit, anddetermining if the first electrocardiogram signal corresponds to the target location. 6. The method of claim 5, wherein the first audio sensor and the first electrocardiogram electrode are operably connected to a first pad member, and wherein the first pad member is positioned at the first site. 7. The method of claim 5, further comprising: detecting a first baseline electrocardiogram signal through the first electrocardiogram sensor,transmitting the first baseline electrocardiogram signal to the electrocardiogram signal processing unit, andcomparing the first baseline electrocardiogram signal to the first electrocardiogram signal. 8. The method of claim 1, wherein the fluid is a saline bolus. 9. The method of claim 1, further comprising: positioning a second audio sensor at a second site on or near the human body,detecting a second audio signal through the second audio sensor,transmitting the second audio signal to the audio signal processing unit, anddetermining if the second audio signal corresponds to the target location. 10. The method of claim 9, wherein the second site is located substantially over an internal jugular vein. 11. The method of claim 9, further comprising: detecting a second baseline audio signal through the second audio sensor,transmitting the second baseline audio signal to the audio signal processing unit, andcomparing the second baseline audio signal to the second audio signal. 12. The method of claim 9, further comprising: determining if the first audio signal and the second audio signal correspond to a non-target location of the catheter tip. 13. The method of claim 9, wherein the first site and the second site are located on opposite sides of and substantially eq... |  | A61B005/00 A61B005/06
US09788893 | B2 | 2017-10-17 | Catheter with soft distal tip for mapping and ablating tubular region | 2014-11-20 | A catheter includes an elongated body having a longitudinal axis, a distal assembly distal the elongated body, the distal assembly having a tapered helical form comprising a larger, electrode-carrying proximal loop and a smaller, softer distal loop, and a shape-memory support member extending through at least the proximal loop. For example, the helical loop subtends at least about 720 radial degrees, with the proximal loop subtending about 360 radial degrees, and the distal loop subtending about 360 radial degrees. The softer distal loop with a straight distal end atraumatically guides the distal assembly into a tubular region so that the larger proximal loop can sit on the ostium of the tubular region with improved electrode and tissue contact. | 1. A catheter comprising: an elongated body having a longitudinal axis;a distal assembly distal the elongated body, the distal assembly having a helical form comprising a proximal loop and a distal loop, and a shape-memory support member extending through at least the proximal loop, the proximal loop and the distal loop being oriented obliquely at an angle relative to the longitudinal axis of the elongated body;at least one irrigated ablation ring electrode mounted on the proximal loop;a control handle proximal the elongated body; anda contraction wire having a proximal end in the control handle and a distal end anchored in the proximal loop, the control handle including a first control member configured to actuate the contraction wire to contract the proximal loop,wherein the proximal loop has a first flexibility and the distal loop has a second flexibility, and the second flexibility is greater than the first flexibility. 2. A catheter of claim 1, wherein the helical form subtends at least about 720 radial degrees. 3. A catheter of claim 1, wherein the helical form subtends about 765 radial degrees. 4. The catheter of claim 1, wherein the proximal loop subtends about 360 radial degrees. 5. The catheter of claim 1, wherein the distal loop subtends about 360 radial degrees. 6. The catheter of claim 1, wherein the distal assembly includes a generally straight distal end distal of the distal loop. 7. The catheter of claim 6, wherein the generally straight distal end is on-axis relative to the longitudinal axis of the elongated body. 8. The catheter of claim 1, wherein the proximal loop has a first radius and the distal loop has a second radius, and the first radius is greater than the second radius. 9. The catheter of claim 1, wherein the proximal loop and the distal loop are on-axis relative to the longitudinal axis of the elongated body. 10. The catheter of claim 1, wherein the oblique angle ranges between about 45 degrees to 105 degrees. 11. The catheter of claim 1, wherein the oblique angle ranges between about 75 to 105 degrees. 12. The catheter of claim 1, wherein the oblique angle is about 90 degrees. 13. The catheter of claim 1, wherein the proximal loop includes between about eight to twenty electrodes. 14. The catheter of claim 1, wherein the proximal loop includes about ten electrodes. 15. The catheter of claim 1, wherein the proximal loop includes about six electrodes spanning about 180 radial degrees. 16. The catheter of claim 1, further comprising at least one impedance recording ring electrode mounted on the proximal loop adapted to measure impedance. 17. The catheter of claim 1, further comprising at least one bi-polar recording ring electrode mounted on the proximal loop adapted to measure PV potentials. 18. The catheter of claim 1, further comprising a deflection wire extending through the elongated body, wherein the control handle includes a second control member configured to actuate the deflection wire to defle... | Biosense Webster (Israel) Ltd. | A61B005/042 A61B018/14 A61B005/053 A61B018/00 A61B017/00 A61B018/12 A61B034/20
US09351669 | B2 | 2016-05-31 | Interconnect for on-body analyte monitoring device | 2010-09-30 | Disclosed herein are systems and methods for providing a compressible interconnect for allowing electrical communication between an electronics unit and an analyte sensor in an on-body analyte monitoring device. In other embodiments, systems and methods are provided for reducing the Z-height of an on-body analyte monitoring device by utilizing novel interconnects. | 1. An electronics unit for receiving one or more signals from an analyte sensor, the electronics unit comprising: a body having a top surface, an opposing bottom surface and a sidewall defined therebetween;a processor and a circuit board disposed within the body; andan interconnect comprising conductive material, the interconnect disposed between the top surface and the opposing bottom surface of the body and having a first end coupled to the circuit board and a second end coupled to an analyte sensor, wherein the first end of the interconnect coupled to the circuit board and the second end of the interconnect coupled to the analyte sensor are disposed within the body. 2. The electronics unit of claim 1, wherein the interconnect further comprises a first contact at the first end of the interconnect, and a second contact at the second end of the interconnect. 3. The electronics unit of claim 2, wherein the second contact at the second end coupled to the analyte sensor establishes electrical communication between the interconnect and the analyte sensor. 4. The electronics unit of claim 1, wherein the second end of the interconnect is frictionally engaged with the circuit board. 5. The electronics unit of claim 1, wherein the analyte sensor comprises a plurality of electrodes including a working electrode, wherein the working electrode comprises an analyte-responsive enzyme and a mediator, wherein at least one of the analyte-responsive enzyme and the mediator is chemically bonded to a polymer disposed on the working electrode, and wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer. 6. The electronics unit of claim 1, further including a data communication component disposed within the body and coupled to the processor and the circuit board, the data communication component configured to communicate signals received from the analyte sensor. 7. The electronics unit of claim 6, wherein the data communication component is configured to wirelessly communicate the signals received from the analyte sensor to a remote location. 8. The electronics unit of claim 6, wherein the data communication component is configured to communicate the signals using one or more of an infrared communication protocol, Bluetooth® communication protocol Zigbee® communication protocol, radio frequency identification (RFID) communication protocol, or 802.11 wireless LAN protocol. 9. An electronics unit for receiving one or more signals from an analyte sensor, the electronics unit comprising: a body having a top surface, an opposing bottom surface and a sidewall defined therebetween;a processor and a circuit board disposed within the body; andan interconnect comprising conductive material, the interconnect disposed between the top surface and the opposing bottom surface of the body and having a first end coupled to the circuit board and a second end coupled to the analyte sensor, wherein the first end of the interconnect ... | ABBOTT DIABETES CARE INC. | A61B005/00 A61B005/145
US09332938 | B2 | 2016-05-10 | Flat lancet immobilization | 2008-09-11 | A lancet enclosed in a sterility sheet forms a lancet packet to maintain the sterility of the lancet and prevent the lancet from unintentionally piercing the sterility sheet prior to lancet actuation. In one form, the lancet is immobilized by hot tack welding a portion of the sterility sheet to the lancet. Alternatively, portions of the sterility sheet are hot tack welded together through an opening in the lancet to limit movement of the lancet. | 1. An integrated disposable, comprising: a test sensor configured to analyze body fluid; anda lancet packet coupled to the test sensor, the lancet packet including: a cover foil folded into two flaps,a lancet disposed between the two flaps of the cover foil wherein the cover foil covers at least a lancet tip of the lancet, andthe cover foil including a hot tack weld that contacts the lancet, wherein the hot tack weld welds to the lancet to immobilize the lancet within the lancet packet. 2. The integrated disposable of claim 1, wherein the hot tack weld secures the cover foil to the lancet. 3. The integrated disposable of claim 1, wherein: the lancet defines an opening; andthe hot tack weld pinches and welds opposing sheets of the cover foil together through the opening in the lancet. 4. The integrated disposable of claim 1, further comprising: a locking mechanism configured to retain the lancet within the lancet packet after the lancet is actuated. 5. The integrated disposable of claim 4, further comprising: the hot tack weld being configured to immobilize the lancet prior to actuation of the lancet;the locking mechanism including a spacer; andone of the lancet or the spacer including an insertion portion and the other of the lancet or spacer defining a receiving end, the insertion portion and the receiving end having complimentary shapes sized to attach the lancet to the spacer after the lancet is actuated. 6. The integrated disposable of claim 1, wherein the cover foil includes a tack spot positioned over the opening in the lancet, wherein the tack spot is melted to form the hot tack weld. 7. A method, comprising: covering at least a lancet tip of a lancet with a cover foil folded into two layers to form a lancet packet by placing at least the lancet tip between the layers; andmelting a portion of the cover foil to form a hot tack weld that contacts the lancet, wherein the hot tack weld welds to the lancet to immobilize the lancet. 8. The method of claim 7, further comprising: forming an integrated disposable by coupling the lancet packet to a test element configured to analyze body fluid. 9. The method of claim 7, wherein the melting the portion of the cover foil includes laser welding the portion of the cover foil to the lancet to form the hot tack weld. 10. The method of claim 7, further comprising: the hot tack weld being configured to immobilize the lancet prior to actuation of the lancet; andforming a locking mechanism on the lancet that immobilizes the lancet after the lancet is actuated. 11. The method of claim 7, further comprising: sterilizing the lancet; andattaching the lancet packet to a test element to form an integrated disposable. 12. The method of claim 7, further comprising: pinching together opposing sheets of the cover foil through an engagement opening of the lancet; andwherein the melting the portion of the cover foil includes melting the pinched portions of the opposing sheets to form the hot tack weld. 13. An... | Roche Diabetes Care, Inc. | A61B005/157 A61B005/15 A61B005/145 A61B005/1468 A61B005/151 A61B017/3209 A61B010/00
US09707393 | B2 | 2017-07-18 | Feedback-control wearable upper-limb electrical stimulation device | 2011-08-26 | A feedback-control wearable upper-limb electrical stimulation device includes a plurality of electrical stimulation electrodes (10), a plurality of electromyography signal sensors (20), an electrical stimulation output unit (30), an electromyography signal acquisition unit (35), an electromyography signal operation unit (40), and a control module (50). Each electrical stimulation electrode (10) is adhered to or fixed in contact with a human trunk and applies an electrical stimulation signal to the neuromuscular system of the human trunk. Each electromyography signal sensor (20) is adhered to or fixed in contact with the human trunk at a corresponding position of the neuromuscular system of the human trunk where each electrical stimulation electrode (10) is disposed in an adhered manner. The electrical stimulation output unit (30) is connected to each electrical stimulation electrode (10) and provides an electrical stimulation signal. The electromyography signal acquisition unit (35) is connected to each electromyography signal sensor (20) and receives a myoelectric signal. The electromyography signal operation unit (40) is connected to the electromyography signal acquisition unit (35). The control module (50) is electrically connected to the electrical stimulation output unit (30) and the electromyography signal operation unit (40). The device first determines, according to the intensity of a myoelectric signal of a human trunk, the intensity of an electrical stimulation signal required for the human trunk to perform a specified action and gives a patient suitable assistance. Therefore, a local disabled limb of a patient can be effectively activated, and a patient can be effectively exercised in controlling a diseased limb. | 1. A wearable upper limb electrical stimulation device with feedback control function comprising: multiple electrical stimulation electrodes;multiple electromyography signal sensors;an electrical stimulation output unit;an electromyography signal acquisition unit;an electromyography signal operation unit; anda control module, wherein:each one of the multiple electrical stimulation electrodes is adapted to be pasted or attached to a human's limb, and an electrical stimulation is adapted to be given to a neuromuscular system of the human's limb by each one of the multiple electrical stimulation electrodes;each one of the multiple electromyography signal sensors is adapted to be pasted or attached to the human's limb and adapted to be set on a corresponding position of the neuromuscular system of the human's limb;each electromyography signal sensor senses and receives a electromyography signal that is adapted to be generated by the human limb, wherein each electromyography signal comprises a volitional electromyography signal, a stimulus electromyography signal and an interference signal;the electrical stimulation output unit is connected with each one of the multiple electrical stimulation electrodes and provides the electrical stimulation;the electromyography signal acquisition unit is connected with the multiple electromyography signal sensors and receives the electromyography signal;the electromyography signal acquisition unit filters and enhances the electromyography signal;the electromyography signal operation unit is connected with the electromyography signal acquisition unit and receives the electromyography signal which has been filtered and enhanced;the control module is connected with the electrical stimulation output unit and the electromyography signal operation unit;the control module receives the electromyography signal filtered and enhanced by the electromyography signal operation unit and controls the electrical stimulation output unit to provide electricity to more than one electrical stimulation electrode of the multiple electrical simulation electrodes, according to a position and a strength of the electromyography signal sensed and received by each electromyography signal sensor;the electromyography signal acquisition unit thereof comprises: an electrical stimulation signal blocking circuit;an instrumentation amplifier;an amplifier circuit;a high-pass and low-pass filter; anda voltage level shifter or a clamp circuit;the electrical stimulation signal blocking circuit preliminarily filters the electromyography signal sensed and received by each electromyography signal sensor from interference of the electrical stimulation;the electromyography signal preliminarily filtered by the electrical stimulation blocking circuit is processed by a front-end amplification through the instrumentation amplifier;the electromyography signal processed by the front-end amplification is enhanced by an amplification ratio through the amplifier circui... | National Yunlin University of Science and Technology | A61N001/00 A61N001/36 A61B005/0488 A61N001/04 A61B005/00 A61B005/04
US10722155 | B2 | 2020-07-28 | Method and device for generating image showing concentration distribution of biological substances in biological tissue | 2014-05-28 | Generating an image of a molar concentration ratio between a first and second biological substance, an absorption spectrum of the first and second biological substance within a predetermined wavelength range including a first, second, third, and fourth isosbestic point, acquiring first imaging data of the tissue using light extracted from white light using a first filter to collectively selectively extract light in a first wavelength range demarcated by the first and second isosbestic point, light in a second wavelength range demarcated by the second and third isosbestic point, and light in a third wavelength range demarcated by the third and fourth isosbestic point; acquiring second imaging data by taking an image of the tissue using light from the white light using a second optical filter to selectively extract light in the second wavelength range; and generating the distribution image on the basis of the first and second imaging data. | 1. A device that generates a distribution image showing a molar concentration ratio between a first biological substance and a second biological substance included in a biological tissue of which an absorption spectrum within a predetermined wavelength range has a first isosbestic point, a second isosbestic point, a third isosbestic point and a fourth isosbestic point, wherein the wavelengths of the first, second, third and fourth isosbestic points are about 528 nm, about 546 nm, about 570 nm, and about 584 nm respectively, the device comprising: a light source which emits white light;a first optical filter configured to selectively transmit, from the white light, light within a wavelength range from the first isosbestic point to the fourth isosbestic point, and not to transmit light outside the wavelength range, the wavelength range including a first wavelength range from the first isosbestic point to the second isosbestic point, a second wavelength range from the second isosbestic point to the third isosbestic point, and a third wavelength range from the third isosbestic point to the fourth isosbestic point;a second optical filter configured to selectively transmit light within the second wavelength range from the white light, and not to transmit light outside the second wavelength range;a controller configured to sequentially insert the first optical filter and the second optical filter one after another, into a light path of the white light emitted by the light source;an image pick-up device configured to take a first image of the biological tissue using the light transmitted through the first optical filter and a second image of the biological tissue using the light transmitted through the second optical filter; andan image processor configured to generate the distribution image on the basis of the first image and the second image;wherein the first and second biological substances are oxyhemoglobin and deoxyhemoglobin,wherein the molar concentration ratio of the first biological substance and the second biological substance included in the biological tissue is oxygen saturation,wherein the first optical filter has a transmission wavelength range of about 528 nm-584 nm and the second optical filter has a transmission wavelength range of about 546 nm-570 nm, andwherein the image processor generates the distribution image on the basis of the first image and the second image by acquiring a first absorbance of the biological tissue in a transmission wavelength range of the first optical filter on the basis of the first image, acquiring a second absorbance of the biological tissue in a transmission wavelength range of the second optical filter on the basis of the second image, and calculating an index X, as X=A1−2kA2 where k is a constant, A1 is the first absorbance, A2 is the second absorbance; and generating the distribution image on the basis of the index X. 2. The device according to claim 1 comprising an endoscope device including an endosco... | HOYA CORPORATION | A61B005/00 A61B005/1455 A61B005/1459 A61B001/06 G06T007/00 A61B001/00 A61B001/04 A61B005/145
US09775520 | B2 | 2017-10-03 | Wearable personal monitoring system | 2015-11-03 | A system includes one or more sensors to detect activities of a mobile object; and a processor coupled to the sensor and the wireless transceiver to classify sequences of motions into groups of similar postures each represented by a model and to apply the models to identify an activity of the object. | 1. A wearable monitoring apparatus, comprising: a housing adapted to be worn by a person on a wrist of the person;an accelerometer disposed in the housing;a radio frequency transceiver disposed in the housing; anda processor disposed in the housing,wherein the processor analyzes signals from the accelerometer to determine whether the signals are indicative of a predefined exercise of the person, andwherein the processor computes calories burned by the person based on the signals from the accelerometer and the predefined exercise of the person, andwherein the processor transmits the calories burned to a remote device via the radio frequency transceiver. 2. The apparatus according to claim 1, further comprising: at least one bioelectric contact disposed on the housing, wherein signals from the at least one bioelectric contact are used by the processor to compute a pulse of the person. 3. The apparatus according to claim 2, wherein the processor generates an audible, tactile or visual alert if the pulse is greater than a first value or lower than a second value. 4. The apparatus according to claim 1, wherein the radio frequency transceiver is a Bluetooth® transceiver. 5. The apparatus according to claim 1, further comprising: a touch-sensitive display disposed on the housing. 6. The apparatus according to claim 1, further comprising: an angular position sensor disposed in the housing. 7. The apparatus according to claim 6, wherein the processor computes the calories burned by the person based on further signals from the angular position sensor. 8. The apparatus according to claim 6, wherein the processor computes sleep data based on movement of the person during sleep using the signals from the accelerometer and further signals from the angular position sensor. 9. The apparatus according to claim 1, further comprising: flash memory disposed in the housing. 10. The apparatus according to claim 1, further comprising: a vibration mechanism disposed on the housing for vibrating the housing. 11. The apparatus according to claim 1, further comprising: a rechargeable battery disposed in the housing. 12. The apparatus according to claim 1, further comprising: a mini-USB port disposed on the housing. 13. The apparatus according to claim 1, further comprising: a speaker and a microphone disposed on the housing. 14. The apparatus according to claim 1, further comprising: a push-button switch disposed on the housing. 15. The apparatus according to claim 1, further comprising: an optical sensor array disposed on the housing, the optical sensor array comprising a photo-detector and at least one of an infra-red light emitting diode and a visible light emitting diode, wherein signals from the optical sensor array are used by the processor to compute a pulse of the person. 16. The apparatus according to claim 15, wherein the processor generates an audible, tactile or visual alert if the pulse is greater than a first value or lower than a second value.... | Empire IP LLC | G08B001/08 A61B005/00 G06F019/00 G08B021/04 A61B005/01 A61B005/0205 A61B005/0402 A61B005/0476 A61B005/0488 A61B005/11 A61B007/04 A61B008/00 G08B021/02 H04M003/51 G08B025/01 A61B005/053 A61B008/06 A61B008/08 A61B005/026 H04W084/18
US09355651 | B2 | 2016-05-31 | System and method for expressive language, developmental disorder, and emotion assessment | 2014-04-29 | In one embodiment, a method for detecting autism in a natural language environment using a microphone, sound recorder, and a computer programmed with software for the specialized purpose of processing recordings captured by the microphone and sound recorder combination, the computer programmed to execute the method, includes segmenting an audio signal captured by the microphone and sound recorder combination using the computer programmed for the specialized purpose into a plurality recording segments. The method further includes determining which of the plurality of recording segments correspond to a key child. The method further includes determining which of the plurality of recording segments that correspond to the key child are classified as key child recordings. Additionally, the method includes extracting phone-based features of the key child recordings; comparing the phone-based features of the key child recordings to known phone-based features for children; and determining a likelihood of autism based on the comparing. | 1. A method comprising: capturing an audio recording from a language environment of a key child;segmenting the audio recording into a plurality of segments using a Minimum Duration Gaussian Mixture Model (MD-GMM) technique, the MD-GMM technique comprising performing a maximum log-likelihood analysis to generate the plurality of segments having a minimum duration constraint;identifying a segment ID for each of the plurality of segments, the segment ID identifying a source for audio in the segment of the plurality of segments;identifying a plurality of key child segments from the plurality of segments, each of the plurality of key child segments having the key child as the segment ID;estimating key child segment characteristics based in part on at least one of the plurality of key child segments, wherein the key child segment characteristics are estimated independent of contents of the plurality of key child segments, wherein the contents are meanings of the plurality of key child segments;determining at least one metric associated with the language environment using the key child segment characteristics; andoutputting the at least one metric. 2. The method of claim 1, further comprising: identifying a plurality of adult segments from the plurality of segments, each of the plurality of adult segments having an adult as the segment ID; andestimating adult segment characteristics based in part on at least one of a plurality of adult segments, wherein the adult segment characteristics are estimated independent of contents of the plurality of adult segments,wherein determining the at least one metric associated with the language environment comprises using the adult segment characteristics. 3. The method of claim 2, wherein adult segment characteristics comprise at least one of: a word count;a duration of speech;a vocalization count; ora parentese count. 4. The method of claim 2, wherein the at least one metric comprises at least one of: a quantity of key child vocalizations in a pre-set time period;a quantity of conversational turns, wherein the conversational turns comprise a sound from one of the adult or the key child, and a response to the sound from the other one of the adult or the key child; ora quantity of adult words directed to the key child in a pre-set time period. 5. The method of claim 1, further comprising: performing a first segmentation using a first MD-GMM of the MD-GMM technique, the first MD-GMM comprising a plurality of models; andgenerating a second MD-GMM of the MD-GMM technique by modifying at least one of the plurality of models,wherein: segmenting the audio recording into the plurality of segments comprises using the second MD-GMM; andidentifying the segment ID comprises using the second MD-GMM. 6. The method of claim 5, wherein: the plurality of models comprise a key child model, an electronic device model, and an adult model;the key child model comprises criteria associated with sounds from a child;the electronic devic... | LENA FOUNDATION | G10L015/14 G10L015/04 G10L017/00 G10L025/66 A61B005/00 A61B005/16 G10L015/06 G10L025/03
US09521977 | B2 | 2016-12-20 | CPR chest compression monitor with reference sensor | 2015-12-28 | Methods and devices for chest compression depth measurement for CPR. | 1. A device for assisting resuscitation of a victim during CPR chest compressions comprising: an anterior sensor disposed on an anterior segment, said anterior segment configured to be held adjacent to an anterior surface of the victim and said anterior sensor operable to produce a first signal corresponding to motion of the anterior surface during compressions;a posterior sensor disposed on a posterior segment, said posterior segment configured to be held adjacent to a posterior surface of the victim and said posterior sensor operable to produce a second signal corresponding to motion of the posterior surface during compressions;a control system operable to receive the first signal corresponding to motion of the anterior surface and the second signal corresponding to motion of the posterior surface and determine compression depth therefrom; anda feedback device configured to provide CPR advisory prompts to a rescuer based on a calculated compression depth, velocity or rate;wherein the anterior segment and the posterior segment are configured to provide active decompression during CPR chest compression cycles. 2. The device of claim 1, wherein the anterior segment is joined to the posterior segment, such that the anterior segment and posterior segment are resiliently biased to open to a configuration larger than a thickness of the victim's chest. 3. The device of claim 1, wherein the anterior segment is joined to the posterior segment, such that the anterior segment and posterior segment are resiliently biased to open to a configuration larger than the circumference of the victim. 4. The device of claim 1 wherein: the posterior sensor is operable to detect acceleration of the posterior surface, and thethe control system is operable to determine depth, velocity or rate of CPR chest compressions from acceleration data obtained from the posterior sensor and data from the anterior sensor during CPR chest compressions. 5. The device of claim 1 wherein: the anterior sensor is operable to detect acceleration of the anterior surface, and the control system is operable to determine depth, velocity or rate of CPR chest compressions from acceleration data obtained from the anterior sensor and data from the posterior sensor during CPR chest compressions. 6. The device of claim 1 wherein: the posterior sensor is operable to detect velocity of the posterior surface, and the control system is operable to determine depth, velocity or rate of CPR chest compressions from velocity data obtained from the posterior sensor and data from the anterior sensor during CPR chest compressions. 7. The device of claim 1 wherein: the anterior sensor is operable to detect velocity of the anterior surface, and the control system is operable to determine depth, velocity or rate of CPR chest compressions from velocity data obtained from the anterior sensor and data from the posterior sensor during CPR chest compressions. 8. The device of claim 1, wherein the control system i... | ZOLL MEDICAL CORPORATION | A61B005/00 A61B005/11 A61H031/00 G01P015/02 G01P003/42 A61N001/39 A61B005/103 A61B005/113
US09693702 | B2 | 2017-07-04 | System and method of potential equalization monitoring for medical devices | 2014-11-14 | A method of potential equalization monitoring includes connecting a plurality of electrical devices to a patient, electrically connecting each of the plurality of electrical devices to a common equipotential grounding hub, the equipotential grounding hub defining a ground, monitoring a flow of current to the ground from at least one of the plurality of electrical devices, detecting a change in the balance of electrical charge between the plurality of devices, and identifying a source of the change in the balance of electrical charge. | 1. A method of potential equalization monitoring, comprising the steps of: connecting a plurality of electrical devices to a patient;electrically connecting each of the plurality of electrical devices to a common equipotential grounding hub, the equipotential grounding hub defining a ground;monitoring a flow of current to the ground from at least two or more of the plurality of electrical devices, via the equipotential grounding hub;detecting a change in the balance of electrical charge between the plurality of devices, via the equipotential grounding hub, andidentifying a source of the change in the balance of electrical charge, via the equipotential grounding hub. 2. The method according to claim 1, wherein: detecting the change in the balance of electrical charge includes mapping the flow of current between the plurality of devices. 3. The method according to claim 2, wherein: detecting the change in the balance of electrical charge includes measuring the flow of electrical charge between the plurality of devices and plotting the flow of electrical charge to obtain change in charge data. 4. The method according to claim 3, further comprising the step of: storing the data in a database for future use. 5. The method according to claim 1, further comprising the step of: generating an audible alarm if the change in the balance of electrical charge exceeds a threshold amount. 6. The method according to claim 1, further comprising the step of: displaying information relating to the change in the balance of electrical charge to a user. 7. The method according to claim 6, wherein: the information includes identifying information indicating a particular device among the plurality of electrical devices that is the source of the change in the balance of electrical charge. 8. The method according to claim 7, wherein: the information includes at least one troubleshooting step to be carried out by a user to resolve the change in the balance of electrical charge. 9. The method according to claim 8, further comprising the step of: storing the information in a database. 10. The method according to claim 8, wherein: the plurality of devices include an ablation device having a catheter having an ablation electrode, and an ECG recording device; anddetecting a change in the balance of electrical charge between the plurality of devices includes detecting a flow of current from the ablation device to the ECG recording device. 11. The method according to claim 10, wherein: the troubleshooting step includes prompting the user to check if a grounding pad is properly attached to the patient. 12. The method according to claim 1, wherein: one of the plurality of electrical devices is an ECG amplifier; andwherein the method further includes the step of monitoring a right leg drive output on the ECG amplifier. 13. A method of potential equalization monitoring, comprising the steps of: electrically connecting a plurality of electrical devices to an equipotential ... | GENERAL ELECTRIC COMPANY | A61B018/00 A61B005/0432 G01R029/24 A61B005/0428 A61B018/12 A61B018/14 A61B005/00
US10564233 | B2 | 2020-02-18 | Magnetic resonance imaging apparatus and control method therefor | 2014-06-18 | The magnetic resonance imaging apparatus includes: a calculation means for calculating the amount of absorption of electromagnetic waves into the object according to the emission of RF pulses in a part of the object or a bed position where imaging is scheduled; a means for setting imaging conditions, in which the calculated amount of absorption satisfies conditions of the specified value of the amount of absorption of electromagnetic waves, from the relationship between the calculated amount of absorption and the specified value of the amount of absorption of electromagnetic waves; and a bed control device that controls a top plate according to the set imaging conditions. | 1. A magnetic resonance imaging apparatus, comprising: a bed including a top plate that moves an object placed thereon;a magnetic generation means which generates a magnetic field in a space in which the object is located;an irradiation coil for irradiating the object with RF pulses;a receiver which detects an NMR generated by the object and images the detected NMR signal;an input and output device for inputting or displaying imaging conditions; anda control device that calculates an amount of absorption of electromagnetic waves caused by the irradiation of the RF pulses to the object based on the input imaging conditions, determines whether or not the calculated amount of absorption satisfies conditions of a specified value of the amount of absorption of electromagnetic waves, and performs movement control of the top plate and irradiation control of the RF pulses in accordance with imaging conditions determined to satisfy the conditions of the specified value of the amount of absorption of electromagnetic waves,wherein the control device further includes a database of an absorption rate of electromagnetic waves that uses a part of the object to be imagined or a bed position where imaging is scheduled as a search parameter,wherein the control device reads an absorption rate of electromagnetic waves from the database with the part of the object to be imaged or the bed position where imaging is scheduled as a search parameter, and calculates the amount of absorption of electromagnetic waves in imaging in the part or the bed position using the read absorption rate,wherein a database of the absorption rate of electromagnetic waves is data indicating an absorption rate of electromagnetic waves of a standard person with a part or a bed position as a search parameter,wherein data indicating the absorption rate of electromagnetic waves of the standard person is read from the database, with the part of the object to be imaged or the bed position where imaging is scheduled as a search parameter, andwherein the amount of absorption of the object is calculated using the read data of the absorption rate of electromagnetic waves of the standard person. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the control device stores position information for imaging in the part of the object or the bed position as information corresponding to the standard person, converts the position information corresponding to the standard person into information corresponding to the object, and controls movement or stop of the top plate of the bed using the converted information corresponding to the object. 3. A control method of a magnetic resonance imaging apparatus including a bed having a top plate that moves an object placed thereon, a magnetic field generation means which generates a magnetic field in a space in which the object is located, an irradiation coil for irradiating the object with RF pulses, a means which detects an NMR signal generated ... | Hitachi, Ltd. | G01R033/00 G01R033/28 A61B005/055 A61B005/00 G01R033/54 G01R033/58
US11058361 | B2 | 2021-07-13 | Signal processing apparatus, imaging apparatus, and signal processing method | 2015-11-02 | In one embodiment, a signal processing apparatus includes a storage circuit and processing circuitry configured to (a) generate detection parameters for detecting a specific signal included in a biosignal relevant to a heartbeat, based on a waveform of the biosignal, (b) store the detection parameters in the storage circuit, (c) detect the specific signal by using the detection parameters, and (d) generate a synchronization signal for performing heartbeat synchronization imaging based on the specific signal. | 1. A signal processing apparatus comprising: a storage circuit; andprocessing circuitry configured to (a) generate two or more waveform templates for detecting an R-wave included in an electrocardiogram (ECG) signal, based on two or more waveforms acquired from a single biosignal sensor,(b) store the two or more waveform templates in the storage circuit,(c) sequentially generate a new waveform template, by extracting a part of the waveform of the ECG signal inputted time-sequentially,(d) calculate a similarity value between the new waveform template and each of the waveform templates stored in the storage circuit,(e) update the waveform templates stored in the storage circuit by storing the new waveform template when the similarity value is smaller than a predetermined value,(f) detect the R-wave by performing matching processing between a waveform of the ECG signal inputted time-sequentially and each of the two or more waveform templates, and(g) generate a synchronization signal for performing heartbeat synchronization imaging based on the R-wave. 2. The signal processing apparatus according to claim 1, wherein the processing circuitry is configured to(a) calculate two or more evaluation values respectively corresponding to the two or more waveform templates, by performing matching processing between the waveform of the ECG signal and the two or more waveform templates, and(b) detect the R-wave by comparing an integrated evaluation value obtained from the calculated two or more evaluation values with a predetermined threshold value. 3. The signal processing apparatus according to claim 2, wherein the two or more evaluation values are respective two or more difference evaluation values based on differences between the waveform of the ECG signal and the two or more waveform templates, and the integrated evaluation value is a smallest value among the two or more difference evaluation values. 4. The signal processing apparatus according to claim 1, wherein the processing circuitry is configured to(a) monitor an operating state of an external imaging apparatus connected to the signal processing apparatus, and(b) determine whether the processing circuitry generates the two or more waveform templates, depending on the operating state of the external imaging apparatus. 5. The signal processing apparatus according to claim 4, wherein the external imaging apparatus is an MRI apparatus, andwherein the processing circuitry is configured to(a) monitor an operating state of the MRI apparatus,(b) update the waveform templates in a period during which the MRI apparatus does not perform imaging, and(c) hold the waveform templates, which are stored in the storage circuit immediately before start of imaging performed by the MRI apparatus, without updating the waveform templates, during the imaging performed by the MRI apparatus. 6. The signal processing apparatus according to claim 1, wherein the processing circuitry is configured to(a) enhance the R-wave in ... | CANON MEDICAL SYSTEMS CORPORATION | A61B005/00 A61B005/0245 A61B005/0456 A61B005/352 A61B006/00 A61B005/055 A61B005/35 A61B005/318
US20160135693 | A1 | 2016-05-19 | COMPRESSION CONTROL DEVICE AND COMPRESSION CONTROL METHOD | 2015-11-12 | A compression control device includes a cuff that is adapted to be wrapped to a first region of a living body, a pressure controller that controls an application pressure to be applied to the first region by the cuff, and a blood pressure obtaining section that obtains a blood pressure value from the living body. The pressure controller controls the application pressure based on the blood pressure value obtained by the blood pressure obtaining section in a state where a prescribed pressure for causing an ischemic state in a part from the first region to a periphery is applied to the first region by the cuff. | 1. A compression control device comprising: a cuff adapted to be wrapped around a first region of a living body;a pressure controller configured to control an application pressure to be applied to the first region by the cuff; anda blood pressure obtaining section configured to obtain a blood pressure value from the living body,wherein the pressure controller controls the application pressure based on the blood pressure value obtained by the blood pressure obtaining section in a state where a prescribed pressure for causing an ischemic state in a part from the first region to a periphery is applied to the first region by the cuff. 2. The compression control device according to claim 1, wherein the blood pressure obtaining section includes: a pulse wave obtaining section attached to a second region of the living body different from the first region and capable of measuring a pulse wave in the second region;an electrocardiogram obtaining section capable of measuring an electrocardiogram of the living body; anda calculating section obtaining a pulse wave transit time based on the pulse wave and the electrocardiogram, and calculating an estimated blood pressure value based on the pulse wave transit time, andthe pressure controller controls the application pressure based on the estimated blood pressure value in a state where the prescribed pressure for causing the ischemic state in the part from the first region to the periphery is applied to the first region by the cuff. 3. The compression control device according to claim 1, wherein the blood pressure obtaining section includes a catheter capable of obtaining an invasive blood pressure value from the living body, andthe pressure controller controls the application pressure based on the invasive blood pressure value in a state where the prescribed pressure for causing the ischemic state in the part from the first region to the periphery is applied to the first region by the cuff. 4. The compression control device according to claim 1, wherein the pressure controller controls the application pressure based on a mean value of blood pressure values obtained by the blood pressure obtaining section during a prescribed time period, in a state where the prescribed pressure for causing the ischemic state in the part from the first region to the periphery is applied to the first region by the cuff. 5. The compression control device according to claim 2, wherein the pressure controller controls the application pressure based on a mean value of blood pressure values obtained by the blood pressure obtaining section during a prescribed time period, in a state where the prescribed pressure for causing the ischemic state in the part from the first region to the periphery is applied to the first region by the cuff. 6. The compression control device according to claim 3, wherein the pressure controller controls the application pressure based on a mean value of blood pressure values obtained by the blood pressure... |  | A61B005/0225 A61B017/135 A61B005/0402 A61B005/0215 A61B005/021
US20160135751 | A1 | 2016-05-19 | SYSTEM AND METHOD FOR DETECTING NEUROMOTOR DISORDER | 2014-06-20 | The present invention provides a system and a method for detecting a neuromotor disorder or condition in a subject using an electronic device that comprises an input device. The method generally involves determining input device usage characteristics of the subject and comparing the result to a reference input device usage characteristic. | 1. A monitoring system for detecting a change in a clinical condition of a subject while using a central processing unit based electronic device comprising an input device that is capable of detecting motion of a subject's hand or finger, said system comprising: a data interception unit configured to intercept inputs from a user, wherein the data interception unit is configured to passively collect an input device usage characteristics;a behavior analysis unit operatively coupled to said data interception unit to receive the passively collected input device usage characteristic, anda behavior comparison unit operatively coupled to said behavior analysis unit,wherein said monitoring system dynamically monitors and passively collects input device usage characteristic information, and translates said input device usage characteristic information into representative data, stores and compares different results, and outputs a result associated with changes in the clinical condition of the subject. 2. The monitoring system of claim 1, wherein the clinical condition comprises hypoglycemia, a neurological disorder, a reaction to mediation, alcohol abuse or withdrawal, mercury poisoning, overactive thyroid, a clinical trial side-effect, peripheral neuropathy, early Parkinson's Disease, early Alzheimer's Disease, or liver failure. 3. The monitoring system of claim 2, wherein the clinical condition is hypoglycemia. 4. The monitoring system of claim 1, wherein the input device is a pointing device. 5. The monitoring system of claim 4, wherein the pointing device comprises mouse, touch screen, track ball, touch pad, joystick, stylus, or a combination thereof. 6. The monitoring system of claim 4, wherein said pointing device usage characteristic comprises: number of changes in direction on the x axis (x-flips); number of changes in direction on the y axis (y-flips); changes in speed; changes in precision; changes in distance; changes in idle time; changes in reaction time; changes in response time; changes in area under the curve; changes in maximum deviation; changes in additional distance; changes in normalized jerk; changes in keystroke transition time; changes in keystroke dwell time; changes in acceleration; changes in click latency; or a combination thereof. 7. The monitoring system of claim 1, wherein said behavior comparison unit is further configured to alert the subject when the result indicates hypoglycemic condition or other condition that influence motor control. 8. The monitoring system of claim 7, wherein said alert comprises a pop-up window, text message to the subject, an automated phone call, a phone call from health care provider, email, or a combination thereof. 9. The monitoring system of claim 1, wherein said monitoring system is configured in a computer, a smartphone, or any other central processing unit based electronic device comprising an input device that is capable of detecting a motion input from the subject's hand or finger... | Arizona Board of Regents for the University of Arizona | A61B005/00 A61B005/11
US10444067 | B2 | 2019-10-15 | Optical sensing apparatus and measuring method thereof | 2014-12-29 | An optical sensing apparatus including a light sensor, a plurality of light-emitting devices, and a controller is provided. The light sensor is disposed on a substrate. The light sensor senses a light reflection signal in a sensing area of the optical sensing apparatus. The light-emitting devices are disposed on the substrate and around the light sensor. The light-emitting devices provide an optical signal to be transmitted into the human tissue. Then, the optical signal is reflected by the human tissue to generate the light reflection signal. The controller determines whether the position of the human tissue has been changed in the sensing area. The controller drives at least one light-emitting device of the light-emitting devices and adjusts the light intensity thereof to provide the appropriate optical signal. Besides, a measuring method of the optical sensing apparatus is proposed. | 1. An optical sensing apparatus, comprising: at least one light sensor, disposed on a substrate and configured to sense a light reflection signal in a sensing area of the optical sensing apparatus;a plurality of light-emitting sources, disposed on the substrate and around the at least one light sensor and configured to provide an optical signal to be transmitted into a human tissue, wherein the optical signal is reflected to generate the light reflection signal in the human tissue;a controller, electrically connected to the at least one light sensor and the light-emitting sources and configured to determine whether a position of the human tissue is changed in the sensing area according to the light reflection signal, wherein the controller drives at least one light-emitting source of the light-emitting sources and adjusts light intensity of the at least one of the light-emitting sources to provide the optical signal according to the position of the human tissue in the sensing area; andat least one light isolation wall, disposed between and spaced apart from the at least one light sensor and each of the light-emitting sources on the substrate, wherein the at least one light isolation wall is configured to isolate any optical signal, generated by any of the plurality of the light-emitting sources and not reflected from the human tissue, from being transmitted into the at least one light sensor. 2. The optical sensing apparatus according to claim 1, wherein the at least one light sensor is located between two light-emitting sources of the light-emitting sources in a first direction and between another two light-emitting sources of the light-emitting sources in a second direction. 3. The optical sensing apparatus according to claim 1, wherein a vertical distance between a surface of the at least one light sensor and a surface of the substrate is equal to a vertical distance between a surface of each of the light-emitting sources and the surface of the substrate. 4. The optical sensing apparatus according to claim 1, wherein a vertical distance between a surface of the light sensor and a surface of the substrate is greater than a vertical distance between a surface of each of the light-emitting sources and the surface of the substrate. 5. The optical sensing apparatus according to claim 1, wherein the at least one light sensor and the light-emitting sources are disposed on a first surface of the substrate, the at least one light isolation wall is further disposed on a second surface of the substrate opposite to the first surface, and the at least one light isolation wall covers the second surface. 6. The optical sensing apparatus according to claim 1, wherein the controller determines whether the position of the human tissue in the sensing area is changed according to at least one of the light reflection signal and an action sensing signal. 7. The optical sensing apparatus according to claim 6, further comprising: an action sensing device, dispo... | INDUSTRIAL TECHNOLOGY RESEARCH INSTITUTE | G01J001/20 G01J001/32 A61B005/00 A61B005/1455
US09339656 | B2 | 2016-05-17 | Methods and systems for identifying reasons for ineffective left ventricular capture in cardiac resynchronization therapy based on EGM morphology | 2014-02-28 | The present disclosure pertains to cardiac pacing methods and systems, and, more particularly, to cardiac resynchronization therapy (CRT). In particular, the present disclosure pertains to determining the efficacy of CRT through use of an effective capture test (ECT). One or more embodiments comprises sensing a signal in response to a ventricular pacing stimulus. Through signal processing, a number of features are parsed from the signal. Exemplary features parsed from the signal include a maximum amplitude, a maximum time associated with the maximum amplitude, a minimum amplitude, and a minimum time associated with the minimum amplitude. The data is evaluated through use of the ECT. By employing the ECT, efficacy of CRT is easily and automatically evaluated. In one or more other embodiments, the reason for ineffective capture is displayed to the user. Exemplary reasons for ineffective capture include pseudo-fusion pacing and lack of capture of the ventricle. | 1. An apparatus for determining a reason for ineffective capture of a ventricle, comprising: (a) delivering means for delivering a ventricular pacing stimulus;(b) sensing means for sensing a signal in response to the ventricular pacing stimulus;(c) processing means for applying an effective capture test to the sensed data, the effective capture test comprising determining whether the ventricular pacing stimulus was effective or ineffective at capturing the ventricle based on one or more features of the signal and timing of the signal features;(d) means for repeatedly activating (a)-(c); and(e) processing means for determining the reason for ineffective capture following the repeated activation of (a)-(c), wherein determining the reason for ineffective capture comprises determining whether at least two features of the signal and timing of the at least two signal features meet one or more conditions of a set of conditions, the reason being one of pseudo-fusion or lack of ventricular capture. 2. An apparatus of claim 1 further comprising: processing means for determining from the signal a maximum amplitude, a maximum time associated with the occurrence of the maximum amplitude (Max), a minimum amplitude (Min), and a minimum time (Tmin) associated with the occurrence of the minimum amplitude for use in determining the reason for ineffective capture. 3. An apparatus of claim 2 wherein the set of conditions comprising: ¦Max−baseline (BL)¦/¦Min−BL¦&lt;0.125 and BL≧¦Min¦/8; (1)¦Max−BL¦/¦Min−BL¦>8; (2)Max≦BL; and (3)0&lt;Tmax−Tmin≦25 ms. (4) 4. An apparatus of claim 3 wherein responsive to determining sensed data satisfies any condition from the set of conditions, determining paced event causes pseudo-fusion. 5. An apparatus of claim 3 wherein responsive to determining sensed data does not satisfy any condition from the set of conditions, determining paced event does not cause ventricular capture. 6. A method for determining a reason for ineffective capture of a ventricle, the method comprising: (a) delivering a ventricular pacing stimulus;(b) sensing a signal in response to the ventricular pacing stimulus;(c) applying an effective capture test to the sensed data, the effective capture test comprising determining whether the ventricular pacing stimulus was effective or ineffective at capturing the ventricle based on one or more features of the signal and timing of the signal features;(d) repeatedly activating (a)-(c); and(e) determining the reason for ineffective capture following the repeated activation of (a)-(c), wherein determining the reason for ineffective capture comprises determining whether at least two features of the signal and timing of the at least two signal features meet one or more conditions of a set of conditions, the reason being one of pseudo-fusion or lack of ventricular capture. 7. A method of claim 6 further comprising: determining from the signal a maximum amplitude, a maximum time associated with the occurrence of the ma... | Medtronic, Inc. | A61N001/37 A61N001/05 A61N001/36 A61B005/00 A61B005/0452 A61N001/372
US09354115 | B2 | 2016-05-31 | Methods and apparatus for coaxial imaging of multiple wavelengths | 2014-07-16 | A hyperspectral/multispectral imager comprising a housing is provided. At least one light source is attached to the housing. An objective lens, in an optical communication path comprising originating and terminating ends, is further attached to the housing and causes light to (i) be backscattered by the tissue of a subject at the originating end and then (ii) pass through the objective lens to a beam steering element at the terminating end of the communication path inside the housing. The beam steering element has a plurality of operating modes each of which causes the element to be in optical communication with a different optical detector in a plurality of optical detectors offset from the optical communication path. Each respective detector filter in a plurality of detector filters covers a corresponding optical detector in the plurality of optical detectors thereby filtering light received by the corresponding detector from the beam steering element. | 1. A hyperspectral/multispectral imaging device, comprising: A) a housing having an exterior and an interior;B) at least one light source disposed on the exterior of the housing;C) at least one objective lens attached to or within the housing, the at least one objective lens disposed in an optical communication path, the optical communication path comprising an originating end and a terminating end, wherein the at least one light source is offset from the optical communication path and positioned so that light from the at least one light source directly illuminates a tissue of a subject positioned at the originating end of the optical communication path and wherein backscattered light from the directly illuminated tissue of the subject is passed from the originating end of the optical communication path, through the at least one objective lens, and to the terminating end of the optical communication path;D) a plurality of optical detectors, wherein each optical detector in the plurality of optical detectors is offset from the optical communication path and within the housing;E) a beam steering element within the interior of the housing, wherein the beam steering element comprises a mirror mounted on an actuator, the actuator configured to position the beam steering element into each operating mode in a plurality of operating modes, the beam steering element in optical communication with the at least one objective lens and positioned at the terminating end of the optical communication path, wherein each respective operating mode in the plurality of operating modes causes the mirror to be in optical communication with a different optical detector in the plurality of optical detectors, wherein: when the beam steering element is in a first operating mode in the plurality of operating modes, the minor is in optical communication with a first optical detector and not in optical communication with a second optical detector in the plurality of optical detectors, andwhen in the beam steering element is in a second operating mode in the plurality of switchable operating modes, the mirror is in optical communication with the second optical detector and not in optical communication with the first optical detector;F) a plurality of detector filters within the housing, each respective detector filter in the plurality of detector filters fixedly covering a corresponding optical detector in the plurality of optical detectors thereby filtering light before it is received by the corresponding optical detector from the beam steering element;G) at least one processor within the interior of the housing, wherein the at least one processor is in electrical communication with the at least one light source, the actuator, and the plurality of optical detectors; andH) a memory within the interior of the housing, wherein at least one program is non-transiently stored in the memory and executable by the at least one processor, the at least one program comprising instructions... | Hypermed Imaging, Inc. | G01J003/28 G01J003/02 G01J003/36 G01N021/25 G01N021/31 G01N021/47 A61B005/00 A61B005/1455 A61B019/00 G01J003/447 G01J003/12
US10143423 | B2 | 2018-12-04 | Systems to monitor body portions for injury after impact | 2014-11-21 | Systems are described for monitoring extremities for injury following an impact. A system embodiment includes, but is not limited to, a device including a deformable substrate integrated with a textile; a sensor assembly coupled to the deformable substrate, the sensor assembly configured to generate one or more sense signals based on detection of a physical impact to a body portion and based on detection of a physiological parameter; circuitry operably coupled to the sensor assembly and configured to receive the one or more sense signals based on detection of the physical impact and to determine whether the physical impact exceeds a threshold impact value, the circuitry configured to instruct the sensor assembly to detect one or more physiological parameters of the body portion when the physical impact exceeds the threshold impact value; and a reporting device operably coupled to the circuitry; and an external device communicatively coupled with the device. | 1. A system, comprising: a device configured for positioning proximate to a body portion, the device including: a deformable substrate configured to interface with the body portion, the deformable substrate including at least one of an elastomeric polymer or a hydrocolloid film;a sensor assembly coupled to the deformable substrate, the sensor assembly configured to generate one or more sense signals based on detection of a physical impact to the body portion, the sensor assembly including a chemical sensor configured to measure a chemical analyte of blood within the body portion and to generate one or more sense signals responsive to detection of the chemical analyte, the chemical analyte of blood including at least one of a blood protein, a blood peptide, a blood cell component, a plasma component, or a platelet component;circuitry operably coupled to the sensor assembly and configured to receive the one or more sense signals based on detection of a physical impact to the body portion and to determine whether the physical impact exceeds a threshold impact value, the circuitry configured to instruct the sensor assembly to begin monitoring for the chemical analyte of blood within the body portion when the physical impact exceeds the threshold impact value; wherein the circuitry is configured to determine whether blood is present in the body portion outside of a blood vessel responsive to detection of the chemical analyte of blood by the chemical sensor; anda reporting device operably coupled to the circuitry and configured to generate one or more communication signals based on instruction by the circuitry; andan external device communicatively coupled with the device, the external device configured to receive the one or more communication signals from the reporting device and to determine a recommended course of action pertaining to at least one of the device and the body portion. 2. The system of claim 1, wherein the deformable substrate includes a conformable structure configured to conform to a skin surface of the body portion. 3. The system of claim 1, wherein the deformable substrate is integrated with a textile configured to interface with the body portion. 4. The system of claim 3, wherein the textile includes an inner surface and an outer surface, wherein at least a portion of the deformable substrate is positioned between the inner surface and the outer surface. 5. The system of claim 3, wherein the textile includes an inner surface and an outer surface, wherein the deformable substrate is attached to one of the inner surface or the outer surface. 6. The system of claim 3, wherein the deformable substrate is integral to a weave of the textile. 7. The system of claim 3, wherein one or more fiber of the textile secures the deformable substrate to at least a portion of the textile. 8. The system of claim 3, wherein the textile includes one or more electronic threads. 9. The system of claim 1, wherein the circuitry includes a comparis... | ELWHA LLC | A61B005/00
US09345415 | B2 | 2016-05-24 | Method and apparatus for measuring multiple ECG leads using a device with embedded leads | 2012-12-26 | An apparatus and a method, the apparatus including a housing, an electrocardiogram (ECG) measuring circuit contained in the housing, and three electrodes embedded in the housing, electrically connected to the ECG measuring circuit, and in spaced apart configuration from each other. | 1. An apparatus comprising: a housing;an electrocardiogram (ECG) measuring circuit contained in the housing;three electrodes embedded in the housing, electrically connected to the ECG measuring circuit, and in spaced apart configuration from each other, wherein two of the three electrodes are arranged on a single surface of the housing and the third electrode is on another surface of the housing; andan ECG amplifier comprising a first input and a second input wherein a first of the three electrodes and a second of the three electrodes are connected to the first input and a third of the three electrodes is connected to the second input, wherein the third of the three electrodes is connected to a negative input of the ECG amplifier and the first and second electrodes of the three electrodes are electrically shorted and connected to a positive input of the ECG amplifier. 2. The apparatus of claim 1, wherein the three electrodes are electrically conductive. 3. The apparatus of claim 1, where the ECG measuring circuit, in cooperation with the three electrodes, measures a plurality of ECG Leads. 4. The apparatus of claim 3, wherein the apparatus measures nine ECG Leads. 5. The apparatus of claim 1, further including a memory to store a measurement obtained by the apparatus. 6. A method comprising, obtaining a first ECG Lead measurement by measuring a differential electrical signal using two of three electrodes of a device having three electrodes embedded in a housing, the electrodes electrically connected to an ECG measuring circuit in the housing, in spaced apart configuration from each other, and held in the hands of a user wherein the housing comprises an ECG amplifier that includes a first input and a second input and wherein a first of the three electrodes and a second of the three electrodes are connected to the first input and a third of the three electrodes is connected to the second input, wherein the third of the three electrodes is connected to a negative input of the ECG amplifier and the first and second electrodes of the three electrodes are electrically shorted and connected to a positive input of the ECG amplifier;measuring a plurality of ECG Lead measurements by measuring a differential electrical signal using two of the three electrodes, one of the electrodes placed in a hand of the user and one of the electrodes placed in contact with a torso portion of the user, wherein two of the three electrodes are arranged on a single surface of the housing and the third electrode is on another surface of the housing. 7. The method of claim 6, further comprising measuring a plurality of ECG Lead measurements by measuring a differential electrical signal using two of the three electrodes, one of the electrodes placed in a hand of the user and one of the electrodes placed in contact with a plurality of contact points on a chest portion of the user. 8. The method of claim 6, wherein the three electrodes are embedded in a body of the device. ... | INTEL CORPORATION | A61B005/0408 A61B005/0432 A61B005/00 A61B005/0404
US20160144162 | A1 | 2016-05-26 | Coupler assembly for catheters | 2015-11-23 | Coupler assemblies and methods are disclosed as the coupler assemblies may be used with a catheter. An exemplary coupler assembly includes a spherical linkage coupler for a catheter. The coupler comprises a first cylinder portion for connecting to a structure, and a second cylinder portion for connecting to a distal end of a body of the catheter. The coupler also comprises a spherical linkage including at least two link arms. Each of the two link arms are connected on one end to the first cylinder portion and on the other end to the second cylinder portion. The two link arms connect a portion of the structure to the distal end of the catheter and enable the structure to move relative to the distal end of the catheter in response to an external force exerted on the structure. | 1.-20. (canceled) 21. A catheter shaft coupler, comprising: a first rigid segment;a second rigid segment;a plurality of link arms extending between the first rigid segment and the second rigid segment, wherein the plurality of link arms are configured to: connect the first rigid segment and the second rigid segment; andallow the second rigid segment to deflect with respect to the first rigid segment. 22. The catheter shaft coupler of claim 21, wherein the first rigid segment is connected to a proximal portion of a catheter shaft and the second rigid segment is connected to a distal portion of the catheter shaft. 23. The catheter shaft coupler of claim 22, wherein the catheter shaft coupler is configured to allow the distal portion of the catheter shaft to deflect with respect to the proximal portion of the catheter shaft. 24. The catheter shaft coupler of claim 23, wherein: a distal end of the second rigid segment is connected with an optical sensor; andthe optical sensor comprises an emitter and a receiver, the emitter configured to emit an optical signal between the distal portion and the proximal portion of the catheter shaft. 25. The catheter shaft coupler of claim 21, wherein a first end of each of the plurality of link arms is connected to the first rigid segment and a second end of each of the plurality of link arms is connected to the second rigid segment. 26. The catheter shaft coupler of claim 25, wherein: the first end of each of the plurality of link arms includes a proximal end that is connected to a distal end of the first rigid segment; andthe second end of each of the plurality of link arms includes a distal end that is connected to a proximal end of the second rigid segment. 27. The catheter shaft coupler of claim 21, wherein the plurality of link arms are configured to allow the second rigid segment to move proximally with respect to the first rigid segment in response to a compressive force being applied to the second rigid segment. 28. The catheter shaft coupler of claim 21, wherein an angle of each of the plurality of link arms with respect to at least one of the first rigid segment and the second rigid segment changes in response to a force being applied to the second rigid segment. 29. The catheter shaft coupler of claim 21, further comprising a plurality of bump stops disposed on at least one of a distal surface of the first rigid segment and a proximal surface of the second rigid segment, wherein the plurality of bump stops are configured to limit a proximal movement of the second rigid segment with respect to the first rigid segment. 30. The catheter shaft coupler of claim 29, wherein an electrode is connected to a distal end of the second rigid segment. 31. The catheter shaft coupler of claim 21, wherein each of the plurality of link arms extend circumferentially around less than an entire circumference of at least one of the first rigid segment and the second rigid segment. 32. A catheter shaft, comprising:... |  | A61M039/10 A61B005/00 A61M025/00
US20160128579 | A1 | 2016-05-12 | PROBE AND SUBJECT INFORMATION ACQUIRING APPARATUS | 2015-11-09 | A probe includes a plurality of element groups, the element group including a first element and a plurality of second elements which are arranged in a first direction, the plurality of element groups being arranged in a second direction intersecting with the first direction. An acoustic lens having a curvature in the first direction is arranged on the first element. When transmission and reception based on a pulse echo technique are performed, the transmission and reception are performed with respect to the subject by using only the first element, and when reception of a photoacoustic wave is performed, the reception from the subject is performed by using the entire element group constituted by the first element and the plurality of second elements. | 1. A probe comprising: a plurality of element groups, the element group comprising a first element and a plurality of second elements which are arranged in a first direction, the plurality of element groups being arranged in a second direction intersecting with the first direction; andan acoustic lens that is arranged on the first element and has a curvature in the first direction,wherein the plurality of element groups is configured so that, when transmission and reception based on a method of transmitting an acoustic wave to a subject and receiving a reflected wave from a subject are performed, the transmission and reception are performed with respect to the subject by using only the first element, and when reception of a photoacoustic wave is performed, the reception from the subject is performed by using the entire element group constituted by the first element and the plurality of second elements. 2. The probe according to claim 1, wherein the first element and the plurality of second elements are connected, for each element, to a detection circuit configured to detect a signal from the element,the probe further comprising an adding circuit configured to perform delay and sum processing on an output signal from the detection circuit to which the first element is connected and output signals from the plurality of detection circuits to which the plurality of second elements are respectively connected, which belong to the same element group. 3. The probe according to claim 2, wherein the output signal from the detection circuit to which the first element is connected is used at the time of the reception based on the method, andwherein the output signal from the adding circuit is used at the time of the reception of the photoacoustic wave. 4. The probe according to claim 2, further comprising a multipurpose terminal configured to output the output signal from the detection circuit to which the first element is connected and the output signal from the adding circuit for each element group. 5. The probe according to claim 4, further comprising a terminal to which a switch switching signal is input which is used for selecting the output signal from the detection circuit to which the first element is connected or the output signal from the adding circuit, to be output from the multipurpose terminal. 6. The probe according to claim 4, further comprising a circuit configured to detect an input of a high voltage transmission signal at the time of the transmission based on the method and select the output signal from the detection circuit to which the first element is connected or the output signal from the adding circuit to be output from the multipurpose terminal on the basis of the detection result. 7. The probe according to claim 2, further comprising a terminal used as both a terminal configured to output the output signal from the detection circuit and a terminal to which a high voltage transmission signal at the time of the transmission bas... |  | A61B005/00
US09339254 | B2 | 2016-05-17 | Object information acquiring apparatus | 2014-01-09 | An object information acquiring apparatus disclosed by this specification includes a detector configured to detect a photoacoustic wave generated from an object on which light is irradiated from a light source, a generator configured to generate, using the photoacoustic wave, an image indicating characteristic information in a photographing region in the object, an acquirer configured to acquire information concerning quantitativity of a measurement value of the photoacoustic wave in the photographing region, and an extractor configured to extract a region having quantitativity in the image on the basis of the information concerning quantitativity. | 1. An object information acquiring apparatus comprising: a holding unit configured to hold an object;a detector configured to output a signal by detecting a photoacoustic wave generated from the object held by the holding unit and irradiated with light from a light source, said detector having a reception region, and said reception region being disposed in contact with said holding unit;a generator configured to generate, using the signal, image data indicating characteristic information about at least a portion in the object;a region acquirer configured to acquire information on a region in the image where said holding unit contacts the object; anda display information acquirer configured to cause a display unit to display display information corresponding to the image data and display information corresponding to the information on the region. 2. The object information acquiring apparatus according to claim 1, wherein first light having a first wavelength and second light having a second wavelength are irradiated from the light source, andsaid generator is configured to generate oxygen saturation as the characteristic information by using measurement values of photoacoustic waves respectively corresponding to the first light and the second light. 3. The object information acquiring apparatus according to claim 1, wherein said generator is configured to generate oxygen saturation distribution as the image data. 4. The object information acquiring apparatus according to claim 1, further comprising a recording unit configured to record the information on the region in association with the image data. 5. The object information acquiring apparatus according to claim 4, wherein said recording unit is configured to record the image data as DICOM image data and the information on the region as a private tag in association with the DICOM image data. 6. The object information acquiring apparatus according to claim 1, further comprising an input unit configured to input a photographing region, wherein said generator is configured to generate the image data in the photographing region. 7. The object information acquiring apparatus according to claim 1, wherein said display information acquirer is configured to cause the display unit to display the display information corresponding to the information on the region such that the region in the image data where said holding unit contacts the object and a region in the image data where said holding unit does not contact the object are distinguishable. 8. The object information acquiring apparatus according to claim 1, wherein the display information corresponding to the information on the region is a boundary line that enables identification of the region in the image data where said holding unit contacts the object. 9. The object information acquiring apparatus according to claim 1, wherein said display information acquirer is configured to cause the display unit to display the region in the image data... | CANON KABUSHIKI KAISHA | A61B005/1455 A61B008/08 A61B005/00 A61B005/145
US10603484 | B2 | 2020-03-31 | System and method for neurostimulation and/or neurorecording | 2015-11-24 | The disclosure describes devices and systems that include a grounding electrode. In one example, a system may include a first module including a pulse generator within a first housing, a second module comprising a switch matrix within a second housing distinct from the first housing, a connecting cable that connects the first module to the second module, a grounding electrode disposed distal of the first module and proximal of the second module, and a plurality of electrodes disposed distal of the second module, wherein each electrode of the plurality of electrodes are selectively coupled to the pulse generator via the switch matrix. These devices or system can be used to provide neurostimulation and/or neurorecording. | 1. A system for neurostimulation, the system comprising: a first module comprising a pulse generator within a first housing;a second module comprising a switch matrix within a second housing distinct from the first housing;a connecting cable that connects the first module to the second module;a grounding electrode disposed distal of the first module and proximal of the second module, wherein the grounding electrode is electrically coupled to the switch matrix of the second module; anda plurality of electrodes disposed distal of the second module, wherein each electrode of the plurality of electrodes are selectively coupled to the pulse generator via the switch matrix. 2. The system of claim 1, wherein the grounding electrode is configured to establish a defined return path for stimulation current delivered by the pulse generator to one or more electrodes of the plurality of electrodes. 3. The system of claim 1, further comprising control electronics for the plurality of electrodes and the grounding electrode, the control electronics configured to at least one of provide neurostimulation current or sense neuro signals via at least one electrode of the plurality of electrodes and the grounding electrode, wherein the control electronics are arranged in at least one of the first module or the second module. 4. The system of claim 1, wherein the grounding electrode is disposed closer to the second module than the first module. 5. The system of claim 1, wherein the grounding electrode is configured to be implanted at least partially outside of a skull of a patient, and wherein the plurality of electrodes are configured to be implanted at least partially within brain tissue. 6. The system of claim 1, wherein the grounding electrode is carried by the connecting cable. 7. The system of claim 1, wherein the switch matrix is configured to select an electrode configuration from the plurality of electrodes and the grounding electrode for delivery of electrical stimulation pulses from the pulse generator. 8. The system of claim 1, wherein the grounding electrode comprises at least one ring electrode. 9. The system of claim 1, wherein the grounding electrode comprises at least one ring segment electrode, the ring segment electrode defining a circumferential arc greater than 180 degrees and less than 360 degrees. 10. The system of claim 1, wherein the grounding electrode comprises a coil electrode. 11. The system of claim 1, wherein the grounding electrode is at least partially formed by at least one of a laser cut-out, a laser pattern, or laser-ablation. 12. The system of claim 1, wherein the grounding electrode is at least partially covered by at least one of an electrically conductive polymer or a noninsulative polymer. 13. The system of claim 1, wherein the grounding electrode is a first grounding electrode, and wherein the system comprises a second grounding electrode coupled to the first grounding electrode via one or more conductors internal o... | Medtronic Bakken Research Center B.V. | A61B005/0478 A61B005/0476 A61N001/05 A61N001/36 A61B005/00 A61N001/04 A61B005/04
US09622672 | B2 | 2017-04-18 | Digitally invertible universal amplifier for recording and processing of bioelectric signals | 2014-06-06 | A system for measuring electrical signals in a biological subject includes a variable gain amplifier with a predetermined transfer function that generates amplified signals corresponding to an input from electrical signals in the biological subject over a predetermined range of frequencies and amplification gain levels, an analog to digital converter generating digital data corresponding to the amplified signals, and a signal processing device receiving the digital data for the plurality of amplified signals. The signal processing device applies an inversion filter with another transfer function that is an inverse of the transfer function of the variable gain amplifier to remove an effect of the transfer function from the digital data, and generates an output signal corresponding to the electrical signals in the subject with reference to the filtered digital data. | 1. A system for measuring electrical signals in a biological subject comprising: a variable gain amplifier configured to generate a plurality of amplified signals corresponding to an input from a plurality of electrical signals measured in the biological subject over a predetermined range of frequencies at a plurality of predetermined amplification gain levels, the variable gain amplifier having a predetermined transfer function;an analog to digital converter (ADC) operatively connected to an output of the variable gain amplifier and configured to generate digital data corresponding to the plurality of amplified signals from the variable gain amplifier;a signal processing device connected to an output of the ADC to receive the digital data corresponding to the plurality of amplified signals, the signal processing device being configured to: apply an inversion filter to the digital data to generate filtered digital data, the inversion filter having another predetermined transfer function that is an inverse of the transfer function of the variable gain amplifier to remove an effect of the transfer function in the variable gain amplifier from the digital data; andgenerate an output signal corresponding to the plurality of electrical signals in the subject with reference to the filtered digital data;a signal generator configured to generate a predetermined signal for input to the variable gain amplifier;a signal acquisition device configured to measure another signal from the output of the variable gain amplifier corresponding to the amplified predetermined signal; andthe signal processing device being connected to the signal acquisition device and further configured to:identify an error between the signal from the variable gain amplifier corresponding to the amplified predetermined signal and an expected amplified signal corresponding to the predetermined input signal, the expected amplified signal being identified with reference to the predetermined transfer function of the amplifier;modify a parameter in the predetermined transfer function of the variable gain amplifier to reduce or eliminate the identified error; andstore the modified parameter in a memory for use in the predetermined transfer function of the inversion filter to reduce an error in the output signal. 2. The system of claim 1, the variable gain amplifier being configured to amplify each of the plurality of signals at each of the plurality of frequencies by an amount that generates the amplified signals with less than a predetermined direct current (DC) saturation level for the variable gain amplifier. 3. The system of claim 1, the variable gain amplifier further comprising: a high-impedance first stage operational amplifier that is configured to receive the plurality of electrical signals measured in the living subject. 4. The system of claim 3, the variable gain amplifier further comprising: a low-noise constant gain second stage operational amplifier configured to generate an ... | Indiana University Research and Technology Corporation | H03M001/18 A61B005/00 A61B005/04 A61B005/0402 A61B005/0476 A61B005/0488
US09352134 | B2 | 2016-05-31 | Segmented balloon catheter | 2013-11-01 | A catheter has an inflatable assembly at its distal portion, including a containment chamber, an axial core and a plurality of longitudinally oriented partitions extending from the axial core to the wall of the chamber to divide the containment chamber into at least four inflatable sectors. Hydraulic valves are connected to respective sectors to enable selective inflation of the sectors by a fluid when the valves are connected to a source of the fluid, and at least one surface electrode is mounted on each of the sectors. When introduced into a heart chamber and diametrically opposed sectors are inflated the assembly is stably fixed against the walls of the heart chamber and readings can be obtained from the surface electrodes of the inflated sectors. | 1. A medical apparatus, comprising: a catheter having an elongated shaft and a distal portion, the distal portion comprising an inflatable assembly,wherein the inflatable assembly comprises a containment chamber having an outer wall, an outer surface, an axial core and a plurality of longitudinally oriented partitions extending from the axial core to the outer wall to divide the containment chamber into at least four inflatable sectors, the sectors being externally delimited by respective bounding portions of the outer surface, the outer surface having perforations formed therethrough;hydraulic valves connected to respective sectors to enable selective inflation of diametrically opposed sectors by a fluid when the valves are connected to a source of the fluid;at least one surface electrode mounted on each of the bounding portions, the at least one surface electrode being deformable, the at least one surface electrode being adapted to receive an electric current and ablate a portion of a heart chamber wall; anda tip electrode disposed on the catheter distal to the inflatable assembly, the tip electrode adapted to provide measurements and further adapted to ablate a portion of a heart chamber wall, anda control processor operative to control the valves,wherein the apparatus is adapted to provide verification of electrical contact between the heart chamber wall and electrodes of diametrically opposed sectors. 2. A medical apparatus, comprising: a catheter having an elongated shaft and a distal portion, the distal portion comprising an inflatable assembly,wherein the inflatable assembly comprises a containment chamber having an outer wall, an outer surface, an axial core and a plurality of longitudinally oriented partitions extending from the axial core to the outer wall to divide the containment chamber into at least four inflatable sectors, the sectors being externally delimited by respective bounding portions of the outer surface, the outer surface having perforations formed therethrough;hydraulic valves connected to respective sectors to enable selective inflation of diametrically opposed sectors by a fluid when the valves are connected to a source of the fluid;at least one surface electrode mounted on each of the bounding portions, the at least one surface electrode being deformable, the at least one surface electrode being adapted to receive an electric current and ablate a portion of a heart chamber wall; anda tip electrode disposed on the catheter distal to the inflatable assembly, the tip electrode adapted to provide measurements and further adapted to ablate a portion of a heart chamber wall,a control processor operative to control the valves, the control processor being adapted to control the valves to admit the fluid into a selected pair of diametrically opposed sectors and to exclude the fluid from sectors other than the selected pair of diametrically opposed sectors, the control processor adapted to cool at least one surface electrode ... | BIOSENSE WEBSTER (ISRAEL) LTD. | A61M025/10 A61B018/00 A61B018/14 A61B005/0408 A61B018/12
US10660556 | B2 | 2020-05-26 | Body fluid sampling element | 2014-06-10 | A sampling element (110) for generating a sample of a body fluid is disclosed. The sampling element (110) comprises a housing (114), the housing having a chamber (122) with at least one puncture element (112) stored therein. A tip (132) of the puncture element (112) is movable through at least one puncture opening (124) of the housing (114) in order to perforate a skin portion of a user. The sampling element (110) further comprises at least one compression element (150), which is adapted to increase a pressure of the body fluid within a body tissue of the user in a region of puncturing when pressed onto the skin portion of the user. The compression element (150) is movably mounted to the housing (114). The sampling element (110) comprises at least one locking mechanism (168) for releasably locking the compression element (150) in at least two positions. The at least two positions comprise a first position (170) and a second position (172), the second position (172) being offset from the first position (170). In a further aspect of the invention, an analytical device (204) is disclosed, the analytical device (204) being adapted for using the sampling element (110) of the invention. | 1. A sampling element for generating a sample of a body fluid, the sampling element comprising: a housing comprising a chamber with at least one puncture element stored therein, wherein a tip of the puncture element is movable through at least one puncture opening of the housing in order to perforate a skin portion of a user;at least one compression element, wherein the compression element is adapted to increase a pressure of a body fluid within a body tissue of the user in a region of puncturing when pressed onto the skin portion of the user, wherein the compression element comprises a first, ring-shaped protrusion which fully or partially surrounds the puncture opening of the housing and is positioned to contact the skin portion of the user,wherein the compression element is movably mounted to the housing, the compression element having a first position prior to use of the sampling element, and a second position offset from the first position after use of the sampling element;a second protrusion positioned to prevent use of the sampling element when the compression element is in the second position; anda locking mechanism configured for releasably locking the compression element in the first position and for non-releasably locking the compression element in the second position. 2. The sampling element of claim 1, wherein the compression element comprises at least one trigger portion adapted to exert a trigger action onto a trigger of an analytical device using the sampling element when the compression element is moved from the first position into the second position. 3. The sampling element of claim 1, wherein the compression element comprises at least one annular protrusion being adapted to be pressed onto the skin portion in the region of puncturing, the compression element further comprising a mounting portion movably mounted to the housing, wherein the mounting portion comprises at least one mounting arm which is slidably guided in a guiding of the housing. 4. The sampling element of claim 1, wherein the locking mechanism comprises a releasable snap fit connection. 5. The sampling element of claim 1, wherein the compression element comprises at least one snap arm which is adapted to be releasably snapped into a first notch when the compression element is in the first position and is adapted to be snapped into a second notch when the compression element is in the second position. 6. The sampling element of claim 1, wherein the sampling element further comprises at least one test chemical adapted for performing at least one detectable detection reaction in the presence of at least one analyte to be detected, wherein the sampling element is adapted to transfer body fluid onto the test chemical. 7. A sampling element generating a sample of a body fluid and configured to be received within an analytical device for use, the sampling element comprising: a housing comprising a chamber with at least one puncture element stored therein, wherein... | ROCHE DIABETES CARE, INC. | A61B005/15 A61B005/151 A61B005/157 A61B005/1477
US10083502 | B2 | 2018-09-25 | Computer aided diagnosis (CAD) apparatus and method | 2015-11-24 | Disclosed are Computer Aided Diagnosis (CAD) apparatus and method to combine information on sequential image frames and to provide a superior classification result for the ROI in the image frame. The CAD apparatus may include a Region of Interest (ROI) detector configured to detect an ROI from image frames, a categorizer configured to create groups of image frames having successive ROI sections from among the image frames based on a result of the detection, a classifier configured to classify an ROI detected from each of the image frames belonging to the groups, and a result combiner configured to combine classification results for the image frames belonging to each group from the groups and to calculate a group result for the each group. | 1. A Computer Aided Diagnosis (CAD) apparatus comprising: a memory configured to store instructions;at least one processor, upon executing the stored instructions, configured to: detect a Region of Interest (ROI) from a plurality of image frames,create at least one group of image frames based on a result of the detection, wherein the at least one group comprises successive image frames in which the ROI is detected, from among the plurality of image frames,extract a feature value from at least one of a contour of the ROI comprised in the successive image frames, an area adjacent to the contour, or pixels inside the contour,classify the ROI comprised in each of the successive image frames belonging to the at least one group by comparing the extracted feature value with a pre-stored diagnostic model,combine classification results for the successive image frames belonging to each group from the at least one group, andcalculate a group result for the each group based on a result of the combination; anda display configured to display the group result on a screen. 2. The CAD apparatus of claim 1, wherein the at least one processor is further configured to group the plurality of successive image frames into a group based on an image frame where no ROI is detected. 3. The CAD apparatus of claim 1, wherein the classification result comprises at least one of BI-RADS Lexicon information or BI-RADS Category information on a corresponding ROI. 4. The CAD apparatus of claim 3, wherein the at least one processor is further configured to: calculate a score for each BI-RADS Lexicon item or each BI-RADS Category item in the each group based on the classification results, andselect a BI-RADS Lexicon item with the highest score or a BI-RADS Category item with the highest score as the group result. 5. The CAD apparatus of claim 4, wherein the at least one processor is further configured to: calculate the score for each BI-RADS Lexicon item or each BI-RADS Category item in the each group by adding up frame scores of image frames corresponding to each BI-RADS Lexicon item or each BI-RADS Category item in the group, andwherein the frame scores represent weights assigned to each of the image frames. 6. The CAD apparatus of claim 5, wherein the frame score is calculated based on a size of an ROI detected from each of the successive image frames in the each group and a model fitness of the each of the successive image frames in the each group. 7. The CAD apparatus of claim 4, wherein the at least one processor is further configured to generate a histogram of the group based on the calculated score for each BI-RADS Lexicon item or each BI-RADS Category item in the each group. 8. The CAD apparatus of claim 4, wherein the at least one processor is further configured to determine whether an ROI of the each group is malignant or benign based on the BI-RADS Lexicon item with the highest score or the BI-RADS Category item with the highest score in the each group. 9. The CA... | Samsung Electronics Co., Ltd. | G06K009/00 G06T007/00 G06K009/62 A61B008/00 A61B008/08 A61B005/00 A61B005/055 A61B006/12 A61B006/00 A61B008/14
US20160120733 | A1 | 2016-05-05 | GARMENT SYSTEM INCLUDING AT LEAST ONE SENSOR AND AT LEAST ONE ACTUATOR RESPONSIVE TO THE SENSOR AND RELATED METHODS | 2014-10-30 | Embodiments disclosed herein relate to a garment system including at least one sensor, and at least one actuator that operates responsive to sensing feedback from the at least one sensor to selectively constrict or selectively dilate at least one flexible compression garment. Such selective constriction or dilation of the at least one flexible compression garment against at least one body part can improve muscle functioning or joint functioning during an activity, such as a sport or other activity. | 1. A garment system, comprising: at least one flexible compression garment configured to be worn on at least one body part of a subject, the at least one flexible compression garment defining an interior space configured to receive the at least one body part;footwear configured to be worn on at least one foot of the subject;one or more sensors supported by the footwear, the one or more sensors configured to sense at least one characteristic associated with movement of the subject or at least one physiological characteristic of the subject, the one or more sensors further configured to output one or more sensing signals indicative of the at least one characteristic;one or more actuators positioned relative to the at least one flexible compression garment and configured to selectively constrict or selectively dilate the at least one flexible compression garment; anda control system operably coupled to the one or more actuators and further operably coupled to the one or more sensors to receive the one or more sensing signals therefrom, the control system including control electrical circuitry configured to direct the one or more actuators to selectively constrict or selectively dilate the at least one flexible compression garment responsive to the one or more sensing signals from the one or more sensors. 2. The garment system of claim 1, wherein the footwear includes at least one shoe, at least one shoe insert, at least one boot, an item of footwear associated with a snow ski, at least one water ski, at least one sock, at least one sandal, or at least one slipper. 3. The garment system of claim 1, wherein the one or more sensors are configured to detect one or more of a change in direction of travel of the subject, a load applied to the one or more sensors by a body part of the subject, pressure applied to the one or more sensors by a body part of the subject, tension applied to the one or more sensors by a body part of the subject, a load on a body part of the subject, pressure on a body part of the subject, tension on a body part of the subject, temperature of a body part of the subject, pulse of a body part of the subject, velocity of a body part of the subject, acceleration of a body part of the subject, oxygenation of a body part of the subject, nerve activity in a body part of the subject, location of the subject, gait of the subject, or pace at which the subject moves. 4. The garment system of claim 1, wherein the one or more sensors include one or more of an accelerometer, a pedometer, a counter, a tension sensor, a pressure sensor, a time-keeper, a pulse sensor, a chemical sensor, an oximeter, or a temperature sensor. 5. The garment of claim 1, wherein the at least one body part is a body part other than the at least one foot of the subject 6. (canceled) 7. (canceled) 8. (canceled) 9. The garment system of claim 1, wherein the footwear includes an interior surface that contacts at least one foot of the subject when worn, and wherein... |  | A61H009/00 A43B017/00 A43B005/04 A43B003/12 A43B003/10 A61B005/1468 A61H001/00 A61B005/103 A61B005/145 A61B005/04 A61B005/0205 A61B005/00 A43B003/00 A41B011/00
US09339207 | B2 | 2016-05-17 | Endovascular devices and methods of use | 2009-01-23 | A method for evaluating flow characteristics in a vessel of a patient includes the steps of positioning a catheter having a balloon at a measuring location within the vessel; transmitting an ultrasound signal into the vessel while the balloon catheter is within the measuring location; evaluating a reflection of the ultrasound signal to determine a flow parameter within the vessel while the catheter is in the measuring position; expanding the balloon within the vessel at the measuring location; and stopping the expanding step when the result of the evaluating step is that the flow through the vessel is substantially stopped. In some embodiments, the method may further comprise the step of detecting an endovascular electrogram signal. The method may be designed for evaluating flow characteristics in a vessel of a patient while enabling the prevention of a balloon from over-expanding and/or over-distending a vessel of a patient. | 1. A method of evaluating flow characteristics in a vessel of a patient, the method comprising: inserting a catheter comprising a balloon, an ultrasound transducer and an endovascular electrogram sensor into the vessel of the patient, wherein the ultrasound transducer is distal of the balloon and the endovascular electrogram sensor when the balloon is expanded;transmitting, from the ultrasound transducer of the catheter, a first non-imaging ultrasound signal into the vessel when the balloon is not expanded;receiving a first reflected non-imaging ultrasound signal;detecting an endovascular electrogram signal using the endovascular electrogram sensor;determining, using a processor, a first blood flow signature pattern within the vessel based on the first reflected non-imaging ultrasound signal, the first blood flow signature pattern being different from the first reflected non-imaging ultrasound signal;determining, using the processor, the location of the catheter within the vessel based on the first blood flow signature pattern and the detected endovascular electrogram signal;initiating expansion of the balloon within the vessel at a measuring location within the vessel, the measuring location within the vessel being different from the determined location within the vessel;transmitting, from the ultrasound transducer of the catheter, a second non-imaging ultrasound signal into the vessel while the balloon is being expanded;receiving a second reflected non-imaging ultrasound signal;determining, using the processor, a second blood flow signature pattern within the vessel based on the second reflected non-imaging ultrasound signal, the second blood flow signature pattern being different from the first blood flow signature pattern and from the second reflected non-imaging ultrasound signal;determining that the second blood flow signature pattern indicates that flow through the vessel has stopped; andstopping the expansion of the balloon while maintaining the stoppage of flow through the vessel in response to the determination that the second blood flow signature pattern indicates that flow through the vessel has stopped. 2. The method of claim 1, wherein the measuring location is within a pulmonary artery. 3. The method of claim 2, wherein the measuring location is within a branch of the pulmonary artery. 4. The method of claim 2, wherein the measuring location is a pulmonary artery wedge position. 5. The method of claim 1, wherein the ultrasound signals are A-mode ultrasound signals. 6. The method of claim 1, wherein the ultrasound signals are Doppler ultrasound signals. 7. The method of claim 1, wherein the blood flow signature patterns are blood flow velocities. 8. The method of claim 1, wherein the blood flow signature patterns are pressure signature patterns. 9. The method of claim 1, further comprising: determining a pressure at the measuring location;determining that the pressure at the measuring location is below a mean pulmonary arter... | VASONOVA, INC. | A61B005/06 A61B005/026 A61B005/042 A61B008/06 A61B008/12 A61B008/00 A61B019/00 G09B023/28 A61B005/0205 A61B007/04 A61B005/0215 A61B005/029 A61B005/0452 A61B005/0472 A61B005/0476 A61B005/1459 A61B008/08 A61B017/00 A61B017/22
US09341569 | B2 | 2016-05-17 | Method for detecting the presence of inhomogeneities in an interior of a turbid medium and device for imaging the interior of turbid media | 2008-12-11 | A method includes performing first and second measurements with a turbid medium to be examined placed in a receiving volume of a device for examining the interior of turbid media. The second measurement is performed after a time interval has passed after the first measurement. Each of the first and second measurements includes subsequently irradiating the turbid medium with light from a light source from a plurality of different source positions and detecting light emanating from the turbid medium in a plurality of different detection positions for each source position, and storing the detected values as measurement results. The method also includes detecting inhomogeneities in the interior of the turbid medium by using the measurement results of one of the first and second measurements as a reference and the measurement results of the respective other of the first and second measurements to determine deviations from the reference. | 1. Method for detecting the presence of inhomogeneities in an interior of a turbid medium, the method comprising the steps: performing a first measurement with the turbid medium to be examined in a receiving volume of a device for examining the interior of turbid media;performing a second measurement with the turbid medium to be examined in the receiving volume of the device for examining the interior of turbid media, wherein the second measurement is performed after a time interval has passed after the first measurement;each of the first and second measurements comprising: irradiating the turbid medium with light from a light source wherein the light from the light source irradiates the turbid medium from a plurality of different source positions, and further wherein the light from the light source irradiates the turbid medium without combining with light from any other source, and for each source position, detecting light emanating from the turbid medium in a plurality of different detection positions, andstoring the detected values as measurement results; anddetecting inhomogeneities in the interior of the turbid medium by using the measurement results of one of the first and second measurements as a reference and the measurement results of the respective other of the first and second measurements to determine deviations from the reference, wherein a multiplier for detecting inhomogenities is calculated, and further wherein computing the multipler comprises, for each combination of a source position and a detection position: computing a ratio between the measurement result from one of the first and second measurements and the corresponding measurement result from the respective other of the first and second measurements,wherein an averaged term &lt;M> is computed according to the equation:〈M〉=1N⁢∑n⁢Φn⁡(t1)Φn⁡(t2)with N being the number of different combinations of source positions and detection positions, n being an index for specific combinations of the source positions and the detection positions, Φ being the detected photon density, t1 being the time of the first measurement, and t2 being the time of the second measurement. 2. The method according to claim 1, wherein an image of the interior of the turbid medium is reconstructed based on the measurement results of the first and second measurements. 3. The method according to claim 2, wherein a first order Born approximation is used for reconstructing the image of the interior of the turbid medium. 4. The method according to claim 1, wherein the presence of an inhomogeneity is detected by generating a histogram based on the measurement results of the first and second measurements. 5. The method according to claim 1, wherein the first measurement and the second measurement are performed over a time interval where an oxygenation level and/or blood content in the turbid medium has changed based on a variation in applied pressure or based on differences within an inhale/exhale breathing cycl... | KONINKLIJKE PHILIPS N.V. | A61B005/05 G01N021/47 A61B005/00 G01N021/64
US09326696 | B2 | 2016-05-03 | Medical accessory proximity testing, detection, and alerting system | 2014-06-24 | Disclosed is a system having an implanted component and external component which are configured to provide a test of wireless communication in order to assess the success or failure of such communication and to store attributes related to such test in a memory log. To provide the communication test the implantable and external components can attempt wireless communication according to communication test parameters which relate to number of times to retry communication, duration of sending communication test signals, durations of waiting for communication test signals and the schedule of the communication tests. The schedule of tests may be period or may change over time in order to become more or less frequency according to a programmable schedule that may also decrease if the communication tests are successful and indicate patient compliance in keeping the external components close by. The communication tests can assist in determining if the patient is maintaining external components within a suggested proximity (e.g. 6 feet) of the patient and may assist to determine if transmission or reception difficulties are the source of communication failure. A physician programmer can provide for programming, conducting, summarizing and assessing the results of communication tests. | 1. A medical device system for monitoring a condition of a patient comprising: an implanted medical device;an external device comprising a transceiver, a random access memory, and a processor coupled to the transceiver and the random access memory, wherein the processor is configured to:periodically activate a wireless communication circuitry in an attempt to receive a plurality of proximity test signals from the implanted medical device;for each of the plurality of activations, determine whether a proximity signal has been received from the implanted device within a certain time after the activation, thereby generating a plurality of proximity test receipt results;apply a test to the plurality of proximity test receipt results to determine whether a proximity failure condition exists;operative when a proximity failure condition exists, take at least one of the following actions: store information regarding the proximity failure in the random access memory or store proximity test failures in the random access memory, determine that the proximity failures meet a selected first proximity test criterion and then send a proximity test failure condition notification to the patient when the proximity test failures in the random access memory meet the selected first proximity test criterion. 2. The medical device system according to claim 1 wherein the processor is further configured to send notification to a remote station when the proximity failures in the random access memory meet a selected second proximity test criterion. 3. The medical device system according to claim 1 wherein the periodic activation occurs every 24 hours. 4. The system of claim 1 further including a proximity test module which operates upon the random access memory to provide a log of the proximity test receipt results including a time of each test and a quantitative value related to a strength of a communication signal. 5. The system of claim 1 further including a proximity test module which operates upon the random access memory to provide a log of the proximity test receipt results including a time of each test and a qualitative value related to a quality of a communication signal. 6. The system of claim 1 wherein the implanted medical device is configured to provide a proximity test in conjunction with the external device and to derive and store proximity test receipt results. 7. The system of claim 6 further including a programmer device configured for downloading the proximity test receipt results of the implanted device and processing these proximity test receipt results to provide a test result profile. 8. The system of claim 6 further including a programmer device configured for downloading the proximity test receipt results of the implanted device and presenting the proximity test receipt results graphically to a user. 9. The system of claim 6 further including a programmer configured for running a diagnostic protocol which causes at least one of the implanted ... | Angel Medical Systems, Inc. | A61B005/0432 A61B005/00 A61B005/055 A61N001/372 G06F019/00
US10813564 | B2 | 2020-10-27 | Low field magnetic resonance methods and apparatus | 2015-11-11 | According to some aspects a system is provided comprising a low-field magnetic resonance (MR) device, at least one electrophysiological device, and at least one controller configured to operate the low-field MR device to obtain MR data and to operate the at least one electrophysiological device to obtain electrophysiological data. | 1. A system, comprising: a low-field magnetic resonance (MR) device comprising: a magnetics system having a plurality of magnetics components configured to produce magnetic fields for performing magnetic resonance imaging, the magnetics system comprising: a B0 magnet configured to produce a B0 field for an imaging region of the low-field MR device, the B0 magnet comprising at least one first B0 magnet and at least one second B0 magnet, which are positioned in a bi-planar arrangement about the imaging region; andat least one laminate panel comprising a plurality of laminate layers, each of the plurality of laminate layers including at least one non-conductive layer and at least one conductive layer patterned to form at least a portion of a gradient coil configured to produce a gradient magnetic field, wherein the at least one laminate panel comprises a first laminate panel comprising a first gradient coil and a second laminate panel comprising a second gradient coil, the first and second gradient coils of the first and second laminate panels being disposed on opposite sides of the imaging region;a power system comprising one or more power components configured to provide power to the magnetics system to operate the low-field MR device to perform image acquisition, the one or more power components comprising at least one power component configured to provide power to the first and second laminate panels;a base that supports the magnetics system and houses the power system; andat least one conveyance mechanism coupled to the base allowing the low-field MR device to be transported to different locations;at least one electrophysiological device; andat least one controller configured to operate the low-field MR device to obtain MR data and to operate the at least one electrophysiological device to obtain electrophysiological data. 2. The system of claim 1, wherein the at least one electrophysiological device comprises an electroencephalography (EEG) device comprising a plurality of electrodes configured to obtain EEG data while a patient is within a field of view of the low-field magnetic resonance device, and wherein the at least one controller is configured to operate the EEG device to obtain the EEG data. 3. The system of claim 2, wherein the at least one controller is configured to adapt at least one characteristic of the low-field magnetic resonance device based, at least in part, on the EEG device. 4. The system of claim 3, wherein the at least one controller is configured to adapt at least one characteristic of the low-field magnetic resonance device based, at least in part, on the EEG data. 5. The system of claim 4, wherein the at least one controller is configured to change a field of view of an MR acquisition process based on the EEG data. 6. The system of claim 4, wherein the at least one controller is configured to change a signal-to-noise ratio of an MR acquisition process based on the EEG data. 7. The system of claim 4, wherein the ... | Hyperfine Research, Inc. | A61B005/04 A61B005/0476 A61B005/00 A61B005/055 G01R033/38 G01R033/381 A61B005/0488 A61B005/0496 A61B005/0402 G01R033/48 G01R033/44 A61B090/00
US09946919 | B2 | 2018-04-17 | Device, system, and method for forensic analysis | 2015-11-19 | A portable, handheld device for forensic analysis includes an input unit configured to receive a forensic input at a scene of forensic analysis, and a processing unit. The processing unit is configured to generate a first image of the forensic input, scale the first image to a predetermined format suitable for forensic analysis, generate a second image of the forensic input in which the first image is scaled to the predetermined format, compare the second image with forensic records, which each respectively include previously collected forensic data associated with one or more individuals or specimens, determine whether the second image matches at least a portion of one of the forensic records, upon determining that the second image matches at least a portion of one of the forensic records, generate a matching notification identifying the record which the second image is determined to match, and output the generated matching notification. | 1. A portable, handheld device for forensic analysis, the device comprising: a memory unit having an executable program recorded therein;an input unit configured to receive a forensic input at a scene of forensic analysis at which the forensic input is collected; anda processing unit configured to, by executing the program recorded in the memory unit at the scene of forensic analysis, generate a first image of the received forensic input,scale the generated first image to a predetermined format suitable for forensic analysis, and generate a second image of the forensic input in which the first image is scaled to the predetermined format,record the second image of the forensic input in the memory unit,compare the second image with forensic records, each forensic record respectively including previously collected forensic data associated with one or more individuals or specimens,determine whether the second image matches at least a portion of at least one of the forensic records based on the comparison,upon determining that the second image matches at least a portion of at least one of the forensic records, generate a matching notification identifying the at least one record which the second image is determined to match, andoutput the generated matching notification on the device,wherein the input unit comprises: a laser light source configured to output at least two parallel laser lights on a surface of an object that includes the forensic input at the scene of forensic analysis, the at least two laser lights being output toward respectively different positions on the surface of the object, anda camera configured to generate the first image of the received forensic input that shows the at least two laser lights positioned on the surface of the object,wherein the processing unit is configured to determine at least one scaling dimension of the forensic input on the surface of the object contained in the first image based on a first positional relationship between the at least two laser lights output on the surface of the object and a second positional relationship between two points of interest in the first image. 2. The device according to claim 1, wherein the forensic records are recorded in the memory unit. 3. The device according to claim 1, comprising: a connection interface configured to receive the forensic records,wherein the processing unit is configured to record the received forensic records in the memory unit. 4. The device according to claim 3, wherein the connection interface includes at least one of: (a) a port configured to receive a wired connection to a computer to receive the forensic records, and(b) a transceiver configured to receive the forensic records via a wireless transmission. 5. The device according to claim 3, wherein the forensic records are received from a law enforcement database. 6. The device according to claim 1, wherein the input unit comprises a scanner configured to obtain a representation of at least one o... | BOOZ ALLEN HAMILTON INC. | G06K009/00 G06F001/16
US10524713 | B2 | 2020-01-07 | Identifying deceptive answers to online questions through human-computer interaction data | 2014-06-18 | The present invention provides a system and a method for eliciting information to sensitive questions and reliably detecting whether one is being deceptive, concealing information, or experiencing a heightened emotional response to the question. In particular, the system and the method of the invention are based on analyzing the user behavioral biometric of using one or more input device(s). | 1. A real-time behavioral biometric-based deception analysis system for determining whether a subject is truthful or deceptive to a question of interest that is presented on a display screen, said system comprising: a data interception unit configured to intercept input in real-time from a subject that is directed to a question presented on a display screen, wherein the data interception unit is configured to passively collect a pointing device usage characteristic in real-time;a behavior analysis unit operatively coupled to said data interception unit to receive the passively collected pointing device usage characteristic; anda behavior comparison unit operatively coupled to said behavior analysis unit, wherein said system dynamically monitors and passively collects behavioral biometric information in real-time, and translates said behavioral biometric information into representative data comprised of cursor trajectory data, wherein the representative data is normalized against population data and the cursor trajectory data is rescaled to account for screen resolution, stores and compares results, and outputs a result associated with truthfulness or deception of the subject to the question of interest presented on the display screen. 2. The behavioral biometric-based deception analysis system of claim 1, wherein said behavior comparison unit is operatively connected to an application for presenting a question on the display screen such that said behavior comparison unit influences the next question presented on the display screen using a decision tree structure based on the result. 3. The behavioral biometric-based deception analysis system of claim 1, wherein said pointing device usage characteristic comprises movement of said pointing device between the starting position of said pointing device and the answer selected by the subject on the display screen, the elapsed time between presentation of the question on the display screen and the selection of the answer by the subject, or a combination thereof. 4. The behavioral biometric-based deception analysis system of claim 3, wherein said pointing device usage characteristic comprises speed, total distance travelled, initial direction of movement, total response time, change in direction on the x-axis, change in direction on the y-axis, acceleration, idle time, area under the curve, amount of deviation, reaction time, changes in angle, or a combination of two or more thereof. 5. The behavioral biometric-based deception analysis system of claim 1, wherein the behavior comparison unit compares the result of the subject's behavioral biometric to a reference behavioral biometric. 6. The behavioral biometric-based deception analysis system of claim 5, wherein said reference behavioral biometric comprises the subject's behavioral biometric to a non-sensitive question. 7. The behavioral biometric-based deception analysis system of claim 5, wherein said reference behavioral biometric comprises an a... | The Arizona Board of Regents on Behalf of the University of Arizona | A61B005/16 G06F021/32 G06F021/62 G06Q010/10 G06F021/31 G06F021/40
US09326732 | B2 | 2016-05-03 | Electrocardiogram sensor and method of processing signals using the same | 2015-01-16 | An electrocardiogram (ECG) sensor includes an analog front end device configured to remove a DC offset in a first ECG signal received from a measuring electrode, and adjust a gain amplification value based on the first ECG signal. The analog front end device is configured to output an adjusted first ECG signal, the adjusted first ECG signal being based on the adjusted gain amplification value and the removed DC offset in the first ECG signal. The ECG sensor includes a digital signal processor configured to analyze and process the adjusted first ECG signal based on an algorithm and to output information. The analog front end device is configured to remove the DC offset in the first ECG signal during a first period, and remove a DC offset in a second ECG signal and simultaneously adjust a gain amplification value for the second ECG signal during a second period. | 1. An electrocardiogram (ECG) sensor comprising: an analog front end device configured to, remove a DC offset in a ECG signal received from a measuring electrode,adjust a gain amplification value based on the ECG signal, andoutput an adjusted ECG signal, the adjusted ECG signal being based on the adjusted gain amplification value and the removed DC offset in the ECG signal; anda digital signal processor configured to detect information by analyzing and processing the adjusted ECG signal based on an algorithm and to output the information,wherein the analog front end device is configured to, remove the DC offset in a first ECG signal during a first period, andremove a DC offset in a second ECG signal and adjust a gain amplification value for the second ECG signal during a second period, the gain amplification value for the second ECG signal being adjusted in response to an enable signal that is activated if the DC offset in the first ECG signal is within a first reference range. 2. The ECG sensor of claim 1, wherein the analog front end device comprises: an amplifier configured to receive the ECG signal from the measuring electrode and amplify the ECG signal based on the gain amplification value for the ECG signal;an analog-to-digital converter configured to perform analog-to-digital conversion on the amplified ECG signal and output a digital ECG signal; andan automatic gain controller configured to detect a peak value of the digital ECG signal and to adjust the gain amplification value for the ECG signal based on the peak value. 3. The ECG sensor of claim 2, wherein the analog front end device further comprises: a DC offset calculator configured to, calculate a mean value of the digital ECG signal,calculate a DC offset with respect to the mean value, andoutput the DC offset calculated with respect to the mean value; anda digital-to-analog converter configured to perform digital-to-analog conversion on the DC offset output from the DC offset calculator to output an analog DC offset, wherein the amplifier is configured to remove the DC offset from the ECG signal based on the analog DC offset. 4. The ECG sensor of claim 3, wherein the DC offset calculator is configured to determine whether the DC offset calculated with respect to the mean value is in the first reference range. 5. The ECG sensor of claim 4, wherein the DC offset calculator is configured to output a bit value corresponding to a changed DC offset with respect to the digital ECG signal if the DC offset calculated with respect to the mean value is not in the first reference range, the bit value being output to the digital-to-analog converter. 6. The ECG sensor of claim 2, wherein the automatic gain controller is configured to detect the peak value of the digital ECG signal and to determine whether the detected peak value is in a second reference range. 7. The ECG sensor of claim 6, wherein the automatic gain controller is configured to output a bit value corresponding to a changed ... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/04 A61B005/00 A61B005/0408
US09332940 | B1 | 2016-05-10 | Compact wearable biological sensor modules | 2015-01-05 | A sensor module can include a battery housing comprising a battery cavity sized and shaped to receive a battery. The battery cavity can be defined at least in part by a wall to be disposed about at least a portion of a periphery of the battery. A package housing can be disposed on the wall of the battery housing, the package housing smaller than the battery housing. An integrated device package can be disposed in or coupled with the package cavity. The integrated device package can include one or more integrated device dies. An interfacing feature can be coupled with the battery housing and extending transverse to the wall. The interfacing feature can be configured to transduce a biological signature into a signal to be processed by the integrated device package. An interconnect assembly can electrically connect the interfacing feature to the integrated device package. | 1. A sensor module comprising: a battery housing comprising a battery cavity sized and shaped to receive a battery, the battery cavity defined at least in part by a wall configured to be disposed about at least a portion of a periphery of the battery;a package housing disposed on the wall of the battery housing, the package housing smaller than the battery housing and comprising a package cavity;an integrated device package disposed in or coupled with the package cavity, the integrated device package comprising one or more integrated device dies;an interfacing feature coupled with the battery housing and extending transverse to the wall, the interfacing feature configured to transduce a biological signature into a signal to be processed by the integrated device package; andan interconnect assembly that electrically connects the interfacing feature to the integrated device package. 2. The sensor module of claim 1, wherein the interfacing feature comprises an electrode. 3. The sensor module of claim 1, wherein the interconnect assembly comprises a contact configured to electrically connect to an external device. 4. The sensor module of claim 3, wherein the contact is exposed on a bottom side of the battery housing by an aperture formed through the battery housing. 5. The sensor module of claim 4, further comprising a cover coupled with the battery housing over the battery cavity to define a top side of the battery housing, the cover comprising the interfacing feature. 6. The sensor module of claim 4, wherein the external device comprises the interfacing feature, the interfacing feature mechanically coupled to a platform on the bottom side of the battery housing, the interfacing feature electrically coupled with the contact through the aperture. 7. The sensor module of claim 1, wherein the package cavity is on a side of the wall opposite the battery cavity. 8. The sensor module of claim 1, further comprising a second interfacing feature coupled with the battery housing and extending transverse to the wall, the second interfacing feature disposed on a side of the battery housing opposite the interfacing feature. 9. The sensor module of claim 8, wherein the interfacing feature comprises a first electrode and the second interfacing feature comprises a second electrode. 10. The sensor module of claim 1, wherein the battery housing and the package housing comprise a unitary body. 11. The sensor module of claim 1, further comprising a third interfacing feature on an exterior surface of the integrated device package, the third interfacing feature comprising a capacitive touch sensor, an antenna, or an optical sensor. 12. The sensor module of claim 1, wherein the one or more integrated device dies comprises at least one of a processor die configured to analyze the signal to determine a heart rate of a user of the sensor module, a communications die configured to transmit data regarding the heart rate to an external computing device, and a motion ... | ANALOG DEVICES, INC. | A61B005/00 A61B005/024 A61B005/11
US09788822 | B2 | 2017-10-17 | Surgical access system and related methods | 2016-01-13 | A surgical access system including a tissue distraction assembly and a tissue retraction assembly, both of which may be equipped with one or more electrodes for use in detecting the existence of (and optionally the distance and/or direction to) neural structures before, during, and after the establishment of an operative corridor to a surgical target site. | 1. A retractor blade for use in a system for accessing a surgical target site, the retractor blade comprising: an elongated retractor blade extending from a proximal end that is dimensioned to couple to a mounting structure to a distal end that is spaced longitudinally from the proximal end, wherein the retractor blade includes a first edge extending between the proximal end and the distal end, a second edge extending between the proximal end and the distal end, an inner surface defined by the retractor blade between the first and second edges and comprising a portion configured to be positioned within an operative corridor established within a patient for accessing the surgical target site and oriented towards a central longitudinal axis of the operative corridor, and an outer face defined by the retractor blade between the first and second edges and comprising a portion configured to be oriented away from the central longitudinal axis, wherein the inner face of the retractor blade defines a receiving portion dimensioned to couple to an engagement portion of a connectable element, and wherein the inner surface of the retractor blade defines a plurality of grooves spaced longitudinally along the portion of the inner surface of the retractor blade and wherein the plurality of grooves are dimensioned to receive a tooth of the connectable element so as to hold the connectable element in place with respect to the retractor blade, wherein the plurality of grooves is on a concave portion of the inner surface of the retractor blade. 2. The retractor blade of claim 1, wherein the grooves defined by the inner surface of the retractor blade comprise a series of ratchet-like grooves configured to longitudinally position the connectable element with respect to the retractor blade. 3. The retractor blade of claim 1, wherein the grooves are positioned in series along a longitudinal center line of the inner surface of the retractor blade. 4. The retractor blade of claim1, wherein the grooves extend in a direction that is substantially perpendicular to a longitudinal axis of the retractor blade. 5. The retractor blade of claim 1, wherein the inner surface is substantially concave and the outer surface is substantially convex. 6. The retractor blade of claim 1, wherein the grooves comprise at least four grooves configured to receive the tooth engaged with a corresponding one of the at least four grooves to position the connectable element with respect to the retractor blade. 7. The retractor blade of claim 1, wherein the retractor blade tapers along at least part of the retractor blade between the first and second edges. 8. The retractor blade of claim 1, further comprising a receiving portion configured to engage with a corresponding engagement element of a shim element or a retractor extender. | NuVasive, Inc. | A61B001/32 A61B017/02 A61B090/00 A61B005/0488 A61B005/0492 A61B005/05 A61B005/11 A61B005/00 A61N001/05 A61B001/06 A61N001/08 A61B017/00
US09933298 | B2 | 2018-04-03 | Monitoring light exposure using a weighting function and light-monitoring system configured for user communication | 2016-01-09 | Light exposure from at least one light source is received with a light detector of a light monitor that includes at least one of (a) an output device and (b) a communication device transported by a user. The light detector converts the light exposure into an electrical signal, and the current time of day at which the light exposure is received is recorded. An instantaneous light exposure value is generated from the electrical signal, and a weighting function is applied to the instantaneous light exposure value as a function of the recorded time of day associated with the light exposure to produce a weighted instantaneous light exposure value. The weighted instantaneous light exposure value is integrated to produce a weighted cumulative luminous exposure value; and the weighted cumulative luminous exposure value is compared with an established luminous-exposure target. | 1. A method for monitoring light exposure, comprising: receiving light exposure from at least one light source with a light detector of a light monitor, wherein the light monitor further includes at least one of (a) an output device and (b) a communication device transported by a user;using the light detector to convert the light exposure into an electrical signal;recording current time of day as the light exposure is received;generating an instantaneous light exposure value from the electrical signal;applying a weighting function to the instantaneous light exposure value as a function of the recorded time of day associated with the light exposure to produce a weighted instantaneous light exposure value;integrating the weighted instantaneous light exposure value to produce a weighted cumulative luminous exposure value;comparing the weighted cumulative luminous exposure value with an established luminous-exposure target; andproviding an indication of the weighted cumulative luminous exposure value in comparison with the established luminous-exposure target. 2. The method of claim 1, wherein the luminous-exposure target is in the range of 12-24 million lux-seconds. 3. The method of claim 1, further comprising applying a minimum light-intensity threshold to the light exposure to exclude illuminance below the minimum light-intensity threshold. 4. The method of claim 1, further comprising: receiving input regarding at least one supplemental factor selected from (a) supplemental treatment received by the user and (b) user physical activity; anddetermining and communicating luminous exposure targets to the user based, in part, on the supplemental factor. 5. The method of claim 1, further comprising separately measuring and recording exposure to radiation in different wavelength ranges. 6. The method of claim 5, further comprising separately measuring and recording exposure to blue light. 7. The method of claim 5, wherein separate detectors respectively measure exposure to light in the visible wavelength range and in the ultraviolet wavelength range. 8. The method of claim 7, further comprising identifying recorded light exposure that is received when ultraviolet radiation exposure is also received as being sourced from natural sunlight. 9. The method of claim 7, further comprising evaluating the measured ultraviolet radiation exposure to estimate production of vitamin D3 by the user due to the user's light exposure. 10. The method of claim 9, further comprising evaluating the measured ultraviolet radiation exposure to estimate risk of the user suffering sunburn due to the user's light exposure. 11. The method of claim 7, further comprising factoring at least one of (a) the user's skin type or (b) the amount of the user's skin that is exposed in calculating effects of the user's light exposure. 12. The method of claim 1, wherein the modification is provided by applying a light-quality weighting factor to the recorded light exposure. 13. The m... | SunSprite | G01J001/42 G01J001/02 A61B005/00 G01J001/44 A61N005/06
US10524650 | B2 | 2020-01-07 | Scanable tray for producing a dental prosthesis | 2014-06-13 | A method for taking an impression for producing a dental prosthesis involves: A) generating an impression of regions of the upper jaw and lower jaw of a patient using an at least two-part tray, whereby the tray is inserted into the oral cavity of the patient;B) fixing the at least two tray parts to each other or against each other;C) determining the position and orientation of the at least two tray parts with respect to each other; andD) taking the tray parts out of the oral cavity of the patient. Also disclosed is the tray for implementing the method, whereby at least one marking is arranged on the surface of each tray part, and a device for intra-oral determination of the orientation and position of the marking and, thus, the at least two tray parts, is provided, whereby the device comprises a computer and a recording facility. | 1. A method for taking an impression of at least partial regions of an upper and lower jaw of a patient for producing a dental prosthesis, said method comprising the steps of: A) generating an impression of at least partial regions of the upper jaw and lower jaw of the patient using an at least two-part tray in an oral cavity of the patient, whereby the tray is inserted into the oral cavity of the patient for this purpose;wherein the at least two-part tray comprises at least one tray part for the upper jaw and at least one tray part for the lower jaw, whereby the tray parts are placed at a distance from each other by means of a screw or a pin, and are attachable at an adjustable distance from each other,wherein the pin, which comprises a mobile pestle that is supported against one of the tray parts as in a bearing in spring-loaded manner, or the screw is arranged on one of the tray parts, whereby the other tray part comprises a recess for accommodation of a tip of the screw or a tip of the pestle, whereby the recess comprises a matching counter-thread for the screw or the recess is a negative fit for the tip of the pestle, whereby the tip of the pestle is arranged appropriately mobile in the recess such that it can be inclined towards all sides,wherein at least one marking is arranged on a surface of each tray part and the markings are arranged on an edge of the tray parts such that the markings of adjacent tray parts that are inserted into the oral cavity of the patient are situated as close to each other such that the markings can be imaged on just one image of a recording device in the oral cavity of the patient,B) fixing the at least two tray parts to each other or against each other in the oral cavity of the patient;C) thereafter determining the position and orientation of the at least two tray parts with respect to each other in the oral cavity of the patient;wherein the position and orientation of the tray parts with respect to each other are determined by means of the at least one marking on each of the tray parts arranged on a labial surface of the tray parts, wherein the at least one marking is a three-dimensional contoured marking that permits tilting of the marking to be recognized and measured; wherein the position and orientation of the markings of the tray parts are determined with respect to each other and data obtained regarding the position and orientation of the markings are used to determine the position and orientation of the tray parts with respect to each other; and wherein the position and orientation of the tray parts are determined with respect to each other by intra-orally imaging or scanning with an optical method the labial surfaces of the tray parts, at least regions thereof, whereby at least two markings of at least two different tray parts are imaged or scanned concurrently, andD) thereafter taking the tray parts out of the oral cavity of the patient. 2. Method according to claim 1, which further comprises determi... | KULZER GMBH | A61B001/24 A61C009/00 A61B005/00 A61C013/00 A61B090/00
US09329758 | B2 | 2016-05-03 | Multiple sensory channel approach for translating human emotions in a computing environment | 2013-11-04 | Sensory inputs of a user can be received by a computing device. At least one of the sensory inputs can include a physiological input providing a physiological measurement from a body of the user. Each sensory input can be processed in a unique one of a set of standards-defined sensory channels, each corresponding to a specific emotion dimension. Processing the sensory inputs can transform the physiological measurement into an emotion dimension value. The emotion dimension values from each of the sensory channels can be aggregated to generate at least one emotion datum value, which is a standards-defined value for an emotional characteristic of the user. Historical data for a user can be optionally collected and used by a learning and calibration component to improve the accuracy of the generated emotion datum values for a specific individual. A programmatic action driven by the emotion datum value can be performed. | 1. A system for incorporating human emotions in a computing environment comprising: one or more processors, and a memory storing computer readable program code;a plurality of discrete sensory channels each for handling sensory input detected via one or more sensors, each of the one or more sensors comprising hardware, wherein different senses are associated with different ones of the discrete sensory channels, wherein each of the discrete sensory channels is a sensory channel corresponding to a specific emotion dimension, wherein sensory input handled within the sensory channels comprises physiological input providing a physiological measurement from a body of the user;a plurality of in-channel processors, each being one of the one or more processors, that process sensory input specific to the channel and that generate emotion dimension values from the sensory input, wherein each emotion dimension value has been transformed to be independent of idiosyncrasies of a sensory capture device from which the sensory input was originally obtained; anda sensory aggregator, implemented in the computer readable program code stored in the memory and executable by at least one of the one or more processors, for aggregating emotion dimension values generated in a per-channel basis by the in-channel processors to generate at least one emotion datum value, which is a value for an emotional characteristic of a user from whom the sensory input was gathered, and is an application independent value that is able to be utilized by a plurality of independent applications to discern emotions of said user and to cause application specific code of the independent applications to be reactive to changes in sensory aggregator generated emotion datum values. 2. The system of claim 1, further comprising: a plurality of sensory capture devices, wherein said sensory capture devices comprise a physiological sensor, an audio or video capture device, and a user input peripheral for manual user input, wherein input obtained from the physiological sensor, from the audio or video capture device, and from the user input peripheral is handled by separately by different ones of the discrete sensory channels, and wherein said physiological sensor obtains a measure from a system of the human body selected from a group of systems consisting of a cardiovascular system, a muscular system, a circulatory system, a nervous system, a respiratory system, an excretory system, an endocrine system, and a digestive system. 3. The system of claim 1, further comprising: a plurality of emotion data consumers each performing a programmatic actions driven by emotion datum values generated by the sensory aggregator. 4. The system of claim 3, wherein at least one of the emotion data consumers comprises a virtual world software application in which a plurality of humans interact though user specific avatars, wherein said user specific avatars express emotions corresponding to emotions of their users driven... | International Business Machines Corporation | G06F003/041 G06F003/0484 A61B005/16 A61B005/00 A61B005/01 A61B005/04 A61B005/053
US09662178 | B2 | 2017-05-30 | System and method for Decrement EvokEd Potential (DEEP) mapping to identify components of the arrythmogenic circuit in cardiac arrhythmias | 2014-05-16 | Various embodiments are described herein for a system and a method for identifying the arrhythmogenic circuit of a patient or subject. In one embodiment, the method comprises obtaining data for electrograms recorded at various locations of the heart while programmed ventricular pacing with extra stimuli was performed, obtaining decrement values for at least two different locations of the heart using the recorded electrograms, generating at least a portion of a decrement map using the decrement values, and identifying the arrhythmogenic circuit based on electrograms having significant decremental properties. | 1. A method for determining a location of a conduction block of a heart, the method comprising: obtaining data for electrograms recorded at various locations of the heart while programmed ventricular pacing with extra stimuli was performed;identifying sharp near-field potentials beyond a QRS complex for both a base train and the extra stimuli for at least two of the recorded electrograms;obtaining decrement values for at least two different locations of the heart where the decrement value for a given location of the heart is obtained by subtracting a time delay of the sharp near-field potential in response to the base train from a time delay of the sharp near-field potential in response to the extra stimuli in the recorded electrograms corresponding to the given location of the heart;generating at least a portion of a decrement map using the decrement values obtained for the different locations of the heart; andidentifying the location of the conduction block based on electrograms having significant decremental properties by comparing the decrement values obtained for the different locations of the heart with other decrement values or a threshold. 2. The method of claim 1, wherein S1 is the base train and the extra stimuli comprise S2 or S2 and S3, and wherein the sharp near-field potentials are identified beyond the QRS complex for both the base train and the extra stimuli. 3. The method of claim 2, wherein the method comprises obtaining decrement values for various locations of the heart by measuring a largest time difference of the sharp near-field potentials in response to at least one of the extra stimuli with respect to the stimuli S1 in the at least two recorded electrograms corresponding to the different locations of the heart. 4. The method of claim 1, wherein before identifying the sharp near-field potential, the method comprises identifying S1 and extra stimuli pacing artifacts in the recorded electrograms. 5. The method of claim 1, wherein the act of obtaining data for electrograms comprises recording one of bipolar electrograms, unipolar electrograms and virtual electrograms. 6. The method of claim 1, wherein the method further comprises identifying significant decremental properties by using decrement values of the decrement map and spatial gradients of the decrement values. 7. The method of claim 1, wherein the method further comprises obtaining decrement values until a largest decrement value is obtained indicating a longest delay. 8. The method of claim 1, wherein one or more unidirectional blocks are located by finding one or more regions on the decrement map with a long decrement that are adjacent to one or more regions that have a short decrement. 9. The method of claim 1, wherein the generated at least a portion of the decrement map is used as a guide for a subsequent cardiac ablation procedure. 10. The method of claim 1, wherein by the decrement values are corrected to a global decrement in surface ECG signals by su... | UNIVERSITY HEALTH NETWORK | A61B005/04 A61B034/10 A61B005/042 A61B005/0472 A61B005/0464 A61B005/0452 A61B005/00
US10588538 | B2 | 2020-03-17 | Medical imaging system having microwave emission/reception | 2014-06-06 | The invention relates to a medical imaging system having microwave emitting antennas and antennas for receiving the electromagnetic field, which are arranged around a space for receiving a human tissue medium to be observed, and comprising: an array (3) of emitting antennas and an array (4) of receiving antennas, wherein said two arrays (3, 4) are independent, and motors capable of angularly rotating and of translating the emitting array (3) and/or the receiving array (4) relative to the space under observation, in order to enable the scanning thereof. | 1. A medical imaging system with microwave transmission antennas and antennas of reception of the electromagnetic field arranged around a volume intended to receive human tissue medium to be observed, characterized in that it comprises: a shell (1) defining the volume intended to receive human tissue medium to be observed, the shell (1) having an external form, the external form of the shell (1) being symmetrical in rotation around a vertical axis, the shell (1) being made of solid dielectric material;an array (3) of transmission antennas (3a) and an array (4) of reception antennas (4a) for microwave imaging, these two arrays (3, 4) being independent, the antennas (3a) of transmission and the antennas (4a) of reception directly coupled to supports (3b, 4b), wherein each of the supports comprise material of dielectric permittivity identical to that of said shell (1); andone or more motors capable of shifting at least one of the transmission array (3) and the reception array (4) relative to the shell (1) in angular rotation around the vertical axis and in translation along the vertical axis to enable scanning of the volume under observation, the two arrays (3, 4) being independent in rotation and translation, the rotation and the translation of each array being independent. 2. The system according to claim 1, characterized in that said supports (3b, 4b) comprise parts made of electromagnetically absorbent materials between the antennas (3a, 4a). 3. The system according to claim 1, characterized in that the supports (3b, 4b) consist of entirely electromagnetically absorbent material. 4. The system according to claim 2, characterized in that said electromagnetically absorbent material further has dielectric permittivity identical to that of the shell (1). 5. The system according to claim 2, characterized in that the shell (1) is arranged in a container (2) made of dielectric material which the arrays (3, 4) enclose at least partially, the materials of the container (2), the shell (1) and the supports (3b, 4b) of arrays being identical. 6. The system according to claim 5, characterized in that the container (2) is externally cylindrical, at least one of the transmission array (3) and the reception array (4) moving in at least one of translation and rotation about the latter. 7. The system according to claim 1, characterized in that it comprises at least two containers (2). 8. The system according to claim 3, characterized in that the shell (1) is detachable, the system being connected to a set of shells of different internal dimensions. 9. The system according to claim 1, characterized in that it comprises at least one of two arrays (3) of reception antennas (3a) and at least two arrays (4) of transmission antennas (4a). 10. The system according to claim 9, characterized in that each of the arrays (3, 4) of antennas is mobile in translation in a limited way. 11. The system according to claim 1, characterized in that the array(s) (3) of recep... | MVG INDUSTRIES OFFICE NATIONAL D'ETUDES ET RECHERCHES AEROSPATIALES (ONERA) | A61B005/00 A61B005/05 G01S013/89
US09355307 | B2 | 2016-05-31 | Method and apparatus for real time performance assessment | 2014-08-27 | An apparatus for monitoring performance comprises two elements: a first element (4) for attaching to a human or animal, containing two or more of physiological, biochemical, kinematic, and/or environmental sensors and a processing means which records in memory signals or derivates of signals from said sensors and contains a means of re-transmitting this data to a second element, and a second element (13, 16) containing a display (12, 17) or audio output to present the processed live or previous recorded data to the user. | 1. A method of monitoring performance of a horse, the method using one or more reference signals derived from an apparatus for monitoring performance, the method comprising providing: (a) providing a girth for attaching to a horse, the girth comprising apparatus having two or more of physiological, biochemical, kinematic, and environmental sensors and a processor to record one or more signals from said sensors, wherein the sensors comprise at least an accelerometer sensor and a position sensor;(b) attaching the girth to the horse;(c) obtaining accelerometer data and position data using said two or more sensors;(d) the apparatus having a means of re-transmitting the data to an external device for storage, display or analysis;(e) transmitting said data to said external device;(f) using said accelerometer data and position data to determine at least a stride length of the horse by combining speed data derived from the position data from the position sensor and accelerometer data derived from the accelerometer sensor;wherein said sensors are only in the girth. 2. The method of claim 1, wherein identification chips are worn both by a human and by the horse, the identification chip possessing therein information about the horse selected from at least one of, name, age, weight, gender or patient ID, the identification chip with the human possessing therein information about the human including name and weight and the identification chip with the horse possessing therein information about the horse selected from at least one of, name, age, weight, gender or patient ID, where data acquisition is commenced when the horse or human passes through an electronic gate. 3. The method of claim 1, wherein identification chips are worn both by a human and by the horse, the identification chip possessing therein information about the horse selected from at least one of, name, age, weight, gender or patient ID, the identification chip with the human possessing therein information about the human including name and weight and the identification chip with the horse possessing therein information about the horse selected from at least one of, name, age, weight, gender or patient ID, where an apparatus for monitoring performance is configured when the horse or human passes through an electronic gate which is only large enough for one horse to pass through at a time which reads the identification chips and transmits that data to the said apparatus. 4. The method of claim 1, wherein the sensors comprise a heart rate sensor that detects electrical signals from the horse using conductive pads which are in contact with the horse's skin when the girth is fitted, wherein at least one of the conductive pads is located at a point substantially midway along the length of the girth another of the conductive pads is located near an end of the girth. 5. The method of claim 1, wherein the sensors comprise at least one of: a heart rate sensor;a skin temperature sensor;a blood gluco... | GMAX TECHNOLOGY LTD. | A63F009/24 A63F013/00 G06F017/00 G06F019/00 G06K009/00 A01K011/00 A01K029/00 A61B005/00 A61B005/11
US10448832 | B2 | 2019-10-22 | Acoustic computing systems for implant and dermal data communication, power supply and energy storage | 2014-11-06 | Embodiments are directed to a computing system that operates at least partially through a medium, which may be a human. The computing system includes a first acoustic computing element implanted in the human. The first acoustic computing element receives input acoustic vibration. The first acoustic computing element converts a first portion of the input acoustic vibration to energy that powers an operation of the first acoustic computing element. The first acoustic computing element converts a second portion of the input acoustic vibration to stored energy of the first acoustic computing element. The first acoustic computing element converts a third portion of the input acoustic vibration to a first input data that is processed by the first acoustic computing element. The first acoustic computing element generates and transmits a first output acoustic vibration based on the first input data. | 1. A computing system configured to operate at least partially through a medium, the computing system comprising: a first acoustic computing element comprising an acoustic input transducer communicatively coupled in parallel to a first rectification circuit, a first signal processor, and a first energy storage element;the first acoustic input transducer configured to receive, through the medium, input modulated acoustic vibration that has been modulated with data and instructions;the first acoustic input transducer further configured to convert the input modulated acoustic vibration to a first set of modulated alternating electric signals;the first acoustic input transducer further configured to communicatively couple the first set of modulated alternating electric signals to the first rectification circuit and the first signal processor;the first rectification circuit configured to convert the first set of modulated alternating electric signals to power and communicatively couple the power to the first signal processor and the first energy storage element;the first signal processor configured to demodulate the first set of modulated alternating electric signals to extract and process the data and instructions; andthe first acoustic computing element further configured to generate and transmit a first output acoustic vibration;wherein a length dimension of the first acoustic computing element and a length dimension of a wavelength of the input acoustic vibration are selected to result in an improvement to coupling efficiency of the input modulated acoustic vibration to the first acoustic computing element. 2. The computing system of claim 1 further comprising: a control system configured to operate external to the medium;wherein the control system is further configured to focus the input modulated acoustic vibration to concentrate focused input modulated acoustic vibration into the medium and toward the first acoustic computing element. 3. The computing system of claim 1, wherein: the medium comprises a live organism; andthe first acoustic element is configured to be implanted within the live organism. 4. The computing system of claim 1, wherein: the medium comprises a live organism; andthe first acoustic element is configured to operate within, under or on a dermis of the live organism. 5. The computing system of claim 1, wherein: the first acoustic computing element further comprises a first sensor communicatively coupled to the first rectification circuit, the first energy storage element and the first signal processor;the first sensor configured to receive the power from at least one of the first rectification circuit and the first energy storage element; andthe first signal processor further configured to control the first sensor to measure a parameter of the medium. 6. The computing system of claim 5, wherein the parameter of the medium comprises at least one of a concentration, a temperature, a pressure, a force, a magnetic field, an ... | INTERNATIONAL BUSINESS MACHINES CORPORATION | A61B005/05 A61B005/00 A61N001/372 A61N001/378 G08C023/02 A61B005/1473 A61B005/01 A61B005/03 A61M005/14
US20160120484 | A1 | 2016-05-05 | PULSE OXIMETERWITH VOICE BROADCAST FUNCTION | 2014-11-03 | A pulse oximeter with a voice broadcast function for voice-broadcasting pulse rate-related information and blood oxygen level-related information is introduced. The pulse oximeter includes a pulse blood oxygen sensor, a signal processing circuit, and a loudspeaker. The pulse blood oxygen sensor senses a user's pulse rate-related information and blood oxygen level-related information and sends a physiological signal carrying the pulse rate-related information and the blood oxygen level-related information. The signal processing circuit is electrically connected to the pulse blood oxygen sensor and adapted to receive and convert the physiological signal into the pulse rate-related information and the blood oxygen level-related information and send the pulse rate-related information and the blood oxygen level-related information. The loudspeaker is electrically connected to the signal processing circuit and adapted to receive and voice-broadcast the pulse rate-related information and the blood oxygen level-related information. | 1. A pulse oximeter with a voice broadcast function, comprising: a pulse blood oxygen sensor for sensing a pulse rate-related information and a blood oxygen level-related information and sending a physiological signal carrying the pulse rate-related information and the blood oxygen level-related information;a signal processing circuit electrically connected to the pulse blood oxygen sensor and adapted to receive the physiological signal, convert the physiological signal into the pulse rate-related information and the blood oxygen level-related information, and send the pulse rate-related information and the blood oxygen level-related information; anda loudspeaker electrically connected to the signal processing circuit and adapted to receive the pulse rate-related information and the blood oxygen level-related information and voice-broadcast the pulse rate-related information and the blood oxygen level-related information. 2. The pulse oximeter of claim 1, wherein the pulse blood oxygen sensor comprises a light emitting unit, a light receiving unit, and an optical signal processing unit, wherein the light emitting unit and the light receiving unit are disposed at two opposite sides of a portion of a user's body, wherein the light emitting unit emits an incident light beam to penetrate the user's body and exit the user's body in form of an outgoing light beam, wherein the light receiving unit receives the outgoing light beam and sends a data indicative of light intensity of the outgoing light beam, wherein the optical signal processing unit is electrically connected to the light receiving unit and adapted to receive and convert the data indicative of the light intensity of the outgoing light beam into the physiological signal. 3. The pulse oximeter of claim 2, wherein the light emitting unit comprises a red light source and an infrared light source, the red light source is a red light-emitting diode, the infrared light source is an infrared light-emitting diode. 4. The pulse oximeter of claim 1, wherein the pulse blood oxygen sensor comprises a light emitting unit, a light receiving unit, and an optical signal processing unit, wherein the light emitting unit and the light receiving unit are disposed on a same side of a portion of a user's body, wherein the light emitting unit emits an incident light beam, wherein the incident light beam falls on the user's body and then reflects off the user's body in form of a reflecting light beam, wherein the light receiving unit receives the reflecting light beam and sends a data indicative of light intensity of the reflecting light beam, wherein the optical signal processing unit is electrically connected to the light receiving unit and adapted to receive and convert the data indicative of the light intensity of the reflecting light beam into a physiological signal. 5. The pulse oximeter of claim 4, wherein the light emitting unit comprises a red light source and an infrared light source, the red light sou... |  | A61B005/00 A61B005/1455
US20160144166 | A1 | 2016-05-26 | MEDICAL LEAD WITH THIN FILM | 2015-11-24 | A medical lead includes an elongated carrier, a thin film attached to the elongated carrier, the thin film including a plurality of electrodes, a plurality of electrical contacts, and a plurality of conducting tracks, each of the plurality of conducting tracks providing an electrical connection between at least one of the plurality of electrodes and one of the plurality of electrical contacts; and a frame element including a fixation zone for the plurality of electrical contacts of the thin film. | 1. A medical device system for at least one of delivery of electrical stimulation pulses or sensing of physiological signals, the system comprising: an elongated carrier;a thin film attached to the elongated carrier, the thin film including a plurality of electrodes, a plurality of electrical contacts, and a plurality of conducting tracks, each of the plurality of conducting tracks providing an electrical connection between at least one of the plurality of electrodes and one of the plurality of electrical contacts; anda frame element including a fixation zone for the plurality of electrical contacts of the thin film. 2. The system of claim 1, wherein the thin film is wound on the elongated carrier in a helical fashion. 3. The system of claim 1, wherein the fixation zone is substantially tangential to an outer cylindrical surface of the elongated carrier. 4. The system of claim 1, wherein the frame element includes a plate portion, wherein the fixation zone is on a surface of the plate portion. 5. The system of claim 1, wherein the frame element includes a mounting portion with which the frame element is mounted on the elongated carrier. 6. The system of claim 5, wherein the mounting portion includes at least one mounting tube portion or at least one mounting tube segment portion, which is at least partially wound around the elongated carrier. 7. The system of claim 1, further comprising an active lead can including a switch matrix and a substantially sealed housing containing the switch matrix, wherein the frame element is mounted to the active lead can. 8. The system of claim 7, wherein the active lead can further includes feedthroughs extending through the substantially sealed housing, the system further comprising an interposer that provides electrical connection paths between the plurality of electrical contacts of the thin film and the feedthroughs. 9. The system of claim 7, further comprising an implantable medical device coupled to the active lead can via a lead. 10. The system of claim 9, wherein N electrical connection tracks are defined between the implantable medical device and the active lead can, and wherein N is less than a total number of individual electrodes of the plurality of electrodes. 11. The system of claim 10, wherein the active lead can includes a pulse generator. 12. The system of claim 10, wherein the implantable medical device includes a pulse generator. 13. The system of claim 1, wherein a proximal end of the thin film is electrically coupled to an interposer and affixed to the fixation zone of the frame element. 14. The system of claim 1, wherein a proximal end of the thin film is not wrapped around the elongated carrier. 15. A method of manufacturing a medical lead, the method comprising: assembling a thin film to an elongated carrier, the thin film including a plurality of electrodes, a plurality of electrical contacts, and a plurality of conducting tracks, each of the plurality of conducting tracks p... |  | A61N001/05 A61B005/00 A61N001/375 A61B005/04
US09820698 | B2 | 2017-11-21 | Actigraphy methods and apparatuses | 2015-11-06 | An actigraphy method includes receiving a physiological parameter signal as a function of time for a physiological parameter other than body motion (such as electrocardiography or a respiration monitor), computing a body motion artifact (BMA) signal as a function of time from the physiological parameter signal (for example, using a local signal power signal, a local variance signal, a short-time Fourier transform, or a wavelet transform over epochs of duration on order a few minutes or less), and computing an actigraphy signal as a function of time from the BMA signal, for example by applying a linear transform to the BMA signal and optionally applying filtering such as median removal and/or high-pass filtering. | 1. A physiological monitoring device comprising: a sensor configured to generate a non-body motion physiological parameter signal as a function of time for a physiological parameter other than velocity, displacement, and acceleration; andan electronic digital signal processing (DSP) device configured to perform operations including: computing a body motion artifact (BMA) signal as a function of time from the non-body motion physiological parameter signal, andcomputing an actigraphy signal as a function of time from the BMA signal. 2. The physiological monitoring device of claim 1 wherein the sensor includes an electrocardiography (ECG) sensor and the physiological parameter includes one or more of (i) at least one ECG trace and (ii) a heart rate. 3. The physiological monitoring device of claim 1 wherein the sensor includes a respiratory sensor and the physiological parameter includes a respiration rate. 4. The physiological monitoring device of claim 1 wherein the sensor includes a Respiratory Inductive Plethysmography (RIP) sensor. 5. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a local signal power signal from the non-body motion physiological parameter signal. 6. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a local signal variance signal from the non-body motion physiological parameter signal. 7. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a Short-Time Fourier Transform (STFT) signal from the non-body motion physiological parameter signal. 8. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a wavelet transform signal from the non-body motion physiological parameter signal. 9. The physiological monitoring device of claim 1 wherein computing a BMA signal as a function of time from the non-body motion physiological parameter signal comprises computing a BMA signal sample for each time window of a succession of time windows of the non-body motion physiological parameter signal. 10. The physiological monitoring device of claim 9 wherein the succession of time windows is a succession of overlapping Hamming time windows. 11. The physiological monitoring device of claim 1 wherein computing an actigraphy signal as a function of time from the BMA signal comprises applying a linear transform to the BMA signal. 12. The physiological monitoring device of claim 1 wherein the DSP device is configured to perform further operations including: filtering the actigraphy signal using a median removal filter. 13. The physiological monitoring ... | Koninklijke Philips N.V. | A61B005/02 A61B005/00 A61B005/0245 A61B005/0402 A61B005/11 A61B005/113 A61B005/08
US10092223 | B2 | 2018-10-09 | Measurement system | 2014-06-24 | This measurement system quantitatively presents the effect of wearing an acoustic device that transmits sound to a user by contacting a vibrating body to a human auricle. In a measurement system including an ear model, which is provided with an artificial auricle and an artificial external ear canal, and an air-conducted sound gauge that measures air-conducted sound in the artificial external ear canal, while an acoustic device is contacted to the ear model of the measurement system, the acoustic device including a vibrating body and transmitting sound to a user by contacting the vibrating body to a human auricle, the measurement system measures, with the air-conducted sound gauge, air-conducted sound generated by the acoustic device and presents the result of measurement together with a characteristic corresponding to when the user wears the acoustic device. | 1. A measurement system for an acoustic device including a vibrating body, wherein the acoustic device is configured to transmit sound to a user by contacting the vibrating body to a human auricle, and the acoustic device is configured to be worn by the user, the measurement system comprising: an ear model including an artificial auricle and an artificial external ear canal; and an air-conducted sound gauge configured to measure air-conducted sound in the artificial external ear canal, wherein the acoustic device is contacted to the ear model of the measurement system, the measurement system takes a measurement, with the air-conducted sound gauge, of air-conducted sound generated by the acoustic device and outputs a result of the measurement together with a characteristic of the acoustic device corresponding to when the user wears the acoustic device, and due to the vibration of the vibrating body, at least an inner wall of the artificial external ear canal vibrates, and artificial external ear canal radiated sound produced by at least the inner wall reaches and is measured by the air-conducted sound gauge. 2. A measurement system for an acoustic device including a vibrating body, wherein the acoustic device is configured to transmit sound to a user by contacting the vibrating body to a human auricle, and the acoustic device is configured to be worn by the user, the measurement system comprising: an ear model including an artificial auricle and an artificial external ear canal; and a vibration sound gauge configured to measure vibration sound in the ear model, wherein while the acoustic device is contacted to the ear model of the measurement system, the measurement system takes a measurement, with the vibration sound gauge, of vibration sound generated by the acoustic device and outputs a result of the measurement together with a characteristic of the acoustic device corresponding to when the user wears the acoustic device, wherein the vibration sound gauge is disposed annularly on an external surface of the artificial external ear canal on an opposite side from the artificial auricle. 3. A measurement system for an acoustic device including a vibrating body, wherein the acoustic device is configured to transmit sound to a user by contacting the vibrating body to a human auricle, and the acoustic device is configured to be worn by the user, the measurement system comprising: an ear model including an artificial auricle and an artificial external ear canal; an air-conducted sound gauge configured to measure air-conducted sound in the artificial external ear canal; and a vibration sound gauge configured to measure vibration sound in the ear model, wherein while the acoustic device is contacted to the ear model of the measurement system, the measurement system takes a measurement, with the air-conducted sound gauge and the vibration sound gauge, of air-conducted sound and vibration sound generated by the acoustic device and outputs combined sound,... | KYOCERA Corporation | A61B005/12 A61B005/00 H04R025/00 H04R011/04 H04R001/10
US20160135700 | A1 | 2016-05-19 | ECG FEATURES FOR TYPE AHEAD EDITING AND AUTOMATIC UPDATE FOR REPORT INTERPRETATION | 2014-06-09 | An electrocardiogram (ECG) device includes a processor (114), a memory (116) coupled to the processor and an interface (120) configured to generate an edit window and configured to permit user input in the edit window. A type ahead module (124) is stored in the memory and is responsive to input entered in the edit window. The type ahead module is configured to predict a next word, phrase, interpretation code or category being input by a user based on a subset of alphanumeric key strokes and a context probability associated with a list of possible ECG entries. A report review and edit module (128) may also be stored in the memory. The report review and edit module is configured to update a severity of an ECG report based upon current ECG interpretation statements and rules in accordance with an access or edit operation by an operator. | 1. An electrocardiogram (ECG) device, comprising: a processor;a memory coupled to the processor;an interface configured to generate an edit window and configured to permit user input in the edit window; anda type ahead module stored in the memory and responsive to input entered in the edit window, the type ahead module being configured to predict a next word, phrase, interpretation code or category being input by a user based on a subset of alphanumeric key strokes and a context probability associated with a list of possible ECG entries. 2. The device as recited in claim 1, wherein the context probability includes a type of test or set of tests being performed as a factor. 3. The device as recited in claim 1, wherein the context probability includes a regression of ECG related features from a collection of records as a factor. 4. The device as recited in claim 1, further comprising adjustable fiducials configured to be moved by a user to recompute the context probability. 5. The device as recited in claim 1, wherein a factor of the context probability includes one of a location of a test and an operator of the test. 6. The device as recited in claim 1, further comprising historic data configured to provide one of frequency and manner of use for the list of possible ECG entries to influence the predictions of the type ahead module. 7. The device as recited in claim 1, wherein the user input includes ECG interpretation statements. 8. An electrocardiogram (ECG) device, comprising: a processor;a memory coupled to the processor;an interface configured to generate an edit window and configured to permit user input in the edit window; anda report review and edit module stored in the memory, the report review and edit module being configured to update a severity of an ECG report based upon current ECG interpretation statements and rules stored in the memory and triggered in accordance with an access or edit operation by an operator. 9. The device as recited in claim 8, wherein the report review and edit module updates the severity upon addition, deletion or change of an ECG statement. 10. The device as recited in claim 8, wherein the rules include assigning a severity based upon an entire ECG interpretation of the current ECG interpretation. 11. The device as recited in claim 8, wherein the rules include assigning a severity based upon a worst severity of the current ECG interpretation. 12. The device as recited in claim 8, wherein the rules include assigning a severity based upon a subset of statements of the current ECG interpretation. 13. The device as recited in claim 8, further comprising a report output from the ECG device including an updated severity based upon the rules. 14. A method for automatic severity updating for an electrocardiogram (ECG), comprising: providing an ECG interpretation based upon ECG testing;triggering a review and edit application upon accessing the ECG interpretation;reviewing the ECG interpretation statement ... |  | A61B005/044 G06F017/27 G06F003/0484 G06F003/0489 G06F003/0481 G06N007/00 G06F017/24
US09326681 | B2 | 2016-05-03 | Non-contact tonometer, control method of the same, and program | 2014-04-07 | An excessive amount of pressurized air is inhibited from being blown to an eye of a patient during measurement of a corneal shape deformation amount. In a non-contact tonometer having a piston, a driving unit to drive the piston, a fluid ejection unit to blow air pressurized by driving the piston toward a cornea of an eye to be inspected, and an intraocular pressure measuring unit to detect a deformed state of the cornea and measure an intraocular pressure, there are provided a piston displacement restricting unit that restricts a displacement amount when the piston pressurizes the air, and a changing unit to change the displacement amount to be restricted by the piston displacement restricting unit. | 1. A non-contact tonometer comprising: a piston;a driving unit which is connected to the piston so as to drive the piston;a fluid ejection unit configured to blow air pressurized by driving the piston, which is disposed in a cylinder, toward a cornea of an eye to be inspected;an intraocular pressure measuring unit configured to measure an intraocular pressure of the eye by optically detecting a deformed state of the cornea; anda stopper that restricts a displacement range of the piston,wherein the restricted displacement range of the piston is changeable according to a measurement range of the intraocular pressure. 2. The non-contact tonometer according to claim 1, wherein the stopper is disposed in a movable range of at least one of the piston and the driving unit. 3. The non-contact tonometer according to claim 1, wherein the restricted displacement range of the piston is changed based on a measurement range of intraocular pressure values that are to be measurable. 4. The non-contact tonometer according to claim 3, wherein the measurement range of the intraocular pressure values that are to be measurable is changed based on a measurement result of a previous time. 5. The non-contact tonometer according to claim 1, wherein the stopper can be moved in a direction of movement of the piston so as to adjust a position where the stopper comes in contact with a driving lever of the driving unit. 6. The non-contact tonometer according to claim 1, wherein the driving unit comprises a rotary solenoid and a driving lever, wherein the fluid ejection unit comprises the cylinder, a transfer tube, and an air chamber, andwherein the intraocular pressure measurement unit comprises a measuring light source and a light receiving element. 7. The non-contact tonometer according to claim 1, wherein the driving unit comprises a solenoid, wherein the fluid ejection unit comprises the cylinder, andwherein the intraocular pressure measurement unit comprises a light source. 8. A non-contact tonometer comprising: a piston;a driving unit which is connected to the piston so as to drive the piston;a fluid ejection unit configured to blow air pressurized by driving the piston, which is disposed in a cylinder, toward a cornea of an eye to be inspected;an intraocular pressure measuring unit configured to measure an intraocular pressure of the eye by optically detecting a deformed state of the cornea; anda piston displacement restricting unit which includes (1) a band having an end portion connected to the piston and (2) a windup portion that changes a length of the band to be played out so as to restrict a displacement range of the piston. 9. A non-contact tonometer comprising: a piston disposed in a cylinder, wherein an intraocular pressure is measured by (1) blowing air pressurized by driving the piston in the cylinder toward a cornea of an eye to be inspected, and (2) detecting a deformed state of the cornea of the eye to be inspected; anda stopper that restricts a d... | Canon Kabushiki Kaisha | A61B005/00 A61B003/16
US09352209 | B2 | 2016-05-31 | Personal activity tracking system | 2014-01-27 | An activity tracking device is defined to detect and record movement of a person to which the activity tracking device is affixed. A secondary electronic device is defined separate from the activity tracking device. The secondary electronic device is defined to receive data from the activity tracking device regarding the detected and recorded movement of the person. A website is defined to communicate with connected devices including the secondary electronic device. The website is defined to provide a user account for the person. The website is defined to convey information regarding the detected and recorded movement of the person through the user account. | 1. A system, comprising: an activity tracking device defined to detect and record physical activity of a person to which the activity tracking device is affixed, the activity tracking device including circuitry configured to detect a first type of movement of the person and a second type of movement of the person, the activity tracking device including circuitry configured to distinguish the first type of movement of the person from the second type of movement of the person, the activity tracking device including circuitry configured to measure an amount of the first type of movement of the person and an amount of the second type of movement of the person, the activity tracking device including a plurality of visual indicators configured to indicate the amount of the first type of movement of the person and the amount of the second type of movement of the person; anda computer system configured to provide an application for execution on a secondary electronic device defined separate from the activity tracking device, the application defined to receive data from the activity tracking device regarding the amount of the first type of movement of the person and the amount of the second type of movement of the person as detected, distinguished, and measured by the activity tracking device, the application defined to operate in part based on the amount of the first type of movement of the person and the amount of the second type of movement of the person, the application defined to perform one or more actions associated with the first type of movement of the person over time, and the application defined to perform one or more actions associated with the second type of movement of the person over time. 2. A system as recited in claim 1, wherein the activity tracking device is connected to a band defined to be worn on a wrist of the person. 3. A system as recited in claim 2, wherein the activity tracking device is removable from the band. 4. A system as recited in claim 1, wherein the activity tracking device includes a motion detection device having circuitry for detecting movement of the activity tracking device. 5. A system as recited in claim 1, wherein the activity tracking device includes a display for visually conveying information regarding the detected and recorded physical activity of the person. 6. A system as recited in claim 5, wherein the plurality of visual indicators are part of the display. 7. A system as recited in claim 1, wherein the activity tracking device includes communication circuitry for communicating data regarding the detected and recorded physical activity of the person from the activity tracking device to the secondary electronic device. 8. A system as recited in claim 7, wherein the communication circuitry is defined to communicate data wirelessly to the secondary electronic device. 9. A system as recited in claim 7, wherein the communication circuitry is defined to communicate data through an electrical contact to... | Fibit, Inc. | A63B071/06 A61B005/11 A61B005/145 A61B005/00 G06Q030/02 G09F003/00 A63F013/20 G06F019/00 G07F017/32 A63B024/00
US09326816 | B2 | 2016-05-03 | Neuromodulation systems having nerve monitoring assemblies and associated devices, systems, and methods | 2013-08-30 | Neuromodulation systems with nerve monitoring assemblies and associated devices, systems, and methods are disclosed herein. A neuromodulation system configured in accordance with some embodiments of the present technology can include, for example, a generator, a nerve monitoring assembly configured to detect electroneurogram (ENG) signals, and a neuromodulation catheter. The neuromodulation catheter can include an elongated shaft with a distal portion and a proximal portion. The distal portion of the shaft can include an array of electrodes configured to detect nerve activity from within a blood vessel of a human. The proximal portion of the shaft can include at least one connector that operably couples the electrodes to the generator and to the nerve monitoring assembly. | 1. A method of monitoring nerve activity associated with treating a nerve, the method comprising: delivering a distal portion of a neuromodulation catheter to a target site within a blood vessel of a human patient, wherein the distal portion comprises a plurality of contacts;detecting electroneurogram (ENG) signals with a nerve monitoring assembly operably coupled to the contacts, wherein the ENG signals are taken from recordings of nerve activity at the contacts;distinguishing electromyogram (EMG) signals from ENG signals with the nerve monitoring assembly;delivering neuromodulation energy to the target site via an energy delivery element;detecting ENG signals with the nerve monitoring assembly proximate to the target site after delivering the neuromodulation energy; andcomparing the ENG signals recorded before and after the neuromodulation energy is delivered, wherein a decrease in a parameter of the ENG signals after neuromodulation indicates sufficient target of nerves proximate to the target site. 2. The method of claim 1 wherein the contacts are electrodes, and wherein: detecting ENG signals comprises recording nerve activity with the electrodes; anddelivering neuromodulation energy to the target site comprises delivering RF energy to the target site via the electrodes. 3. The method of claim 1 wherein the detecting ENG signals comprises: recording neural signals at a plurality of locations proximate to the target site; andaveraging the neural signals taken at the plurality of locations. 4. The method of claim 1, further comprising: stimulating the nerves proximate to the target site with a first contact of the plurality of contacts, wherein the nerves are stimulated with a non-therapeutic level of energy; andrecording nerve activity resulting from the stimulation with at least a second contact and a third contact of the plurality of contacts. 5. The method of claim 1 wherein distinguishing EMG signals from ENG signals with the nerve monitoring assembly comprises distinguishing the EMG signals from the ENG signals using at least one amplifier coupled to the contacts in a quasi-tripole (QT) circuit. 6. The method of claim 1 wherein distinguishing EMG signals from ENG signals with the nerve monitoring assembly comprises distinguishing the EMG signals from the ENG signals using a plurality of amplifiers coupled to the contacts in a true-tripole (TT) circuit. 7. The method of claim 1 wherein distinguishing EMG signals from ENG signals with the nerve monitoring assembly comprises distinguishing the EMG signals from the ENG signals using a plurality of amplifiers coupled to the contacts in an adaptive tripole (AT) circuit. 8. The method of claim 1 wherein the energy delivery element comprises an array of electrodes, wherein the electrodes define the contacts, and wherein delivering neuromodulation energy to the target site via the energy delivery element comprises delivering RF energy via the electrodes to a vessel wall at the target site.... | Medtronic Ardian Luxembourg S.a.r.l. | A61B018/14 A61B017/32 A61B018/02 A61B005/00 A61B005/04 A61N001/05 A61N001/36 A61B017/00 A61B018/00
US10278651 | B2 | 2019-05-07 | Early confirmation of a prolonged aperiodic measurement | 2014-06-18 | A medical device (10), medical method and graphical user interface (GUI) (18) displaying data regarding measurement of a physiological parameter. A window (22) for the physiological parameter is displayed using a display device (20). The window (22) includes an indicator (24) identifying a measured value of the physiological parameter. In response to automatic trigger of measurement of the physiological parameter, the window (22) is updated to further include a first indicator (34) indicating measurement of the physiological parameter is ongoing. In response to manual trigger of measurement of the physiological parameter, the window (22) is updated to replace the indicator (24) identifying the measured value with a second indicator (36) indicating measurement of the physiological parameter is ongoing. The second indicator (36) is prominent relative to the first indicator (34). | 1. A medical device displaying data regarding measurement of a physiological parameter, said medical device comprising: a display device; andat least one processor programmed to: display a window for the physiological parameter using the display device, the window including a third indicator identifying a measured value of the physiological parameter;in response to an automatic trigger of measurement of the physiological parameter, update the window to further include a first indicator, having a first format and being static over time, indicating measurement of the physiological parameter is ongoing; andin response to a manual trigger of measurement of the physiological parameter, update the window to replace the third indicator identifying the measured value with a second indicator, having a second format, indicating measurement of the physiological parameter is ongoing;wherein the format of the second indicator is different from the format of the first indicator. 2. The medical device according to claim 1, wherein the second indicator is dynamic over time. 3. The medical device according to claim 1, wherein the at least one processor is further programmed to: during an automatically triggered measurement of the physiological parameter, update the window to include an indicator identifying a preliminary value estimating the physiological parameter. 4. The medical device according to claim 3, wherein the at least one processor is further programmed to: during the automatically triggered measurement of the physiological parameter, receive acceptance of the preliminary value;in response to receiving acceptance of the preliminary value, end the automatically triggered measurement and update the window to replace the second indicator indicating measurement of the physiological parameter is ongoing with the indicator identifying the measured value, wherein the measured value is the preliminary value. 5. The medical device according to claim 3, wherein the at least one processor is further programmed to: in response to the automatically triggered measurement failing to meet predetermined criteria within a predetermined amount of time, suspend the automatically triggered measurement and wait up to a predetermined amount of time for a user of the medical device to accept the preliminary value;in response to failing to receive acceptance of the preliminary value within the predetermined amount of time, abort the automatically triggered measurement and update the window to replace the second indicator indicating measurement of the physiological parameter is ongoing with an indicator indicating the automatically triggered measurement was aborted. 6. The medical device according to claim 1, wherein the at least one processor is further programmed to: automatically trigger measurement of the physiological parameter according to a timing sequence; andin response to completion of a measurement of the physiological parameter, update the window include an in... | KININKLIJKE PHILIPS N.V. | A61B005/00 G06F019/00 A61B005/01 A61B005/021 A61B005/022 A61B005/08
US09924298 | B2 | 2018-03-20 | Wireless ambulatory care network | 2014-10-31 | A method for connecting wireless devices using Bluetooth® low energy can be used to connect wireless computing devices with a medical device in an ambulatory care setting. The wireless computing devices can receive measured vital sign data and transmit the vital sign data to a patient's electronic medical record. The wireless computing devices can be configured to include a user-friendly name and the transmitter power in one or more advertising broadcasts. The medical device can receive advertising broadcasts from one or more wireless computing devices and sort the devices by proximity, last use, number of times used, or alphabetically. After the devices exchange information over Bluetooth® low energy, the devices can connect in Bluetooth® Classic. Systems for implementing the method are also disclosed. | 1. A method for connecting with a wireless device, comprising: receiving, at a host medical computing device, a Bluetooth® low-energy broadcast from one or more wireless devices that are in a range of the host medical computing device, wherein the host medical computing device is configured to receive vital signs data of a patient, the vital signs data being obtained from at least one of: a thermometer, a heart rate monitor, a pulse oximeter, a non-invasive blood pressure monitor, and a respiration rate monitor; andwherein the Bluetooth® low-energy broadcast includes a user-friendly alias that corresponds to a media access control (MAC) address of the one or more wireless devices;identifying the one or more wireless devices by associating the MAC address with the user-friendly alias;sorting the one or more wireless devices;pre-pairing the one or more wireless devices with the host medical computing device;on a device selection display, displaying a name for each of the one or more wireless devices, wherein the name displayed corresponds to the user-friendly alias of the one or more wireless devices; andconnecting to one of the one or more wireless devices as a background process, including: closing the device selection display; andrecording vital sign measurements while connecting to the one or more wireless devices. 2. The method for connecting with a wireless device of claim 1, further comprising receiving a transmitter power from at least one of the one or more wireless devices. 3. The method for connecting with a wireless device of claim 2, wherein the one or more wireless devices are sorted by the received transmitter power from at least one of the one or more wireless devices. 4. The method for connecting with a wireless device of claim 1, further comprising: determining the signal strength of the one or more wireless devices. 5. The method for connecting with a wireless device of claim 4, further comprising: receiving a transmitter power from at least one of the one or more wireless devices; anddetermining a distance of the one or more wireless devices from the host medical computing device using the signal strength and the transmitter power of the one or more wireless devices. 6. The method for connecting with a wireless device of claim 5, wherein the one or more wireless devices are sorted using the distance from each wireless device to the host medical computing device. 7. The method for connecting with a wireless device of claim 1, wherein the name comprises up to 28 bytes. 8. The method for connecting with a wireless device of claim 1, wherein the one or more wireless devices are sorted by the order in which the wireless devices were last connected to the host medical computing device; and further comprising pre-selecting one of the one or more wireless devices based on the wireless device most recently connected to the host medical computing device. 9. A system for communicating vital signs data wirelessly, comprising: a host... | Welch Allyn, Inc. | H04B007/00 H04W004/00 H04W076/02 A61B005/00
US09351670 | B2 | 2016-05-31 | Glycemic risk determination based on variability of glucose levels | 2012-12-31 | A system and method for determining glycemic risks based on an analysis of glucose data includes visualization of hypoglycemia, variability, and hyperglycemia with a control grid, increasing the accuracy of glucose estimates using a combination of CGM and HbA1c, calculating glycemic risk by applying a probability distribution, and tailoring SMBG test schedules based on CGM use/wear. | 1. A system for displaying glycemic information and risk, the system comprising: a non-volatile memory in which is stored a glucose data processing program configured to program a processor to analyze received glucose data of a person and from the analysis, produce a display;a display device; anda processor, coupled with the nonvolatile memory and the display device, the processor being configured to access the memory to load and run in the processor the glucose data processing program to analyze glucose data, wherein the processor is programmed to: analyze the received glucose data to determine a plurality of glucose medians, low range variability of glucose, and hypoglycemia risk; andcause the display device to visually present a first display and a second display at the same time,wherein the first display comprises a graph having a vertical axis representing glucose levels, a horizontal axis representing a plurality of times of the day, and a plot of the glucose median of received glucose data across the plurality of times of the day, andwherein the second display comprises a plurality of hypoglycemia risk indicators and a plurality of variability risk indicators,wherein each hypoglycemia risk indicator indicates an amount of risk that hypoglycemia will occur in a different segment of time within the plurality of times of the day and each hypoglycemia risk indicator has a horizontal position within the second display that corresponds to the position of that different segment of time within the plurality of times of the day, andwherein each variability risk indicator indicates an amount of low range glucose variability in a different segment of time within the plurality of times of the day and each variability risk indicator has a horizontal position within the second display that corresponds to the position of that different segment of time within the plurality of times of the day. 2. The system of claim 1, wherein the glucose data processing program further programs the processor to control the display device to visually present, at a different time from the first and second displays, a control grid on which is shown a hypoglycemia risk curve, a glucose variability curve, and risk reduction distances for hypoglycemia and glucose variability. 3. The system of claim 2, wherein the glucose data processing program further programs the processor to control the display device to visually present on the control grid a treatment recommendation point. 4. The system of claim 2, wherein the glucose data processing program further programs the processor to control the display device to visually present on the control grid an uncertainty boundary. 5. The system of claim 2, wherein the glucose data processing program further programs the processor to control the display device to visually present on the control grid a hyperglycemia curve and a risk reduction distance for hyperglycemia. 6. The system of claim 1, wherein the system further comprises a t... | ABBOTT DIABETES CARE INC. | A61B005/145 A61B005/00 G06F019/00
US09338819 | B2 | 2016-05-10 | Variable data usage personal medical system and method | 2013-05-29 | A variable data usage personal medical system including a self-care device attached to a patient, and operable to generate self-care device data and to transmit the self-care device data at a fixed interval; a cellular communication device operable to receive and store the transmitted self-care device data, to register a count at an interval counter for each of the fixed intervals in which the transmitted self-care device data is received, to generate a data cellular packet from overhead plus the stored self-care device data when the interval counter equals a fixed interval index, and to transmit the data packet; and a cloud infrastructure operably connected to the cellular communication device over a cellular network, and operable to receive, process, and store the transmitted data packet. The cloud infrastructure is operable to transmit a value for the fixed interval index to the cellular communication device for storage. | 1. A variable data usage personal medical system for use with a patient comprising: a self-care device attached to the patient, the self-care device being operable to generate self-care device data and to transmit the self-care device data at a fixed interval;a cellular communication device operably connected to the self-care device, the cellular communication device being operable to receive and store the transmitted self-care device data, to store an interval counter and to increment the interval counter for each of the fixed intervals in which the transmitted self-care device data is received, and in response to determining that the interval counter equals a fixed interval index, to generate a data cellular packet from overhead plus the stored self-care device data and to transmit the data packet, wherein the fixed interval index controls how often data packets are transmitted from the cellular communication device; anda cloud infrastructure operably connected to the cellular communication device over a cellular network, the cloud infrastructure being operable to receive, process, and store the transmitted data packet;wherein the cellular communication device is operable to store the fixed interval index and the cloud infrastructure is operable to transmit a value for the fixed interval index to the cellular communication device for storage. 2. The system of claim 1 wherein the cloud infrastructure is further operable to generate and transmit display data from the transmitted data packet, and the system further comprises a display device operably connected to the cloud infrastructure, the display device being operable to present the transmitted display data. 3. The system of claim 2 wherein: the display device is further operable to generate a first display device data request when the display device is enabled and transmit the first display device data request to the cloud infrastructure;the cloud infrastructure is further operable to send a wake up message to the cellular communication device in response to the first display device data request;the display device is further operable to generate repeated display device data requests and transmit the repeated device data requests to the cloud infrastructure;the cloud infrastructure is further operable to send repeated cloud infrastructure data requests to the cellular communication device;the cellular communication device is further operable to send data cellular packets to the cloud infrastructure in response to each of the repeated cloud infrastructure data requests, one of the data cellular packets being sent each time the self-care device sends self-care device data to the cellular communication device;the cloud infrastructure is further operable to generate continuous display data from the data cellular packets and transmit the continuous display data to the display device; andthe display device is further operable to display the continuous display data. 4. The system of claim 2 wherei... | MEDTRONIC MINIMED, INC. | H04W088/00 A61M005/142 H04W004/00 G06F019/00 A61B005/145 A61M005/172
US09330561 | B2 | 2016-05-03 | Remote communication systems and methods for communicating with a building gateway control to control building systems and elements | 2013-07-31 | A wearable monitoring device, worn by a user, has one or more sensors which acquire at least one of a user's activities, behaviors and habit information, an antenna and a unique user ID. The monitoring device includes a wireless user interface with a one or more input selection elements, which are accessible by the user to control at least a portion of one or more controllable devices housed in a building. One or more controllable systems or devices are at the building. At least a first portion of the one or more controllable systems or devices have an interface with a receiver in communication with the monitoring device that enables the monitoring device to communicate with the receiver. | 1. A system to interact with one or more controllable systems and one or more controllable devices at a building, comprising: a wearable monitoring device that includes a microphone, an RF transmitter and sensors to determine air quality, sound level, sound quality, light quality and ambient temperature near an individual, the RF transmitter serving as a communication system;the wearable monitoring device including a wireless user interface with a one or more input selection elements, the one or more input selection elements being accessible by the individual to control at least a portion of the one or more controllable devices housed in the building;an accelerometer configured to detect an individual's movement information, the accelerometer and the wearable monitoring device configured to assist to determine sleep information and sleep behavior information of the individual, the microphone configured to record individual movement sounds detected by the accelerometer, the accelerometer configured to cause the microphone to stop recording the individual movement sounds when the movement sounds are not directed to a sleep related parameter; andthe one or more controllable systems or the one or more controllable devices at the building, at least a first portion of the one or more controllable systems or the one or more controllable devices having an interface with a receiver in communication with the wearable monitoring device that enables the wearable monitoring device to communicate with the receiver, the first portion of the one or more controllable systems or the one or more controllable devices are directly controllable by the wearable monitoring device and operable to make a change to a setting of that one or more controllable systems or one or more controllable devices at the building. 2. The system of claim 1, further comprising: ID circuitry at the wearable monitoring device. 3. The system of claim 1, further comprising: a building gateway control system at the building that includes a broadband communication interface, a remote data network and a wireless controllable device interface to control a setting for at least the first portion of the one or more controllable devices. 4. The system of claim 3, wherein the building gateway control system includes a monitoring device interface to receive remote commands from the wearable monitoring device, via a wireless data protocol. 5. The system of claim 4, wherein the wireless data protocol is used selected from at least one of 802.11, ultra-wideband (UWB), and wireless technology standard for exchanging data. 6. The system of claim 3, wherein the broadband interface communicates with the remote data network. 7. The system of claim 4, wherein the wireless data protocol is used selected from at least one of 802.11, ultra-wideband (UWB), and wireless technology standard for exchanging data. 8. The system of claim 3, wherein the building gateway control system is positioned at an interior o... | Hello Inc. | G08C019/16 G08C017/02 G10L021/00 H02J007/02 H02J017/00 A61B005/00 A61B005/11 H02J007/00 A61B005/0205 H04L029/08
US09327117 | B2 | 2016-05-03 | Bladder sensing using impedance and posture | 2010-04-20 | This disclosure provides techniques for bladder sensing. In accordance with the techniques described in this disclosure, a device may measure the impedance of a bladder, determine the posture of a patient, and determine a status of the bladder based on the impedance and posture. | 1. A method comprising: receiving, by a processor in a medical device system, a measured impedance value of a urinary bladder of a patient;receiving, by the processor, a determined posture of the patient; anddetermining, with the processor, a filling status of the urinary bladder of the patient based on the measured impedance value of the bladder of the patient and the determined posture of the patient, wherein the measured impedance value is associated with a plurality of filling statuses, the plurality of filling statuses including the determined filling status. 2. The method of claim 1, further comprising: measuring the impedance value of the bladder of the patient with an implanted medical device; and wherein receiving a determined posture of the patient comprises: sensing a signal indicative of a posture of the patient with a posture sensor; anddetermining, with the processor, the posture of the patient based on the sensed signal. 3. The method of claim 2, wherein the posture sensor comprises one or more accelerometers, and wherein determining the posture comprises determining the posture based on one or more signals from the one or more accelerometers. 4. The method of claim 2, wherein measuring the impedance value of the bladder of the patient comprises: emitting an electrical signal via a first electrode implanted proximate to a wall of the bladder at a first location,receiving the electrical signal via a second electrode implanted proximate to the wall of the bladder at a second location, andmeasuring the impedance value of the bladder based on the electrical signal. 5. The method of claim 4, wherein measuring the impedance value of the bladder of the patient further comprises: measuring a voltage between a third electrode implanted adjacent to the first electrode and a fourth electrode implanted adjacent to the second electrode; andmeasuring the impedance value of the bladder based on the electrical signal and the measured voltage. 6. The method of claim 1, wherein the filling status comprises at least one of a volumetric indication of an amount of urine in the bladder or a fill level category. 7. The method of claim 1, wherein determining the filling status of the bladder comprises determining the filling status from a plurality of filling statuses, wherein the filling status is associated with the measured impedance value and the determined posture. 8. The method of claim 1, further comprising generating an alert if the determined filling status indicates a fill level of the bladder above a threshold level. 9. The method of claim 1, further comprising controlling an implanted medical device to deliver therapy to the patient based on the determined filling status. 10. The method of claim 9, wherein the therapy comprises electrical stimulation therapy delivered to at least one of a pelvic floor nerve or a pelvic floor muscle based on the determined filling status. 11. The method of claim 7, wherein each of the plurality of fil... | Medtronic, Inc. | A61B005/04 A61B005/05 A61B005/00 A61N001/36 A61B005/053 A61B005/11 A61B005/20
US09345428 | B2 | 2016-05-24 | Hyperspectral imaging of angiogenesis | 2012-11-15 | The invention is directed to methods and systems of hyperspectral and multispectral imaging of medical tissues. In particular, the invention is directed to new devices, tools and processes for the detection and evaluation of diseases and disorders such as, but not limited to diabetes and peripheral vascular disease and cancer, that incorporate hyperspectral or multispectral imaging. | 1. A multispectral or hyperspectral medical imaging system comprising: a plurality of LED lights, each respective LED light in the plurality of LED lights configured to emit radiation at a particular spectral band in a plurality of spectral bands used by a diagnostic protocol module;a first stage optic configured to receive radiation projected, from one or more respective LED lights in the plurality of LED lights, off of an in vivo tissue located at a region of interest on a subject;one or more polarizers in optical communication with the first stage optic;an imaging sensor for recording an image of the radiation projected off of the in vivo tissue;a diagnostic protocol module adapted to detect a particular characteristic of the in-vivo tissue; anda diagnostic processor configured to: switch one or more respective LED light in the plurality of LED lights off and on based on the diagnostic protocol module,instruct the imaging sensor to record a plurality of images of the region of interest based on the diagnostic protocol module, each respective image in the plurality of images corresponding to radiation emitted at a particular spectral band by one or more respective LED light in the plurality of LED lights,obtain a multispectrally or hyperspectrally resolved image based on the plurality of images, andobtain a pseudo-color image of the region of interest based on the multispectrally or hyperspectrally resolved image, the pseudo-color image enhancing the visibility of the particular characteristic of the in vivo tissue present in the region of interest; wherein the pseudo-color image comprises: (i) a color hue plane that represents a first characteristic in the group consisting of (a) the oxyhemoglobin concentration, (b) the deoxyhemoglobin concentration and (c) a mathematical combination of the oxyhemoglobin concentration and the deoxyhemoglobin concentration,(ii) a color saturation plane that represents a second characteristic in the group consisting of (a) the oxyhemoglobin concentration, (b) the deoxyhemoglobin concentration and (c) a mathematical combination of the oxyhemoglobin concentration and the deoxyhemoglobin concentration, with the proviso that the second characteristic is other than the first characteristic, and(iii) a color intensity plane that represents reflectance of radiation having a spectral band contained within a region that is from about 450 nm to about 580 nm. 2. The multispectral or hyperspectral medical imaging system of claim 1, wherein the imaging system weighs less than 25 pounds. 3. The multispectral or hyperspectral medical imaging system of claim 1, wherein the plurality of LED lights is configured in a circular array about the first stage optic. 4. The multispectral or hyperspectral medical imaging system of claim 1, wherein at least one respective LED light in the plurality of LED lights is configured to emit radiation in the visible spectrum. 5. The multispectral or hyperspectral medical imaging system of cla... | Hypermed Imaging, Inc. | A61B005/00 G01N021/31 A61B005/1455 G01N021/359
US09717438 | B2 | 2017-08-01 | Medical device for detecting a ventricular arrhythmia event | 2015-10-29 | A medical device and method for detecting a ventricular arrhythmia event is disclosed. The medical device includes input circuitry configured to receive an electrocardiogram (ECG) signal, processing circuitry coupled to the input circuitry and configured to identify at least one fiducial point of a first heartbeat signature and at least fiducial point of a second heartbeat signature of the ECG signal, and feature extraction circuitry coupled to the processing circuitry. The feature extraction circuitry is configured to determine at least one difference between the at least one fiducial point of the first heartbeat signal and the at least one fiducial point of the second heartbeat signal. Machine learning circuitry is coupled to the feature extraction circuitry and is configured to select a ventricular arrhythmia class based on the at least one difference. | 1. A method for a medical device for detecting a ventricular arrhythmia event comprising: receiving an electrocardiogram (ECG) signal by way of input circuitry;identifying a first P onset, a first Q onset, and a first S offset within a first heartbeat signal, a second P onset, a second Q onset, and a second S offset within a second heartbeat signal, and a third P onset, a third Q onset, and a third S offset within a third heartbeat signal of the ECG signal by way of processing circuitry coupled to the input circuitry, wherein the first heartbeat signal, the second heartbeat signal, and the third heartbeat signal are generated from three consecutive heartbeats;determining QP interval variability from the first Q onset, the second P onset, the second Q onset, and the third P onset and SP interval variability from the first S offset, the second P onset, the second S offset, and the third P onset by way of feature extraction circuitry coupled to the processing circuitry; anddetecting a ventricular arrhythmia based on the QP interval variability and the SP interval variability by way of machine learning circuitry coupled to the feature extraction circuitry. 2. The method for the medical device of claim 1 wherein: identifying further comprises identifying a first R peak and a first T onset within the first heartbeat signal and a second R peak and a second T onset within the second heartbeat signal by way of the processing circuitry;determining further comprises determining RT interval variability from the first R peak, the first T onset, the second R peak, and the second T onset by way of the feature extraction circuitry; anddetecting further comprises detecting ventricular arrhythmia based upon the QP interval variability, the SP interval variability, and the RT interval variability by way of the machine learning circuitry. 3. The method for the medical device of claim 1 wherein: identifying further comprises identifying a first R peak and a first T onset within the first heartbeat signal, a second R peak and a second T onset within the second heartbeat signal, and a third R peak and a third T onset within the third heartbeat signal by way of the processing circuitry;determining further comprises determining RT interval variability from the first R peak, the first T onset, the second R peak, and the second T onset and determining TR interval variability from the first T onset, the second R peak, the second T onset, and the third R peak by way of the feature extraction circuitry; anddetecting further comprises detecting ventricular arrhythmia based upon the QP interval variability, the SP interval variability, the RT interval variability, and the TR interval variability by way of the machine learning circuitry. 4. The method for the medical device of claim 1 wherein: identifying further comprises identifying a first R peak and a first T onset within the first heartbeat signal, a second R peak and a second T onset within the second heartbeat signal, ... | Khalifa University of Science and Technology | A61B005/0464 A61B005/0468 A61B005/0428 A61B005/0245 A61B005/00 A61B005/04 A61B005/024 A61B005/0456 A61B005/0402
US09345398 | B2 | 2016-05-24 | Atherectomy catheter drive assemblies | 2013-03-15 | A drive assembly for driving an imaging catheter has a rotatable fiber and a rotatable drive shaft. The drive assembly includes a fiber optic rotating junction and a motor configured to rotate the rotatable portion of the fiber optic rotating junction. In some embodiments, the drive assembly includes a sensor configured to detect a rotational position of the fiber optic rotating junction and a processor configured to obtain the detected rotational position and stop the motor only when the fiber optic rotating junction is in a predetermined rotational position. In some embodiments, the motor includes a hollow shaft through which at least a portion of the fiber optic rotating junction extends. | 1. A drive assembly for driving an imaging catheter having a rotatable fiber and a rotatable drive shaft, the drive assembly comprising: a fiber optic rotating junction having a stationary portion with a stationary fiber therein and a rotatable portion with a rotatable fiber therein;a first optical connection configured to connect the stationary fiber with a light source;a motor configured to rotate the rotatable portion of the fiber optic rotating junction;a second optical connection configured to connect the rotatable portion of the fiber optic rotating junction with both the drive shaft and the rotatable fiber of the imaging catheter so as to transmit torque from the motor to the drive shaft and the rotatable fiber of the catheter and so as to transmit light from the light source to the rotatable fiber of the catheter;a sensor configured to detect a rotational position of the fiber optic rotating junction; anda processor configured to obtain the detected rotational position and stop the motor only when the fiber optic rotating junction is in a predetermined rotational position. 2. The drive assembly of claim 1, wherein the sensor is a slot sensor configured to detect a flag on the fiber optic rotating junction. 3. The drive assembly of claim 1, further comprising a locking mechanism configured to lock a handle of the imaging catheter to the drive assembly. 4. The drive assembly of claim 3, wherein the locking mechanism includes mechanical features to physically align the handle with respect to the drive assembly. 5. The drive assembly of claim 4, wherein the locking mechanism is configured such that physical alignment of the catheter handle with respect to the drive assembly further aligns an optical connection of the handle with the predetermined rotational position of the fiber optic rotating junction. 6. The drive assembly of claim 1, wherein the stationary and rotatable fibers are configured to transmit an optical coherence tomography signal. 7. A method of driving an imaging catheter having a rotatable fiber and a rotatable drive shaft, the method comprising: connecting a stationary fiber of a stationary portion of a fiber optic rotating junction in a drive assembly with a light source;connecting a rotatable fiber of the fiber optic rotating junction with the drive shaft and the rotatable fiber of the imaging catheter;rotating the rotatable portion of a fiber optic rotating junction with a motor in the drive assembly such that both the drive shaft and the rotatable fiber of the imaging catheter rotate and such that light is transmitted from the light source to the rotatable fiber of the imaging catheter;sensing a position of the fiber optic rotating junction; andstopping the motor based upon the sensed position only when the fiber optic rotating junction is in a predetermined rotational position. 8. The method of claim 7, wherein the position is sensed with a slot sensor. 9. The method of claim 7, further comprising locking a hand... | Avinger, Inc. | A61B005/05 A61B001/313 A61B001/04 A61B001/00 A61B001/07 A61B017/3207 A61B005/00
US10335067 | B2 | 2019-07-02 | Optical unit and optical analysis device | 2014-05-22 | An optical unit (100) includes a light-emitting portion (10) that irradiates an analysis target with light, light-receiving portions (20) that receive light that has been reflected or diffused by the analysis target, a mounting substrate (30) on which these portions are mounted, and a plate-shaped member (40) having optical transparency. The plate-shaped member (40) is arranged so as to cover the light-emitting portion (10) and the light-receiving portions (20) that are mounted on the mounting substrate (30). Positions of the light-emitting portion (10), the light-receiving portions (20), and the plate-shaped member (40) are set such that, in the case where the analysis target is in contact with the plate-shaped member (40), light emitted from the light-emitting portion (10) is incident on two or more of the light-receiving portions (20) after being reflected or diffused by the analysis target. | 1. An optical unit comprising: a light emitter that irradiates an analysis target with light;a plurality of light receivers that receive light that has been reflected or diffused by the analysis target;a mounting substrate on which the light emitter and the plurality of light receivers are mounted; anda plate-shaped member having optical transparency,wherein the plate-shaped member is arranged so as to cover the light emitter and the plurality of light receivers that are mounted on the mounting substrate, andpositions of the light emitter, the light receivers, and the plate-shaped member are respectively set such that, in a case where the analysis target is in contact with the plate-shaped member, light emitted from the light emitter is incident on two or more of the light receivers after being reflected or diffused by the analysis target,wherein the optical unit further comprises a cooler,the light emitter comprises, between a light-emitting diode and the mounting substrate, a heat transmitter for transmitting heat generated in the light-emitting diode to the mounting substrate,the cooler is arranged in contact with a surface of the mounting substrate on which the light emitter and each light receiver are not mounted, andthe mounting substrate comprises:a conduction channel that is connected to the light-emitting diode and a light receiving photodiode of each light receiver;a wiring that is formed inside the mounting substrate and is connected to the conduction channel; anda heat channel that connects the wiring and the cooler. 2. The optical unit according to claim 1, comprising: a plurality of the light emitters, the plurality of light emitters being mounted on the mounting substrate,wherein positions of the light emitters, the light receivers, and the plate-shaped member are respectively set such that, in a case where the analysis target is in contact with the plate-shaped member, light emitted from each light emitter of the plurality of light emitters is incident on two or more of the light receivers after being reflected or diffused by the analysis target. 3. The optical unit according to claim 2, wherein the plurality of light emitters are mounted on the mounting substrate so as to respectively surround an arbitrary point on the mounting substrate. 4. The optical unit according to claim 3, wherein the plurality of light emitters are constituted by a plurality of groups having mutually different peak wavelengths of light that is emitted, and each of the plurality of groups has two light emitters having the same peak wavelength, andthe two light emitters belonging to the same group are mounted at positions opposite each other across the arbitrary point. 5. The optical unit according to claim 2, wherein the plurality of light emitters are mounted on the mounting substrate in a state in which the plurality of light emitters are arranged linearly. 6. The optical unit according to claim 2, wherein each light emitter of the plurality of ligh... | NEC SOLUTION INNOVATORS, LTD. | A61B005/1455 G01N021/49 A61B005/145 G01N021/31
US10058253 | B2 | 2018-08-28 | System, method, and article for heart rate variability monitoring | 2015-11-10 | What is provided is a system, method, and article of manufacture for detecting, monitoring, and tracking whether a user is in a more relaxed or stressed state by measuring the user's heart rate variability. The system has a heart rate monitoring device with at least one sensor that is configured to measure user heart rate. The system also has a mobile computing device, which may be configured to interact with the heart rate monitoring device for monitoring, tracking, and displaying a user's heart rate and mental stress level information. The method measures the user's heart rate information to calculate a measure of heart rate variability and to analyze whether peaks corresponding to activation levels of the user's parasympathetic nervous system (PNS) and sympathetic nervous system (SNS) exceed a predetermined threshold level within a predetermined time period. | 1. A system for monitoring heart rate variability and mental stress level of a user comprising: a heart rate monitoring device with at least one sensor configured to measure a sample of the user's heart rate and to transmit an output signal containing the user's heart rate sample data;a mobile computing device that communicates with the heart rate monitoring device, the mobile computing device comprising at least one receiver configured to convert the output signal from the at least one sensor into a digital format;the mobile computing device further comprising a processor, the processor configured to determine heart beat interval data from the output signal;the processor further configured to calculate a measure of heart rate variability by filtering heart beat interval data from the output signal;the processor further configured to predict an activation level of the user's parasympathetic nervous system (PNS) and an activation level of the user's sympathetic nervous system (SNS) based on the measure of heart rate variability;the processor further configured to identify whether a first peak corresponding to the activation level of the user's PNS exceeds a predetermined threshold level and whether a second peak corresponding to the activation level of the user's SNS exceeds the predetermined threshold level, wherein the frequency of the second peak exceeding the predetermined threshold level is indicative of higher user mental stress levels; andthe mobile computing device further comprising a graphical user interface configured for displaying the user's heart rate and mental stress level information. 2. The system of claim 1, the mobile computing device is configured to include the heart rate monitoring device. 3. The system of claim 1, the at least one sensor is a photoplethysmographic sensor configured to measure the user's heart rate by detecting the amount of light emitted by a light source after interacting with the user's skin. 4. The system of claim 1, the processor is further configured to compute a score based on the frequency of the first peak exceeding the predetermined threshold level and the frequency of the second peak exceeding the predetermined threshold level, the score corresponds to the user's mental stress levels. 5. The system of claim 4, the processor transmits a notification alert to the user based on the frequency of the second peak exceeding the predetermined threshold level during a predetermined time period. 6. The system of claim 5, the processor transmits the notification alert to the user when the frequency of the second peak exceeding the predetermined threshold level is at least three and the predetermined time period is less than or equal to twenty minutes. 7. The system of claim 4, the score increases when the frequency of the second peak does not exceed the predetermined threshold level during a predetermined time period. 8. The system of claim 4, the score increases when the frequency of the first peak e... | Zenmark, LLC | A61B005/024 A61B005/00 A61B005/16
US10492887 | B2 | 2019-12-03 | Medicine ingestion state management method, medicine and medicine ingestion state management device | 2016-01-26 | Disclosed is a medicine ingestion state management method capable of objectively managing medicine ingestion states of patients. The medicine ingestion state management method involves an operation of prescribing medicine (10) encapsulating, together with a medicament, a medicine information transmitting unit (20) having a function of transmitting medicine information capable of specifying a type and a quantity of the medicament to each individual patient, and an operation of grasping the medicine ingestion state of each patient by collecting the medicine information from each medicine information transmitting unit (20) in each patient. | 1. A capsule comprising; a medicament; andan IC chip which transmits information to outside of a body of a patient who ingests the medicament, the IC chip being covered with a first material which is not dissolved within the patient's body, andwherein the IC chip and the medicament are combined with a second material dissolvable within the patient's stomach,wherein the IC chip is connected to an electrode in the capsule and starts a given operation using power generated by exposure of the electrode to gastric juice, andwherein the exposure of the electrode to the gastric juice is caused by dissolution of the second material by the gastric juice. 2. The capsule according to claim 1, wherein the IC chip transmits information indicating a type of the medicament. 3. The capsule according to claim 1, wherein the IC chip includes a first module that counts a length of time elapsed from when the capsule dissolved within the patient's body, and a second module that transmits information calculated from the length of time which is counted by the first module. 4. The capsule according to claim 3, wherein the first module starts counting the length of time elapsed when power is supplied to the IC chip as a result of dissolution of the capsule within the patient's stomach. 5. The capsule according to claim 3, further comprising an electrode connected to the IC chip, wherein the electrode functions as a battery by immersion into body fluid. 6. The capsule according to claim 3, wherein the second module transmits medicine information containing information representing a medicine effective period of the medicament. 7. The capsule according to claim 1, wherein the IC chip includes: a first IC chip; anda second IC chip,wherein the first IC chip is covered with the first material, and the second IC chip is partly covered with a second material which responds to body fluid within the patient's body. 8. The capsule according to claim 7, wherein the first IC chip transmits first identifying information, the second IC chip transmits second identifying information, and the first identifying information and second identifying information include the same information about the medicament. 9. A capsule which can be swallowed and encapsulates a medicament, the capsule being made from a first material dissolvable within a stomach, the capsule comprising: an IC chip covered with a second material which does not dissolve in a human body and encapsulated with the medicament in the capsule,wherein the IC chip is connected to an electrode in the capsule and includesa memory;a module that writes information into the memory, wherein the module starting writing after power is supplied from the electrode to the IC chip as a result of dissolution of the capsule in the stomach; anda transmitting module that transmits information stored in the memory, using the power,wherein the power is generated by exposure of the electrode to gastric juice, andwherein the exposure of the el... | FUJITSU LIMITED | A61B090/98 G06Q010/10 G16H020/13 A61B005/07 A61B005/00 G06F019/00 G16H010/60 A61J003/00
US10182726 | B2 | 2019-01-22 | Device for physiological monitoring and measurement | 2016-01-14 | Continuous monitoring of a plurality of physiological data may be used for health and fitness improvements for a user. To this end, a physiological monitoring and measurement device may include a wearable strap that receives heart rate data for a user including a time series of heart rate data, a maximum heart rate, and a resting heart rate. A processor may transform the heart rate data into a time series of heart rate reserve data that is weighted, e.g., to account for cardiovascular efficiencies at different intensity levels, to provide a weighted time series of heart rate reserve values. An intensity score that provides an indicator of cardiovascular intensity for an exercise routine may be generated from the weighted time series of heart rate reserve values and displayed to the user on the wearable strap. | 1. A device for physiological monitoring and measurement comprising: a continuously wearable physiological measurement device including one or more sensors configured to continuously monitor and determine a heart rate of a user;a display device coupled to the continuously wearable physiological measurement device;a strap attached to the continuously wearable physiological measurement device and couplable to an appendage of the user; anda processor in the continuously wearable physiological measurement device, the processor coupled to the one or more sensors and the processor configured to: programmatically receive heart rate data for the user from the continuously wearable physiological measurement device, the heart rate data including a time series of heart rate data for the user, a maximum heart rate of the user, and a resting heart rate of the user;transform the heart rate data including the time series of heart rate data, the maximum heart rate and the resting heart rate to a time series of heart rate reserve data;weight the heart rate reserve data according to a dynamic weighting scheme to provide a dynamically weighted time series of heart rate reserve values, the dynamic weighting scheme based on historical heart rate reserve data;sum and normalize the dynamically weighted time series of heart rate reserve values for a predetermined unit interval to provide a summed and normalized value;programmatically generate an intensity score providing an indicator of cardiovascular intensity based on the dynamically weighted time series of heart rate reserve values, wherein generating the intensity score includes scaling the summed and normalized value to a predetermined distribution in a scale, and further wherein a path of the intensity score over time is used to dynamically determine and adjust the dynamic weighting scheme applied to the heart rate reserve data; anddisplay the intensity score on a user interface rendered on the display device. 2. The device of claim 1, wherein the dynamic weighting scheme is created for the user by a machine learning algorithm trained to correlate heart rate data to cardiovascular intensities. 3. The device of claim 2, wherein the machine learning algorithm is retrained to adjust perceived difficulties of exercise routines as a fitness of the user improves. 4. The device of claim 1, wherein the dynamic weighting scheme accounts for cardiovascular efficiencies at different intensity levels. 5. The device of claim 1, wherein the processor is further configured to display qualitative information associated with the intensity score. 6. The device of claim 5, wherein the qualitative information includes an indication of whether the user exceeded an anaerobic threshold of the user during an exercise routine. 7. The device of claim 5, wherein the qualitative information includes an indication of whether the user is likely to experience muscle soreness. 8. The device of claim 5, wherein the qualitative information ... | Whoop, Inc. | A61B005/11 A61B005/0205 A61B005/024 A61B005/00 A63B024/00 G06F001/16 A61B005/0255 A61B005/053 G06F019/00
US10993650 | B2 | 2021-05-04 | System for determining nerve direction to a surgical instrument | 2015-09-25 | A system performs surgical procedures and assessments. The system includes the use of neurophysiology-based monitoring to determine nerve proximity and nerve direction to surgical instruments employed in accessing a surgical target site. | 1. A system for finding a direction of a nerve from a surgical instrument comprising: a first surgical accessory having at least a first stimulation electrode on a distal region and a second stimulation electrode spaced apart from the first stimulation electrode, the stimulation electrodes each configured to deliver a first set of at least one electrical stimulation signal;at least one sensor configured to detect an evoked response from the nerve depolarized by the electrical stimulation signals; anda control unit in communication with the first surgical accessory and the at least one sensor, wherein the control unit includes a touch screen display and a base, wherein the touch screen display and the base include at least a processor in communication with a memory, the memory including software instructions, which when accessed by the processor causes the processor to execute the steps of: (a) electrically stimulate the first stimulation electrode with the first set of at least one electrical stimulation signal until a first initial bracket within which a first threshold stimulation current level must lie is determined, (b) electrically stimulate the second stimulation electrode until a second initial bracket within which a second threshold stimulation current level must lie is determined, (c) process the determined first and second initial brackets to find the direction of the nerve from the surgical instrument, said processing including determining the first threshold stimulation current level by narrowing the first initial bracket to a first final bracket and determining the second threshold stimulation current level by narrowing the second initial bracket to a second final bracket, (d) display on the touch screen display an arc indicating a general direction of the nerve from the surgical instrument, the general direction based on the determinations of (i) said first initial bracket and (ii) said second initial bracket, and (e) narrow a length of the arc based on the determinations of (i) said first threshold stimulation current level and (ii) said second threshold stimulation current level, wherein the step of narrowing the length of the arc is indicative of the direction of the nerve from the surgical instrument. 2. The system of claim 1, wherein the first surgical accessory is configured to advance through tissue to create an operative corridor to a surgical target site. 3. The system of claim 2, wherein the first surgical accessory configured to create the operative corridor to the surgical target site is configured to access a spinal target site. 4. The system of claim 3, wherein the first surgical accessory configured to create the operative corridor to the surgical target site is configured to establish the operative corridor via a lateral, trans-psoas approach. 5. The system of claim 1, wherein the surgical accessory comprises at least one of a series of sequential dilation cannulae, and a K-wire. 6. The system of claim 5, wherei... | NuVasive, Inc. | A61B005/389 A61B005/05 A61B005/00 A61B005/296 A61B090/00
US09448164 | B2 | 2016-09-20 | Systems and methods for noninvasive blood glucose and other analyte detection and measurement using collision computing | 2016-01-12 | In a noninvasive system for detection/measurement of glucose and other analytes in a medium such as tissue, spectra from the medium are deconstructed into features. Conditioned features, which contain frequency components specific to glucose or the other analytes, are derived from one or more features by modulating a carrier kernel with the feature. The conditioned features are computationally collided with one or more Zyotons that are co-dependent with the conditioned features. One or more collisions amplify a property of the analyte e.g., energy absorbed by glucose in tissue from radiation directed to the skin. A gradient of several values of the amplified property, each value corresponding to a particular radiation pattern according to a spectroscopic tomographic sequence, is used to select a suitable projector curve, with which a representative amplified value is projected to an accurate estimate of the concentration of glucose or the other analytes, without needing personalized calibration. | 1. A method for determining concentration of glucose in a tissue sample, the method comprising the steps of: (a) obtaining energy absorbed by glucose from energy directed to the tissue sample as a result of a collision between two co-dependent waveforms; and(b) projecting, via a first processor, the energy absorbed by glucose to a glucose concentration. 2. The method of claim 1, further comprising: (c) receiving a feature corresponding to a specified wavelength range of a spectral signal and representing at least one of: (i) absorption of energy, within the specified wavelength range, by at least one of: (A) glucose in the tissue sample, and (B) at least one confounder in the tissue sample, and(ii) loss of energy within the specified wavelength range due to scattering; and(d) generating a conditioned feature by modulating a carrier kernel using the received feature. 3. The method of claim 2, wherein receiving the feature comprises: directing near-infrared radiation to the tissue sample;collecting radiation from the tissue sample;generating an absorption spectrum from the collected radiation; andselecting a region of the absorption spectrum bounded by the wavelength range as the feature. 4. The method of claim 3, wherein collecting the radiation comprises at least one of receiving diffusely reflected radiation from the tissue sample and receiving radiation transmitted through the tissue sample. 5. The method of claim 2, wherein: one of the two co-dependent waveforms comprises an original Zyoton constructed to amplify energy absorbed by the feature; andthe carrier kernel is constructed such that the original Zyoton and the conditioned feature obtained by modulating the carrier kernel using the feature are co-dependent. 6. The method of claim 5, wherein colliding the original Zyoton with the conditioned feature comprises, in a first iteration: colliding the original Zyoton with the conditioned feature to obtain a modified Zyoton; andrenormalizing the modified Zyoton to obtain a renormalized Zyoton. 7. The method of claim 6, further comprising: in each of −1 additional iterations, >1: colliding the original Zyoton with a renormalized Zyoton from an immediately preceding iteration, to obtain a remodified Zyoton; andrenormalizing the remodified Zyoton to obtain a renormalized modified Zyoton; andcomputing energy gain of the modified Zyoton obtained after N iterations relative to energy of the original Zyoton to determine the energy absorbed by glucose. 8. The method of claim 7, wherein the number of iterations is selected such that accuracy of the glucose concentration measurement, as determined by Absolute Relative Deviation (ARD) relative to a glucose concentration measurement obtained using a reference method, is less than 15%. 9. The method of claim 7, wherein the number of iterations is selected such that accuracy of a plurality of glucose concentration measurements, as determined by Mean Absolute Relative Deviation (MARD) relative to a ... | Zyomed Corp. | G01J005/02 G01N021/3577 A61B005/024 G01N021/359 G01N021/47 G01N021/49 G01N033/49 A61B005/00 A61B005/145 A61B005/1455 A61B005/1495
US10602970 | B2 | 2020-03-31 | Glucose test device | 2015-08-31 | A glucose test device includes a casing, a carrying unit, consumables, a first push rod, a locking component, a first driving unit and a control unit. The carrying unit is disposed at the casing. The consumables are disposed in the carrying unit. The first push rod is movably disposed at the casing. The first push rod is suitable for moving to push one consumable out of the carrying unit. The locking component is movably disposed at the casing. The first driving unit is connected to the locking component. The control unit is electrically connected to the first driving unit. The control unit is suitable for controlling the first driving unit to drive the locking component to lock the first push rod as unmovable. The control unit is suitable for controlling the first driving unit to drive the locking component to release the first push rod. | 1. A glucose test device, comprising: a casing;a carrying unit and a plurality of consumables, the carrying unit being disposed at the casing, the consumables being disposed in the carrying unit;a first push rod movably disposed at the casing and aligned with the carrying unit, wherein the first push rod is suitable for moving to push one of the consumables out of the carrying unit;a locking component movably disposed at the casing;an elastic component, wherein the elastic component is connected between the first push rod and the casing, the first push rod is suitable for resisting an elastic force of the elastic component to move from an initial position to a loading position, and for moving from the loading position toward the carrying unit by the elastic force of the elastic component so as to push one of the consumables out of the carrying unit;a pulling component, wherein the pulling component is movably disposed at the casing and has a pulling portion, and when the pulling component moves relative to the casing, the pulling component resists the elastic force of the elastic component to pull an end of the first push rod by the pulling portion, so that the first push rod moves from the initial position to the loading position;a first driving unit connected to the locking component; anda control unit electrically connected to the first driving unit, wherein the control unit is suitable for controlling the first driving unit to drive the locking component to be operated to a locking position so as to lock the first push rod as unmovable, and the control unit is suitable for controlling the first driving unit to drive the locking component to be operated to a releasing position so as to release the first push rod,wherein the pulling component has a positioning trench, the locking component has a positioning portion, and when the locking component is located at the locking position, the positioning portion extends into the positioning trench to position the pulling component, so as to prevent the pulling component from pulling the first push rod,wherein the consumables surround an axis, and the carrying unit is rotatably disposed around the axis to be driven by a second driving unit to align any one of the consumables with the first push rod. 2. The glucose test device of claim 1, wherein the first driving unit comprises a motor and a gear set, the gear set is connected between the motor and the locking component, and the motor is suitable for driving the gear set to be operated so as to drive the locking component to be operated to the locking position or the releasing position. 3. The glucose test device of claim 1, comprising a second push rod, wherein the second push rod is movably disposed at the casing and is aligned with the carrying unit, the second push rod is connected to the first driving unit, and when the consumables comprise a plurality of lancets and a plurality of test strips, the first push rod is suitable for moving to push o... | SKYLA CORPORATION HSINCHU SCIENCE PARK BRANCH | A61B005/151 A61B005/15 A61B005/157 H01H013/14
US09345889 | B2 | 2016-05-24 | Patient specific data driven safety interlocks for medical devices | 2014-07-25 | A medical device is provided that includes an input/output, at least one sensor, a memory, a controller and at least one delivery member. The input/output is configured to provide a communication path to and from the medical device. The at least one sensor is used to monitor at least one patient function. The memory is used to store patient specific data from the at least one sensor and operating parameters of the medical device. The controller is used to control operations of the medical device. The controller is in communication with the at least one sensor, the memory and the input/output. The controller is configured to deny device operational change requests received via the input/output based at least in part on the patient specific data sensed by the at least one sensor. The at least one delivery member is under the control of the controller and is configured to provide a therapeutic function of the medical device. | 1. A method of generating patient specific interlocks for a medical device, the method comprising: sensing at least one patient function of a patient with a medical device;collecting sensed patient function data from the sensing of the at least one patient function;analyzing the collected patient function data; andbased at least in part on the analyzed collected patient function data, generating at least one patient specific interlock that denies specific operational change requests to the medical device. 2. The method of claim 1 further comprising: dynamically changing the at least one patient specific interlock based on then current patient function data. 3. The method of claim 1, wherein generating the at least one patient specific interlock further comprises: using externally provided patient specific data along with the analyzed collected patient function data. 4. The method of claim 1, wherein generating the at least one patient specific interlock, further comprises: setting prohibited operating parameter requests for the medical device. 5. The method of claim 4, wherein the prohibited operating parameter requests further comprise at least one of prohibiting specific pacing rate requests, pacing mode requests, threshold test requests, rates of response requests, tracking modes requests, permanent output setting requests that lose capture, output to super-threshold value requests, requests that effect the longevity of the medical device and symptomatic tracking requests. 6. The method of claim 1, wherein measuring the at least one patient function of the patient with a medical device further comprises: measuring electrical activity of the patient's heart. 7. The method of claim 1, wherein the at least one patient specific interlock that denies specific operational change requests is created to be unique to at least one of who is requesting the operational change of the medical device and the location of the medical device when the request is received. 8. The method of claim 1 wherein analyzing data from the collected patient function data further comprises at least one of determining if the patient function data indicates a ventricular pacing percent that is above a select percentage, determining if the patient function data indicates an atrial pacing percent that is above a select percentage, determining if the patient function data indicates that the patient is in atrial fibrillation and determining if the patient function data indicates the patient is showing signs of heart failure. 9. A method of operating a medical device, the method comprising: measuring at least one patient function with the medical device;when a request is received at an input to the medical device to change at least one operating parameter of a medical device, applying at least one patient specific interlock that is based at least in part on the at least one measured patient function to determine if the requested change to the at least one operating paramete... | Medtronic, Inc. | A61N001/37 A61B005/02 A61N001/08 A61N001/36 A61N001/365 A61N001/372
US09326702 | B2 | 2016-05-03 | Medical device navigation system | 2014-03-13 | A system for navigating a medical device is provided. In one embodiment, a magnetic field generator assembly generates a magnetic field. Position sensors on the medical device, on an imaging system and on the body generate signals indicative of the positions within the magnetic field. The generator assembly and reference sensors are arranged such that a correlation exists between them and the positions of the body and of a radiation emitter and a radiation detector of the imaging system. An electronic control unit (ECU) determines, responsive to signals generated by the sensors, a position of the medical device, a position of one of the radiation emitter and detector and a distance between the emitter and detector. Using this information, the ECU can, for example, register images from the imaging system in a coordinate system and superimpose an image of the device on the image from the imaging system. | 1. A system for navigating a medical device relative to a body, comprising: a magnetic field generator assembly configured to generate a magnetic field, said medical device comprising a device position sensor configured to generate a device position signal responsive to a position of said device position sensor within said magnetic field;first and second reference position sensors, said first reference position sensor configured to be affixed to an imaging system of a type comprising a structure movable relative to said body and comprising an arm, a radiation emitter coupled to said arm and a radiation detector coupled to said arm and movable relative to said arm, said second reference position sensor configured to be affixed to said imaging system or said body, said first and second reference position sensors generating first and second reference position signals, respectively, responsive to positions of said first and second reference position sensors within said magnetic field wherein a position of a first member of said magnetic field generator assembly, said first reference position sensor and said second reference position sensor is correlated to a position of said radiation emitter, a position of a second member of said magnetic field generator assembly, said first reference position sensor and said second reference position sensor is correlated to a position of said radiation detector and a position of a third member of said magnetic field generator assembly, said first reference position sensor and said second reference position sensor is correlated to a position of said body; and,an electronic control unit configured toreceive said device position signal and said first and second reference position signals; anddetermine: a device position of said medical device within a first coordinate system;a reference position of one of said radiation detector and said radiation emitter within said first coordinate system; and,a distance between said radiation detector and said radiation emitter. 2. The system of claim 1 wherein said magnetic field generator assembly is configured to be attached to said imaging system. 3. The system of claim 1 wherein said reference position includes correlation to an orientation of said one of said radiation detector and said radiation emitter within said first coordinate system. 4. The system of claim 1 wherein said electronic control unit is further configured to register an image generated by said imaging system in said first coordinate system. 5. The system of claim 4 wherein said electronic control unit is further configured to superimpose an image of said medical device on said image responsive to said device position signal. 6. The system of claim 1 wherein said electronic control unit is further configured to display said reference position on a display. 7. The system of claim 1 wherein said electronic control unit is further configured to compare said reference position with a position of said magnet... | MediGuide Ltd. | A61B005/05 A61B005/06 A61B006/00 A61B019/00 A61B006/10 A61B018/14
US20160135739 | A1 | 2016-05-19 | DETERMINING TREATMENT COMPLIANCE USING COMBINED PERFORMANCE INDICATORS | 2015-11-12 | Methods and systems for monitoring compliance of a patient with a prescribed treatment regimen are described. Patient activity is detected unobtrusively with an activity sensor at the patient location, and activity data is transmitted to a monitoring location. Patient speech detected during use of a communication system such as a mobile telephone by the patient may also be used as an activity signal. Patient activity and/or speech is processed at the patient location or monitoring location to identify activity parameters or patterns that indicate whether the patient has complied with the prescribed treatment regimen. The activity sensor and other components at the patient location may be incorporated into, or associated with, a cell phone, computing system, game system, or vehicle system, for example. The system may provide a report to an interested party, for example a medical care provider or insurance company, regarding patient compliance with the prescribed treatment regimen. | 1. A system comprising: at least one audio sensor in a communication system for sensing at least one audio signal including patient speech from a patient at a patient location during use of the communication system, the patient having a brain-related disorder and a prescribed treatment regimen for treating at least one aspect of the brain-related disorder;at least one first activity sensor for sensing at least one first activity signal indicative of a first activity of the patient;signal processing circuitry configured to process the at least one first activity signal and at least one second activity signal indicative of a second activity of the patient to generate at least one activity data signal, the activity data signal containing activity data indicative of whether the patient has complied with the treatment regimen; andat least one transmitting device at the patient location for transmitting the at least one activity data signal and at least one audio data signal based on the at least one audio signal to a receiving device at a monitoring location. 2. The system of claim 1, including at least one second activity sensor for sensing the at least one second activity signal indicative of the second activity of the patient. 3.-26. (canceled) 27. The system of claim 1, wherein the at least one first activity sensor is configured to transduce an activity signal associated with performance of a non-speech activity. 28.-33. (canceled) 34. The system of claim 1, wherein the first activity of the patient and the second activity of the patient are both non-speech activities. 35.-36. (canceled) 37. The system of claim 1, including patient identification circuitry configured to determine a presence of the patient from at least one identity signal sensed at the patient location;wherein the signal processing circuitry is configured to detect an activity of the patient based at least in part on the determination of the presence of the patient by the patient identification circuitry. 38.-55. (canceled) 56. The system of claim 1, wherein the signal processing circuitry includes activity analysis circuitry for processing the at least one first activity signal and the at least one second activity signal to determine at least one activity pattern of the patient. 57.-65. (canceled) 66. The system of claim 1, including notification circuitry for generating a notification indicative of whether the patient has complied with the treatment regimen. 67.-70. (canceled) 71. The system of claim 66, including threat detection circuitry for determining based upon at least one of the at least one first activity signal and the at least one second activity signal that the patient poses a threat;wherein the notification circuitry is configured to generate a notification indicative of the threat. 72. A method comprising: sensing with at least one audio sensor in a communication system at least one audio signal including patient speech from a patient at a patient lo... | Elwha LLC, a limited liability company of the State of Delaware | A61B005/00 A61B005/16 A61B005/1171 A61B005/0205 A61B007/00
US20160136357 | A1 | 2016-05-19 | SINGLE NEEDLE INTEGRATED ARTIFICIAL PANCREAS | 2014-01-27 | Disclosed is a single needle integrated artificial pancreas, wherein the single needle integrated artificial pancreas comprises a fluid reservoir unit; a fluid driving unit connected to the drug reservoir unit; an indwelling unit connected to the fluid reservoir unit and comprising a glucose sensor, an insulin indwelling cannula, a puncture needle and an inserter; wherein the glucose sensor and the indwelling cannula are simultaneously implanted in the same subcutaneous tissue in a patient via the inserter under the aid of the puncture needle; and the puncture needle is a steel puncture needle or groove puncture needle; and a glucose monitoring and insulin delivery control unit is respectively connected to the fluid reservoir unit, the fluid driving unit and the indwelling unit. The single needle integrated artificial pancreas reduces the chance of a diabetic patient being infected and is simple in installation and convenient to use. | 1. A single needle integrated artificial pancreas, comprising: a fluid reservoir unit configured to store insulin;a fluid driving unit connected to the fluid reservoir unit and configured to deliver the insulin stored in the fluid reservoir unit to the subcutaneous tissue of a patient;an indwelling unit connected to the fluid reservoir unit, comprising a glucose sensor, an indwelling cannula, a puncture needle and an inserter; anda glucose monitoring and insulin delivery control unit, which are connected to the fluid reservoir unit, the fluid driving unit and the indwelling unit, and configured to process effective glucose signals output by the glucose sensor, and to control the fluid reservoir unit, the fluid driving unit and the indwelling unit;wherein the glucose sensor and the indwelling cannula are simultaneously implanted in the same subcutaneous tissue in the patient via the inserter with the aid of the puncture needle;wherein a distance between an end of the insulin-indwelling hose and an end of the glucose sensor ranges from 1 mm to 8 mm;wherein the puncture needle is a steel puncture needle or a groove puncture needle. 2. The single needle integrated artificial pancreas according to claim 1, wherein the glucose sensor comprises an electrode sensor, a fibre optic sensor, or a combination thereof. 3. The single needle integrated artificial pancreas according to claim 2, wherein the electrode sensor is integrated in an outer surface of the indwelling cannula, and the steel puncture needle is located inside the indwelling cannula. 4. The single needle integrated artificial pancreas according to claim 2, wherein the electrode sensor is integrated in the outer surface of the indwelling cannula, and the electrode sensor together with the indwelling cannula are located in the groove of the groove puncture needle. 5. The single needle integrated artificial pancreas according to claim 3, wherein the electrode of the electrode sensor is ring-shaped or fan-shaped, and an electrode contact point of the electrode sensor is ring-shaped or fan-shaped. 6. The single needle integrated artificial pancreas according to claim 2, wherein the electrode sensor has an electrode coating structure comprising an anti-interference layer, an enzyme layer, a regulation layer and a protective layer. 7. The single needle integrated artificial pancreas according to claim 2, wherein the electrode sensor has an electrode coating structure comprising an enzyme layer, a regulation layer and a protective layer. 8. The single needle integrated artificial pancreas according to claim 2, wherein the electrode sensor has an electrode coating structure comprising an anti-interference layer, a bonding layer, an enzyme layer, a bonding layer, a regulation layer and a protective layer. 9. The single needle integrated artificial pancreas according to claim 1, wherein the indwelling cannula and the glucose sensor are arranged in parallel, and located in a groove of the groove p... |  | A61M005/172 A61M005/142 A61B005/145
US20160136057 | A1 | 2016-05-19 | SYSTEM AND APPARATUS FOR CONTROLLING FOOD INTAKE | 2014-11-19 | A controlled food intake system includes an apparatus having an inlet tube configured to connect to a lower esophageal stoma, a flow controller, and at least one container configured to detachably attach to the inlet tube and to administer nutrition liquids to a patient at a flow rate based on a position of the flow controller. The apparatus also includes an outlet tube configured to connect to an upper esophageal stoma and a waste container that detachably attaches to the outlet tube. Processing circuitry of the apparatus is configured to control the flow controller to control the flow rate of the nutrition liquids being administered to the patient based on prescribed nutrition intake order. The controlled food intake system also includes at least one server configured to generate reports based on the nutrition liquids administered to the patient and substance consumed by the patient. | 1. A system for controlled food intake, comprising: a controlled food intake apparatus, having an inlet tube configured to connect to a lower esophageal stoma,a flow controller,at least one container configured to detachably attach to the inlet tube and to administer nutrition liquids to a patient at a flow rate based on a position of the flow controller,an outlet tube configured to connect to an upper esophageal stoma,at least one waste container configured to detachably attach to the outlet tube, andprocessing circuitry configured to control the flow controller to control the flow rate of the nutrition liquids being administered to the patient based on prescribed nutrition intake order; andat least one server configured to generate reports based on at least one of the nutrition liquids administered to the patient and substance consumed by the patient. 2. The system of claim 1, wherein the inlet tube includes an inlet flowmeter downstream of the flow controller to detect a measured flow rate of fluid passing through the inlet tube. 3. The system of claim 2, wherein the processing circuitry is further configured to output an alarm when a difference between the ordered flow rate and the measured flow rate is greater than a predetermined threshold. 4. The system of claim 1, wherein the processing circuitry is further configured to determine a quantity of the nutrition liquids that have been administered to the patient. 5. The system of claim 1, wherein the at least one server is further configured to receive inputs from a medical provider regarding the prescribed nutrition intake orders. 6. The system of claim 1, wherein one or more inflators are detachably attached to proximal ends of the inlet tube and the outlet tube to establish seals with an esophagus of the patient when the inlet tube is inserted into the lower esophageal stoma and the outlet tube is inserted into the upper esophageal stoma. 7. The system of claim 6, wherein the processing circuitry is further configured to detect when a pressure in the inflators is less than a predetermined pressure. 8. The system of claim 1, wherein the processing circuitry is further configured to determine a minimum starting level for the at least one container to provide a prescribed quantity of the nutrition liquids to the patient. 9. The system of claim 8, wherein the processing circuitry is further configured to output an alarm when an initial level of the at least one container is less than the minimum starting level. 10. The system of claim 1, wherein the processing circuitry is further configured to output an alarm when the at least one container is at a predetermined empty value. 11. The system of claim 1, wherein the processing circuitry is further configured to determine when a prescribed quantity of nutrition liquids has been administered to the patient. 12. The system of claim 1, wherein the processing circuitry is further configured to override the prescribed nutrition intake order... | UMM AL-QURA UNIVERSITY | A61J015/00 A61B005/00
US20160135758 | A1 | 2016-05-19 | Patient Care Device with Audible Alarm Sensing and Backup | 2014-06-14 | An alarm sensing and backup system for patient devices that senses when the normal alarm sensing system is not sounding properly and sounds an independent backup alarm. | 1. A patient care device; comprising: a. an audible alarm sensor system to detect critical conditions of the patient or of the patient care device;b. one or more audible alarms that are triggered to sound by the audible alarm sensor system;c. an independent backup alarm for the patient care device;d. an alarm verification system comprising a sound sensing device to sense and verify whether the one or more audible alarms have actually sounded;e. wherein if the sound sensing device does not detect a correct audible alarm the alarm verification system activates the independent backup alarm. 2. The patient care device of claim 1 wherein the patient care device is an infant warming center. 3. The patient care device of claim 2 further comprising a vertical column structure mounted on the infant warming center that supports a radiant heater head, containing a radiant heater assembly that is one of the associated warming mechanisms. 4. The patient care device of claim 3 wherein the sound sensing device to sense the audible alarm and the independent backup alarm are mounted within the patient care device. 5. The patient care device of claim 4 wherein the sound sensing device to sense the audible alarm and the independent backup alarm are mounted in the vertical column structure. 6. The patient care device of claim 1 wherein the patient care device is an anesthesia machine. 7. The patient care device of claim 1 wherein the patient care device is an infusion pump. 8. The patient care device of claim 1 wherein the patient care device is a ventilator. 9. The patient care device of claim 1 wherein the patient care device is a monitoring device. 10. The patient care device of claim 1 wherein the patient care device is in a hospital or medical setting. 11. The patient care device of claim 1 wherein the one or more audible alarms can vary in pattern for differing critical conditions. 12. The patient care device of claim 11 wherein the sound-sensing device can sense and verify whether a correct alarm pattern is sounded for each differing critical condition. 13. The patient care device of claim 12 wherein the sound-sensing device can sense and verify whether the correct alarm pattern is sounded with sufficient volume. 14. The patient care device of claim 12 wherein the sound-sensing device can sense and distinguish at least between correct alarm patterns based on Pulse, Pulse Frequency, Harmonic Components, Pulse Duration, Rise/Fall Time, or Tempo. 15. The patient care device of claim 12 wherein the sound-sensing device can sense a voice alarm. 16. The patient care device of claim 1 wherein the sound sensing device is a microphone. 17. The patient care device of claim 1 wherein the one or more audible alarms are sounded by speakers. 18. A method for providing an independent backup alarm for a patient care device having an audible alarm sensor system for detecting critical conditions and one or more audible alarms comprising: a. sensing the one or ... | Segars California Partners, LP | A61B005/00
US20160143579 | A1 | 2016-05-26 | Wearable sports monitoring equipment and method for characterizing sports performances or sportspersons | 2015-11-16 | The invention concerns wearable electronic devices, systems and methods for sports performance monitoring. In one embodiment, the invention provides a device comprising a heartbeat sensor for providing a heartbeat signal, a motion sensor for providing a motion signal and processing means adapted to calculate at least one performance parameter depicting said sports performance and/or the person using temporal characteristics of periodic features of the heartbeat signal compared with temporal characteristics of periodic features in the motion signal. The invention allows for utilization of an existing relation between cadence and heart rate for characterizing the performance or the person in a novel way. | 1. Wearable electronic equipment for monitoring a sports performance of a person, comprising: a heartbeat sensor for providing a heartbeat signal in response to heartbeat of the person,a motion sensor for providing a motion signal in response to movement of the person, anda processing means adapted to calculate at least one performance parameter depicting said sports performance and/or the person based on temporal characteristics of periodic features of the heartbeat signal compared with temporal characteristics of periodic features of the motion signal. 2. The wearable electronic equipment according to claim 1, wherein the processing means is further adapted to: calculate at least one correlation factor dependent on the temporal characteristics of periodic features of the heartbeat signal compared with the temporal characteristics of periodic features in the motion signal, andusing said correlation factor, calculate said at least one performance parameter. 3. The wearable electronic equipment according to claim 1, wherein the processing means is further adapted to calculate at least a parameter depicting intensity of the performance, strain of the performance, fatigue of the person, fitness level of the person or a combination thereof as said at least one performance parameter. 4. The wearable electronic equipment according to claim 1, wherein the processing means is further adapted to calculate at least the estimated anaerobic heart rate threshold level of the person, the present deviation from anaerobic heart rate threshold level, the ratio of cadence and heart rate as an index, fitness index, fatigue index, energy consumption or a combination thereof as said at least one performance parameter. 5. The wearable electronic equipment according to claim 1, wherein the processing means is further adapted to calculate at least a maximum speed and/or step length of the person at-in an aerobic ranges as said at least one performance parameter. 6. The wearable electronic equipment according to claim 1, wherein the processing means is further adapted to determine the heart rate of said person using the heartbeat signal and a frequency of periodic features of the motion signal, and wherein the performance parameter is dependent on the difference between the heart rate and frequency of the periodic features of the motion signal. 7. The wearable electronic equipment according to claim 1, wherein the processing means is further adapted to: detect individual heartbeats from the heartbeat signal and to associate first time stamps with the individual heartbeats, anddetect periodic movements of the person relating to the sports performance from the motion signal and to associate second time stamps with the periodic movements, andto utilize said time stamps for calculating the performance parameter. 8. The wearable electronic equipment according to claim 7, wherein the processing means is further adapted to detect and/or estimate, using said first and sec... |  | A61B005/00 A61B005/0402 A61B005/0205
US09622698 | B2 | 2017-04-18 | System and method for detecting cancerous tissue from a thermal image | 2014-11-19 | What is disclosed is a system and method for the detection of cancerous tissue by analyzing blocks of pixels in a thermal image of a region of exposed skin tissue. In one embodiment, matrices are received which have been derived from vectors of temperature values associated with pixels in blocks of pixels which have been isolated from a plurality of thermal images of both cancerous and non-cancerous tissue. The vectors are rearranged to form matrices. A thermal image of a subject is received. Blocks of pixels which reside within a region of exposed skin tissue are identified and isolated. For each identified pixel block, an image vector comprising temperature values associated with these pixels is formed. The vector is provided to a classifier which uses the matrices to classify tissue associated with this block of pixels as being either cancerous or non-cancerous tissue. | 1. A method for detecting cancerous tissue from a thermal image, the method comprising: receiving, with a processor, matrices AH of size m×H and AN of size m×N comprising vectors of size m×1 derived from temperature values associated with pixels in blocks of size p×q, where m=pq, which have been isolated from a plurality of thermal images of cancerous and non-cancerous tissues, where H and N represent a number of cancerous and non-cancerous blocks of pixels, respectively, said vectors being rearranged to form said matrices;receiving a thermal image of a subject;identifying a block of pixels in a region of exposed skin tissue in said received thermal image;forming a vector y of size m×1 derived from temperature values associated with pixels in said identified block of pixels; andproviding said vector to a classifier, said classifier using said matrices to classify tissue in said thermal image represented by said block of pixels as being one of: cancerous and a non-cancerous tissue. 2. The method of claim 1, wherein vector y is represented by: y=[AHAN]x where x=[xHTxNT]T is a coefficient vector of size (N+H)×1, xH is a vector of size H×1 and xN is a vector of size N×1 obtained from blocks of pixels from said plurality of thermal images of cancerous and non-cancerous tissues, respectively, and T is a transpose operation. 3. The method of claim 2, further comprising: performing minimization using vector y and matrices AH and AN to obtain a sparse coefficient vector x*, where x*=[(xH*)T(xN*)T]T; anddetermining residues represented by: RN=∥y−ANxN*∥2; andRH=∥y−AHxH*∥2 where RH and RN are residues of cancerous and non-cancerous tissues, respectively. 4. The method of claim 3, wherein said minimization is a constrained minimization with convex optimization. 5. The method of claim 3, wherein, in response to RH&lt;RN, classifying tissue in said block of pixels as cancerous. 6. The method of claim 3, wherein, in response tolog⁡(RNRH)>0,classifying tissue in said block of pixels as cancerous. 7. The method of claim 1, wherein said classifier is any of: Support Vector Machine (SVM), a neural network, a Bayesian network, a Logistic regression, Naïve Bayes, Randomized Forests, Decision Trees, Boosted Decision Trees, K-nearest neighbor, Restricted Boltzmann Machine (RBM), and a hybrid system comprising any combination hereof. 8. The method of claim 1, wherein, in advance of identifying a block of pixels in said thermal image for processing, further comprising enhancing a spatial resolution of said thermal image. 9. The method of claim 1, wherein said matrices AH and AN are one of; basis matrices, product of sampling matrices and basis matrices, dictionary matrices, and random matrices. 10. The method claim 1, wherein said vector y comprises temperature values containing a number of rows which is less than a number of rows in matrices AH and AN. 11. The method of claim 1, wherein, in response to said subject being classified as having cancerous tissue, p... | XEROX CORPORATION | A61B005/00 A61B005/01 A61B010/00
US09345909 | B2 | 2016-05-24 | Skin ulcer treatment | 2011-12-22 | A system and method for facilitating the healing of a skin ulcer, such as a diabetic ulcer, on a treatment region of a body. Ultrasound waves are transmitted to the treatment region. Interferential electrical stimulation is applied to the treatment region, simultaneously with the ultrasound transmission. The treatment region may be massaged with a massaging device during the treatment session. A gel may be applied onto the treatment region to facilitate the ultrasound transmission and electrical stimulation. A measurement system may be used to obtain accurate measurements of the skin ulcer, to provide a quantitative determination of the healing progression. | 1. A method for facilitating healing of a skin ulcer on a treatment region of a body, the method comprising: measuring size of the skin ulcer;applying a gel onto skin at the treatment region;transmitting ultrasound waves to the treatment region;positioning electrodes of an electrical stimulation apparatus onto the skin at the treatment region in proximity to the skin ulcer;applying interferential electrical stimulation through the electrodes to the treatment region simultaneously with transmission of the ultrasound waves; andmassaging the treatment region simultaneously with transmitting ultrasound waves and applying interferential electrical stimulation. 2. The method according to claim 1, wherein measuring size of the skin ulcer comprises: positioning a reference measurement ruler adjacent the skin ulcer;imaging the skin ulcer and the reference measurement ruler; andcalculating at least one measurement of a surface area of a skin ulcer bed and a surface area of a skin ulcer region based on a comparative reference scale established using the reference measurement ruler. 3. The method according to claim 1, further including quantitating healing progression of the skin ulcer, the quantitating comprising: measuring the skin ulcer before beginning a treatment session by calculating at least one measurement of a surface area of a skin ulcer bed and a surface area of a skin ulcer region; andcomparing measurements of the surface areas calculated following successive treatment sessions. 4. The method according to claim 1, wherein the ultrasound waves are transmitted to the treatment region at an operating frequency between 1 and 3 MHz and at an operating intensity between 1.5-2.1 W/cm2. 5. The method according to claim 1, wherein an operating frequency of the interferential electrical stimulation is varied during a treatment session, the varying in accordance with a predetermined pattern selected from the group consisting of: applying a first frequency for a fixed amount of time before switching to a second frequency;gradually changing frequencies from a first frequency to a second frequency over various durations of time; andintermittently applying extreme frequencies within a frequency range. 6. The method according to claim 1, wherein the method does not comprise applying a vacuum or suction. |  | A61B008/00 A61N007/00 A61N001/04 A61N001/32 A61B005/00
US20160132122 | A1 | 2016-05-12 | Gesture Feedback for Non-Sterile Medical Displays | 2013-05-29 | The present disclosure relates to a feedback system for control or input gestures in a medical environment, which includes a presence detector which generates a threshold and acquires detection data which describe whether a medical input gesture reaches and/or crosses the threshold; and the extent to which the gesture has passed the threshold; and a computer which is connected to the presence detector and triggers a sterility-preserving response to a positive detection by the presence detector including information how dangerous the situation is becoming with regard to non-sterility. The invention also relates to a medical image display system comprising such a feedback system and to a method of providing sterility-preserving feedback with respect to control or input gestures in a medical system. | 1. A feedback system for control or input gestures in a medical environment, comprising: a presence detector which generates a threshold and acquires detection data which describe whether a medical input gesture crosses the threshold and the extent to which the gesture has passed the threshold; anda computer which is connected to the presence detector and triggers a sterility-preserving response to a positive detection by the presence detector including information related to a danger level. 2. The feedback system according to claim 1, wherein the presence detector comprises a sensor unit which monitors the presence of an object in a predetermined portion of space which includes the threshold. 3. The feedback system according to claim 1, wherein the presence detector monitors a substantially planar sensing area which includes a threshold plane. 4. The feedback system according to claim 1, wherein the computer is connected to a response indicator which issues a sterility-preserving signal. 5. The feedback system according to claim 1, further comprising a medical display in front of which the threshold is generated. 6. The feedback system of claim 5, wherein the threshold is located at a predetermined or selectable distance from the display which represents a minimum distance to be observed when a gesture or gesturing object approaches the display. 7. The feedback system of claim 5, wherein a positioning aid assists in establishing and/or maintaining a predetermined or selectable distance between the medical display and the threshold and/or between the medical display and the presence detector. 8. The feedback system of claim 5, further comprising a display moving unit for moving the display away from a gesture or gesturing object which has reached or crossed the threshold or breached a predetermined or selected distance from the display. 9. A method of providing sterility-preserving feedback with respect to control or input gestures in a medical system, the method comprising the steps of: generating a threshold by using a presence detector and acquiring detection data which describe whether a medical input gesture crosses the threshold and the extent to which the gesture has passed the threshold; andtriggering a sterility-preserving response to a positive detection including information related to how dangerous the situation is becoming with regard to non-sterility, by the presence detector by using a computer which is connected to the presence detector. 10. (canceled) 11. The method according to claim 9, wherein the presence detector acquires data which describe the extent to which the gesture or gesturing object has entered or passed the threshold, using one of a continuous, a non-continuous, or in an incremental, distance measurement. 12. The method according to claim 9, wherein the computer's response involves providing user guidance, which comprises one or more of the following: a warning message, in particular if actual contact wi... | Brainlab AG | G06F003/01 A61B005/00 G06F003/0481 G06F003/041 G06F003/16
US09332926 | B2 | 2016-05-10 | MRI imaging probe | 2011-11-25 | A magnetic resonance imaging (MRI) probe includes a coil section having an imaging coil, and a handle section connected to the coil section. The handle section has a phase shifter circuit with inductors, capacitors, and coax line electrically connected and is configured to provide appropriate phase shift. The handle section further has a coaxial cable winding electrically connected to the imaging coil, and wound cylindrically, and has a slot therethrough between the ends of the cylindrical winding. The handle section further has a pre-amp circuit mounted on a substrate and electrically connected to the cylindrical coax winding. | 1. An endocavity MRI imaging probe for use in a magnetic resonance imager (MRI) which generates a main magnetic field, the probe comprising: a coil section having an imaging coil, the coil section being configured for insertion into an orifice of a patient;a handle section connected to the coil section, the handle section being configured to be grasped manually by a clinician during positioning of the handle section having: a phase shifter circuit comprising a set of inductors and capacitors electrically connected, a coaxial cable winding including a coaxial cable electrically connected to the imaging coil and wound around a cylinder; anda pre-amp circuit mounted on a substrate and electrically connected to the coaxial cable winding, wherein at least a portion of the substrate passes through and is supported by the coaxial cable winding. 2. The probe of claim 1, wherein the imaging coil includes a passive tuning circuit. 3. The probe of claim 2, further comprising a fiducial marker disposed in the coil section. 4. The probe of claim 3, wherein the coil section and the handle section are cylindrical and further including: a cylindrical neck section connecting the coil section to the handle section, the neck section having a smaller diameter than that of the coil and handle sections, the neck section, the coil section, the handle section and the neck section being aligned along a longitudinal axis. 5. The probe of claim 4, wherein the substrate is planar and is parallel to the longitudinal axis. 6. The probe of claim 1, comprising: a second imaging coil in the coil section, the second imaging coil having a second passive tuning circuit and electrically connected to the first imaging coil through a decoupling circuit;a second set of inductors and capacitors electrically connected, a second coaxial cable winding electrically connected to the second imaging coil and wound cylindrically around the slot of the first cylindrical winding; anda second pre-amp circuit mounted on the substrate electrically connected to the second cylindrical winding. 7. The probe of claim 6, wherein the first and second coaxial cylindrical windings, the coaxial cables of the first and second coaxial cylindrical windings being wound on the cylinder radially relative to each other. 8. The probe of claim 6, wherein the first and second coaxial cylindrical windings, the coaxial cables of the first and second coaxial cylindrical windings being wound side-by-side on the cylinder. 9. The probe of claim 8, wherein the second coaxial cable windings are wound in opposite directions around the cylinder. 10. The probe of claim 4, further comprising a coil shell surrounding the coil section and the neck section. 11. The probe of claim 10, wherein the coil section further has a passage therethrough, and the coil shell defines a slot therethrough aligned with the passage, the slot and passage permitting passage of a medical instrument to pass through the coil section of the probe... | Invivo Corporation | A61B005/055 A61B005/00 G01R033/34 G01R033/28 A61B017/00 A61B017/34 A61B019/00
US09332911 | B2 | 2016-05-10 | System and method for prediction and detection of circulatory shock | 2012-08-13 | Systems and methods for prediction and detection of circulatory shock using estimates or measurements of arterial blood pressure, heart rate, stroke volume, cardiac output, total peripheral resistance, cardiac ejection fraction, cardiac contractility and ventricular end-diastolic volume are provided. These estimates and measurements are used to determine a type of circulatory shock. In some embodiments, the type of circulatory shock is determined to be one of septic shock, hypovolemic shock, anaphylactic shock, hemorrhagic shock, and cardiogenic shock. | 1. A system for displaying information about circulatory shock, comprising: a display;a processor; anda memory storing computer-executable instructions which, when executed by the processor, cause the processor to: receive arterial blood pressure information from a blood pressure measuring device,compute a time series of estimates, based on the received arterial blood pressure information, of each of total peripheral resistance, ejection fraction, and ventricular end-diastolic volume,select a type of circulatory shock, from a plurality of types of circulatory shock, based on the computed time series of estimates,cause the display of at least one of the time series of estimates, andcause the display of the selected type of circulatory shock. 2. The system of claim 1, wherein the at least one of the time series of estimates and the selected type of circulatory shock are displayed simultaneously. 3. The system of claim 1, wherein the time series of estimates are updated as additional arterial blood pressure information is received. 4. The system of claim 3, wherein the time series of estimates are updated upon an onset time of a cardiac beat. 5. The system of claim 1, further comprising a user interface, and wherein the memory stores computer-executable instructions which, when executed by the processor, cause the processor to: receive, from the user interface, a user selection of at least one of the time series of estimates, and wherein causing the display of the at least one of the time series of estimates comprises causing the display of the selected at least one of the time series of estimates. 6. The system of claim 1, further comprising a user interface, and wherein the memory stores computer-executable instructions which, when executed by the processor, cause the processor to: receive, from the user interface, a user selection of a scale of a display axis, and wherein causing the display of the at least one of the time series of estimates comprises causing the display of the at least one of the time series of estimates according to the selected scale. 7. The system of claim 6, wherein the scale is a temporal scale. 8. The system of claim 1, wherein the memory stores computer-executable instructions which, when executed by the processor, cause the processor to: compute an additional time series of estimates, based on the received arterial blood pressure information, of at least one of cardiac output and cardiac contractility, andcause the display of the additional time series. 9. The system of claim 1, wherein the memory stores computer-executable instructions which, when executed by the processor, cause the processor to: cause the display of one or more time series of drug infusion information. 10. The system of claim 9, wherein the one or more time series of drug infusion information and the at least one of the time series of estimates are displayed simultaneously. 11. A computer-implemented method for displaying information about ... | Massachusetts Institute of Technology | A61B005/00 A61B005/021 A61B005/0402 G06F019/00 A61B005/029
US09332890 | B2 | 2016-05-10 | Processor for electronic endoscope and electronic endoscope apparatus | 2011-08-25 | A processor for an electronic endoscope includes first and second image storing units and a controller, which in a first mode, image data is converted sequentially into a video signal and is outputted while being stored sequentially in the first and second image storing units. In a second mode, image data stored in the first image storing unit is outputted while generated image data is sequentially stored in the second image storing unit. In a third mode, image data stored in the second image storing unit is outputted while generated image data is sequentially stored in the first image storing unit. In the second and third modes, it is possible to switch between a first playback mode that continuously outputs the multiple stored image data and a second playback mode that repeatedly outputs one of the multiple stored image data. | 1. A processor for an electronic endoscope for processing an image signal from the electronic endoscope and displaying an image on a monitor, comprising: an image data generator that generates a plurality of frames of image data from the image signal;first and second image storages, each of which stores the plurality of frames of image data;a signal processor that converts the plurality of frames of image data into a video signal which is displayed on the monitor; anda controller that controls the signal processor and the first and second image storages,wherein the controller performs control according to one of:a first mode, in which the plurality of frames of image data generated by the image data generator are sequentially converted into the video signal and are outputted, and are sequentially stored in the first and second image data storages;a second mode, in which the plurality of frames of image data stored in the first image storage are converted into the video signal and are outputted, and the plurality of frames of image data generated by the image data generator are sequentially stored in the second image storage; anda third mode, in which the plurality of frames of image data stored in the second image storage are converted into the video signal and are outputted, and the plurality of frames of image data generated by the image data generator are sequentially stored in the first image storage,wherein:in the second mode, a first reproduction mode and a second reproduction mode are switched from one to another, wherein, in the first reproduction mode, the plurality of frames of image data stored in the first image storage are successively converted into the video signal and are outputted, and, in the second reproduction mode, one of the plurality of frames of image data stored in the first image storage is repeatedly converted into the video signal and is outputted; andin the third mode, the first reproduction mode and the second reproduction mode are switched from one to another, wherein, in the first reproduction mode, plurality of frames of image data stored in the second mage storage are successively converted into the video signal and are outputted, and, in the second reproduction mode, one of the plurality of frames of image data stored in the second image storage is repeatedly converted into the video signal and is outputted;the processor for an electronic endoscope further comprising:a judger that receives an input of a control signal for switching from the first mode to the second or the third mode and for switching from the third mode or second mode to the first mode, and makes a judgment on which of the second mode and the third mode the controller switches to, when the control signal is inputted in the first mode,wherein the controller switches from the second mode to the first mode, when the control signal is inputted in the second mode,wherein, in response to the controller switching from the second mode to the first mode,... | HOYA CORPORATION | A61B001/00 A61B001/04 H04N005/225 H04N007/18 H04N001/46 G02B023/24 A61B001/045 A61B001/05 A61B005/00 H04N005/232
US09345884 | B2 | 2016-05-24 | Methods and systems for controlling a state of a neurostimulator | 2012-09-28 | A method and system is described for ensuring a state of an active implantable medical device based on the presence and persistence of a magnetic field. The output of a magnetic field sensor is monitored. The active implantable medical device is maintained in a first state, for so long as the presence of a magnetic field is detected by the magnetic field sensor, until a first interval is surpassed. If the first interval is surpassed, then a determination is made as to whether a second interval has been surpassed. If it is determined that the second interval has not been surpassed, then the active implantable medical device is transitioned into a second state. If it is determined that the second interval has been surpassed, then it is ensured that the active implantable medical device is in a predetermined one of the first and second states. | 1. A method for selecting a state of operation of an active implantable medical device, the method comprising: outputting a signal from a magnet sensor of the implantable medical device when a magnetic field is detected relative to the implantable medical device by the magnet sensor, the implantable medical device operating in a current state at the time of a first detection of the magnetic field;initiating a timer of the implantable medical device that tracks an elapsed-magnet-detection time period over which the signal is continuously output;comparing by a processor of the implantable medical device, the elapsed-magnet-detection time period to a first interval characterized by a first interval end point and a second interval characterized by a second interval end point that is different from the first interval end point;if the elapsed-magnet-detection time period exceeds the second interval end point, outputting a signal from the processor that forces the state of operation of the active implantable medical device to a first state regardless of the current state, wherein the first state corresponds to a therapy disabled state;if the elapsed-magnet-detection time period terminates between the first interval end point and the second interval end point, outputting a signal from the processor that toggles the state of the active implantable medical device from the current state to either the first state or to a different state;if the elapsed-magnet-detection time period terminates before the first interval end point, maintaining the state of the active implantable medical device in the current state;monitoring a number of times the state of operation of the implantable medical device has changed from a therapy enabled state to the therapy disabled state within a period of time due to terminations of the elapsed-magnetic-detection time period; andpreventing the state of the active implantable medical device from being changed from the therapy disabled state to the therapy enabled state when the elapsed-magnetic-detection time period terminates between the first interval end point and the second interval end point, and the number of times exceeds a threshold value. 2. The method of claim 1, wherein each of the states of the active implantable medical device corresponds to a level of activity of the active implantable medical device, and the level of activity of the first state is different from the level of activity of any other state. 3. The method of claim 1, wherein outputting a signal from the processor that toggles the state of the active implantable medical device further comprises: when the current state is the first state, outputting a signal from the processor that toggles the current state to a second state; andwhen the current state is the second state, outputting a signal from the processor that toggles the current state to the first state. 4. The method of claim 1, further comprising: generating a first alert at the time of a first de... | NeuroPace, Inc. | A61N001/36 A61B005/00 A61B005/0476 A61N001/37
US09345412 | B2 | 2016-05-24 | System and method for multiclass discrimination of neural response data | 2012-10-02 | Systems and methods are described herein for analyzing neural response data that can be assigned to multiple classes. The systems and methods begin with a set of training data from which optimal weight factors are derived. The derived weight factors are used in a classifier which is then applied to test data from test subjects. The classifier filters out the effects of less relevant data in the test data and provides a result in the form of probabilities associated with classes for the test data. | 1. A computer system for analyzing neural response data of a subject to determine the subject's choice when the choice is unknown, the system comprising: a processor, a computer-readable memory, and a computer-readable storage media;the computer-readable storage media comprising: a training module that uses a loss function to determine relevant features in a set of neural response training data, wherein the training data is collected when the subject's choice is known; identifies a set of optimal weights that produce a minimum for the loss function according to an iterative procedure of setting at least one of the optimal weights to zero and identifying features based on partial derivatives of the loss function; and creates a classifier using the set of optimal weights, wherein the training module is stored on the computer-readable storage media for execution by the processor via the computer-readable memory;a classifier module for applying the classifier to a subject's test data to determine the subject's choice when the choice is unknown, wherein the classifier module is stored on the computer-readable storage media for execution by the processor via the computer-readable memory. 2. A computer system according to claim 1, wherein the training module divides the set of neural response training data into a first subset of neural response training data and a second subset of neural response training data, uses the loss function to identify a first subset of optimal weights for the first subset of neural response training data, and identifies a second subset of optimal weights for the second subset of neural response training data. 3. A computer system according to claim 2, wherein the training module compares the first subset of optimal weights and the second subset of optimal weights for verification. 4. A computer system according to claim 1, wherein the new set of neural response data is collected by either EEG equipment or fMRI equipment. 5. A computer system according to claim 1, further comprising a classifier module that applies the classifier to a set of neural response test data. 6. A computer system according to claim 1, wherein the classifier module determines, for a data point in the set of neural response test data, a probability that the data point is associated with an available choice. 7. A computer system according to claim 6, wherein the available choice is one of three or more available choices. 8. A computer-implemented method for analyzing neural response data, the method comprising: identifying, by a training module, a set of optimal weights that produce a minimum for a loss function given a set of neural response training data according to an iterative procedure of setting at least one of the optimal weights to zero and identifying features based on partial derivatives of the loss function; andcreating, by the training module, a classifier using the set of optimal weights. 9. A method according to claim 8, further co... | Leidos, Inc. | G06F015/00 G06F015/18 A61B005/04 A61B005/0476 A61B005/055 A61B005/16 G06F019/00 G01R033/48
US09349301 | B2 | 2016-05-24 | Sensor-based movement guidance | 2011-01-28 | Technologies are generally described for methods, instructions, and applications for a user-guidance system. In some examples, the system may include sensors to measure current weight distribution of a user of the user-guidance system and to transmit resulting measured weight distribution data, a guidance apparatus to calculate a subsequent user movement based on at least the measured weight distribution data and to transmit resulting user movement instructions, and an actuator to produce actuating signals, based on at least the user movement instructions, that provide instructions for the subsequent user movement. | 1. A user-guidance system, comprising: sensors to measure current weight distribution on one of a user's legs;a visual sensor to gather visual information regarding the user's current geographic location and physical surroundings including any obstacles;a guidance apparatus to: receive the measured current weight distribution from the sensors,receive the gathered visual information regarding the user's current geographic location and physical surroundings including any obstacles from the visual sensor, andcalculate a subsequent stride for one of the user's legs based on at least the received weight distribution data of the one of the user's legs and the received visual information, wherein the calculated subsequent stride for one of the user's legs includes a horizontal stride direction, a vertical stride direction, a stride length, a stride angle, a stride depth, a stride height, and a stride rate for the subsequent stride for one of the user's legs; andan actuator to provide timely guidance instructions for an appropriate subsequent stride for the user by: producing actuating signals based on at least the calculated subsequent stride for one of the user's legs, andtransmitting the actuating signals to instruct the subsequent stride for one of the user's legs. 2. The user-guidance system of claim 1, wherein the sensors are configured to be attached to at least one of the user's shoes. 3. The user-guidance system of claim 1, wherein the sensors are configured to be user-adorned and transmit the measured weight distribution data via a short-range communication system. 4. The user-guidance system of claim 1, wherein the guidance apparatus is configured to be embedded in a smart phone. 5. The user-guidance system of claim 1, wherein the guidance apparatus is configured to be user-retained and transmit the actuating signals to instruct the subsequent stride for one of the user's legs via a short-range communication system. 6. The user-guidance system of claim 1, wherein the actuator is configured to be worn across upper torso of the user. 7. The user-guidance system of claim 1, wherein the actuator is configured to be attached to the sensors. 8. The user-guidance system of claim 1, wherein the actuating signals differ in length, number, or intensity. 9. A guidance system, comprising: at least one weight sensor to determine current weight distribution on one of a user's legs;a visual sensor to provide visual data regarding the user's physical surroundings including any obstacles;a guiding device to: receive the determined current weight distribution from the at least one weight sensor,receive the visual data regarding the user's physical surroundings including any obstacles from the visual sensor, andtranslate a subsequent stride for one of the user's legs in response to the received current weight distribution on one of the user's legs and the received visual data into instructions to direct the user's physical stride for one of the user's le... | Empire Technology Development LLC | G09B021/00 G08B023/00 G08G001/095 A63B023/08 A61H003/06 A61B005/11 G01S015/93 A61B005/00 A61B005/103
US09326695 | B1 | 2016-05-03 | Electrode harness and method of taking biopotential measurements | 2004-11-12 | The present invention relates to an electrode harness and more particularly to an electrode harness with various features, which enhance the use and performance of the electrode harness. The present invention further relates to a method of taking biopotential measurements. The electrode harness and methods of the present invention allow for use with most applications where biopotential measurements are taken. The electrode harness can be used in ECG (or EKG), EEG, EMG, and other such biopotential measurement applications. Because of the versatility of various embodiments of the present invention, preferably the electrode harness can be adjusted for different applications or for application to various sized and shaped subjects. The electrode harness is further preferably part of a system, which includes either wireless or tethered bridges between the electrode harness and a monitor, and preferably includes various forms of processors for analyzing the biopotential signal. | 1. An electrode harness for physiological monitoring of a subject, the electrode harness comprising material operable to interconnect at least two electrodes; andat least two electrode connectors attached to the materialwherein the at least two electrode connectors are non-releasable connectors for attaching to electrodes with male connectors. 2. The electrode harness in claim 1, wherein the electrode harness further includes a wireless transmitter or transceiver, which is electrically connected to the at least two electrode connectors. 3. The electrode harness in claim 1, wherein the material comprises at least two expandable arms, each of the at least two expandable arms corresponding to the at least two electrodes. 4. The electrode harness in claim 1, wherein the male connectors are button connectors. 5. The electrode harness in claim 1, wherein the material comprises an electrical pathway from at least one of the at least two electrode connectors, the electrical pathway being electrically shielded from large defibrillator voltages. 6. The electrode harness in claim 1, the electrode harness further comprising a marker which can be used to line up the electrode harness with an anatomical feature of the subject. 7. An electrode harness for physiological monitoring of a subject, the electrode harness comprising: at least two electrodes;material operable to mechanically and electrically interconnect the at least two electrodes; anda track system attached or integral to the materialwherein at least one of the at least two electrodes is movable along the track system and one of the at least two electrodes is a dry electrode. 8. An electrode harness for physiological monitoring of a subject, the electrode harness comprising at least two dry electrodes; andmaterial operable to mechanically and electrically interconnect the at least two dry electrodes, the material comprising an electrical pathway from at least one of the at least two dry electrodes, the electrical pathway being shielded on a side of the pathway such that the shield is adapted to be positioned between the subject and the electrical pathway from large defibrillator voltages;wherein the at least two dry electrodes are attached to the material, the electrode harness further comprises a track system attached or integral to the material with at least one electrode movable along the track system with a connector, and the electrode movable along the track has a spring type connector, which moves with the electrode along the track and keeps the electrode connected to a conductive layer of the material. | Orbital Research Inc | A61B005/0408
US09344777 | B2 | 2016-05-17 | Wireless communication authentication for medical monitoring device | 2012-12-20 | Methods, systems, and devices for short-range low-power wireless communication of analyte information are provided. In some implementations, short-range low-power wireless communication of analyte information may include receiving an electromagnetic wireless communication signal and harvesting energy from the electromagnetic wireless communication signal. In some implementations, short-range low-power wireless communication of analyte information may include enabling capabilities associated with an external sensor in response to detecting the external sensor. In some implementations, short-range low-power wireless communication of analyte information may include detecting an analyte sample; determining an analyte concentration associated with the detected analyte sample; and transmitting an indication of the analyte concentration to an external device. | 1. An apparatus, comprising: a housing;a memory unit configured to store computer executable instructions;a monitor adapted to analyze fluid on a test strip;a transceiver; anda processor configured to execute the computer executable instructions stored in the memory unit to: enable a sensor interface and search for an external sensor with the sensor interface while user interface (UI) capabilities of the apparatus associated with the sensor interface are disabled, the search for the external sensor includes transmission of a signal over a determined period of time,in response to not detecting the external sensor within the determined period of time, terminate transmission of the signal and disable the enabled sensor interface, andin response to detecting the external sensor, enable the disabled UI capabilities associated with the sensor interface, activate the external sensor, and determine an analyte concentration from sample information received from the external sensor over the sensor interface. 2. The apparatus of claim 1, wherein the transceiver is configured to perform short-range low-power wireless communication, and the processor is configured to search for the external sensor via short-range low-power wireless communication. 3. The apparatus of claim 2, wherein the short-range low-power wireless communication includes a near field communication protocol or a radio frequency identification protocol. 4. The apparatus of claim 1, wherein, when disabled, the UI capabilities associated with the sensor interface are hidden from a user of the apparatus. 5. The apparatus of claim 1, wherein capabilities associated with the monitor adapted to analyze the fluid on the test strip are enabled while the UI capabilities associated with the external sensor are disabled. 6. A method comprising: enabling a sensor interface of an apparatus, wherein the apparatus comprises a monitor adapted to analyze fluid on a test strip;searching for an external sensor with the sensor interface while user interface (UT) capabilities of the apparatus associated with the sensor interface are disabled, the searching for the external sensor includes transmitting a signal for a determined period of time;in response to not detecting the external sensor within the determined period of time, terminating the signal being transmitted and disabling the enabled sensor interface; andin response to detecting the external sensor, enabling the disabled UI capabilities associated with the sensor interface, activating the external sensor, and determining an analyte concentration from sample information received from the external sensor over the sensor interface. 7. The method of claim 6, wherein searching for the external sensor includes using short-range low-power wireless communication. 8. The method of claim 7, wherein using short-range low-power wireless communication includes using a near field communication protocol or a radio frequency identification protocol. 9. The metho... | ABBOTT DIABETES CARE INC. | G08C017/02 H04Q009/00 A61B005/00 A61B005/145
US09351681 | B2 | 2016-05-31 | One repetition maximum (1RM) estimating apparatus and method | 2013-11-18 | A one repetition maximum (1RM) estimating apparatus includes a potential signal receiving unit configured to receive a potential signal of a muscle from a sensor attached to a user, and a 1RM estimating unit configured to estimate a 1RM of the user based on the potential signal. | 1. An apparatus for estimating a one repetition maximum (1RM), the apparatus comprising: a potential signal receiving unit configured to receive a potential signal of a muscle from an electromyography (EMG) sensor attached to a user; anda 1RM estimating unit configured to estimate a 1RM of the user based on the potential signal. 2. The apparatus of claim 1, wherein the 1RM estimating unit comprises: an envelope determining unit configured to determine an envelope of amplitudes of the potential signal;a frequency spectrum determining unit configured to determine a mean frequency of a frequency spectrum of the potential signal; anda gradient determining unit configured to determine a gradient based on the envelope and the mean frequency of the frequency spectrum. 3. The apparatus of claim 2, wherein the envelope determining unit is further configured to determine the envelope based on an absolute value of the potential signal. 4. The apparatus of claim 2, wherein the frequency spectrum determining unit is further configured to determine the mean frequency of the frequency spectrum based on the potential signal sampled at predetermined intervals based on a maximum value of the envelope. 5. The apparatus of claim 2, wherein the gradient determining unit is further configured to determine the gradient based on a plurality of feature points in the mean frequency of the frequency spectrum corresponding to peaks of the envelope. 6. An apparatus for estimating a one repetition maximum (1RM), the apparatus comprising: an electromyography (EMG) sensing unit configured to sense a potential signal of a muscle of a user; anda 1RM estimating unit configured to estimate a 1RM of the user based on the potential signal. 7. The apparatus of claim 6, wherein the 1RM estimating unit comprises: an envelope determining unit configured to determine an envelope of amplitudes of the potential signal;a frequency spectrum determining unit configured to determine a mean frequency of a frequency spectrum of the potential signal; anda gradient determining unit configured to determine a gradient based on the envelope and the mean frequency of the frequency spectrum. 8. The apparatus of claim 7, wherein the envelope determining unit is further configured to determine the envelope based on an absolute value of the potential signal. 9. The apparatus of claim 7, wherein the frequency spectrum determining unit is further configured to determine the mean frequency of the frequency spectrum based on the potential signal sampled at predetermined intervals based on a maximum value of the envelope. 10. The apparatus of claim 7, wherein the gradient determining unit is further configured to determine the gradient based on a plurality of feature points in the mean frequency of the frequency spectrum corresponding to peaks of the envelope. 11. A method of estimating a one repetition maximum (1RM), the method comprising: receiving a potential signal of a muscle from an electromyogr... | Samsung Electronics Co., Ltd. | A63B024/00 A61B005/22 A61B005/0488 A61B005/00
US09326660 | B2 | 2016-05-03 | Endoscope system with insertion support apparatus | 2014-09-16 | An endoscope system includes: an image storage section that stores image information associated with position information of a luminal organ of a subject; a virtual endoscopic image generating section that generates a virtual endoscopic image endoscopically depicted; an image pickup section that picks up an image in the luminal organ; a position registration processing section that compares the virtual endoscopic image and an endoscopic image, and extracts an virtual endoscopic image similar to the endoscopic image; a position estimation section that estimates a distal end position of an endoscope insertion portion; a feature region determination section that determines whether or not the distal end position exists in a feature region; a storage section that stores information for presentation; and a candidate information presentation control section that presents the virtual endoscopic image based on the information in the storage section, to a display section as candidate information. | 1. An endoscope system comprising: an image storage section that stores image information associated with position information of a luminal organ of a subject;a virtual endoscopic image generating section that generates a virtual endoscopic image endoscopically depicted from an arbitrary view point position based on the image information;an image pickup section that is provided in an endoscope and picks up an image in the luminal organ of the subject;a position registration processing section that compares the virtual endoscopic image generated by the virtual endoscopic image generating section and an endoscopic image generated by image pickup by the image pickup section, and extracts an virtual endoscopic image similar to the endoscopic image;a position estimation section that estimates a distal end position of an endoscope insertion portion in the luminal organ of the subject based on position information associated with the virtual endoscopic image extracted by the position registration processing section;a feature region determination section that determines whether or not the distal end position of the endoscope insertion portion estimated by the position estimation section exists in a feature region related to a branch in the luminal organ;a storage section that stores, when it is determined that the distal end position of the endoscope insertion portion exists in the feature region by the feature region determination section, the distal end position of the endoscope insertion portion estimated by the position estimation section; anda candidate information presentation control section that, when the position registration processing section fails to extract a virtual endoscopic image similar to the endoscopic image, presents the virtual endoscopic image corresponding to the distal end position of the endoscope insertion portion which is determined to exist within the feature region by the feature region determination section based on the distal end position of the endoscope insertion portion stored in the storage section, to a display section as candidate information. 2. An endoscope system according to claim 1, further comprising: a luminal organ extracting section that extracts a specified luminal organ from the image information and generates a three-dimensional luminal organ image; anda route data generating section that generates route data from an insertion start position to a target position in at least the specified luminal organ based on the image information,wherein the candidate information presentation control section superimposes the route data on the three-dimensional luminal organ image or a two-dimensional tomographic image which is obtained by cutting the image information in an arbitrary direction, and presents the candidate information upon the route data or in a vicinity of the route data. 3. An endoscope system according to claim 2, comprising a storage limiting section that limits storage of the distal end position of... | OLYMPUS CORPORATION | A61B001/04 A61B001/267 A61B001/00 A61B005/06 A61B006/03 A61B006/00 A61B001/05
US09351678 | B2 | 2016-05-31 | Lancet holder and lancet device including the same | 2011-09-20 | A lancet holder includes a holder body having a lancet seat at one end to mount a lancet and receiving the lancet and elastic fixing portions arranged in pairs to face each other on the inner side of the holder body so that the lancet received in the holder body is vertically held, and elastically fixing the lancet. | 1. A lancet device comprising: a housing;a lancet holder mounted to the housing, and comprising: a holder body having a cavity defining a lancet seat at a first end of the holder body for receiving a lancet needle;a plurality of elastic fixing portions arranged in pairs to face each other on an inner side of the cavity of the holder body operable to engage the lancet needle so that the lancet needle received in the cavity of the holder body is vertically held, and elastically fixing the lancet needle;a lancet plate operable to engage the lancet needle, the lancet plate disposed within the cavity of the holder body;an elastic support member disposed within the cavity of the holder body, under the lancet plate, and configured to compensate for a variable length of the lancet needle; anda movable shaft coupled to an outer surface of the holder body for releasing and moving the lancet needle. 2. The lancet device of claim 1, wherein the plurality of elastic fixing portions are arranged at predetermined angles around the inner side of the cavity of the holder body. 3. The lancet device of claim 2, wherein the plurality of elastic fixing portions include: a plurality of first fixing members arranged at predetermined angles around the inner side of the cavity of the holder body and adjacent to an open end of the cavity of the holder body; anda plurality of second fixing members arranged at predetermined angles around the inner side of the cavity of the holder body, wherein the plurality of second fixing members are positioned proximally to the plurality of first fixing members. 4. The lancet device of claim 3, wherein the plurality of elastic fixing portions are plates protruding inward. 5. The lancet device claim 1, further comprising a guide wing elongated longitudinally along the outer surface of the holder body and operable to move along a releasing guide groove formed on an inner side of the housing. 6. The lancet device of claim 2, wherein the plurality of elastic fixing portions are plates protruding inward. | I-SEN, INC. | A61B017/34 A61B005/151 A61B005/15
US10939869 | B2 | 2021-03-09 | Imaging system and method for enhanced visualization of near surface vascular structures | 2014-06-06 | A sub-dermal structure visualization system and method is disclosed. The system may use an illumination module including: a near-infrared (NIR) light source adapted to substantially uniformly illuminate an imaged area including sub-dermal regions thereof with NIR light; and a first optical system including at least one optical element for controlling at least one of spectral and polarization properties of the NIR light prior to illuminating the imaged area. An imaging module rejects unwanted signals from an imaged area while passing desired optical signals that are to be received by an image acquisition module. The desired optical signals can comprise a vein visualization signal to assist in visualizing a vascular structure below a skin layer of a patient. | 1. A sub-dermal structure visualization system, comprising: a stationary illumination subsystem, comprising:a near-infrared (NIR) light source including at least one of an LED, a laser or a laser diode, adapted to illuminate an imaged area of interest while the NIR light source projects an uniform and constant illumination over the imaged area of interest, and wherein the illuminated imaged area of interest is substantially larger than a vein, the NIR light source having a specific wavelength and an illumination bandwidth determined upon at least one of a skin pigmentation and a presence of body fat in a region of a subject being imaged; anda first optical system configured with at least one polarizing element for controlling spectral and polarization properties of the NIR light source directed to the illuminated imaged area;a stationary imaging subsystem, comprising:a second optical system including an optical filter configured to reject unwanted one or more optical components of an optical signal returning from the illuminated imaged area while passing a part of the returning optical signal having one or more desired spectral and polarization properties;a lens system arranged to optically relay an image as provided by a configuration selected from: a predetermined magnification and focus configuration and an adjustable magnification and focus configuration; andan image acquisition subsystem including a two dimensional array detector configured to collect the image from the stationary imaging subsystem and to select therefrom the passing part of the returning optical signal, wherein the passing part of the returning optical signal comprises image intensity histogram containing a vein visualization signal representing the vein;an image enhancing subsystem including signal electronics configured to select for display from the image intensity histogram only a sub-range of intensities forming a subportion of the passing part of the returning optical signal that comprises only the vein visualization signal, to thus enhance the vein visualization signal; andan image display module configured with at least one of an electronic visual display and an image projector that displays the image with at least one display property selected from: an aspect ratio, a desired resolution, and an image contrast that enhances the image provided by the image enhancing subsystem. 2. The sub-dermal structure visualization system of claim 1, wherein the imaging subsystem includes means for generating image display fiducials transposed from the illuminated imaged area to the image, wherein the illuminated imaged area forms an area of interest (AOI) to enhance spatial correlation. 3. The sub-dermal structure visualization system of claim 1, wherein the illumination produced is at least one of: a uniform illumination; ora non-uniform illumination which is corrected by a preset correction function to produce a uniform illumination. 4. The sub-dermal structure visualization... | Lawrence Livermore National Security, LLC | A61B005/00 A61B005/02
US09339228 | B2 | 2016-05-17 | Adjustable thoracic garments useable by patients having different thoracic dimensions for locating EM transducer(s) in proximity to predefined thoracic anatomic features | 2013-12-18 | A thoracic garment for bringing an EM transducer to contact with a thoracic surface area of a wearer. The thoracic garment comprises a front thoracic garment piece and a back thoracic garment piece the front and/or back thoracic garment pieces having EM transducer(s), the front thoracic garment piece having an arrangement of passages formed therein or thereon, a handhold mechanically connected to the front thoracic garment piece, a positioning marker, and a plurality of straps each threaded in the arrangement of passages. The thoracic garment is configured to be adapted by the wearer which uses a first hand to hold the handhold for mounting the front thoracic garment piece against or close to the chest such that the positioning marker is positioned at the predefined position while using a second hand for pulling the straps. | 1. A thoracic garment for bringing an EM transducer to contact with a thoracic surface area of a wearer, comprising: a front thoracic garment piece and a back thoracic garment piece at least one of said front and back thoracic garment pieces having at least one EM transducer or at least one mechanical connector for connecting said at least one EM transducer, said front thoracic garment piece having an arrangement of passages formed therein or thereon;a handhold mechanically connected to said front thoracic garment piece;a positioning marker located on or in said front thoracic garment piece such that when said thoracic garment is worn by a human wearer and said front thoracic garment piece is placed against or close to the chest of said human wearer, said positioning marker located in a predefined position in relation to one or more anatomical features of the wearer;a plurality of straps each threaded in said arrangement of passages;wherein said thoracic garment is configured to be adapted by said wearer using a first hand to hold said handhold for holding said front thoracic garment piece against or close to said chest such that said positioning marker is positioned at said predefined position while using a second hand for pulling said plurality of straps. 2. The thoracic garment of claim 1, further comprising at least one brake unit to lock the movement of each of said plurality of straps. 3. The thoracic garment of claim 1, further comprising at least one pulling element associated with each of said plurality of straps and adapted to be pulled by said second hand for pulling said plurality of straps said plurality of. 4. The thoracic garment of claim 3, wherein said arrangement of passages is symmetrically arranged in or on said front thoracic garment piece so that said at least one pulling element is configured to pull symmetrically said plurality of straps along the left and right sides of said human wearer. 5. The thoracic garment of claim 1, wherein said positioning marker is integrated with said handhold. 6. The thoracic garment of claim 1, wherein when the garment is configured to be adapted by said wearer said handhold is located in a midway between armpits and below the bottom tip of the anatomical feature. 7. The thoracic garment of claim 1, wherein said anatomical feature is the jugular notch of said human wearer. 8. The thoracic garment of claim 1, wherein said front thoracic garment piece comprises a non-stretchable breast plate. 9. The thoracic garment of claim 8, wherein said breast plate has said at least one EM transducer mounted thereon or therein. 10. The thoracic garment of claim 1, wherein said back thoracic garment piece comprises a non-stretchable back plate mechanically connected to at least some of said plurality of straps. 11. The thoracic garment of claim 1, wherein at least one of said front and back thoracic garment pieces has at least one mechanical connector for at least one EM transducer configured to d... | Sensible Medical Innovations Ltd. | A61B005/103 A61B005/00 A61B005/05 A61B005/053 A61B005/01 A61B005/08
US10448997 | B2 | 2019-10-22 | Measurement device for skin properties and non-invasive treatment device | 2014-06-25 | The application provides a non-invasive measurement device (30) and method for the measurement of skin properties using laser light, the device comprising a probe module (70) and an imaging module (50). The application also provides a non-invasive treatment device (130, 230, 330, 430) and method comprising the measurement device/method. The probe module (70) provides a probe light beam (82) that enters the skin along the probe axis (71). The probe light beam (82) is separate from any treatment radiation beam (22), so that the probe light beam (82) can be optimized for the measurement. A more reliable skin measurement system is provided because it measures a plurality of positions along the probe axis (71), within a probe region (95). The measurement device and method may also comprise a treatment module (110) or treatment step. The skin parameters measured may then be directly used to control the treatment parameters. This results in a skin treatment device (10) being provided that is both effective and delivers reproducible results. | 1. A non-invasive measurement device for measurement of skin properties using laser light, the device comprising: a probe module comprising: a laser light source configured to generate a probe light beam along a probe axis and impinge on an outer surface of a skin to be treated; anda first optical system configured to direct the probe light beam to a probe region inside the skin; andan imaging module, comprising: a second optical system comprising: an imaging optical axis that intersects with the probe axis; andan optical detector array, disposed along a detector axis in an image plane, comprising a plurality of light detection elements arranged in a plurality of rows and a plurality of columns, wherein the plurality of columns are positioned along the detector axis;the second optical system configured to: combine outputs of the plurality of light detection elements within one of: a same row and a same column received at a same time, wherein the outputs of the plurality of light detection elements are combined separately from the outputs of the plurality of light detection elements in another one of said plurality of rows or said plurality of columns, the outputs corresponding to a plurality of probe positions distributed along the probe axis within the probe region;form an image based on the combined outputs; andgenerating at least one control parameter associated with at least one of: the probe module and the imaging module based on the formed image, wherein said at least one control parameter altering at least one of: the laser light and the forming of said image. 2. A measurement device according to claim 1, wherein the detector axis is comprised in a plane that comprises the probe axis and the imaging optical axis. 3. The measurement device according to claim 1, wherein an angle enclosed by the probe axis and the imaging optical axis is in a range of 20 to 90 degrees. 4. The measurement device according to claim 1, wherein an angle enclosed by the probe axis and the skin outer surface is in a range of 45 to 90 degrees. 5. The measurement device according to claim 1, wherein an angle enclosed by the detector axis and the imaging optical axis is corrected according to a Scheimpflug principle. 6. The measurement device according to claim 1, wherein the optical detector array comprises: a matrix of Y light detection elements extending along the detector axis and X light detection elements extending along an axis perpendicular to the detector axis, wherein a ratio of said Y light detector elements to said X light detector elements is greater than 5 to 1. 7. A non-invasive treatment device for treatment of skin using electromagnetic treatment radiation, the device comprising: a probe module comprising: a laser light source configured to generate a probe light beam along a probe axis and impinge on an outer surface of a skin to be treated; anda first optical system configured to direct the probe light beam to a probe region inside the skin; a... | KONINKLIJKE PHILIPS N.V. | A61B018/20 A61N005/06 A61B005/00 A61B018/00
US09382107 | B2 | 2016-07-05 | Hydration monitoring system | 2016-01-27 | An activity and volume sensing beverage container cap system for a beverage container including a cap that couples with the beverage container, a processor, a timer and at least one activity level sensor, such as an inclinometer, coupled with the processor. The activity level sensor detects container orientation. When the container is tilted, the timer measures an amount of time in an orientation. Based on the current and previous orientation and the amount of time the container is in different orientations, the processor determines an amount of time of activity of the user, and a level of activity of the user, such as number of steps or moves per time. The inclinometer is also used to determine the amount of volume dispensed from the container when tilted over a threshold indicative of drinking. Thus one inclinometer may be utilized as both to determine activity level and volume of fluid dispensed. | 1. A hydration monitoring system comprising: a housing configured to couple with a beverage container wherein said housing comprises an opening that dispenses contents of said beverage container;a moveable seal that closes said opening; and,one or more sensors coupled with said housing or that extend into said beverage container;a wireless communication interface coupled with said housing;a memory coupled with said housing;a processor coupled with said memory and said wireless communication interface and situated in said housing wherein said processor is configured toreceive sensor data from said one or more sensors;calculate or estimate an amount of fluid inside of, an amount of fluid dispensed from or an amount of fluid added to said beverage container from said sensor data;store said amount of fluid in said memory; and,transmit said amount of fluid dispensed from said beverage container when a wireless communication channel is available to an external device; and,a display coupled to said processor and configured to indicate said amount of fluid dispensed from said beverage container. 2. The system of claim 1, wherein said processor is further configured to synchronize personal profile parameters when said wireless communication channel is available to said external device. 3. The system of claim 1, wherein said processor is further configured to operate autonomously when said wireless communication channel is not available. 4. The system of claim 1, wherein said external device is a smartphone, a tablet computer, a laptop, or a cloud-based system. 5. The system of claim 1, wherein said one or more sensors comprise one or more of a capacitance probe, a resistance probe, an ultrasonic sensor, an optical sensor, a mechanical float, a pressure sensor or a weight sensor. 6. The system of claim 1, wherein said one or more sensors comprise a tilt sensor or inclinometer and further comprising a valve and a timer that measures how long said valve is open and wherein said processor is further configured to estimate an amount of fluid dispensed from said beverage container based on data from said tilt sensor or inclinometer and said timer in combination with a predetermined size of said opening in said housing configured to dispense said liquid. 7. The system of claim 6, wherein said tilt sensor or inclinometer is also utilized to calculate an activity level based on changes in tilt over time. 8. The system of claim 1, wherein said processor resets said amount of fluid dispensed from said beverage container at a regular time interval. 9. The system of claim 1, wherein said one or more sensors comprise one or more of an accelerometer, a gyroscope, a GPS system, a pedometer, a thermometer, a physiological monitor, a heart rate sensor, a respiration rate sensor, a blood pressure sensor, a blood oxygen level sensor, a hydration sensor, a humidity sensor, an altitude sensor, and a barometric pressure sensor. 10. The system of claim 9, wherein said d... | IQHYDR8, LLC | B67D003/00 A61B005/02 A61B005/11 G01F023/00 A61B005/024 B67D001/00
US09332916 | B2 | 2016-05-10 | Systems, devices, and methods for determining fractional flow reserve in the presence of a catheter | 2014-04-15 | Systems, devices, and methods for obtaining fractional flow reserve in the presence of a catheter. In a method of determining a fractional flow reserve in the presence of a catheter, the method comprises the steps of obtaining measurements of an inner luminal organ diameter proximal to, at, and distal to a stenosis and a length of the stenosis, obtaining a pressure drop measurement at the stenosis, calculating a volumetric flow of fluid through the inner luminal organ at the stenosis, and determining a stenotic pressure drop at the stenosis corresponding to dimensions of the guidewire as a function of the calculated volumetric flow of fluid through the inner luminal organ at the stenosis, wherein the stenotic pressure drop is indicative of a fractional flow reserve at or near the stenosis. | 1. A method of determining a fractional flow reserve in the presence of a catheter, the method comprising the steps of: measuring a pressure gradient change within a luminal organ at a stenosis using a catheter having a known size; anddetermining a fractional flow reserve at or near the stenosis based upon the pressure gradient change, luminal organ size measurements determined based upon conductance measurements obtained using an impedance wire positioned at least partially within a lumen of the catheter, and a difference between the known size of the catheter and a known size of the impedance wire. 2. The method of claim 1, wherein the luminal organ size measurements comprise a luminal organ size measurement proximal to the stenosis and a luminal organ size measurement distal to the stenosis. 3. The method of claim 2, wherein the luminal organ size measurements further comprise a luminal organ size measurement at the stenosis. 4. The method of claim 3, wherein the step of determining is also based upon an identified length of the stenosis. 5. The method of claim 1, wherein the step of determining is also based upon a calculated volumetric flow of fluid through the luminal organ at the stenosis. 6. The method of claim 1, wherein the step of determining is performed by: (a) obtaining the luminal organ size measurements using the impedance wire, the luminal organ size measurements comprising a luminal organ size measurement proximal to the stenosis and a luminal organ size measurement distal to the stenosis;(b) determining a length of the stenosis based on various conductance measurements obtained using the impedance wire at or near the stenosis;(c) calculating a volumetric flow of fluid through the luminal organ at the stenosis as a function of the luminal organ size measurements, the length of the stenosis, and the difference between the known size of the catheter and the known size of the impedance wire; and(d) determining the fractional flow reserve based in part upon the calculated volumetric flow;wherein (a) and (b) are performed together or in any order. 7. The method of claim 6, wherein the determination in (d) is further based in part upon a mean aortic pressure determined using the catheter. 8. The method of claim 1, further comprising the step of: determining a functional assessment of stenosis severity using the fractional flow reserve. 9. The method of claim 1, wherein the step of measuring the pressure gradient change is performed using a procedure selected from the group consisting of (i) angiography and intravascular ultrasound and (ii) angiography and optical coherence tomography. 10. The method of claim 1, wherein the step of measuring the pressure gradient change is performed using an apparatus selected from the group consisting of a pressure catheter and the catheter having a pressure sensor. 11. A system for determining a fractional flow reserve, comprising: an impedance wire configured to obtain conductance measurem... | 3DT Holdings, LLC | A61B005/026 A61B008/06 A61B005/00 A61B006/00 A61B005/0215 A61B008/12 A61B005/02 A61B005/027 A61B005/053
US09642967 | B2 | 2017-05-09 | Catheter monitor integration with patient support, hygiene and healthcare communication systems | 2015-08-19 | Catheter monitor integration with patient support systems, such as patient beds, and with other healthcare communication systems including hand hygiene systems is provided. A patient bed that transmits bed data and catheter monitor data is disclosed. When catheter tubing is connected to a catheter monitor integrated into a patient bed, the bed automatically implements a falls prevention protocol. A countdown timer pertaining to removal of catheter tubing from a patient is displayed on one or more display devices. Monitoring for caregiver handwashing compliance in connection with use of a catheter monitor or a catheter kit is also provided. | 1. A patient bed comprising a frame configured to support a patient,a barrier coupled to the frame and configured to inhibit patient egress from the frame,a catheter monitor, the barrier having a cavity shaped to receive at least a portion of the catheter monitor, andbed control circuitry carried by the frame and configured to control at least one function of the patient bed, the catheter monitor having monitor circuitry in communication with the bed control circuitry. 2. The patient bed of claim 1, further comprising at least one graphical user interface (GUI) coupled to the bed control circuitry, the GUI being operable to display bed information pertaining to the at least one function of the patient bed and to display monitor information pertaining to at least one function of the catheter monitor. 3. The patient bed of claim 1, wherein the bed control circuitry is configured to communicate to a network of a healthcare facility bed information pertaining to the at least one function of the patient bed and monitor information pertaining to at least one function of the catheter monitor. 4. The patient bed of claim 1, wherein the bed control circuitry is configured to communicate to a network of a healthcare facility bed identification (ID) data and catheter monitor identification (ID) data. 5. The patient bed of claim 4, wherein the bed control circuitry is configured to communicate to the network of the healthcare facility patient identification (ID) data. 6. The patient bed of claim 1, further comprising an additional piece of patient care equipment, the cavity being sized to receive at least a portion of the additional piece of patient care equipment when the catheter monitor is removed from the cavity. 7. The patient bed of claim 6, wherein the additional piece of patient care equipment comprises a pulse oximeter or a sequential compression device. 8. A patient support apparatus comprising a patient support structure to support a patient,control circuitry to control at least one function of the patient support structure, the control circuitry being operable to implement a falls prevention protocol, anda catheter monitor coupled to the patient support structure, the control circuitry implementing the falls prevention protocol in response to a signal from the catheter monitor indicating at least one of the following: the catheter monitor being turned on, a catheter tube being attached to the catheter monitor, and urine flowing through the catheter monitor. 9. The patient support apparatus of claim 8, wherein the falls prevention protocol comprises an alert being generated to indicate at least one or more of the following: that casters of the patient support apparatus are not braked or that at least one siderail of the patient support apparatus is not in a raised position or that an upper frame of the patient support apparatus is not in a lowered position relative to a base frame of the patient support apparatus. 10. The patient support ... | Hill-Rom Services, Inc. | H01C003/10 A61M005/172 A61H009/00 A61B005/20 A61G007/018 A61M025/00 A61M005/50 A61G007/05 A61G007/08 G06F019/00 A61B005/00
US10238299 | B2 | 2019-03-26 | Photoacoustic image-generating apparatus and light source control method | 2016-01-20 | At least a tip portion of a puncture needle 15 is inserted into a subject. The puncture needle 15 includes a light guide member which guides light from a laser unit 13, and a light emission unit which is provided in the vicinity of the tip portion and emits light guided by the light guide member, and generates a photoacoustic wave caused by light from the light emission unit in the tip portion. A probe 11 detects the photoacoustic wave from the puncture needle 15 in a state where the puncture needle 15 is inserted into the subject. A Photoacoustic image generation unit 25 generates a photoacoustic image based on the detected photoacoustic wave. A Light source control unit 30 controls the amount of light emitted from the light emission unit based on the photoacoustic image. | 1. A photoacoustic image-generating apparatus comprising: a light source;an inserted object, of which at least a tip portion is adapted to be inserted into a subject, the inserted object including a light guide member which guides light from the light source, and an end surface of the light guide member in light advancing direction from the light source that is provided in a vicinity of the tip portion and emits the light guided by the light guide member, and generating a photoacoustic wave caused by the light from the end surface of the light guide member in the light advancing direction from the light source in the tip portion;a probe that detects the photoacoustic wave from the inserted object in a state where at least the tip of the inserted object is inserted into the subject; anda processor, the processor configured to generate a photoacoustic image based on the detected photoacoustic wave,analyze the photoacoustic image to determine pixel values in the photoacoustic image, where the analysis comprises generating a distribution graph of the pixel values in the photoacoustic image; andbased on the distribution graph, determining an area of a pixel having a pixel value of a peak in the photoacoustic image equal to or greater than a first threshold value and a pixel value around the peak equal to or greater than a second threshold value; andcontrol an updated set amount of light emitted from the end surface of the light guide member in the light advancing direction from the light source based on the the determined area, wherein the updated set amount of light is controlled such that the determined area is equal to or less than a predetermined area. 2. The photoacoustic image-generating apparatus according to claim 1, wherein the inserted object further includes a light absorption member which absorbs light from the end surface of the light guide member in the light advancing direction from the light source to generate the photoacoustic wave. 3. The photoacoustic image-generating apparatus according to claim 1, wherein the inserted object further includes a light absorption member which absorbs light from the end surface of the light guide member in the light advancing direction from the light source unit to generate the photoacoustic wave. 4. The photoacoustic image-generating apparatus according to claim 2, wherein the inserted object is a puncture needle which has a communication lumen, has an opening at the tip, and punctures the subject, andat least one of the light guide member and the light absorption member becomes an inner needle which seals at least a part of the lumen of the puncture needle. 5. The photoacoustic image-generating apparatus according to claim 3, wherein the inserted object is a puncture needle which has a communication lumen, has an opening at the tip, and punctures the subject, andat least one of the light guide member and the light absorption member becomes an inner needle which seals at least a part of the lumen... | FUJIFILM Corporation | A61B005/00 A61B008/00 A61B008/08
US10517513 | B2 | 2019-12-31 | Hand function diagnostic and therapeutic system | 2015-10-30 | Systems, devices and methods for quantitatively assessing hand function strength and range of motion via various instrumentations are disclosed. In one example embodiment, a hand function diagnostic and therapeutic system includes distal and proximal finger flexion measurement apparatuses used to automate quantitative assessment of hand function based on a combination of strength, range of motion, and time performance metrics. | 1. A proximal finger flexion measurement apparatus, comprising: a base platform;a palm platform perpendicularly coupled to the base platform, the palm platform having opposing first and second surfaces where the first and second surfaces are substantially perpendicular to the base platform;a hand positioning member perpendicularly coupled to the first surface of the palm platform, and extending a distance away from the first surface of the palm platform,wherein the hand positioning member is coupled to the first surface of the palm platform at a location above the base platform sufficient to allow a user to position a hand between the hand positioning member and the base platform with the palm of the hand in contact with the first surface and adjacent to an edge of the palm platform at a location such that when a user positions the palm of a hand on the first surface of the palm platform with the hand positioning member touching a base of a thumb of the hand between the thumb and an index finger of the hand, the edge of the palm platform is located beneath at least one knuckle of the hand, thereby permitting the user to pivot at least one finger about the at least one knuckle while the palm rests on the first surface of the palm platform;a proximal finger flexion surface rotatably coupled via a pivoting member to and perpendicular with the base platform at a location adjacent to the edge of the palm platform such that the proximal finger flexion surface extends away from the second surface of the palm platform, such that when a user positions the palm of the hand on the first surface of the palm platform with the hand positioning member touching the base of the thumb of the hand between thumb and index finger of the hand, at least one finger of the hand interfaces with the proximal finger flexion surface, andwhere the proximal finger flexion surface is biased to remain in a resting position relative to the palm platform, wherein in the resting position the proximal finger flexion surface forms an angle with the first surface of the palm platform, wherein the angle is substantially equal to or less than 90 degrees, and when the user applies a force to the proximal finger flexion surface via the at least one finger, the proximal finger flexion surface is rotated toward the second surface of the palm platform such that the change in angle between the proximal finger flexion surface and the first surface of the palm platform increases and is measured by an angular displacement sensor in communication with the pivoting member. 2. The apparatus of claim 1, wherein the angular displacement sensor is an optical rotary sensor. 3. The apparatus of claim 1, wherein the hand positioning member is cylindrical. 4. The apparatus of claim 1, wherein the hand positioning member extends at least 1 cm away from the first surface of the palm platform. 5. The apparatus of claim 1, wherein the proximal finger flexion measurement apparatus is configured to measure ... | Oregon Health & Science University | A61B005/00 A61B005/11
US10849509 | B2 | 2020-12-01 | Patient signal filtering | 2014-11-21 | Disclosed herein is a framework for facilitating patient signal filtering. In accordance with one aspect, the framework performs a signature cycle matching pursuit method to remove a first signal component from a combination patient signal and generate a first output signal. Sub-bandwidth filtering of the first output signal may be performed to remove a second signal component and generate a second output signal. The framework may further remove a third signal component from the second output signal to generate a third output signal. A signal of interest may then be reconstructed based on the third output signal. | 1. A non-transitory computer readable medium having embodied thereon computer-executable instructions which, when executed, cause a computer to perform steps for signal filtering, the steps comprising: receiving an electrocardiogram signal, wherein the electrocardiogram signal includes an electrocardiogram signal component, noise, a respiration signal component and a blood pressure signal component;removing the electrocardiogram signal component from the electrocardiogram signal to generate a first output signal by performing an electrocardiogram signature cycle matching pursuit method;eliminating the noise from the first output signal to generate a second output signal by performing sub-bandwidth filtering of the first output signal;removing the respiration signal component from the second output signal to generate a third output signal; andreconstructing a time-domain blood pressure signal based on the third output signal by performing inverse fast Fourier transform (IFFT) on integrated frequency components of the third output signal. 2. The non-transitory computer readable medium of claim 1 having embodied thereon further computer-executable instructions which, when executed, cause the computer to perform the electrocardiogram signature cycle matching pursuit method by deriving a signature cycle from the electrocardiogram signal and removing the electrocardiogram signal component associated with the signature cycle from the electrocardiogram signal. 3. The non-transitory computer readable medium of claim 2 having embodied thereon further computer-executable instructions which, when executed, cause the computer to perform the sub-bandwidth filtering of the first output signal by performing steps comprising: separating, via multiple filters, the first output signal into multiple frequency sub-bandwidths to generate the second output signal; andadaptively adjusting, in response to changes in the second output signal, filter parameters associated with the filters to remove the noise. 4. The non-transitory computer readable medium of claim 1 having embodied thereon further computer-executable instructions which, when executed, cause the computer to remove the respiration signal component by low pass filtering. 5. The non-transitory computer readable medium of claim 1 having embodied thereon further computer-executable instructions which, when executed, cause the computer to clean up the blood pressure signal by applying a Savitzky-Golay smoothing filter. 6. A method of signal filtering, comprising: receiving an electrocardiogram signal that includes first, second, third signal components and a signal component of interest;removing the first signal component from the electrocardiogram signal to generate a first output signal by performing a signature cycle matching pursuit method based on the electrocardiogram signal, wherein performing the signature cycle matching pursuit method includes selecting, from the electrocardiogram signal, reference ... | Siemens Healthcare GmbH | A61B005/00 A61B005/021 A61B005/04 A61B005/0402
US10292881 | B2 | 2019-05-21 | Dynamic apnea therapy surface | 2015-09-09 | Dynamic support surface technologies configure a person support surface in various ways, including configurations in which the person support surface at least temporarily assumes an apnea therapy position, such as a progressive lateral tilt or a ramp-like position. | 1. A dynamic person support system comprising: a person support surface;a lateral rotation apparatus coupled to the person support surface, the lateral rotation apparatus comprising a plurality of independently rotatable longitudinally arranged support planes and a lateral rotation actuator operably coupled to one or more of the support planes; anda control unit comprising a processor and a non-transitory machine readable storage medium comprising a dynamic therapy routine, the dynamic therapy routine comprising instructions executable by the processor to cause the control unit to control the operation of the lateral rotation apparatus by:determining a maximum supine position duration;monitoring the actual supine position duration of a human subject positioned on the person support surface; andcontrolling the lateral rotation actuator to maintain the actual supine position duration below the maximum supine position duration. 2. The dynamic person support system of claim 1, wherein the lateral rotation actuator comprises an electromechanical device configured to drive lateral rotation of the independently rotatable support planes. 3. The dynamic person support system of claim 1, wherein the lateral rotation actuator comprises a plurality of inflatable bladders supporting the independently rotatable support planes and an air supply operably coupled to the inflatable bladders. 4. The dynamic person support system of claim 1, wherein the control unit is configured to compute the maximum supine position duration as a function of an apnea-hypopnea index (AHI) value of the monitored human subject. 5. The dynamic person support system of claim 1, wherein the control unit is configured to compute the maximum supine position duration based on a first apnea-hypopnea index (AHI) value and a second AHI value, wherein the first AHI value is determined while the human subject is in a supine position and the second AHI value is determined while the human subject is in a non-supine position. 6. The dynamic person support system of claim 1, comprising a sensor in communication with the control unit, wherein the control unit is configured to receive a sensed value from the sensor and determine the maximum supine position duration based on the sensed value. 7. The dynamic person support system of claim 6, wherein the sensed value is indicative of an apnea-hypopnea index (AHI) of the monitored human subject. 8. The dynamic person support system of claim 6, wherein the sensed value is indicative of a sleep state of the monitored human subject. 9. The dynamic person support system of claim 6, wherein the control unit is configured to adjust the maximum supine position duration in response to the sensed value. 10. The dynamic person support system of claim 7, wherein the control unit is configured to increase the maximum supine position duration in response to the sensed value being below a threshold value. 11. The dynamic person support system of claim 10, wh... | Hill-Rom Services, Inc. | A61G007/015 A61G007/018 A61G007/008 A61B005/00 A61G007/00 A61M021/02 A61G007/057
US09332908 | B1 | 2016-05-10 | Hybrid operating room | 2014-12-29 | A hybrid operating room with support structure. The hybrid operating room includes an interior for use in the imaging of and performance of surgery on a patient, the interior including a ceiling. The hybrid operating room also includes a support assembly positioned above the unit for supporting equipment used in conjunction with such imaging and surgery so that the equipment extends downwardly from the ceiling into the interior. The support assembly is designed to support a load of at least 3,500 pounds, and more in certain implementations. | 1. A hybrid operating room for performing imaging and surgery, comprising: a structure enclosing an interior having a first size that is large enough to function as a hybrid operating room, said structure including a central section, a first side section communicating with said central section and a second side section communicating with said central section opposite said first side section, said central section, first side section and second side section together defining said interior; anda support assembly for supporting equipment suspended within said interior of said structure, said support assembly comprising a plurality of support columns and a plurality of support members positioned above said structure and supported by said support columns, said plurality of support members having a weight and said support columns extending between said plurality of support members and a surface so that the weight of said plurality of support members and any equipment attached thereto is transmitted to the surface. 2. A hybrid operating room according to claim 1, further including a plurality of attachment assemblies for connecting equipment used in a hybrid operating room with said support assembly so that said equipment may be suspended in said interior. 3. A hybrid operating room according to claim 1, wherein said central section is sized and configured to be transported as a single section on roadways. 4. A hybrid operating room according to claim 1, wherein said central section, said first side section and said second side section define contiguous sections of a hybrid operating room and said surface comprises a plurality of surfaces in a hospital building. 5. A hybrid operating room according to claim 1, wherein said central section, said first side section and said second side section define contiguous sections of a hybrid operating room on the floor of a hospital building above which there are no additional floors and said surface comprises a plurality of surfaces in the hospital building. 6. A hybrid operating room according to claim 1, wherein said support assembly is capable of supporting a load of at least 25,000 pounds. 7. A hybrid operating room according to claim 1, wherein said support assembly is capable of supporting a load of at least 3,500 pounds. 8. A hybrid operating room according to claim 1, further including at least one attachment assembly for connecting equipment used in a hybrid operating room with said support assembly such that said equipment may be suspended within said interior. 9. A hybrid operating room, comprising: an interior in which imaging and surgery may be performed, said interior having a ceiling, wherein no floors are positioned above said ceiling;a plurality of support columns; anda support structure for use in suspending equipment used in imaging, said support structure comprising a plurality of support members arranged to be supported by said plurality of support columns, wherein said support structur... | AMOHS, INC. | A61G003/00 A61B005/00 A61B019/00 A61B006/00 E04B001/343 E04B001/18 E04H003/08
US09345421 | B2 | 2016-05-24 | Magnetic resonance spectroscopy pulse sequence, acquisition, and processing system and method | 2015-02-19 | Systems and methods are provided for processing a set of multiple serially acquired magnetic resonance spectroscopy (MRS) free induction decay (FID) frames from a multi-frame MRS acquisition series from a region of interest (ROI) in a subject, and for providing a post-processed MRS spectrum. Processing parameters are dynamically varied while measuring results to determine the optimal post-processed results. Spectral regions opposite water from chemical regions of interest are evaluated and used in at least one processing operation. Frequency shift error is estimated via spectral correlation between free induction decay (FID) frames and a reference spectrum. Multiple groups of FID frames within the acquired set are identified to different phases corresponding with a phase step cycle of the acquisition. Baseline correction is also performed via rank order filter (ROF) estimate and a polynomial fit. Sections of the ROF may be excluded from the polynomial fit, such as for example sections determined to be associated with relevant spectral peaks. | 1. A method for processing a set of multiple serially acquired magnetic resonance spectroscopy (MRS) free induction decay (FID) frames from a multi-frame MRS acquisition series from a region of interest (ROI) in a subject, and for providing a post-processed MRS spectrum, comprising at least one of the following: determining a frequency shift error of an FID frame by performing spectral cross-correlation between the FID frame and a reference frame;adjusting a phase of an FID frame by a phase shift configured to reduce an extent of the post-processed MRS spectrum having a negative value;estimating a line width between two opposite walls of a water peak region of an FID frame of the series, comparing the estimated line width against a threshold value, and qualifying and including, or disqualifying and excluding, the FID frame from the set for further processing to provide the post-processed MRS spectrum based at least in part upon the comparison;associating first and second groups of FID frames within the set with first and second different relative phases, respectively, of a phase step cycle along the MRS acquisition series, processing the first and second groups separately, averaging the first and second groups separately to form first and second average interim spectra, and averaging the first and second average interim spectra; andgenerating a first baseline estimate using a rank order filter (ROF) on a range of interest of an interim MRS spectrum partially post-processed from the set, fitting a polynomial to at least a portion of the first baseline estimate to generate a baseline curve, and subtracting the baseline curve from the interim MRS spectrum to provide the post-processed MRS spectrum. 2. The method of claim 1, wherein the method comprises: determining a frequency shift error of an FID frame by performing spectral cross-correlation between the FID frame and a reference frame. 3. The method of claim 2, further comprising: shifting the FID frame by the frequency shift error so as to increase coherence of the FID frame with the reference spectrum. 4. The method of claim 2, further comprising: forming the reference spectrum by averaging multiple FID frames from the acquisition series. 5. The method of claim 2, further comprising: performing an initial frequency shift error correction operation by estimating a frequency shift error for each FID frame of the series by determining a location of a water peak by locating maximum peak value in a range around an expected location for the water peak, calculating a difference between the determined location and the expected location, and adjusting the FID frame by the difference; andaveraging the respectively shifted FID frames to form the reference spectrum. 6. The method of claim 2, wherein the determining is performed on each FID frame of the series. 7. The method of claim 6, wherein: the series comprises first and second groups of FID frames;the FID frames of the first group comprise a fir... | Nocimed, Inc. | G06K009/72 A61B005/055 G01R033/46 A61B005/00 G01R033/48 G01R033/56 G06T007/00 G01R033/483
US10092198 | B2 | 2018-10-09 | Medical pendant system | 2014-06-10 | A medical pendant system, comprising: a movable component, an execution component and a bus (1); the movable component comprises a key control module (3); the key control module (3) comprises a braking system control key (K1, K4) and a pendant lifting/lowering control key (K2, K3); the execution component comprises an output control part (4), the output control part (4) being connected respectively with a bus power supply (5), a pendant motor (9) and a braking system (10); the movable component and the execution component are connected respectively to the bus (1); the bus (1) is used for transmitting power supply and also for transmitting instructions from the movable component to the execution component. In the vertical direction the medical pendant system can be adjusted steplessly without having to alter the lengths of the cables. The number of cable cores is reduced, circuit design is optimized, installation and maintenance are easy, functional integration and extension are convenient, and reliability is improved. | 1. A medical pendant system, characterized in that the medical pendant system comprises a movable component, an execution component and a bus; the movable component comprises a key control module; the key control module comprises a braking system control key and a pendant lifting/lowering control key; the execution component comprises an output control part; the output control part is connected respectively with a bus power supply, a pendant motor, and a braking system; the movable component and the execution component are connected respectively to the bus; the bus is used for transmitting power supply, and also used for transmitting instructions from the movable component to the execution component; wherein the bus is disposed in an inner groove inside the pendant cabinet or on the exterior surface thereof and extends along the lengthwise direction of the pendant cabinet; andwherein the bus comprises only two lines, each of the two lines comprising an insulator accommodated inside the inner groove and a conductor enclosed inside the insulator. 2. The medical pendant system of claim 1, wherein the medical pendant system further comprises a relay module located between the movable component and the bus, and a relay module located between the execution component and the bus. 3. The medical pendant cabinet of claim 1, wherein the inner groove is substantially filled when the insulator is accommodated inside the inner groove and the conductor is enclosed inside the insulator. 4. The medical pendant system of claim 1, wherein the cross section of the conductor is substantially Y-shaped, Ω-shaped, V-shaped, W-shaped, C-shaped, or T-shaped. 5. The medical pendant system of claim 1, wherein the output control part is further connected with a pendant protecting sensor. 6. The medical pendant system of claim 1, wherein the output control part is further connected with a height limiting unit for limiting the lifting and lowering range of the pendant. 7. The medical pendant system of claim 1, wherein the output control part is further connected with a pendant motor power supply. 8. The medical pendant system of claim 1, wherein the movable component further comprises an environmental lamp module. 9. The medical pendant system of claim 1, wherein the movable component further comprises a display module, and the execution component further comprises a display data line module. 10. The medical pendant system of claim 1, wherein the movable component further comprises a pulse sensor module, and the execution component further comprises a pulse diagnostic instrument module. 11. The medical pendant system of claim 1, wherein the key control module further comprises an input conversion interface. | Maquet (Suzhou) Co. Ltd. | A61B005/024 E04F017/08 A61B005/00 H01H047/00 A61G012/00
US09351711 | B2 | 2016-05-31 | Filter for fine needle biopsy | 2014-11-20 | A biopsy needle comprises a first lumen open to a tissue penetrating distal tip of the needle for receiving a tissue sample therein and a second lumen in combination with a filter separating the first and second lumens from one another, the filter being configured to permit fluids to pass therethrough while preventing the sampled tissue from passing therethrough. | 1. A method for performing a needle biopsy, comprising: advancing a needle into a target portion of tissue to capture a sample of the target tissue within a first lumen of a longitudinal body of the needle, the needle including a stylet slidably received within the first lumen, the stylet including a second lumen extending therethrough, the stylet including a distal tip at a distal end thereof, the distal tip of the stylet including a filter separating the first lumen from the second lumen; andreducing a pressure within the second lumen of the stylet to pass fluids through the filter from the first lumen to the second lumen, the filter preventing the sampled tissue from passing therethrough. 2. The method according to claim 1, wherein a cross-sectional shape of the first and second lumens is symmetrical. 3. The method according to claim 2, wherein the first lumen surrounds the second lumen. 4. The method according to claim 1, wherein the filter comprises a filter material disposed in a distal end of the second lumen. 5. The method according to claim 2, wherein, when the stylet is moved proximally relative to the longitudinal body to a second position, a distal portion of the first lumen is opened to receive the sampled tissue therein. 6. The method according to claim 5, wherein a proximal end of the second lumen includes a proximal port coupleable to a source of negative pressure to draw tissue into the distal portion of the first lumen toward the distal tip of the stylet and to aspirate fluids in the distal portion of the first lumen into the second lumen through the filter. 7. The method according to claim 6, wherein the source of negative pressure is syringe. 8. The method according to claim 1, wherein the filter permits only particulate matter of less than a predetermined diameter to pass therethough. 9. The method according to claim 8, wherein the filter permits red blood cells to pass therethrough. 10. The method according to claim 4, wherein the filter material is a membrane. 11. The method according to claim 4, wherein the filter material is a plug. | BOSTON SCIENTIFIC SCIMED, INC. | A61B005/00 A61B010/02
US09351671 | B2 | 2016-05-31 | Multiplexed pathlength resolved noninvasive analyzer apparatus and method of use thereof | 2014-09-22 | A noninvasive analyzer apparatus and method of use thereof is described using one or a plurality of sample illumination zones coupled to at least one two-dimensional detector array monitoring a plurality of detection zones. Control of illumination times and/or patterns along with selected detection zones yields pathlength resolved groups of spectra. Sectioned pixels and/or zones of the detector are optionally filtered for different light throughput as a function of wavelength. The pathlength resolved groups of spectra are subsequently analyzed to determine an analyte property. Optionally, in the mapping and/or collection phase, incident light is controllably varied in time in terms of any of: sample probe position, incident light solid angle, incident light angle, depth of focus, energy, intensity, and/or detection angle. Optionally, one or more physiological property and/or model property related to a physiological property is used in the analyte property determination. | 1. An apparatus for noninvasive estimation of an analyte concentration of a human subject, comprising: a near-infrared noninvasive analyzer, comprising: a near-infrared source;a photon transport system configured to transport photons from said near-infrared source to an illumination zone;a two-dimensional planar transmittance filter array comprising: a first filter comprising a first fifty percent cut-on transmittance inflection point at a first wavelength in a range of 1200 to 2500 nanometers;a second filter comprising a second fifty-percent cut-on transmittance inflection at a second wavelength, said second wavelength at least one hundred nanometers shorter than said first wavelength,said first filter positioned closer to the illumination zone than said second filter; anda detector, said detector configured to detect a sub-set of the photons after the photons undergo diffuse reflection and arrive at a detection zone, both the illumination zone and the detection zone in proximate contact with the subject during use,said two-dimensional planar transmittance filter orientated substantially parallel to a line from the illumination zone to the detection zone. 2. The apparatus of claim 1, said two-dimensional planar transmittance filter array further comprising: a third filter comprising a third cut-on transmittance inflection point at a third wavelength in a range of 1100 to 1900 nanometers, said third wavelength at least one hundred nanometers shorter than said second wavelength,said second filter positioned closer to the illumination zone than said third filter. 3. The apparatus of claim 1, said analyzer further comprising: at least one two-dimensional near-infrared detector array configured to detect diffusely reflected photons from a detection zone proximate the subject interface zone, said two-dimensional near-infrared detector array within ten centimeters of the illumination zone during use; anda processor configured to convert signals from said two-dimensional near-infrared detector array into a vibrational spectroscopy reading. 4. The apparatus of claim 3, said two-dimensional near-infrared detector array further comprising: a two-dimensional array of detector wells, wherein a first photon facing element of said array of detector wells comprises a first total surface area, wherein a second photon facing element of said array of detector wells comprises a second total surface area, said second total surface area at least fifty percent larger than said first total surface area. 5. The apparatus of claim 3, said two-dimensional near-infrared detector array comprising: a first detector positioned at a first radial distance from a center of the illumination zone, said first detector comprising a first surface area; anda second detector positioned at a second distance from the center of the illumination zone, said second detector comprising a second surface area,said second distance greater than said first distance, andsaid second surface area... |  | A61B005/00 A61B005/1455 A61B005/145 A61M005/172
US09334520 | B2 | 2016-05-10 | Identification of mycoplasm contamination in biotechnology production using Raman spectroscopy | 2014-02-27 | Mycoplasma contamination of a known cell line is detected by collecting a Raman spectrum of a targeted volume within a sample, the targeted volume containing a known cell line of interest, obtaining a reference spectrum uniquely associated with the known cell line where the obtained reference spectrum is known to be free of mycoplasma and comparing, using a processing device, the reference spectrum to the collected spectrum. Mycoplasma is further detected by identifying whether there are unnatural molecular compositions within the collected spectrum based upon the comparison of the reference spectrum to the collected spectrum and providing an indication as to whether mycoplasma is detected in the collected Raman spectrum based upon whether unnatural molecular compositions are identified within the collected spectrum. | 1. A method of detecting mycoplasma in a bioprocess manufacturing application that uses recombinant technology to modify materials within bacteria comprising: collecting a sample from a cell culture used by the bioprocess manufacturing application to catalyze a biochemical reaction within a microorganism;controlling a Raman spectrometer, by a processor, to collect a Raman spectrum of a targeted volume within the sample so as to collect a Raman spectrum of a single cell of a known cell line of interest;obtaining a reference spectra uniquely associated with the known cell line, the reference spectra comprising at least two of: spectral measurements of mycoplasma by itself, a contaminated cell line, and a pure cell line;comparing, using a processing device, the reference spectra to the collected spectrum;identifying whether there are unnatural molecular compositions within the collected spectrum based upon the comparison of the reference spectra to the collected spectrum;providing an indication as to whether mycoplasma is detected in the collected Raman spectrum based upon whether unnatural molecular compositions are identified within the collected spectrum; andstopping the bioprocess manufacturing application when mycoplasma is detected in the collected Raman spectrum. 2. The method according to claim 1, further comprising: identifying whether the targeted volume contains a cell from a select one of a Chinese hamster ovarian line and Escherichia coli line. 3. The method according to claim 1, wherein comparing the reference spectra to the collected spectrum comprises computing by the processing device, a difference spectrum as the difference between the reference spectra and the collected spectrum. 4. The method according to claim 3, wherein identifying whether there are unnatural molecular compositions within the collected spectrum based upon the comparison of the reference spectra to the collected spectrum comprises identifying unnatural molecules based upon an analysis of the difference spectrum. 5. The method according to claim 1, wherein the targeted volume comprises a single cell. 6. The method according to claim 5, wherein the collected Raman spectrum contains sufficient spectral content to examine at least substantially the entirety of the contents of the cell within the targeted volume. 7. The method according to claim 1, wherein the collected Raman spectrum is a superposition of all the molecular material of a single cell being illuminated by a laser used to collect the Raman spectrum. 8. The method according to claim 1, further comprising: scanning the sample of interest to identify at least one targeted volume as a potential host for mycoplasma where the targeted volume is identified as a belonging to a known line of interest; andevaluating at least one targeted volume determined to be a potential host for mycoplasma by: collecting a Raman spectrum of the targeted volume within a sample under evaluation, the targeted volume contai... | BATTELLE MEMORIAL INSTITUTE | G01J003/44 C12Q001/04 C12M003/00 A61B005/00 G01N033/48 G01N021/65 G01N033/569 G01N021/85
US09582916 | B2 | 2017-02-28 | Method and system for unsupervised cross-modal medical image synthesis | 2015-09-30 | A method and apparatus for unsupervised cross-modal medical image synthesis is disclosed, which synthesizes a target modality medical image based on a source modality medical image without the need for paired source and target modality training data. A source modality medical image is received. Multiple candidate target modality intensity values are generated for each of a plurality of voxels of a target modality medical image based on corresponding voxels in the source modality medical image. A synthesized target modality medical image is generated by selecting, jointly for all of the plurality of voxels in the target modality medical image, intensity values from the multiple candidate target modality intensity values generated for each of the plurality of voxels. The synthesized target modality medical image can be refined using coupled sparse representation. | 1. A method for unsupervised cross-modal synthesis of a target modality medical image from a source modality medical image, comprising: receiving the source modality medical image;generating multiple candidate target modality intensity values for each of a plurality of voxels of the target modality medical image based on corresponding voxels in the source modality medical image; andgenerating a synthesized target modality medical image by selecting, jointly for all of the plurality of voxels in the target modality medical image, intensity values from the multiple candidate target modality intensity values generated for each of the plurality of voxels. 2. The method of claim 1, wherein generating multiple candidate target modality intensity values for each of a plurality of voxels of the target modality medical image based on corresponding voxels in the source modality medical image comprises: generating the multiple candidate target modality intensity values for each of the plurality of voxels independently using a cross-modal nearest neighbor search between a corresponding portion of the source modality medical image and a plurality of target modality training images. 3. The method of claim 1, wherein generating multiple candidate target modality intensity values for each of a plurality of voxels of the target modality medical image based on corresponding voxels in the source modality medical image comprises: extracting from the source modality medical image a plurality of image patches, each centered at a respective one of a plurality of voxels of the source modality medical image;obtaining, for each of the plurality of image patches extracted from the source modality medical image, K nearest image patches from a plurality of target modality training images using a cross-modal nearest neighbor search based on a cross-modal similarity measure between the image patches extracted from the source modality medical image and image patches of the target modality training images; andgenerating, for each of the plurality of image patches extracted from the source image, K candidate target modality intensity values for a corresponding set of voxels in the target modality medical image using intensity values of a corresponding set of voxels in each of the K nearest image patches from the plurality of target modality training images, wherein the corresponding set of voxels in the target modality medical image includes a voxel corresponding to the voxel at which the image patch extracted from the source image is centered and neighboring voxels and the corresponding set of voxels in each of the K nearest image patches includes a center voxel and neighboring voxels in each of the K nearest image patches. 4. The method of claim 3, wherein the cross-modal similarity measure between the image patches extracted from the source modality medical image and the image patches of the target modality training images is a voxel intensity-based mutual information measu... | SIEMENS HEALTHCARE GMBH | G06K009/00 G06T011/60 G06T007/00 A61B005/055 A61B005/00 G01R033/50 G06K009/62 G01R033/56
US10524681 | B2 | 2020-01-07 | Lead and a system for medical applications | 2015-11-24 | In one example, a medical device system for at least one of delivery of electrical stimulation pulses or sensing of physiological signals, the system including an elongated carrier; and a thin film wound around the elongated carrier, wherein the thin film includes a plurality of electrical contacts, a plurality of electrodes located distal to the plurality of electrical contacts, and a plurality of conducting tracks, each of the plurality of conducting tracks providing an electrical connection between at least one of the plurality of electrodes and one of the plurality of electrical contacts, and wherein the thin film is wound around the elongated carrier from the distal end of the elongated film to the proximal end of the elongated thin film. | 1. A medical device system for at least one of delivery of electrical stimulation pulses or sensing of physiological signals, the system comprising: an elongated carrier;a thin film helically wound around the elongated carrier, wherein the thin film includes a first plurality of electrical contacts, a plurality of electrodes located distal to the first plurality of electrical contacts, and a plurality of conducting tracks, each of the plurality of conducting tracks providing an electrical connection between at least one of the plurality of electrodes and one of the first plurality of electrical contacts, wherein the thin film is wound around the elongated carrier from a distal end of the thin film to a proximal end of the thin film; andan interposer including a second plurality of electrical contacts, wherein each contact of the second plurality of electrical contacts is electrically coupled to at least one of the first plurality of electrical contacts of the thin film, andwherein a distal portion of the interposer including the second plurality of electrical contacts is helically wound around a proximal portion of the elongated carrier at least one of over the first plurality of electrical contacts of the thin film or under the first plurality of electrical contacts of the thin film such that the second plurality of electrical contacts of the interposer align with the first plurality of electrical contacts of the thin film. 2. The system of claim 1, wherein no part of the thin film extends or protrudes from the elongated carrier at both a proximal end and a distal end of the elongated carrier. 3. The system of claim 1, wherein the thin film has a substantially consistent width over its length from the distal end to the proximal end of the thin film. 4. The system of claim 1, wherein the first plurality of electrical contacts are located at the proximal end of the thin film. 5. The system of claim 1, wherein the interposer comprises a third plurality of electrical contacts located proximal the second plurality of electrical contacts, wherein the third plurality of electrical contacts are configured to electrically couple the second plurality of electrical contacts to an electrical connector of an active lead can. 6. The system of claim 5, further comprising the active lead can including a switch matrix, and wherein the active lead can includes a substantially sealed housing containing the switch matrix. 7. The system of claim 6, further comprising an implantable medical device coupled to the active lead can via a lead. 8. The system of claim 7, wherein the active lead can includes a pulse generator. 9. The system of claim 7, wherein the implantable medical device includes a pulse generator. 10. The system of claim 7, wherein N electrical connection tracks are defined between the implantable medical device and the active lead can, and wherein N is less than a total number of individual electrodes of the plurality of electrodes. 11. The s... | Medtronic Bakken Research Center B.V. | A61N001/05 A61B005/0478 A61B005/00
US09486153 | B2 | 2016-11-08 | Medical device system having an implanted medical device and an external device | 2016-02-01 | Disclosed is a system having an implanted component and external component which are configured to provide a test of wireless communication in order to assess the success or failure of such communication and to store attributes related to such test in a memory log. To provide the communication test the implantable and external components can attempt wireless communication according to communication test parameters which relate to number of times to retry communication, duration of sending communication test signals, durations of waiting for communication test signals and the schedule of the communication tests. The schedule of tests may be period or may change over time in order to become more or less frequency according to a programmable schedule that may also decrease if the communication tests are successful and indicate patient compliance in keeping the external components close by. The communication tests can assist in determining if the patient is maintaining external components within a suggested proximity (e.g. 6 feet) of the patient and may assist to determine if transmission or reception difficulties are the source of communication failure. A physician programmer can provide for programming, conducting, summarizing and assessing the results of communication tests. | 1. A medical device system, the system comprising: an implanted medical device, andan external device comprising a transceiver, a random access memory, and a processor coupled to the transceiver and the random access memory, wherein the processor is configured to:periodically activate wireless communication circuitry in an attempt to communicate a plurality of proximity test signals with the implanted medical device;for each of the plurality of activations, determine whether a proximity signal has been received from the implanted device within a certain time after the activation, thereby generating a plurality of proximity test receipt results;apply a test to the plurality of proximity test receipt results to determine whether a proximity failure condition exists, thereby generating a plurality of proximity test results;operative when a proximity failure condition exists, take at least one of the following actions: store information regarding the proximity failure in the random access memory or send a proximity test failure condition notification; andwhen the proximity failures in the random access memory meet a selected first proximity test criterion, the processor is further configured to send notification to the patient. 2. The medical device system according to claim 1 wherein the processor is further configured to send notification to a remote station when the proximity failures in the random access memory meet a selected second proximity test criterion. 3. The medical device system according to claim 1 wherein the periodic activation occurs every 24 hours. 4. The medical device system of claim 1 further including a proximity test module which operates upon the random access memory to provide a log of the plurality of proximity test results including a time of each test and a quantitative value related to a strength of each of the plurality of the proximity test signals. 5. The medical device system of claim 1 further including a proximity test module which operates upon the random access memory to provide a log of the plurality of proximity test results including a time of each test and a qualitative value related to a quality of each of the plurality of the proximity test signals. 6. The medical device system of claim 1 wherein the implantable medical device is configured to provide a proximity test in conjunction with the external device and to derive and store a second plurality of proximity test results. 7. The medical device system of claim 6 further including a programmer device configured for downloading the second plurality of proximity test results from the implanted device and processing these test results to provide a test result profile. 8. The medical device system of claim 6 further including a programmer device configured for downloading the second plurality of proximity test results from the implanted device and presenting the test results graphically to a user. 9. The medical device system of claim 6 further includi... | Angel Medical Systems, Inc. | A61N001/00 A61B005/0432 A61B005/00 A61B005/055 A61N001/372 G06F019/00
US10966628 | B2 | 2021-04-06 | Ultrasound diagnosis apparatus | 2015-10-23 | According to one embodiment, an ultrasound diagnosis apparatus is configured to generate ultrasound image data of a subject through an ultrasound probe. The ultrasound diagnosis apparatus includes a specifying unit, a transmission/reception condition change unit, and an image generator. The specifying unit specifies a characteristic site, which is included in a scanning range related to the ultrasound image data and has specific acoustic characteristics, based on three-dimensional image data of the subject generated in advance. The transmission/reception condition change unit changes transmission/reception conditions of ultrasound waves based on the acoustic characteristics of the characteristic site. The image generator generates the ultrasound image data based on changed transmission/reception conditions. | 1. An ultrasound diagnosis apparatus configured to generate ultrasound image data of a subject through an ultrasound probe, the ultrasound diagnosis apparatus comprising processing circuitry configured to: specify a characteristic site, which is included in a scanning range related to the ultrasound image data and has specific acoustic characteristics, based on three-dimensional image data of the subject generated in advance;obtain a weighting coefficient based on an attenuation rate of the characteristic site;obtain a delay time based on the acoustic characteristics of the characteristic site;change reception conditions of ultrasound waves by multiplying each of reception signals, which has been delayed by the delay time, by the weighting coefficient with respect to each channel, and adding up the reception signals to obtain a reception signal that has been subjected to reception beam forming; andgenerate the ultrasound image data based on the reception signal that has been subjected to reception beam forming under changed reception conditions. 2. The ultrasound diagnosis apparatus of claim 1, further comprising a detector configured to detect either or both position and angle of the ultrasound probe, wherein the processing circuitry is further configured to: specify a new characteristic site based on the position and the angle of the ultrasound probe when the ultrasound probe has moved; andchange the reception conditions based on acoustic characteristics of the new characteristic site. 3. The ultrasound diagnosis apparatus of claim 1, wherein the processing circuitry is further configured to: determine whether to newly change the reception conditions based on contrast of the ultrasound image data; andchange the reception conditions when reception conditions are determined to be changed. 4. The ultrasound diagnosis apparatus of claim 1, further comprising an operation unit configured to receive an input to newly change the reception conditions, wherein the processing circuitry is further configured to change the reception conditions based on the input. 5. The ultrasound diagnosis apparatus of claim 1, wherein the ultrasound probe is a two-dimensional array probe, andthe processing circuitry is further configured to change the reception conditions related to azimuth direction and elevation direction of the two-dimensional array probe. 6. The ultrasound diagnosis apparatus of claim 1, wherein the three-dimensional image data is generated by at least one of an X-ray CT apparatus, an MRI apparatus, an ultrasound diagnosis apparatus, and an X-ray diagnosis apparatus. 7. The ultrasound diagnosis apparatus of claim 1, wherein the characteristic site is a site including bone. 8. The ultrasound diagnosis apparatus of claim 7, wherein the site including bone is skull. 9. The ultrasound diagnosis apparatus of claim 1, wherein the processing circuitry is further configured to specify reference image data in the three-dimensional image data of the su... | Canon Medical Systems Corporation | A61B005/055 A61B008/00 A61B008/08 A61B006/03 A61B008/14
US10123725 | B2 | 2018-11-13 | Hearing prosthesis efficacy altering and/or forecasting techniques | 2015-11-13 | A method of forecasting subjective efficacy of a hearing prosthesis implanted in a recipient, including subjecting a recipient to a plurality of temporally spaced aural training tasks, the aural training tasks evoking hearing percepts with the hearing prosthesis, obtaining first data indicative of the recipient's ability to hear with the hearing prosthesis based on performance by the recipient of the temporally spaced aural training tasks, and forecasting second data based on the first data, wherein the second data are indicative of the recipient's ability to hear with the hearing prosthesis at one or more temporal locations in the future. | 1. A method of forecasting subjective efficacy of a hearing prosthesis implanted in a recipient, comprising: subjecting a recipient to a plurality of temporally spaced aural tests, the aural tests evoking hearing percepts with the hearing prosthesis;obtaining first data indicative of the recipient's ability to hear with the hearing prosthesis based on performance by the recipient of the temporally spaced aural tests; andforecasting second data based on the first data, wherein the second data are indicative of the recipient's ability to hear with the hearing prosthesis at one or more temporal locations in the future. 2. The method of claim 1, wherein: the first data are based on measurements of the ability of the recipient to hear with the hearing prosthesis taken in effective immediacy at least one of before or after respective temporally spaced aural training tasks. 3. The method of claim 1, wherein: the first data are based on measurements of the ability of the recipient to hear with the hearing prosthesis taken at least one of before or after at least one of the aural training tasks without any cognitive stressors being located between the measurements and the respective aural training tasks. 4. The method of claim 1, wherein: the first data are based on measurements of the ability of the recipient to hear with the hearing prosthesis taken before and after, both in close temporal proximity to, respective temporally spaced aural training tasks. 5. The method of claim 1, further comprising forecasting third data, wherein the third data are data indicating an estimated ability to hear with the prosthesis if temporally spaced aural training tasks are executed in the future with different temporal spacing between the tasks relative to that which was the case for temporally spaced aural training tasks temporally linked to respective temporally spaced aural tests of the plurality of temporally spaced aural tests. 6. The method of claim 5, further comprising: comparing second data with the third data; andindicating to the recipient the effects of the different temporal spacing between the tasks relative to that which would be the case if temporally spaced aural training tasks are executed in the future with the same temporal spacing between the tasks relative to that which was the case for the temporally spaced aural training tasks temporally linked to the respective temporally spaced aural tests of the plurality of temporally spaced aural tests. 7. A method of altering a recipient's ability to hear with a hearing prosthesis implanted in a recipient, comprising: subjecting a recipient who has participated in a plurality of respective temporally spaced aural training tasks to a plurality of respective temporally spaced aural tests, the respective aural tests evoking hearing percepts evoked by the hearing prosthesis;influencing the recipient's commitment to performing subsequent aural training tasks based on first data relating to the aural tests;... | Cochlear Limited | A61B005/12 H04R025/00 G09B021/00
US09330202 | B2 | 2016-05-03 | Exercise detection apparatus and control method for exercise detection apparatus | 2012-12-11 | An exercise detection apparatus determines body movements based on data detected by a detection unit, judges movement states for each body movement based on data detected by the detection unit, corrects the movement states based on a predetermined rule, calculates an exercise state of a user based on a result of the judgment, and stores the calculated exercise state in a storage unit. Accordingly, it is possible to provide an exercise detection apparatus capable of reducing erroneous judgments of the exercise states of the user. | 1. An exercise detection apparatus that is configured to be worn or carried by a user, the exercise detection apparatus comprising: a main body unit, the main body unit including a display unit and a display switch;a storage unit;a detection unit; anda control unit configured to: judge whether predetermined movements of a body of the user are occurring based on data detected by the detection unit;judge a movement state for the body across a predetermined unit of body movements based on data detected by the detection unit;calculate an exercise state of the user during the predetermined unit of body movements based on the respective judgment results;store the exercise state calculated by the control unit in the storage unit; andcorrect a target exercise state based on a predetermined rule, wherein the predetermined rule is a rule for correcting the target exercise state based on a transition of the exercise state;the transition occurs between the target exercise state and exercise states before and after the target exercise state; andthe display unit displays the corrected target exercise state. 2. The exercise detection apparatus according to claim 1, wherein the predetermined rule is a rule that, for transitions of the exercise state between the target exercise state and the exercise states before and after the target exercise state, corrects an abnormal transition to a non-abnormal transition. 3. The exercise detection apparatus according to claim 2, wherein for the exercise states calculated by the control unit for each unit of body movements in a range of a predetermined number of units of body movements from before to after the unit of body movements that is the target of the exercise state calculation,the abnormal transition is a transition in which the exercise states of the first and last units of body movements in the range are the same, and exercise states that differ from the exercise states of the first and last units of body movements in the range make up less than half of the exercise states of units of body movements in the range, andthe non-abnormal transition is a transition in which the exercise states of the first and last units of body movements in the range are the same and the exercise states of the other units of body movements in the range are the same. 4. The exercise detection apparatus according to claim 1, wherein the exercise state is a function of both the identity of the movements of the body and the movement state. 5. The exercise detection apparatus according to claim 1, wherein the detection unit detects an acceleration value in at least one axial direction; andthe control unit judges whether predetermined movements of the body are occurring based on the acceleration value detected by the detection unit. 6. The exercise detection apparatus according to claim 1, wherein the detection unit detects absolute pressure values; andthe control unit judges the movement state based on the absolute pressure values detec... | OMRON HEALTHCAE CO., LTD. OMRON CORPORATION | G01B003/02 G06F017/40 A61B005/11 A61B005/22 A61B005/00 G01C022/00 A61B005/02
US09333365 | B2 | 2016-05-10 | Antenna for an implantable medical device | 2010-07-30 | This disclosure describes antenna structures for use in an implantable medical device. The implantable medical device may include a housing that hermetically encloses electronic components of the implantable medical device and a fixation mechanism that affixes the implantable medical device to a target location, such as a wall of a vessel. The fixation mechanism functions as a radiating element of an antenna of the implantable medical device. The housing of the implantable medical device may include a conductive loop that electrically couples to a telemetry module and magnetically couples to the fixation mechanism. The telemetry module may provide signals to be transmitted to the inner loop and those signals are magnetically coupled between the inner loop and the fixation mechanism, which radiates the signals. | 1. An apparatus comprising: a housing that includes: a telemetry module; andan inner conductive feed loop that is electrically coupled to the telemetry module, wherein at least a portion of the inner conductive feed loop forms a portion of an antenna; anda fixation mechanism located outside the housing and mechanically coupled to the housing, and configured to affix the apparatus to a target location, wherein at least a portion of the fixation mechanism forms a radiating portion of the antenna, the fixation mechanism being magnetically coupled through the housing to the conductive feed loop to obtain the communication signals to be transmitted by the telemetry module and radiate the signals from the telemetry module to another device. 2. The apparatus of claim 1, wherein the fixation mechanism is a cylindrical stent-like structure that is configured to lodge against a vessel wall. 3. The apparatus of claim 1, wherein the fixation mechanism includes a plurality of struts arranged to form the fixation mechanism, wherein a first portion of the plurality of struts form a ring and a second portion of the plurality of struts mechanically couple the ring to the housing. 4. The apparatus of claim 3, wherein the first portion of the plurality of struts form a ring having one of a zig-zag shape and a sinusoidal shape. 5. The apparatus of claim 1, wherein the fixation mechanism includes a plurality of struts arranged to form the fixation mechanism, wherein a first portion of the plurality of struts form a plurality of rings that are joined in series to form the cylindrical body and a second portion of the plurality of struts mechanically couple the cylindrical body to the housing. 6. The apparatus of claim 1, wherein the fixation mechanism includes at least one fixation loop mechanically coupled to the housing. 7. The apparatus of claim 6, wherein at least a portion of the conductive loop within the housing is located within the circumference of the fixation loop. 8. The apparatus of claim 1, wherein a shape of the fixation mechanism defines a lumen and at least a portion of the conductive loop is located within the lumen defined by the fixation mechanism. 9. The apparatus of claim 8, wherein the entire housing is located within the lumen defined by the fixation mechanism. 10. The apparatus of claim 1, further comprising a sensor within the housing to measure one or more parameters. 11. The apparatus of claim 1, wherein the sensor comprises one of a pressure sensor, a pH sensor, an oxygen sensor, a temperature sensor, and an electrode. 12. The apparatus of claim 1, wherein the housing further includes: a non-conductive header portion that includes the conductive loop;a feed through via which the conductive loop is electrically coupled to the telemetry module. 13. An apparatus comprising: a housing that includes: a telemetry module; andan inner conductive feed loop that is electrically coupled to the telemetry module, wherein at least a portion o... | Medtronic, Inc. | A61N001/08 A61N001/372 A61B005/00 A61B005/0215 A61B005/11
US09333333 | B2 | 2016-05-10 | Offset electrode | 2012-02-22 | An electrode including a non-conductive substrate having a top surface and at least one channel extending therethrough, an electrically conductive trace material positioned adjacent a portion of the top surface of the non-conductive substrate and extending through the channel, and adapted for electrically coupling to a power source, and a second electrically conductive material that is inert or more corrosion resistant than the trace material. The second material is positioned adjacent to and entirely covering a top surface of the trace material. The electrode further includes a conductive hydrogel laterally offset from the trace material, the hydrogel may be positioned adjacent to a portion of a top surface of the second electrically conductive material. | 1. An electrode comprising: a non-conductive substrate having a top surface and a channel extending therethrough; the top surface of the non-conductive substrate having a recessed portion defined therein;an electrically conductive trace material positioned adjacent at least a portion of the top surface of the non-conductive substrate entirely within said recessed portion and said channel; and the electrically conductive trace material extending through said channel, and being adapted to be electrically coupled to a power source; anda second electrically conductive material that is inert or more corrosion resistant than the trace material, the second electrically conductive material being positioned adjacent to and entirely covering a top surface of the trace material; the second electrically conductive material being positioned entirely within said recessed portion and the channel;wherein at least a portion of an outer surface of each of the electrically conductive trace material and the second electrically conductive material is in direct physical contact with the non-conductive substrate. 2. The electrode according to claim 1, wherein the trace material is comprised of copper. 3. The electrode according to claim 1, wherein the second electrically conductive material is comprised of gold or gold/nickel. 4. The electrode according to claim 1, wherein each of the electrically conductive trace material and the second electrically conductive material extend from said recessed portion through said channel to a bottom surface of the non-conductive substrate. 5. The electrode according to claim 1, wherein the electrically conductive trace material encircles the second electrically conductive material in a portion of said channel extending from said recessed portion to a bottom surface of the non-conductive substrate, the electrically conductive trace material does not extend to the top surface of the non-conductive substrate. 6. An electrode comprising; a non-conductive substrate having a top surface and a bottom surface and a channel extending therethrough, the non-conductive substrate being comprised of a flexible material; the top surface of the non-conductive substrate haying a recessed portion defined therein;a copper trace material positioned adjacent at least a portion of the top surface of the non-conductive substrate entirely within said recessed portion and the channel; and the electrically conductive trace material extending through said channel, said trace material being adapted to be electrically coupled to an electrode power source; anda second conductive material comprised of gold or gold/nickel, and positioned adjacent to and covering a top surface of said trace material; and the second electrically conductive material being positioned entirely within said recessed portion and the channel;wherein at least a portion of an outer surface of each of the electrically conductive trace material and the second electrically conductive mater... | ETHICON, INC. | A61N001/00 A61N001/04 A61B005/0408 A61N001/36
US20160136450 | A1 | 2016-05-19 | LASER SYSTEM AND METHOD FOR TREATING DIABETES | 2015-10-28 | A method of treating diabetes including activating a laser system, wherein the laser system emits a composite laser beam with more than one wavelength, and directing the composite laser beam over at least one of a pancreas, a thyroid, a foot, and a thoracic spine to treat diabetes. | 1. A diabetes treatment system, comprising: a glucose sensor configured to detect a glucose level in a patient; anda laser system coupled to the glucose sensor, wherein the laser system isconfigured to treat diabetes by emitting a composite laser beam with more than one wavelength in response to the detected glucose level. 2. The system of claim 1, wherein the composite laser beam comprises a first laser beam with a first wavelength of approximately 532-594 nm. 3. The system of claim 1, wherein the composite laser beam comprises a second laser beam with a second wavelength of approximately 808-980 nm. 4. The system of claim 1, wherein the composite laser beam comprises a third laser beam with a third wavelength of approximately 1064-1300 nm. 5. The system of claim 1, comprising a controller configured to control the laser system in response to a signal from the glucose sensor. 6. The system of claim 5, wherein the controller comprises a processor, and a memory configured to store one or more treatment regimens, and wherein the controller selects a treatment regimen based on the signal from the glucose sensor. 7. The system of claim 1, wherein the laser system comprises one or more laser diodes. 8. The system of claim 1, wherein the diabetes treatment system is configured to couple to a user. 9. A method of treating diabetes, comprising: activating a laser system, wherein the laser system emits a composite laser beam with more than one wavelength; anddirecting the composite laser beam over at least one of a pancreas, a thyroid, a foot, and a thoracic spine to treat diabetes. 10. The method of claim 9, wherein the composite laser beam comprises a first laser beam with a first wavelength of approximately 532-594 nm. 11. The method of claim 9, wherein the composite laser beam comprises a second laser beam with a second wavelength of approximately 808-980 nm. 12. The method of claim 9, wherein the composite user beam comprises a third laser beam with a third wavelength of approximately 1064-1300 nm. 13. The method of claim 9, comprising pulsing the composite laser beam. 14. The method of claim 9, comprising moving the composite laser beam in a pattern over at least one of the pancreas, the thyroid, the foot, and the thoracic spine. 15. The method of claim 9, comprising directing the composite laser beam for a first duration over at least one of the pancreas, the thyroid, the foot, and the thoracic spine. 16. A method of treating diabetes, comprising: receiving a signal from a glucose sensor; andcontrolling a laser system in response to the signal from the glucose sensor, wherein the laser system emits a composite laser beam with a first laser beam, a second laser beam, and a third laser beam, wherein the first laser beam has a first wavelength, the second laser beam has a second wavelength, and the third laser beam has a third wave length, and wherein the first wavelength, the second wavelength, and the third wavelength are different fr... | THERAPEUTIC LASER APPLICATIONS, LLC | A61N005/06 A61B005/00 A61B005/145
US09355217 | B2 | 2016-05-31 | Interactive medical diagnosing with portable consumer devices | 2014-09-10 | Medical tests and examinations are performed with consumer computing devices. The medical tests and examinations are selected and customized in accordance with conditions present in the testing environment, as well as physical characteristics of a user taking the test. The tests can include visual acuity tests, colorblindness tests and other medical tests, such as hearing tests. | 1. A computer program product comprising one or more hardware storage devices having stored thereon computer-executable instructions that are executable by one or more processors of a computing device to cause the computing device to conduct a medical examination, and by at least causing the computing device to implement the following: the computing device accessing a remote computing system, the remote computing system having one or more medical tests provided by one or more medical professionals for a medical patient;the computing device accessing a patient's medical information;the computing device identifying the one or more medical tests to perform based at least in part on the patient's medical information;the computing device conducting the medical examination, including presenting one or more test elements of the one or more medical tests, remotely from the one or more medical professionals;the computing device receiving user input responsive to the one or more test elements; andthe computing device sending one or more results of the one or more medical tests to at least one of the remote computing system or another remote computing system for analysis by the one or more medical professionals. 2. The computer program product of claim 1, further comprising: the computing device identifying one or more contextual conditions for configuring the one or more medical tests. 3. The computer program product of claim 2, wherein presenting the one or more test elements of the one or more medical tests comprises the computing device adjusting at least one test element based on the identified contextual condition. 4. The computer program product of claim 2, wherein the computing device includes a portable consumer device that detects the one or more contextual conditions based on user input, a microphone, a camera, or a proximity sensor. 5. The computer program product of claim 2, wherein the one or more results are sent in real-time in response to receiving the user input. 6. The computer program product of claim 1, wherein the one or more medical tests include a vision test. 7. The computer program product of claim 1, wherein the one or more medical tests include a hearing test. 8. The computer program product of claim 1, wherein the computing device is a wireless mobile computing device. 9. A computing system, comprising: one or more processor; andone or more storage medium having stored computer-executable instructions which are executable by the one or more processor to cause the computing system to implement a method for causing the computing system to conduct a medical examination, and by causing the computing system to at least perform the following: the computing system accessing a remote computing system, the remote computing system having one or more medical tests provided by one or more medical professionals for a medical patient;the computing system accessing a patient's medical information;the computing system identifying the on... | ORCA HEALTH, INC. | A61B003/02 A61B003/00 G06F019/00 A61B003/18 A61B005/12
US09345430 | B2 | 2016-05-24 | Imaging apparatus and imaging method thereof, image processing apparatus and image processing method thereof, and program | 2014-07-30 | An imaging apparatus includes: an unpolarized light-emitting portion configured to emit light having an unpolarized component; a polarized light-emitting portion configured to emit light having a predetermined polarized component via a first polarization filter; and an imaging element configured to image a subject, which is irradiated with light by one of the unpolarized light-emitting portion and the polarized light-emitting portion, through a second polarization filter, the first polarization filter and the second polarization filter having polarization directions in an orthogonal relationship to each other, the imaging element being further configured to image the subject irradiated with light by the unpolarized light-emitting portion and the subject irradiated with light by the polarized light-emitting portion in a time division manner and output an unpolarized light image and an orthogonal polarized light image that are obtained as a result of imaging. | 1. An imaging apparatus, comprising: an unpolarized light-emitting portion configured to emit light having an unpolarized component;a polarized light-emitting portion configured to emit light having a predetermined polarized component via a first polarization filter; andan imaging element configured to image a subject, which is irradiated with light by one of the unpolarized light-emitting portion and the polarized light-emitting portion, through a second polarization filter, the first polarization filter and the second polarization filter having polarization directions in an orthogonal relationship to each other, the imaging element being further configured to image the subject irradiated with light by the unpolarized light-emitting portion and the subject irradiated with light by the polarized light-emitting portion in a time division manner and output an unpolarized light image and an orthogonal polarized light image that are obtained as a result of imaging. 2. The imaging apparatus according to claim 1, wherein the unpolarized light-emitting portion and the polarized light-emitting portion include a plurality of unpolarized light-emitting portions and a plurality of polarized light-emitting portions that are arranged in a point symmetrical manner with the imaging element being a center. 3. The imaging apparatus according to claim 1, wherein the unpolarized light-emitting portion and the polarized light-emitting portion include a plurality of unpolarized light-emitting portions and a plurality of polarized light-emitting portions that are arranged in an annular manner with the imaging element being a center. 4. An imaging method of an imaging apparatus, the imaging apparatus including an unpolarized light-emitting portion configured to emit light having an unpolarized component,a polarized light-emitting portion configured to emit light having a predetermined polarized component through the first polarization filter, andan imaging element configured to image a subject, which is irradiated with light by one of the unpolarized light-emitting portion and the polarized light-emitting portion, through a second polarization filter, the first polarization filter and the second polarization filter having polarization directions in an orthogonal relationship to each other, the method comprising by the imaging apparatus:imaging the subject irradiated with light by the unpolarized light-emitting portion and the subject irradiated with light by the polarized light-emitting portion in a time division manner; andoutputting an unpolarized light image and an orthogonal polarized light image. 5. An image processing apparatus, comprising a specular reflection image generator configured to acquire an unpolarized light image and an orthogonal polarized light image that are captured by an imaging apparatus including an unpolarized light-emitting portion configured to emit light having an unpolarized component,a polarized light-emitting portion configured to em... | SONY CORPORATION | A61B005/00 G06T007/00 G06K009/00 G01N021/21
US09615764 | B2 | 2017-04-11 | Real-time coloring of electrophysiological map | 2014-11-03 | Catheterization of the heart is carried out with a probe having a plurality of electrodes and sensors by displaying an electroanatomical map of the heart on a monitor. During a time interval that does not exceed a duration of a cardiac cycle of the heart the following steps are performed: reading data from at least one of the electrodes and sensors, and invoking a processor to perform an algorithm on the data. The data is one of a plurality of inputs of the algorithm, and the result of the algorithm includes a transformation of the data. The method is further carried out by rendering the result of the algorithm on the monitor to modify the electroanatomical map. | 1. A method of catheterization, comprising the steps of: inserting a probe into a heart of a living subject, the probe having a plurality of electrodes and sensors;displaying an electroanatomical map of the heart on a monitor; andduring a time interval that does not exceed a duration of a cardiac cycle of the heart, performing the further steps of:reading data from at least one of the electrodes and sensors;invoking a processor to perform an algorithm on the data, the algorithm having a plurality of inputs and a result, wherein the data comprises one of the plurality of inputs and the result comprises a transformation of the data; andrendering the result on the monitor to modify the electroanatomical map; andwherein the algorithm comprises the steps of:selecting first electrodes from the electrodes wherein the first electrodes are in contact with tissue of the heart;identifying relative motion of the first electrodes with the tissue; selecting second electrodes from the first electrodes, wherein the second electrodes exhibit a predetermined amount of the relative motion;selecting third electrodes from the second electrodes, wherein respective locations of the third electrodes in successive beats have a correlation score below a predetermined threshold; andgenerating the result from data obtained from the third electrodes. 2. The method according to claim 1, further comprising the step of eliminating data that lies outside a portion of a respiratory cycle prior to generating the result. 3. The method according to claim 1, wherein the time interval does not exceed 500 ms. 4. The method according to claim 1, wherein the sensors comprise a temperature sensor. 5. The method according to claim 1, wherein the sensors comprise a position sensor. 6. The method according to claim 1, wherein the sensors comprise a contact force sensor. 7. The method according to claim 6, wherein the algorithm comprises an integration of data from the contact force sensor and position sensors and rendering comprises presenting a distorted surface on the electroanatomical map. 8. The method according to claim 1, wherein the sensors comprise optical sensors. 9. The method according to claim 8, wherein the data comprises optical reflectance data from the optical sensors and the algorithm comprises an analysis of fluctuations in the optical reflectance data and rendering comprises displaying an indication of an imminent steam pop event. 10. The method according to claim 1, wherein the algorithm comprises a calculation of wavefront propagation and rendering comprises updating wavefront annotations on the electroanatomical map. 11. The method according to claim 10, wherein the algorithm comprises a calculation of wavefront propagation and rendering comprises updating vectors that representing the wavefront annotations on the electroanatomical map. 12. The method according to claim 1, wherein the algorithm comprises an iterative analysis of temperature data from a plural... | BIOSENSE WEBSTER (ISRAEL) LTD | A61B005/04 A61B005/044 A61B005/042 A61B005/06 A61B005/0205 A61B005/00 A61B005/11 A61B005/01
US09345429 | B2 | 2016-05-24 | Image processing device, image processing system, image processing method, and program | 2012-12-13 | The present technique relates to an image processing device, an image processing system, an image processing method, and a program that enable simple and accurate detection of regions showing a detection object such as a body hair in images.;A body hair region detection unit includes: a projective transform unit that generates a projected image by projectively transforming a first image generated by photographing a body hair into the coordinate system of a second image generated by photographing the body hair from a different direction from the first image; a difference image generation unit that generates a difference image between the second image and the projected image; and a region detection unit that detects a candidate region formed with pixels having difference values equal to or larger than a predetermined threshold value in the difference image, and divides the region corresponding to the candidate region in the second image into a body hair region showing the body hair and a non-body-hair region outside the body hair region. The present technique can be applied to devices that analyze skin conditions, for example. | 1. An image processing device comprising a detection object region detection unit configured to detect a region showing a detection object in an image generated by photographing the detection object,wherein the detection object region detection unit includes: a projective transform unit configured to generate a projected image by projectively transforming a first image generated by photographing the detection object into a coordinate system of a second image generated by photographing the detection object from a different direction from the first image;a difference image generation unit configured to generate a difference image between the second image and the projected image; anda region detection unit configured to detect a candidate region formed with a pixel having a difference value equal to or larger than a predetermined threshold value in the difference image, and divide a region corresponding to the candidate region in the second image into a detection object region showing the detection object and a non-detection-object region outside the detection object region,in which the detection object is a body hair of a person such that the detection object region detection unit is configured to detect the region in the image which shows the body hair. 2. An image processing device comprising: a detection object region detection unit configured to detect a region showing a detection object in an image generated by photographing the detection object,wherein the detection object region detection unit includes: a projective transform unit configured to generate a projected image by projectively transforming a first image generated by photographing the detection object into a coordinate system of a second image generated by photographing the detection object from a different direction from the first image;a difference image generation unit configured to generate a difference image between the second image and the projected image; anda region detection unit configured to detect a candidate region formed with a pixel having a difference value equal to or larger than a predetermined threshold value in the difference image, and divide a region corresponding to the candidate region in the second image into a detection object region showing the detection object and a non-detection-object region outside the detection object region,further comprising a detection object removal unit configured to remove the detection object from an image generated by photographing the detection object,wherein the detection object removal unit includes: the detection object region detection unit; anda removal unit configured to remove the detection object from one of the first image and the second image by projecting at least pixels in a region corresponding to a region showing the detection object in the one of the first image and the second image onto pixels in the one of the images to replace the corresponding pixels based on a result of the detection performed by the dete... | Sony Corporation | G06K009/00 G06K009/46 G06K009/66 A61B005/00 G06T005/00 G06T007/00
US10271792 | B2 | 2019-04-30 | Biological information measuring apparatus | 2015-11-19 | A biological information measuring apparatus (measuring apparatus) includes a band which fixes a case unit to a living body. The band is provided with a recessed groove part on a side facing the living body. The groove part has a depth of 1020 μm or more and 1140 μm or less. | 1. A biological information measuring apparatus comprising: a detection unit which detects biological information;a case unit which houses the detection unit; anda band which fixes the case unit to a living body;wherein the band is provided with a recessed groove part on a side facing the living body, andthe groove part has a depth of 1020 μm or more and 1140 μm or less. 2. The biological information measuring apparatus according to claim 1, wherein the depth of the groove part is 1050 μm or more and 1100 μm or less. 3. The biological information measuring apparatus according to claim 1, wherein the depth of the groove part is 1060 μm or more and 1080 μm or less. 4. The biological information measuring apparatus according to claim 1, wherein the depth of the groove part at an end of the band is deeper than the depth of the groove part in the other parts of band. 5. The biological information measuring apparatus according to claim 1, wherein the groove part is provided along a direction of extension of the band. 6. The biological information measuring apparatus according to claim 5, wherein the groove part is provided to reach the case unit. 7. The biological information measuring apparatus according to claim 1, wherein the groove part is provided along a direction intersecting with a direction of extension of the band. 8. The biological information measuring apparatus according to claim 1, wherein the band has a hole part penetrating the band from the side facing the living body to the opposite side, and the groove part is provided to continue to the hold part. 9. The biological information measuring apparatus according to claim 1, wherein the band includes a first band part extending on one side of the case unit, and a second band part extending on the opposite side via the case unit, and the groove part is provided on the first band part and the second band part. 10. The biological information measuring apparatus according to claim 1, wherein the band is made of a material that is deformable along the living body. 11. The biological information measuring apparatus according to claim 10, wherein the band is made of an elastic resin material. 12. The biological information measuring apparatus according to claim 1, wherein the band includes a first band and a second band, the first band including the groove part and a buckle comprised with a flame and a buckle tongue, and the second band includes a plurality of holes to be inserted by the buckle tongue. 13. A biological information measuring apparatus comprising: a detection unit which detects biological information;a case unit which houses the detection unit; anda band which fixes the case unit to a living body;wherein the band is provided with a recessed groove part on a side facing the living body, andthe groove part has a width of 910 μm or more and 2300 μm or less. 14. The biological information measuring apparatus according to claim 13, wherein the width of the groove part is 95... | SEIKO EPSON CORPORATION | A61B005/00 A61B005/024
US09339215 | B2 | 2016-05-17 | Method and apparatus for monitoring movement and breathing of multiple subjects in a common bed | 2012-05-24 | The invention relates to methods and apparatus for monitoring movement and breathing of two or more subjects occupying common bedding. Particularly, the invention relates to a method for monitoring movement of subjects located in a common bedding, the method comprising the steps of: —imaging the bedding by an optical sensor; —performing a motion estimation by producing motion vectors indicating the local displacement of corresponding image blocks between consecutive images, or images that are several frames apart, received from said optical sensor; —calculating motion clusters by measuring spatial and temporal correlations of the motion vectors; —segmenting the calculated motion clusters by assignment of each motion cluster to a corresponding subject, wherein the assignment of the motion clusters to the corresponding subject is based on the spatial and/or temporal similarity of the motion clusters among each other and on previous segmentation results. | 1. A method for monitoring movement of subjects located in a common bedding, said subjects being persons, animals, or a combination of persons and animals, the method comprising the steps of: imaging the bedding by an optical sensor;performing a motion estimation by producing motion vectors indicating the local displacement of corresponding image blocks between consecutive images, or images that are several frames apart, received from said optical sensor;calculating motion clusters by measuring spatial and temporal correlations of the motion vectors;segmenting the calculated motion clusters by assigning each motion cluster to a corresponding subject,wherein an assignment of the motion clusters to the corresponding subject is based on any one or a combination of the spatial similarity and temporal similarity of the motion clusters among each other, andwherein a respiration analysis is performed for each subject in the common bedding based on data of the consecutive images, or several images apart, provided by the optical sensor in that the analysis of the breathing of two or more of the subjects in the common bedding, is performed by using the largest motion areas in said consecutive images, or images that are several frames apart, which are identified during times of no limb movement by the subjects. 2. The method according to claim 1, wherein the segmentation results of prior images are taken into consideration for the assignment of motion clusters to a corresponding subject in a subsequent image. 3. The method according to claim 1, wherein said images are segmented into boundaries of dominant chest motion. 4. The method according to claim 1, wherein an individual breathing analysis per subject is performed by executing an independent respiration analysis per identified said largest motion area. 5. The method according to claim 4, wherein a light pattern is projected on said common bedding and said pattern of projected light is deformed such that the micro movements due to breathing by one subject or the other subject can be extracted from the images. 6. The method according to claim 1, wherein the optical sensor additionally images also the environment of the bedding in addition to said two or more subjects sleeping in said common bedding. 7. The method according to claim 6, wherein the motion vectors resulting from the motion estimations are used to track the subjects within said bedding environment. 8. The method according to claim 1, wherein said motion estimation is performed by at least one algorithm from the group consisting of block-matching algorithm, optical flow estimation, maximum a posteriori probability estimation, Markov random field estimation, mean square error calculation, sum of absolute difference calculation, mean absolute differences calculation, sum of squared error calculation, or sum of absolute transformed differences calculation. 9. An apparatus for monitoring movement of subjects located in a common bedding en... | Koninklijke Philips N.V. | A61B005/113 G06T007/20 A61B005/08 A61B005/00
US10350389 | B2 | 2019-07-16 | Intravascular devices, systems, and methods having a radiopaque patterned flexible tip | 2015-11-02 | Intravascular devices, systems, and methods are disclosed. In some instances, the intravascular device is a guide wire with distal tip element having a radiopaque marker pattern. For example, in some implementations a sensing guide wire includes a flexible elongate member; a sensing element coupled to the flexible elongate member; a flexible tip element coupled to and extending distally from the sensing element, the flexible tip element having a first radiopaque section, a second radiopaque section, and a non-radiopaque section positioned between the first and second radiopaque sections. Methods of making, manufacturing, and/or assembling such intravascular devices and associated systems are also provided. | 1. A sensing guide wire, comprising: a flexible elongate member;a sensing element coupled to the flexible elongate member; anda flexible tip element coupled to and extending distally from the sensing element, the flexible tip element having a plurality of radiopaque sections each defined by a coil formed of radiopaque material and a plurality of non-radiopaque sections each defined by a coil formed of non-radiopaque material;wherein the plurality of radiopaque sections includes only a first radiopaque section, a second radiopaque section, and a third radiopaque section, the third radiopaque section being at a distal end of the flexible tip element;wherein the plurality of non-radiopaque sections includes only a first non-radiopaque section positioned between the first and second radiopaque sections and a second non-radiopaque section positioned between the second and third radiopaque sections;wherein the first radiopaque section, the second radiopaque section, the first non-radiopaque section, and the second non-radiopaque section are each shorter than the third radiopaque section;wherein a distal end of the coil defining the second radiopaque section is welded to a winding of the second non-radiopaque section such that the distal end of the coil defining the second radiopaque section is spaced apart circumferentially from a proximal end of the coil defining the second non-radiopaque section, the second non-radiopaque section being positioned adjacent to the second radiopaque section without windings of the second non-radiopaque section being intertwined with windings of the second radiopaque section; andwherein a proximal end of the coil defining the second radiopaque section is welded to a winding of the first non-radiopaque section such that the proximal end of the coil defining the second radiopaque section is spaced apart circumferentially from a distal end of the coil defining the first non-radiopaque section, the first non-radiopaque section being positioned adjacent to the second radiopaque section without windings of the first non-radiopaque section being intertwined with windings of the second radiopaque section. 2. The guide wire of claim 1, wherein the sensing element includes at least one of a pressure sensor or a flow sensor. 3. The guide wire of claim 1, wherein the first radiopaque section is adjacent a sensor housing in which the sensing element is disposed. 4. The guide wire of claim 1, wherein the radiopaque material is selected from the group consisting of platinum, tungsten, and gold. 5. The guide wire of claim 1, wherein the non-radiopaque material is selected from the group consisting of stainless steel, MP35N, and ninitol alloys. 6. The guide wire of claim 1, wherein the first and second radiopaque sections are welded to the first non-radiopaque section, and wherein the second and third radiopaque sections are welded to the second non-radiopaque section. 7. The guide wire of claim 1, wherein the flexible tip element ... | VOLCANO CORPORATION | A61B005/00 A61B005/026 A61M025/09 A61B005/0215
US09352131 | B2 | 2016-05-31 | Twisted ribbon wire guidewire coil | 2015-02-12 | A medical guidewire having distal tip coil that includes a twisted ribbon wire construction. The distal tip coil is mounted to the distal end of the distal core section of the guidewire. The coil is made from a flat ribbon wire with a polygonal, preferably rectangular, cross-section that is twisted along the length of the wire. The twisted, flat ribbon wire creates a tip coil that has a roughened, undulating surface at the outer circumference that, whether coated or bare, is perceptible to the physician or cardiologist when pushing or torquing the guidewire in an interventional procedure. | 1. A method for forming an intraluminal medical guidewire, comprising: providing a stationary, cylindrical core having a longitudinal axis A;forming a coil by wrapping onto the stationary cylindrical core a single strand of a wire from a wire supply reel having an axis B;anchoring a free end of the wire to the stationary cylindrical core;moving the wire supply spool continuously about axis A of the stationary cylindrical core to form 360° coils of the wire onto the stationary cylindrical core; andrevolve the wire supply spool about an axis C to impart a twisted shape to each coil on the stationary cylindrical core. 2. The method for forming an intraluminal medical guidewire of claim 1, wherein axis C is perpendicular to axis B. 3. The method for forming an intraluminal medical guidewire of claim 2, wherein the wire has a polygonal cross-sectional shape so that the coils have a shape in the form of a helical twist. 4. The method for forming an intraluminal medical guidewire of claim 3, wherein the polygonal cross-sectional shape of the wire is a rectangular cross-sectional shape. 5. The method for forming an intraluminal medical guidewire of claim 3, wherein the polygonal cross-sectional shape of the wire includes a cross-sectional shape selected from the group of cross-sectional shapes consisting of a rectangle, a square, a pentagon, or a hexagon. 6. The method for forming an intraluminal medical guidewire of claim 1, wherein the 360° coils are arranged to be at least one of overlapping, tightly packed, and spaced apart. 7. The method for forming an intraluminal medical guidewire of claim 1, wherein each 360° coil undergoes 1 to 10 revolutions of the wire supply reel about axis C. 8. The method for forming an intraluminal medical guidewire of claim 2, wherein the coil wire has a constant outside diameter. | ABBOTT CARDIOVASCULAR SYSTEMS INC. | A61B005/00 A61M025/00 A61M025/09
US09326726 | B2 | 2016-05-03 | Wireless system for epilepsy monitoring and measurement | 2014-05-16 | A wireless system for brain monitoring/mapping of neurological-disorder patients includes a plurality of electrodes each configured for surface abutment of brain tissue and main circuitry for placement outside a body of a patient and configured to transmit power at radio frequencies and send and receive data using infrared energy. Remote circuitry is provided for subcutaneous implantation in a head of the patient. The remote circuitry is connected to the plurality of electrodes and includes a multiplexer sampling signals from the plurality of electrodes. The multiplexer outputs electrode signals to an amplifier and A/D converter for transmission to the main circuitry. The remote circuitry is configured to (a) receive transmitted power at radio frequencies from the main circuitry, (b) capture and digitize full-bandwidth EEG signals from each of the electrodes, and (c) send data to and receive data from the main circuitry using infrared energy. | 1. A digital EEG system for signal processing circuitry for brain monitoring/mapping of neurological-disorder patients, comprising: a plurality of scalp electrodes;signal processing circuitry connected to the plurality of scalp electrodes, the signal processing circuitry including:a multiplexer sampling signals from the plurality of scalp electrodes, the multiplexer outputting electrode signals; an amplifier amplifying electrode signals from the plurality of scalp electrodes; andan A/D converter converting the electrode signals to digital for transmission to an EEG recording device. 2. The digital EEG system of claim 1 wherein the multiplexer include an array of series L and shunt capacitors at each channel input. 3. The digital EEG system of claim 1 wherein the multiplexer includes a low pass filter at each channel input. | YALE UNIVERSITY ITN ENERGY SYSTEMS, INC. | A61N001/00 A61B005/04 A61B005/00 A61B005/0476 A61B005/0478
US20160120445 | A1 | 2016-05-05 | DEVICES, SYSTEMS AND METHODS FOR CONTEXTUALIZED RECORDING OF BIOMETRIC MEASUREMENTS | 2015-11-04 | Described are apparatuses (e.g., devices, systems, software), and methods for monitoring the cardiac health of a patient which incorporates secondary information. The apparatus may be a heart rate monitor, a blood pressure monitor, an ECG or other biometric device in communication with an electronic device having communication ability. The devices detect a cardiac or other biometric measurement and correlate the detected information with secondary information, such as GPS data, to produce a contexualized measurement. | 1. An apparatus adapted to correlate a blood pressure measurement with GPS data comprising: a cuff for wrapping around a limb of a subject the cuff further having an inflatable bladder formed integrally therein, wherein the cuff is substantially rectangular in shape having a first surface, a second surface, a first end, a second end, and two sides, and a securement mechanism;one or more tubes positioned within the cuff in communication with the inflatable bladder;an exhaust valve;a housing connected to one of a first end and a second end of the cuff wherein the housing includes one or more of a controller, a memory, a pump, a power supply, and a motor;a GPS sensor in communication with the apparatus; anda transmitter for wirelessly transmitting blood pressure data to a second device. 2. The apparatus of claim 1 wherein the blood pressure monitor device is configurable to be in communication with a communication network; and a computer executable instruction that, when executed by a processor analyzes a received blood pressure data and GPS data. 3. The apparatus of claim 2 wherein the computer executable instruction analyzes one or more of patient data and electronic device data. 4. The apparatus of claim 1 wherein the historical user data and historical GPS data is accessible via a communication network. 5. The apparatus of claim 1 further comprising one or more of an accelerometer, a barometer, a thermometer, a UV sensor, a microphone, a galvanic skin response sensor, a body posture sensor, and skin hydration sensor in communication with the apparatus. 6. The apparatus of claim 1 further comprising a pressure detector. 7. The apparatus of claim 1 further comprising a pulse detector. 8. The apparatus of claim 1 wherein patient data and electronic device data is accessible via a communication network. 9. The apparatus of claim 1 wherein the blood pressure monitoring device is in wireless communication with one or more displays. 10. An apparatus adapted to correlate a blood pressure measurement with data recorded by one or more of a barometer, an accelerometer, and a thermometer comprising: a cuff for wrapping around a limb of a subject the cuff further having an inflatable bladder formed integrally therein, wherein the cuff is substantially rectangular in shape having a first surface, a second surface, a first end, a second end, and two sides, and a securement mechanism;one or more tubes positioned within the cuff in communication with the inflatable bladder;an exhaust valve;a housing connected to one of a first end and a second end of the cuff wherein the housing includes one or more of a controller, a memory, a pump, a power supply, and a motor; anda transmitter for wirelessly transmitting blood pressure data to a second device;wherein the one or more of the barometer, the accelerometer and the thermometer are in communication with the apparatus. 11. The apparatus of claim 10 wherein the blood pressure monitor device is configurable to... |  | A61B005/11 A61B005/1455 A61B005/022 A61B005/0235 A61B005/0205 A61B005/00
US10357200 | B2 | 2019-07-23 | Scanning laser vein contrast enhancer having releasable handle and scan head | 2015-11-03 | The present invention is a Miniature Vein Enhancer that includes a Miniature Projection Head. The Miniature Projection Head may be operated in one of three modes, AFM, DBM, and RTM. The Miniature Projection Head of the present invention projects an image of the veins of a patient, which aids the practitioner in pinpointing a vein for an intravenous drip, blood test, and the like. The Miniature projection head may have a cavity for a power source or it may have a power source located in a body portion of the Miniature Vein Enhancer. The Miniature Vein Enhancer may be attached to one of several improved needle protectors, or the Miniature Vein Enhancer may be attached to a body similar to a flashlight for hand held use. The Miniature Vein Enhancer of the present invention may also be attached to a magnifying glass, a flat panel display, and the like. | 1. A handheld vein-image-enhancing system comprising: a detachable scan head, said detachable scan head comprising: a first laser configured to emit a beam of light at a first wavelength;a scanner configured to transmit said beam of light from said first laser onto a target area;a photodetector configured to receive an image of the veins formed by differential absorption of said selective wavelength of light by the veins and the surrounding tissue; said photodetector further configured to output a signal representing said image;a second laser configured to receive said signal from said photodetector and to use said signal to emit a beam of light at a second wavelength, said scanner further configured to scan said beam of light at said second wavelength from said second laser to project said received image onto the target area;a housing, an interior of said housing configured to house said first laser, said second laser, said photodetector, and said scanner, said housing comprising: an exit window for said scanned laser light; anda scan head connector comprising: means for mechanically coupling; andmeans for electrically coupling;anda handle, said handle comprising: a housing; said housing comprising an opening into a battery compartment configured to receive a battery;a handle connector comprising:means for mechanically coupling said handle connector to said means for mechanically coupling of said scan head connector; andmeans for electrically coupling said handle connector to said means for electrically coupling of said scan head connector;a cover for said battery compartment:means for latching said cover in a closed position; andmeans for disengaging said means for mechanically coupling of said scan head from said means for mechanically coupling of said handle. 2. The handheld vein-image-enhancing system according to claim 1 further comprising: an electrical switch;a trigger, said trigger molded into said housing and configured to flex when depressed to toggle said electrical switch. 3. The handheld vein-image-enhancing system according to claim 1 further comprising: means for cooling said interior of said scan head. | AccuVein, Inc. | A61B005/00 A61M005/42
US09340818 | B2 | 2016-05-17 | Diagnostic apparatus and diagnostic method using the same | 2011-10-05 | Provided is a diagnostic apparatus. The diagnostic apparatus includes a microfluidic chip including first and second measurement parts for respectively measuring an amount of hemoglobin and an active degree of an enzyme within a blood sample. The second measurement part of the microfluidic chip analyzes the active degree of the enzyme within the blood sample using voltammetry. | 1. A diagnostic apparatus comprising a microfluidic chip comprising first and second measurement parts for respectively measuring an amount of hemoglobin and an active degree of an enzyme within a blood sample, wherein the microfluidic chip comprises:a sample inlet port for inputting the blood sample;a sample dissolution part for converting the blood sample into lysate;a passage through which the lysate is moved into the first and second measurement parts; andlines electrically connected to the first and second measurement parts, respectively, whereinthe sample dissolution part comprises detergent, andthe second measurement part of the microfluidic chip comprises an electrode and an electron transport medium. 2. The diagnostic apparatus of claim 1, wherein the first measurement part of the microfluidic chip measures the amount of hemoglobin within the blood sample using photometry or voltammetry. 3. The diagnostic apparatus of claim 2, wherein the first measurement part of the microfluidic chip is provided in plurality to measure the amount of hemoglobin within the blood sample using a plurality of measuring methods selected from the photometry, the voltammetry, or the colorimetry. 4. The diagnostic apparatus of claim 1, wherein the second measurement of the microfluidic chip measures the active degree of the enzyme within the blood sample using voltammetry. 5. The diagnostic apparatus of claim 4, wherein the enzyme is glucose-6-phosphate dehydrogenase (G6PD). 6. The diagnostic apparatus of claim 1, further comprising: a loading part on which the microfluidic chip is loaded; anda socket part circuitally connected to the loading part. 7. The diagnostic apparatus of claim 1, further comprising first and second analysis parts for analyzing values respectively measured by the first and second measurement parts of the microfluidic chip. 8. The diagnostic apparatus of claim 1, further comprising an amplifier for amplifying values measured by the first and second measurement parts of the microfluidic chip. 9. The diagnostic apparatus of claim 1, further comprising an analog-digital converter for converting values measured by the first and second measurement parts of the microfluidic chip into digital signals. 10. The diagnostic apparatus of claim 1, further comprising a display part for converting values analyzed by the first and second analysis parts into images. 11. The diagnostic apparatus of claim 1, further comprising a control part for controlling the microfluidic chip, the socket part, the first and second analysis parts, and the display part. 12. The diagnostic apparatus of claim 1, wherein the sample dissolution part constitutes the first and second measurement parts, and the branched passage is directly connected to the sample inlet port. 13. The diagnostic apparatus of claim 1, wherein the detergent comprises at least one selected from soap, sulfonate, sulfate, phosphate, phosphate, alkyl glycoside, bide acid, glucamide, polyoxyet... | ACCESS BIO, INC. | G01N033/00 C12Q001/32 A61B005/145 A61B005/1455 A61B005/1468 A61B005/1486 G01N027/327
US09332904 | B2 | 2016-05-10 | Quantifying ocular counter roll | 2012-05-30 | This document provides methods and materials for quantifying ocular counter roll and detecting vestibular otolith damage. For example, devices for quantifying ocular counter roll, methods of making devices for quantifying ocular counter roll, and methods for using a device to quantify (e.g., quantify indirectly) ocular counter roll are provided. | 1. A device for quantifying ocular counter roll, wherein said device comprises: (a) a viewing tube configured to have a hollow inner lumen;(b) a calibration line configured to be fixed with respect to said viewing tube;(c) an eye piece component configured to have a hollow inner lumen wherein said eye piece component allows a clear viewing path through said viewing tube; and(d) an inclinometer. 2. The device of claim 1, wherein said device is mechanically fixed to viewing goggles, glasses, or a handle. 3. The device of claim 1, wherein said viewing tube is configured to fit over an end of said eye piece component. 4. The device of claim 1, wherein said device comprises two inclinometers. 5. A method for quantifying ocular counter roll using a device comprising a viewing tube, a calibration line configured to be fixed with respect to said viewing tube, an eye piece component configured to allow a clear viewing path through said viewing tube, and an inclinometer to view an external target line placed within about 1 to 6 meters of the subject, wherein said method comprises: (a) obtaining a baseline measurement from said inclinometer when a subject perceives said calibration line and said external target line are parallel;(b) adjusting the angle of said external target line by 20° or less;(c) measuring the angle of said subject's head tilt with said inclinometer when said subject perceives said calibration line and said external target line are parallel; and(d) quantifying ocular counter roll as the difference between the angle of said external target line and the measured angle of said subject's head tilt. 6. The method of claim 5, wherein said device comprises two inclinometers. 7. A method for quantifying ocular counter roll using a device comprising a viewing tube, a calibration line configured to be fixed with respect to said viewing tube, an eye piece component configured to allow a clear viewing path through said viewing tube, and an inclinometer to view an external target line placed within about 1 to 6 meters of the subject, wherein said method comprises: (a) a human subject looking through said eye piece component and said viewing tube to visualize said calibration line and said external target line;(b) said subject, optionally, calibrating said device by manually rotating said viewing tube about the viewing axis until said calibration line and said external target line are parallel;(c) said subject tilting said subject's head with said device held fixed in relation to said subject's head until said subject perceives said calibration line and said external target line are parallel after the angle of said external target line is tilted by 20° or less; and(d) quantifying ocular counter roll as the difference between said subject's head tilt and external target line tilt. 8. The method of claim 7, wherein said subject's vision from the opposite eye is precluded. 9. A method for quantifying ocular counter roll using a device comprising ... | Mayo Foundation for Medical Education and Research | A61B003/14 A61B003/113 A61B005/12 A61B005/11 A61B005/00
US09345438 | B2 | 2016-05-24 | Weaning from ventilation using capnography | 2014-02-24 | Devices and systems for monitoring weaning of a subject from a respiratory ventilator including a processing logic configured to characterize distinct patterns in a series of CO2 waveforms, the distinct patterns indicative of the effectiveness of a weaning process; and to provide an indication relating to the effectiveness of the weaning process. | 1. A method for monitoring effectiveness of weaning of a subject from a respiratory ventilator, the method comprising: obtaining a series of CO2 waveforms from expired breath of a subject undergoing respiratory ventilation weaning;identifying a distinct pattern in the series of CO2 waveforms;determining a frequency and/or an extent of the distinct pattern in the series of CO2 waveforms; andproviding an indication relating to the effectiveness of the weaning process, based on the determined frequency and/or extent of the distinct pattern;wherein an increase in the frequency and/or the extent of the distinct pattern, as compared to a baseline, is indicative of a deterioration in the weaning effectiveness. 2. The method of claim 1, wherein said distinct pattern is selected from a group consisting of “sigh events”, “spike events” and “pools”. 3. The method of claim 1, further comprising providing a signal to a ventilator at least partially based on the determined frequency and/or extent of the distinct patterns in the series of CO2 waveforms and thus inducing a change in one or more parameters of the ventilator. 4. The method of claim 3, wherein said change in one or more parameters of the ventilator comprises a change in the level of support to the ventilation, allowing the patient to rest from weaning attempts, changing a mode/program of weaning or any combination thereof. 5. The method of claim 3, wherein providing a signal to the ventilator is further based on one or more additional parameters selected from the group consisting of: age of the subject, medical condition of the subject, medical history of the subject and medications administered to the subject. 6. The method of claim 3, wherein providing a signal to the ventilator is further based on one or more additional parameters selected from the group consisting of: respiration rate, tidal volumes, minute ventilation (the total lung ventilation per minute), inspiration rate and expiration rate at ambient pressure in the absence of ventilator assistance, ventilator mode and ventilator setup parameters. 7. The method of claim 1, wherein characterizing distinct patterns in a series of CO2 waveforms comprises segmenting the series of CO2 waveforms into breaths (waveforms), and calculating for each breath one or more features. 8. The method of claim 1, wherein the one or more features comprise: area under the curve (AUC), breath duration (inhalation, expiration or both), I to E ratio, maximal CO2 value, minimal CO2 value or any combination thereof. 9. The method of claim 1, wherein characterizing distinct patterns in the series of CO2 waveforms comprises determining the ratios between the distinct patterns. | ORIDION MEDICAL 1987 LTD. | A61B005/083 A61B005/08 A61M016/00 A61B005/00 A61M016/08
US09355215 | B2 | 2016-05-31 | Connector interface system for data acquisition | 2015-04-24 | A data acquisition system includes a receptacle and a data acquisition device. The receptacle has a housing, sensor inputs to receive data signals from sensors coupled to an object, and a rib to block insertion of a standard Universal Serial Bus (USB) plug and facilitate insertion of a modified USB plug having a slot that mates with the rib. The data acquisition device includes circuitry to receive, store and process data, a USB plug having pins operatively coupled to the circuitry, a first subset of pins configured to receive data signals from the receptacle and a second subset of pins configured to support standard USB communication with USB-compliant devices, and a slot formed in the USB plug such that the slot facilitates interconnection of the USB plug both with standard USB-compliant devices and with the receptacle, the slot mating with the rib to facilitate interconnection. | 1. A data acquisition device comprising: circuitry configured to receive data acquired from one or more sensors having electrode leads coupled to a living subject;a data connector having a plurality of pins operatively coupled to the circuitry and configured to: receive data signals from the one or more sensors having electrode leads coupled to the living subject; andsupport communication through a standard interface with devices capable of communicating through the standard interface; anda modification to the data connector configured such that the modification facilitates interconnection of the data connector both with the standard interface and with a living subject-connected receptacle that is configured to prevent interconnection with the standard interface. 2. The device of claim 1 in which the modification formed in the data connector comprises a slot that mates with an accommodating feature in the living subject-connected receptacle. 3. The device of claim 2 in which the data connector has top and bottom sides, and the slot is formed symmetrically in at least one of the data connector's top and bottom sides. 4. The device of claim 1 in which the device is configured to acquire biometric data from the living subject. 5. The device of claim 1 further comprising circuitry to store, process and/or transmit the acquired data. 6. A data acquisition receptacle comprising: a plurality of sensor inputs configured to receive data signals from one or more sensors having electrode leads coupled to a living subject; anda connection prevention mechanism that is configured to prevent insertion of a standard data connector and to facilitate insertion of and electrical connection with a modified data connector, the standard data connector and the modified data connector each supporting communication through a standard interface, the modified data connector having a feature that defeats the connection prevention mechanism. 7. The receptacle of claim 6 in which the connection prevention mechanism comprises a rib. 8. The receptacle of claim 7 in which the feature of the modified data connector that defeats the receptacle's connection prevention mechanism comprises a slot that mates with the rib when the modified data connector is inserted into the receptacle. 9. The receptacle of claim 7 in which the rib is formed along a center portion of the receptacle. 10. The receptacle of claim 7 further comprising a substrate configured to facilitate secure connection of the modified data connector while inserted in the receptacle. 11. The receptacle of claim 6 further comprising a recess to receive the modified data connector. 12. The receptacle of claim 6 in which the receptacle is formed in an unshrouded configuration that lacks a recess. 13. A data acquisition system comprising: (i) a receptacle including: (a) a housing;(b) a plurality of sensor inputs configured to receive data signals from one or more sensors having electrode leads coupled to a living... | BRAEMAR MANUFACTURING, LLC | G06F003/06 G06F019/00 A61B005/024 H01R013/642 G06F013/38 G06F013/40 H01R013/44 G06F013/00 H01R027/00
US09339213 | B2 | 2016-05-17 | Measuring device and method for indicating level of fatigue | 2013-03-07 | Gradually strengthening electric stimulation impulses are sent to the median nerve at the wrist. Thereby an action potential caused on some level of the stimulation signal proceeds along the sensory nerve to the spinal cord, from where the nerve impulse (Hoffmann's reflex) returns via the alpha motor neuron pool along a motor nerve to a muscle and causes the movement of the muscle. This movement is detected by an accelerometer. The intensity of the stimulation signal and the time from the stimulation signal to the motoric response are saved, and by means of these the level of fatigue is evaluated. A device comprises a cuff with two electrodes and a current generator for forming impulses for the electrodes. The cuff can be placed on the wrist such that at least one of the electrodes is on the volar side of the hand at the point of the median nerve. The measuring device further comprises an accelerometer, that is provided to produce a signal, when the accelerometer detects a movement. The measuring device further comprises a control arrangement, which is provided for guiding the current generator to produce impulses and to monitor the signal of the accelerometer. The control arrangement stops the production of impulses, when the signal created by the accelerometer is detected, and saves the intensity value of at least the latest impulse and the time ΔT, which is the difference between the points of time of the signal produced by the accelerometer and said latest impulse. From these the value describing the level of fatigue can be calculated. | 1. A measuring device for indicating level of fatigue, the measuring device comprising a cuff with two electrodes, anode and cathode, and a current generator which is arranged to form impulses for electrodes, wherein the cuff is provided in a way that it can be placed on the wrist such that at least one of the electrodes is on the volar side of the hand at the point of the median nerve, and the measuring device further comprises an accelerometer that is provided to produce a signal when the accelerometer detects a movement, and the measuring device further comprises a control arrangement that is provided to guide the current generator to produce impulses and to monitor the signal of the accelerometer,to grow the intensity of the impulses in regard to time,to stop producing impulses, when the signal produced by the accelerometer is detected,to save at least the intensity value of the latest impulse and time ΔT, which is a difference between a time of the signal produced by the accelerometer and a time of the latest impulse, andto calculate a level of fatigue based on at least one of the intensity value of the latest impulse and the time ΔT. 2. The measuring device according to claim 1, wherein the current produced by the current generator is alternating current. 3. The measuring device according to claim 1, wherein the current produced by the current generator is direct current. 4. The measuring device according to claim 1, wherein the signals produced by the current generator are of a constant length. 5. The measuring device according to claim 1, wherein the signals produced by the current generator have a constant frequency, i.e. the time difference between two successive signals is essentially constant. 6. The measuring device according to claim 1, wherein the cuff is provided to go around the wrist, and the electrodes are placed in a way, that one electrode is on the dorsal side of the hand. 7. The measuring device according to claim 1, wherein the electrodes are placed in the cuff such that when the cuff is in its position, both electrodes are on the volar side of the hand essentially one after another in direction of the median nerve. 8. The measuring device according to claim 1, wherein the measuring device is provided such that it comprises a part that extends to a finger of the hand that the cuff is attached to, and the accelerometer is in this part. 9. The measuring device according to claim 8, wherein the part is configured to extend to a thumb, a forefinger or a middle finger of the hand. 10. The measuring device according to claim 1, wherein the signal produced by the accelerometer is the H-reflex caused by stimulation of the median nerve. 11. The measuring device according to claim 1, wherein the control arrangement comprises a calculation unit for calculating a value of fatigue based on at least the saved intensity value of the impulse or the time Δ,T or both, and for comparing it/them with a reference value, and the diffe... | Juno Medical LLC | A61B005/00 A61B005/11 A61N001/04 A61B005/16
US10185395 | B2 | 2019-01-22 | Input device, information processing device, and input method | 2015-12-24 | To provide a device that efficiently selects desired data out of a large number of selectable data.;A pulse meter 10 includes a body motion sensor 60 that detects body motion information concerning a limb, an inclination-angle calculating unit 24 that calculates an inclination angle of the limb on the basis of the body motion information, a display unit 50 that displays an input candidate value corresponding to the inclination angle, and a determination instructing unit 40 that decides the input candidate value as an input value. | 1. An input device comprising: a case configured to be worn on the wrist of a user:a body-motion-information detecting unit configured to detect body motion information concerning the wrist;an inclination-angle calculating unit configured to calculate an inclination angle of the limb wrist on the basis of the body motion information;a candidate-value determining unit configured to determine an input candidate value, from a plurality of displayed data selectable as the input candidate value, by changing the input candidate value from a first input candidate value to a second input candidate value using the calculated inclination angle, wherein the first and second input candidate values correspond to individual information about the user used for calculating biological information relating to the user:a display unit configured to change a display portion displaying the first input candidate value to the second input candidate value in response to the second input candidate value being the determined input candidate value; anda setting unit configured to receive a user input while the determined input candidate value is displayed on the display portion and set the determined input candidate value as an input value in response to receiving the user input, wherein the user input is one of a depression of a button or a tap on a case housing the input device,wherein the body-motion-information detecting unit, the inclination-angle calculating unit, the candidate-value determining unit, the display unit, and the setting unit are housed within the case. 2. The input device according to claim 1, wherein the body-motion-information detecting unit includes an acceleration sensor configured to output an acceleration signal, andthe inclination-angle calculating unit calculates a displacement amount from the inclination angle in a predetermined posture of the limb on the basis of the acceleration signal. 3. The input device according to claim 2, further comprising: a storing unit configured to store the first input candidate value in the predetermined posture as a reference value, whereinthe candidate-value determining unit determines an increase/decrease value on the basis of the inclination angle and calculates the second input candidate value using the reference value and the increase/decrease value. 4. The input device according to claim 3, wherein the storing unit stores an upper limit value and a lower limit value of the inclination angle and an upper limit value and a lower limit value of the input candidate value, andwhen the inclination angle is one of the upper limit value and the lower limit value, the candidate-value determining unit determines the input candidate value as one of the upper limit value and the lower limit value. 5. The input device according to claim 3, wherein the candidate-value determining unit increases a rate of change of the increase/decrease value according to an increase in the displacement amount. 6. The input device a... | Seiko Epson Corporation | G06F003/01 A61B005/00 A61B005/021 A61B005/024 A61B005/11 A61B005/0205 G06F003/0346 G06F001/16
US09354493 | B2 | 2016-05-31 | Projection system | 2015-03-06 | A projection system includes a light source, an imaging unit, a control unit, and a projecting unit. The light source applies a light beam having a predetermined wavelength. The imaging unit captures a subject to which the light beam is applied. The control unit generates image data for projection based on the captured image. The projecting unit projects a projection image onto the subject. The imaging unit captures the projection image together with an image of a region responding to the light beam. The control unit corrects the image data for projection to additionally display the projection image on a region which responds to the light beam and does not overlap a region onto which the projection image is projected, and to erase the projection image in a region onto which the projection image is projected and does not overlap the region responding to the light beam. | 1. A projection system comprising: a light source for applying a light beam having a predetermined wavelength;an imaging unit for capturing a subject to which the light beam having the predetermined wavelength is applied;a control unit for generating image data for projection based on the image captured by the imaging unit;a projecting unit for projecting, onto the subject, a projection image based on the image data for projection, whereinthe imaging unit captures the projection image together with an image of a region responding to the light beam having the predetermined wavelength, andthe control unitcorrects the image data for projection to additionally display the projection image on a region which responds to the light beam having the predetermined wavelength and does not overlap a region onto which the projection image is projected, andcorrects the image data for projection to erase the projection image in a region onto which the projection image is projected and does not overlap the region responding to the light beam having the predetermined wavelength. 2. The projection system according to claim 1, wherein the image data for projection generated by the control unit is an image representing the region responding to the light beam having the predetermined wavelength. 3. The projection system according to claim 1, further comprising: a storage unit for storing plural pieces of figure information each representing image data of an object having a predetermined figure; whereinthe control unitselects a piece of figure information from the plural pieces of figure information stored in the storage unit, based on a degree of similarity with the image of the region responding to the light beam having the predetermined wavelength captured by the imaging unit, andgenerates the image data for projection such that an image having a figure indicated by the selected piece of figure information is projected onto the region responding to the light beam having the predetermined wavelength. 4. The projection system according to claim 1, wherein the subject includes an affected part of a patient or a medical device. 5. The projection system according to claim 1, wherein the subject includes an object to be captured which a fluorescent substance responding to the light beam having the predetermined wavelength is applied thereon, is kneaded therein, or is poured therein. | PANASONIC INTELLECTUAL PROPERTY MANAGEMENT CO., LTD. | H04N005/222 G03B021/14 A61B005/06 A61B019/00 G06T001/00 G06F003/01 H04N009/31 H04N005/232 H04N005/74 G03B015/14 G03B017/54 A61B005/00 H04N007/18
US09333355 | B2 | 2016-05-10 | Methods, systems, and devices for pairing vagus nerve stimulation with motor therapy in stroke patients | 2014-12-04 | A method of treating motor deficits in a stroke patient, comprising assessing a patient's motor deficits, determining therapeutic goals for the patient, based on the patient's motor deficits, selecting therapeutic tasks based on the therapeutic goals, performing each of the selected therapeutic tasks repetitively, observing the performance of the therapeutic tasks, initiating the stimulation of the vagus nerve manually at approximately a predetermined moment during the performance of the therapeutic tasks, stimulating the vagus nerve of the patient during the performance of the selected therapeutic tasks, and improving the patient's motor deficits. | 1. A stimulation control device, comprising: a stimulation initiation system configured to output stimulation initiation signals to a vagus nerve stimulator;a manual input system configured to be engaged by an operator to cause output of a first stimulation initiation signal;a motion detection system configured to be engaged by the operator to cause output of a second stimulation initiation signal; anda selection input system configured to enable the operator to select between using the manual input system and the motion detection system,wherein the stimulation control device is configured such that engaging the manual input system causes the output of the first stimulation initiation signal from the stimulation initiation system to the vagus nerve stimulator, andthe stimulation control device is configured such that engaging the motion detection system enables the output of the second stimulation initiation signal from the stimulation initiation system to the vagus nerve stimulator. 2. The stimulation control device of claim 1, wherein: the motion detection system is configured such that engaging the motion detection system enables automatic output of the second stimulation initiation signal from the stimulation initiation system to the vagus nerve stimulator. 3. The stimulation control device of claim 1, wherein: the motion detection system is configured such that engaging the motion detection system enables automatic output of the second stimulation initiation signal from the stimulation initiation system to the vagus nerve stimulator based on a detection of a movement performed by a patient to which the vagus nerve stimulator provides stimulation. 4. The stimulation control device of claim 1, further comprising: a visual output system configured to provide data to the operator. 5. The stimulation control device of claim 4, wherein: the visual output system is configured to enable input by the operator, thereby enabling the operator to engage the selection input system to select between using the manual input system and the motion detection system to respectively cause the output of the first stimulation initiation signal or the output of the second stimulation initiation signal from the stimulation initiation system to the vagus nerve stimulator. 6. The stimulation control device of claim 1, wherein: the manual input system includes a manual switch configured to be engaged by an operator to cause the output of the first stimulation initiation signal. 7. A system, comprising: a vagus nerve stimulator;a stimulation control device communicably connected to the vagus nerve stimulator including a stimulation initiation output to provide stimulation initiation signals to the vagus nerve stimulator;a manual control system configured to be engaged to provide a stimulation initiation signal from the stimulation control device upon a manual input;a motion detection system configured to be engaged to provide a stimulation initiation signal from th... | MICROTRANSPONDER, INC. THE BOARD OF REGENTS, THE UNIVERSITY OF TEXAS SYSTEM | A61N001/00 A61N001/36 A61N001/372 G09B023/28 A61B005/11
US10365343 | B2 | 2019-07-30 | Multi-echo length and offset varied saturation (Me-LOVARS) method for chemical exchange saturation transfer (CEST) MR imaging | 2014-06-03 | A novel approach for CEST MR imaging is called Multi-echo Length and Offset VARied Saturation (Me-LOVARS) CEST. This method allows efficient collection of additional CEST data without penalty in scan time, which could be useful for enhancing the contrast, increasing the specificity or improving quantification of exchange. As CEST-MRI has shown promise at both the pre-clinical and clinical levels, including for detecting and grading brain tumors and evaluating ischemia, using either endogenous CEST contrast or through administration of CEST contrast agents. This fast and robust imaging method is suitable for accelerating image collection and widening the scope of applications for CEST-MRI. | 1. A method for magnetic resonance imaging of a subject comprising: generating a magnetization using an magnetic resonance imaging machine;applying “N” number of iterative modules to achieve multiple readouts;wherein the “N” number of iterative modules each comprise a saturation preparation component, a readout component, and a flip back component;processing the multiple readouts into an image of the subject. 2. The method of claim 1 further comprising using N equals approximately 3 to 8 modules. 3. The method of claim 1 further comprising using modules with a length of tsat/N. 4. The method of claim 1 further comprising using modules with a length of approximately 0.3 seconds to 1 second. 5. The method of claim 1 further comprising using the readout component comprising a low flip angle fast gradient echo readout sequence. 6. The method of claim 5 further comprising using a flip angle of α. 7. The method of claim 1 further comprising using a flip back pulse for retaining the magnetization. 8. The method of claim 1 further comprising defining magnetization as longitudinal magnetization:MNz,sat=xs⁢ɛ⁢⁢ksw⁢T1,w·MNz·e-TsatN·T1,w=b·MNz·e-TsatN·T1,win the absence of saturated protons. 9. The method of claim 8 further comprising defining longitudinal magnetization after the Nth module asMNz,-α=⁢MNz,α·cos⁢⁢α+MNx,α·e-TET2*·sin⁢⁢α=⁢b·MNz·e-TsatN·T1,w[1-sin2⁢α⁡(1-e-TET2*)]. 10. The method of claim 1 further comprising defining magnetization as longitudinal and transverse magnetization. 11. The method of claim 1 further comprising defining magnetization as MNz,α=MNz,sat·cos α a)MNx,α=MNz,sat·sin α b)when applying the flip angle component. 12. The method of claim 1 further comprising defining an iterative relationship between modules asMNz,-α=⁢b·MN-1z,-α·e-TsatN·T1,w[1-sin2⁢α⁡(1-e-TET2*)]=⁢b·MN-2z,-α·e-2⁢TsatN·T1,w[1-sin2⁢α⁡(1-e-TET2*)]2⁢…⁢⁢…=⁢b·M0z,-α·e-TsatT1,w[1-sin2⁢α⁡(1-e-TET2*)]N. 13. The method of claim 1 further comprising using a T2 decay term. 14. The method of claim 1 further comprising executing the method using a non-transitory computer readable medium. | The Johns Hopkins University Kennedy Krieger Institute, Inc. | G01V003/00 G01R033/561 A61B005/055 A61B005/00 G01R033/56
US09829302 | B2 | 2017-11-28 | System and method for measuring a non-bone tendon | 2015-11-13 | There is disclosed a system and method for measuring the diameter or thickness of a tendon. In an embodiment, the device includes an adjustable tube having a variably sized cross-sectional diameter in a direction perpendicular to an axis extending between the first end and the second end, the adjustable tube being positionable from an open configuration in which the variably sized cross-sectional diameter is larger than a diameter of a tendon so as to allow loading of the tendon, to a closed configuration in which the variably sized cross-sectional diameter approximates the diameter of the tendon so as to allow measurement of the tendon. The device includes a measurement gradient in operable configuration with the adjustable tube so as to allow measurement of the tendon in the closed configuration. Other embodiments are also disclosed. | 1. A tendon measurement device, comprising: an adjustable tube having a first end, a second end, and an inner region having a variably sized cross-sectional diameter in a direction perpendicular to an axis extending between said first end and said second end, said adjustable tube positionable between an open configuration in which said inner region of said variably sized cross-sectional diameter is larger than a diameter of a tendon so as to allow loading of the tendon within the inner region and a closed configuration in which said inner region of said variably sized cross-sectional diameter approximates the diameter of the tendon within said inner region so as to allow measurement of the tendon within said inner region; anda measurement gradient operably associated with said adjustable tube so as to allow measurement of the tendon within said inner region when said adjustable tube is positioned in said closed configuration. 2. The device of claim 1, further comprising an outer tube encompassing said adjustable tube, said outer tube affixed to a first edge of said adjustable tube, and said adjustable tube adjustably positionable with respect to said outer tube. 3. The device of claim 2, wherein said measurement gradient provides an indication of the diameter of the tendon. 4. The device of claim 2, further comprising a ratcheting mechanism operably configured to position said adjustable tube between said open configuration and said closed configuration. 5. The device of claim 4, wherein said ratcheting mechanism comprises a rotatable tube, said rotatable tube affixed to a second edge of said adjustable tube and positioned between said adjustable tube and said outer tube when said adjustable tube is in said open configuration. 6. The device of claim 1, wherein said adjustable tube has a length from said first end to said second end, and wherein said length of said adjustable tube exceeds a length of the tendon under measurement. 7. The device of claim 1, wherein said adjustable tube has a length from said first end to said second end, and wherein said length of said adjustable tube approximates a length of the tendon under measurement. 8. The device of claim 1, wherein said adjustable tube has a length from said first end to said second end, and wherein said length of said adjustable tube equals a length of the tendon under measurement. 9. The device of claim 1, wherein said adjustable tube has a length from said first end to said second end, and wherein said length of said adjustable tube is at least half a length of the tendon under measurement. 10. A system for measuring a diameter of a length of tendon, comprising: a cinchable tube defining an inner region, said cinchable tube configured to move between an open position and a closed position, wherein when in said open position, said cinchable tube is configured to receive at least half the length of the tendon within said inner region and when in said closed position, said cinchable ... | Allosource | A61B005/107 G01B005/08 A61B019/00 A61B005/00
US10881311 | B2 | 2021-01-05 | Apparatus and method for the automated measurement of sural nerve conduction velocity and amplitude | 2015-11-03 | Apparatus for measuring sural nerve conduction velocity and amplitude, the apparatus including a housing; stimulation means for electrically stimulating a human sural nerve; a biosensor comprising a plurality of electrodes for detecting a sural nerve response evoked by the stimulation means; acquisition means electrically connected to the biosensor for electrically acquiring the sural nerve response detected by the biosensor; processing means electrically connected to the acquisition means for digitizing, processing and storing the acquired sural nerve response; calculation means electrically connected to the processing means for calculating the conduction velocity and amplitude of the processed sural nerve response; and display means for displaying the sural nerve conduction velocity and amplitude; wherein the stimulation means and the biosensor are designed to be placed on a patient's anatomy, in the vicinity of a sural nerve. | 1. Apparatus for detecting a nerve response, the apparatus comprising: a biosensor, the biosensor comprising: a substrate;a plurality of electrodes mounted to the substrate for detecting the nerve response;a biosensor reuse code carried by the substrate for determining reuse of the biosensor, wherein the biosensor reuse code is one from a set of biosensor reuse codes, and further wherein the biosensor reuse code carried by the substrate is not unique to that biosensor; anda control unit for detecting reuse of the biosensor when the biosensor is connected to the control unit, wherein the control unit determines the likelihood of biosensor reuse by comparing the biosensor reuse code associated with the biosensor connected to the control unit with biosensor reuse codes acquired by the control unit from previous connections between the control unit and biosensors to determine whether the biosensor reuse code of the biosensor connected to the control unit and the biosensor reuse codes acquired by the control unit from previous connections between the control unit and biosensors form a random pattern conforming with a random distribution of biosensor reuse codes. 2. Apparatus according to claim 1 wherein the biosensor reuse code is encoded on the biosensor by means of electrically passive elements. 3. Apparatus according to claim 2 wherein the passive elements comprise at least two conductive members carried by the substrate, and further wherein the biosensor reuse code is encoded by selectively connecting or not connecting the at least two conductive members. 4. Apparatus according to claim 1 further comprising a processor unit for processing the nerve response acquired by the electrodes. 5. Apparatus according to claim 4 wherein the control unit permanently disables the processor unit when the likelihood of biosensor reuse exceeds a reuse threshold. 6. Apparatus according to claim 4 wherein the control unit permanently disables the processor unit when the biosensor reuse code of the biosensor connected to the control unit and the biosensor reuse codes acquired by the control unit from previous connections between the control unit and biosensors do not form a random pattern conforming with a random distribution of biosensor reuse codes. 7. Apparatus according to claim 1 wherein the biosensor reuse code is randomly assigned to the biosensor. 8. Apparatus according to claim 1 wherein the control unit is configured to update the pattern of biosensor reuse codes with the biosensor reuse code of the biosensor currently connected to the control unit whenever a connection between a biosensor and the control unit is made. 9. A kit comprising: a plurality of biosensors for detecting nerve responses, wherein each of the biosensors comprises a substrate, a plurality of electrodes mounted to the substrate for detecting a nerve response, and a biosensor reuse code carried by the substrate for determining reuse of the biosensor;wherein the biosensor reuse c... | NeuroMetrix, Inc. | A61B005/04 A61B005/05 A61B005/00 A61B005/01
US09351664 | B2 | 2016-05-31 | Implantable device for external urinary control | 2008-10-11 | The present invention relates to an implantable apparatus for obtaining urinary control and emptying of the urinary bladder. The apparatus operates with a powered member (100) operating from the outside of the urinary bladder assisted by a support structure to discharge urine from the urinary bladder. A control device (200) controls the operation of the powered member. The control device further comprises a source of energy for operating the powered member and other energy consuming parts of the apparatus and a control assembly. | 1. An apparatus for treating urinary retention of a mammal patient by discharging urine from the urinary bladder in order to empty the urinary bladder, comprising: an implantable powered member comprising: a first part comprising a urinary bladder surface contact part adapted to exert a force from the outside of the urinary bladder on a selected surface part of the urinary bladder in order to actively compress the urinary bladder for reducing the volume of the urinary bladder to discharge urine from the urinary bladder to empty the urinary bladder, wherein the powered member is adapted to, via the urinary bladder surface contact part, empty the urinary bladder by exerting the force, the force being strong enough to force urine out from the urinary bladder through a urethra to thereby empty the urinary bladder, anda second part being at least a part of a support structure adapted to be fixed to human tissue, wherein the force of the powered member exerted towards the urinary bladder is held at least partially against the support structure, said support structure being adapted to be supported against at least one of the pelvic bone, the pubic bone, sacrum, the spinal cord, the peritoneum, the abdominal wall, the pelvic wall, and the urinary bladder, wherein the force exerted by the urinary bladder contact part to empty the urinary bladder is configured to be held against human tissue via the second part and the support structure adapted to be fixed to human tissue, when the apparatus is implanted, anda control device for controlling the operation of the powered member to discharge urine from the urinary bladder to empty the urinary bladder, thereby control the emptying of the urinary bladder, andwherein the powered member is hydraulically operated to provide compression or release of the urinary bladder. 2. The apparatus according to claim 1, wherein the control device comprises an implanted source of energy for operating the powered member and other energy consuming parts of the apparatus. 3. The apparatus according to claim 2, wherein said control device is adapted to be implanted at least partly subcutaneously or in the abdomen or in the pelvic region. 4. The apparatus according to claim 2, wherein the control device is adapted to be implanted both subcutaneously and in the abdominal cavity, said control device comprising at least two parts adapted to be connected, when implanted. 5. The apparatus of claim 1, comprising an operable pressurizer comprising a reservoir for hydraulic fluid, and wherein the urinary bladder surface contact part comprises an expandable cavity hydraulically connected to the reservoir. 6. The apparatus according to claim 5, wherein the operable pressurizer comprises a pump for transporting the hydraulic fluid form the reservoir to expand the expandable cavity thereby compressing the urinary bladder. 7. The apparatus according to claim 6, wherein the operable pressurizer is adapted to have the hydraulic fluid transp... |  | A61F002/02 A61B005/07 A61B005/02 A61B005/00 A61B017/12 A61B017/135 A61F002/00 A61B017/00 A61B019/00 A61F002/30 A61F002/04
US09326700 | B2 | 2016-05-03 | Catheter display showing tip angle and pressure | 2008-12-23 | A method for displaying information includes receiving a measurement with respect to an invasive probe inside a body of a subject of at least one probe parameter, selected from a group of parameters consisting of a bend angle of the probe and a pressure on the probe. Responsively to the measurement, an icon is displayed on a display screen representing the at least one probe parameter for viewing by an operator of the probe. | 1. A method for displaying information, comprising: receiving a measurement with respect to an invasive probe comprising an insertion tube inside a body of a subject of at least one probe parameter, selected from a group of parameters consisting of a bend angle of the probe and a pressure on the probe, wherein the probe further comprises a resilient joint at a distal end of the insertion tube and a distal tip, wherein the resilient joint comprises a spring-loaded joint configured to permit axial displacement and angular deflection of the distal tip, and wherein the distal tip is configured to bend at the spring-loaded joint, and a magnetic position sensing assembly located within the probe comprising coils on opposite sides of the spring-loaded joint and wherein receiving the measurement comprises measuring the axial displacement and angular deflection of the distal tip due to engagement of tissue in the body by the distal tip by using the magnetic position sensing assembly to sense the position of the distal tip relative to the distal end of the insertion tube; andresponsively to the measurement, displaying an icon on a display screen representing the at least one probe parameter for viewing by an operator of the probe, wherein displaying the icon shows both the bend angle and the pressure, the pressure being an indication and in proportion of both the axial displacement and angular deflection of the distal tip, and wherein the icon is articulated to show the bend angle and is colored to represent whether the pressure is within a pre-determined pressure range. 2. The method according to claim 1, wherein the probe further comprises a catheter, which is inserted into a chamber of a heart of the subject and engages myocardial tissue. 3. The method according to claim 2, further comprising applying energy via the distal tip so as to ablate the myocardial tissue, wherein the operator controls application of the energy responsively to the icon. 4. The method according to claim 1, wherein measuring the deformation further comprises sensing both the deformation of the joint and a position of the probe within the body by transmitting and receiving one or more magnetic fields. 5. The method according to claim 1, wherein displaying the icon further comprises positioning the icon on the display screen so as to represent a location of the probe within the body. 6. The method according to claim 5, wherein positioning the icon further comprises locating the icon on the display screen relative to a map of a surface of an organ of the body, and wherein receiving the measurement comprises measuring the pressure between the probe and the surface. 7. An apparatus for performing a medical procedure, comprising: invasive probe comprising an insertion tube, a resilient joint at a distal end of the insertion tube, and a distal tip, wherein the resilient joint comprises a spring-loaded joint that is configured to permit axial displacement and angular deflection of ... | Biosense Webster (Israel) Ltd. | A61B018/14 A61B005/05 A61B005/06 A61B005/00 A61B019/00
US09330092 | B2 | 2016-05-03 | Systems, methods, apparatus and computer-accessible-medium for providing polarization-mode dispersion compensation in optical coherence tomography | 2012-07-19 | Exemplary systems, apparatus, methods and computer-accessible medium for generating information regarding at least one sample can be provided. For example, it is possible to receiving first data which is based on at least one first radiation provided to the sample(s) and at least one second radiation provided from the sample(s) that is/are associated with the first radiation(s) It is also possible to generate second data by reducing the influence of first optical effects induced on the first radiation(s) prior to reaching the sample(s), and second optical effects induced on the second radiation(s) after leaving the sample(s). | 1. An apparatus for generating information regarding at least one sample, comprising: at least one polarization related data receiver first arrangement which is configured to: receive first data, second data, third data and fourth data, whereini. the first data is associated with a first radiation provided to the at least one sample, a second radiation from the at least one sample, and at least one further radiation provided from the at least one sample or a reference,ii. the second data is associated with a third radiation provided to the at least one sample, the second radiation, and the at least one further radiation,iii. the third data is associated with the first radiation, a fourth radiation from the at least one sample, and the at least one further radiation, andiii. the fourth data is associated with the third radiation, the fourth radiation from the at least one sample, and the at least one further radiation,iv. wherein each of the first and third radiations and the second and fourth radiations have polarizations states are different from one another; anda second arrangement including a computer which is configured to generate fifth data by combining at least two of the first, second, third and fourth data, wherein the combination reduces the influence on the fifth data of a polarization mode dispersion acting on any of the first, second, or third radiations. 2. The apparatus according to claim 1, further comprising at least one optical frequency shifter/delay third arrangement which is configured to generate at least one of (i) an optical frequency shift on the first radiation in relation to the third radiation, (ii) an optical frequency shift on the second radiation in relation to the fourth radiation, (iii) an optical delay of the first radiation in relation to the third radiation, or (iv) an optical delay of the second radiation in relation to the fourth radiation. 3. The apparatus according to claim 1, wherein the at least one first arrangement is further configured to measure characteristics of the at least one sample that influence an optical polarization of a radiation within the at least one sample based on the information. 4. The apparatus according to claim 1, wherein the at least one first arrangement is further configured to receive sixth data which is associated with signals generated by at least one optical device that is provided in an optical path that excludes the at least one sample, and wherein the at least one second arrangement reduces the influences using the sixth data. 5. The apparatus according to claim 1, wherein the at least one further radiation is provided from the reference. 6. The apparatus according to claim 1, wherein the at least one first arrangement is further configured to resolve the first, second, third and fourth data as a function of an optical wavelengths of at least one of the first, second, third or fourth radiations. 7. A non-transitory computer accessible medium which includes instruct... | The General Hospital Corporation | G06F017/30 G01B009/02 A61B003/10 G02B027/48 A61B005/00
US09336355 | B2 | 2016-05-10 | Apparatus and method for generating a condition indication | 2012-02-10 | An apparatus for generating a condition indication using a time sequence of data values, each data value representing a physiological measure of a condition of a subject at a time, includes: a transformer for transforming the time sequence of data values into a transformed sequence of data values using a transform rule, wherein the transform rule is such that a certain characteristic in a time course of the physiological measure is represented by the transformed sequence of data values is more linear than the time course before the transform; a rate of change calculator for calculating an estimated rate of change for the transformed sequence of data values; and a processor for processing the estimated rate of change to output the condition indication. | 1. Apparatus for generating a condition indication using a time sequence of data values, each data value representing a physiological measure of a condition of a subject at a time, wherein the physiological measure is a blood glucose value comprising: a transformer configured to transform the time sequence of data values into a transformed sequence of data values using a transform rule, wherein the transform rule is such that a certain characteristic in a time course of the physiological measure as represented by the transformed sequence of data values is more linear than the time course before the transform, said transformer comprising: an estimated probability function calculator configured to calculate an estimated cumulative distribution function associated with a set of data values; anda transform calculator configured to calculate the transform rule using a predetermined linearly increasing target cumulative distribution function indicating the same probability for a certain data value within the certain characteristic;a rate of change calculator configured to calculate an estimated rate of change for the transformed sequence of data values; anda processor configured to process the estimated rate of change to output the condition indication,wherein at least one of the transformer, rate of change calculator and processor include a hardware implementation. 2. Apparatus in accordance with claim 1, in which the processor is configured to generate a graphic, audible or tactile display of the estimated rate of change or to generate an electrical, magnetic or electro-magnetic signal representing the estimated rate of change. 3. Apparatus in accordance with claim 1, in which the transformed sequence of data values has a transformed value for a current time instant, and wherein the processor is configured to predict a data value for a later time instance from the estimated rate of change, in which the processor is configured to evaluate, whether the predicted data value exceeds a high alarm threshold or is below a low alarm threshold, and in which the processor is configured to generate a visual, audible, tactile, mechanical, electric or magnetic alarm indication. 4. Apparatus in accordance with claim 1, in which the transformer comprises a memory configured to store the transform rule as parameters for a parameterized function, or as a look-up table in which non-transform data values have associated transformed data values. 5. Apparatus in accordance with claim 1, in which the time sequence of data values comprises glycemic data values indicating a glycemic condition for a time instance of the subject. 6. Apparatus in accordance with claim 1, in which the certain characteristic is represented by a characteristic data record, the characteristic data record comprising value data derived by average rates of change for different value ranges in a measurement range, where the value data have associated time instants, andin which the rate of change ... | DIABETES TOOLS SWEDEN AB | G01N033/48 G06F019/00 A61B005/145 A61B005/00
US09351705 | B2 | 2016-05-31 | Miniaturized photoacoustic imaging apparatus including a rotatable reflector | 2010-01-08 | A reflection-mode photoacoustic endoscope includes a light source configured to emit a light pulse, a signal detection or transmission unit configured to detect or emit an ultrasonic pulse, and a rotatable reflector. The rotatable reflector is configured to reflect at least one of the light pulse and the ultrasonic pulse into a target area of an object, and reflect a response signal to the signal detection unit. The response signal is one of a photoacoustic wave generated by the object responsive to the light pulse and an ultrasonic pulse echo generated by the object responsive to the ultrasonic pulse. | 1. A reflection-mode photoacoustic endoscope comprising: a tube;a light source configured to emit a light pulse suitable for photoacoustic imaging;a signal detection or detection and transmission unit configured to detect or emit and detect an ultrasonic pulse; anda reflector rotatable relative to said tube and configured to: reflect the light pulse and the ultrasonic pulse into a target area of an object; andreflect a response signal to said signal detection unit, the response signal being one of a photoacoustic wave generated by the object responsive to the light pulse and an ultrasonic pulse echo generated by the object responsive to the ultrasonic pulse. 2. A reflection-mode photoacoustic endoscope in accordance with claim 1, further comprising a micromotor coupled at a first end to said reflector, said micromotor configured to impart a rotational force on said reflector. 3. A reflection-mode photoacoustic endoscope in accordance with claim 2, wherein said micromotor comprises a galvanometer lead. 4. A reflection-mode photoacoustic endoscope in accordance with claim 2, wherein said micromotor comprises a plurality of gears. 5. A reflection-mode photoacoustic endoscope in accordance with claim 2, wherein said micromotor is coupled at a second end to a gear mechanism. 6. A reflection-mode photoacoustic endoscope in accordance with claim 1, further comprising a micromotor and a magnet assembly, said magnet assembly coupled to said reflector. 7. A reflection-mode photoacoustic endoscope in accordance with claim 6, wherein said magnet assembly comprises a first magnet coupled to said micromotor and a second magnet positioned with respect to said first magnet and coupled to said reflector, said micromotor configured to generate torque, said first magnet configured to convert the torque into a magnetic field, said second magnet configured to rotate said reflector based on the magnetic field. 8. A reflection-mode photoacoustic endoscope in accordance with claim 1, wherein said signal detection unit comprises an ultrasonic transducer, said reflector comprises a parabolic mirror configured to reflect and focus the response signal onto said ultrasonic transducer. 9. A reflection-mode photoacoustic endoscope in accordance with claim 1, wherein said reflector comprises a plurality of reflective surfaces. 10. A reflection-mode photoacoustic endoscope in accordance with claim 1, wherein said endoscope is sized to fit within a catheter. 11. An endoscopy system comprising: a reflection-mode photoacoustic endoscope comprising: a tube;a light source configured to emit a light pulse suitable for photoacoustic imaging;a signal detection or transmission unit configured to detect or emit an ultrasonic pulse; anda mirror rotatable relative to said tube configured to reflect the light pulse and the ultrasonic pulse into a target area of an object, and reflect a response signal to said signal detection unit, the response signal being one of a photoacoustic w... | Washington University | A61B008/08 A61B008/00 A61B005/00 A61B008/12 A61B005/08
US20160120443 | A1 | 2016-05-05 | Single Chamber Volume Measurement Apparatus and Methods of Making and Using the Same | 2015-09-28 | A single chamber system or apparatus configured to measure a volume of a person, animal, and/or object, and methods of manufacturing and using the same, are disclosed. The system/apparatus includes a chamber having a door thereon or affixed thereto, and a fixed volume therein when the door is closed; an oscillating membrane or diaphragm in a wall of the chamber; a pressure sensor and/or a pressure transducer configured to measure pressure fluctuations in the chamber; and an oscillation amplitude detector coupled to the oscillating membrane or diaphragm. The chamber generally has a volume sufficient to enclose the person, animal, or object therein. The volume of the chamber generally does not vary with relatively small pressure changes or fluctuations. | 1. A system or apparatus, comprising: a) a chamber having a door thereon or affixed thereto, and a fixed volume therein when the door is closed;b) an oscillating membrane or diaphragm in a wall of the chamber or door;c) a pressure sensor and/or a pressure transducer configured to measure pressure fluctuations in the chamber; andd) an oscillation amplitude detector coupled to the oscillating membrane or diaphragm. 2. The system or apparatus of claim 1, further comprising a computer or logic configured to determine a value proportional to the pressure amplitude in the chamber resulting from oscillations of the oscillating membrane or diaphragm and a value proportional to the volume displacement from oscillation or position amplitudes of the oscillating membrane or diaphragm with and without a subject or object in the chamber. 3. The system or apparatus of claim 1, wherein the chamber has a volume sufficient to enclose a person, animal, or object therein, and the system or apparatus is configured to measure a volume of the person, animal, or object. 4. The system or apparatus of claim 1, wherein the chamber comprises one or more transparent walls. 5. The system or apparatus of claim 1, wherein the chamber comprises (i) a first section comprising the door and (ii) a second, stationary section configured to house or contain a subject or object. 6. The system or apparatus of claim 3, wherein the volume of the chamber does not vary with relatively small pressure changes or fluctuations. 7. The system or apparatus of claim 1, wherein the oscillating membrane or diaphragm comprises a rigid oscillating membrane or diaphragm. 8. The system or apparatus of claim 1, wherein the oscillating membrane or diaphragm oscillates at a frequency of from 1 to 100 Hz. 9. A system or apparatus of claim 1, wherein the oscillating membrane or diaphragm has an area of from 10 to 5000 cm2. 10. A system or apparatus of claim 1, wherein the oscillation amplitude detector comprises (i) an optical light source and a light detector or (ii) an accelerometer. 11. A system or apparatus of claim 1, wherein the oscillation amplitude detector is configured to measure movement and/or displacement of the oscillating membrane or diaphragm. 12. A system or apparatus of claim 1, wherein the system or apparatus is configured to calculate a percentage or other proportion of body fat of a person and/or animal. 13. The system or apparatus of claim 2, wherein the computer or logic is on the exterior of the apparatus, and is configured to control operations of the apparatus and perform calculations to determine body volume, body fat content, and/or weight of a person or animal. 14. A method of determining a volume of a person, animal, or object in a single chamber volume measuring apparatus, comprising: a) measuring a first oscillation or position amplitude of an oscillating membrane or diaphragm in the apparatus without the person, animal, or object therein;b) placing the person, an... |  | A61B005/107 A61B005/00
US10478088 | B2 | 2019-11-19 | Disposable contamination prevention liners for MRI machines | 2015-11-12 | A pack of disposable multi-layered liner sheets suitable for covering non-planar surfaces of medical scanners having a long, hard to clean, bore or tunnel, particularly MRI scanners, to establish a sterile patient environment. Liner sheets are connected end-to-end in a stack by an adhesive or a linear array of perforations to respective succeeding and preceding sheets of the pack. In one embodiment, the pack is secured inside the bore by pegs on the top side of the pack engaging strategically placed clips attached to the bore wall. In other embodiments, the pack is secured inside the bore by opposing sides of a backing layer extending laterally beyond the liner sheets to engage a groove in guide rails affixed inside the bore. | 1. A liner pack for mounting on a medical device having a bore with a curved bore wall surface and an open front end, the liner pack comprising: a backing layer having a top side and a bottom side;a stack of disposable multi-layered liner sheets secured to the bottom side of the backing layer; andmounting means for supporting the liner pack within the bore in a generally curved configuration to substantially cover the bore wall surface,wherein the stack of disposable multi-layered liner sheets comprises a series of end-to-end connected sheets which are configured to conform to and cover the curved bore wall surface substantially in its entirety, andwherein each sheet of the series of end-to-end connected liner sheets is removably attached at a leading end to a leading end of respective succeeding and preceding liner sheets in the stack by a straight line of perforations and at a trailing end to a trailing end of respective succeeding and preceding liner sheets in the stack by an adhesive. 2. The liner pack of claim 1, wherein the mounting means comprises attachment devices on the top side of the backing layer configured to cooperate with locking receptacles strategically secured to the curved bore wall surface to support the liner pack substantially flush against the bore wall surface in its generally curved configuration. 3. The liner pack of claim 2, wherein the liner pack has a leading end and a trailing end; andwherein the attachment devices are positioned on the top side of the backing layer adjacent the leading end and the trailing end of the liner pack. 4. The liner pack of claim 2, wherein the backing layer has a leading end and a trailing end,wherein the trailing end has a notch-shaped indentation extending inwardly from a trailing end outer edge to a trailing end inner edge, andwherein the attachment devices are positioned on the top side of the backing layer adjacent an edge of the leading end and adjacent the outer and inner edges of the trailing end. 5. The liner pack of claim 1, wherein the backing layer has a trailing end which includes a notch-shaped indentation extending inwardly from a trailing end outer edge to a trailing end inner edge, andwherein segments of the trailing end on opposite sides of the indentation define a pair of opposing flaps extending from the inner edge to the outer edge and configured to engage a front face of the medical device adjacent the open front end of the bore. 6. The liner pack of claim 5, wherein the mounting means comprises an attachment device secured to a top side of each flap adjacent its outer edge and configured to cooperate with a locking receptacle strategically attached to a front face of the medical device adjacent the open front end of the bore. 7. The liner pack of claim 1, wherein the backing layer comprises a trailing end and a pair of opposing flaps extending outwardly from the trailing end, andwherein the mounting means comprises a locking device attached to a top surface of... | - | A61B006/00 A61B005/055 A61B046/10 A61B046/20
US09351661 | B2 | 2016-05-31 | Devices, systems, and methods to measure parallel tissue conductance, luminal cross-sectional areas, fluid velocity, and/or determine plaque vulnerability using temperature | 2014-04-22 | Devices, systems, and methods to measure parallel tissue conductance, luminal cross-sectional areas, fluid velocity, and/or determine plaque vulnerability using temperature. In at least one embodiment of a method to obtain parallel tissue conductance, the method comprises the steps of inserting at least part of a detection device into a luminal organ, applying current thereto, obtaining a native temperature measurement, injecting a solution of a known conductivity into the luminal organ, detecting a temperature change indicative of the fluid within the luminal organ, measuring an output conductance, and calculating a parallel tissue conductance based upon the output conductance and the conductivity of the injected solution. | 1. A method to obtain conductance data within a luminal organ, the method comprising the steps of: operating a detection device at least partially positioned within a luminal organ at a first location to obtain a first fluid temperature measurement;detecting a temperature change at the first location using the detection device, the temperature change indicative of an injection of a bolus of a solution having a known conductivity into the luminal organ; andmeasuring multiple output conductances using at least two impedance electrodes of the detection device prior to dilution of the bolus at the first location at least until a threshold temperature is reached. 2. The method of claim 1, further comprising the step of: calculating a parallel tissue conductance at the first location based in part upon the temperature change, at least one output conductance of the multiple output conductances, and the conductivity of the injected solution. 3. The method of claim 1, further comprising the steps of: calculating a cross-sectional area of the luminal organ at the first location based in part upon at least one output conductance of the multiple output conductances;selecting an appropriately-sized stent based upon the cross-sectional area; andimplanting the stent into the luminal organ at the first location. 4. The method of claim 1, wherein the first fluid temperature measurement is a measurement of blood obtained using a thermistor of the detection device. 5. The method of claim 1, wherein the temperature change is based upon a temperature measurement at least 5° C. lower than the first fluid temperature measurement. 6. The method of claim 1, wherein the threshold temperature is within 2° C. of the first fluid temperature measurement. 7. The method of claim 1, further comprising the step of: obtaining a native conductance of fluid present within the luminal organ prior to injection of the bolus. 8. The method of claim 1, further comprising the step of: calibrating the detection device using at least one phantom immersed in a solution bath of the same solution injected into the luminal organ, wherein the solution within the solution bath is at a known temperature. 9. The method of claim 1, wherein the detection device comprises an inflatable balloon along a longitudinal axis of the detection device. 10. The method of claim 9, further comprising the step of: inflating the balloon to breakup materials causing a stenosis within the luminal organ. 11. The method of claim 9, wherein the detection device further comprises a stent located over the balloon, the stent capable of being distended to a desired lumen size and implanted into the luminal organ. 12. A device configured to obtain conductance data within a luminal organ, comprising: an elongated body having a longitudinal axis extending from a proximal end to a distal end;at least two electrodes located on the elongated body, the at least two electrodes configured to obtain conductance data within... | 3DT Holdings, LLC | A61B005/05 A61B005/053 A61B005/02 A61B005/028 A61B005/00 G01K013/00 A61B005/01 A61B005/107 A61F002/958
US10631821 | B2 | 2020-04-28 | Rib blockage delineation in anatomically intelligent echocardiography | 2014-06-18 | An interactive visual guidance tool for an imaging acquisition and display system and configured for user navigation with respect to a blockage of a field of view (216) detects, and spatially defines, the blockage. It also integrates, with the image for joint visualization, an indicium (244) that visually represents the definition. The indicium is moved dynamically according to movement, relative to the blockage, of the field of view. The indicium can be shaped like a line segment, or two indicia (244, 248) can be joined in a “V” shape to frame a region of non-blockage. The defining maybe based on determining whether ultrasound beams in respective directions are blocked. Included, for deriving the image, in some embodiments are imaging channels for receiving image data for which a metric of coherence, i.e., similarity among channel data, is computed. The determination for a direction is based on the metric for locations in that direction. One application is navigating an ultrasound probe between blocking ribs (208, 212) to achieve a standard cardiac view. | 1. An ultrasound system, comprising: an ultrasound probe comprising a transducer array, wherein the transducer array is configured to obtain ultrasound data representative of an anatomy within a field of view; anda processor in communication with a display and the ultrasound probe,wherein the processor is configured to receive, via a plurality of data channels, the ultrasound data obtained by the ultrasound probe,wherein the processor is configured to generate an ultrasound image of the anatomy within the field of view based on the ultrasound data,wherein the processor is configured to detect and spatially define a blockage region of the ultrasound image comprising a portion of the field of view and associated with an anatomical structure within the anatomy based on the ultrasound data,wherein the processor is configured to generate a graphical representation based on detecting and spatially defining the blockage region,wherein the processor is configured to integrate the graphical representation with the ultrasound image to form a joint visualization,wherein the processor is configured to output, to the display, the joint visualization of the ultrasound image and the graphical representation,wherein the ultrasound image includes the field of view,wherein the joint visualization comprises the graphical representation overlaid on the anatomy in the ultrasound image to identify the portion of the field of view corresponding to the blockage region in the ultrasound image,wherein the processor is configured to move the joint visualization dynamically in response to a movement of the field of view relative to the blockage,wherein at least one of the ultrasound probe or the processor is configured to apply delays to the ultrasound data received via the plurality of data channels, andwherein the processor is configured to detect and spatially define the blockage region by computing a metric of similarity for the ultrasound data to which the delays have been applied. 2. The system of claim 1, wherein the processor is further configured to modify the graphical representation such that the graphical representation continues to identify the blockage region through the movement of the field of view. 3. The system of claim 1, wherein transducer array is configured to obtain the ultrasound data by activating a plurality of apertures; wherein the metric of similarity is representative of a correlation between beamsummed data of different apertures; andwherein the processor is configured to divide the apertures among the plurality of data channels. 4. The system of claim 3, wherein at least a portion of the apertures are interleaving complementary apertures. 5. The system of claim 4, wherein the interleaving complementary apertures are associated with a randomly divided portion of the plurality of data channels. 6. The system of claim 4, wherein the processor is configured to generate the ultrasound image from a first subset of the ultrasound data, wherein ... | KONINKLIJKE PHILIPS N.V. | A61B008/00 A61B008/08 G01S007/52 A61B006/00 G01S015/89 A61B005/044
US09326722 | B2 | 2016-05-03 | Methods of predicting and monitoring labor | 2012-03-21 | The invention includes, in part, methods and apparatus for determining the status of labor in a pregnant subject. In some embodiments of the invention, organ-level uterine function is determined as a measure of the status of labor such as non-labor (e.g. false labor), latent phase labor, active phase labor, or post-delivery. | 1. A method of determining labor status in a pregnant subject, the method comprising: receiving, at an electromyography (EMG) recorder, a plurality of signals, each of the plurality of signals comprising an EMG signal generated by different regions of a subject's uterine wall and measured with a respective one of a plurality of electrode pairs, each of the electrode pairs positioned on the subject's abdomen in a manner minimizing or eliminating overlap of the signal of the plurality of signals measured by two or more of the positioned electrode pairs;detecting three or more uterine muscle contractions, each of the three or more detected uterine muscle contractions associated with a respective one of the plurality of signals;calculating a synchronization factor of the three or more uterine muscle contractions based on a ratio of a summed duration of each of the three or more detected uterine muscle contractions to a total time spanned by the three or more detected uterine muscle contractions; andidentifying, responsive to the synchronization factor crossing a predetermined threshold, a labor status of the subject. 2. The method of claim 1, further comprising comparing the synchronization factor of the three or more uterine muscle contractions to a control regional synchronization pattern. 3. The method of claim 1, wherein the method further comprises: measuring a strength and a frequency of the three or more uterine muscle contractions; andcalculating the synchronization factor responsive to the strength and the frequency of the three or more uterine muscle contractions. 4. The method of claim 1, further comprising positioning each of the plurality of electrode pairs on a surface of the subject's abdomen. 5. The method of claim 1, wherein each of the plurality of electrode pairs comprise a ground electrode and a recording electrode. 6. The method of claim 1, wherein the labor status is absent, latent phase labor, active phase labor, delivery of the baby, or afterbirth-delivery. 7. The method of claim 1, wherein the subject has been, is currently, or will be administered one or more labor-enhancing treatments. 8. The method of claim 1, wherein the subject has been, is currently, or will be administered one or more labor-inhibiting treatments. 9. The method of claim 1, further comprising receiving the plurality of signals from a remote device. 10. The method of claim 1, wherein the labor status is one of true labor and false labor. 11. The method of claim 1, further comprising diagnosing a medical condition in the subject responsive to the labor status. 12. The method of claim 11, wherein the medical condition is false labor. 13. The method of claim 11, wherein the medical condition is premature labor. 14. The method of claim 1, further comprising selecting a treatment for the subject based at least in part on the identified labor status. | University of Vermont and State Agricultural College | A61B005/00 A61B005/0488 A61B005/0492
US20160128590 | A1 | 2016-05-12 | ELECTRONIC APPARATUS HAVING MEASUREMENT ELECTRODES | 2015-03-02 | A measurement device includes a housing, an electronic unit included in the housing and configured to measure a biological state of a subject, and an electrode to be placed on the subject for measuring the biological state. The electrode is embedded in the housing and has a surface that is exposed through an opening in the housing. | 1. A measurement device comprising: a housing;an electronic unit included in the housing and configured to measure a biological state of a subject; andan electrode to be placed on the subject for measuring the biological state, whereinthe electrode is embedded in the housing and has a surface that is exposed through an opening in the housing. 2. The measurement device according to claim 1, wherein the electrode has a first portion and a second portion that has a smaller outer perimeter than the first portion, andthe exposed surface is a surface of the second portion. 3. The measurement device according to claim 1, wherein the exposed surface is smoother than the other surfaces of the electrode. 4. The measurement device according to claim 1, wherein the electrode has a concave portion that recedes from the exposed surface. 5. The measurement device according to claim 1, wherein the exposed surface is flush with the surface of the housing. 6. The measurement device according to claim 1, further comprising: a second electrode to be placed on the subject for measuring the biological state, wherein the second electrode is electrically connected to the electronic unit and embedded in the housing and has a surface that is exposed through another opening in the housing. 7. The measurement device according to claim 6, wherein the surface of the housing has a first corner region, a second corner region, a third corner region, and a fourth corner region that is diagonal to the first corner region, andthe first electrode is located closest to the first corner region and the second electrode is located closest to the fourth corner region. 8. The measurement device according to claim 1, further comprising: a substrate having a first surface on which the electronic unit is disposed and a second surface that is opposite to the first surface and with which the first and second electrodes are in electrical contact. 9. The measurement device according to claim 1, wherein the electronic unit and the substrate are enclosed entirely within the housing. 10. The measurement device according to claim 1, wherein the housing is formed of a flexible material. 11. A measurement device comprising: a housing having an opening;an electronic unit included in the housing and configured to measure a biological state of a subject; andan electrode to be placed on the subject for measuring the biological state, wherein the electrode is disposed in the opening of the housing. 12. The measurement device according to claim 11, wherein a surface of the electrode is exposed to an outside of the measurement device on a surface of the housing. 13. The measurement device according to claim 12, wherein the electrode has a first portion and a second portion that has a smaller outer perimeter than the first portion, andthe exposed surface is a surface of the second portion. 14. The measurement device according to claim 12, wherein the exposed surface is smoother than the other sur... |  | A61B005/0408 H01M004/00
US10485464 | B2 | 2019-11-26 | Spectrometric apparatus and storage case | 2015-10-29 | A biological body inspection apparatus includes a measurement module that acquires an amount of light having a wavelength under measurement contained in light incident on the measurement module and an enclosure that accommodates the measurement module and has a window that transmits light traveling toward the measurement module. An adhesive member is provided on a surface of the enclosure at least in an area thereof that surrounds the window. The adhesive member has a light blocking section that is located in an area outside the window and surrounds the window in a plan view in the direction along the optical axis of the light traveling toward the measurement module and blocks light that belongs to a measurement wavelength region within which at least the wavelength under measurement is present. | 1. A spectrometric apparatus comprising: a measurement module that acquires an amount of measurement light having a measurement wavelength contained in light incident on the measurement module;an enclosure that accommodates the measurement module and has a window that transmits the light traveling toward the measurement module; andan adhesive member that is provided on a surface of the enclosure at least in a first area that surrounds the window,wherein the adhesive member has a light blocking section that is located in a second area in the first area, the second area is located outside the window and surrounds the window in a plan view, and the light blocking section blocks the measurement light, andthe light blocking section further comprises a conductive ground electrode of the spectrometric apparatus. 2. The spectrometric apparatus according to claim 1, wherein the light blocking section is provided inside an outer circumferential edge of the adhesive member in the plan view. 3. The spectrometric apparatus according to claim 2, wherein the light blocking section is provided along an outer edge of the window in the plan view. 4. The spectrometric apparatus according to claim 1, wherein the adhesive member has a light transmissive section that transmits at least the measurement light, and the light blocking section and the light transmissive section are formed integrally with each other. 5. The spectrometric apparatus according to claim 1, wherein the adhesive member has a light transmissive section that is provided in a position where the light transmissive section overlaps with the window in the plan view and transmits at least the measurement light,the window has an opening provided in the enclosure and a light transmissive member disposed in the opening, andthe light transmissive section has a refractive index that differs from a refractive index of the light transmissive member by an amount smaller than or equal to a predetermined threshold. 6. The spectrometric apparatus according to claim 1, further comprising a circuit substrate provided with a controller that controls the spectrometric apparatus,wherein the light blocking section contains carbon-based conductive fillers so that the light blocking section has conductivity, andthe enclosure has a wiring electrically connecting between the light blocking section and the circuit substrate. 7. The spectrometric apparatus according to claim 1, wherein the measurement module further includes: a light source configured to irradiate a subject with light;a spectroscopic filter configured to select light of a predetermined wavelength from reflected light from the subject so as to transmit the selected light therethrough, and the spectroscopic filter is configured to change a wavelength of the light to be outputted; anda light reception sensor configured to receive the light from the spectroscopic filter. | Seiko Epson Corporation | A61B005/1455 G01J003/02 A61B005/00 G01J003/26 G01N021/01 G01N021/31
US09351679 | B2 | 2016-05-31 | Cartridge and meter for receiving the cartridge | 2013-01-09 | The specification describes a cartridge for insertion into a meter, the interior of the meter comprising a boss having a main body and a protrusion extending from the main body. The cartridge comprises an orientation member provided at one end of the cartridge, the orientation member comprising a recess having a cross-sectional shape corresponding to that of the boss and configured to receive the boss therein, the orientation member further comprising a ramp surrounding at least part of the portion of the recess corresponding to the main body, the portion of the recess which corresponds to the protrusion adjoining the lower end of the ramp via a step. | 1. A cartridge for insertion into a meter, the interior of the meter comprising a boss having a main body and a protrusion extending from the main body, the cartridge comprising: an orientation member provided at one end of the cartridge, the orientation member comprising a recess having a cross-sectional shape corresponding to that of the boss and configured to receive the boss therein, the orientation member further comprising a ramp surrounding a majority portion of the recess corresponding to the main body, a second portion of the recess corresponds to the protrusion and adjoins the lower end of the ramp via a first step and adjoins the higher end of the ramp via a second step that is a greater height than the first step, wherein a second portion of the recess has a base that is defined by a distance between the first step and second step such that the protrusion is received in the second portion when the cartridge is fully inserted into the meter. 2. The cartridge of claim 1, further comprising a shaft, the orientation member being provided at one end of the shaft. 3. The cartridge of claim 2, further comprising a housing, the shaft being provided within the housing and being rotatable relative to the housing. 4. A system comprising: a meter comprising a drive shaft with a boss at one end of the drive shaft, the boss having a main body and a protrusion; anda cartridge comprising an orientation member provided at one end of the cartridge, the orientation member comprising a recess having a cross-sectional shape corresponding to that of the boss and configured to receive the boss therein, the orientation member further comprising a ramp surrounding a majority portion of the recess corresponding to a main body and a second portion of the recess which corresponds to the protrusion and adjoins a lower end of the ramp via a step. 5. A system comprising: a cartridge; anda meter for receiving the cartridge therein, wherein one of the meter and the cartridge comprises a boss having a main body and a protrusion extending from the main body, and the other one of the meter and the cartridge comprises an orientation member comprising a recess having a cross-sectional shape corresponding to that of the boss and configured to receive the boss therein, the orientation member further comprising a ramp surrounding a majority portion of the recess corresponding to the main body, the portion of the recess which corresponds to the protrusion adjoining the lower end of the ramp via a first step, wherein one of the boss and the orientation member is provided at one end of the cartridge and the other one of the boss and the orientation member is provided in the interior of the meter, the portion of the recess which corresponds to the protrusion adjoining the higher end of the ramp via a second step that is a greater height than the first step, wherein a second portion of the recess has a base that is defined by a distance between the first step and second step ... | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/151 A61B005/15 A61B005/145
US09339263 | B2 | 2016-05-17 | Dilation system and method | 2014-01-03 | A dilator system includes a series of sequential dilators. A first dilator has an elongated body having a major axis and a minor axis, the minor axis being substantially normal to the major axis. A second dilator has an elongated body having a major axis and a minor axis, the minor axis of the second dilator being substantially normal to the major axis of the second dilator. A bore of the second dilator is configured to matingly receive the first dilator in such a way that the major axis of the second dilator is aligned with the minor axis of the first dilator and the minor axis of the second dilator is aligned with the major axis of the first dilator. | 1. A dilation system, comprising: a first dilator including an elongated body having longitudinal axis, a proximal end, a distal end, and an outer surface extending between the proximal end and the distal end, the first dilator having a major axis and a minor axis, the minor axis being substantially normal to the major axis; anda second dilator including an elongated body having longitudinal axis, a proximal end, a distal end, an outer surface extending between the proximal end and the distal end, and a bore extending from the proximal end to the distal end, the second dilator having a major axis and a minor axis, the minor axis of the second dilator being substantially normal to the major axis of the second dilator, the bore of the second dilator configured to matingly receive the first dilator in such a way that the major axis of the second dilator is aligned with the minor axis of the first dilator and the minor axis of the second dilator is aligned with the major axis of the first dilator. 2. The dilation system of claim 1, wherein the elongated body of the first dilator has a bore extending from the proximal end to the distal end, and wherein the bore of the first dilator is aligned with the longitudinal axis of the first dilator. 3. The dilation system of claim 2, wherein the bore of the second dilator is aligned with the longitudinal axis of the second dilator. 4. The dilation system of claim 1, wherein the elongated body of the first dilator has a bore extending from the proximal end to the distal end, and wherein the bore of the first dilator is offset relative to the longitudinal axis of the first dilator. 5. The dilation system of claim 4, wherein the bore of the second dilator is aligned with the longitudinal axis of the second dilator. 6. The dilation system of claim 1, wherein the elongated body of the first dilator has at least one channel for receiving a stimulating probe, the channel extending from the proximal end to the distal end and intersecting the major axis of the first dilator. 7. The dilation system of claim 6, wherein the elongated body of the second dilator has at least one channel for receiving a stimulating probe, the channel extending from the proximal end to the distal end and intersecting the major axis of the second dilator. 8. The dilation system of claim 1 further comprising: a third dilator including an elongated body having longitudinal axis, a proximal end, a distal end, an outer surface extending between the proximal end and the distal end, and a bore extending from the proximal end to the distal end, the third dilator having a substantially circular transverse cross section, the bore of the third dilator configured to matingly receive the second dilator. 9. The dilation system of claim 8, wherein the elongated body of the third dilator has at least one channel for receiving a stimulating probe, the channel extending from the proximal end to the distal end. 10. The dilation system of claim 1, where... | DePuy Synthes Products, Inc. | A61B017/00 A61B017/02 A61B001/018 A61B001/32 A61B005/00 A61B017/34 A61B005/04 A61M029/00
US09333411 | B2 | 2016-05-10 | Exercise supporting device, exercise supporting method and exercise supporting program | 2013-08-27 | An exercise supporting device including a detecting section which repeatedly detects motion data relative to the exercise motion status of a user during the exercise motion; an analyzing and judging section which obtains, based on the motion data, a first motion information value corresponding to the user's moving speed and a second motion information value corresponding to the user's footstep count per unit time or footstep width at each time the detecting section detects the motion data, and judges whether the first motion information value is out of a first numerical value range and whether the second motion information value is out of a second numerical value range; and an output section which, when at least one of the motion information values is judged as being out of the relevant numerical value range, performs an informing operation regarding this motion information value. | 1. An exercise supporting device comprising: a detecting section which repeatedly detects motion data related to an exercise motion status of a user who is making an exercise motion by a moving motion;an analyzing and judging section which obtains, based on the detected motion data, a movement distance from a moving motion start point of the user, a first motion information value corresponding to a moving speed of the user and a second motion information value corresponding to a footstep count of the user per unit time or a footstep width of the user, at each time the detecting section detects the motion data, wherein a plurality of movement distance ranges respectively corresponding to different movement distances have a plurality of respective first numerical value ranges in each of which a plurality of values are set as an allowable range of the first motion information value and a plurality of respective second numerical value ranges in each of which a plurality of values are set as an allowable range of the second motion information value,wherein the analyzing and judging section judges whether or not the first motion information value is out of a specific first numerical value range set for a specific movement distance range in which the movement distance when the first motion information value is obtained is included from among the plurality of first numerical value ranges, at each time the first motion information value is obtained, andwherein the analyzing and judging section judges whether or not the second motion information value is out of a specific second numerical value range set for the specific movement distance range from among the plurality of second numerical value ranges, at each time the second motion information value is obtained; andan output section which, when at least one of the first motion information value and the second motion information value is judged by the analyzing and judging section as being out of a relevant one of the specific first numerical value range and the specific second numerical value range, performs an informing operation regarding the first motion information value or the second motion information value judged as being out of the relevant one of the specific first numerical value range and the specific second numerical value range, during the exercise motion of the user,wherein the output section includes a display section which displays an image including character information, and wherein the output section: performs, as the informing operation, a standard display on the display section with a first display format in which a display color of a background portion of the display section is set to a first background color, and character information included in a first image corresponding to the first motion information value and the second motion information value is displayed on the display section with a first size, when the analyzing and judging section judges that the first motion information... | CASIO COMPUTER CO., LTD. | G06F017/00 A63B071/06 A61B005/11 G06Q010/06 A61B005/00
US09339241 | B2 | 2016-05-17 | Assessment and prediction of cardiovascular status during cardiac arrest and the post-resuscitation period using signal processing and machine learning | 2012-05-25 | Real-time, short-term analysis of ECG, by using multiple signal processing and machine learning techniques, is used to determine counter shock success in defibrillation. Combinations of measures when used with machine learning algorithms readily predict successful resuscitation, guide therapy and predict complications. In terms of guiding resuscitation, they may serve as indicators and when to provide counter shocks and at what energy levels they should be provided as well as to serve as indicators of when certain drugs should be provided (in addition to their doses). For cardiac arrest, the system is meant to run in real time during all current resuscitation procedures including post-resuscitation care to detect deterioration for guiding care such as therapeutic hypothermia. | 1. A computer-implemented method for automated monitoring and online assessment of chances of survival for a patient in cardiac arrest, the method comprising: obtaining an ECG signal from the patient;preprocessing the ECG signal to remove high frequency noise and baseline jumps caused by noise and interference;performing, in a first processor, non-linear characterization of the preprocessed ECG signal and calculating a prototype distance of the preprocessed ECG signal;performing, in a second processor, feature extraction of the preprocessed ECG signal with a complex wavelet transform;performing, in a third processor, attribute extraction from the preprocessed ECG signal;performing, in a fourth processor, attribute extraction from an end tidal CO2 (ETCO2) signal;receiving distance values from non-linear characterization of the preprocessed ECG signal, extracted features of the preprocessed time-series ECG signal, attributes extracted from the ETCO2 signal, and attributes extracted from Dual-Tree Complex Wavelet Decomposition of the pre-processed ECG signal, and performing a feature selection using the received data with a predictive model;using machine learning to classify results of the feature selection process;generating a shock success prediction, which results in return of spontaneous circulation (ROSC);generating decompensation and re-arrest prediction; andrecommending therapeutic alternatives and medications for guiding therapy. 2. The method of claim 1, further comprising integrating categorical data, comprising at least one of demographics, medical history and medication information, into the feature database. 3. The method of claim 1, further comprising performing attribute extraction from other continuous physiologic signals, wherein the continuous physiologic signals comprises at least one of tissue impedance, vascular waveform data from piezoelectric sensors or other devices capable of obtaining vascular waveform data, tissue oxygenation signals from near infrared spectroscopy and other tissue oxygenation devices, and wherein the step of receiving includes receiving attributes extracted from other physiologic signals and categorical variables. 4. A system for automated monitoring and online assessment of chances of survival for a patient in cardiac arrest, the system comprising: an ECG device for providing an ECG signal from the patient;a filter for preprocessing the ECG signal to remove high frequency noise and baseline jumps caused by noise and interference;a first signal processor for performing non-linear characterization of the preprocessed ECG signal and calculating a prototype distance of the preprocessed ECG signal;a second signal processor for performing feature extraction of the preprocessed ECG signal with complex wavelet transform;a third signal processor for performing attribute extraction from the preprocessed ECG signal;a fourth signal processor for performing attribute extraction from an end tidal CO2 (ETCO2) signal... | Virginia Commonwealth University | A61B005/00 A61N001/37 A61B005/083 G06F019/00 A61B005/0452 A61B005/0205 A61B005/04 A61B005/053 A61B005/1455 A61B005/0402 A61N001/39
US09345420 | B2 | 2016-05-24 | Magnetic resonance imaging apparatus | 2013-02-14 | A magnetic resonance imaging apparatus according to an embodiment is a magnetic resonance imaging apparatus that performs a contrast medium imaging, comprising a time management part that measures and manages a lapse time after injection of a contrast medium, a setting part that sets a start time of a second-half imaging in the form of a lapse time after injection of contrast medium, the second-half imaging being performed a predetermined idle time after a first-half imaging and the first-half imaging being performed after injection of the contrast medium, and a display part that displays at least the lapse time measured by the time management part and the start time set by the setting part. | 1. A magnetic resonance imaging apparatus that performs a contrast medium imaging, comprising: circuitry configured to measure and manage a lapse time after injection of a contrast medium,divide an original imaging series performing the contrast medium imaging and including (i) a first-half imaging performed as an early-phase contrast imaging after the injection of the contrast medium, (ii) an idle time provided after the first-half imaging, and (iii) a second-half imaging performed as a late-phase contrast imaging after the idle time, into a first-half imaging series and a second-half imaging series excluding the idle time, the first-half imaging series and the second-half imaging series respectively corresponding to the first-half imaging and the second-half imaging in the original imaging series,set a start time for the second-half imaging series in the form of the lapse time after injection of the contrast medium,insert another imaging series , different from the original imaging series, between the first-half imaging series and the second-half imaging series, by setting a start time for the another imaging series in the form of the lapse time such that the start time for the another imaging series is earlier than the start time of the second-half imaging series, andperform control to start the another imaging series after the first-half imaging series in accordance with the start time set for the another imaging series, and to start the second-half imaging series after the another imaging series in accordance with the start time set for the second-half imaging series; anda display that displays at least the measured lapse time and the set start time,wherein the first-half imaging series and the second-half imaging series are performed on a first patient and the another imaging series is performed on a second patient. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the set start time of the second-half imaging series is changeable by a user operation. 3. The magnetic resonance imaging apparatus according to claim 1, wherein the display displays the start time of the another imaging and the start time of the another imaging is changeable by a user operation. 4. The magnetic resonance imaging apparatus according to claim 1, wherein the circuitry manages the first-half imaging series and the second-half imaging series resulting from the division as one group, and automatically reflects a modification to an imaging condition of any one of imagings in the group in the imagings other than the imaging whose imaging condition is modified. 5. The magnetic resonance imaging apparatus according to claim 1, wherein the display provides an alarm indication if an imaging time of the another imaging is longer than the idle time. 6. The magnetic resonance imaging apparatus according to claim 1, wherein, when an imaging time of the another imaging is longer than the idle time, the circuitry automatically adjusts a value of a par... | KABUSHIKI KAISHA TOSHIBA TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/05 A61B005/055 A61B005/00 G01R033/563
US09328371 | B2 | 2016-05-03 | Sensor head for use with implantable devices | 2013-07-16 | The present invention provides a sensor head for use in an implantable device that measures the concentration of an analyte in a biological fluid which includes: a non-conductive body; a working electrode, a reference electrode and a counter electrode, wherein the electrodes pass through the non-conductive body forming an electrochemically reactive surface at one location on the body and forming an electronic connection at another location on the body, further wherein the electrochemically reactive surface of the counter electrode is greater than the surface area of the working electrode; and a multi-region membrane affixed to the nonconductive body and covering the working electrode, reference electrode and counter electrode. In addition, the present invention provides an implantable device including at least one of the sensor heads of the invention and methods of monitoring glucose levels in a host utilizing the implantable device of the invention. | 1. An implantable continuous glucose sensor system configured for implantation in a host, the system comprising: an electrode comprising an electroactive surface, wherein the electrode is configured for implantation and is configured to generate in vivo a signal indicative of glucose concentration in a host;a membrane comprising: a first domain configured to reduce a flux of glucose therethrough;a second domain comprising an enzyme configured to react with glucose to produce a measured species, wherein the second domain is located less distant to the electroactive surface than the first domain; anda third domain comprising an architecture that permits access to the electroactive surface only for species that have a molecular weight less than or about equal to a molecular weight of the measured species, wherein the third domain is located less distant to the electroactive surface than the second domain; andsensor electronics operably connected to the electrode and configured to measure the signal from the electrode. 2. The implantable continuous glucose sensor of claim 1, wherein the molecular weight of the architecture limits diffusion of species that have a molecular weight more than 34 Daltons. 3. The glucose sensor of claim 1, wherein the first domain comprises a polyurethane. 4. The glucose sensor of claim 1, wherein the first domain is an outermost domain of the membrane. 5. The glucose sensor of claim 1, wherein the first domain is configured to interface with biological fluid. 6. The glucose sensor of claim 1, wherein the electrode is a working electrode, wherein the system further comprises a reference electrode and a counter electrode. 7. The glucose sensor of claim 1, wherein the counter electrode comprises an electroactive surface, wherein the counter electrode electroactive surface has a larger area than an area of the working electrode electroactive surface. | DexCom, Inc. | A61B005/05 A61B005/00 C12Q001/00 A61B005/145 A61B005/1486
US09351659 | B2 | 2016-05-31 | Biomedical electrode configuration for suppressing movement artifact | 2009-08-24 | A biomedical sensor for detecting EMG signals may include a non-conducting fixed framework supporting a symmetrical arrangement of four electrode surfaces, configured as two signal detection contacts, each with a respective associated signal reference contact. The mechanical and electrical configuration act together to suppress movement artifact. | 1. A biomedical sensor for sensing phenomena of a muscle, comprising: a substrate; a plurality of conductive areas arranged to be spaced along a length of the substrate and configured as signal detection contacts each with a proximate, distinct respective associated signal reference contact, forming a plurality of signal detection/reference contact pairs spaced along the length, each one in the plurality of signal detection/reference contact pairs being aligned in a direction transversely to the length and in parallel with respect to each other so that each signal detection/reference contact pair produces movement induced artifacts in unison in response to relative movement between the plurality of signal detection/reference contact pairs and the muscle; and a differential sensing circuit that receives input from each of the signal detection/reference contact pairs and provides an output that is an arithmetic difference between the signals detected at detection contacts measured with respect to the reference contacts, the reference contacts being directly electrically coupled together and to a reference input of the differential sensing circuit, and the detection contacts being provided at separate inputs to the differential sensing circuit, wherein each signal detection/reference contact pair is configured to react in response to the relative movement, wherein at least one common electrical component caused by the relative movement is cancelled out by the differential sensing circuit, and wherein the at least one common electrical component caused by the relative movement includes the movement induced artifacts that are suppressed by the differential sensing circuit. 2. The biomedical sensor as in claim 1, wherein the substrate is rigid. 3. The biomedical sensor as in claim 1, further comprising: a framework coupled to the substrate, wherein the substrate is a body-directed electrically insulating bottom surface of the framework. 4. The biomedical sensor as in claim 1, wherein the plurality of signal detection/reference contact pairs are spaced apart so that the signal detection contact is aligned to be on the same side of each pair. 5. The biomedical sensor as in claim 1, wherein each of the conductive areas is formed in the shape of a conductive rod aligned in parallel on a plane of the substrate, and projecting from the substrate. 6. The biomedical sensor as in claim 1, wherein each of the conductive areas is formed in the shape of a disc attached to and coplanar with the substrate. 7. The biomedical sensor as in claim 1, wherein each of the conductive areas is formed in a shape of a dome whose cross-sectional profile exhibits a concave curved shape attached to and projecting from the substrate. 8. The biomedical sensor as in claim 1, wherein each of the conductive contact areas are formed of substantially identical material composition, size and shape. 9. The biomedical sensor as claim 1, wherein each of the conductive areas is forme... | Altec, Inc. | A61B005/04 A61B005/0488 A61B005/0492 A61B005/00
US09339227 | B2 | 2016-05-17 | Cognitive function assessment in a patient | 2011-06-17 | A method for analysis of the extent of conscious awareness and likelihood of recovery of a patient includes the steps of applying to the patient a sensory stimulus sequence which is typically auditory; carrying out an EEG to generate waveform signals to record changes in the electromagnetic fields generated by the patient's neural activity; using software provided in a processor to process the waveform signals in order to locate waveform peaks, identify the event-related potential (ERP) components obtained in the waveform and to obtain quantitative measures of those components; and using the software to generate and communicate scores indicative of the extent of conscious awareness and likelihood of recovery of the patient. | 1. A method for analysis of the extent of conscious awareness and/or likelihood of recovery of a patient comprising: applying to the patient a sensory stimulus sequence generated by a stimulator;while applying the stimulus sequence, carrying out an EEG or MEG on the patient to record waveform signals from an array of sensors on, in, or near the head of the patient;wherein the stimulus sequence comprises a compressed stimulus sequence that elicits at least five indicators across a spectrum of brain functions comprising sensory processing, perceptual processing, attention/alerting mechanisms, memory retrieval, and language;using software provided in a processor to process the waveform signals in order to locate waveform peaks, identify the evoked responses contained in the waveform and obtain quantitative measures of these evoked responses;wherein the software uses an automated peak localization method that performs a mathematical decomposition of an averaged waveform to identify all peaks in the waveform signals; and wherein the software then applies identification criteria to choose a candidate peak of interest;and using the software to generate and communicate scores based on the quantitative measures that are indicative of the extent of conscious awareness and/or likelihood of recovery of the patient. 2. The method according to claim 1 wherein the compressed stimulus sequence takes less than 5 minutes. 3. The method according to claim 1 wherein the software generates scores comprising a diagnostic score and wherein points in the diagnostic score are allotted based on said evoked responses related to said plurality of brain functions. 4. The method according to claim 1 wherein the software generates scores comprising a diagnostic score and wherein points in the diagnostic score are allotted based on the statistical assessment of differences between quantitative measures of the patient's evoked responses and normative values from a database. 5. The method according to claim 1 wherein the software generates scores comprising a prognostic score and wherein points in the prognostic score are allotted based on the statistical relationships between averaged waveform features and historical outcomes of patients diagnosed with the same condition. 6. The method according to claim 1, wherein the peak localization method identifies all peaks by determining the zero crossings of the averaged waveform's first and second derivatives, yielding peak and inflection points of the waveform's curvature. 7. The method according to claim 1, wherein the peak localization method employs an amplitude thresholding method whereby small peaks are rejected from further analysis. 8. The method according to claim 1, wherein the identification criteria are applied that describe generic ERP/ERF component characteristics based on several basic features, such as experimental condition, latency, and polarity. 9. The method according to claim 1, wherein the peak localizatio... | National Research Council of Canada | A61B005/05 A61B005/00 A61B005/0484 A61B005/16 A61B005/04
US09451898 | B2 | 2016-09-27 | Method and system for extracting ventricular fibrillation signals in electrocardiogram using spline interpolation with uniform phase ensembles | 2016-01-12 | This invention discloses a system and a method for extracting VF signal in ECG recorded during uninterrupted CPR. The present invention provides a method for extracting a Ventricular fibrillation (VF) signal in Electrocardiography (ECG), comprising: receiving an ECG signal; adding a plurality of shadowing functions to the ECG signal, to obtain a plurality of modification signals; decomposing the plurality of modification signals by using an Empirical Mode Decomposition (EMD) method, to generate a plurality of Intrinsic Mode Functions (IMFs); calculating the sum of IMFs in different frequency regions based on time sequence, dividing by a number of the shadowing signal, to obtain a plurality of modification intrinsic mode functions; combining the plurality of modification IMFs with the same property, to obtain a shape function; modeling the shape functions to obtain a compression signal; and subtracting the compression signal from the ECG signal based on time sequence, to obtain the VF signal. | 1. A computer-implemented method for extracting a Ventricular fibrillation (VF) signal in Electrocardiography (ECG), the method comprising: (A) receiving an ECG signal;(B) adding a plurality of shadowing functions to the ECG signal, to obtain a plurality of modification signals, wherein each pair of shadowing functions is conjugate and has same phase difference;(C) decomposing the plurality of modification signals by using an Empirical Mode Decomposition (EMD) method, to generate a plurality of Intrinsic Mode Functions (IMFs);(D) calculating the sum of IMFs in different frequency regions based on time sequence, dividing by a number of the shadowing signal, to obtain a plurality of modification intrinsic mode functions;(E) combining the plurality of modification IMFs with the same property, to obtain a shape function;(F) modeling the shape function, to obtain a compression signal; and(G) subtracting the compression signal from the ECG signal based on time sequence, to obtain the VF signal. 2. The method according to claim 1, wherein the ECG signal is a function of amplitude and time. 3. The method according to claim 1, the steps further comprises: (B1) adding a first shadowing signal to the ECG signal, to obtain a first modification signal;(B2) generating a plurality of first intrinsic mode functions by using the EMD method, wherein the first intrinsic mode function is an amplitude expression of the first modification signal in the different frequency regions;(B3) adding a second shadowing signal to the ECG signal to obtain a second modification signal;(B4) generating a plurality of second intrinsic mode functions by using the EMD method, wherein the second intrinsic mode function is an amplitude expression of the second modification signal in the different frequency regions; and(B5) repeating to add a np-1 shadowing signal and a np shadowing signal to the ECG signal, to obtain the plurality of intrinsic mode functions, wherein the phase difference of the shadowing signals are constant. 4. The method according to claim 3, wherein the first shadowing function and the second shadowing function are conjugate. 5. The method according to claim 1, the steps further comprises: (C1) selecting the intrinsic mode functions corresponding to a first frequency region, calculating the sum of the intrinsic mode functions based on the time sequence, then dividing by the a number of the shadowing signal, to obtain a first modification intrinsic mode function; and(C2) repeating step C1 until to obtain the modification intrinsic mode functions from all frequency regions. 6. The method according to claim 3, wherein a function of the shadowing signal is calculated according to the following expression: wi(t)=ε*cos(2πt/Tw−2πi/np)wherein c is the amplitude of the shadowing signal, Tw is a period of the shadowing signal, np is the number of the shadowing signal and i is any number into 0 to np-1. 7. The method according to claim 1, the steps further comprises: (B1)... | NATIONAL CENTRAL UNIVERSITY | A61B005/04 A61B005/046 A61B005/00 A61B005/11
US09332930 | B2 | 2016-05-10 | Use of a support device | 2014-01-16 | A method and system for determining a pose of a subject. The subject is placed onto a surface of a platform having a first radio-frequency transceiver and a second radio-frequency transceiver. A first radio-frequency transponder is placed onto an upper extremity of the subject. A second radio-frequency transponder is placed onto a lower extremity of the subject. A first signal is transmitted from the first transceiver to the first transponder, resulting in the first transponder sending a second signal. A third signal is transmitted from the second transceiver to the second transponder, resulting in the second transponder sending a fourth signal. The second signal is received by the first transceiver. The fourth signal is received by the second transceiver. A pose of the subject is determined based on receiving the second signal and the fourth signal. | 1. A method for determining a pose of a subject, said method comprising: placing transponders on the subject disposed on a surface of a platform, wherein a totality of the transponders consist of a first radio-frequency transponder and a second radio-frequency transponder placed onto an upper extremity and a lower extremity, respectively, of the subject, and wherein a first radio-frequency transceiver is located at a first location on the surface and a second radio-frequency transceiver is located at a second location on the surface;transmitting a first radio-frequency signal from the first radio-frequency transceiver to the first radio-frequency transponder on the upper extremity of the subject, resulting in the first radio-frequency transponder sending a second radio-frequency signal to the first radio-frequency transceiver;receiving, by the first radio-frequency transceiver, the second radio-frequency signal;receiving, by the second radio-frequency transceiver, a fourth radio-frequency signal transmitted by the second radio-frequency transponder in response to the second radio-frequency transponder having previously received a third radio-frequency signal from the second radio-frequency transceiver; anddetermining the pose of the subject based on said receiving the second radio-frequency signal and said receiving the fourth radio-frequency signal. 2. The method of claim 1, wherein the method further comprises storing data in the first radio-frequency transponder, and wherein the second radio-frequency signal sent by the first radio-frequency transponder to the first radio-frequency transceiver comprises the stored data. 3. The method of claim 2, wherein the stored data comprises information about the first radio-frequency transponder and a location of the first radio-frequency transponder. 4. The method of claim 2, wherein the stored data comprises patient information pertaining to the subject. 5. The method of claim 4, wherein the patient information comprises a name of the subject, an address of the subject, a government issued identification number of the subject, and a medical condition of the subject. 6. The method of claim 5, wherein the patient information further comprises a blood type of the subject. 7. The method of claim 2, wherein the subject is a patient, and wherein the stored data comprises a series of numbers corresponding to the patient, the first radio-frequency transponder, and a body extremity of the patient. 8. The method of claim 7, wherein the stored data comprises a series of numbers corresponding to the patient, the first radio-frequency transponder, and a body extremity of the patient. 9. The method of claim 1, wherein the subject is a human being. 10. The method of claim 9, wherein the upper extremity and lower extremity respectively comprise a head and a leg of the subject, and wherein the first radio-frequency transponder and the second radio-frequency transponder are respectively placed onto the head and ... | International Business Machines Corporation | A61B005/00 A61B005/11
US09326735 | B2 | 2016-05-03 | Method and apparatus for monitoring cardiac output of a patient in a home environment | 2013-04-23 | There is provided a method of determining a measure of the cardiac output, CO, of a patient, the method comprising obtaining measurements of one or more physiological characteristics of the patient, the physiological characteristics including at least the heart rate, HR, of the patient, the systolic blood pressure, S, of the patient and the diastolic blood pressure, D, of the patient; and processing the measurements to determine a measure of the cardiac output of the patient; wherein the measure of the cardiac output of the patient is derived using the relationship;CO=K1⁡(HR⁢S-DS+D),CO=K2⁢HR⁡(S-D)⁢(PTT)2,and⁢/⁢or⁢CO=K2⁢HR⁡(S-D)⁢(PAT-PEP)2 where K1 and K2 are patient-specific calibration factors, PAT is the pulse arrival time, PEP is the pre-ejection period and PTT is the pulse transit time, the time taken for a pulse wave to travel between two points in the body of the patient. | 1. A method of determining a measure of the cardiac output, CO, of a patient with a monitoring apparatus, the monitoring apparatus comprising a hardware processing unit and one or more physiological sensors, the method comprising: obtaining, with the processing unit, measurements of one or more physiological characteristics of the patient from the one or more physiological sensors, the physiological characteristics including at least the heart rate, HR, of the patient, the systolic blood pressure, S, of the patient and the diastolic blood pressure, D, of the patient; andprocessing, with the processing unit, the measurements to determine a measure of the cardiac output of the patient;wherein the measure of the cardiac output of the patient is derived using the relationship:CO=K1⁡(HR⁢S-DS+D),⁢CO=K2⁢HR⁡(S-D)⁢(PTT)2,⁢and⁢/⁢orCO=K2⁢HR⁡(S-D)⁢(PAT-PEP)2,where K1 and K2 are patient-specific calibration factors, PAT is the pulse arrival time, PEP is the pre-ejection period and PTT is the pulse transit time, the time taken for a pulse wave to travel between two points in the body of the patient. 2. A method as claimed in claim 1, wherein K1 and/or K2 are determined by the processing unit from a reference measurement of the cardiac output of the patient. 3. A method as claimed in claim 1, wherein the measure of the cardiac output of the patient is a measure of the change in the cardiac output from an earlier cardiac output measurement or a reference measurement for cardiac output, and the measure of the change in the cardiac output is given by⁢COCOref=(HR⁢S-DS+D)/(HR⁢S-DS+D)ref,⁢⁢COCOref=(HR⁡(S-D)⁢PTT2)/(HR⁡(S-D)⁢PTT2)ref,⁢⁢and⁢/⁢orCOCOref=(HR⁡(S-D)⁢(PAT-PEP)2)/(HR⁡(S-D)⁢(PAT-PEP)2)refwhere ( . . . )ref denotes the earlier cardiac output measurement or the reference measurement for cardiac output. 4. A method as claimed in claim 1, wherein the step of obtaining measurements comprises obtaining, from the one or more physiological sensors, an electrocardiogram signal, a photoplethysmogram signal, and at least one of heart sounds and an impedance cardiogram signal; and wherein the step of processing comprises processing the electrocardiogram signal and photoplethysmogram signal to determine the pulse arrival time, and processing the at least one of the heart sounds or the impedance cardiogram signal and the electrocardiogram signal to determine the pre-ejection period. 5. A method as claimed in claim 1, further comprising the steps of: comparing, with the processing unit, the measure of the cardiac output of the patient or measure of the change in the cardiac output of the patient to one or more threshold values; andproviding, with the processing unit, an output indicating the result of the step of comparing. 6. A method as claimed in claim 1, the method further comprising the steps of: obtaining, with the processing unit, measurements of acceleration acting on the patient during the measurement of the physiological characteristics of the patient;processi... | Koninklijke Philips N.V. | A61B005/02 A61B005/00 A61B005/029 A61B005/0205 A61B005/026 A61B005/0402 A61B005/11 A61B007/04 A61B007/00 A61B005/024 A61B005/021
US09326721 | B2 | 2016-05-03 | Labor management devices for pelvic floor support | 2015-07-22 | The present disclosure provides devices for decreasing the incidence of Cesarean childbirth by managing the labor process. In one aspect, a tactile feedback device is positioned adjacent the perianal tissues. A perianal support device includes a perianal support member having a pressure surface configured for engagement with tissue adjacent an anal orifice and a pressure detecting system associated with the perianal support member to detect pressure indicative of pressure on the perianal tissue of a patient. | 1. A child birth labor pelvic floor support apparatus, comprising: a support member comprising: a pressure surface configured to apply pressure on tissue adjacent a pelvic floor of a patient to support the pelvic floor during child birth;a compression element extending in a direction from the pressure surface and configured to convey applied loading toward the pressure surface to maintain the pressure surface in a position that applies the pressure on the tissue to support the pelvic floor; anda pressure detecting system configured to detect pressure on the tissue adjacent the pelvic floor of the patient. 2. The childbirth labor pelvic floor support apparatus of claim 1, wherein the pressure detecting system comprises: a pressure detecting element configured to detect pressure applied on the tissue adjacent the pelvic floor by the pressure surface; anda user interface configured to display information relating to the pressure detected by the pressure detecting element. 3. The childbirth labor pelvic floor support apparatus of claim 2, wherein the user interface is configured to wirelessly communicate with the pressure detecting element. 4. The childbirth labor pelvic floor support apparatus of claim 1, wherein the pressure detecting system comprises a compliant pad on a pressure surface of the childbirth labor pelvic floor support apparatus and the compliant pad is configured to change appearance when the pressure exceeds a threshold pressure. 5. The childbirth labor pelvic floor support apparatus of claim 1, wherein the tissue adjacent the pelvic floor includes at least a portion of a perineum between an anus and a vaginal opening of the patient. 6. The childbirth labor pelvic floor support apparatus of claim 1, wherein the pressure detecting system comprises: a securing member extending from the childbirth labor pelvic floor support apparatus, the securing member being configured to convey information indicative of the pressure applied on the tissue adjacent the pelvic floor adjacent an anal orifice of the patient by the pressure surface. 7. The childbirth labor pelvic floor support apparatus of claim 6, wherein the securing member comprises an element configured to change appearance when pressure changes. 8. The childbirth labor pelvic floor support apparatus of claim 1, further comprising a counter associated with the pressure detecting system, the counter configured to count pressure sequences above a first threshold indicative of effective pushes of the patient during child birth labor. 9. The childbirth labor pelvic floor support apparatus of claim 1, wherein the compression element comprises: a first rigid compression element surface disposed to face a first buttock in a gluteal cleft of the patient in response to the pressure surface being disposed adjacent an anal orifice of the patient, the first compression element surface extending away from the pressure surface in a first direction out of the gluteal cleft; anda second rigid... | STETRIX, INC. | A61B005/00 A61B017/42 A61F005/00 A61B019/00 A61F006/08 A61B017/00
US09345879 | B2 | 2016-05-24 | Device and implantation system for electrical stimulation of biological systems | 2014-03-07 | The present specification discloses devices and methodologies for the treatment of transient lower esophageal sphincter relaxations (tLESRs). Individuals with tLESRs may be treated by implanting a stimulation device within the patient's lower esophageal sphincter and applying electrical stimulation to the patient's lower esophageal sphincter, in accordance with certain predefined protocols. The presently disclosed devices have a simplified design because they do not require sensing systems capable of sensing when a person is engaged in a wet swallow and have improved energy storage requirements. | 1. A system for treating a gastrointestinal condition of a patient, comprising: a pulse generator in electrical communication with at least one electrode, wherein said pulse generator is adapted to be implanted such that said electrode is configured to be in electrical communication with a lower esophageal sphincter (LES) of the patient, and wherein said pulse generator generates a pulse stream in accordance with a preset period;an energy storage component; anda non-volatile memory for storing programmatic code which, when executed by a processor, causes said pulse generator to generate the pulse stream for the preset period,wherein said pulse stream comprises a pulse width having a range of 30 μsec to 5 msec, a pulse amplitude having a range of 2 to 15 mAmp, a pulse frequency having a range of 2 to 80 Hz, an on period having a range of 1 second to 23 hours, 59 minutes, and 59 seconds, an off period having a range of 1 second to 23 hours, 59 minutes, and 59 seconds, and a duty cycle having a range of 1 to 100%, wherein values for each of said pulse width, pulse amplitude, and pulse frequency are defined such that swallowing of the patient is not hindered by said stimulation while said stimulation is occurring. 2. The system of claim 1, further comprising an accelerometer wherein said system includes a pre-programmed stimulation mode adapted to automatically stimulate the patient's LES during time periods in which the patient is in a supine position as determined by said accelerometer, said supine position being defined by a level of inclination relative to a vertical baseline wherein 0 degrees is a vertical position and wherein said level is programmable and has a range of 50 to 80 degrees. 3. The system of claim 2, wherein said automatic stimulation occurs after the patient has been in the supine position for a predetermined period of time and wherein said predetermined period of time is programmable and has a range of 1 minute to 60 minutes. 4. The system of claim 3, wherein said automatic stimulation occurs after a minimum percentage of data points occur within said predetermined period of time and wherein said minimum percentage of data points is programmable and has a range of 50 to 90 percent. 5. The system of claim 1, wherein said system does not have a sensor to determine when the patient is swallowing and does not modify said stimulation based on said swallowing. 6. A system for treating a gastrointestinal condition of a patient, comprising: a pulse generator in electrical communication with at least one electrode, wherein said pulse generator is adapted to be implanted such that said electrode is configured to be in electrical communication with a lower esophageal sphincter (LES) of the patient, and wherein said pulse generator generates a pulse stream;an energy storage component; andprogrammatic code which, when executed by a processor, causes said pulse generator to generate the pulse stream, wherein the programmatic code is not ... | EndoStim, Inc. | A61N001/36 A61B005/03 A61B005/145 A61N001/05